U--D | mg/L | Uranium Dissolved | MSEA | Metals in Seawater (undigested) by (CRC) ICP-MS | This method describes the multi-elemental determination of trace elements by Collision/Reaction Cell (CRC) ICP-MS. The method measures ions produced by a radio-frequency inductively coupled plasma. Analyte species originating in a liquid are pneumatically nebulized and the resulting aerosol transported by argon gas into the plasma torch. The ions produced are entrained in the plasma gas and extracted, by means of a differentially pumped vacuum interface, into a mass spectrometer. The ions flow through a Collision/Reaction Cell (CRC) where depending on the mode of analysis either the ions passes through the cell unhindered (no gas - classical ICP-MS mode), or collisions/reactions chemistry takes place between the ions and a neutral gas such as hydrogen or helium. Any polyatomic ions which give interference on the analytes are removed. The ions are then sorted according to their mass-to-charge ratios by a quadrupole mass spectrometer. The ions are then sorted according to their mass-to-charge ratios by a quadrupole mass spectrometer. The ions transmitted through the quadrupole are quantified by a channel electron multiplier and the ion information processed by a data handling system. Although most polyatomic interference is removed, any additional interference must be assessed and valid corrections applied or the data flagged to indicate problems. Interference correction must include compensation for background ions contributed by the plasma gas, reagents, and constituents of the sample matrix. Instrument drift as well as suppressions or enhancements of instrument response caused by the sample matrix must be corrected for by the use of internal standards. | 6e-06 | Y | Y | Y | 4 | 4/12/2023 | 4/29/2024 | 5/16/2024 2:00:14 PM | 2009-09-08 |  | RDGIBSON |  | Concentration | mg/L | Y | (fl. conc.) |  | Uranium Dissolved (fl. conc.) | 7440-61-1 | Metals - Dissolved; | Metals - Dissolved; EMS code: U--D | Chemical | Dissolved | LAB
U--D | mg/L | Uranium Dissolved | MWID | Metals in Water by ICPMS (digested) | Metals in Water by ICPMS (digested) Based on EPA 200.8 | 0.0001 | Y | Y | Y | 31767 | 10/26/2018 5:20:00 PM | 10/26/2018 5:20:00 PM | 7/31/2024 1:46:41 PM | 2009-12-24 |  | RDGIBSON |  | Concentration | mg/L | Y | (fl. conc.) |  | Uranium Dissolved (fl. conc.) | 7440-61-1 | Metals - Dissolved; | Metals - Dissolved; EMS code: U--D | Chemical | Dissolved | LAB
U--D | mg/L | Uranium Dissolved | METC | MET-D-NP-CCMS-VA | Ultra trace metals in water are analyzed by CRC ICPMS, based on US EPA Method 6020B (July 2014). This procedure is intended for pristine field-filtered acid-preserved water samples. The detection limits (LOR) for this test are based on lab instrumental analysis only, not including filtration. ALS-supplied field filtration equipment does not support these LOR. Therefore, because of the high probability of false positives due to filtration, it is strongly recommended that a filtration blank be analysed to aid in data interpretation. | 2e-06 | Y | Y | Y | 1776 | 4/3/2012 | 8/15/2024 | 8/23/2024 9:50:35 AM | 2019-08-20 |  | KATFRASE |  | Concentration | mg/L | Y | (fl. conc.) |  | Uranium Dissolved (fl. conc.) | 7440-61-1 | Metals - Dissolved; | Metals - Dissolved; EMS code: U--D | Chemical | Dissolved | LAB
U--D | mg/L | Uranium Dissolved | X170 | ICP | This code is a catch-all for a large variety of ICP methods including water and also solids digested into solutions as an intermediate result. Different procedures will have greatly different detection limits. (a) For air particulates, preliminary results are reported in mg/L and loading results as mg/dm2/d. (b) For Hi Vol samples preliminiary results are same as reported in mg/L and loading results as ug/m3. (c) Note that for Rain samples there is no filtration nor digestion steps. | 1e-05 | Y | Y | Y | 3221 | 4/24/2002 | 5/19/2015 | 10/31/2016 5:30:52 PM | 1997-06-23 | 6/3/1998 6:51:47 PM | NPEPPIN |  | Concentration | mg/L | Y | (fl. conc.) |  | Uranium Dissolved (fl. conc.) | 7440-61-1 | Metals - Dissolved; | Metals - Dissolved; EMS code: U--D | Chemical | Dissolved | LAB
U--D | mg/L | Uranium Dissolved | F032 | Filt.,Acid,ICPMS | There is no acid digestion required for this analysis. A water sample is collected and preserved in 0.2 percent nitric acid. To analyse the extractable metals the sample is shaken, allowed to sit for at least twenty four hours, and a decanted portion is introduced to the spectrometer. When the dissolved metals are required, the samples must first be filtered through a 0.45 micron cellulose acetate filter paper and then followed by preservation in 0.2 percent nitric acid. It is important to note that different laboratories may have very different detection limits for some variables analyzed by ICPMS. | 0.0002 | Y | Y | Y | 74252 | 12/19/1996 8:15:00 AM | 8/26/2024 | 8/26/2024 4:53:05 PM | 1997-01-16 | 8/15/2019 11:01:03 AM | MPAK | KATFRASE | Concentration | mg/L | Y | (fl. conc.) |  | Uranium Dissolved (fl. conc.) | 7440-61-1 | Metals - Dissolved; | Metals - Dissolved; EMS code: U--D | Chemical | Dissolved | LAB
U--D | mg/L | Uranium Dissolved | 0600 | LIF, U Std Addition | LIF, U Std Addition | 0.0001 | Y | Y | Y | 95 | 12/1/2010 | 12/1/2010 | 12/9/2010 3:55:27 PM | 1996-04-12 | 6/3/1998 6:51:40 PM | EMS |  | Concentration | mg/L | Y | (fl. conc.) |  | Uranium Dissolved (fl. conc.) | 7440-61-1 | Metals - Dissolved; | Metals - Dissolved; EMS code: U--D | Chemical | Dissolved | LAB
U--D | mg/L | Uranium Dissolved | X562 | Dissolved metals: FF, FA, HNO3, HR-ICPMS | Ultra trace metals in water are analyzed by high resolution inductively coupled plasma mass spectrometry (HR-ICPMS), based on US EPA Method 200.8, (Oct. 1999). This procedure is intended for pristine field-filtered acid-preserved water samples. The detection limits (LOR) for this test are based on lab instrumental analysis only, not including filtration. ALS-supplied field filtration equipment does not support these LOR. Therefore, because of the high probability of false positives due to filtration, it is strongly recommended that a filtration blank be analyzed to aid in data interpretation. | 2e-06 | Y | Y | Y | 1264 | 6/8/2015 | 10/19/2023 | 11/9/2023 7:41:27 AM | 2015-10-22 |  | RDGIBSON |  | Concentration | mg/L | Y | (fl. conc.) |  | Uranium Dissolved (fl. conc.) | 7440-61-1 | Metals - Dissolved; | Metals - Dissolved; EMS code: U--D | Chemical | Dissolved | LAB
U--E | mg/L | Uranium Extractable | F068 | HNO3dig,TeflonTube/Autoclave,ICP-MS | This method is applicable to determiniation of dissolved, extractable, and total recoverable metals in ground waters, surface waters, drinking waters, and wastewaters. Dissolved, extractable, and total recoverable sample fractions are acidified to pH <1.5 with analytical grade nitric acid (HNO3) to prevent the metals from adsorbing to the container material or precipitating. Extractable and total recoverable water samples are collected in 250 ml acid washed polyethylene bottles and are preserved with Nitric acid (0.5 ml acid for 100 mLs of sample). All preserved samples are stored at room temperature and can be held for up to 6 months. Unfiltered, un-preserved samples are stored at 4C for no more than 30 days. For dissolved and extractable water fractions, no sample preparation is necessary. Samples are simply poured in autosampler tubes (i.e. 15 ml centrifuge tubes), and placed on the autosampler and analyzed. Samples for total recoverable metals are digested with Nitric acid in special Teflon tubes in an autoclave prior to analyses. | 2e-06 | Y | Y | Y | 9 | 3/26/2012 3:01:00 AM | 3/26/2012 3:01:00 AM | 3/28/2012 10:45:30 AM | 1999-06-28 |  | NPEPPIN |  | Concentration | mg/L | Y | (fl. conc.) |  | Uranium Extractable (fl. conc.) | 7440-61-1 | Metals - Extractable; | Metals - Extractable; EMS code: U--E | Chemical | Extractable | LAB
U--E | mg/L | Uranium Extractable | X553 | Metals in water, extractable, ICP-SFMS | The whole water sample is collected in a HDPE/Teflon bottle. In a Class 100 workstation it is preserved with concentrated Nitric Acid at concentration of 0.2 percent (v/v). The acid preservative extracts the metals from material in the sample and a decanted portion is introduced directly to an inductively coupled argon plasma-sector field mass spectrometer (ICP-SFMS) by pneumatic nebulization. Each respective element is measured at a specific mass to charge ratio, m/z value expressed in amu (atomic mass units), and at a specified instrument resolution. LR: Low Resolution ~300, MR: Mid Resolution ~4000 and HR: High Resolution ~10000. Ref: Method number 02-2061, National Laboratory for Environmental Testing, CCIW, Environment Canada, Burlington. Requested April 2003. | 5e-07 | Y | Y | Y | 2598 | 5/22/2003 | 9/21/2007 | 11/28/2023 2:11:53 PM | 2004-01-19 |  | UMULLER |  | Concentration | mg/L | Y | (fl. conc.) |  | Uranium Extractable (fl. conc.) | 7440-61-1 | Metals - Extractable; | Metals - Extractable; EMS code: U--E | Chemical | Extractable | LAB
U--E | mg/L | Uranium Extractable | F047 | 24hr HNO3 dil'n,ICP-MS | There is no acid digestion required for this analysis. A water sample is collected and preserved in 0.2 percent nitric acid. To analyse the extractable metals the sample is shaken, allowed to sit for at least twenty four hours, and a decanted portion is introduced to the spectrometer. When the dissolved metals are required, the samples must first be filtered through a 0.45 micron cellulose acetate filter paper and then followed by preservation in 0.2 percent nitric acid. It is important to note that different laboratories may have very different detection limits for some variables analyzed by ICP/MS. | 1e-06 | Y | Y | Y | 2691 | 4/16/1998 10:21:00 AM | 12/13/2011 12:02:00 PM | 9/11/2019 8:46:18 AM | 1997-10-20 | 6/3/1998 6:51:47 PM | NPEPPIN |  | Concentration | mg/L | Y | (fl. conc.) |  | Uranium Extractable (fl. conc.) | 7440-61-1 | Metals - Extractable; | Metals - Extractable; EMS code: U--E | Chemical | Extractable | LAB
U--L | mg/L | Uranium Leachable | MEND | MEND (mod) EPA 300.1 (mod) | MEND (mod) EPA 300.1 (mod). Inorganic anions are analyzed by obtaining an extract from a soil sample that has been added in a 3:1 ratio of deionized water to soil; then leached and the extract filtered prior to analysis; which is then analyzed by Ion Chromatography with conductivity and/or UV detection. | 1e-05 | Y | Y | Y | 0 |  |  |  | 2024-01-29 |  | RDGIBSON |  | Concentration | mg/L | Y | (fl. conc.) |  | Uranium Leachable (fl. conc.) | 7440-61-1 | Metals - Extractable; | Metals - Extractable; EMS code: U--L | Chemical |  | LAB
U--T | mg/L | Uranium Total | MWID | Metals in Water by ICPMS (digested) | Metals in Water by ICPMS (digested) Based on EPA 200.8 | 0.0001 | Y | Y | Y | 36811 |  |  | 7/31/2024 1:46:41 PM | 2009-12-24 |  | RDGIBSON |  | Concentration | mg/L | Y | (fl. conc.) |  | Uranium Total (fl. conc.) | 7440-61-1 | Metals - Total; | Metals - Total; EMS code: U--T | Chemical | Total | LAB
U--T | ppm (S) | Uranium Total | 6041 | Aqua Regia by ICPMS | These were sediment/soils digested with Aqua-Regia (Nitric Acid & Hydrochloric Acids) and then analysis by ICPMS. Results would be reported in mg/kg and/or %. | 0.005 | Y | Y | Y | 21 | 10/13/2019 | 8/5/2020 | 8/10/2020 3:52:07 PM | 2019-09-06 | 10/21/2019 11:56:48 AM | KATFRASE | KATFRASE | DimensionlessRatio | ppm | N | (ratio) |  | Uranium Total (ratio) | 7440-61-1 | Metals - Total; | Metals - Total; EMS code: U--T | Chemical | Total | LAB
U--T | mg/L | Uranium Total | TCLP | Toxicity Characteristic Leaching Procedure | The TCLP procedure is designed to simulate the climatic leaching action | 0.1 | Y | Y | Y | 38 | 12/8/2009 7:49:00 PM | 9/28/2011 5:53:00 PM | 8/24/2016 2:59:13 PM | 2009-03-06 |  | JCHUDAK |  | Concentration | mg/L | Y | (fl. conc.) |  | Uranium Total (fl. conc.) | 7440-61-1 | Metals - Total; | Metals - Total; EMS code: U--T | Chemical | Total | LAB
U--T | mg/L | Uranium Total | F108 | SPR,IDA Chelation,ICPMS | SPR;IDA Chelation;ICPMS |  | Y | Y | Y | 4 | 11/30/2000 10:32:00 AM | 11/30/2000 10:32:00 AM | 9/10/2002 9:29:49 AM | 2001-03-08 |  | NPEPPIN |  | Concentration | mg/L | Y | (fl. conc.) |  | Uranium Total (fl. conc.) | 7440-61-1 | Metals - Total; | Metals - Total; EMS code: U--T | Chemical | Total | LAB
U--T | mg/L | Uranium Total | F033 | Acid Dig,Centrifuge,ICPMS | For the determination of total recoverable elements in water or waste water, take 100 ml aliquot from a well mixed, acid preserved, sample containing not more than 0.25% (w/v) total solids and transfer to a 250 ml Griffin beaker. (If total solids are greater than 0.25% reduce the size of the aliquot by a proportionate amount). Add 1 ml of conc. nitric acid and 0.5 ml conc. hydrochloric acid. Heat on a hot plate until the volume has been reduced to 15-20ml, ensuring that the sample does not boil. Allow to cool and transfer to a 50 ml volumetric flask. Dilute to volume with ASTM type I water. Centrifuge the sample or allow to stand overnight to separate insoluble material. Prior analysis, pipette 20 ml into a 50 ml volumetric flask and dilute to volume with ASTM type I water. Analysis should be performed within seventy two hours of the completed preparation. Note that different laboratories may have very different field preservation, digestion procedures and detection limits for this method, so these should be documented. | 1e-05 | Y | Y | Y | 37047 | 12/16/1996 7:44:00 AM | 8/22/2024 3:24:00 AM | 8/25/2024 9:47:36 PM | 1997-01-16 | 6/3/1998 6:51:46 PM | MPAK |  | Concentration | mg/L | Y | (fl. conc.) |  | Uranium Total (fl. conc.) | 7440-61-1 | Metals - Total; | Metals - Total; EMS code: U--T | Chemical | Total | LAB
U--T | ug/g wet | Uranium Total | X560 | Metals in Biota (wet), total, acid digestion, ICP-MS | Five grams of wet, homogenized biota is weighed in a centrifuge tube, and 2.5 mL of Nitric Acid and 2.5 mL of Hydrochloric acid are added. The sample is digested at room temperature for one hour and then digested at 90-95 degrees C in a water bath or heat block for one more hour. The sample is cooled and the volume is made to 50 mL with DI. The digestate is then diluted at ten fold and introduced to a collision/reaction cell inductively coupled argon plasma mass spectrometer (CRC ICPMS) by pneumatic nebulization. Each respective element is measured at a specific mass to charge ratio, m/z value expressed in amu (atomic mass units). - Maxxam Analytics (BC). Added to EMS Dec 10, 2008. | 0.05 | Y | Y | Y | 191 | 11/15/2008 4:00:00 AM | 10/30/2023 | 11/28/2023 2:21:33 PM | 2008-12-11 | 3/11/2015 4:17:10 PM | RDGIBSON | EMS | MassPerMass | µg/g | Y | (mass conc.) | (wet) | Uranium Total (wet) (mass conc.) | 7440-61-1 | Metals - Total; | Metals - Total; EMS code: U--T | Chemical | Total | LAB
U--T | mg/L | Uranium Total | X555 | Metals in water, total, "in-bottle" digestion, ICP-SFMS | The whole water sample collected in a HDPE/Teflon bottle, is de-capped in a Class 100 workstation, acidified to 2% (v/v) HNO3 and 1% (v/v) HCl and recapped. The sample is "in-bottle" digested in a convection oven, set at 60 degrees C for 16 hours. The sample is cooled and introduced to an inductively coupled argon plasma-sector field mass spectrometer (ICP-SFMS) by pneumatic nebulization. Each respective element is measured at a specific mass to charge ratio, m/z value expressed in amu (atomic mass units), and at resolution. LR: Low Resolution ~300, MR: Mid Resolution ~4000 and HR: High a specified instrument Resolution ~10000. Ref: Method number 02-2003, National Laboratory for Environmental Testing, CCIW, Environment Canada, Burlington. Requested April 2003. | 5e-07 | Y | Y | Y | 17324 | 5/20/2003 | 12/6/2023 2:08:00 AM | 2/14/2024 9:05:32 AM | 2004-01-19 |  | UMULLER |  | Concentration | mg/L | Y | (fl. conc.) |  | Uranium Total (fl. conc.) | 7440-61-1 | Metals - Total; | Metals - Total; EMS code: U--T | Chemical | Total | LAB
U--T | ug/g | Uranium Total | PB02 | PBM,HNO3/HCl dig,dry60C,100mesh, | Performance Based Method. Samples are dried at 60 C and sieved using a .149 mm (100 mesh) sieve. Solid samples are digested with the combination of nitric and hydrochloric acids with a 1:1 ratio of acids. For environmental purposes, strong acid leaches are used routinely since they have the advantage of speed, low cost, and superio reproducibility. All metals trapped within the refractory matrix are unlikely to have environmental consequences. Instrumental analysis of extracts produced using this digestion method can be performed by a variety of techniques. |  | Y | Y | Y | 52 | 5/24/2001 3:42:00 PM | 11/26/2001 10:23:00 AM | 5/26/2003 2:28:44 PM | 2001-07-18 |  | NPEPPIN |  | MassPerMass | mg/kg | Y | (mass conc.) |  | Uranium Total (mass conc.) | 7440-61-1 | Metals - Total; | Metals - Total; EMS code: U--T | Chemical | Total | LAB
U--T | mg/L | Uranium Total | X344 | HNO3 Dig,Laser-induced fluorescen | Unfiltered sample; HNO3 digestion. Uranyl ion, complexed with pyrophosphate, is excited at 337 nm with a pulsed nitrogen laser and the resulting fluorescence at 494, 516 or 540 nm is monitored. The method of standard additions is used to minimize interferences from matrix effects. | 0.0002 | Y | Y | Y | 49 |  |  | 10/15/2001 1:09:32 PM | 1996-04-12 | 6/3/1998 6:51:40 PM | EMS |  | Concentration | mg/L | Y | (fl. conc.) |  | Uranium Total (fl. conc.) | 7440-61-1 | Metals - Total; | Metals - Total; EMS code: U--T | Chemical | Total | LAB
U--T | mg/L | Uranium Total | 1349 | Tot Metals in Seawater by CRC ICPMS (BC) | Tot Metals in Seawater by CRC ICPMS (BC) | 5e-05 | Y | Y | Y | 98 | 8/16/2019 | 12/5/2021 | 12/13/2021 12:34:17 PM | 2019-09-30 |  | KATFRASE |  | Concentration | mg/L | Y | (fl. conc.) |  | Uranium Total (fl. conc.) | 7440-61-1 | Metals - Total; | Metals - Total; EMS code: U--T | Chemical | Total | LAB
U--T | ppm (S) | Uranium Total | MS81 | ALS Minerals: ME-MS81 | ALS Minerals: ME-MS81. A prepared sample (0.100 g) is added to lithium metaborate/lithium tetraborate flux; mixed well and fused in a furnace at 1025°C. The resulting melt is then cooled and dissolved in an acid mixture containing nitric; hydrochloric and hydrofluoric acids. This solution is then analyzed by inductively coupled plasma - mass spectrometry. | 0.05 | Y | Y | Y | 0 |  |  |  | 2024-01-29 |  | EMS |  | DimensionlessRatio | ppm | N | (ratio) |  | Uranium Total (ratio) | 7440-61-1 | Metals - Total; | Metals - Total; EMS code: U--T | Chemical | Total | LAB
U--T | mg/L | Uranium Total | 0610 | HNO3 Dig: LIF, U Std Add | HNO3 Dig: LIF, U Std Add | 0.0002 | Y | Y | Y | 495 |  |  | 12/18/2000 11:37:22 AM | 1996-04-12 | 6/3/1998 6:51:40 PM | EMS |  | Concentration | mg/L | Y | (fl. conc.) |  | Uranium Total (fl. conc.) | 7440-61-1 | Metals - Total; | Metals - Total; EMS code: U--T | Chemical | Total | LAB
U--T | ppm (S) | Uranium Total | MS61 | ALS Minerals ME-MS61 | ALS Minerals ME-MS61. A prepared sample is digested with HNO3; HF; and HClO4 in a graphite heating block to near dryness. After cooling; the sample is boiled in diluted HCl. Bulk to volume with deionized water; mixed and analyzed by ICP-AES and ICP-MS for all elements. | 0.1 | Y | Y | Y | 0 |  |  |  | 2024-01-29 |  | EMS |  | DimensionlessRatio | ppm | N | (ratio) |  | Uranium Total (ratio) | 7440-61-1 | Metals - Total; | Metals - Total; EMS code: U--T | Chemical | Total | LAB
U--T | ug/g wet | Uranium Total | BPSY | MET-WET-BPSY-CCMS-VA | Metals in Animal Tissue and Vegetation (Biota) - Prescriptive. EPA 200.3/6020B This method is conducted following British Columbia Lab Manual method "Metals in Animal Tissue and Vegetation (Biota) - Prescriptive". Tissue samples are digested with nitric and hydrochloric acids; in combination with addition of hydrogen peroxide. Instrumental analysis is by collision cell inductively coupled plasma - mass spectrometry (modified from EPA Method 6020B). | 0.0004 | Y | Y | Y | 2 | 11/7/2019 | 11/7/2019 | 6/10/2021 1:57:10 PM | 2021-06-04 |  | EMS |  | MassPerMass | µg/g | Y | (mass conc.) | (wet) | Uranium Total (wet) (mass conc.) | 7440-61-1 | Metals - Total; | Metals - Total; EMS code: U--T | Chemical | Total | LAB
U--T | ug/m3 | Uranium Total | F088 | Aliquot,HNO3/HCl dig, ICPMS,loading | Aliquot microwave digested with 3:1 HNO3/HCL. Analysis by ICPMS. For HiVol samples preliminary result are reported in mg/L and loading results are reported in ug/m3 |  | Y | Y | Y | 565 | 8/17/2000 10:46:00 AM | 12/14/2006 | 12/30/2006 1:01:38 AM | 2000-08-16 |  | NPEPPIN |  | AirConcentration | µg/m³ | Y | (air conc.) |  | Uranium Total (air conc.) | 7440-61-1 | Metals - Total; | Metals - Total; EMS code: U--T | Chemical | Total | LAB
U--T | ug/g | Uranium Total | ICPM | Metals in Soil by CRC ICPMS | This analysis is carried out using procedures from CSR Analytical Method: "Strong Acid Leachable Metals (SALM) in Soil", BC Ministry of Environment, 26 June 2009, and procedures adapted from EPA Method 200.2. The sample is manually homogenized, dried at 60 degrees Celsius, sieved through a 63 um (230 mesh) sieve, and a representative subsample of the dry material is weighed. The sample is then digested at 95 degrees Celsius for 2 hours by block digester using concentrated nitric and hydrochloric acids. Instrumental analysis is by collision cell inductively coupled plasma - mass spectrometry (modifed from EPA Method 6020A). Method Limitation: This method is not a total digestion technique. It is a very strong acid digestion that is intended to dissolve those metals that may be environmentally available. By design, elements bound in silicate structures are not normally dissolved by this procedure as they are not usually mobile in the environment. | 0.05 | Y | Y | Y | 103 | 9/19/2017 | 9/24/2023 | 1/12/2024 11:16:52 AM | 2018-01-09 |  | JKROGH |  | MassPerMass | mg/kg | Y | (mass conc.) |  | Uranium Total (mass conc.) | 7440-61-1 | Metals - Total; | Metals - Total; EMS code: U--T | Chemical | Total | LAB
U--T | ug/g | Uranium Total | X551 | Metals in Biota (dry), total, acid digestion, ICP-SFMS | 0.1 gram of dried, homogenized biota is weighed in a microwave TFM vessel, and 8mL of Nitric Acid and 1mL of Hydrogen Peroxide are added. The sample is digested at 200 degrees C for 15 minutes in a closed vessel high pressure microwave oven. The sample is cooled, the volume adjusted to 100mL and introduced to an iductively coupled argon plasma-sector field mass spectrometer (ICP-SFMS) by pneumatic nebulization. Each respective element is measured at a specific mass to charge ratio, m/z value expressed in amu (atomic mass units), and at a specified instrument resolution. LR: Low Resolution ~300, MR: Mid Resolution ~4000 and HR: Hich Resolution ~10000. Ref: Method number 02-2705, National Laboratory for Environmental Testing, CCIW, Environment Canada, Burlington. Requested April 2003. | 0.001 | Y | Y | Y | 46 | 8/11/2015 | 10/30/2023 | 11/28/2023 2:21:33 PM | 2004-01-19 |  | UMULLER |  | MassPerMass | mg/kg | Y | (mass conc.) |  | Uranium Total (mass conc.) | 7440-61-1 | Metals - Total; | Metals - Total; EMS code: U--T | Chemical | Total | LAB
U--T | mg/L | Uranium Total | MSEA | Metals in Seawater (undigested) by (CRC) ICP-MS | This method describes the multi-elemental determination of trace elements by Collision/Reaction Cell (CRC) ICP-MS. The method measures ions produced by a radio-frequency inductively coupled plasma. Analyte species originating in a liquid are pneumatically nebulized and the resulting aerosol transported by argon gas into the plasma torch. The ions produced are entrained in the plasma gas and extracted, by means of a differentially pumped vacuum interface, into a mass spectrometer. The ions flow through a Collision/Reaction Cell (CRC) where depending on the mode of analysis either the ions passes through the cell unhindered (no gas - classical ICP-MS mode), or collisions/reactions chemistry takes place between the ions and a neutral gas such as hydrogen or helium. Any polyatomic ions which give interference on the analytes are removed. The ions are then sorted according to their mass-to-charge ratios by a quadrupole mass spectrometer. The ions are then sorted according to their mass-to-charge ratios by a quadrupole mass spectrometer. The ions transmitted through the quadrupole are quantified by a channel electron multiplier and the ion information processed by a data handling system. Although most polyatomic interference is removed, any additional interference must be assessed and valid corrections applied or the data flagged to indicate problems. Interference correction must include compensation for background ions contributed by the plasma gas, reagents, and constituents of the sample matrix. Instrument drift as well as suppressions or enhancements of instrument response caused by the sample matrix must be corrected for by the use of internal standards. | 6e-06 | Y | Y | Y | 441 | 8/12/2009 12:08:00 AM | 4/29/2024 | 6/7/2024 10:03:16 AM | 2009-09-08 |  | RDGIBSON |  | Concentration | mg/L | Y | (fl. conc.) |  | Uranium Total (fl. conc.) | 7440-61-1 | Metals - Total; | Metals - Total; EMS code: U--T | Chemical | Total | LAB
U--T | mg/L | Uranium Total | EQ01 | EQUIS (mg/L) | EQUIS (mg/L) | 0.0001 | N | Y | Y | 38 |  |  | 10/22/1996 11:20:29 PM | 1996-04-12 | 6/3/1998 6:51:40 PM | EMS |  | Concentration | mg/L | Y | (fl. conc.) |  | Uranium Total (fl. conc.) | 7440-61-1 | Metals - Total; | Metals - Total; EMS code: U--T | Chemical | Total | LAB
U--T | mg/L | Uranium Total | F068 | HNO3dig,TeflonTube/Autoclave,ICP-MS | This method is applicable to determiniation of dissolved, extractable, and total recoverable metals in ground waters, surface waters, drinking waters, and wastewaters. Dissolved, extractable, and total recoverable sample fractions are acidified to pH <1.5 with analytical grade nitric acid (HNO3) to prevent the metals from adsorbing to the container material or precipitating. Extractable and total recoverable water samples are collected in 250 ml acid washed polyethylene bottles and are preserved with Nitric acid (0.5 ml acid for 100 mLs of sample). All preserved samples are stored at room temperature and can be held for up to 6 months. Unfiltered, un-preserved samples are stored at 4C for no more than 30 days. For dissolved and extractable water fractions, no sample preparation is necessary. Samples are simply poured in autosampler tubes (i.e. 15 ml centrifuge tubes), and placed on the autosampler and analyzed. Samples for total recoverable metals are digested with Nitric acid in special Teflon tubes in an autoclave prior to analyses. | 2e-06 | Y | Y | Y | 1779 | 6/1/1999 1:33:00 PM | 7/4/2008 9:48:00 AM | 7/30/2024 6:32:12 AM | 1999-06-28 |  | NPEPPIN |  | Concentration | mg/L | Y | (fl. conc.) |  | Uranium Total (fl. conc.) | 7440-61-1 | Metals - Total; | Metals - Total; EMS code: U--T | Chemical | Total | LAB
U--T | mg/dm2/d | Uranium Total | DUST | HNO3 dig.(microwave or hot block),ICPMS | HNO3 dig.(microwave or hot block);ICPMS |  | Y | Y | Y | 267 | 7/19/2016 1:11:00 PM | 7/31/2024 | 7/31/2024 11:23:13 AM | 2008-02-07 |  | RDGIBSON |  | MassFlux | mg/dm²/d | Y | (mass flux) |  | Uranium Total (mass flux) | 7440-61-1 | Metals - Total; | Metals - Total; EMS code: U--T | Chemical | Total | LAB
U--T | mg/L | Uranium Total | X170 | ICP | This code is a catch-all for a large variety of ICP methods including water and also solids digested into solutions as an intermediate result. Different procedures will have greatly different detection limits. (a) For air particulates, preliminary results are reported in mg/L and loading results as mg/dm2/d. (b) For Hi Vol samples preliminiary results are same as reported in mg/L and loading results as ug/m3. (c) Note that for Rain samples there is no filtration nor digestion steps. | 1e-05 | Y | Y | Y | 3114 | 5/3/2002 | 5/19/2015 | 10/24/2022 10:58:25 AM | 1997-06-23 | 6/3/1998 6:51:47 PM | NPEPPIN |  | Concentration | mg/L | Y | (fl. conc.) |  | Uranium Total (fl. conc.) | 7440-61-1 | Metals - Total; | Metals - Total; EMS code: U--T | Chemical | Total | LAB
U--T | ug/g wet | Uranium Total | MTIW | MET-WET-CCMS-MID-VA | Metals in Tissue by CRC ICPMS (WET) Methodology: This method is conducted following British Columbia Lab Manual method "Metals in Animal Tissue and Vegetation (Biota) - Prescriptive". Tissue samples are homogenized and sub-sampled prior to hotblock digestion with nitric and hydrochloric acids, in combination with addition of hydrogen peroxide. Instrumental analysis is by collision cell inductively coupled plasma - mass spectrometry (modified from EPA Method 6020A). Method Limitation: This method employs a strong acid/peroxide digestion, and is intended to provide a conservative estimate of bio-available metals. Near complete recoveries are achieved for most toxicologically important metals, but elements associated with recalcitrant minerals may be only partially recovered. | 0.0004 | Y | Y | Y | 24 | 11/25/2018 | 12/6/2018 | 7/21/2021 2:55:08 PM | 2018-12-24 |  | EMS |  | MassPerMass | µg/g | Y | (mass conc.) | (wet) | Uranium Total (wet) (mass conc.) | 7440-61-1 | Metals - Total; | Metals - Total; EMS code: U--T | Chemical | Total | LAB
U--T | ug/m3 | Uranium Total | 6040 | Aqua Regia Dig:ICP | Aqua regia digestion is used to bring the metals into solution prior to ICP analysis. For Hi-Vol samples, preliminary results are reported as mg/L and24 hr loadings as ug/m3 | 5e-05 | Y | Y | Y | 437 | 2/13/2002 | 2/9/2005 | 2/23/2005 1:13:48 AM | 2004-02-16 |  | UMULLER |  | AirConcentration | µg/m³ | Y | (air conc.) |  | Uranium Total (air conc.) | 7440-61-1 | Metals - Total; | Metals - Total; EMS code: U--T | Chemical | Total | LAB
U--T | mg/L | Uranium Total | METN | MET-T-NP-CCMS-VA | Ultra trace metals in water are analyzed by CRC ICPMS, based on US EPA Method 6020B (July 2014). The detection limits provided can only be met for undigested samples. This procedure is intended for colorless, non-turbid, acid-preserved water samples (i.e. pristine water samples), having turbidity < 1 NTU and no odor. Where turbidity exceeds 1 NTU, and/or the sample is colored and has an odor, results may be biased low compared to true Total Metals concentrations. ALS recommends that turbidity analysis be requested on samples submitted for this test to aid with interpretation of results. Where turbidity is <1NTU, undigested metals are equivalent to total metals concentrations. | 2e-06 | Y | Y | Y | 3648 | 4/3/2012 | 8/19/2024 | 8/26/2024 1:03:06 PM | 2019-08-20 |  | KATFRASE |  | Concentration | mg/L | Y | (fl. conc.) |  | Uranium Total (fl. conc.) | 7440-61-1 | Metals - Total; | Metals - Total; EMS code: U--T | Chemical | Total | LAB
U--T | ppm (S) | Uranium Total | 6042 | Aqua Regia by ICPMS | These were sediment/soils digested with Aqua-Regia (Nitric Acid & Hydrochloric Acids) and then analysis by ICPMS. Results would be reported in ppm s. | 0.005 | Y | Y | Y | 71 | 5/14/2022 | 7/27/2024 | 7/29/2024 9:13:08 AM | 2022-06-28 |  | EMS |  | DimensionlessRatio | ppm | N | (ratio) |  | Uranium Total (ratio) | 7440-61-1 | Metals - Total; | Metals - Total; EMS code: U--T | Chemical | Total | LAB
U--T | mg/L | Uranium Total | F121 | Acid,ICPMS | Acidified, no digestion, ICPMS | 2e-06 | Y | Y | Y | 8466 | 3/12/2010 7:57:00 PM | 8/15/2024 9:10:00 PM | 8/26/2024 8:06:03 AM | 2002-07-22 | 7/16/2012 3:32:17 PM | UMULLER | RDGIBSON | Concentration | mg/L | Y | (fl. conc.) |  | Uranium Total (fl. conc.) | 7440-61-1 | Metals - Total; | Metals - Total; EMS code: U--T | Chemical | Total | LAB
U--T | ppm (S) | Uranium Total | MS41 | ALS Minerals ME-MS41 | ALS Minerals ME-MS41. A prepared sample is digested with aqua regia in a graphite heating block. After cooling; the resulting solution is diluted to with deionized water; mixed and analyzed by ICP-AES. Samples are then analyzed by ICP-MS for the remaining suite of elements. | 0.05 | Y | Y | Y | 0 |  |  |  | 2024-01-29 |  | EMS |  | DimensionlessRatio | ppm | N | (ratio) |  | Uranium Total (ratio) | 7440-61-1 | Metals - Total; | Metals - Total; EMS code: U--T | Chemical | Total | LAB
U--T | mg/L | Uranium Total | F082 | Aliquot,HNO3/HCL dig,ICPMS | Aliquot microwave digested with 3:1 HNO3/HCL. Analysis by ICPMS. For HiVol samples preliminary result are reported in mg/L and loading results are reported in ug/m3 | 0.0002 | Y | Y | Y | 26817 | 7/26/2005 7:04:00 AM | 8/23/2024 | 8/26/2024 4:53:05 PM | 2000-06-13 | 8/15/2019 11:12:50 AM | NPEPPIN | KATFRASE | Concentration | mg/L | Y | (fl. conc.) |  | Uranium Total (fl. conc.) | 7440-61-1 | Metals - Total; | Metals - Total; EMS code: U--T | Chemical | Total | LAB
U--T | ug/g wet | Uranium Total | X552 | Metals in Biota (wet), total, acid digestion, ICP-SFMS | One gram of wet, homogenized biota is weighed in a microwave TFM vessel, and 8mL of Nitric Acid and 1mL of Hydrogen Peroxide are added. The sample is digested at 200 degrees C for 15 minutes in a closed vessel high pressure microwave oven. The sample is cooled, the volume adjusted to 100mL and introduced to an inductively coupled argon plasma-sector field mass spectrometer (ICP-SFMS) by pneumatic nebulization. Each respective element is measured at a specific mass to charge ratio, m/z value expressed in amu (atomic mass units), and at a specified instrument resolution. LR: Low Resolution ~300, MR: Mid Resolution ~4000. Ref: 02-2704, National Laboratory for Environmental Testing, CCIW, Environment Canada, Burlington. Requested April 2003. | 0.0001 | Y | Y | Y | 37 | 8/9/2016 | 10/16/2023 | 11/28/2023 2:21:33 PM | 2004-01-19 | 4/19/2021 1:22:00 PM | UMULLER | EMS | MassPerMass | µg/g | Y | (mass conc.) | (wet) | Uranium Total (wet) (mass conc.) | 7440-61-1 | Metals - Total; | Metals - Total; EMS code: U--T | Chemical | Total | LAB
U--T | mg/L | Uranium Total | X561 | Total metals: FA, HNO3, HR-ICPMS (No Prep)Prep) | Ultra trace metals in water are analyzed by high resolution inductively coupled plasma mass spectrometry (HR-ICPMS), modified from US EPA 200.8, (Revision 5.5). The detection limits provided can only be met for undigested samples. This procedure is intended for colorless, non-turbid, acid-preserved water samples (i.e. pristine water samples), having turbidity < 1 NTU and no odor. Where turbidity exceeds 1 NTU, and/or the sample is colored and has an odor, results may be biased low compared to true Total Metals concentrations. ALS recommends that turbidity analysis be requested on samples submitted for this test to aid with interpretation of results. | 2e-06 | Y | Y | Y | 2756 | 6/8/2015 | 10/5/2023 | 1/31/2024 2:37:41 PM | 2015-10-22 |  | RDGIBSON |  | Concentration | mg/L | Y | (fl. conc.) |  | Uranium Total (fl. conc.) | 7440-61-1 | Metals - Total; | Metals - Total; EMS code: U--T | Chemical | Total | LAB
U--T | ug/g | Uranium Total | FS82 | PBM,dry60C,10mesh,SALM-ICP-MS | PBM;dry60C;10mesh;SALM-ICP-MS | 0.05 | Y | Y | Y | 1212 | 6/20/2005 | 7/24/2024 | 7/29/2024 5:17:09 PM | 2002-11-12 |  | UMULLER |  | MassPerMass | mg/kg | Y | (mass conc.) |  | Uranium Total (mass conc.) | 7440-61-1 | Metals - Total; | Metals - Total; EMS code: U--T | Chemical | Total | LAB
U--T | ug/g | Uranium Total | MTID | MET-DRY-CCMS-VA | Metals in Tissue by CRC ICPMS (DRY) Methodology: This method is conducted following British Columbia Lab Manual method "Metals in Animal Tissue and Vegetation (Biota) - Prescriptive". Tissue samples are homogenized and sub-sampled prior to hotblock digestion with nitric and hydrochloric acids, in combination with addition of hydrogen peroxide. Instrumental analysis is by collision cell inductively coupled plasma - mass spectrometry (modified from EPA Method 6020A). Method Limitation: This method employs a strong acid/peroxide digestion, and is intended to provide a conservative estimate of bio-available metals. Near complete recoveries are achieved for most toxicologically important metals, but elements associated with recalcitrant minerals may be only partially recovered. | 0.002 | Y | Y | Y | 122 | 8/13/2014 | 9/8/2022 | 10/26/2022 7:14:15 AM | 2015-03-09 | 3/11/2015 4:12:05 PM | RDGIBSON | EMS | MassPerMass | mg/kg | Y | (mass conc.) |  | Uranium Total (mass conc.) | 7440-61-1 | Metals - Total; | Metals - Total; EMS code: U--T | Chemical | Total | LAB
U-TCLP | mg/L | Uranium TCLP (Toxicity Characteristic Leaching Procedure) | FS82 | PBM,dry60C,10mesh,SALM-ICP-MS | PBM;dry60C;10mesh;SALM-ICP-MS | 0.02 | Y | Y | Y | 0 |  |  |  | 2024-01-24 |  | RDGIBSON |  | Concentration | mg/L | Y | (fl. conc.) |  | Uranium TCLP (Toxicity Characteristic Leaching Procedure) (fl. conc.) | 7440-61-1 | Metals - Extractable; | Metals - Extractable; EMS code: U-TCLP | Chemical |  | LAB
UND1 | mg/L | Undecanoic Acid | RFAS | Separation of resin and fatty acids (RFAs) in water by GC/MS | Scientific Principle The separation of resin and fatty acids (RFAs) in water samples is performed using gas chromatography/mass spectrometry (GC/MS) in the Selected Ion Monitoring (SIM) mode. The separation is based on compound size; boiling point and polarity. Analyte concentration is calculated using the internal standard method. RFA compounds are extracted from water samples using a liquid-liquid solvent extraction with DCM. Extraction occurs after the sample is brought to a pH of ~ 2 and placed in a rotary tumbler (at approx. 40 rpm) for 1 hour. The resulting extract is transferred to a vial; dried with acidified baked sodium sulfate; methylated with diazomethane and concentrated to 1mL under nitrogen prior to analysis. The resulting extracts are vialed and fortified with labeled internal standard compounds prior to analysis by gas chromatography with a (50%-Phenyl)-methylpolysiloxane (DB-17) column and a mass spectrometer detector in SIM Mode. | 0.006 | Y | Y | Y | 1 | 9/17/2021 7:49:00 PM | 9/17/2021 7:49:00 PM | 2/1/2022 2:38:14 PM | 2022-01-19 |  | EMS |  | Concentration | mg/L | Y | (fl. conc.) |  | Undecanoic Acid (fl. conc.) | 112-37-8 | Organics; | Organics; EMS code: UND1 | Chemical |  | LAB
UNH3 | mg/L | Un-ionized Ammonia | CALN | Un-ionized Ammonia Calculation | The concentration of unionized ammonia in the water referred to in subsection (1) must be determined in accordance with the following formula: total ammonia × 1 ÷ (1 +10^(pKa pH)) where total ammonia is the concentration of total ammonia in that water namely, unionized ammonia (NH3) plus ionized ammonia (NH4+) determined in accordance with subsection (4), expressed in mg/L as nitrogen (N) pKa is 0.09018 + 2729.92/T, where T is the ambient water temperature in kelvin and pH is the pH of that water. REF: from the new National Wastewater Systems Effluent Regulations | 0.001 | Y | Y | Y | 1411 | 2/8/2013 10:26:00 AM | 8/21/2024 1:37:00 PM | 8/26/2024 3:51:07 PM | 2013-02-25 |  | RDGIBSON |  | Concentration | mg/L | Y | (fl. conc.) |  | Un-ionized Ammonia (fl. conc.) | 7664-41-7 | Nutrients; | Nutrients; EMS code: UNH3 | Chemical |  | LAB
UV01 | AU/cm | UV Absorbance 250nm | F126 | UV Measurements of DOC | This procedure is used to estimate the amount of dissolved organic carbon (DOC) in waters and wastewaters. Some organic compounds regularly found in waters have strong UV absorptions. Thus there is a strong relationship between the UV absorption and the organic carbon content of the water. DOC is usually measured at 254nm, though other wavelengths are often chosen to minimize any interference in the sample. In this method an absorption scan of 250nm, 254nm, 310nm, 340nm, 360nm, and 365nm is read. In an average sample the absorption should decrease with the increase of the wavelength. However, irregularities in this trend could indicate some source of UV absorbing interference. | 0.01 | Y | Y | Y | 410 | 11/25/2002 | 8/7/2008 | 10/7/2008 2:21:42 PM | 2004-04-26 |  | UMULLER |  | Apperance | AU/cm | N | (app.) |  | UV Absorbance 250nm (app.) |  | Physical Properties; | Physical Properties; EMS code: UV01 | Physical |  | LAB
UV02 | AU/cm | UV Absorbance 254nm | F126 | UV Measurements of DOC | This procedure is used to estimate the amount of dissolved organic carbon (DOC) in waters and wastewaters. Some organic compounds regularly found in waters have strong UV absorptions. Thus there is a strong relationship between the UV absorption and the organic carbon content of the water. DOC is usually measured at 254nm, though other wavelengths are often chosen to minimize any interference in the sample. In this method an absorption scan of 250nm, 254nm, 310nm, 340nm, 360nm, and 365nm is read. In an average sample the absorption should decrease with the increase of the wavelength. However, irregularities in this trend could indicate some source of UV absorbing interference. | 0.01 | Y | Y | Y | 426 | 11/25/2002 | 2/12/2021 | 2/16/2021 3:07:09 PM | 2004-04-26 |  | UMULLER |  | Apperance | AU/cm | N | (app.) |  | UV Absorbance 254nm (app.) |  | Physical Properties; | Physical Properties; EMS code: UV02 | Physical |  | LAB
UV02 | % Transmittance | UV Absorbance 254nm | UVTTU | UVT Test, unfiltered, spectrophotometer at 254nm | UV-absorbing organic constituents in a sample absorb light in proportion to their concentration. The transmissivity (%) of these samples are measured at a standard UV wavelength of 254nm. Samples are unfiltered. The APHA method source is 5910.B Ultraviolet Absorption Method. | 0.1 | Y | Y | Y | 5 | 8/30/2022 4:51:00 PM | 7/27/2024 1:42:00 PM | 8/2/2024 5:01:07 PM | 2022-09-06 | 9/6/2022 12:05:02 PM | EMS | EMS | DimensionlessRatio | % | N | (ratio) | (transmittance) | UV Absorbance 254nm (transmittance) (ratio) |  | Physical Properties; | Physical Properties; EMS code: UV02 | Physical |  | LAB
UV02 | % Transmittance | UV Absorbance 254nm | UVTT | UVT Test, filtered, spectrophotometer at 254nm | UV-absorbing organic constituents in a sample absorb light in proportion to their concentration. The transmissivity (%) of these samples are measured at a standard UV wavelength of 254nm. Samples are filtered to control variations in UV absorption caused by particulates. The APHA method source is 5910.B Ultraviolet Absorption Method. | 0.1 | Y | Y | Y | 728 | 7/15/2010 4:57:00 PM | 7/27/2024 1:42:00 PM | 8/2/2024 5:01:07 PM | 2010-12-23 | 12/23/2010 9:39:54 AM | RDGIBSON | RDGIBSON | DimensionlessRatio | % | N | (ratio) | (transmittance) | UV Absorbance 254nm (transmittance) (ratio) |  | Physical Properties; | Physical Properties; EMS code: UV02 | Physical |  | LAB
UV03 | AU/cm | UV Absorbance 310nm | F126 | UV Measurements of DOC | This procedure is used to estimate the amount of dissolved organic carbon (DOC) in waters and wastewaters. Some organic compounds regularly found in waters have strong UV absorptions. Thus there is a strong relationship between the UV absorption and the organic carbon content of the water. DOC is usually measured at 254nm, though other wavelengths are often chosen to minimize any interference in the sample. In this method an absorption scan of 250nm, 254nm, 310nm, 340nm, 360nm, and 365nm is read. In an average sample the absorption should decrease with the increase of the wavelength. However, irregularities in this trend could indicate some source of UV absorbing interference. | 0.01 | Y | Y | Y | 410 | 11/25/2002 | 8/7/2008 | 10/7/2008 2:21:42 PM | 2004-04-26 |  | UMULLER |  | Apperance | AU/cm | N | (app.) |  | UV Absorbance 310nm (app.) |  | Physical Properties; | Physical Properties; EMS code: UV03 | Physical |  | LAB
UV04 | AU/cm | UV Absorbance 340nm | F126 | UV Measurements of DOC | This procedure is used to estimate the amount of dissolved organic carbon (DOC) in waters and wastewaters. Some organic compounds regularly found in waters have strong UV absorptions. Thus there is a strong relationship between the UV absorption and the organic carbon content of the water. DOC is usually measured at 254nm, though other wavelengths are often chosen to minimize any interference in the sample. In this method an absorption scan of 250nm, 254nm, 310nm, 340nm, 360nm, and 365nm is read. In an average sample the absorption should decrease with the increase of the wavelength. However, irregularities in this trend could indicate some source of UV absorbing interference. | 0.01 | Y | Y | Y | 410 | 11/25/2002 | 8/7/2008 | 10/7/2008 2:21:42 PM | 2004-04-26 |  | UMULLER |  | Apperance | AU/cm | N | (app.) |  | UV Absorbance 340nm (app.) |  | Physical Properties; | Physical Properties; EMS code: UV04 | Physical |  | LAB
UV05 | AU/cm | UV Absorbance 360nm | F126 | UV Measurements of DOC | This procedure is used to estimate the amount of dissolved organic carbon (DOC) in waters and wastewaters. Some organic compounds regularly found in waters have strong UV absorptions. Thus there is a strong relationship between the UV absorption and the organic carbon content of the water. DOC is usually measured at 254nm, though other wavelengths are often chosen to minimize any interference in the sample. In this method an absorption scan of 250nm, 254nm, 310nm, 340nm, 360nm, and 365nm is read. In an average sample the absorption should decrease with the increase of the wavelength. However, irregularities in this trend could indicate some source of UV absorbing interference. | 0.01 | Y | Y | Y | 410 | 11/25/2002 | 8/7/2008 | 10/7/2008 2:21:42 PM | 2004-04-26 |  | UMULLER |  | Apperance | AU/cm | N | (app.) |  | UV Absorbance 360nm (app.) |  | Physical Properties; | Physical Properties; EMS code: UV05 | Physical |  | LAB
UV06 | AU/cm | UV Absorbance 365nm | F126 | UV Measurements of DOC | This procedure is used to estimate the amount of dissolved organic carbon (DOC) in waters and wastewaters. Some organic compounds regularly found in waters have strong UV absorptions. Thus there is a strong relationship between the UV absorption and the organic carbon content of the water. DOC is usually measured at 254nm, though other wavelengths are often chosen to minimize any interference in the sample. In this method an absorption scan of 250nm, 254nm, 310nm, 340nm, 360nm, and 365nm is read. In an average sample the absorption should decrease with the increase of the wavelength. However, irregularities in this trend could indicate some source of UV absorbing interference. | 0.01 | Y | Y | Y | 467 | 11/25/2002 | 8/7/2008 | 1/30/2019 4:06:39 PM | 2004-04-26 |  | UMULLER |  | Apperance | AU/cm | N | (app.) |  | UV Absorbance 365nm (app.) |  | Physical Properties; | Physical Properties; EMS code: UV06 | Physical |  | LAB
UV07 | AU/cm | UV Absorbance 510 nm | F126 | UV Measurements of DOC | This procedure is used to estimate the amount of dissolved organic carbon (DOC) in waters and wastewaters. Some organic compounds regularly found in waters have strong UV absorptions. Thus there is a strong relationship between the UV absorption and the organic carbon content of the water. DOC is usually measured at 254nm, though other wavelengths are often chosen to minimize any interference in the sample. In this method an absorption scan of 250nm, 254nm, 310nm, 340nm, 360nm, and 365nm is read. In an average sample the absorption should decrease with the increase of the wavelength. However, irregularities in this trend could indicate some source of UV absorbing interference. | 0.001 | Y | Y | Y | 1 | 6/23/2018 | 6/23/2018 | 8/9/2018 2:31:11 PM | 2018-08-09 |  | RDGIBSON |  | Apperance | AU/cm | N | (app.) |  | UV Absorbance 510 nm (app.) |  | Physical Properties; | Physical Properties; EMS code: UV07 | Physical |  | LAB
UVSS | mg/L | Unidentified Volatile Selenium Species (As Se) | E611 | E611S: EPA 8260D (mod) | E611S: Dimethyl selenide and dimethyl diselenide are analyzed by static headspace GC-MS. Samples are prepared in headspace vials and are heated and agitated on the headspace autosampler; causing VOCs to partition between the aqueous phase and the headspace in accordance with Henry`s law. | 2e-06 | Y | Y | Y | 96 | 10/31/2023 | 8/10/2024 | 8/19/2024 12:19:03 PM | 2023-12-12 |  | EMS |  | Concentration | mg/L | Y | (fl. conc.) |  | Unidentified Volatile Selenium Species (As Se) (fl. conc.) |  | Metals; | Metals; EMS code: UVSS | Chemical |  | LAB
UVT5 | %T/cm | UV Transmittance 510 nm | CALT | Calculation (%T/cm) | Calculation (% Transmittance per centimetre) | 1 | Y | Y | Y | 1 | 7/9/2018 | 7/9/2018 | 8/9/2018 2:31:11 PM | 2018-08-09 |  | RDGIBSON |  | Apperance | %T/cm | N | (app.) |  | UV Transmittance 510 nm (app.) |  | Physical Properties; | Physical Properties; EMS code: UVT5 | Physical |  | LAB
UVTR | %T/cm | UV Transmittance 254nm | UVTT | UVT Test, filtered, spectrophotometer at 254nm | UV-absorbing organic constituents in a sample absorb light in proportion to their concentration. The transmissivity (%) of these samples are measured at a standard UV wavelength of 254nm. Samples are filtered to control variations in UV absorption caused by particulates. The APHA method source is 5910.B Ultraviolet Absorption Method. | 1 | Y | Y | Y | 1 | 5/19/2024 | 5/19/2024 | 5/30/2024 1:43:07 PM | 2024-05-29 |  | RDGIBSON |  | Apperance | %T/cm | N | (app.) |  | UV Transmittance 254nm (app.) |  | Physical Properties; | Physical Properties; EMS code: UVTR | Physical |  | LAB
UVTR | %T/cm | UV Transmittance 254nm | CALT | Calculation (%T/cm) | Calculation (% Transmittance per centimetre) | 0 | Y | Y | Y | 2 | 4/22/2018 | 5/27/2018 | 8/9/2018 2:50:24 PM | 2017-08-29 |  | RDGIBSON |  | Apperance | %T/cm | N | (app.) |  | UV Transmittance 254nm (app.) |  | Physical Properties; | Physical Properties; EMS code: UVTR | Physical |  | LAB
V--D | mg/L | Vanadium Dissolved | METC | MET-D-NP-CCMS-VA | Ultra trace metals in water are analyzed by CRC ICPMS, based on US EPA Method 6020B (July 2014). This procedure is intended for pristine field-filtered acid-preserved water samples. The detection limits (LOR) for this test are based on lab instrumental analysis only, not including filtration. ALS-supplied field filtration equipment does not support these LOR. Therefore, because of the high probability of false positives due to filtration, it is strongly recommended that a filtration blank be analysed to aid in data interpretation. | 0.0002 | Y | Y | Y | 1776 | 4/3/2012 | 8/15/2024 | 8/23/2024 9:50:35 AM | 2019-08-20 |  | KATFRASE |  | Concentration | mg/L | Y | (fl. conc.) |  | Vanadium Dissolved (fl. conc.) | 7440-62-2 | Metals - Dissolved; | Metals - Dissolved; EMS code: V--D | Chemical | Dissolved | LAB
V--D | mg/L | Vanadium Dissolved | EQ01 | EQUIS (mg/L) | EQUIS (mg/L) | 0.001 | N | Y | Y | 922 |  |  | 10/23/1996 12:07:04 AM | 1996-04-12 | 6/3/1998 6:51:40 PM | EMS |  | Concentration | mg/L | Y | (fl. conc.) |  | Vanadium Dissolved (fl. conc.) | 7440-62-2 | Metals - Dissolved; | Metals - Dissolved; EMS code: V--D | Chemical | Dissolved | LAB
V--D | mg/L | Vanadium Dissolved | X562 | Dissolved metals: FF, FA, HNO3, HR-ICPMS | Ultra trace metals in water are analyzed by high resolution inductively coupled plasma mass spectrometry (HR-ICPMS), based on US EPA Method 200.8, (Oct. 1999). This procedure is intended for pristine field-filtered acid-preserved water samples. The detection limits (LOR) for this test are based on lab instrumental analysis only, not including filtration. ALS-supplied field filtration equipment does not support these LOR. Therefore, because of the high probability of false positives due to filtration, it is strongly recommended that a filtration blank be analyzed to aid in data interpretation. | 0.0002 | Y | Y | Y | 1264 | 6/8/2015 | 10/19/2023 | 11/9/2023 7:41:27 AM | 2015-10-22 |  | RDGIBSON |  | Concentration | mg/L | Y | (fl. conc.) |  | Vanadium Dissolved (fl. conc.) | 7440-62-2 | Metals - Dissolved; | Metals - Dissolved; EMS code: V--D | Chemical | Dissolved | LAB
V--D | mg/L | Vanadium Dissolved | MSEA | Metals in Seawater (undigested) by (CRC) ICP-MS | This method describes the multi-elemental determination of trace elements by Collision/Reaction Cell (CRC) ICP-MS. The method measures ions produced by a radio-frequency inductively coupled plasma. Analyte species originating in a liquid are pneumatically nebulized and the resulting aerosol transported by argon gas into the plasma torch. The ions produced are entrained in the plasma gas and extracted, by means of a differentially pumped vacuum interface, into a mass spectrometer. The ions flow through a Collision/Reaction Cell (CRC) where depending on the mode of analysis either the ions passes through the cell unhindered (no gas - classical ICP-MS mode), or collisions/reactions chemistry takes place between the ions and a neutral gas such as hydrogen or helium. Any polyatomic ions which give interference on the analytes are removed. The ions are then sorted according to their mass-to-charge ratios by a quadrupole mass spectrometer. The ions are then sorted according to their mass-to-charge ratios by a quadrupole mass spectrometer. The ions transmitted through the quadrupole are quantified by a channel electron multiplier and the ion information processed by a data handling system. Although most polyatomic interference is removed, any additional interference must be assessed and valid corrections applied or the data flagged to indicate problems. Interference correction must include compensation for background ions contributed by the plasma gas, reagents, and constituents of the sample matrix. Instrument drift as well as suppressions or enhancements of instrument response caused by the sample matrix must be corrected for by the use of internal standards. | 0.0002 | Y | Y | Y | 4 | 4/12/2023 | 4/29/2024 | 5/16/2024 2:00:14 PM | 2009-09-08 |  | RDGIBSON |  | Concentration | mg/L | Y | (fl. conc.) |  | Vanadium Dissolved (fl. conc.) | 7440-62-2 | Metals - Dissolved; | Metals - Dissolved; EMS code: V--D | Chemical | Dissolved | LAB
V--D | mg/L | Vanadium Dissolved | X556 | Metals in water, dissolved, "in-bottle" digestion, ICP-SFMS | The whole water sample is filtered through a 0.45 micron cellulose acetate filter paper and the filtrate is stored in a HDPE/Teflon bottle. The container is de-capped in a Class 100 workstation, acidified to 2% (v/v) HNO3 and 1% (v/v) HCl and recapped. The sample is "in-bottle" digested in a convection oven, set at 60 degrees C for 16 hours. The sample is cooled and introduced to an inductively coupled argon plasma-sector field mass spectrometer (ICP-SFMS) by pneumatic nebulization. Each respective element is measured at a specific mass to charge ratio, m/z value expressed in amu (atomic mass units), and at a specified intrument resolution. LR: Low Resolution ~300, MR: Mid Resolution ~4000 and HR: High Resolution ~10000. Ref: Method number 02-2003, National Laboratory for Environmental Testing, CCIW, Environment Canada, Burlington, Requested April 2003. | 5e-06 | Y | Y | Y | 4234 | 6/11/2003 | 12/6/2023 2:04:00 AM | 2/14/2024 9:05:32 AM | 2004-01-19 |  | UMULLER |  | Concentration | mg/L | Y | (fl. conc.) |  | Vanadium Dissolved (fl. conc.) | 7440-62-2 | Metals - Dissolved; | Metals - Dissolved; EMS code: V--D | Chemical | Dissolved | LAB
V--D | mg/L | Vanadium Dissolved | 0150 | Ext(Cupf.):AA Anal. | Ext(Cupf.):AA Anal. | 0.001 | Y | Y | Y | 47 |  |  | 10/24/1996 11:57:35 AM | 1996-04-12 | 6/3/1998 6:51:46 PM | EMS |  | Concentration | mg/L | Y | (fl. conc.) |  | Vanadium Dissolved (fl. conc.) | 7440-62-2 | Metals - Dissolved; | Metals - Dissolved; EMS code: V--D | Chemical | Dissolved | LAB
V--D | mg/L | Vanadium Dissolved | 0160 | Dig:Ext(Cupf.):AA Anal. | Dig:Ext(Cupf.):AA Anal. | 0.001 | Y | Y | Y | 8 | 9/29/2000 | 12/1/2010 | 12/9/2010 3:55:27 PM | 1996-04-12 | 6/3/1998 6:51:40 PM | EMS |  | Concentration | mg/L | Y | (fl. conc.) |  | Vanadium Dissolved (fl. conc.) | 7440-62-2 | Metals - Dissolved; | Metals - Dissolved; EMS code: V--D | Chemical | Dissolved | LAB
V--D | mg/L | Vanadium Dissolved | X170 | ICP | This code is a catch-all for a large variety of ICP methods including water and also solids digested into solutions as an intermediate result. Different procedures will have greatly different detection limits. (a) For air particulates, preliminary results are reported in mg/L and loading results as mg/dm2/d. (b) For Hi Vol samples preliminiary results are same as reported in mg/L and loading results as ug/m3. (c) Note that for Rain samples there is no filtration nor digestion steps. | 0.03 | Y | Y | Y | 12653 | 4/11/1996 6:02:00 PM | 11/6/2018 | 6/12/2020 1:21:12 PM | 1996-04-12 | 6/3/1998 6:51:40 PM | EMS |  | Concentration | mg/L | Y | (fl. conc.) |  | Vanadium Dissolved (fl. conc.) | 7440-62-2 | Metals - Dissolved; | Metals - Dissolved; EMS code: V--D | Chemical | Dissolved | LAB
V--D | mg/L | Vanadium Dissolved | F039 | Field filt + acid, ICP-AES | Dissolved and extractable water samples are ready for analysis once collected. Total samples are autoclaved with HNO3 and hydrochloric acid (HCI) in capped Teflon vials. Analysis is performed via Inductively Coupled Argon Plasma Atomic Emission Spectrometry (ICP-AES)(simultaneous multielement analysis), Graphite Furnace Atomic Absorption Spectrometry (GF-AAS)(low level analysis) and/or Flame Atomic Absorption Spectrometry (F-AAS). | 0.01 | Y | Y | Y | 5962 | 7/29/1996 6:03:00 PM | 7/5/2019 1:21:00 PM | 9/11/2019 8:47:50 AM | 1997-05-08 | 6/3/1998 6:51:47 PM | MPAK |  | Concentration | mg/L | Y | (fl. conc.) |  | Vanadium Dissolved (fl. conc.) | 7440-62-2 | Metals - Dissolved; | Metals - Dissolved; EMS code: V--D | Chemical | Dissolved | LAB
V--D | mg/L | Vanadium Dissolved | MWID | Metals in Water by ICPMS (digested) | Metals in Water by ICPMS (digested) Based on EPA 200.8 | 0.005 | Y | Y | Y | 31767 | 10/26/2018 5:20:00 PM | 10/26/2018 5:20:00 PM | 7/31/2024 1:46:41 PM | 2009-12-24 |  | RDGIBSON |  | Concentration | mg/L | Y | (fl. conc.) |  | Vanadium Dissolved (fl. conc.) | 7440-62-2 | Metals - Dissolved; | Metals - Dissolved; EMS code: V--D | Chemical | Dissolved | LAB
V--D | mg/L | Vanadium Dissolved | X285 | HNO3 Dig:ICP Analysis | HNO3 Dig:ICP Analysis | 0.01 | Y | Y | Y | 650 | 1/30/2002 | 11/13/2002 | 2/19/2004 4:22:44 PM | 1998-05-22 | 6/3/1998 6:51:49 PM | NPEPPIN |  | Concentration | mg/L | Y | (fl. conc.) |  | Vanadium Dissolved (fl. conc.) | 7440-62-2 | Metals - Dissolved; | Metals - Dissolved; EMS code: V--D | Chemical | Dissolved | LAB
V--D | mg/L | Vanadium Dissolved | F032 | Filt.,Acid,ICPMS | There is no acid digestion required for this analysis. A water sample is collected and preserved in 0.2 percent nitric acid. To analyse the extractable metals the sample is shaken, allowed to sit for at least twenty four hours, and a decanted portion is introduced to the spectrometer. When the dissolved metals are required, the samples must first be filtered through a 0.45 micron cellulose acetate filter paper and then followed by preservation in 0.2 percent nitric acid. It is important to note that different laboratories may have very different detection limits for some variables analyzed by ICPMS. | 0.0005 | Y | Y | Y | 78007 | 3/10/1998 2:08:00 PM | 8/26/2024 | 8/26/2024 4:53:05 PM | 1997-10-20 | 8/15/2019 11:01:34 AM | NPEPPIN | KATFRASE | Concentration | mg/L | Y | (fl. conc.) |  | Vanadium Dissolved (fl. conc.) | 7440-62-2 | Metals - Dissolved; | Metals - Dissolved; EMS code: V--D | Chemical | Dissolved | LAB
V--D | mg/L | Vanadium Dissolved | F068 | HNO3dig,TeflonTube/Autoclave,ICP-MS | This method is applicable to determiniation of dissolved, extractable, and total recoverable metals in ground waters, surface waters, drinking waters, and wastewaters. Dissolved, extractable, and total recoverable sample fractions are acidified to pH <1.5 with analytical grade nitric acid (HNO3) to prevent the metals from adsorbing to the container material or precipitating. Extractable and total recoverable water samples are collected in 250 ml acid washed polyethylene bottles and are preserved with Nitric acid (0.5 ml acid for 100 mLs of sample). All preserved samples are stored at room temperature and can be held for up to 6 months. Unfiltered, un-preserved samples are stored at 4C for no more than 30 days. For dissolved and extractable water fractions, no sample preparation is necessary. Samples are simply poured in autosampler tubes (i.e. 15 ml centrifuge tubes), and placed on the autosampler and analyzed. Samples for total recoverable metals are digested with Nitric acid in special Teflon tubes in an autoclave prior to analyses. | 5e-05 | Y | Y | Y | 1398 | 6/1/1999 1:22:00 PM | 4/17/2015 1:06:00 PM | 8/11/2021 12:04:11 PM | 1999-06-28 |  | NPEPPIN |  | Concentration | mg/L | Y | (fl. conc.) |  | Vanadium Dissolved (fl. conc.) | 7440-62-2 | Metals - Dissolved; | Metals - Dissolved; EMS code: V--D | Chemical | Dissolved | LAB
V--D | mg/L | Vanadium Dissolved | X085 | LF,LA,HNO3,AA | Samples are lab-filtered (0.45 um), and lab-acidified (HNO3); analysis by AA. | 0.01 | Y | Y | Y | 83 | 9/27/2007 | 9/18/2015 | 10/28/2015 9:54:16 AM | 1996-04-12 | 6/3/1998 6:51:40 PM | EMS |  | Concentration | mg/L | Y | (fl. conc.) |  | Vanadium Dissolved (fl. conc.) | 7440-62-2 | Metals - Dissolved; | Metals - Dissolved; EMS code: V--D | Chemical | Dissolved | LAB
V--E | mg/L | Vanadium Extractable | F069 | SEM extr,6N HCl,ICP/OES | 6.0 N Hydrochloric acid (HCI) is added to an aliquot of the sediment sample. The extract produced from the addition of the acid is then analyzed for simultaneously extracted metals (SEM) using inductively coupled plasma - optical emission spectrophotometry. The detection limits depend on the metal parameters in question and the instrumentation used. | 0.03 | Y | Y | Y | 10 | 3/26/2012 3:01:00 AM | 3/26/2012 3:01:00 AM | 11/25/2021 1:50:15 PM | 1999-07-09 |  | NPEPPIN |  | Concentration | mg/L | Y | (fl. conc.) |  | Vanadium Extractable (fl. conc.) | 7440-62-2 | Metals - Extractable; | Metals - Extractable; EMS code: V--E | Chemical | Extractable | LAB
V--E | mg/L | Vanadium Extractable | F047 | 24hr HNO3 dil'n,ICP-MS | There is no acid digestion required for this analysis. A water sample is collected and preserved in 0.2 percent nitric acid. To analyse the extractable metals the sample is shaken, allowed to sit for at least twenty four hours, and a decanted portion is introduced to the spectrometer. When the dissolved metals are required, the samples must first be filtered through a 0.45 micron cellulose acetate filter paper and then followed by preservation in 0.2 percent nitric acid. It is important to note that different laboratories may have very different detection limits for some variables analyzed by ICP/MS. | 1e-05 | Y | Y | Y | 2676 | 4/16/1998 10:21:00 AM | 6/26/2003 | 9/11/2019 8:46:18 AM | 1997-10-20 | 6/3/1998 6:51:47 PM | NPEPPIN |  | Concentration | mg/L | Y | (fl. conc.) |  | Vanadium Extractable (fl. conc.) | 7440-62-2 | Metals - Extractable; | Metals - Extractable; EMS code: V--E | Chemical | Extractable | LAB
V--E | mg/L | Vanadium Extractable | X553 | Metals in water, extractable, ICP-SFMS | The whole water sample is collected in a HDPE/Teflon bottle. In a Class 100 workstation it is preserved with concentrated Nitric Acid at concentration of 0.2 percent (v/v). The acid preservative extracts the metals from material in the sample and a decanted portion is introduced directly to an inductively coupled argon plasma-sector field mass spectrometer (ICP-SFMS) by pneumatic nebulization. Each respective element is measured at a specific mass to charge ratio, m/z value expressed in amu (atomic mass units), and at a specified instrument resolution. LR: Low Resolution ~300, MR: Mid Resolution ~4000 and HR: High Resolution ~10000. Ref: Method number 02-2061, National Laboratory for Environmental Testing, CCIW, Environment Canada, Burlington. Requested April 2003. | 5e-06 | Y | Y | Y | 2597 | 5/22/2003 | 9/21/2007 | 11/28/2023 2:11:53 PM | 2004-01-19 |  | UMULLER |  | Concentration | mg/L | Y | (fl. conc.) |  | Vanadium Extractable (fl. conc.) | 7440-62-2 | Metals - Extractable; | Metals - Extractable; EMS code: V--E | Chemical | Extractable | LAB
V--E | mg/L | Vanadium Extractable | F040 | Field Acid, lab filt, ICP-AES | Dissolved and extractable water samples are ready for analysis once collected. Total samples are autoclaved with HNO3 and hydrochloric acid (HCI) in capped Teflon vials. Analysis is performed via Inductively Coupled Argon Plasma Atomic Emission Spectrometry (ICP-AES)(simultaneous multielement analysis), Graphite Furnace Atomic Absorption Spectrometry (GF-AAS)(low level analysis) and/or Flame Atomic Absorption Spectrometry (F-AAS). | 0.01 | Y | Y | Y | 2350 | 8/5/1997 6:06:00 PM | 3/14/2008 | 9/11/2019 8:47:41 AM | 1997-05-08 | 6/3/1998 6:51:47 PM | MPAK |  | Concentration | mg/L | Y | (fl. conc.) |  | Vanadium Extractable (fl. conc.) | 7440-62-2 | Metals - Extractable; | Metals - Extractable; EMS code: V--E | Chemical | Extractable | LAB
V--T | mg/L | Vanadium Total | X170 | ICP | This code is a catch-all for a large variety of ICP methods including water and also solids digested into solutions as an intermediate result. Different procedures will have greatly different detection limits. (a) For air particulates, preliminary results are reported in mg/L and loading results as mg/dm2/d. (b) For Hi Vol samples preliminiary results are same as reported in mg/L and loading results as ug/m3. (c) Note that for Rain samples there is no filtration nor digestion steps. | 0.003 | Y | Y | Y | 6106 | 3/28/1996 6:03:00 PM | 11/7/2018 | 10/24/2022 10:58:25 AM | 1996-04-12 | 6/3/1998 6:51:40 PM | EMS |  | Concentration | mg/L | Y | (fl. conc.) |  | Vanadium Total (fl. conc.) | 7440-62-2 | Metals - Total; | Metals - Total; EMS code: V--T | Chemical | Total | LAB
V--T | ug/g | Vanadium Total | X230 | Hom,HNO3/HCl,ICP | The sample is dried, ground to pass 100 mesh, and digested with a nitric-hydrochloric acid mixture to solubilize the solid matter and to remove the organic material by oxidation and volatilization. The completed digests are submitted to metals instrumentation personnel for metals analysis by Inductively Coupled Atomic Emission Spectrophotometry (ICP-AES), or hydride analysis by atomic absorption. | 1 | Y | Y | Y | 159 | 2/12/2002 | 12/30/2005 | 2/3/2006 1:13:13 AM | 1996-04-12 | 6/3/1998 6:51:40 PM | EMS |  | MassPerMass | mg/kg | Y | (mass conc.) |  | Vanadium Total (mass conc.) | 7440-62-2 | Metals - Total; | Metals - Total; EMS code: V--T | Chemical | Total | LAB
V--T | ug/m3 | Vanadium Total | F088 | Aliquot,HNO3/HCl dig, ICPMS,loading | Aliquot microwave digested with 3:1 HNO3/HCL. Analysis by ICPMS. For HiVol samples preliminary result are reported in mg/L and loading results are reported in ug/m3 |  | Y | Y | Y | 565 | 8/17/2000 10:46:00 AM | 12/14/2006 | 12/30/2006 1:01:38 AM | 2000-08-16 |  | NPEPPIN |  | AirConcentration | µg/m³ | Y | (air conc.) |  | Vanadium Total (air conc.) | 7440-62-2 | Metals - Total; | Metals - Total; EMS code: V--T | Chemical | Total | LAB
V--T | ppm (S) | Vanadium Total | 6041 | Aqua Regia by ICPMS | These were sediment/soils digested with Aqua-Regia (Nitric Acid & Hydrochloric Acids) and then analysis by ICPMS. Results would be reported in mg/kg and/or %. | 0.1 | Y | Y | Y | 21 | 10/13/2019 | 8/5/2020 | 8/10/2020 3:52:07 PM | 2019-09-06 | 10/21/2019 11:56:51 AM | KATFRASE | KATFRASE | DimensionlessRatio | ppm | N | (ratio) |  | Vanadium Total (ratio) | 7440-62-2 | Metals - Total; | Metals - Total; EMS code: V--T | Chemical | Total | LAB
V--T | mg/L | Vanadium Total | TCLP | Toxicity Characteristic Leaching Procedure | The TCLP procedure is designed to simulate the climatic leaching action | 0.1 | Y | Y | Y | 22 | 12/8/2009 7:49:00 PM | 9/28/2011 5:53:00 PM | 11/3/2011 11:10:21 PM | 2009-03-06 |  | JCHUDAK |  | Concentration | mg/L | Y | (fl. conc.) |  | Vanadium Total (fl. conc.) | 7440-62-2 | Metals - Total; | Metals - Total; EMS code: V--T | Chemical | Total | LAB
V--T | mg/L | Vanadium Total | 0042 | GI HNO3 Dig: ICP | Sample aliquot taken from unfiltered unpreserved General Ions bottle; HNO3 digestion; ICP analysis. | 0.003 | Y | Y | Y | 6261 | 9/20/1996 | 7/5/2019 3:21:00 PM | 7/12/2019 6:49:13 PM | 1996-04-12 | 6/3/1998 6:51:46 PM | EMS |  | Concentration | mg/L | Y | (fl. conc.) |  | Vanadium Total (fl. conc.) | 7440-62-2 | Metals - Total; | Metals - Total; EMS code: V--T | Chemical | Total | LAB
V--T | mg/L | Vanadium Total | MWID | Metals in Water by ICPMS (digested) | Metals in Water by ICPMS (digested) Based on EPA 200.8 | 0.005 | Y | Y | Y | 36812 |  |  | 7/31/2024 1:46:41 PM | 2009-12-24 |  | RDGIBSON |  | Concentration | mg/L | Y | (fl. conc.) |  | Vanadium Total (fl. conc.) | 7440-62-2 | Metals - Total; | Metals - Total; EMS code: V--T | Chemical | Total | LAB
V--T | mg/L | Vanadium Total | F038 | HNO3/HCl dig,ICP-AES | Dissolved and extractable water samples are ready for analysis once collected. Total samples are autoclaved with HNO3 and hydrochloric acid (HCI) in capped Teflon vials. Analysis is performed via Inductively Coupled Argon Plasma Atomic Emission Spectrometry (ICP-AES)(simultaneous multielement analysis), Graphite Furnace Atomic Absorption Spectrometry (GF-AAS) (low level analysis) and/or Flame Atomic Absorption Spectrometry (F-AAS). | 0.01 | Y | Y | Y | 5397 | 8/1/1996 6:04:00 PM | 5/6/2008 | 9/11/2019 8:47:43 AM | 1997-05-08 | 6/3/1998 6:51:47 PM | MPAK |  | Concentration | mg/L | Y | (fl. conc.) |  | Vanadium Total (fl. conc.) | 7440-62-2 | Metals - Total; | Metals - Total; EMS code: V--T | Chemical | Total | LAB
V--T | ppm (S) | Vanadium Total | MS81 | ALS Minerals: ME-MS81 | ALS Minerals: ME-MS81. A prepared sample (0.100 g) is added to lithium metaborate/lithium tetraborate flux; mixed well and fused in a furnace at 1025°C. The resulting melt is then cooled and dissolved in an acid mixture containing nitric; hydrochloric and hydrofluoric acids. This solution is then analyzed by inductively coupled plasma - mass spectrometry. | 5 | Y | Y | Y | 0 |  |  |  | 2024-01-29 |  | EMS |  | DimensionlessRatio | ppm | N | (ratio) |  | Vanadium Total (ratio) | 7440-62-2 | Metals - Total; | Metals - Total; EMS code: V--T | Chemical | Total | LAB
V--T | ug/m3 | Vanadium Total | 6040 | Aqua Regia Dig:ICP | Aqua regia digestion is used to bring the metals into solution prior to ICP analysis. For Hi-Vol samples, preliminary results are reported as mg/L and24 hr loadings as ug/m3 | 0.003 | Y | Y | Y | 437 | 2/13/2002 | 2/9/2005 | 2/23/2005 1:13:48 AM | 1998-05-05 | 6/3/1998 6:51:48 PM | NPEPPIN |  | AirConcentration | µg/m³ | Y | (air conc.) |  | Vanadium Total (air conc.) | 7440-62-2 | Metals - Total; | Metals - Total; EMS code: V--T | Chemical | Total | LAB
V--T | mg/L | Vanadium Total | F068 | HNO3dig,TeflonTube/Autoclave,ICP-MS | This method is applicable to determiniation of dissolved, extractable, and total recoverable metals in ground waters, surface waters, drinking waters, and wastewaters. Dissolved, extractable, and total recoverable sample fractions are acidified to pH <1.5 with analytical grade nitric acid (HNO3) to prevent the metals from adsorbing to the container material or precipitating. Extractable and total recoverable water samples are collected in 250 ml acid washed polyethylene bottles and are preserved with Nitric acid (0.5 ml acid for 100 mLs of sample). All preserved samples are stored at room temperature and can be held for up to 6 months. Unfiltered, un-preserved samples are stored at 4C for no more than 30 days. For dissolved and extractable water fractions, no sample preparation is necessary. Samples are simply poured in autosampler tubes (i.e. 15 ml centrifuge tubes), and placed on the autosampler and analyzed. Samples for total recoverable metals are digested with Nitric acid in special Teflon tubes in an autoclave prior to analyses. | 5e-05 | Y | Y | Y | 1779 | 6/1/1999 1:34:00 PM | 7/4/2008 9:48:00 AM | 7/30/2024 6:32:12 AM | 1999-06-28 |  | NPEPPIN |  | Concentration | mg/L | Y | (fl. conc.) |  | Vanadium Total (fl. conc.) | 7440-62-2 | Metals - Total; | Metals - Total; EMS code: V--T | Chemical | Total | LAB
V--T | mg/L | Vanadium Total | F082 | Aliquot,HNO3/HCL dig,ICPMS | Aliquot microwave digested with 3:1 HNO3/HCL. Analysis by ICPMS. For HiVol samples preliminary result are reported in mg/L and loading results are reported in ug/m3 | 0.0005 | Y | Y | Y | 29448 | 7/26/2005 7:04:00 AM | 8/23/2024 | 8/26/2024 4:53:05 PM | 2000-06-13 | 8/15/2019 11:13:17 AM | NPEPPIN | KATFRASE | Concentration | mg/L | Y | (fl. conc.) |  | Vanadium Total (fl. conc.) | 7440-62-2 | Metals - Total; | Metals - Total; EMS code: V--T | Chemical | Total | LAB
V--T | ug/g wet | Vanadium Total | X560 | Metals in Biota (wet), total, acid digestion, ICP-MS | Five grams of wet, homogenized biota is weighed in a centrifuge tube, and 2.5 mL of Nitric Acid and 2.5 mL of Hydrochloric acid are added. The sample is digested at room temperature for one hour and then digested at 90-95 degrees C in a water bath or heat block for one more hour. The sample is cooled and the volume is made to 50 mL with DI. The digestate is then diluted at ten fold and introduced to a collision/reaction cell inductively coupled argon plasma mass spectrometer (CRC ICPMS) by pneumatic nebulization. Each respective element is measured at a specific mass to charge ratio, m/z value expressed in amu (atomic mass units). - Maxxam Analytics (BC). Added to EMS Dec 10, 2008. | 2 | Y | Y | Y | 541 | 11/15/2008 4:00:00 AM | 10/30/2023 | 11/28/2023 2:21:33 PM | 2008-12-11 | 3/11/2015 4:17:10 PM | RDGIBSON | EMS | MassPerMass | µg/g | Y | (mass conc.) | (wet) | Vanadium Total (wet) (mass conc.) | 7440-62-2 | Metals - Total; | Metals - Total; EMS code: V--T | Chemical | Total | LAB
V--T | mg/L | Vanadium Total | F033 | Acid Dig,Centrifuge,ICPMS | For the determination of total recoverable elements in water or waste water, take 100 ml aliquot from a well mixed, acid preserved, sample containing not more than 0.25% (w/v) total solids and transfer to a 250 ml Griffin beaker. (If total solids are greater than 0.25% reduce the size of the aliquot by a proportionate amount). Add 1 ml of conc. nitric acid and 0.5 ml conc. hydrochloric acid. Heat on a hot plate until the volume has been reduced to 15-20ml, ensuring that the sample does not boil. Allow to cool and transfer to a 50 ml volumetric flask. Dilute to volume with ASTM type I water. Centrifuge the sample or allow to stand overnight to separate insoluble material. Prior analysis, pipette 20 ml into a 50 ml volumetric flask and dilute to volume with ASTM type I water. Analysis should be performed within seventy two hours of the completed preparation. Note that different laboratories may have very different field preservation, digestion procedures and detection limits for this method, so these should be documented. | 0.001 | Y | Y | Y | 36986 | 11/20/1997 9:32:00 AM | 8/22/2024 3:24:00 AM | 8/25/2024 9:47:36 PM | 1998-05-14 | 6/3/1998 6:51:49 PM | NPEPPIN |  | Concentration | mg/L | Y | (fl. conc.) |  | Vanadium Total (fl. conc.) | 7440-62-2 | Metals - Total; | Metals - Total; EMS code: V--T | Chemical | Total | LAB
V--T | ug/g | Vanadium Total | PB02 | PBM,HNO3/HCl dig,dry60C,100mesh, | Performance Based Method. Samples are dried at 60 C and sieved using a .149 mm (100 mesh) sieve. Solid samples are digested with the combination of nitric and hydrochloric acids with a 1:1 ratio of acids. For environmental purposes, strong acid leaches are used routinely since they have the advantage of speed, low cost, and superio reproducibility. All metals trapped within the refractory matrix are unlikely to have environmental consequences. Instrumental analysis of extracts produced using this digestion method can be performed by a variety of techniques. |  | Y | Y | Y | 52 | 5/24/2001 3:42:00 PM | 11/26/2001 10:23:00 AM | 5/26/2003 2:28:49 PM | 2001-07-18 |  | NPEPPIN |  | MassPerMass | mg/kg | Y | (mass conc.) |  | Vanadium Total (mass conc.) | 7440-62-2 | Metals - Total; | Metals - Total; EMS code: V--T | Chemical | Total | LAB
V--T | ug/g wet | Vanadium Total | MTIW | MET-WET-CCMS-MID-VA | Metals in Tissue by CRC ICPMS (WET) Methodology: This method is conducted following British Columbia Lab Manual method "Metals in Animal Tissue and Vegetation (Biota) - Prescriptive". Tissue samples are homogenized and sub-sampled prior to hotblock digestion with nitric and hydrochloric acids, in combination with addition of hydrogen peroxide. Instrumental analysis is by collision cell inductively coupled plasma - mass spectrometry (modified from EPA Method 6020A). Method Limitation: This method employs a strong acid/peroxide digestion, and is intended to provide a conservative estimate of bio-available metals. Near complete recoveries are achieved for most toxicologically important metals, but elements associated with recalcitrant minerals may be only partially recovered. | 0.02 | Y | Y | Y | 24 | 11/25/2018 | 12/6/2018 | 7/21/2021 2:55:08 PM | 2018-12-24 |  | EMS |  | MassPerMass | µg/g | Y | (mass conc.) | (wet) | Vanadium Total (wet) (mass conc.) | 7440-62-2 | Metals - Total; | Metals - Total; EMS code: V--T | Chemical | Total | LAB
V--T | mg/dm2/d | Vanadium Total | DUST | HNO3 dig.(microwave or hot block),ICPMS | HNO3 dig.(microwave or hot block);ICPMS |  | Y | Y | Y | 267 | 7/19/2016 1:11:00 PM | 7/31/2024 | 7/31/2024 11:23:13 AM | 2008-02-07 |  | RDGIBSON |  | MassFlux | mg/dm²/d | Y | (mass flux) |  | Vanadium Total (mass flux) | 7440-62-2 | Metals - Total; | Metals - Total; EMS code: V--T | Chemical | Total | LAB
V--T | ug/g | Vanadium Total | MTID | MET-DRY-CCMS-VA | Metals in Tissue by CRC ICPMS (DRY) Methodology: This method is conducted following British Columbia Lab Manual method "Metals in Animal Tissue and Vegetation (Biota) - Prescriptive". Tissue samples are homogenized and sub-sampled prior to hotblock digestion with nitric and hydrochloric acids, in combination with addition of hydrogen peroxide. Instrumental analysis is by collision cell inductively coupled plasma - mass spectrometry (modified from EPA Method 6020A). Method Limitation: This method employs a strong acid/peroxide digestion, and is intended to provide a conservative estimate of bio-available metals. Near complete recoveries are achieved for most toxicologically important metals, but elements associated with recalcitrant minerals may be only partially recovered. | 0.1 | Y | Y | Y | 122 | 8/13/2014 | 9/8/2022 | 10/26/2022 7:14:15 AM | 2015-03-09 | 3/11/2015 4:12:05 PM | RDGIBSON | EMS | MassPerMass | mg/kg | Y | (mass conc.) |  | Vanadium Total (mass conc.) | 7440-62-2 | Metals - Total; | Metals - Total; EMS code: V--T | Chemical | Total | LAB
V--T | mg/L | Vanadium Total | X285 | HNO3 Dig:ICP Analysis | HNO3 Dig:ICP Analysis | 0.01 | Y | Y | Y | 594 | 8/10/1993 | 10/5/1993 | 4/30/2004 11:32:39 AM | 1998-05-22 | 6/3/1998 6:51:49 PM | NPEPPIN |  | Concentration | mg/L | Y | (fl. conc.) |  | Vanadium Total (fl. conc.) | 7440-62-2 | Metals - Total; | Metals - Total; EMS code: V--T | Chemical | Total | LAB
V--T | mg/L | Vanadium Total | F121 | Acid,ICPMS | Acidified, no digestion, ICPMS | 0.0002 | Y | Y | Y | 8436 | 3/12/2010 7:57:00 PM | 8/15/2024 9:10:00 PM | 8/26/2024 8:06:03 AM | 2002-07-22 | 7/16/2012 3:32:35 PM | UMULLER | RDGIBSON | Concentration | mg/L | Y | (fl. conc.) |  | Vanadium Total (fl. conc.) | 7440-62-2 | Metals - Total; | Metals - Total; EMS code: V--T | Chemical | Total | LAB
V--T | ppm (S) | Vanadium Total | MS61 | ALS Minerals ME-MS61 | ALS Minerals ME-MS61. A prepared sample is digested with HNO3; HF; and HClO4 in a graphite heating block to near dryness. After cooling; the sample is boiled in diluted HCl. Bulk to volume with deionized water; mixed and analyzed by ICP-AES and ICP-MS for all elements. | 1 | Y | Y | Y | 0 |  |  |  | 2024-01-29 |  | EMS |  | DimensionlessRatio | ppm | N | (ratio) |  | Vanadium Total (ratio) | 7440-62-2 | Metals - Total; | Metals - Total; EMS code: V--T | Chemical | Total | LAB
V--T | ug/g | Vanadium Total | 0192 | Hom, Nitric/Perchlor,ICP | Sample already in a dry state so does not require drying before homogenization and nitric perchlorid digestion followed by ICP analysis. | 1 | Y | Y | Y | 708 |  |  | 10/24/1996 12:50:25 PM | 1996-04-12 | 6/3/1998 6:51:46 PM | EMS |  | MassPerMass | mg/kg | Y | (mass conc.) |  | Vanadium Total (mass conc.) | 7440-62-2 | Metals - Total; | Metals - Total; EMS code: V--T | Chemical | Total | LAB
V--T | ug/g wet | Vanadium Total | X552 | Metals in Biota (wet), total, acid digestion, ICP-SFMS | One gram of wet, homogenized biota is weighed in a microwave TFM vessel, and 8mL of Nitric Acid and 1mL of Hydrogen Peroxide are added. The sample is digested at 200 degrees C for 15 minutes in a closed vessel high pressure microwave oven. The sample is cooled, the volume adjusted to 100mL and introduced to an inductively coupled argon plasma-sector field mass spectrometer (ICP-SFMS) by pneumatic nebulization. Each respective element is measured at a specific mass to charge ratio, m/z value expressed in amu (atomic mass units), and at a specified instrument resolution. LR: Low Resolution ~300, MR: Mid Resolution ~4000. Ref: 02-2704, National Laboratory for Environmental Testing, CCIW, Environment Canada, Burlington. Requested April 2003. | 0.001 | Y | Y | Y | 37 | 8/9/2016 | 10/16/2023 | 11/28/2023 2:21:33 PM | 2004-01-19 | 4/19/2021 1:22:00 PM | UMULLER | EMS | MassPerMass | µg/g | Y | (mass conc.) | (wet) | Vanadium Total (wet) (mass conc.) | 7440-62-2 | Metals - Total; | Metals - Total; EMS code: V--T | Chemical | Total | LAB
V--T | mg/L | Vanadium Total | X561 | Total metals: FA, HNO3, HR-ICPMS (No Prep)Prep) | Ultra trace metals in water are analyzed by high resolution inductively coupled plasma mass spectrometry (HR-ICPMS), modified from US EPA 200.8, (Revision 5.5). The detection limits provided can only be met for undigested samples. This procedure is intended for colorless, non-turbid, acid-preserved water samples (i.e. pristine water samples), having turbidity < 1 NTU and no odor. Where turbidity exceeds 1 NTU, and/or the sample is colored and has an odor, results may be biased low compared to true Total Metals concentrations. ALS recommends that turbidity analysis be requested on samples submitted for this test to aid with interpretation of results. | 0.0002 | Y | Y | Y | 2756 | 6/8/2015 | 10/5/2023 | 1/31/2024 2:37:41 PM | 2015-10-22 |  | RDGIBSON |  | Concentration | mg/L | Y | (fl. conc.) |  | Vanadium Total (fl. conc.) | 7440-62-2 | Metals - Total; | Metals - Total; EMS code: V--T | Chemical | Total | LAB
V--T | mg/L | Vanadium Total | EQ01 | EQUIS (mg/L) | EQUIS (mg/L) | 0.001 | N | Y | Y | 1800 |  |  | 10/23/1996 12:08:27 AM | 1996-04-12 | 6/3/1998 6:51:40 PM | EMS |  | Concentration | mg/L | Y | (fl. conc.) |  | Vanadium Total (fl. conc.) | 7440-62-2 | Metals - Total; | Metals - Total; EMS code: V--T | Chemical | Total | LAB
V--T | ug/g | Vanadium Total | FS89 | Homo,HNO3 dig,ICP-AES | Homo;HNO3 dig;ICP-AES | 2 | Y | Y | Y | 4 | 1/5/2001 | 8/20/2001 12:00:00 PM | 10/23/2002 1:42:53 PM | 2001-03-08 |  | NPEPPIN |  | MassPerMass | mg/kg | Y | (mass conc.) |  | Vanadium Total (mass conc.) | 7440-62-2 | Metals - Total; | Metals - Total; EMS code: V--T | Chemical | Total | LAB
V--T | mg/L | Vanadium Total | MSEA | Metals in Seawater (undigested) by (CRC) ICP-MS | This method describes the multi-elemental determination of trace elements by Collision/Reaction Cell (CRC) ICP-MS. The method measures ions produced by a radio-frequency inductively coupled plasma. Analyte species originating in a liquid are pneumatically nebulized and the resulting aerosol transported by argon gas into the plasma torch. The ions produced are entrained in the plasma gas and extracted, by means of a differentially pumped vacuum interface, into a mass spectrometer. The ions flow through a Collision/Reaction Cell (CRC) where depending on the mode of analysis either the ions passes through the cell unhindered (no gas - classical ICP-MS mode), or collisions/reactions chemistry takes place between the ions and a neutral gas such as hydrogen or helium. Any polyatomic ions which give interference on the analytes are removed. The ions are then sorted according to their mass-to-charge ratios by a quadrupole mass spectrometer. The ions are then sorted according to their mass-to-charge ratios by a quadrupole mass spectrometer. The ions transmitted through the quadrupole are quantified by a channel electron multiplier and the ion information processed by a data handling system. Although most polyatomic interference is removed, any additional interference must be assessed and valid corrections applied or the data flagged to indicate problems. Interference correction must include compensation for background ions contributed by the plasma gas, reagents, and constituents of the sample matrix. Instrument drift as well as suppressions or enhancements of instrument response caused by the sample matrix must be corrected for by the use of internal standards. | 0.0002 | Y | Y | Y | 441 | 8/12/2009 12:08:00 AM | 4/29/2024 | 6/7/2024 10:03:16 AM | 2009-09-08 |  | RDGIBSON |  | Concentration | mg/L | Y | (fl. conc.) |  | Vanadium Total (fl. conc.) | 7440-62-2 | Metals - Total; | Metals - Total; EMS code: V--T | Chemical | Total | LAB
V--T | ug/g | Vanadium Total | FB04 | Freeze Dry,Hom,Acid Dig,ICP | Sample moisture is determined by weighing samples before and after freeze-drying. Freeze-dried samples are ground in a Braun blender before being weighed into Teflon PFA liners and digested with concentrated HNO3, concentrated HCl, and 50% H2O2 in a microwave oven; which results in a complete sample dissolution. Samples are collected and volumized with Type I water. Analysis is performed via Inductively Coupled Argon Plasma Atomic Emission Spectrometry (ICP-AES) (simultaneous multielement analysis), Graphite Furnace Atomic Absorption Spectrometry (GF-AAS) (low level Cd and Pb analysis), Cold Vapor Atomic Fluorescence Spectrometry (CV-AFS) (Mercury analysis) and/or Flame Atomic Absorption Spectrometry (F-AAS) (K analysis). | 0.8 | Y | Y | Y | 245 | 8/22/1996 6:11:00 PM | 12/1/2010 | 12/9/2010 3:55:27 PM | 1997-02-27 | 6/3/1998 6:51:46 PM | MPAK |  | MassPerMass | mg/kg | Y | (mass conc.) |  | Vanadium Total (mass conc.) | 7440-62-2 | Metals - Total; | Metals - Total; EMS code: V--T | Chemical | Total | LAB
V--T | mg/L | Vanadium Total | 1349 | Tot Metals in Seawater by CRC ICPMS (BC) | Tot Metals in Seawater by CRC ICPMS (BC) | 0.0005 | Y | Y | Y | 98 | 8/16/2019 | 12/5/2021 | 12/13/2021 12:34:17 PM | 2019-09-30 |  | KATFRASE |  | Concentration | mg/L | Y | (fl. conc.) |  | Vanadium Total (fl. conc.) | 7440-62-2 | Metals - Total; | Metals - Total; EMS code: V--T | Chemical | Total | LAB
V--T | ug/g wet | Vanadium Total | BPSY | MET-WET-BPSY-CCMS-VA | Metals in Animal Tissue and Vegetation (Biota) - Prescriptive. EPA 200.3/6020B This method is conducted following British Columbia Lab Manual method "Metals in Animal Tissue and Vegetation (Biota) - Prescriptive". Tissue samples are digested with nitric and hydrochloric acids; in combination with addition of hydrogen peroxide. Instrumental analysis is by collision cell inductively coupled plasma - mass spectrometry (modified from EPA Method 6020B). | 0.02 | Y | Y | Y | 2 | 11/7/2019 | 11/7/2019 | 6/10/2021 1:57:10 PM | 2021-06-04 |  | EMS |  | MassPerMass | µg/g | Y | (mass conc.) | (wet) | Vanadium Total (wet) (mass conc.) | 7440-62-2 | Metals - Total; | Metals - Total; EMS code: V--T | Chemical | Total | LAB
V--T | mg/L | Vanadium Total | F048 | Acid dig,conc'n,nebul.,ICP-OES | This method provides an automated procedure for the Inductively Coupled Argon Plasma Optical Emission Spectrophotometric (ICP-OES) determination and quantitation of trace amounts of metals in surface, ground and precipitation waters after a manual acid digestion and concentration procedure and ultrasonic nebulisation. Organically-bound metals and metal complexes are converted to metal ions by digestion with nitric acid [method details says HCl]. To increase the signal detected by the ICP, the samples are concentrated 10 times. The liquid sample is pumped onto the piezoelectric transducer, where the ultrasonic action efficiently converts it to a fine, dense aerosol. The nebulizer carrier gas flow then transports the wet aerosol through a heated U-tube, where the solvent is vaporized. Subsequently, most of the solvent vapour is removed from the aerosol stream by passing the stream through a chilled condenser. The result is a dry, analyte-laden aerosol which is introduced to the plasma. There, the high temperature desolvates the aerosol and dissociates the sample material into individual atoms and ions, which are excited to emit light at wave lengths characteristic of the elements present. Dissolved samples are field filtered through 0.45 micron cellulose acetate filter prior to acid preservation. Samples for extractable analysis are acidified with 2 mL concentrated HNO3 per litre and allowed to stand a minimum of 24 hours before analysis. Clients providing these samples are advised to preserve samples at the collection site. | 0.0001 | Y | Y | Y | 6159 | 7/4/1997 2:36:00 PM | 4/23/2003 | 9/11/2019 8:46:38 AM | 1997-10-20 | 6/3/1998 6:51:47 PM | NPEPPIN |  | Concentration | mg/L | Y | (fl. conc.) |  | Vanadium Total (fl. conc.) | 7440-62-2 | Metals - Total; | Metals - Total; EMS code: V--T | Chemical | Total | LAB
V--T | ug/g | Vanadium Total | 0196 | D&G: LTA/HCl: ICP | D&G: LTA/HCl: ICP | 1 | Y | Y | Y | 27 |  |  | 10/24/1996 2:08:13 PM | 1996-04-12 | 6/3/1998 6:51:40 PM | EMS |  | MassPerMass | mg/kg | Y | (mass conc.) |  | Vanadium Total (mass conc.) | 7440-62-2 | Metals - Total; | Metals - Total; EMS code: V--T | Chemical | Total | LAB
V--T | ug/g | Vanadium Total | FS82 | PBM,dry60C,10mesh,SALM-ICP-MS | PBM;dry60C;10mesh;SALM-ICP-MS | 2 | Y | Y | Y | 3624 | 5/21/2002 | 7/24/2024 | 7/29/2024 5:17:09 PM | 2002-11-12 |  | UMULLER |  | MassPerMass | mg/kg | Y | (mass conc.) |  | Vanadium Total (mass conc.) | 7440-62-2 | Metals - Total; | Metals - Total; EMS code: V--T | Chemical | Total | LAB
V--T | mg/L | Vanadium Total | X555 | Metals in water, total, "in-bottle" digestion, ICP-SFMS | The whole water sample collected in a HDPE/Teflon bottle, is de-capped in a Class 100 workstation, acidified to 2% (v/v) HNO3 and 1% (v/v) HCl and recapped. The sample is "in-bottle" digested in a convection oven, set at 60 degrees C for 16 hours. The sample is cooled and introduced to an inductively coupled argon plasma-sector field mass spectrometer (ICP-SFMS) by pneumatic nebulization. Each respective element is measured at a specific mass to charge ratio, m/z value expressed in amu (atomic mass units), and at resolution. LR: Low Resolution ~300, MR: Mid Resolution ~4000 and HR: High a specified instrument Resolution ~10000. Ref: Method number 02-2003, National Laboratory for Environmental Testing, CCIW, Environment Canada, Burlington. Requested April 2003. | 5e-06 | Y | Y | Y | 17331 | 5/20/2003 | 12/6/2023 2:08:00 AM | 2/14/2024 9:05:32 AM | 2004-01-19 |  | UMULLER |  | Concentration | mg/L | Y | (fl. conc.) |  | Vanadium Total (fl. conc.) | 7440-62-2 | Metals - Total; | Metals - Total; EMS code: V--T | Chemical | Total | LAB
V--T | mg/L | Vanadium Total | 0200 | 20%A.R.Dig:AA Anal. | 20%A.R.Dig:AA Anal. | 0.1 | Y | Y | Y | 70 |  |  | 6/25/2003 9:22:39 AM | 1996-04-12 | 6/3/1998 6:51:40 PM | EMS |  | Concentration | mg/L | Y | (fl. conc.) |  | Vanadium Total (fl. conc.) | 7440-62-2 | Metals - Total; | Metals - Total; EMS code: V--T | Chemical | Total | LAB
V--T | mg/L | Vanadium Total | METN | MET-T-NP-CCMS-VA | Ultra trace metals in water are analyzed by CRC ICPMS, based on US EPA Method 6020B (July 2014). The detection limits provided can only be met for undigested samples. This procedure is intended for colorless, non-turbid, acid-preserved water samples (i.e. pristine water samples), having turbidity < 1 NTU and no odor. Where turbidity exceeds 1 NTU, and/or the sample is colored and has an odor, results may be biased low compared to true Total Metals concentrations. ALS recommends that turbidity analysis be requested on samples submitted for this test to aid with interpretation of results. Where turbidity is <1NTU, undigested metals are equivalent to total metals concentrations. | 0.0002 | Y | Y | Y | 3648 | 4/3/2012 | 8/19/2024 | 8/26/2024 1:03:06 PM | 2019-08-20 |  | KATFRASE |  | Concentration | mg/L | Y | (fl. conc.) |  | Vanadium Total (fl. conc.) | 7440-62-2 | Metals - Total; | Metals - Total; EMS code: V--T | Chemical | Total | LAB
V--T | ug/m3 | Vanadium Total | 5312 | Tef:N/P Dig:ICP Analysis | A nitric/perchloric acid digestion if used to bring the metals into solution prior to ICP analysis. For Hi-Vol samples preliminary results are reported as mg/L and 24 hour loading results as ug/m3. | 0.003 | Y | Y | Y | 1191 | 11/7/1997 | 5/15/2000 | 11/18/2000 1:29:40 AM | 1996-10-13 | 6/3/1998 6:51:46 PM | EMS |  | AirConcentration | µg/m³ | Y | (air conc.) |  | Vanadium Total (air conc.) | 7440-62-2 | Metals - Total; | Metals - Total; EMS code: V--T | Chemical | Total | LAB
V--T | ppm (S) | Vanadium Total | 6042 | Aqua Regia by ICPMS | These were sediment/soils digested with Aqua-Regia (Nitric Acid & Hydrochloric Acids) and then analysis by ICPMS. Results would be reported in ppm s. | 0.1 | Y | Y | Y | 71 | 5/14/2022 | 7/27/2024 | 7/29/2024 9:13:08 AM | 2022-06-28 |  | EMS |  | DimensionlessRatio | ppm | N | (ratio) |  | Vanadium Total (ratio) | 7440-62-2 | Metals - Total; | Metals - Total; EMS code: V--T | Chemical | Total | LAB
V--T | ug/g | Vanadium Total | EQ02 | EQUIS (ug/g) | EQUIS (ug/g) | 1 | N | Y | Y | 760 |  |  | 10/23/1996 12:08:28 AM | 1996-04-12 | 6/3/1998 6:51:40 PM | EMS |  | MassPerMass | mg/kg | Y | (mass conc.) |  | Vanadium Total (mass conc.) | 7440-62-2 | Metals - Total; | Metals - Total; EMS code: V--T | Chemical | Total | LAB
V--T | mg/L | Vanadium Total | X349 | FA:HNO3:ICP | Samples are field-acidified with HNO3; HNO3 digestion; analysis by ICP/AES. | 0.01 | Y | Y | Y | 30 |  |  | 10/30/2014 8:31:19 AM | 1998-05-22 | 6/3/1998 6:51:49 PM | NPEPPIN |  | Concentration | mg/L | Y | (fl. conc.) |  | Vanadium Total (fl. conc.) | 7440-62-2 | Metals - Total; | Metals - Total; EMS code: V--T | Chemical | Total | LAB
V--T | mg/L | Vanadium Total | 0270 | HNO3 Dig:Ext(Cupferon)AA | HNO3 Dig:Ext(Cupferon)AA | 0.001 | Y | Y | Y | 120 |  |  | 10/24/1996 11:57:35 AM | 1996-04-12 | 6/3/1998 6:51:40 PM | EMS |  | Concentration | mg/L | Y | (fl. conc.) |  | Vanadium Total (fl. conc.) | 7440-62-2 | Metals - Total; | Metals - Total; EMS code: V--T | Chemical | Total | LAB
V--T | ug/g | Vanadium Total | 0190 | D&G:N/P Dig:ICP Anal. | Sample is dried and grind-mixed before nitric perchloric digestion followed by ICP analysis. | 1 | Y | Y | Y | 810 |  |  | 10/24/1996 3:06:16 PM | 1996-04-12 | 6/3/1998 6:51:40 PM | EMS |  | MassPerMass | mg/kg | Y | (mass conc.) |  | Vanadium Total (mass conc.) | 7440-62-2 | Metals - Total; | Metals - Total; EMS code: V--T | Chemical | Total | LAB
V--T | mg/L | Vanadium Total | 0040 | HNO3 Dig: ICP Analysis | HNO3 Dig: ICP Analysis | 0.01 | Y | Y | Y | 15123 | 6/25/2010 3:15:00 PM | 6/25/2010 3:15:00 PM | 7/22/2010 3:51:36 PM | 1996-04-12 | 6/3/1998 6:51:40 PM | EMS |  | Concentration | mg/L | Y | (fl. conc.) |  | Vanadium Total (fl. conc.) | 7440-62-2 | Metals - Total; | Metals - Total; EMS code: V--T | Chemical | Total | LAB
V--T | ug/g | Vanadium Total | FS01 | D&G,Nitric/HCl:ICP | Samples are dried at 60 C, then either sieved to -100mesh (150 um) and rolled to homogenize, or finely ground with a mortar and pestle or a shatter box laboratory mill and rolled (depending on the submitter's requirements). The samples are then weighed into Teflon digestion liners and digested with HNO3 and HCl in a microwave oven. Samples are cooled, volumized, and allowed to settle overnight before being decanted. Analysis is then performed via Inductively Coupled Argon Plasma Atomic Emission Spectrometry (ICP-AES) (Simultaneous multi element analysis). In ICP Spectrometry, the liquid sample is pumped to the nebulizer, where it is convertd into an aerosol. the nebulizer carrier gas flow than transports the wet aerosol past the impact bead, where the large droplets condense out and flow to the waste container. The remaining fine aerosol is carried into the plasma. There, the high temperature desolvates the aerosol and dissociates the sample material into individual atoms and ions, which are excited to emit light at wavelengths characteristic of the elements present. A diffraction grating in the spectrometer diffracts the various wavelengths of light through exit slits to the photo-multiplier tubes, which convert the light energy to electronic signals. | 2 | Y | Y | Y | 2642 | 8/30/1996 1:52:00 PM | 10/10/2007 | 5/24/2008 1:15:03 AM | 1996-11-20 | 6/3/1998 6:51:46 PM | MPAK |  | MassPerMass | mg/kg | Y | (mass conc.) |  | Vanadium Total (mass conc.) | 7440-62-2 | Metals - Total; | Metals - Total; EMS code: V--T | Chemical | Total | LAB
V--T | ppm (S) | Vanadium Total | MS41 | ALS Minerals ME-MS41 | ALS Minerals ME-MS41. A prepared sample is digested with aqua regia in a graphite heating block. After cooling; the resulting solution is diluted to with deionized water; mixed and analyzed by ICP-AES. Samples are then analyzed by ICP-MS for the remaining suite of elements. | 1 | Y | Y | Y | 0 |  |  |  | 2024-01-29 |  | EMS |  | DimensionlessRatio | ppm | N | (ratio) |  | Vanadium Total (ratio) | 7440-62-2 | Metals - Total; | Metals - Total; EMS code: V--T | Chemical | Total | LAB
V--T | ug/g | Vanadium Total | ICPM | Metals in Soil by CRC ICPMS | This analysis is carried out using procedures from CSR Analytical Method: "Strong Acid Leachable Metals (SALM) in Soil", BC Ministry of Environment, 26 June 2009, and procedures adapted from EPA Method 200.2. The sample is manually homogenized, dried at 60 degrees Celsius, sieved through a 63 um (230 mesh) sieve, and a representative subsample of the dry material is weighed. The sample is then digested at 95 degrees Celsius for 2 hours by block digester using concentrated nitric and hydrochloric acids. Instrumental analysis is by collision cell inductively coupled plasma - mass spectrometry (modifed from EPA Method 6020A). Method Limitation: This method is not a total digestion technique. It is a very strong acid digestion that is intended to dissolve those metals that may be environmentally available. By design, elements bound in silicate structures are not normally dissolved by this procedure as they are not usually mobile in the environment. | 0.2 | Y | Y | Y | 103 | 9/19/2017 | 9/24/2023 | 1/12/2024 11:16:52 AM | 2018-01-09 |  | JKROGH |  | MassPerMass | mg/kg | Y | (mass conc.) |  | Vanadium Total (mass conc.) | 7440-62-2 | Metals - Total; | Metals - Total; EMS code: V--T | Chemical | Total | LAB
V--T | ug/g | Vanadium Total | X551 | Metals in Biota (dry), total, acid digestion, ICP-SFMS | 0.1 gram of dried, homogenized biota is weighed in a microwave TFM vessel, and 8mL of Nitric Acid and 1mL of Hydrogen Peroxide are added. The sample is digested at 200 degrees C for 15 minutes in a closed vessel high pressure microwave oven. The sample is cooled, the volume adjusted to 100mL and introduced to an iductively coupled argon plasma-sector field mass spectrometer (ICP-SFMS) by pneumatic nebulization. Each respective element is measured at a specific mass to charge ratio, m/z value expressed in amu (atomic mass units), and at a specified instrument resolution. LR: Low Resolution ~300, MR: Mid Resolution ~4000 and HR: Hich Resolution ~10000. Ref: Method number 02-2705, National Laboratory for Environmental Testing, CCIW, Environment Canada, Burlington. Requested April 2003. | 0.01 | Y | Y | Y | 49 | 8/11/2015 | 10/30/2023 | 11/28/2023 2:21:33 PM | 2004-01-19 |  | UMULLER |  | MassPerMass | mg/kg | Y | (mass conc.) |  | Vanadium Total (mass conc.) | 7440-62-2 | Metals - Total; | Metals - Total; EMS code: V--T | Chemical | Total | LAB
V001 | mg/L | Vinclozolin | ZN01 | DCM+Ac.Extr:GC/MS | DCM+Ac.Extr:GC/MS | 0.0005 | Y | Y | Y | 28 | 7/28/2005 7:03:00 AM | 3/29/2012 10:15:00 PM | 5/3/2012 9:27:52 AM | 1996-04-12 | 11/16/1999 4:06:28 PM | EMS | NPEPPIN | Concentration | mg/L | Y | (fl. conc.) |  | Vinclozolin (fl. conc.) | 50471-44-8 | Pesticides; | Pesticides; EMS code: V001 | Chemical |  | LAB
VALA | mg/L | Valeric Acid | VFAD | Volatile Fatty Acids by GC/FID | Volatile Fatty Acids by GC/FID; Standard Methods for the Examination of Water and Wastewater; 21st Edition, Method 5560 D; pages 5-59 to 5-61. 5560 D. Gas Chromatographic Method Principle: This gas chromatographic procedure may be usedto determine the individual concentrations of the following fattyacids: acetic; propionic; butyric; isobutyric; valeric; andisovaleric acids in digester sludge samples. Caproic and heptanoicacids also may be determined; but with less certaintybecause of their reduced solubility in water. Volatile fatty acidsare generally soluble in digester sludge and exist in either theprotonated or unprotonated form; depending on the pH of thesludge sample. Acetic and propionic acids usually are in higherconcentrations than the other fatty acids in digester sludge. Allacids must be converted to the protonated (volatile) form beforeinjection into the gas chromatograph (GC) so vaporization mayoccur. The fatty acids must also be separated from the solidsmaterial in the sample before introduction into the GC to minimizedegradation of the GC column. Digester sludge samples areprepared by acidification; centrifugation; and filtration. The sampleis analyzed by direct injection into a gas chromatographequipped with a flame ionization detector after both centrifugationand filtration.b. Interference: The blank amount may be high relative tosample analyte. High blanks may be caused by a buildup ofcontaminants in the injector and guard column or by samplecarry-over. The analyst must be aware of the blank levels andcorrect unacceptably high blank levels that are above the reportinglimits.c. Minimum detectable concentration: The method detectionlevel has been determined by the USEPA method1 in reagentwater to be about 3 mg/L for acetic acid and about 1 mg/L for allother target compounds. | 2 | Y | Y | Y | 5 | 3/14/2017 10:30:00 AM | 11/19/2017 5:10:00 PM | 6/20/2018 9:02:11 AM | 2018-06-19 |  | RDGIBSON |  | Concentration | mg/L | Y | (fl. conc.) |  | Valeric Acid (fl. conc.) | 109-52-4 | Organics; | Organics; EMS code: VALA | Chemical |  | LAB
VDST | mg/dm2/d | Volatile Dustfall | X192 | Gravimetric (Sol+Insol) | Gravimetric (Sol+Insol) |  | Y | Y | Y | 1607 | 9/15/2015 4:14:00 PM | 8/19/2024 | 8/19/2024 4:02:08 PM | 2008-06-04 |  | RDGIBSON |  | MassFlux | mg/dm²/d | Y | (mass flux) |  | Volatile Dustfall (mass flux) |  | Particulates; | Particulates; EMS code: VDST | Chemical |  | LAB
VDST | mg/dm2/d | Volatile Dustfall | DUST | HNO3 dig.(microwave or hot block),ICPMS | HNO3 dig.(microwave or hot block);ICPMS |  | Y | Y | Y | 13 | 1/31/2020 | 1/31/2020 | 1/16/2023 3:51:10 PM | 2008-02-07 |  | RDGIBSON |  | MassFlux | mg/dm²/d | Y | (mass flux) |  | Volatile Dustfall (mass flux) |  | Particulates; | Particulates; EMS code: VDST | Chemical |  | LAB
VFR- | mg/L | Residue: VolatileFilt (TDS; Volatile) | CS01 | System Calculation | System Calculation (mg/L) |  | N | Y | N | 2161 |  |  | 3/19/2002 1:14:40 PM | 1996-04-12 | 6/3/1998 6:51:42 PM | EMS |  | Concentration | mg/L | Y | (fl. conc.) |  | Residue: VolatileFilt (TDS; Volatile) (fl. conc.) |  | Physical Properties; | Physical Properties; EMS code: VFR- | Physical |  | LAB
VFR- | mg/L | Residue: VolatileFilt (TDS; Volatile) | XCAL | Calculation (mg/L) | Calculation (mg/L) | 3 | Y | Y | Y | 49 | 8/11/2014 | 6/14/2024 | 6/27/2024 3:13:06 PM | 2015-03-11 |  | RDGIBSON |  | Concentration | mg/L | Y | (fl. conc.) |  | Residue: VolatileFilt (TDS; Volatile) (fl. conc.) |  | Physical Properties; | Physical Properties; EMS code: VFR- | Physical |  | LAB
VHC- | ug/g | Hydrocarbons; Volatile | F084 | MeOH extrac,GC/FID,nC6-nC10 | Wet solids samples are extracted with methanol using a mechanical shaker. Extracts are directly analyzed by a capillary column gas chromatography with flame ionization detection. Published as a Performance Based Method circa Dec. 2000. |  | Y | Y | Y | 185 | 11/16/2004 | 10/28/2022 5:04:00 PM | 1/7/2023 3:12:10 PM | 2000-10-25 |  | NPEPPIN |  | MassPerMass | mg/kg | Y | (mass conc.) |  | Hydrocarbons; Volatile (mass conc.) |  | Petroleum Hydrocarbons; | Petroleum Hydrocarbons; EMS code: VHC- | Chemical |  | LAB
VHC- | mg/L | Hydrocarbons; Volatile | F083 | He purge+trap,GC/FID,nC6-nC10 | Volatile organic compounds (nC6-nC10) are purged from water samples with helium, adsorbed onto a sorbent trap, and then thermally desorbed to a capillary column gas chromatograph equipped with a flame ionization detector. |  | Y | Y | Y | 464 | 9/6/2000 9:56:00 AM | 2/1/2024 11:07:00 AM | 2/15/2024 8:42:43 AM | 2000-10-25 | 11/22/2000 1:38:40 PM | NPEPPIN | NPEPPIN | Concentration | mg/L | Y | (fl. conc.) |  | Hydrocarbons; Volatile (fl. conc.) |  | Petroleum Hydrocarbons; | Petroleum Hydrocarbons; EMS code: VHC- | Chemical |  | LAB
VHC- | ug/g | Hydrocarbons; Volatile | MS03 | Static Headspace GC/MS (Soil) | (Soil) A static headspace procedure is used as a means of driving the analytes out of the matrix into a headspace. Once equilibrium is achieved, a fixed volume of headspeace is taken for subsequent determination by Gas Chromatography - Mass Spectroscopy (GC/MS) using Selected Ion Monitoring. | 10 | Y | Y | Y | 15 | 1/14/2022 | 2/4/2022 | 1/11/2024 1:07:08 PM | 2010-01-13 | 1/13/2010 10:23:23 AM | RDGIBSON | RDGIBSON | MassPerMass | mg/kg | Y | (mass conc.) |  | Hydrocarbons; Volatile (mass conc.) |  | Petroleum Hydrocarbons; | Petroleum Hydrocarbons; EMS code: VHC- | Chemical |  | LAB
VHC- | mg/L | Hydrocarbons; Volatile | MS04 | Static Headspace GC/MS (Water) | (Water) A static headspace procedure is used as a means of driving the analytes out of the matrix into a headspace. Once equilibrium is achieved, a fixed volume of headspeace is taken for subsequent determination by Gas Chromatography - Mass Spectroscopy (GC/MS) using Selected Ion Monitoring. | 0.01 | Y | Y | Y | 527 | 8/13/2009 | 8/12/2024 | 8/15/2024 11:32:04 AM | 2010-01-13 | 1/13/2010 10:24:10 AM | RDGIBSON | RDGIBSON | Concentration | mg/L | Y | (fl. conc.) |  | Hydrocarbons; Volatile (fl. conc.) |  | Petroleum Hydrocarbons; | Petroleum Hydrocarbons; EMS code: VHC- | Chemical |  | LAB
VHC- | mg/L | Hydrocarbons; Volatile | MS01 | Purge&Trap GC/MS (VOC-W) | Purge&Trap GC/MS (VOC-W) |  | Y | Y | Y | 29 | 5/3/2002 | 6/22/2007 | 7/9/2012 2:09:49 PM | 2002-08-12 |  | NPEPPIN |  | Concentration | mg/L | Y | (fl. conc.) |  | Hydrocarbons; Volatile (fl. conc.) |  | Petroleum Hydrocarbons; | Petroleum Hydrocarbons; EMS code: VHC- | Chemical |  | LAB
VOFR | m3/min A | Volume Flow Rate-Stk | 5094 | Volumetric Flow Calc-Stack | Volumetric Flow Calc-Stack | 0.001 | Y | Y | Y | 59 |  |  | 12/1/2009 6:08:55 AM | 1996-04-12 | 6/3/1998 6:51:46 PM | EMS |  | Flow | m³/min | Y | (flow) |  | Volume Flow Rate-Stk (flow) |  |  | EMS code: VOFR | Chemical |  | LAB
VOFR | m3/min A | Volume Flow Rate-Stk | X191 | S-Type Pitot Tube | S-Type Pitot Tube | 0 | Y | Y | Y | 44227 | 2/18/2014 8:00:00 AM | 6/29/2024 8:00:00 AM | 8/22/2024 1:49:14 PM | 1996-04-12 | 6/3/1998 6:51:40 PM | EMS |  | Flow | m³/min | Y | (flow) |  | Volume Flow Rate-Stk (flow) |  |  | EMS code: VOFR | Chemical |  | LAB
VOFR | m3/min A | Volume Flow Rate-Stk | CAL8 | Calculation (m3/min A) | Calculation (m3/min A) | 0 | Y | Y | Y | 259 |  |  | 8/1/2024 3:52:42 PM | 1996-04-12 | 6/3/1998 6:51:46 PM | EMS |  | Flow | m³/min | Y | (flow) |  | Volume Flow Rate-Stk (flow) |  |  | EMS code: VOFR | Chemical |  | LAB
VPH- | mg/L | Hydrocarbons; Volatile Petroleum | F099 | H20,Calc'n,VH-MAHs,nC6-nC10 | Volatile Petroleum Hydrocarbons (nC6 to nC10) are calculated using the results from selected methods. The calculation procedure for VPH requires that both Volatile Hydrocarbons (VH) and Monocyclic Aromatic Hydrocarbons (MAHs, including BTEX and styrene) be analyzed using methodologies which have been approved by the Director. Selected MAHs are subtracted from VH results to produce VPH values. The Procedure section lists the MAHs to be excluded from both waters and soils. Performance Based Method published circa Dec. 2000. | 0.1 | Y | Y | Y | 530 | 9/6/2000 9:56:00 AM | 2/1/2024 11:07:00 AM | 2/15/2024 8:42:43 AM | 2000-07-14 |  | NPEPPIN |  | Concentration | mg/L | Y | (fl. conc.) |  | Hydrocarbons; Volatile Petroleum (fl. conc.) |  | Petroleum Hydrocarbons; | Petroleum Hydrocarbons; EMS code: VPH- | Chemical |  | LAB
VPH- | ug/g | Hydrocarbons; Volatile Petroleum | MS03 | Static Headspace GC/MS (Soil) | (Soil) A static headspace procedure is used as a means of driving the analytes out of the matrix into a headspace. Once equilibrium is achieved, a fixed volume of headspeace is taken for subsequent determination by Gas Chromatography - Mass Spectroscopy (GC/MS) using Selected Ion Monitoring. | 10 | Y | Y | Y | 62 | 3/5/2013 | 2/7/2022 | 1/11/2024 1:07:08 PM | 2010-01-13 | 1/13/2010 10:28:59 AM | RDGIBSON | RDGIBSON | MassPerMass | mg/kg | Y | (mass conc.) |  | Hydrocarbons; Volatile Petroleum (mass conc.) |  | Petroleum Hydrocarbons; | Petroleum Hydrocarbons; EMS code: VPH- | Chemical |  | LAB
VPH- | mg/L | Hydrocarbons; Volatile Petroleum | MS04 | Static Headspace GC/MS (Water) | (Water) A static headspace procedure is used as a means of driving the analytes out of the matrix into a headspace. Once equilibrium is achieved, a fixed volume of headspeace is taken for subsequent determination by Gas Chromatography - Mass Spectroscopy (GC/MS) using Selected Ion Monitoring. | 0.01 | Y | Y | Y | 317 | 5/8/2008 | 7/8/2024 | 7/22/2024 3:17:33 PM | 2010-01-13 | 1/13/2010 10:29:29 AM | RDGIBSON | RDGIBSON | Concentration | mg/L | Y | (fl. conc.) |  | Hydrocarbons; Volatile Petroleum (fl. conc.) |  | Petroleum Hydrocarbons; | Petroleum Hydrocarbons; EMS code: VPH- | Chemical |  | LAB
VPH- | mg/L | Hydrocarbons; Volatile Petroleum | X530 | Purge&Trap, gc/FID(nC5-nC11) | Volatile petroleum hydrocarbons (nC5-nC11) are extracted from samples using methanol. The methanol is analyzed using a purge and trap technique. Using purge and trap analysis, the volatile compounds are purged with an inert gas, adsorbed on a sorbent trap, and then desorbed into a capillary gas chromatograph with flame ionization detector. VPH results are corrected for BETX concentrations and percent moisture. Superceded Dec. 2000. | 0.1 | Y | Y | Y | 36 | 1/24/2001 | 4/11/2001 | 7/27/2001 11:21:29 AM | 1999-07-29 |  | NPEPPIN |  | Concentration | mg/L | Y | (fl. conc.) |  | Hydrocarbons; Volatile Petroleum (fl. conc.) |  | Petroleum Hydrocarbons; | Petroleum Hydrocarbons; EMS code: VPH- | Chemical |  | LAB
VPH- | mg/L | Hydrocarbons; Volatile Petroleum | XCAL | Calculation (mg/L) | Calculation (mg/L) | 0.1 | Y | Y | Y | 471 | 8/6/2017 | 8/13/2024 | 8/15/2024 11:32:04 AM | 2018-08-15 |  | RDGIBSON |  | Concentration | mg/L | Y | (fl. conc.) |  | Hydrocarbons; Volatile Petroleum (fl. conc.) |  | Petroleum Hydrocarbons; | Petroleum Hydrocarbons; EMS code: VPH- | Chemical |  | LAB
VPH- | ug/g | Hydrocarbons; Volatile Petroleum | F100 | Solids,Cal'n,VH-MAHs,nC6-nC10 | Volatile Petroleum Hydrocarbons (nC6 to nC10) are calculated using the results from selected methods. The calculation procedure for VPH requires that both Volatile Hydrocarbons (VH) and Monocyclic Aromatic Hydrocarbons (MAHs, including BTEX and styrene) be analyzed using methodologies which have been approved by the Director. Selected MAHs are subtracted from VH results to produce VPH values. The Procedure section lists the MAHs to be excluded from both waters and soils. Performance Based Method published circa Dec. 2000. | 100 | Y | Y | Y | 442 | 1/3/2001 9:09:00 AM | 7/24/2018 | 7/19/2024 8:20:35 AM | 2000-07-14 |  | NPEPPIN |  | MassPerMass | mg/kg | Y | (mass conc.) |  | Hydrocarbons; Volatile Petroleum (mass conc.) |  | Petroleum Hydrocarbons; | Petroleum Hydrocarbons; EMS code: VPH- | Chemical |  | LAB
VPH- | ug/g | Hydrocarbons; Volatile Petroleum | X529 | MeOH extr'n,purge&trap,gc/FID(nC5-nC11) | Volatile petroleum hydrocarbons (nC5-nC11) are extracted from samples using methanol. The methanol is analyzed using a purge and trap technique. Using purge and trap analysis, the volatile compounds are purged with an inert gas, adsorbed on a sorbent trap, and then desorbed into a capillary gas chromatograph with flame ionization detector. VPH results are corrected for BETX concentrations and percent moisture. Superceded Dec. 2000. | 100 | Y | Y | Y | 1 | 10/28/2022 5:04:00 PM | 10/28/2022 5:04:00 PM | 1/7/2023 3:12:10 PM | 1999-07-29 |  | NPEPPIN |  | MassPerMass | mg/kg | Y | (mass conc.) |  | Hydrocarbons; Volatile Petroleum (mass conc.) |  | Petroleum Hydrocarbons; | Petroleum Hydrocarbons; EMS code: VPH- | Chemical |  | LAB
W--D | mg/L | Tungsten Dissolved | MWID | Metals in Water by ICPMS (digested) | Metals in Water by ICPMS (digested) Based on EPA 200.8 | 0.001 | Y | Y | Y | 123 |  |  | 1/27/2023 10:37:54 AM | 2009-12-24 |  | RDGIBSON |  | Concentration | mg/L | Y | (fl. conc.) |  | Tungsten Dissolved (fl. conc.) | 7440-33-7 | Metals - Dissolved; | Metals - Dissolved; EMS code: W--D | Chemical | Dissolved | LAB
W--D | mg/L | Tungsten Dissolved | F049 | Filt,HNO3 dig,conc'n,nebul'ICP-OES | This method provides an automated procedure for the Inductively Coupled Argon Plasma Optical Emission Spectrophotometric (ICP-OES) determination and quantitation of trace amounts of metals in surface, ground and precipitation waters after a manual acid digestion and concentration procedure and ultrasonic nebulisation. Organically-bound metals and metal complexes are converted to metal ions by digestion with nitric acid. To increase the signal detected by the ICP, the samples are concentrated 10 times. The liquid sample is pumped onto the piezoelectric transducer, where the ultrasonic action efficiently converts it to a fine, dense aerosol. The nebulizer carrier gas flow then transports the wet aerosol through a heated U-tube, where the solvent is vaporized. Subsequently, most of the solvent vapour is removed from the aerosol stream by passing the stream through a chilled condenser. The result is a dry, analyte-laden aerosol which is introduced to the plasma. There, the high temperature desolvates the aerosol and dissociates the sample material into individual atoms and ions, which are excited to emit light at wave lengths characteristic of the elements present. Dissolved samples are field filtered through 0.45 micron cellulose acetate filter prior to acid preservation. Samples for extractable analysis are acidified with 2 mL concentrated HNO3 per litre and allowed to stand a minimum of 24 hours before analysis. Clients providing these samples are advised to preserve samples at the collection site. | 0.1 | Y | Y | Y | 72 | 5/9/2003 9:40:00 AM | 9/16/2003 3:18:00 PM | 6/7/2024 5:30:17 PM | 1998-11-17 |  | NPEPPIN |  | Concentration | mg/L | Y | (fl. conc.) |  | Tungsten Dissolved (fl. conc.) | 7440-33-7 | Metals - Dissolved; | Metals - Dissolved; EMS code: W--D | Chemical | Dissolved | LAB
W--D | mg/L | Tungsten Dissolved | X556 | Metals in water, dissolved, "in-bottle" digestion, ICP-SFMS | The whole water sample is filtered through a 0.45 micron cellulose acetate filter paper and the filtrate is stored in a HDPE/Teflon bottle. The container is de-capped in a Class 100 workstation, acidified to 2% (v/v) HNO3 and 1% (v/v) HCl and recapped. The sample is "in-bottle" digested in a convection oven, set at 60 degrees C for 16 hours. The sample is cooled and introduced to an inductively coupled argon plasma-sector field mass spectrometer (ICP-SFMS) by pneumatic nebulization. Each respective element is measured at a specific mass to charge ratio, m/z value expressed in amu (atomic mass units), and at a specified intrument resolution. LR: Low Resolution ~300, MR: Mid Resolution ~4000 and HR: High Resolution ~10000. Ref: Method number 02-2003, National Laboratory for Environmental Testing, CCIW, Environment Canada, Burlington, Requested April 2003. | 5e-06 | Y | Y | Y | 4159 | 6/27/2004 | 12/6/2023 2:04:00 AM | 2/14/2024 9:05:32 AM | 2008-06-12 |  | RDGIBSON |  | Concentration | mg/L | Y | (fl. conc.) |  | Tungsten Dissolved (fl. conc.) | 7440-33-7 | Metals - Dissolved; | Metals - Dissolved; EMS code: W--D | Chemical | Dissolved | LAB
W--D | mg/L | Tungsten Dissolved | MSEA | Metals in Seawater (undigested) by (CRC) ICP-MS | This method describes the multi-elemental determination of trace elements by Collision/Reaction Cell (CRC) ICP-MS. The method measures ions produced by a radio-frequency inductively coupled plasma. Analyte species originating in a liquid are pneumatically nebulized and the resulting aerosol transported by argon gas into the plasma torch. The ions produced are entrained in the plasma gas and extracted, by means of a differentially pumped vacuum interface, into a mass spectrometer. The ions flow through a Collision/Reaction Cell (CRC) where depending on the mode of analysis either the ions passes through the cell unhindered (no gas - classical ICP-MS mode), or collisions/reactions chemistry takes place between the ions and a neutral gas such as hydrogen or helium. Any polyatomic ions which give interference on the analytes are removed. The ions are then sorted according to their mass-to-charge ratios by a quadrupole mass spectrometer. The ions are then sorted according to their mass-to-charge ratios by a quadrupole mass spectrometer. The ions transmitted through the quadrupole are quantified by a channel electron multiplier and the ion information processed by a data handling system. Although most polyatomic interference is removed, any additional interference must be assessed and valid corrections applied or the data flagged to indicate problems. Interference correction must include compensation for background ions contributed by the plasma gas, reagents, and constituents of the sample matrix. Instrument drift as well as suppressions or enhancements of instrument response caused by the sample matrix must be corrected for by the use of internal standards. | 1e-05 | Y | Y | Y | 4 | 4/12/2023 | 4/29/2024 | 5/16/2024 2:00:14 PM | 2009-09-08 |  | RDGIBSON |  | Concentration | mg/L | Y | (fl. conc.) |  | Tungsten Dissolved (fl. conc.) | 7440-33-7 | Metals - Dissolved; | Metals - Dissolved; EMS code: W--D | Chemical | Dissolved | LAB
W--D | mg/L | Tungsten Dissolved | X562 | Dissolved metals: FF, FA, HNO3, HR-ICPMS | Ultra trace metals in water are analyzed by high resolution inductively coupled plasma mass spectrometry (HR-ICPMS), based on US EPA Method 200.8, (Oct. 1999). This procedure is intended for pristine field-filtered acid-preserved water samples. The detection limits (LOR) for this test are based on lab instrumental analysis only, not including filtration. ALS-supplied field filtration equipment does not support these LOR. Therefore, because of the high probability of false positives due to filtration, it is strongly recommended that a filtration blank be analyzed to aid in data interpretation. | 1e-05 | Y | Y | Y | 22 | 10/20/2015 | 10/19/2023 | 11/9/2023 7:41:27 AM | 2016-03-24 |  | RDGIBSON |  | Concentration | mg/L | Y | (fl. conc.) |  | Tungsten Dissolved (fl. conc.) | 7440-33-7 | Metals - Dissolved; | Metals - Dissolved; EMS code: W--D | Chemical | Dissolved | LAB
W--D | mg/L | Tungsten Dissolved | METC | MET-D-NP-CCMS-VA | Ultra trace metals in water are analyzed by CRC ICPMS, based on US EPA Method 6020B (July 2014). This procedure is intended for pristine field-filtered acid-preserved water samples. The detection limits (LOR) for this test are based on lab instrumental analysis only, not including filtration. ALS-supplied field filtration equipment does not support these LOR. Therefore, because of the high probability of false positives due to filtration, it is strongly recommended that a filtration blank be analysed to aid in data interpretation. | 1e-05 | Y | Y | Y | 394 | 3/30/2020 | 4/9/2024 | 4/29/2024 4:53:06 PM | 2022-10-07 |  | EMS |  | Concentration | mg/L | Y | (fl. conc.) |  | Tungsten Dissolved (fl. conc.) | 7440-33-7 | Metals - Dissolved; | Metals - Dissolved; EMS code: W--D | Chemical | Dissolved | LAB
W--D | mg/L | Tungsten Dissolved | F032 | Filt.,Acid,ICPMS | There is no acid digestion required for this analysis. A water sample is collected and preserved in 0.2 percent nitric acid. To analyse the extractable metals the sample is shaken, allowed to sit for at least twenty four hours, and a decanted portion is introduced to the spectrometer. When the dissolved metals are required, the samples must first be filtered through a 0.45 micron cellulose acetate filter paper and then followed by preservation in 0.2 percent nitric acid. It is important to note that different laboratories may have very different detection limits for some variables analyzed by ICPMS. | 1 | Y | Y | Y | 20919 | 4/3/2015 | 8/26/2024 | 8/26/2024 4:51:04 PM | 2010-11-16 |  | RDGIBSON |  | Concentration | mg/L | Y | (fl. conc.) |  | Tungsten Dissolved (fl. conc.) | 7440-33-7 | Metals - Dissolved; | Metals - Dissolved; EMS code: W--D | Chemical | Dissolved | LAB
W--E | mg/L | Tungsten Extractable | X553 | Metals in water, extractable, ICP-SFMS | The whole water sample is collected in a HDPE/Teflon bottle. In a Class 100 workstation it is preserved with concentrated Nitric Acid at concentration of 0.2 percent (v/v). The acid preservative extracts the metals from material in the sample and a decanted portion is introduced directly to an inductively coupled argon plasma-sector field mass spectrometer (ICP-SFMS) by pneumatic nebulization. Each respective element is measured at a specific mass to charge ratio, m/z value expressed in amu (atomic mass units), and at a specified instrument resolution. LR: Low Resolution ~300, MR: Mid Resolution ~4000 and HR: High Resolution ~10000. Ref: Method number 02-2061, National Laboratory for Environmental Testing, CCIW, Environment Canada, Burlington. Requested April 2003. | 0.001 | Y | Y | Y | 1715 | 6/27/2004 | 9/21/2007 | 4/18/2018 9:28:02 AM | 2009-05-14 | 5/15/2009 11:26:37 AM | RDGIBSON | RDGIBSON | Concentration | mg/L | Y | (fl. conc.) |  | Tungsten Extractable (fl. conc.) | 7440-33-7 | Metals - Extractable; | Metals - Extractable; EMS code: W--E | Chemical | Extractable | LAB
W--T | mg/L | Tungsten Total | MSEA | Metals in Seawater (undigested) by (CRC) ICP-MS | This method describes the multi-elemental determination of trace elements by Collision/Reaction Cell (CRC) ICP-MS. The method measures ions produced by a radio-frequency inductively coupled plasma. Analyte species originating in a liquid are pneumatically nebulized and the resulting aerosol transported by argon gas into the plasma torch. The ions produced are entrained in the plasma gas and extracted, by means of a differentially pumped vacuum interface, into a mass spectrometer. The ions flow through a Collision/Reaction Cell (CRC) where depending on the mode of analysis either the ions passes through the cell unhindered (no gas - classical ICP-MS mode), or collisions/reactions chemistry takes place between the ions and a neutral gas such as hydrogen or helium. Any polyatomic ions which give interference on the analytes are removed. The ions are then sorted according to their mass-to-charge ratios by a quadrupole mass spectrometer. The ions are then sorted according to their mass-to-charge ratios by a quadrupole mass spectrometer. The ions transmitted through the quadrupole are quantified by a channel electron multiplier and the ion information processed by a data handling system. Although most polyatomic interference is removed, any additional interference must be assessed and valid corrections applied or the data flagged to indicate problems. Interference correction must include compensation for background ions contributed by the plasma gas, reagents, and constituents of the sample matrix. Instrument drift as well as suppressions or enhancements of instrument response caused by the sample matrix must be corrected for by the use of internal standards. | 1e-05 | Y | Y | Y | 20 | 4/13/2023 | 4/29/2024 | 6/7/2024 10:03:16 AM | 2009-09-08 |  | RDGIBSON |  | Concentration | mg/L | Y | (fl. conc.) |  | Tungsten Total (fl. conc.) | 7440-33-7 | Metals - Total; | Metals - Total; EMS code: W--T | Chemical | Total | LAB
W--T | mg/L | Tungsten Total | MWID | Metals in Water by ICPMS (digested) | Metals in Water by ICPMS (digested) Based on EPA 200.8 | 0.001 | Y | Y | Y | 46 |  |  | 1/19/2023 12:03:00 PM | 2009-12-24 |  | RDGIBSON |  | Concentration | mg/L | Y | (fl. conc.) |  | Tungsten Total (fl. conc.) | 7440-33-7 | Metals - Total; | Metals - Total; EMS code: W--T | Chemical | Total | LAB
W--T | mg/L | Tungsten Total | F082 | Aliquot,HNO3/HCL dig,ICPMS | Aliquot microwave digested with 3:1 HNO3/HCL. Analysis by ICPMS. For HiVol samples preliminary result are reported in mg/L and loading results are reported in ug/m3 | 0.0001 | Y | Y | Y | 9144 | 4/6/2015 | 8/23/2024 | 8/26/2024 3:51:07 PM | 2015-07-30 |  | RDGIBSON |  | Concentration | mg/L | Y | (fl. conc.) |  | Tungsten Total (fl. conc.) | 7440-33-7 | Metals - Total; | Metals - Total; EMS code: W--T | Chemical | Total | LAB
W--T | ppm (S) | Tungsten Total | MS41 | ALS Minerals ME-MS41 | ALS Minerals ME-MS41. A prepared sample is digested with aqua regia in a graphite heating block. After cooling; the resulting solution is diluted to with deionized water; mixed and analyzed by ICP-AES. Samples are then analyzed by ICP-MS for the remaining suite of elements. | 0.05 | Y | Y | Y | 0 |  |  |  | 2024-01-29 |  | EMS |  | DimensionlessRatio | ppm | N | (ratio) |  | Tungsten Total (ratio) | 7440-33-7 | Metals - Total; | Metals - Total; EMS code: W--T | Chemical | Total | LAB
W--T | mg/L | Tungsten Total | 1349 | Tot Metals in Seawater by CRC ICPMS (BC) | Tot Metals in Seawater by CRC ICPMS (BC) | 0.001 | Y | Y | Y | 98 | 8/16/2019 | 12/5/2021 | 12/13/2021 12:34:17 PM | 2019-09-30 |  | KATFRASE |  | Concentration | mg/L | Y | (fl. conc.) |  | Tungsten Total (fl. conc.) | 7440-33-7 | Metals - Total; | Metals - Total; EMS code: W--T | Chemical | Total | LAB
W--T | mg/L | Tungsten Total | METN | MET-T-NP-CCMS-VA | Ultra trace metals in water are analyzed by CRC ICPMS, based on US EPA Method 6020B (July 2014). The detection limits provided can only be met for undigested samples. This procedure is intended for colorless, non-turbid, acid-preserved water samples (i.e. pristine water samples), having turbidity < 1 NTU and no odor. Where turbidity exceeds 1 NTU, and/or the sample is colored and has an odor, results may be biased low compared to true Total Metals concentrations. ALS recommends that turbidity analysis be requested on samples submitted for this test to aid with interpretation of results. Where turbidity is <1NTU, undigested metals are equivalent to total metals concentrations. | 1e-05 | Y | Y | Y | 280 | 8/6/2019 | 4/4/2024 | 4/29/2024 4:53:06 PM | 2021-06-16 |  | EMS |  | Concentration | mg/L | Y | (fl. conc.) |  | Tungsten Total (fl. conc.) | 7440-33-7 | Metals - Total; | Metals - Total; EMS code: W--T | Chemical | Total | LAB
W--T | mg/L | Tungsten Total | F033 | Acid Dig,Centrifuge,ICPMS | For the determination of total recoverable elements in water or waste water, take 100 ml aliquot from a well mixed, acid preserved, sample containing not more than 0.25% (w/v) total solids and transfer to a 250 ml Griffin beaker. (If total solids are greater than 0.25% reduce the size of the aliquot by a proportionate amount). Add 1 ml of conc. nitric acid and 0.5 ml conc. hydrochloric acid. Heat on a hot plate until the volume has been reduced to 15-20ml, ensuring that the sample does not boil. Allow to cool and transfer to a 50 ml volumetric flask. Dilute to volume with ASTM type I water. Centrifuge the sample or allow to stand overnight to separate insoluble material. Prior analysis, pipette 20 ml into a 50 ml volumetric flask and dilute to volume with ASTM type I water. Analysis should be performed within seventy two hours of the completed preparation. Note that different laboratories may have very different field preservation, digestion procedures and detection limits for this method, so these should be documented. | 1 | Y | Y | Y | 10866 | 11/1/2017 10:00:00 AM | 8/22/2024 3:24:00 AM | 8/25/2024 9:47:36 PM | 2010-11-16 |  | RDGIBSON |  | Concentration | mg/L | Y | (fl. conc.) |  | Tungsten Total (fl. conc.) | 7440-33-7 | Metals - Total; | Metals - Total; EMS code: W--T | Chemical | Total | LAB
W--T | ppm (S) | Tungsten Total | 6041 | Aqua Regia by ICPMS | These were sediment/soils digested with Aqua-Regia (Nitric Acid & Hydrochloric Acids) and then analysis by ICPMS. Results would be reported in mg/kg and/or %. | 0.001 | Y | Y | Y | 21 | 10/13/2019 | 8/5/2020 | 8/10/2020 3:52:07 PM | 2019-09-06 | 10/21/2019 11:56:54 AM | KATFRASE | KATFRASE | DimensionlessRatio | ppm | N | (ratio) |  | Tungsten Total (ratio) | 7440-33-7 | Metals - Total; | Metals - Total; EMS code: W--T | Chemical | Total | LAB
W--T | ppm (S) | Tungsten Total | MS61 | ALS Minerals ME-MS61 | ALS Minerals ME-MS61. A prepared sample is digested with HNO3; HF; and HClO4 in a graphite heating block to near dryness. After cooling; the sample is boiled in diluted HCl. Bulk to volume with deionized water; mixed and analyzed by ICP-AES and ICP-MS for all elements. | 0.1 | Y | Y | Y | 0 |  |  |  | 2024-01-29 |  | EMS |  | DimensionlessRatio | ppm | N | (ratio) |  | Tungsten Total (ratio) | 7440-33-7 | Metals - Total; | Metals - Total; EMS code: W--T | Chemical | Total | LAB
W--T | mg/L | Tungsten Total | X555 | Metals in water, total, "in-bottle" digestion, ICP-SFMS | The whole water sample collected in a HDPE/Teflon bottle, is de-capped in a Class 100 workstation, acidified to 2% (v/v) HNO3 and 1% (v/v) HCl and recapped. The sample is "in-bottle" digested in a convection oven, set at 60 degrees C for 16 hours. The sample is cooled and introduced to an inductively coupled argon plasma-sector field mass spectrometer (ICP-SFMS) by pneumatic nebulization. Each respective element is measured at a specific mass to charge ratio, m/z value expressed in amu (atomic mass units), and at resolution. LR: Low Resolution ~300, MR: Mid Resolution ~4000 and HR: High a specified instrument Resolution ~10000. Ref: Method number 02-2003, National Laboratory for Environmental Testing, CCIW, Environment Canada, Burlington. Requested April 2003. | 1e-06 | Y | Y | Y | 15776 | 6/27/2004 | 12/6/2023 2:08:00 AM | 2/14/2024 9:05:32 AM | 2009-03-31 |  | JCHUDAK |  | Concentration | mg/L | Y | (fl. conc.) |  | Tungsten Total (fl. conc.) | 7440-33-7 | Metals - Total; | Metals - Total; EMS code: W--T | Chemical | Total | LAB
W--T | ppm (S) | Tungsten Total | MS81 | ALS Minerals: ME-MS81 | ALS Minerals: ME-MS81. A prepared sample (0.100 g) is added to lithium metaborate/lithium tetraborate flux; mixed well and fused in a furnace at 1025°C. The resulting melt is then cooled and dissolved in an acid mixture containing nitric; hydrochloric and hydrofluoric acids. This solution is then analyzed by inductively coupled plasma - mass spectrometry. | 1 | Y | Y | Y | 0 |  |  |  | 2024-01-29 |  | EMS |  | DimensionlessRatio | ppm | N | (ratio) |  | Tungsten Total (ratio) | 7440-33-7 | Metals - Total; | Metals - Total; EMS code: W--T | Chemical | Total | LAB
W--T | mg/L | Tungsten Total | X561 | Total metals: FA, HNO3, HR-ICPMS (No Prep)Prep) | Ultra trace metals in water are analyzed by high resolution inductively coupled plasma mass spectrometry (HR-ICPMS), modified from US EPA 200.8, (Revision 5.5). The detection limits provided can only be met for undigested samples. This procedure is intended for colorless, non-turbid, acid-preserved water samples (i.e. pristine water samples), having turbidity < 1 NTU and no odor. Where turbidity exceeds 1 NTU, and/or the sample is colored and has an odor, results may be biased low compared to true Total Metals concentrations. ALS recommends that turbidity analysis be requested on samples submitted for this test to aid with interpretation of results. | 1e-05 | Y | Y | Y | 20 | 10/20/2015 | 10/5/2023 | 1/31/2024 2:37:41 PM | 2016-03-24 |  | RDGIBSON |  | Concentration | mg/L | Y | (fl. conc.) |  | Tungsten Total (fl. conc.) | 7440-33-7 | Metals - Total; | Metals - Total; EMS code: W--T | Chemical | Total | LAB
W--T | ppm (S) | Tungsten Total | 6042 | Aqua Regia by ICPMS | These were sediment/soils digested with Aqua-Regia (Nitric Acid & Hydrochloric Acids) and then analysis by ICPMS. Results would be reported in ppm s. | 0.001 | Y | Y | Y | 71 | 5/14/2022 | 7/27/2024 | 7/29/2024 9:13:08 AM | 2022-06-28 |  | EMS |  | DimensionlessRatio | ppm | N | (ratio) |  | Tungsten Total (ratio) | 7440-33-7 | Metals - Total; | Metals - Total; EMS code: W--T | Chemical | Total | LAB
W--T | ug/g | Tungsten Total | FS82 | PBM,dry60C,10mesh,SALM-ICP-MS | PBM;dry60C;10mesh;SALM-ICP-MS | 1 | Y | Y | Y | 178 | 11/10/2017 4:09:00 PM | 7/24/2024 | 7/29/2024 5:17:09 PM | 2010-11-16 |  | RDGIBSON |  | MassPerMass | mg/kg | Y | (mass conc.) |  | Tungsten Total (mass conc.) | 7440-33-7 | Metals - Total; | Metals - Total; EMS code: W--T | Chemical | Total | LAB
W--T | mg/L | Tungsten Total | F048 | Acid dig,conc'n,nebul.,ICP-OES | This method provides an automated procedure for the Inductively Coupled Argon Plasma Optical Emission Spectrophotometric (ICP-OES) determination and quantitation of trace amounts of metals in surface, ground and precipitation waters after a manual acid digestion and concentration procedure and ultrasonic nebulisation. Organically-bound metals and metal complexes are converted to metal ions by digestion with nitric acid [method details says HCl]. To increase the signal detected by the ICP, the samples are concentrated 10 times. The liquid sample is pumped onto the piezoelectric transducer, where the ultrasonic action efficiently converts it to a fine, dense aerosol. The nebulizer carrier gas flow then transports the wet aerosol through a heated U-tube, where the solvent is vaporized. Subsequently, most of the solvent vapour is removed from the aerosol stream by passing the stream through a chilled condenser. The result is a dry, analyte-laden aerosol which is introduced to the plasma. There, the high temperature desolvates the aerosol and dissociates the sample material into individual atoms and ions, which are excited to emit light at wave lengths characteristic of the elements present. Dissolved samples are field filtered through 0.45 micron cellulose acetate filter prior to acid preservation. Samples for extractable analysis are acidified with 2 mL concentrated HNO3 per litre and allowed to stand a minimum of 24 hours before analysis. Clients providing these samples are advised to preserve samples at the collection site. | 0.1 | Y | Y | Y | 71 | 5/8/2003 9:40:00 AM | 9/12/2003 3:16:00 PM | 6/7/2024 5:30:17 PM | 1998-11-13 |  | NPEPPIN |  | Concentration | mg/L | Y | (fl. conc.) |  | Tungsten Total (fl. conc.) | 7440-33-7 | Metals - Total; | Metals - Total; EMS code: W--T | Chemical | Total | LAB
W--T | mg/L | Tungsten Total | F121 | Acid,ICPMS | Acidified, no digestion, ICPMS | 1 | Y | Y | Y | 1 |  |  | 6/7/2024 5:30:15 PM | 2010-11-16 |  | RDGIBSON |  | Concentration | mg/L | Y | (fl. conc.) |  | Tungsten Total (fl. conc.) | 7440-33-7 | Metals - Total; | Metals - Total; EMS code: W--T | Chemical | Total | LAB
X001 | ug/g | Xylene | CAL2 | Calculation (ug/g) | Calculation (ug/g) | 0 | Y | Y | Y | 110 | 3/29/2006 | 2/4/2022 | 1/11/2024 1:07:09 PM | 2002-07-11 |  | NPEPPIN |  | MassPerMass | mg/kg | Y | (mass conc.) |  | Xylene (mass conc.) | 1330-20-7 | Volatile Organic Compounds (VOCs); | Volatile Organic Compounds (VOCs); EMS code: X001 | Chemical |  | LAB
X001 | mg/L | Xylene | MS01 | Purge&Trap GC/MS (VOC-W) | Purge&Trap GC/MS (VOC-W) | 0.4 | Y | Y | Y | 891 | 7/29/2005 7:02:00 AM | 6/4/2024 10:24:00 AM | 6/13/2024 5:01:27 PM | 1996-04-12 | 3/18/2003 2:49:40 PM | EMS | UMULLER | Concentration | mg/L | Y | (fl. conc.) |  | Xylene (fl. conc.) | 1330-20-7 | Volatile Organic Compounds (VOCs); | Volatile Organic Compounds (VOCs); EMS code: X001 | Chemical |  | LAB
X001 | mg/L | Xylene | X384 | Purge and trap GC/MS | The sample is analyzed by purge and trap gas chromatography with detection by mass spectrometry. The procedure involves purging the volatile compounds from the sample with an inert gas, and trapping them on a solid sorbent. When purging of the sample is complete, the trap is heated and the compounds are transferred to the gas chromatographic column. Analysis is then accomplished by separation of the components by gas chromatography with detection by mass spectrometry. | 0.5 | Y | Y | Y | 54 | 1/24/2001 | 4/30/2002 | 3/19/2004 8:18:44 AM | 1996-04-12 | 6/3/1998 6:51:40 PM | EMS |  | Concentration | mg/L | Y | (fl. conc.) |  | Xylene (fl. conc.) | 1330-20-7 | Volatile Organic Compounds (VOCs); | Volatile Organic Compounds (VOCs); EMS code: X001 | Chemical |  | LAB
X001 | ug/g | Xylene | MS02 | Purge&Trap GC/MS (VOC-S) | Purge&Trap GC/MS (VOC-S) | 0.002 | Y | Y | Y | 5 | 10/28/2022 5:04:00 PM | 10/28/2022 5:04:00 PM | 1/7/2023 3:12:10 PM | 1996-04-12 | 6/3/1998 6:51:40 PM | EMS |  | MassPerMass | mg/kg | Y | (mass conc.) |  | Xylene (mass conc.) | 1330-20-7 | Volatile Organic Compounds (VOCs); | Volatile Organic Compounds (VOCs); EMS code: X001 | Chemical |  | LAB
X001 | mg/L | Xylene | 8260D | EPA 8260D (mod) | EPA 8260D (mod) VOLATILE ORGANIC COMPOUNDS BY GAS CHROMATOGRAPHY/MASS SPECTROMETRY This method is used to determine volatile organic compounds (VOCs) in a variety of solid waste matrices. | 0.0075 | Y | Y | Y | 0 |  |  |  | 2024-01-25 |  | RDGIBSON |  | Concentration | mg/L | Y | (fl. conc.) |  | Xylene (fl. conc.) | 1330-20-7 | Volatile Organic Compounds (VOCs); | Volatile Organic Compounds (VOCs); EMS code: X001 | Chemical |  | LAB
X001 | mg/L | Xylene | XCAL | Calculation (mg/L) | Calculation (mg/L) | 0 | Y | Y | Y | 1415 | 5/4/2002 | 8/14/2024 | 8/16/2024 4:05:21 PM | 2002-07-11 |  | NPEPPIN |  | Concentration | mg/L | Y | (fl. conc.) |  | Xylene (fl. conc.) | 1330-20-7 | Volatile Organic Compounds (VOCs); | Volatile Organic Compounds (VOCs); EMS code: X001 | Chemical |  | LAB
X002 | mg/L | o-Xylene | MS01 | Purge&Trap GC/MS (VOC-W) | Purge&Trap GC/MS (VOC-W) | 0.0005 | Y | Y | Y | 205 | 4/15/2002 | 4/1/2011 | 7/9/2012 2:09:49 PM | 2002-08-12 |  | NPEPPIN |  | Concentration | mg/L | Y | (fl. conc.) |  | o-Xylene (fl. conc.) | 95-47-6 | Volatile Organic Compounds (VOCs); | Volatile Organic Compounds (VOCs); EMS code: X002 | Chemical |  | LAB
X002 | ug/g | o-Xylene | MS02 | Purge&Trap GC/MS (VOC-S) | Purge&Trap GC/MS (VOC-S) | 0.05 | Y | Y | Y | 15 | 1/14/2022 | 2/4/2022 | 1/11/2024 1:07:09 PM | 2022-07-27 |  | EMS |  | MassPerMass | mg/kg | Y | (mass conc.) |  | o-Xylene (mass conc.) | 95-47-6 | Volatile Organic Compounds (VOCs); | Volatile Organic Compounds (VOCs); EMS code: X002 | Chemical |  | LAB
X002 | ug/g | o-Xylene | FS04 | Static Headspace:HRGC/LRMS | Static Headspace anaylsis is a method for the determination of volatile compounds in liquids and/or solids by analysing the vapour phase of a sample in a closed system; in this case a sealed vial. Static Headspace is sometimes referred to as Equilibrium Headspace because the vapour phase above the sample in a sealed vial is brought into thermodynamic equilibrium with the sample prior to analysis. Volatile organic compounds are thermally extracted from the sample matrix by Static Headspace sampling and introduced directly to a thick-film megabore GC column. Compounds eluting from the GC column are introduced into a Mass Spectrometer and target compounds are identified by a comparison of both retention times and mass spectra with those of reference standards stored in a computerized data system. Reference retention times and mass spectra for analytes are obtained by the measurement of calibration standards under precisely the same extraction conditions used for samples. The concentration of each identified component is measured by relating the response of the sample quantitation mass ions to the response of the quantitation mass ions of a series of external standards. A deuterated surrogate recovery standard is added to provide information on matrix effects. | 0.01 | Y | Y | Y | 120 | 7/2/1999 2:40:00 PM | 3/7/2013 3:50:00 AM | 3/19/2013 1:15:56 PM | 1996-09-17 | 6/3/1998 6:51:45 PM | MPAK |  | MassPerMass | mg/kg | Y | (mass conc.) |  | o-Xylene (mass conc.) | 95-47-6 | Volatile Organic Compounds (VOCs); | Volatile Organic Compounds (VOCs); EMS code: X002 | Chemical |  | LAB
X002 | mg/L | o-Xylene | X384 | Purge and trap GC/MS | The sample is analyzed by purge and trap gas chromatography with detection by mass spectrometry. The procedure involves purging the volatile compounds from the sample with an inert gas, and trapping them on a solid sorbent. When purging of the sample is complete, the trap is heated and the compounds are transferred to the gas chromatographic column. Analysis is then accomplished by separation of the components by gas chromatography with detection by mass spectrometry. | 0.0005 | Y | Y | Y | 73 | 2/7/2002 10:55:00 AM | 2/17/2006 | 3/31/2006 1:01:04 AM | 1996-04-12 | 6/3/1998 6:51:41 PM | EMS |  | Concentration | mg/L | Y | (fl. conc.) |  | o-Xylene (fl. conc.) | 95-47-6 | Volatile Organic Compounds (VOCs); | Volatile Organic Compounds (VOCs); EMS code: X002 | Chemical |  | LAB
X002 | mg/L | o-Xylene | MS04 | Static Headspace GC/MS (Water) | (Water) A static headspace procedure is used as a means of driving the analytes out of the matrix into a headspace. Once equilibrium is achieved, a fixed volume of headspeace is taken for subsequent determination by Gas Chromatography - Mass Spectroscopy (GC/MS) using Selected Ion Monitoring. | 0.0005 | Y | Y | Y | 61 | 8/13/2009 | 6/28/2024 12:29:00 AM | 7/9/2024 3:31:02 PM | 2009-12-23 |  | RDGIBSON |  | Concentration | mg/L | Y | (fl. conc.) |  | o-Xylene (fl. conc.) | 95-47-6 | Volatile Organic Compounds (VOCs); | Volatile Organic Compounds (VOCs); EMS code: X002 | Chemical |  | LAB
X002 | mg/L | o-Xylene | 8260D | EPA 8260D (mod) | EPA 8260D (mod) VOLATILE ORGANIC COMPOUNDS BY GAS CHROMATOGRAPHY/MASS SPECTROMETRY This method is used to determine volatile organic compounds (VOCs) in a variety of solid waste matrices. | 0.005 | Y | Y | Y | 0 |  |  |  | 2024-01-25 |  | RDGIBSON |  | Concentration | mg/L | Y | (fl. conc.) |  | o-Xylene (fl. conc.) | 95-47-6 | Volatile Organic Compounds (VOCs); | Volatile Organic Compounds (VOCs); EMS code: X002 | Chemical |  | LAB
X002 | mg/L | o-Xylene | F007 | Static Headspace:HRGC/LRMS | Static headspace analysis is a method for the determination of volatile compounds in liquids and/or solids by analysing the vapour phase of a sample in a closed system; in this case a sealed vial. Static Headspace is sometimes referred to as Equilibrium Headspace because the vapour phase above the sample in a sealed vial is brought into thermodynamic equilibrium with the sample prior to analysis. Volatile organic compounds are thermally extracted from the sample matrix by Static Headspace sampling and introduced directly to a thick-film megabore GC column. Compounds eluting from the GC column are introduced into a Mass Spectrometer and target compounds are identified by a comparison of both retention times and mass spectra with those of reference standards stored in a computerized data system. Reference retention times and mass spectra for analytes are obtained by the measurement of calibration standards under precisely the same extraction conditions used for samples. The concentration of each identified component is measured by relating the response of the sample quantitation mass ions to the response of the quantitation mass ions of a series of external standards. A deuterated surrogate recovery standard is added to provide information on matrix effects. | 0.0005 | Y | Y | Y | 1408 | 4/21/1998 2:34:00 PM | 8/14/2024 | 8/16/2024 4:05:21 PM | 1996-09-17 | 6/3/1998 6:51:45 PM | MPAK |  | Concentration | mg/L | Y | (fl. conc.) |  | o-Xylene (fl. conc.) | 95-47-6 | Volatile Organic Compounds (VOCs); | Volatile Organic Compounds (VOCs); EMS code: X002 | Chemical |  | LAB
X002 | ug/g | o-Xylene | MEOH | MeOH extrac,GC/FID,nC6-nC10 | MeOH extrac;GC/FID;nC6-nC10. EPA8260B; 5021; 5035; BC MOE | 0.05 | Y | Y | Y | 49 | 5/21/2015 | 8/11/2019 | 8/13/2019 10:56:08 AM | 2015-05-26 |  | RDGIBSON |  | MassPerMass | mg/kg | Y | (mass conc.) |  | o-Xylene (mass conc.) | 95-47-6 | Volatile Organic Compounds (VOCs); | Volatile Organic Compounds (VOCs); EMS code: X002 | Chemical |  | LAB
X003 | ug/g | M;P-Xylene | FS04 | Static Headspace:HRGC/LRMS | Static Headspace anaylsis is a method for the determination of volatile compounds in liquids and/or solids by analysing the vapour phase of a sample in a closed system; in this case a sealed vial. Static Headspace is sometimes referred to as Equilibrium Headspace because the vapour phase above the sample in a sealed vial is brought into thermodynamic equilibrium with the sample prior to analysis. Volatile organic compounds are thermally extracted from the sample matrix by Static Headspace sampling and introduced directly to a thick-film megabore GC column. Compounds eluting from the GC column are introduced into a Mass Spectrometer and target compounds are identified by a comparison of both retention times and mass spectra with those of reference standards stored in a computerized data system. Reference retention times and mass spectra for analytes are obtained by the measurement of calibration standards under precisely the same extraction conditions used for samples. The concentration of each identified component is measured by relating the response of the sample quantitation mass ions to the response of the quantitation mass ions of a series of external standards. A deuterated surrogate recovery standard is added to provide information on matrix effects. | 0.01 | Y | Y | Y | 89 | 7/2/1999 2:40:00 PM | 3/7/2013 3:50:00 AM | 3/19/2013 1:15:56 PM | 1996-09-17 | 6/3/1998 6:51:45 PM | MPAK |  | MassPerMass | mg/kg | Y | (mass conc.) |  | M;P-Xylene (mass conc.) | 108-38-3, 106-42-3 | Volatile Organic Compounds (VOCs); | Volatile Organic Compounds (VOCs); EMS code: X003 | Chemical |  | LAB
X003 | ug/g | M;P-Xylene | X263 | Purge and trap GC/MS | Purge and trap GC/MS | 0.001 | Y | Y | Y | 31 | 11/19/2004 | 4/4/2006 | 4/27/2006 1:02:37 AM | 1996-04-12 | 6/3/1998 6:51:42 PM | EMS |  | MassPerMass | mg/kg | Y | (mass conc.) |  | M;P-Xylene (mass conc.) | 108-38-3, 106-42-3 | Volatile Organic Compounds (VOCs); | Volatile Organic Compounds (VOCs); EMS code: X003 | Chemical |  | LAB
X003 | mg/L | M;P-Xylene | MS01 | Purge&Trap GC/MS (VOC-W) | Purge&Trap GC/MS (VOC-W) | 0.0005 | Y | Y | Y | 205 | 4/15/2002 | 4/1/2011 | 7/9/2012 2:09:49 PM | 2002-08-12 | 8/12/2002 4:16:32 PM | NPEPPIN | NPEPPIN | Concentration | mg/L | Y | (fl. conc.) |  | M;P-Xylene (fl. conc.) | 108-38-3, 106-42-3 | Volatile Organic Compounds (VOCs); | Volatile Organic Compounds (VOCs); EMS code: X003 | Chemical |  | LAB
X003 | mg/L | M;P-Xylene | F007 | Static Headspace:HRGC/LRMS | Static headspace analysis is a method for the determination of volatile compounds in liquids and/or solids by analysing the vapour phase of a sample in a closed system; in this case a sealed vial. Static Headspace is sometimes referred to as Equilibrium Headspace because the vapour phase above the sample in a sealed vial is brought into thermodynamic equilibrium with the sample prior to analysis. Volatile organic compounds are thermally extracted from the sample matrix by Static Headspace sampling and introduced directly to a thick-film megabore GC column. Compounds eluting from the GC column are introduced into a Mass Spectrometer and target compounds are identified by a comparison of both retention times and mass spectra with those of reference standards stored in a computerized data system. Reference retention times and mass spectra for analytes are obtained by the measurement of calibration standards under precisely the same extraction conditions used for samples. The concentration of each identified component is measured by relating the response of the sample quantitation mass ions to the response of the quantitation mass ions of a series of external standards. A deuterated surrogate recovery standard is added to provide information on matrix effects. | 0.0005 | Y | Y | Y | 1408 | 4/21/1998 2:34:00 PM | 8/14/2024 | 8/16/2024 4:05:21 PM | 1996-09-17 | 6/3/1998 6:51:45 PM | MPAK |  | Concentration | mg/L | Y | (fl. conc.) |  | M;P-Xylene (fl. conc.) | 108-38-3, 106-42-3 | Volatile Organic Compounds (VOCs); | Volatile Organic Compounds (VOCs); EMS code: X003 | Chemical |  | LAB
X003 | mg/L | M;P-Xylene | X384 | Purge and trap GC/MS | The sample is analyzed by purge and trap gas chromatography with detection by mass spectrometry. The procedure involves purging the volatile compounds from the sample with an inert gas, and trapping them on a solid sorbent. When purging of the sample is complete, the trap is heated and the compounds are transferred to the gas chromatographic column. Analysis is then accomplished by separation of the components by gas chromatography with detection by mass spectrometry. | 0.0005 | Y | Y | Y | 76 | 2/7/2002 10:55:00 AM | 2/17/2006 | 3/31/2006 1:01:04 AM | 1996-04-12 | 6/3/1998 6:51:41 PM | EMS |  | Concentration | mg/L | Y | (fl. conc.) |  | M;P-Xylene (fl. conc.) | 108-38-3, 106-42-3 | Volatile Organic Compounds (VOCs); | Volatile Organic Compounds (VOCs); EMS code: X003 | Chemical |  | LAB
X003 | mg/L | M;P-Xylene | MS04 | Static Headspace GC/MS (Water) | (Water) A static headspace procedure is used as a means of driving the analytes out of the matrix into a headspace. Once equilibrium is achieved, a fixed volume of headspeace is taken for subsequent determination by Gas Chromatography - Mass Spectroscopy (GC/MS) using Selected Ion Monitoring. | 0.001 | Y | Y | Y | 76 | 8/13/2009 | 6/28/2024 12:29:00 AM | 7/9/2024 3:31:02 PM | 2009-12-23 |  | RDGIBSON |  | Concentration | mg/L | Y | (fl. conc.) |  | M;P-Xylene (fl. conc.) | 108-38-3, 106-42-3 | Volatile Organic Compounds (VOCs); | Volatile Organic Compounds (VOCs); EMS code: X003 | Chemical |  | LAB
X003 | ug/g | M;P-Xylene | MS02 | Purge&Trap GC/MS (VOC-S) | Purge&Trap GC/MS (VOC-S) | 0.05 | Y | Y | Y | 15 | 1/14/2022 | 2/4/2022 | 1/11/2024 1:07:09 PM | 2022-07-27 |  | EMS |  | MassPerMass | mg/kg | Y | (mass conc.) |  | M;P-Xylene (mass conc.) | 108-38-3, 106-42-3 | Volatile Organic Compounds (VOCs); | Volatile Organic Compounds (VOCs); EMS code: X003 | Chemical |  | LAB
X003 | ug/g | M;P-Xylene | MEOH | MeOH extrac,GC/FID,nC6-nC10 | MeOH extrac;GC/FID;nC6-nC10. EPA8260B; 5021; 5035; BC MOE | 0.05 | Y | Y | Y | 49 | 5/21/2015 | 8/11/2019 | 8/13/2019 10:56:08 AM | 2015-05-26 |  | RDGIBSON |  | MassPerMass | mg/kg | Y | (mass conc.) |  | M;P-Xylene (mass conc.) | 108-38-3, 106-42-3 | Volatile Organic Compounds (VOCs); | Volatile Organic Compounds (VOCs); EMS code: X003 | Chemical |  | LAB
X003 | mg/L | M;P-Xylene | 8260D | EPA 8260D (mod) | EPA 8260D (mod) VOLATILE ORGANIC COMPOUNDS BY GAS CHROMATOGRAPHY/MASS SPECTROMETRY This method is used to determine volatile organic compounds (VOCs) in a variety of solid waste matrices. | 0.005 | Y | Y | Y | 0 |  |  |  | 2024-01-25 |  | RDGIBSON |  | Concentration | mg/L | Y | (fl. conc.) |  | M;P-Xylene (fl. conc.) | 108-38-3, 106-42-3 | Volatile Organic Compounds (VOCs); | Volatile Organic Compounds (VOCs); EMS code: X003 | Chemical |  | LAB
X006 | ug/g | Xylenes (Total) | CAL2 | Calculation (ug/g) | Calculation (ug/g) | 0.075 | Y | Y | Y | 83 |  |  | 7/19/2024 8:20:35 AM | 2021-04-29 | 4/30/2021 11:21:46 AM | RDGIBSON | EMS | MassPerMass | mg/kg | Y | (mass conc.) |  | Xylenes (Total) (mass conc.) | 1330-20-7 | Volatile Organic Compounds (VOCs); | Volatile Organic Compounds (VOCs); EMS code: X006 | Chemical | Total | LAB
X006 | mg/L | Xylenes (Total) | 8260D | EPA 8260D (mod) | EPA 8260D (mod) VOLATILE ORGANIC COMPOUNDS BY GAS CHROMATOGRAPHY/MASS SPECTROMETRY This method is used to determine volatile organic compounds (VOCs) in a variety of solid waste matrices. | 0.0075 | Y | Y | Y | 0 |  |  |  | 2024-01-25 |  | RDGIBSON |  | Concentration | mg/L | Y | (fl. conc.) |  | Xylenes (Total) (fl. conc.) | 1330-20-7 | Volatile Organic Compounds (VOCs); | Volatile Organic Compounds (VOCs); EMS code: X006 | Chemical | Total | LAB
X006 | mg/L | Xylenes (Total) | MS04 | Static Headspace GC/MS (Water) | (Water) A static headspace procedure is used as a means of driving the analytes out of the matrix into a headspace. Once equilibrium is achieved, a fixed volume of headspeace is taken for subsequent determination by Gas Chromatography - Mass Spectroscopy (GC/MS) using Selected Ion Monitoring. | 0.1 | Y | Y | Y | 54 | 12/9/2010 | 6/28/2024 12:29:00 AM | 7/30/2024 6:32:12 AM | 2010-11-16 |  | RDGIBSON |  | Concentration | mg/L | Y | (fl. conc.) |  | Xylenes (Total) (fl. conc.) | 1330-20-7 | Volatile Organic Compounds (VOCs); | Volatile Organic Compounds (VOCs); EMS code: X006 | Chemical | Total | LAB
Y--D | mg/L | Yttrium Dissolved | X556 | Metals in water, dissolved, "in-bottle" digestion, ICP-SFMS | The whole water sample is filtered through a 0.45 micron cellulose acetate filter paper and the filtrate is stored in a HDPE/Teflon bottle. The container is de-capped in a Class 100 workstation, acidified to 2% (v/v) HNO3 and 1% (v/v) HCl and recapped. The sample is "in-bottle" digested in a convection oven, set at 60 degrees C for 16 hours. The sample is cooled and introduced to an inductively coupled argon plasma-sector field mass spectrometer (ICP-SFMS) by pneumatic nebulization. Each respective element is measured at a specific mass to charge ratio, m/z value expressed in amu (atomic mass units), and at a specified intrument resolution. LR: Low Resolution ~300, MR: Mid Resolution ~4000 and HR: High Resolution ~10000. Ref: Method number 02-2003, National Laboratory for Environmental Testing, CCIW, Environment Canada, Burlington, Requested April 2003. | 5e-06 | Y | Y | Y | 4158 | 6/27/2004 | 12/6/2023 2:04:00 AM | 2/14/2024 9:05:32 AM | 2008-06-12 |  | RDGIBSON |  | Concentration | mg/L | Y | (fl. conc.) |  | Yttrium Dissolved (fl. conc.) | 7440-65-5 | Metals - Dissolved; | Metals - Dissolved; EMS code: Y--D | Chemical | Dissolved | LAB
Y--D | mg/L | Yttrium Dissolved | F032 | Filt.,Acid,ICPMS | There is no acid digestion required for this analysis. A water sample is collected and preserved in 0.2 percent nitric acid. To analyse the extractable metals the sample is shaken, allowed to sit for at least twenty four hours, and a decanted portion is introduced to the spectrometer. When the dissolved metals are required, the samples must first be filtered through a 0.45 micron cellulose acetate filter paper and then followed by preservation in 0.2 percent nitric acid. It is important to note that different laboratories may have very different detection limits for some variables analyzed by ICPMS. | 0.001 | Y | Y | Y | 48 | 7/30/2020 | 12/19/2023 | 1/2/2024 4:32:04 PM | 1998-06-22 |  | NPEPPIN |  | Concentration | mg/L | Y | (fl. conc.) |  | Yttrium Dissolved (fl. conc.) | 7440-65-5 | Metals - Dissolved; | Metals - Dissolved; EMS code: Y--D | Chemical | Dissolved | LAB
Y--D | mg/L | Yttrium Dissolved | X356 | LF,LA:HNO3:ICP | Samples are lab-filtered (0.45 um); lab-acidified with HNO3; analysis by ICP/AES. | 0.001 | Y | Y | Y | 391 |  |  | 3/23/2010 5:08:09 PM | 1998-05-22 | 6/3/1998 6:51:49 PM | NPEPPIN |  | Concentration | mg/L | Y | (fl. conc.) |  | Yttrium Dissolved (fl. conc.) | 7440-65-5 | Metals - Dissolved; | Metals - Dissolved; EMS code: Y--D | Chemical | Dissolved | LAB
Y--D | mg/L | Yttrium Dissolved | X562 | Dissolved metals: FF, FA, HNO3, HR-ICPMS | Ultra trace metals in water are analyzed by high resolution inductively coupled plasma mass spectrometry (HR-ICPMS), based on US EPA Method 200.8, (Oct. 1999). This procedure is intended for pristine field-filtered acid-preserved water samples. The detection limits (LOR) for this test are based on lab instrumental analysis only, not including filtration. ALS-supplied field filtration equipment does not support these LOR. Therefore, because of the high probability of false positives due to filtration, it is strongly recommended that a filtration blank be analyzed to aid in data interpretation. | 5e-06 | Y | Y | Y | 10 | 10/20/2015 | 10/24/2018 | 10/24/2018 12:26:14 PM | 2016-03-24 |  | RDGIBSON |  | Concentration | mg/L | Y | (fl. conc.) |  | Yttrium Dissolved (fl. conc.) | 7440-65-5 | Metals - Dissolved; | Metals - Dissolved; EMS code: Y--D | Chemical | Dissolved | LAB
Y--D | mg/L | Yttrium Dissolved | METC | MET-D-NP-CCMS-VA | Ultra trace metals in water are analyzed by CRC ICPMS, based on US EPA Method 6020B (July 2014). This procedure is intended for pristine field-filtered acid-preserved water samples. The detection limits (LOR) for this test are based on lab instrumental analysis only, not including filtration. ALS-supplied field filtration equipment does not support these LOR. Therefore, because of the high probability of false positives due to filtration, it is strongly recommended that a filtration blank be analysed to aid in data interpretation. | 5e-06 | Y | Y | Y | 394 | 3/30/2020 | 4/9/2024 | 4/29/2024 4:53:06 PM | 2022-10-07 |  | EMS |  | Concentration | mg/L | Y | (fl. conc.) |  | Yttrium Dissolved (fl. conc.) | 7440-65-5 | Metals - Dissolved; | Metals - Dissolved; EMS code: Y--D | Chemical | Dissolved | LAB
Y--D | mg/L | Yttrium Dissolved | MSEA | Metals in Seawater (undigested) by (CRC) ICP-MS | This method describes the multi-elemental determination of trace elements by Collision/Reaction Cell (CRC) ICP-MS. The method measures ions produced by a radio-frequency inductively coupled plasma. Analyte species originating in a liquid are pneumatically nebulized and the resulting aerosol transported by argon gas into the plasma torch. The ions produced are entrained in the plasma gas and extracted, by means of a differentially pumped vacuum interface, into a mass spectrometer. The ions flow through a Collision/Reaction Cell (CRC) where depending on the mode of analysis either the ions passes through the cell unhindered (no gas - classical ICP-MS mode), or collisions/reactions chemistry takes place between the ions and a neutral gas such as hydrogen or helium. Any polyatomic ions which give interference on the analytes are removed. The ions are then sorted according to their mass-to-charge ratios by a quadrupole mass spectrometer. The ions are then sorted according to their mass-to-charge ratios by a quadrupole mass spectrometer. The ions transmitted through the quadrupole are quantified by a channel electron multiplier and the ion information processed by a data handling system. Although most polyatomic interference is removed, any additional interference must be assessed and valid corrections applied or the data flagged to indicate problems. Interference correction must include compensation for background ions contributed by the plasma gas, reagents, and constituents of the sample matrix. Instrument drift as well as suppressions or enhancements of instrument response caused by the sample matrix must be corrected for by the use of internal standards. | 0.0005 | Y | Y | Y | 4 | 4/12/2023 | 4/29/2024 | 5/16/2024 2:00:14 PM | 2023-05-26 |  | EMS |  | Concentration | mg/L | Y | (fl. conc.) |  | Yttrium Dissolved (fl. conc.) | 7440-65-5 | Metals - Dissolved; | Metals - Dissolved; EMS code: Y--D | Chemical | Dissolved | LAB
Y--E | mg/L | Yttrium Extractable | X553 | Metals in water, extractable, ICP-SFMS | The whole water sample is collected in a HDPE/Teflon bottle. In a Class 100 workstation it is preserved with concentrated Nitric Acid at concentration of 0.2 percent (v/v). The acid preservative extracts the metals from material in the sample and a decanted portion is introduced directly to an inductively coupled argon plasma-sector field mass spectrometer (ICP-SFMS) by pneumatic nebulization. Each respective element is measured at a specific mass to charge ratio, m/z value expressed in amu (atomic mass units), and at a specified instrument resolution. LR: Low Resolution ~300, MR: Mid Resolution ~4000 and HR: High Resolution ~10000. Ref: Method number 02-2061, National Laboratory for Environmental Testing, CCIW, Environment Canada, Burlington. Requested April 2003. | 0.001 | Y | Y | Y | 1717 | 9/25/2003 | 9/21/2007 | 4/18/2018 9:28:02 AM | 2009-05-14 | 5/15/2009 11:26:53 AM | RDGIBSON | RDGIBSON | Concentration | mg/L | Y | (fl. conc.) |  | Yttrium Extractable (fl. conc.) | 7440-65-5 | Metals - Extractable; | Metals - Extractable; EMS code: Y--E | Chemical | Extractable | LAB
Y--T | ppm (S) | Yttrium Total | MS81 | ALS Minerals: ME-MS81 | ALS Minerals: ME-MS81. A prepared sample (0.100 g) is added to lithium metaborate/lithium tetraborate flux; mixed well and fused in a furnace at 1025°C. The resulting melt is then cooled and dissolved in an acid mixture containing nitric; hydrochloric and hydrofluoric acids. This solution is then analyzed by inductively coupled plasma - mass spectrometry. | 0.1 | Y | Y | Y | 0 |  |  |  | 2024-01-29 |  | EMS |  | DimensionlessRatio | ppm | N | (ratio) |  | Yttrium Total (ratio) | 7440-65-5 | Metals - Total; | Metals - Total; EMS code: Y--T | Chemical | Total | LAB
Y--T | ppm (S) | Yttrium Total | MS61 | ALS Minerals ME-MS61 | ALS Minerals ME-MS61. A prepared sample is digested with HNO3; HF; and HClO4 in a graphite heating block to near dryness. After cooling; the sample is boiled in diluted HCl. Bulk to volume with deionized water; mixed and analyzed by ICP-AES and ICP-MS for all elements. | 0.1 | Y | Y | Y | 0 |  |  |  | 2024-01-29 |  | EMS |  | DimensionlessRatio | ppm | N | (ratio) |  | Yttrium Total (ratio) | 7440-65-5 | Metals - Total; | Metals - Total; EMS code: Y--T | Chemical | Total | LAB
Y--T | mg/L | Yttrium Total | X555 | Metals in water, total, "in-bottle" digestion, ICP-SFMS | The whole water sample collected in a HDPE/Teflon bottle, is de-capped in a Class 100 workstation, acidified to 2% (v/v) HNO3 and 1% (v/v) HCl and recapped. The sample is "in-bottle" digested in a convection oven, set at 60 degrees C for 16 hours. The sample is cooled and introduced to an inductively coupled argon plasma-sector field mass spectrometer (ICP-SFMS) by pneumatic nebulization. Each respective element is measured at a specific mass to charge ratio, m/z value expressed in amu (atomic mass units), and at resolution. LR: Low Resolution ~300, MR: Mid Resolution ~4000 and HR: High a specified instrument Resolution ~10000. Ref: Method number 02-2003, National Laboratory for Environmental Testing, CCIW, Environment Canada, Burlington. Requested April 2003. | 5e-06 | Y | Y | Y | 15778 | 6/27/2004 | 12/6/2023 2:08:00 AM | 2/14/2024 9:05:32 AM | 2009-03-31 |  | JCHUDAK |  | Concentration | mg/L | Y | (fl. conc.) |  | Yttrium Total (fl. conc.) | 7440-65-5 | Metals - Total; | Metals - Total; EMS code: Y--T | Chemical | Total | LAB
Y--T | ppm (S) | Yttrium Total | MS41 | ALS Minerals ME-MS41 | ALS Minerals ME-MS41. A prepared sample is digested with aqua regia in a graphite heating block. After cooling; the resulting solution is diluted to with deionized water; mixed and analyzed by ICP-AES. Samples are then analyzed by ICP-MS for the remaining suite of elements. | 0.05 | Y | Y | Y | 0 |  |  |  | 2024-01-29 |  | EMS |  | DimensionlessRatio | ppm | N | (ratio) |  | Yttrium Total (ratio) | 7440-65-5 | Metals - Total; | Metals - Total; EMS code: Y--T | Chemical | Total | LAB
Y--T | ppm (S) | Yttrium Total | 6042 | Aqua Regia by ICPMS | These were sediment/soils digested with Aqua-Regia (Nitric Acid & Hydrochloric Acids) and then analysis by ICPMS. Results would be reported in ppm s. | 0.003 | Y | Y | Y | 71 | 5/14/2022 | 7/27/2024 | 7/29/2024 9:13:08 AM | 2022-06-28 |  | EMS |  | DimensionlessRatio | ppm | N | (ratio) |  | Yttrium Total (ratio) | 7440-65-5 | Metals - Total; | Metals - Total; EMS code: Y--T | Chemical | Total | LAB
Y--T | mg/L | Yttrium Total | F082 | Aliquot,HNO3/HCL dig,ICPMS | Aliquot microwave digested with 3:1 HNO3/HCL. Analysis by ICPMS. For HiVol samples preliminary result are reported in mg/L and loading results are reported in ug/m3 | 5e-06 | Y | Y | Y | 69 | 7/30/2020 | 8/16/2024 | 8/21/2024 11:26:07 AM | 2021-11-03 |  | RDGIBSON |  | Concentration | mg/L | Y | (fl. conc.) |  | Yttrium Total (fl. conc.) | 7440-65-5 | Metals - Total; | Metals - Total; EMS code: Y--T | Chemical | Total | LAB
Y--T | mg/L | Yttrium Total | X561 | Total metals: FA, HNO3, HR-ICPMS (No Prep)Prep) | Ultra trace metals in water are analyzed by high resolution inductively coupled plasma mass spectrometry (HR-ICPMS), modified from US EPA 200.8, (Revision 5.5). The detection limits provided can only be met for undigested samples. This procedure is intended for colorless, non-turbid, acid-preserved water samples (i.e. pristine water samples), having turbidity < 1 NTU and no odor. Where turbidity exceeds 1 NTU, and/or the sample is colored and has an odor, results may be biased low compared to true Total Metals concentrations. ALS recommends that turbidity analysis be requested on samples submitted for this test to aid with interpretation of results. | 5e-06 | Y | Y | Y | 19 | 10/20/2015 | 10/24/2018 | 10/24/2018 12:26:14 PM | 2016-03-24 |  | RDGIBSON |  | Concentration | mg/L | Y | (fl. conc.) |  | Yttrium Total (fl. conc.) | 7440-65-5 | Metals - Total; | Metals - Total; EMS code: Y--T | Chemical | Total | LAB
Y--T | ppm (S) | Yttrium Total | 6041 | Aqua Regia by ICPMS | These were sediment/soils digested with Aqua-Regia (Nitric Acid & Hydrochloric Acids) and then analysis by ICPMS. Results would be reported in mg/kg and/or %. | 0.003 | Y | Y | Y | 21 | 10/13/2019 | 8/5/2020 | 8/10/2020 3:52:07 PM | 2019-09-06 | 10/21/2019 11:57:01 AM | KATFRASE | KATFRASE | DimensionlessRatio | ppm | N | (ratio) |  | Yttrium Total (ratio) | 7440-65-5 | Metals - Total; | Metals - Total; EMS code: Y--T | Chemical | Total | LAB
Y--T | mg/L | Yttrium Total | METN | MET-T-NP-CCMS-VA | Ultra trace metals in water are analyzed by CRC ICPMS, based on US EPA Method 6020B (July 2014). The detection limits provided can only be met for undigested samples. This procedure is intended for colorless, non-turbid, acid-preserved water samples (i.e. pristine water samples), having turbidity < 1 NTU and no odor. Where turbidity exceeds 1 NTU, and/or the sample is colored and has an odor, results may be biased low compared to true Total Metals concentrations. ALS recommends that turbidity analysis be requested on samples submitted for this test to aid with interpretation of results. Where turbidity is <1NTU, undigested metals are equivalent to total metals concentrations. | 5e-06 | Y | Y | Y | 280 | 8/6/2019 | 4/4/2024 | 4/29/2024 4:53:06 PM | 2021-06-16 |  | EMS |  | Concentration | mg/L | Y | (fl. conc.) |  | Yttrium Total (fl. conc.) | 7440-65-5 | Metals - Total; | Metals - Total; EMS code: Y--T | Chemical | Total | LAB
Y--T | mg/L | Yttrium Total | X285 | HNO3 Dig:ICP Analysis | HNO3 Dig:ICP Analysis | 0.001 | Y | Y | Y | 431 |  |  | 3/23/2010 5:08:09 PM | 1998-05-22 | 6/3/1998 6:51:49 PM | NPEPPIN |  | Concentration | mg/L | Y | (fl. conc.) |  | Yttrium Total (fl. conc.) | 7440-65-5 | Metals - Total; | Metals - Total; EMS code: Y--T | Chemical | Total | LAB
Y--T | mg/L | Yttrium Total | MSEA | Metals in Seawater (undigested) by (CRC) ICP-MS | This method describes the multi-elemental determination of trace elements by Collision/Reaction Cell (CRC) ICP-MS. The method measures ions produced by a radio-frequency inductively coupled plasma. Analyte species originating in a liquid are pneumatically nebulized and the resulting aerosol transported by argon gas into the plasma torch. The ions produced are entrained in the plasma gas and extracted, by means of a differentially pumped vacuum interface, into a mass spectrometer. The ions flow through a Collision/Reaction Cell (CRC) where depending on the mode of analysis either the ions passes through the cell unhindered (no gas - classical ICP-MS mode), or collisions/reactions chemistry takes place between the ions and a neutral gas such as hydrogen or helium. Any polyatomic ions which give interference on the analytes are removed. The ions are then sorted according to their mass-to-charge ratios by a quadrupole mass spectrometer. The ions are then sorted according to their mass-to-charge ratios by a quadrupole mass spectrometer. The ions transmitted through the quadrupole are quantified by a channel electron multiplier and the ion information processed by a data handling system. Although most polyatomic interference is removed, any additional interference must be assessed and valid corrections applied or the data flagged to indicate problems. Interference correction must include compensation for background ions contributed by the plasma gas, reagents, and constituents of the sample matrix. Instrument drift as well as suppressions or enhancements of instrument response caused by the sample matrix must be corrected for by the use of internal standards. | 0.0005 | Y | Y | Y | 20 | 4/13/2023 | 4/29/2024 | 6/7/2024 10:03:16 AM | 2023-05-26 |  | EMS |  | Concentration | mg/L | Y | (fl. conc.) |  | Yttrium Total (fl. conc.) | 7440-65-5 | Metals - Total; | Metals - Total; EMS code: Y--T | Chemical | Total | LAB
Y--T | mg/L | Yttrium Total | 1349 | Tot Metals in Seawater by CRC ICPMS (BC) | Tot Metals in Seawater by CRC ICPMS (BC) | 0.0005 | Y | Y | Y | 98 | 8/16/2019 | 12/5/2021 | 12/13/2021 12:34:17 PM | 2019-09-30 |  | KATFRASE |  | Concentration | mg/L | Y | (fl. conc.) |  | Yttrium Total (fl. conc.) | 7440-65-5 | Metals - Total; | Metals - Total; EMS code: Y--T | Chemical | Total | LAB
YB-T | ppm (S) | Ytterbium | MS81 | ALS Minerals: ME-MS81 | ALS Minerals: ME-MS81. A prepared sample (0.100 g) is added to lithium metaborate/lithium tetraborate flux; mixed well and fused in a furnace at 1025°C. The resulting melt is then cooled and dissolved in an acid mixture containing nitric; hydrochloric and hydrofluoric acids. This solution is then analyzed by inductively coupled plasma - mass spectrometry. | 0.03 | Y | Y | Y | 0 |  |  |  | 2024-01-29 |  | EMS |  | DimensionlessRatio | ppm | N | (ratio) |  | Ytterbium (ratio) |  | Metals; | Metals; EMS code: YB-T | Chemical |  | LAB
ZN-D | mg/L | Zinc Dissolved | 0050 | Graphite Furnace AA | Graphite Furnace AA | 0.005 | Y | Y | Y | 2289 |  |  | 10/24/1996 12:33:08 PM | 1996-04-12 | 6/3/1998 6:51:41 PM | EMS |  | Concentration | mg/L | Y | (fl. conc.) |  | Zinc Dissolved (fl. conc.) | 7440-66-6 | Metals - Dissolved; | Metals - Dissolved; EMS code: ZN-D | Chemical | Dissolved | LAB
ZN-D | mg/L | Zinc Dissolved | 0090 | Ext.(APDC):AA Anal. | Ext.(APDC):AA Anal. | 0.005 | Y | Y | Y | 271 |  |  | 10/24/1996 1:46:03 PM | 1996-04-12 | 6/3/1998 6:51:41 PM | EMS |  | Concentration | mg/L | Y | (fl. conc.) |  | Zinc Dissolved (fl. conc.) | 7440-66-6 | Metals - Dissolved; | Metals - Dissolved; EMS code: ZN-D | Chemical | Dissolved | LAB
ZN-D | mg/L | Zinc Dissolved | 0020 | HNO3 Dig:AA Analysis | HNO3 Dig:AA Analysis | 0.02 | Y | Y | Y | 6 |  |  | 3/4/1997 7:02:37 PM | 1996-04-12 | 6/3/1998 6:51:41 PM | EMS |  | Concentration | mg/L | Y | (fl. conc.) |  | Zinc Dissolved (fl. conc.) | 7440-66-6 | Metals - Dissolved; | Metals - Dissolved; EMS code: ZN-D | Chemical | Dissolved | LAB
ZN-D | mg/L | Zinc Dissolved | X556 | Metals in water, dissolved, "in-bottle" digestion, ICP-SFMS | The whole water sample is filtered through a 0.45 micron cellulose acetate filter paper and the filtrate is stored in a HDPE/Teflon bottle. The container is de-capped in a Class 100 workstation, acidified to 2% (v/v) HNO3 and 1% (v/v) HCl and recapped. The sample is "in-bottle" digested in a convection oven, set at 60 degrees C for 16 hours. The sample is cooled and introduced to an inductively coupled argon plasma-sector field mass spectrometer (ICP-SFMS) by pneumatic nebulization. Each respective element is measured at a specific mass to charge ratio, m/z value expressed in amu (atomic mass units), and at a specified intrument resolution. LR: Low Resolution ~300, MR: Mid Resolution ~4000 and HR: High Resolution ~10000. Ref: Method number 02-2003, National Laboratory for Environmental Testing, CCIW, Environment Canada, Burlington, Requested April 2003. | 5e-05 | Y | Y | Y | 4233 | 6/11/2003 | 12/6/2023 2:04:00 AM | 2/14/2024 9:05:32 AM | 2004-01-19 |  | UMULLER |  | Concentration | mg/L | Y | (fl. conc.) |  | Zinc Dissolved (fl. conc.) | 7440-66-6 | Metals - Dissolved; | Metals - Dissolved; EMS code: ZN-D | Chemical | Dissolved | LAB
ZN-D | mg/L | Zinc Dissolved | F039 | Field filt + acid, ICP-AES | Dissolved and extractable water samples are ready for analysis once collected. Total samples are autoclaved with HNO3 and hydrochloric acid (HCI) in capped Teflon vials. Analysis is performed via Inductively Coupled Argon Plasma Atomic Emission Spectrometry (ICP-AES)(simultaneous multielement analysis), Graphite Furnace Atomic Absorption Spectrometry (GF-AAS)(low level analysis) and/or Flame Atomic Absorption Spectrometry (F-AAS). | 0.002 | Y | Y | Y | 6631 | 7/29/1996 6:03:00 PM | 7/5/2019 1:21:00 PM | 9/11/2019 8:47:50 AM | 1997-05-08 | 6/3/1998 6:51:47 PM | MPAK |  | Concentration | mg/L | Y | (fl. conc.) |  | Zinc Dissolved (fl. conc.) | 7440-66-6 | Metals - Dissolved; | Metals - Dissolved; EMS code: ZN-D | Chemical | Dissolved | LAB
ZN-D | mg/L | Zinc Dissolved | X203 | FF:FA,HNO3,Dir. Flame,AA | Samples are field-filtered (0.45 um), and field-acidified (HNO3); analysis by AA. | 0.005 | Y | Y | Y | 323 |  |  | 10/24/1996 12:04:45 PM | 1996-04-12 | 6/3/1998 6:51:41 PM | EMS |  | Concentration | mg/L | Y | (fl. conc.) |  | Zinc Dissolved (fl. conc.) | 7440-66-6 | Metals - Dissolved; | Metals - Dissolved; EMS code: ZN-D | Chemical | Dissolved | LAB
ZN-D | mg/L | Zinc Dissolved | X085 | LF,LA,HNO3,AA | Samples are lab-filtered (0.45 um), and lab-acidified (HNO3); analysis by AA. | 0.005 | Y | Y | Y | 1227 | 12/12/1996 | 9/18/2015 | 10/28/2015 9:54:16 AM | 1996-04-12 | 6/3/1998 6:51:41 PM | EMS |  | Concentration | mg/L | Y | (fl. conc.) |  | Zinc Dissolved (fl. conc.) | 7440-66-6 | Metals - Dissolved; | Metals - Dissolved; EMS code: ZN-D | Chemical | Dissolved | LAB
ZN-D | mg/L | Zinc Dissolved | X113 | HNO3/HCl/Perchlor,AA | HNO3/HCl/Perchlor;AA | 0.005 | Y | Y | Y | 2870 |  |  | 11/15/2000 1:53:53 PM | 1996-04-12 | 6/3/1998 6:51:41 PM | EMS |  | Concentration | mg/L | Y | (fl. conc.) |  | Zinc Dissolved (fl. conc.) | 7440-66-6 | Metals - Dissolved; | Metals - Dissolved; EMS code: ZN-D | Chemical | Dissolved | LAB
ZN-D | mg/L | Zinc Dissolved | X095 | FF,FA,HNO3,ICP | FF;FA;HNO3;ICP | 0.005 | Y | Y | Y | 9051 |  |  | 8/26/2024 2:41:32 PM | 1996-04-12 | 6/3/1998 6:51:41 PM | EMS |  | Concentration | mg/L | Y | (fl. conc.) |  | Zinc Dissolved (fl. conc.) | 7440-66-6 | Metals - Dissolved; | Metals - Dissolved; EMS code: ZN-D | Chemical | Dissolved | LAB
ZN-D | mg/L | Zinc Dissolved | X544 | FF,FA,HNO3,ICPMS | Analytes nebulized followed by ion entrained in plasma gas and extracted by means of a differentially pumped vacuum interface into a mass spectrometer | 0.001 | Y | Y | Y | 30 |  |  | 10/30/2014 8:31:19 AM | 2001-09-18 |  | NPEPPIN |  | Concentration | mg/L | Y | (fl. conc.) |  | Zinc Dissolved (fl. conc.) | 7440-66-6 | Metals - Dissolved; | Metals - Dissolved; EMS code: ZN-D | Chemical | Dissolved | LAB
ZN-D | mg/L | Zinc Dissolved | X027 | Ext,APDC,AA | Ext;APDC;AA | 0.005 | Y | Y | Y | 100 |  |  | 10/24/1996 2:40:25 PM | 1996-04-12 | 6/3/1998 6:51:41 PM | EMS |  | Concentration | mg/L | Y | (fl. conc.) |  | Zinc Dissolved (fl. conc.) | 7440-66-6 | Metals - Dissolved; | Metals - Dissolved; EMS code: ZN-D | Chemical | Dissolved | LAB
ZN-D | t/d | Zinc Dissolved | CALC | Calculation (t/d) | Calculation (t/d) | 0 | Y | Y | Y | 5483 |  |  | 7/3/2024 12:59:26 PM | 1996-04-12 | 6/3/1998 6:51:41 PM | EMS |  | MassFlow | t/d | Y | (mass flow) |  | Zinc Dissolved (mass flow) | 7440-66-6 | Metals - Dissolved; | Metals - Dissolved; EMS code: ZN-D | Chemical | Dissolved | LAB
ZN-D | mg/L | Zinc Dissolved | 5138 | AA Analysis | AA Analysis | 0.02 | Y | Y | Y | 17 |  |  | 10/24/1996 9:37:08 AM | 1996-04-12 | 6/3/1998 6:51:41 PM | EMS |  | Concentration | mg/L | Y | (fl. conc.) |  | Zinc Dissolved (fl. conc.) | 7440-66-6 | Metals - Dissolved; | Metals - Dissolved; EMS code: ZN-D | Chemical | Dissolved | LAB
ZN-D | mg/L | Zinc Dissolved | EQ01 | EQUIS (mg/L) | EQUIS (mg/L) | 0.005 | N | Y | Y | 7589 |  |  | 10/24/1996 3:16:55 PM | 1996-04-12 | 6/3/1998 6:51:41 PM | EMS |  | Concentration | mg/L | Y | (fl. conc.) |  | Zinc Dissolved (fl. conc.) | 7440-66-6 | Metals - Dissolved; | Metals - Dissolved; EMS code: ZN-D | Chemical | Dissolved | LAB
ZN-D | mg/L | Zinc Dissolved | X109 | LF, AA | LF; AA | 0.01 | Y | Y | Y | 28 |  |  | 10/24/1996 11:59:43 AM | 1996-04-12 | 6/3/1998 6:51:41 PM | EMS |  | Concentration | mg/L | Y | (fl. conc.) |  | Zinc Dissolved (fl. conc.) | 7440-66-6 | Metals - Dissolved; | Metals - Dissolved; EMS code: ZN-D | Chemical | Dissolved | LAB
ZN-D | mg/L | Zinc Dissolved | X099 | FF,FA,HNO3,AA | FF;FA;HNO3;AA | 0.005 | Y | Y | Y | 1126 |  |  | 10/24/1996 12:45:23 PM | 1996-04-12 | 6/3/1998 6:51:41 PM | EMS |  | Concentration | mg/L | Y | (fl. conc.) |  | Zinc Dissolved (fl. conc.) | 7440-66-6 | Metals - Dissolved; | Metals - Dissolved; EMS code: ZN-D | Chemical | Dissolved | LAB
ZN-D | mg/L | Zinc Dissolved | X562 | Dissolved metals: FF, FA, HNO3, HR-ICPMS | Ultra trace metals in water are analyzed by high resolution inductively coupled plasma mass spectrometry (HR-ICPMS), based on US EPA Method 200.8, (Oct. 1999). This procedure is intended for pristine field-filtered acid-preserved water samples. The detection limits (LOR) for this test are based on lab instrumental analysis only, not including filtration. ALS-supplied field filtration equipment does not support these LOR. Therefore, because of the high probability of false positives due to filtration, it is strongly recommended that a filtration blank be analyzed to aid in data interpretation. | 0.0001 | Y | Y | Y | 1264 | 6/8/2015 | 10/19/2023 | 11/9/2023 7:41:27 AM | 2015-10-22 |  | RDGIBSON |  | Concentration | mg/L | Y | (fl. conc.) |  | Zinc Dissolved (fl. conc.) | 7440-66-6 | Metals - Dissolved; | Metals - Dissolved; EMS code: ZN-D | Chemical | Dissolved | LAB
ZN-D | mg/L | Zinc Dissolved | X170 | ICP | This code is a catch-all for a large variety of ICP methods including water and also solids digested into solutions as an intermediate result. Different procedures will have greatly different detection limits. (a) For air particulates, preliminary results are reported in mg/L and loading results as mg/dm2/d. (b) For Hi Vol samples preliminiary results are same as reported in mg/L and loading results as ug/m3. (c) Note that for Rain samples there is no filtration nor digestion steps. | 0.005 | Y | Y | Y | 15767 | 4/11/1996 6:02:00 PM | 11/6/2018 | 4/12/2024 1:44:31 PM | 1996-04-12 | 8/15/2019 1:24:32 PM | EMS | KATFRASE | Concentration | mg/L | Y | (fl. conc.) |  | Zinc Dissolved (fl. conc.) | 7440-66-6 | Metals - Dissolved; | Metals - Dissolved; EMS code: ZN-D | Chemical | Dissolved | LAB
ZN-D | mg/L | Zinc Dissolved | F023 | HNO3,GFAA | Dissolved and extractable water samples are ready for analysis once collected. Total samples are autoclaved with HNO3 and hydrochloric acid (HCl) in capped Teflon vials. Analysis is performed via Inductively Coupled Argon Plasma Atomic Emission Spectrometry (ICP-AES) (simultaneous multielement analysis), Graphite Furnace Atomic Absorption Spectrometry (GF-AAS) (low level analysis) and/or Flame Atomic Absorption Spectrometry (F-AAS). Note that extractable samples are acidified with 0.5 mL HNO3 per 100 mL of sample at time of sample collection. | 0.001 | Y | Y | Y | 78 |  |  | 2/19/2002 10:02:58 AM | 2001-02-12 |  | NPEPPIN |  | Concentration | mg/L | Y | (fl. conc.) |  | Zinc Dissolved (fl. conc.) | 7440-66-6 | Metals - Dissolved; | Metals - Dissolved; EMS code: ZN-D | Chemical | Dissolved | LAB
ZN-D | mg/L | Zinc Dissolved | X356 | LF,LA:HNO3:ICP | Samples are lab-filtered (0.45 um); lab-acidified with HNO3; analysis by ICP/AES. | 0.001 | Y | Y | Y | 125 | 10/17/2000 | 5/1/2002 | 1/20/2004 12:07:44 PM | 2001-02-12 |  | NPEPPIN |  | Concentration | mg/L | Y | (fl. conc.) |  | Zinc Dissolved (fl. conc.) | 7440-66-6 | Metals - Dissolved; | Metals - Dissolved; EMS code: ZN-D | Chemical | Dissolved | LAB
ZN-D | mg/L | Zinc Dissolved | F068 | HNO3dig,TeflonTube/Autoclave,ICP-MS | This method is applicable to determiniation of dissolved, extractable, and total recoverable metals in ground waters, surface waters, drinking waters, and wastewaters. Dissolved, extractable, and total recoverable sample fractions are acidified to pH <1.5 with analytical grade nitric acid (HNO3) to prevent the metals from adsorbing to the container material or precipitating. Extractable and total recoverable water samples are collected in 250 ml acid washed polyethylene bottles and are preserved with Nitric acid (0.5 ml acid for 100 mLs of sample). All preserved samples are stored at room temperature and can be held for up to 6 months. Unfiltered, un-preserved samples are stored at 4C for no more than 30 days. For dissolved and extractable water fractions, no sample preparation is necessary. Samples are simply poured in autosampler tubes (i.e. 15 ml centrifuge tubes), and placed on the autosampler and analyzed. Samples for total recoverable metals are digested with Nitric acid in special Teflon tubes in an autoclave prior to analyses. | 0.0001 | Y | Y | Y | 2132 | 6/1/1999 1:22:00 PM | 4/17/2015 1:06:00 PM | 7/18/2024 8:07:20 AM | 1999-06-28 |  | NPEPPIN |  | Concentration | mg/L | Y | (fl. conc.) |  | Zinc Dissolved (fl. conc.) | 7440-66-6 | Metals - Dissolved; | Metals - Dissolved; EMS code: ZN-D | Chemical | Dissolved | LAB
ZN-D | mg/L | Zinc Dissolved | X168 | Solvent Extract: AA | Solvent Extract: AA | 0.001 | Y | Y | Y | 1 |  |  | 10/24/1996 12:09:42 PM | 1996-04-12 | 6/3/1998 6:51:41 PM | EMS |  | Concentration | mg/L | Y | (fl. conc.) |  | Zinc Dissolved (fl. conc.) | 7440-66-6 | Metals - Dissolved; | Metals - Dissolved; EMS code: ZN-D | Chemical | Dissolved | LAB
ZN-D | mg/L | Zinc Dissolved | MSEA | Metals in Seawater (undigested) by (CRC) ICP-MS | This method describes the multi-elemental determination of trace elements by Collision/Reaction Cell (CRC) ICP-MS. The method measures ions produced by a radio-frequency inductively coupled plasma. Analyte species originating in a liquid are pneumatically nebulized and the resulting aerosol transported by argon gas into the plasma torch. The ions produced are entrained in the plasma gas and extracted, by means of a differentially pumped vacuum interface, into a mass spectrometer. The ions flow through a Collision/Reaction Cell (CRC) where depending on the mode of analysis either the ions passes through the cell unhindered (no gas - classical ICP-MS mode), or collisions/reactions chemistry takes place between the ions and a neutral gas such as hydrogen or helium. Any polyatomic ions which give interference on the analytes are removed. The ions are then sorted according to their mass-to-charge ratios by a quadrupole mass spectrometer. The ions are then sorted according to their mass-to-charge ratios by a quadrupole mass spectrometer. The ions transmitted through the quadrupole are quantified by a channel electron multiplier and the ion information processed by a data handling system. Although most polyatomic interference is removed, any additional interference must be assessed and valid corrections applied or the data flagged to indicate problems. Interference correction must include compensation for background ions contributed by the plasma gas, reagents, and constituents of the sample matrix. Instrument drift as well as suppressions or enhancements of instrument response caused by the sample matrix must be corrected for by the use of internal standards. | 0.0005 | Y | Y | Y | 4 | 4/12/2023 | 4/29/2024 | 5/16/2024 2:00:14 PM | 2009-09-08 |  | RDGIBSON |  | Concentration | mg/L | Y | (fl. conc.) |  | Zinc Dissolved (fl. conc.) | 7440-66-6 | Metals - Dissolved; | Metals - Dissolved; EMS code: ZN-D | Chemical | Dissolved | LAB
ZN-D | mg/L | Zinc Dissolved | F032 | Filt.,Acid,ICPMS | There is no acid digestion required for this analysis. A water sample is collected and preserved in 0.2 percent nitric acid. To analyse the extractable metals the sample is shaken, allowed to sit for at least twenty four hours, and a decanted portion is introduced to the spectrometer. When the dissolved metals are required, the samples must first be filtered through a 0.45 micron cellulose acetate filter paper and then followed by preservation in 0.2 percent nitric acid. It is important to note that different laboratories may have very different detection limits for some variables analyzed by ICPMS. | 4e-05 | Y | Y | Y | 83037 | 8/16/1996 2:35:00 PM | 8/26/2024 | 8/26/2024 4:53:05 PM | 1997-01-16 | 4/29/2011 8:52:00 AM | MPAK | RDGIBSON | Concentration | mg/L | Y | (fl. conc.) |  | Zinc Dissolved (fl. conc.) | 7440-66-6 | Metals - Dissolved; | Metals - Dissolved; EMS code: ZN-D | Chemical | Dissolved | LAB
ZN-D | mg/L | Zinc Dissolved | METB | MET-D-NP-BCMOE-MS-VA | Ultra trace metals in water are analyzed by CRC ICPMS, based on US EPA Method 6020B (July 2014). This procedure is intended for pristine field-filtered acid-preserved water samples. The detection limits (LOR) for this test are based on lab instrumental analysis only, not including filtration. ALS-supplied field filtration equipment does not support these LOR. Therefore, because of the high probability of false positives due to filtration, it is strongly recommended that a filtration blank be analysed to aid in data interpretation. BC Ministry of Environment version. | 0.0002 | Y | Y | Y | 596 | 6/11/2019 | 8/15/2024 | 8/23/2024 9:50:35 AM | 2019-08-20 |  | KATFRASE |  | Concentration | mg/L | Y | (fl. conc.) |  | Zinc Dissolved (fl. conc.) | 7440-66-6 | Metals - Dissolved; | Metals - Dissolved; EMS code: ZN-D | Chemical | Dissolved | LAB
ZN-D | mg/L | Zinc Dissolved | X028 | FA,HNO3,AA Analysis | FA;HNO3;AA Analysis | 0.005 | Y | Y | Y | 537 |  |  | 10/29/2001 11:34:44 AM | 1996-04-12 | 6/3/1998 6:51:41 PM | EMS |  | Concentration | mg/L | Y | (fl. conc.) |  | Zinc Dissolved (fl. conc.) | 7440-66-6 | Metals - Dissolved; | Metals - Dissolved; EMS code: ZN-D | Chemical | Dissolved | LAB
ZN-D | mg/L | Zinc Dissolved | METC | MET-D-NP-CCMS-VA | Ultra trace metals in water are analyzed by CRC ICPMS, based on US EPA Method 6020B (July 2014). This procedure is intended for pristine field-filtered acid-preserved water samples. The detection limits (LOR) for this test are based on lab instrumental analysis only, not including filtration. ALS-supplied field filtration equipment does not support these LOR. Therefore, because of the high probability of false positives due to filtration, it is strongly recommended that a filtration blank be analysed to aid in data interpretation. | 0.0005 | Y | Y | Y | 1179 | 4/3/2012 | 4/9/2024 | 4/29/2024 4:53:06 PM | 2021-04-14 |  | RDGIBSON |  | Concentration | mg/L | Y | (fl. conc.) |  | Zinc Dissolved (fl. conc.) | 7440-66-6 | Metals - Dissolved; | Metals - Dissolved; EMS code: ZN-D | Chemical | Dissolved | LAB
ZN-D | mg/L | Zinc Dissolved | MWID | Metals in Water by ICPMS (digested) | Metals in Water by ICPMS (digested) Based on EPA 200.8 | 0.01 | Y | Y | Y | 31873 | 10/26/2018 5:20:00 PM | 10/26/2018 5:20:00 PM | 7/31/2024 1:46:41 PM | 2009-12-24 |  | RDGIBSON |  | Concentration | mg/L | Y | (fl. conc.) |  | Zinc Dissolved (fl. conc.) | 7440-66-6 | Metals - Dissolved; | Metals - Dissolved; EMS code: ZN-D | Chemical | Dissolved | LAB
ZN-D | mg/L | Zinc Dissolved | X072 | FA,HNO3,Graph.Furnace,AA | Samples are field-acidified with HNO3; HNO3 digestion; analysis by GFAA. | 0.005 | Y | Y | Y | 109 |  |  | 10/24/1996 12:10:43 PM | 1996-04-12 | 6/3/1998 6:51:41 PM | EMS |  | Concentration | mg/L | Y | (fl. conc.) |  | Zinc Dissolved (fl. conc.) | 7440-66-6 | Metals - Dissolved; | Metals - Dissolved; EMS code: ZN-D | Chemical | Dissolved | LAB
ZN-D | mg/L | Zinc Dissolved | F049 | Filt,HNO3 dig,conc'n,nebul'ICP-OES | This method provides an automated procedure for the Inductively Coupled Argon Plasma Optical Emission Spectrophotometric (ICP-OES) determination and quantitation of trace amounts of metals in surface, ground and precipitation waters after a manual acid digestion and concentration procedure and ultrasonic nebulisation. Organically-bound metals and metal complexes are converted to metal ions by digestion with nitric acid. To increase the signal detected by the ICP, the samples are concentrated 10 times. The liquid sample is pumped onto the piezoelectric transducer, where the ultrasonic action efficiently converts it to a fine, dense aerosol. The nebulizer carrier gas flow then transports the wet aerosol through a heated U-tube, where the solvent is vaporized. Subsequently, most of the solvent vapour is removed from the aerosol stream by passing the stream through a chilled condenser. The result is a dry, analyte-laden aerosol which is introduced to the plasma. There, the high temperature desolvates the aerosol and dissociates the sample material into individual atoms and ions, which are excited to emit light at wave lengths characteristic of the elements present. Dissolved samples are field filtered through 0.45 micron cellulose acetate filter prior to acid preservation. Samples for extractable analysis are acidified with 2 mL concentrated HNO3 per litre and allowed to stand a minimum of 24 hours before analysis. Clients providing these samples are advised to preserve samples at the collection site. | 0.0002 | Y | Y | Y | 17500 |  |  | 5/19/2015 3:10:36 PM | 1997-10-20 | 6/3/1998 6:51:47 PM | NPEPPIN |  | Concentration | mg/L | Y | (fl. conc.) |  | Zinc Dissolved (fl. conc.) | 7440-66-6 | Metals - Dissolved; | Metals - Dissolved; EMS code: ZN-D | Chemical | Dissolved | LAB
ZN-D | mg/L | Zinc Dissolved | X116 | FF,FA,HNO3,Graph.Furn,AA | Samples are field-filtered (0.45 um), field-acidified with HNO3; analysis by GFAA. | 0.005 | Y | Y | Y | 229 |  |  | 11/17/1998 12:06:58 PM | 1996-04-12 | 6/3/1998 6:51:41 PM | EMS |  | Concentration | mg/L | Y | (fl. conc.) |  | Zinc Dissolved (fl. conc.) | 7440-66-6 | Metals - Dissolved; | Metals - Dissolved; EMS code: ZN-D | Chemical | Dissolved | LAB
ZN-D | mg/L | Zinc Dissolved | X285 | HNO3 Dig:ICP Analysis | HNO3 Dig:ICP Analysis | 0.001 | Y | Y | Y | 2163 | 8/26/1993 | 9/16/2003 3:18:00 PM | 10/15/2018 12:58:52 PM | 1996-04-12 | 6/3/1998 6:51:41 PM | EMS |  | Concentration | mg/L | Y | (fl. conc.) |  | Zinc Dissolved (fl. conc.) | 7440-66-6 | Metals - Dissolved; | Metals - Dissolved; EMS code: ZN-D | Chemical | Dissolved | LAB
ZN-E | mg/L | Zinc-Extractable | F047 | 24hr HNO3 dil'n,ICP-MS | There is no acid digestion required for this analysis. A water sample is collected and preserved in 0.2 percent nitric acid. To analyse the extractable metals the sample is shaken, allowed to sit for at least twenty four hours, and a decanted portion is introduced to the spectrometer. When the dissolved metals are required, the samples must first be filtered through a 0.45 micron cellulose acetate filter paper and then followed by preservation in 0.2 percent nitric acid. It is important to note that different laboratories may have very different detection limits for some variables analyzed by ICP/MS. | 4e-05 | Y | Y | Y | 2676 | 4/16/1998 10:21:00 AM | 6/26/2003 | 9/11/2019 8:46:18 AM | 1997-10-20 | 6/3/1998 6:51:47 PM | NPEPPIN |  | Concentration | mg/L | Y | (fl. conc.) |  | Zinc-Extractable (fl. conc.) | 7440-66-6 | Metals - Extractable; | Metals - Extractable; EMS code: ZN-E | Chemical | Extractable | LAB
ZN-E | mg/L | Zinc-Extractable | X553 | Metals in water, extractable, ICP-SFMS | The whole water sample is collected in a HDPE/Teflon bottle. In a Class 100 workstation it is preserved with concentrated Nitric Acid at concentration of 0.2 percent (v/v). The acid preservative extracts the metals from material in the sample and a decanted portion is introduced directly to an inductively coupled argon plasma-sector field mass spectrometer (ICP-SFMS) by pneumatic nebulization. Each respective element is measured at a specific mass to charge ratio, m/z value expressed in amu (atomic mass units), and at a specified instrument resolution. LR: Low Resolution ~300, MR: Mid Resolution ~4000 and HR: High Resolution ~10000. Ref: Method number 02-2061, National Laboratory for Environmental Testing, CCIW, Environment Canada, Burlington. Requested April 2003. | 5e-05 | Y | Y | Y | 2597 | 5/22/2003 | 9/21/2007 | 11/28/2023 2:11:53 PM | 2004-01-19 |  | UMULLER |  | Concentration | mg/L | Y | (fl. conc.) |  | Zinc-Extractable (fl. conc.) | 7440-66-6 | Metals - Extractable; | Metals - Extractable; EMS code: ZN-E | Chemical | Extractable | LAB
ZN-E | mg/L | Zinc-Extractable | F069 | SEM extr,6N HCl,ICP/OES | 6.0 N Hydrochloric acid (HCI) is added to an aliquot of the sediment sample. The extract produced from the addition of the acid is then analyzed for simultaneously extracted metals (SEM) using inductively coupled plasma - optical emission spectrophotometry. The detection limits depend on the metal parameters in question and the instrumentation used. | 0.005 | Y | Y | Y | 10 | 3/26/2012 3:01:00 AM | 3/26/2012 3:01:00 AM | 11/25/2021 1:50:15 PM | 1999-07-09 |  | NPEPPIN |  | Concentration | mg/L | Y | (fl. conc.) |  | Zinc-Extractable (fl. conc.) | 7440-66-6 | Metals - Extractable; | Metals - Extractable; EMS code: ZN-E | Chemical | Extractable | LAB
ZN-E | ug/g | Zinc-Extractable | X230 | Hom,HNO3/HCl,ICP | The sample is dried, ground to pass 100 mesh, and digested with a nitric-hydrochloric acid mixture to solubilize the solid matter and to remove the organic material by oxidation and volatilization. The completed digests are submitted to metals instrumentation personnel for metals analysis by Inductively Coupled Atomic Emission Spectrophotometry (ICP-AES), or hydride analysis by atomic absorption. | 0.25 | Y | Y | Y | 74 | 12/13/2011 12:02:00 PM | 12/13/2011 12:02:00 PM | 2/9/2012 12:15:57 PM | 1996-04-12 | 6/3/1998 6:51:41 PM | EMS |  | MassPerMass | mg/kg | Y | (mass conc.) |  | Zinc-Extractable (mass conc.) | 7440-66-6 | Metals - Extractable; | Metals - Extractable; EMS code: ZN-E | Chemical | Extractable | LAB
ZN-E | mg/L | Zinc-Extractable | F040 | Field Acid, lab filt, ICP-AES | Dissolved and extractable water samples are ready for analysis once collected. Total samples are autoclaved with HNO3 and hydrochloric acid (HCI) in capped Teflon vials. Analysis is performed via Inductively Coupled Argon Plasma Atomic Emission Spectrometry (ICP-AES)(simultaneous multielement analysis), Graphite Furnace Atomic Absorption Spectrometry (GF-AAS)(low level analysis) and/or Flame Atomic Absorption Spectrometry (F-AAS). | 0.002 | Y | Y | Y | 2361 | 8/5/1997 6:06:00 PM | 3/14/2008 | 9/11/2019 8:47:41 AM | 1997-05-08 | 6/3/1998 6:51:47 PM | MPAK |  | Concentration | mg/L | Y | (fl. conc.) |  | Zinc-Extractable (fl. conc.) | 7440-66-6 | Metals - Extractable; | Metals - Extractable; EMS code: ZN-E | Chemical | Extractable | LAB
ZN-E | umol/g | Zinc-Extractable | FSEM | AVS extraction, SEM Analysis ICPMS | Determination of simultaneous extractable metals (SEM) in sediment after sample extracted with hydrochloric acid for acid volatile sulfide (AVS) | 5 | Y | Y | Y | 19 | 3/12/2020 | 3/24/2020 | 2/22/2023 12:51:04 PM | 2010-11-16 |  | RDGIBSON |  | ParticlesPerMass | µmol/g | Y | (part. /m) |  | Zinc-Extractable (part. /m) | 7440-66-6 | Metals - Extractable; | Metals - Extractable; EMS code: ZN-E | Chemical | Extractable | LAB
ZN-E | mg/L | Zinc-Extractable | X512 | HNO3,pHto7.5,MIK extr'n,AA | HNO3 Digestion;pH to 7.5; MIK solvent extraction;AA | 0.001 | Y | Y | Y | 875 |  |  | 8/23/1999 3:43:46 PM | 1999-04-22 |  | NPEPPIN |  | Concentration | mg/L | Y | (fl. conc.) |  | Zinc-Extractable (fl. conc.) | 7440-66-6 | Metals - Extractable; | Metals - Extractable; EMS code: ZN-E | Chemical | Extractable | LAB
ZN-I | mg/dm2/d | Zinc-Insoluble | EQ15 | EQUIS (mg/dm2/d) | EQUIS (mg/dm2/d) | 0.0001 | N | Y | Y | 168 |  |  | 10/22/1996 10:18:48 PM | 1996-04-12 | 6/3/1998 6:51:41 PM | EMS |  | MassFlux | mg/dm²/d | Y | (mass flux) |  | Zinc-Insoluble (mass flux) | 7440-66-6 | Metals - Insoluble; | Metals - Insoluble; EMS code: ZN-I | Chemical | Insoluble | LAB
ZN-I | mg/dm2/d | Zinc-Insoluble | 5039 | N/P Dig, ICP | Aliquot of prepared sample is digested first with nitric acid then with perchloric acid. Analysis by ICP. (a) For air particulates, preliminary results are reported in mg/L and loading results as mg/dm2/d. (b) For Hi Vol samples preliminary results are same as reported in mg/L and loading results as ug/m3. | 0.001 | Y | Y | Y | 204 | 12/5/1997 | 5/15/2000 | 11/18/2000 1:30:14 AM | 1996-04-12 | 6/3/1998 6:51:41 PM | EMS |  | MassFlux | mg/dm²/d | Y | (mass flux) |  | Zinc-Insoluble (mass flux) | 7440-66-6 | Metals - Insoluble; | Metals - Insoluble; EMS code: ZN-I | Chemical | Insoluble | LAB
ZN-S | mg/dm2/d | Zinc-Soluble | EQ15 | EQUIS (mg/dm2/d) | EQUIS (mg/dm2/d) | 0.0005 | N | Y | Y | 174 |  |  | 10/22/1996 10:18:48 PM | 1996-04-12 | 6/3/1998 6:51:41 PM | EMS |  | MassFlux | mg/dm²/d | Y | (mass flux) |  | Zinc-Soluble (mass flux) | 7440-66-6 | Metals - Soluble; | Metals - Soluble; EMS code: ZN-S | Chemical | Soluble | LAB
ZN-S | mg/dm2/d | Zinc-Soluble | 5045 | ICP Analysis | ICP a) For Air particulates analyzed without an acid digestion step, preliminary results are reported in mg/L and loading results as mg/dm2/d. b) For HiVol samples, preliminary results are reported as mg/L and loading results as ug/m3. | 0.001 | Y | Y | Y | 39 | 11/28/1997 | 5/15/2000 | 11/18/2000 1:30:15 AM | 1998-10-28 |  | NPEPPIN |  | MassFlux | mg/dm²/d | Y | (mass flux) |  | Zinc-Soluble (mass flux) | 7440-66-6 | Metals - Soluble; | Metals - Soluble; EMS code: ZN-S | Chemical | Soluble | LAB
ZN-S | umol/g | Zinc-Soluble | FSEM | AVS extraction, SEM Analysis ICPMS | Determination of simultaneous extractable metals (SEM) in sediment after sample extracted with hydrochloric acid for acid volatile sulfide (AVS) | 5 | Y | Y | Y | 0 |  |  |  | 2024-02-02 |  | EMS |  | ParticlesPerMass | µmol/g | Y | (part. /m) |  | Zinc-Soluble (part. /m) | 7440-66-6 | Metals - Soluble; | Metals - Soluble; EMS code: ZN-S | Chemical | Soluble | LAB
ZN-S | mg/dm2/d | Zinc-Soluble | X170 | ICP | This code is a catch-all for a large variety of ICP methods including water and also solids digested into solutions as an intermediate result. Different procedures will have greatly different detection limits. (a) For air particulates, preliminary results are reported in mg/L and loading results as mg/dm2/d. (b) For Hi Vol samples preliminiary results are same as reported in mg/L and loading results as ug/m3. (c) Note that for Rain samples there is no filtration nor digestion steps. | 0.001 | Y | Y | Y | 179 |  |  | 2/13/1998 2:28:23 AM | 1996-04-12 | 6/3/1998 6:51:41 PM | EMS |  | MassFlux | mg/dm²/d | Y | (mass flux) |  | Zinc-Soluble (mass flux) | 7440-66-6 | Metals - Soluble; | Metals - Soluble; EMS code: ZN-S | Chemical | Soluble | LAB
ZN-T | ug/g | Zinc Total | FB04 | Freeze Dry,Hom,Acid Dig,ICP | Sample moisture is determined by weighing samples before and after freeze-drying. Freeze-dried samples are ground in a Braun blender before being weighed into Teflon PFA liners and digested with concentrated HNO3, concentrated HCl, and 50% H2O2 in a microwave oven; which results in a complete sample dissolution. Samples are collected and volumized with Type I water. Analysis is performed via Inductively Coupled Argon Plasma Atomic Emission Spectrometry (ICP-AES) (simultaneous multielement analysis), Graphite Furnace Atomic Absorption Spectrometry (GF-AAS) (low level Cd and Pb analysis), Cold Vapor Atomic Fluorescence Spectrometry (CV-AFS) (Mercury analysis) and/or Flame Atomic Absorption Spectrometry (F-AAS) (K analysis). | 0.2 | Y | Y | Y | 241 | 8/22/1996 6:11:00 PM | 2/1/2002 2:30:00 PM | 6/23/2003 1:15:35 PM | 1997-02-27 | 6/3/1998 6:51:46 PM | MPAK |  | MassPerMass | mg/kg | Y | (mass conc.) |  | Zinc Total (mass conc.) | 7440-66-6 | Metals - Total; | Metals - Total; EMS code: ZN-T | Chemical | Total | LAB
ZN-T | ug/g | Zinc Total | 0190 | D&G:N/P Dig:ICP Anal. | Sample is dried and grind-mixed before nitric perchloric digestion followed by ICP analysis. | 1 | Y | Y | Y | 945 |  |  | 10/24/1996 3:06:16 PM | 1996-04-12 | 6/3/1998 6:51:41 PM | EMS |  | MassPerMass | mg/kg | Y | (mass conc.) |  | Zinc Total (mass conc.) | 7440-66-6 | Metals - Total; | Metals - Total; EMS code: ZN-T | Chemical | Total | LAB
ZN-T | mg/L | Zinc Total | F082 | Aliquot,HNO3/HCL dig,ICPMS | Aliquot microwave digested with 3:1 HNO3/HCL. Analysis by ICPMS. For HiVol samples preliminary result are reported in mg/L and loading results are reported in ug/m3 |  | Y | Y | Y | 28601 | 7/26/2005 7:04:00 AM | 8/23/2024 | 8/26/2024 4:53:05 PM | 2000-06-13 |  | NPEPPIN |  | Concentration | mg/L | Y | (fl. conc.) |  | Zinc Total (fl. conc.) | 7440-66-6 | Metals - Total; | Metals - Total; EMS code: ZN-T | Chemical | Total | LAB
ZN-T | mg/L | Zinc Total | X561 | Total metals: FA, HNO3, HR-ICPMS (No Prep)Prep) | Ultra trace metals in water are analyzed by high resolution inductively coupled plasma mass spectrometry (HR-ICPMS), modified from US EPA 200.8, (Revision 5.5). The detection limits provided can only be met for undigested samples. This procedure is intended for colorless, non-turbid, acid-preserved water samples (i.e. pristine water samples), having turbidity < 1 NTU and no odor. Where turbidity exceeds 1 NTU, and/or the sample is colored and has an odor, results may be biased low compared to true Total Metals concentrations. ALS recommends that turbidity analysis be requested on samples submitted for this test to aid with interpretation of results. | 0.0001 | Y | Y | Y | 2756 | 6/8/2015 | 10/5/2023 | 1/31/2024 2:37:41 PM | 2015-10-22 |  | RDGIBSON |  | Concentration | mg/L | Y | (fl. conc.) |  | Zinc Total (fl. conc.) | 7440-66-6 | Metals - Total; | Metals - Total; EMS code: ZN-T | Chemical | Total | LAB
ZN-T | ug/g | Zinc Total | X492 | Homo,HCl-HNO3 dig,ICP-OES,dry calc | Homogenize samples and take sub-samples. For Total Recoverable Metals Digest with Aqua Regia. Analyses by ICPOES (Inductively Coupled Plasma Optical Emission Spectroscopy). Perform moisture analysis on separate sub-sample and calculate dry value for trace elements. | 1 | Y | Y | Y | 1 |  |  | 8/24/2016 2:59:13 PM | 1998-11-13 |  | NPEPPIN |  | MassPerMass | mg/kg | Y | (mass conc.) |  | Zinc Total (mass conc.) | 7440-66-6 | Metals - Total; | Metals - Total; EMS code: ZN-T | Chemical | Total | LAB
ZN-T | ug/m3 | Zinc Total | X136 | HNO3/HCl/Perchlor,AA | HNO3/HCl/Perchlor;AA | 0 | Y | Y | Y | 412 |  |  | 2/19/2002 10:06:15 AM | 1996-04-12 | 6/3/1998 6:51:41 PM | EMS |  | AirConcentration | µg/m³ | Y | (air conc.) |  | Zinc Total (air conc.) | 7440-66-6 | Metals - Total; | Metals - Total; EMS code: ZN-T | Chemical | Total | LAB
ZN-T | ug/g | Zinc Total | FS89 | Homo,HNO3 dig,ICP-AES | Homo;HNO3 dig;ICP-AES | 0.3 | Y | Y | Y | 11 | 1/5/2001 | 10/31/2022 | 11/22/2022 2:40:39 PM | 2001-03-08 |  | NPEPPIN |  | MassPerMass | mg/kg | Y | (mass conc.) |  | Zinc Total (mass conc.) | 7440-66-6 | Metals - Total; | Metals - Total; EMS code: ZN-T | Chemical | Total | LAB
ZN-T | ug/g wet | Zinc Total | X560 | Metals in Biota (wet), total, acid digestion, ICP-MS | Five grams of wet, homogenized biota is weighed in a centrifuge tube, and 2.5 mL of Nitric Acid and 2.5 mL of Hydrochloric acid are added. The sample is digested at room temperature for one hour and then digested at 90-95 degrees C in a water bath or heat block for one more hour. The sample is cooled and the volume is made to 50 mL with DI. The digestate is then diluted at ten fold and introduced to a collision/reaction cell inductively coupled argon plasma mass spectrometer (CRC ICPMS) by pneumatic nebulization. Each respective element is measured at a specific mass to charge ratio, m/z value expressed in amu (atomic mass units). - Maxxam Analytics (BC). Added to EMS Dec 10, 2008. | 0.1 | Y | Y | Y | 541 | 11/15/2008 4:00:00 AM | 10/30/2023 | 11/28/2023 2:21:33 PM | 2008-12-11 | 3/11/2015 4:17:10 PM | RDGIBSON | EMS | MassPerMass | µg/g | Y | (mass conc.) | (wet) | Zinc Total (wet) (mass conc.) | 7440-66-6 | Metals - Total; | Metals - Total; EMS code: ZN-T | Chemical | Total | LAB
ZN-T | ug/g | Zinc Total | E496 | CSSS (1993) 11 | CSSS (1993) 11: Plant available micro nutrients (Cu; Fe; Mn; Zn) are determined by extraction using 0.005 M DTPA followed by analysis using ICPOES. | 0.2 | Y | Y | Y | 0 |  |  |  | 2024-02-27 |  | RDGIBSON |  | MassPerMass | mg/kg | Y | (mass conc.) |  | Zinc Total (mass conc.) | 7440-66-6 | Metals - Total; | Metals - Total; EMS code: ZN-T | Chemical | Total | LAB
ZN-T | mg/L | Zinc Total | X072 | FA,HNO3,Graph.Furnace,AA | Samples are field-acidified with HNO3; HNO3 digestion; analysis by GFAA. | 0.005 | Y | Y | Y | 439 |  |  | 10/29/2001 11:34:44 AM | 1996-04-12 | 6/3/1998 6:51:41 PM | EMS |  | Concentration | mg/L | Y | (fl. conc.) |  | Zinc Total (fl. conc.) | 7440-66-6 | Metals - Total; | Metals - Total; EMS code: ZN-T | Chemical | Total | LAB
ZN-T | mg/L | Zinc Total | 0200 | 20%A.R.Dig:AA Anal. | 20%A.R.Dig:AA Anal. | 0.02 | Y | Y | Y | 10 |  |  | 10/24/1996 4:16:44 AM | 1996-04-12 | 6/3/1998 6:51:41 PM | EMS |  | Concentration | mg/L | Y | (fl. conc.) |  | Zinc Total (fl. conc.) | 7440-66-6 | Metals - Total; | Metals - Total; EMS code: ZN-T | Chemical | Total | LAB
ZN-T | ug/g | Zinc Total | 0031 | ICP | ICP | 1 | Y | Y | Y | 12 | 5/6/1997 4:39:00 PM | 5/7/2024 | 7/11/2024 10:43:43 AM | 1996-04-12 | 6/3/1998 6:51:42 PM | EMS |  | MassPerMass | mg/kg | Y | (mass conc.) |  | Zinc Total (mass conc.) | 7440-66-6 | Metals - Total; | Metals - Total; EMS code: ZN-T | Chemical | Total | LAB
ZN-T | mg/L | Zinc Total | X555 | Metals in water, total, "in-bottle" digestion, ICP-SFMS | The whole water sample collected in a HDPE/Teflon bottle, is de-capped in a Class 100 workstation, acidified to 2% (v/v) HNO3 and 1% (v/v) HCl and recapped. The sample is "in-bottle" digested in a convection oven, set at 60 degrees C for 16 hours. The sample is cooled and introduced to an inductively coupled argon plasma-sector field mass spectrometer (ICP-SFMS) by pneumatic nebulization. Each respective element is measured at a specific mass to charge ratio, m/z value expressed in amu (atomic mass units), and at resolution. LR: Low Resolution ~300, MR: Mid Resolution ~4000 and HR: High a specified instrument Resolution ~10000. Ref: Method number 02-2003, National Laboratory for Environmental Testing, CCIW, Environment Canada, Burlington. Requested April 2003. | 5e-05 | Y | Y | Y | 17336 | 5/20/2003 | 12/6/2023 2:08:00 AM | 2/14/2024 9:05:32 AM | 2004-01-19 |  | UMULLER |  | Concentration | mg/L | Y | (fl. conc.) |  | Zinc Total (fl. conc.) | 7440-66-6 | Metals - Total; | Metals - Total; EMS code: ZN-T | Chemical | Total | LAB
ZN-T | mg/L | Zinc Total | F033 | Acid Dig,Centrifuge,ICPMS | For the determination of total recoverable elements in water or waste water, take 100 ml aliquot from a well mixed, acid preserved, sample containing not more than 0.25% (w/v) total solids and transfer to a 250 ml Griffin beaker. (If total solids are greater than 0.25% reduce the size of the aliquot by a proportionate amount). Add 1 ml of conc. nitric acid and 0.5 ml conc. hydrochloric acid. Heat on a hot plate until the volume has been reduced to 15-20ml, ensuring that the sample does not boil. Allow to cool and transfer to a 50 ml volumetric flask. Dilute to volume with ASTM type I water. Centrifuge the sample or allow to stand overnight to separate insoluble material. Prior analysis, pipette 20 ml into a 50 ml volumetric flask and dilute to volume with ASTM type I water. Analysis should be performed within seventy two hours of the completed preparation. Note that different laboratories may have very different field preservation, digestion procedures and detection limits for this method, so these should be documented. | 0.001 | Y | Y | Y | 40834 | 10/15/1996 9:27:00 AM | 8/22/2024 3:24:00 AM | 8/25/2024 9:47:36 PM | 1997-01-16 | 6/3/1998 6:51:46 PM | MPAK |  | Concentration | mg/L | Y | (fl. conc.) |  | Zinc Total (fl. conc.) | 7440-66-6 | Metals - Total; | Metals - Total; EMS code: ZN-T | Chemical | Total | LAB
ZN-T | ug/g wet | Zinc Total | X217 | Hom,Nitric/Perchl,ICP | Sample is wet but not dried before homogenization and nitric perchloric digestion followed by ICP analysis. | 0.05 | Y | Y | Y | 74 |  |  | 10/24/1996 10:01:25 AM | 1996-04-12 | 6/3/1998 6:51:42 PM | EMS |  | MassPerMass | µg/g | Y | (mass conc.) | (wet) | Zinc Total (wet) (mass conc.) | 7440-66-6 | Metals - Total; | Metals - Total; EMS code: ZN-T | Chemical | Total | LAB
ZN-T | ug/m3 | Zinc Total | F088 | Aliquot,HNO3/HCl dig, ICPMS,loading | Aliquot microwave digested with 3:1 HNO3/HCL. Analysis by ICPMS. For HiVol samples preliminary result are reported in mg/L and loading results are reported in ug/m3 |  | Y | Y | Y | 565 | 8/17/2000 10:46:00 AM | 12/14/2006 | 12/30/2006 1:01:38 AM | 2000-08-16 |  | NPEPPIN |  | AirConcentration | µg/m³ | Y | (air conc.) |  | Zinc Total (air conc.) | 7440-66-6 | Metals - Total; | Metals - Total; EMS code: ZN-T | Chemical | Total | LAB
ZN-T | ug/g | Zinc Total | ICPM | Metals in Soil by CRC ICPMS | This analysis is carried out using procedures from CSR Analytical Method: "Strong Acid Leachable Metals (SALM) in Soil", BC Ministry of Environment, 26 June 2009, and procedures adapted from EPA Method 200.2. The sample is manually homogenized, dried at 60 degrees Celsius, sieved through a 63 um (230 mesh) sieve, and a representative subsample of the dry material is weighed. The sample is then digested at 95 degrees Celsius for 2 hours by block digester using concentrated nitric and hydrochloric acids. Instrumental analysis is by collision cell inductively coupled plasma - mass spectrometry (modifed from EPA Method 6020A). Method Limitation: This method is not a total digestion technique. It is a very strong acid digestion that is intended to dissolve those metals that may be environmentally available. By design, elements bound in silicate structures are not normally dissolved by this procedure as they are not usually mobile in the environment. | 1 | Y | Y | Y | 103 | 9/19/2017 | 9/24/2023 | 1/12/2024 11:16:52 AM | 2018-01-09 |  | JKROGH |  | MassPerMass | mg/kg | Y | (mass conc.) |  | Zinc Total (mass conc.) | 7440-66-6 | Metals - Total; | Metals - Total; EMS code: ZN-T | Chemical | Total | LAB
ZN-T | ug/g | Zinc Total | FS01 | D&G,Nitric/HCl:ICP | Samples are dried at 60 C, then either sieved to -100mesh (150 um) and rolled to homogenize, or finely ground with a mortar and pestle or a shatter box laboratory mill and rolled (depending on the submitter's requirements). The samples are then weighed into Teflon digestion liners and digested with HNO3 and HCl in a microwave oven. Samples are cooled, volumized, and allowed to settle overnight before being decanted. Analysis is then performed via Inductively Coupled Argon Plasma Atomic Emission Spectrometry (ICP-AES) (Simultaneous multi element analysis). In ICP Spectrometry, the liquid sample is pumped to the nebulizer, where it is convertd into an aerosol. the nebulizer carrier gas flow than transports the wet aerosol past the impact bead, where the large droplets condense out and flow to the waste container. The remaining fine aerosol is carried into the plasma. There, the high temperature desolvates the aerosol and dissociates the sample material into individual atoms and ions, which are excited to emit light at wavelengths characteristic of the elements present. A diffraction grating in the spectrometer diffracts the various wavelengths of light through exit slits to the photo-multiplier tubes, which convert the light energy to electronic signals. | 0.3 | Y | Y | Y | 2642 | 8/30/1996 1:52:00 PM | 10/10/2007 | 5/24/2008 1:15:03 AM | 1996-11-20 | 6/3/1998 6:51:46 PM | MPAK |  | MassPerMass | mg/kg | Y | (mass conc.) |  | Zinc Total (mass conc.) | 7440-66-6 | Metals - Total; | Metals - Total; EMS code: ZN-T | Chemical | Total | LAB
ZN-T | mg/L | Zinc Total | 0100 | Dig:Ext.(APDC):AA Anal. | Dig:Ext.(APDC):AA Anal. | 0.005 | Y | Y | Y | 1067 |  |  | 6/16/2000 4:19:26 PM | 1996-04-12 | 6/3/1998 6:51:41 PM | EMS |  | Concentration | mg/L | Y | (fl. conc.) |  | Zinc Total (fl. conc.) | 7440-66-6 | Metals - Total; | Metals - Total; EMS code: ZN-T | Chemical | Total | LAB
ZN-T | ug/g wet | Zinc Total | MTIW | MET-WET-CCMS-MID-VA | Metals in Tissue by CRC ICPMS (WET) Methodology: This method is conducted following British Columbia Lab Manual method "Metals in Animal Tissue and Vegetation (Biota) - Prescriptive". Tissue samples are homogenized and sub-sampled prior to hotblock digestion with nitric and hydrochloric acids, in combination with addition of hydrogen peroxide. Instrumental analysis is by collision cell inductively coupled plasma - mass spectrometry (modified from EPA Method 6020A). Method Limitation: This method employs a strong acid/peroxide digestion, and is intended to provide a conservative estimate of bio-available metals. Near complete recoveries are achieved for most toxicologically important metals, but elements associated with recalcitrant minerals may be only partially recovered. | 0.2 | Y | Y | Y | 24 | 11/25/2018 | 12/6/2018 | 7/21/2021 2:55:08 PM | 2018-12-24 |  | EMS |  | MassPerMass | µg/g | Y | (mass conc.) | (wet) | Zinc Total (wet) (mass conc.) | 7440-66-6 | Metals - Total; | Metals - Total; EMS code: ZN-T | Chemical | Total | LAB
ZN-T | mg/L | Zinc Total | X028 | FA,HNO3,AA Analysis | FA;HNO3;AA Analysis | 0.005 | Y | Y | Y | 227 |  |  | 4/17/1998 3:45:56 PM | 1996-04-12 | 6/3/1998 6:51:41 PM | EMS |  | Concentration | mg/L | Y | (fl. conc.) |  | Zinc Total (fl. conc.) | 7440-66-6 | Metals - Total; | Metals - Total; EMS code: ZN-T | Chemical | Total | LAB
ZN-T | mg/L | Zinc Total | 0020 | HNO3 Dig:AA Analysis | HNO3 Dig:AA Analysis | 0.02 | Y | Y | Y | 1099 |  |  | 6/25/2003 9:22:40 AM | 1996-04-12 | 6/3/1998 6:51:41 PM | EMS |  | Concentration | mg/L | Y | (fl. conc.) |  | Zinc Total (fl. conc.) | 7440-66-6 | Metals - Total; | Metals - Total; EMS code: ZN-T | Chemical | Total | LAB
ZN-T | ug/m3 | Zinc Total | 5312 | Tef:N/P Dig:ICP Analysis | A nitric/perchloric acid digestion if used to bring the metals into solution prior to ICP analysis. For Hi-Vol samples preliminary results are reported as mg/L and 24 hour loading results as ug/m3. | 0.1 | Y | Y | Y | 1704 | 11/7/1997 | 5/15/2000 | 11/18/2000 1:29:41 AM | 1996-04-12 | 6/3/1998 6:51:41 PM | EMS |  | AirConcentration | µg/m³ | Y | (air conc.) |  | Zinc Total (air conc.) | 7440-66-6 | Metals - Total; | Metals - Total; EMS code: ZN-T | Chemical | Total | LAB
ZN-T | mg/L | Zinc Total | X073 | FA,HNO3,Direct Flame,AA | Samples are field-acidified with HNO3; HNO3 digestion; analysis by AA. | 0.005 | Y | Y | Y | 262 | 2/12/1998 | 1/27/2003 | 2/5/2003 9:12:45 AM | 1996-04-12 | 6/3/1998 6:51:41 PM | EMS |  | Concentration | mg/L | Y | (fl. conc.) |  | Zinc Total (fl. conc.) | 7440-66-6 | Metals - Total; | Metals - Total; EMS code: ZN-T | Chemical | Total | LAB
ZN-T | mg/L | Zinc Total | X145 | Dig, ICP | Dig; ICP | 0.005 | Y | Y | Y | 241 | 5/8/2003 9:40:00 AM | 9/12/2003 3:16:00 PM | 12/1/2003 5:30:26 PM | 1996-04-12 | 6/3/1998 6:51:41 PM | EMS |  | Concentration | mg/L | Y | (fl. conc.) |  | Zinc Total (fl. conc.) | 7440-66-6 | Metals - Total; | Metals - Total; EMS code: ZN-T | Chemical | Total | LAB
ZN-T | mg/L | Zinc Total | 0040 | HNO3 Dig: ICP Analysis | HNO3 Dig: ICP Analysis | 0.01 | Y | Y | Y | 14280 | 10/17/2000 | 12/1/2010 | 10/18/2023 8:55:31 AM | 1996-04-12 | 6/3/1998 6:51:41 PM | EMS |  | Concentration | mg/L | Y | (fl. conc.) |  | Zinc Total (fl. conc.) | 7440-66-6 | Metals - Total; | Metals - Total; EMS code: ZN-T | Chemical | Total | LAB
ZN-T | ug/g | Zinc Total | MTID | MET-DRY-CCMS-VA | Metals in Tissue by CRC ICPMS (DRY) Methodology: This method is conducted following British Columbia Lab Manual method "Metals in Animal Tissue and Vegetation (Biota) - Prescriptive". Tissue samples are homogenized and sub-sampled prior to hotblock digestion with nitric and hydrochloric acids, in combination with addition of hydrogen peroxide. Instrumental analysis is by collision cell inductively coupled plasma - mass spectrometry (modified from EPA Method 6020A). Method Limitation: This method employs a strong acid/peroxide digestion, and is intended to provide a conservative estimate of bio-available metals. Near complete recoveries are achieved for most toxicologically important metals, but elements associated with recalcitrant minerals may be only partially recovered. | 0.5 | Y | Y | Y | 122 | 8/13/2014 | 9/8/2022 | 10/26/2022 7:14:15 AM | 2015-03-09 | 3/11/2015 4:12:05 PM | RDGIBSON | EMS | MassPerMass | mg/kg | Y | (mass conc.) |  | Zinc Total (mass conc.) | 7440-66-6 | Metals - Total; | Metals - Total; EMS code: ZN-T | Chemical | Total | LAB
ZN-T | ppm (S) | Zinc Total | 6041 | Aqua Regia by ICPMS | These were sediment/soils digested with Aqua-Regia (Nitric Acid & Hydrochloric Acids) and then analysis by ICPMS. Results would be reported in mg/kg and/or %. | 0.1 | Y | Y | Y | 21 | 10/13/2019 | 8/5/2020 | 8/10/2020 3:52:07 PM | 2019-09-06 | 10/21/2019 11:57:03 AM | KATFRASE | KATFRASE | DimensionlessRatio | ppm | N | (ratio) |  | Zinc Total (ratio) | 7440-66-6 | Metals - Total; | Metals - Total; EMS code: ZN-T | Chemical | Total | LAB
ZN-T | mg/L | Zinc Total | F068 | HNO3dig,TeflonTube/Autoclave,ICP-MS | This method is applicable to determiniation of dissolved, extractable, and total recoverable metals in ground waters, surface waters, drinking waters, and wastewaters. Dissolved, extractable, and total recoverable sample fractions are acidified to pH <1.5 with analytical grade nitric acid (HNO3) to prevent the metals from adsorbing to the container material or precipitating. Extractable and total recoverable water samples are collected in 250 ml acid washed polyethylene bottles and are preserved with Nitric acid (0.5 ml acid for 100 mLs of sample). All preserved samples are stored at room temperature and can be held for up to 6 months. Unfiltered, un-preserved samples are stored at 4C for no more than 30 days. For dissolved and extractable water fractions, no sample preparation is necessary. Samples are simply poured in autosampler tubes (i.e. 15 ml centrifuge tubes), and placed on the autosampler and analyzed. Samples for total recoverable metals are digested with Nitric acid in special Teflon tubes in an autoclave prior to analyses. | 0.0001 | Y | Y | Y | 2443 | 6/1/1999 1:34:00 PM | 7/4/2008 9:48:00 AM | 7/30/2024 6:32:12 AM | 1999-06-28 |  | NPEPPIN |  | Concentration | mg/L | Y | (fl. conc.) |  | Zinc Total (fl. conc.) | 7440-66-6 | Metals - Total; | Metals - Total; EMS code: ZN-T | Chemical | Total | LAB
ZN-T | mg/L | Zinc Total | 0042 | GI HNO3 Dig: ICP | Sample aliquot taken from unfiltered unpreserved General Ions bottle; HNO3 digestion; ICP analysis. | 0.01 | Y | Y | Y | 6857 | 9/20/1996 | 7/5/2019 3:21:00 PM | 7/12/2019 6:49:13 PM | 1996-04-12 | 6/3/1998 6:51:41 PM | EMS |  | Concentration | mg/L | Y | (fl. conc.) |  | Zinc Total (fl. conc.) | 7440-66-6 | Metals - Total; | Metals - Total; EMS code: ZN-T | Chemical | Total | LAB
ZN-T | ug/g | Zinc Total | PB02 | PBM,HNO3/HCl dig,dry60C,100mesh, | Performance Based Method. Samples are dried at 60 C and sieved using a .149 mm (100 mesh) sieve. Solid samples are digested with the combination of nitric and hydrochloric acids with a 1:1 ratio of acids. For environmental purposes, strong acid leaches are used routinely since they have the advantage of speed, low cost, and superio reproducibility. All metals trapped within the refractory matrix are unlikely to have environmental consequences. Instrumental analysis of extracts produced using this digestion method can be performed by a variety of techniques. |  | Y | Y | Y | 52 | 5/24/2001 3:42:00 PM | 11/26/2001 10:23:00 AM | 5/26/2003 2:28:53 PM | 2001-07-18 |  | NPEPPIN |  | MassPerMass | mg/kg | Y | (mass conc.) |  | Zinc Total (mass conc.) | 7440-66-6 | Metals - Total; | Metals - Total; EMS code: ZN-T | Chemical | Total | LAB
ZN-T | ppm (S) | Zinc Total | MS41 | ALS Minerals ME-MS41 | ALS Minerals ME-MS41. A prepared sample is digested with aqua regia in a graphite heating block. After cooling; the resulting solution is diluted to with deionized water; mixed and analyzed by ICP-AES. Samples are then analyzed by ICP-MS for the remaining suite of elements. | 2 | Y | Y | Y | 0 |  |  |  | 2024-01-29 |  | EMS |  | DimensionlessRatio | ppm | N | (ratio) |  | Zinc Total (ratio) | 7440-66-6 | Metals - Total; | Metals - Total; EMS code: ZN-T | Chemical | Total | LAB
ZN-T | mg/L | Zinc Total | X351 | LA:HNO3:AA | Samples are lab-acidified (HNO3); digested HNO3; analysis by AA. | 0.005 | Y | Y | Y | 23 |  |  | 1/19/2007 8:30:58 AM | 1996-04-12 | 6/3/1998 6:51:41 PM | EMS |  | Concentration | mg/L | Y | (fl. conc.) |  | Zinc Total (fl. conc.) | 7440-66-6 | Metals - Total; | Metals - Total; EMS code: ZN-T | Chemical | Total | LAB
ZN-T | mg/L | Zinc Total | METT | MET-T-NP-BCMOE-MS-VA | Ultra trace metals in water are analyzed by CRC ICPMS, based on US EPA Method 6020B (July 2014). The detection limits provided can only be met for undigested samples. This procedure is intended for colorless, non-turbid, acid-preserved water samples (i.e. pristine water samples), having turbidity < 1 NTU and no odor. Where turbidity exceeds 1 NTU, and/or the sample is colored and has an odor, results may be biased low compared to true Total Metals concentrations. ALS recommends that turbidity analysis be requested on samples submitted for this test to aid with interpretation of results. Where turbidity is <1NTU, undigested metals are equivalent to total metals concentrations. BC Ministry of Environment version. | 0.0002 | Y | Y | Y | 2589 | 6/5/2019 | 8/18/2024 | 8/26/2024 1:03:06 PM | 2019-08-20 |  | KATFRASE |  | Concentration | mg/L | Y | (fl. conc.) |  | Zinc Total (fl. conc.) | 7440-66-6 | Metals - Total; | Metals - Total; EMS code: ZN-T | Chemical | Total | LAB
ZN-T | ug/g | Zinc Total | X230 | Hom,HNO3/HCl,ICP | The sample is dried, ground to pass 100 mesh, and digested with a nitric-hydrochloric acid mixture to solubilize the solid matter and to remove the organic material by oxidation and volatilization. The completed digests are submitted to metals instrumentation personnel for metals analysis by Inductively Coupled Atomic Emission Spectrophotometry (ICP-AES), or hydride analysis by atomic absorption. | 0.5 | Y | Y | Y | 267 | 2/12/2002 | 12/30/2005 | 2/3/2006 1:13:13 AM | 1996-04-12 | 6/3/1998 6:51:41 PM | EMS |  | MassPerMass | mg/kg | Y | (mass conc.) |  | Zinc Total (mass conc.) | 7440-66-6 | Metals - Total; | Metals - Total; EMS code: ZN-T | Chemical | Total | LAB
ZN-T | ppm (S) | Zinc Total | MS61 | ALS Minerals ME-MS61 | ALS Minerals ME-MS61. A prepared sample is digested with HNO3; HF; and HClO4 in a graphite heating block to near dryness. After cooling; the sample is boiled in diluted HCl. Bulk to volume with deionized water; mixed and analyzed by ICP-AES and ICP-MS for all elements. | 2 | Y | Y | Y | 0 |  |  |  | 2024-01-29 |  | EMS |  | DimensionlessRatio | ppm | N | (ratio) |  | Zinc Total (ratio) | 7440-66-6 | Metals - Total; | Metals - Total; EMS code: ZN-T | Chemical | Total | LAB
ZN-T | mg/dm2/d | Zinc Total | DUST | HNO3 dig.(microwave or hot block),ICPMS | HNO3 dig.(microwave or hot block);ICPMS |  | Y | Y | Y | 453 | 7/19/2016 1:11:00 PM | 7/31/2024 | 7/31/2024 11:23:13 AM | 2008-02-07 |  | RDGIBSON |  | MassFlux | mg/dm²/d | Y | (mass flux) |  | Zinc Total (mass flux) | 7440-66-6 | Metals - Total; | Metals - Total; EMS code: ZN-T | Chemical | Total | LAB
ZN-T | mg/L | Zinc Total | F121 | Acid,ICPMS | Acidified, no digestion, ICPMS | 0.0001 | Y | Y | Y | 8440 | 3/12/2010 7:57:00 PM | 8/15/2024 9:10:00 PM | 8/26/2024 8:06:03 AM | 2002-07-22 | 7/16/2012 3:32:47 PM | UMULLER | RDGIBSON | Concentration | mg/L | Y | (fl. conc.) |  | Zinc Total (fl. conc.) | 7440-66-6 | Metals - Total; | Metals - Total; EMS code: ZN-T | Chemical | Total | LAB
ZN-T | mg/L | Zinc Total | F038 | HNO3/HCl dig,ICP-AES | Dissolved and extractable water samples are ready for analysis once collected. Total samples are autoclaved with HNO3 and hydrochloric acid (HCI) in capped Teflon vials. Analysis is performed via Inductively Coupled Argon Plasma Atomic Emission Spectrometry (ICP-AES)(simultaneous multielement analysis), Graphite Furnace Atomic Absorption Spectrometry (GF-AAS) (low level analysis) and/or Flame Atomic Absorption Spectrometry (F-AAS). | 0.002 | Y | Y | Y | 5540 | 8/1/1996 6:04:00 PM | 5/6/2008 | 9/11/2019 8:47:43 AM | 1997-05-08 | 6/3/1998 6:51:47 PM | MPAK |  | Concentration | mg/L | Y | (fl. conc.) |  | Zinc Total (fl. conc.) | 7440-66-6 | Metals - Total; | Metals - Total; EMS code: ZN-T | Chemical | Total | LAB
ZN-T | ug/m3 | Zinc Total | EQ40 | EQUIS (mg/m3) | EQUIS (mg/m3) | 0 | N | Y | Y | 3069 |  |  | 10/23/1996 12:01:36 AM | 1996-04-12 | 6/3/1998 6:51:41 PM | EMS |  | AirConcentration | µg/m³ | Y | (air conc.) |  | Zinc Total (air conc.) | 7440-66-6 | Metals - Total; | Metals - Total; EMS code: ZN-T | Chemical | Total | LAB
ZN-T | ug/m3 | Zinc Total | 6040 | Aqua Regia Dig:ICP | Aqua regia digestion is used to bring the metals into solution prior to ICP analysis. For Hi-Vol samples, preliminary results are reported as mg/L and24 hr loadings as ug/m3 | 0.1 | Y | Y | Y | 437 | 2/13/2002 | 2/9/2005 | 2/23/2005 1:13:48 AM | 1998-05-05 | 6/3/1998 6:51:48 PM | NPEPPIN |  | AirConcentration | µg/m³ | Y | (air conc.) |  | Zinc Total (air conc.) | 7440-66-6 | Metals - Total; | Metals - Total; EMS code: ZN-T | Chemical | Total | LAB
ZN-T | mg/L | Zinc Total | MWID | Metals in Water by ICPMS (digested) | Metals in Water by ICPMS (digested) Based on EPA 200.8 | 0.01 | Y | Y | Y | 36975 |  |  | 7/31/2024 1:46:41 PM | 2009-12-24 |  | RDGIBSON |  | Concentration | mg/L | Y | (fl. conc.) |  | Zinc Total (fl. conc.) | 7440-66-6 | Metals - Total; | Metals - Total; EMS code: ZN-T | Chemical | Total | LAB
ZN-T | mg/L | Zinc Total | TCLP | Toxicity Characteristic Leaching Procedure | The TCLP procedure is designed to simulate the climatic leaching action | 0.1 | Y | Y | Y | 38 | 12/8/2009 7:49:00 PM | 9/28/2011 5:53:00 PM | 8/24/2016 2:59:13 PM | 2009-03-06 |  | JCHUDAK |  | Concentration | mg/L | Y | (fl. conc.) |  | Zinc Total (fl. conc.) | 7440-66-6 | Metals - Total; | Metals - Total; EMS code: ZN-T | Chemical | Total | LAB
ZN-T | mg/L | Zinc Total | X349 | FA:HNO3:ICP | Samples are field-acidified with HNO3; HNO3 digestion; analysis by ICP/AES. | 0.005 | Y | Y | Y | 6683 |  |  | 8/26/2024 2:41:32 PM | 1996-04-12 | 6/3/1998 6:51:41 PM | EMS |  | Concentration | mg/L | Y | (fl. conc.) |  | Zinc Total (fl. conc.) | 7440-66-6 | Metals - Total; | Metals - Total; EMS code: ZN-T | Chemical | Total | LAB
ZN-T | mg/L | Zinc Total | F048 | Acid dig,conc'n,nebul.,ICP-OES | This method provides an automated procedure for the Inductively Coupled Argon Plasma Optical Emission Spectrophotometric (ICP-OES) determination and quantitation of trace amounts of metals in surface, ground and precipitation waters after a manual acid digestion and concentration procedure and ultrasonic nebulisation. Organically-bound metals and metal complexes are converted to metal ions by digestion with nitric acid [method details says HCl]. To increase the signal detected by the ICP, the samples are concentrated 10 times. The liquid sample is pumped onto the piezoelectric transducer, where the ultrasonic action efficiently converts it to a fine, dense aerosol. The nebulizer carrier gas flow then transports the wet aerosol through a heated U-tube, where the solvent is vaporized. Subsequently, most of the solvent vapour is removed from the aerosol stream by passing the stream through a chilled condenser. The result is a dry, analyte-laden aerosol which is introduced to the plasma. There, the high temperature desolvates the aerosol and dissociates the sample material into individual atoms and ions, which are excited to emit light at wave lengths characteristic of the elements present. Dissolved samples are field filtered through 0.45 micron cellulose acetate filter prior to acid preservation. Samples for extractable analysis are acidified with 2 mL concentrated HNO3 per litre and allowed to stand a minimum of 24 hours before analysis. Clients providing these samples are advised to preserve samples at the collection site. | 0.0002 | Y | Y | Y | 6159 | 7/4/1997 2:36:00 PM | 4/23/2003 | 9/11/2019 8:46:38 AM | 1997-10-20 | 6/3/1998 6:51:47 PM | NPEPPIN |  | Concentration | mg/L | Y | (fl. conc.) |  | Zinc Total (fl. conc.) | 7440-66-6 | Metals - Total; | Metals - Total; EMS code: ZN-T | Chemical | Total | LAB
ZN-T | mg/L | Zinc Total | X285 | HNO3 Dig:ICP Analysis | HNO3 Dig:ICP Analysis | 0.001 | Y | Y | Y | 3091 | 8/10/1993 | 10/5/1993 | 4/30/2004 1:54:39 PM | 1996-04-12 | 6/3/1998 6:51:41 PM | EMS |  | Concentration | mg/L | Y | (fl. conc.) |  | Zinc Total (fl. conc.) | 7440-66-6 | Metals - Total; | Metals - Total; EMS code: ZN-T | Chemical | Total | LAB
ZN-T | mg/L | Zinc Total | 0105 | 5%AR Dig:APDC Ext:AA Anl | 5%AR Dig:APDC Ext:AA Anl | 0.005 | Y | Y | Y | 418 |  |  | 10/24/1996 2:12:47 PM | 1996-04-12 | 6/3/1998 6:51:41 PM | EMS |  | Concentration | mg/L | Y | (fl. conc.) |  | Zinc Total (fl. conc.) | 7440-66-6 | Metals - Total; | Metals - Total; EMS code: ZN-T | Chemical | Total | LAB
ZN-T | mg/L | Zinc Total | MSEA | Metals in Seawater (undigested) by (CRC) ICP-MS | This method describes the multi-elemental determination of trace elements by Collision/Reaction Cell (CRC) ICP-MS. The method measures ions produced by a radio-frequency inductively coupled plasma. Analyte species originating in a liquid are pneumatically nebulized and the resulting aerosol transported by argon gas into the plasma torch. The ions produced are entrained in the plasma gas and extracted, by means of a differentially pumped vacuum interface, into a mass spectrometer. The ions flow through a Collision/Reaction Cell (CRC) where depending on the mode of analysis either the ions passes through the cell unhindered (no gas - classical ICP-MS mode), or collisions/reactions chemistry takes place between the ions and a neutral gas such as hydrogen or helium. Any polyatomic ions which give interference on the analytes are removed. The ions are then sorted according to their mass-to-charge ratios by a quadrupole mass spectrometer. The ions are then sorted according to their mass-to-charge ratios by a quadrupole mass spectrometer. The ions transmitted through the quadrupole are quantified by a channel electron multiplier and the ion information processed by a data handling system. Although most polyatomic interference is removed, any additional interference must be assessed and valid corrections applied or the data flagged to indicate problems. Interference correction must include compensation for background ions contributed by the plasma gas, reagents, and constituents of the sample matrix. Instrument drift as well as suppressions or enhancements of instrument response caused by the sample matrix must be corrected for by the use of internal standards. | 0.0005 | Y | Y | Y | 432 | 8/12/2009 12:08:00 AM | 4/29/2024 | 6/7/2024 10:03:16 AM | 2009-09-08 |  | RDGIBSON |  | Concentration | mg/L | Y | (fl. conc.) |  | Zinc Total (fl. conc.) | 7440-66-6 | Metals - Total; | Metals - Total; EMS code: ZN-T | Chemical | Total | LAB
ZN-T | mg/dm2/d | Zinc Total | EQ15 | EQUIS (mg/dm2/d) | EQUIS (mg/dm2/d) | 0.0006 | N | Y | Y | 205 |  |  | 10/22/1996 10:18:48 PM | 1996-04-12 | 6/3/1998 6:51:41 PM | EMS |  | MassFlux | mg/dm²/d | Y | (mass flux) |  | Zinc Total (mass flux) | 7440-66-6 | Metals - Total; | Metals - Total; EMS code: ZN-T | Chemical | Total | LAB
ZN-T | mg/L | Zinc Total | 0060 | HNO3 Dig: HGA Analysis | HNO3 Dig: HGA Analysis | 0.005 | Y | Y | Y | 3133 |  |  | 10/22/1997 1:13:43 PM | 1996-04-12 | 6/3/1998 6:51:41 PM | EMS |  | Concentration | mg/L | Y | (fl. conc.) |  | Zinc Total (fl. conc.) | 7440-66-6 | Metals - Total; | Metals - Total; EMS code: ZN-T | Chemical | Total | LAB
ZN-T | ug/g | Zinc Total | EQ02 | EQUIS (ug/g) | EQUIS (ug/g) | 1 | N | Y | Y | 955 |  |  | 10/23/1996 12:08:28 AM | 1996-04-12 | 6/3/1998 6:51:41 PM | EMS |  | MassPerMass | mg/kg | Y | (mass conc.) |  | Zinc Total (mass conc.) | 7440-66-6 | Metals - Total; | Metals - Total; EMS code: ZN-T | Chemical | Total | LAB
ZN-T | ppm (S) | Zinc Total | 6042 | Aqua Regia by ICPMS | These were sediment/soils digested with Aqua-Regia (Nitric Acid & Hydrochloric Acids) and then analysis by ICPMS. Results would be reported in ppm s. | 0.1 | Y | Y | Y | 71 | 5/14/2022 | 7/27/2024 | 7/29/2024 9:13:08 AM | 2022-06-28 |  | EMS |  | DimensionlessRatio | ppm | N | (ratio) |  | Zinc Total (ratio) | 7440-66-6 | Metals - Total; | Metals - Total; EMS code: ZN-T | Chemical | Total | LAB
ZN-T | mg/L | Zinc Total | 1349 | Tot Metals in Seawater by CRC ICPMS (BC) | Tot Metals in Seawater by CRC ICPMS (BC) | 0.003 | Y | Y | Y | 98 | 8/16/2019 | 12/5/2021 | 12/13/2021 12:34:17 PM | 2019-09-30 |  | KATFRASE |  | Concentration | mg/L | Y | (fl. conc.) |  | Zinc Total (fl. conc.) | 7440-66-6 | Metals - Total; | Metals - Total; EMS code: ZN-T | Chemical | Total | LAB
ZN-T | mg/L | Zinc Total | X113 | HNO3/HCl/Perchlor,AA | HNO3/HCl/Perchlor;AA | 0.005 | Y | Y | Y | 1373 |  |  | 11/15/2000 1:51:17 PM | 1996-04-12 | 6/3/1998 6:51:41 PM | EMS |  | Concentration | mg/L | Y | (fl. conc.) |  | Zinc Total (fl. conc.) | 7440-66-6 | Metals - Total; | Metals - Total; EMS code: ZN-T | Chemical | Total | LAB
ZN-T | mg/L | Zinc Total | X545 | FA,HNO3,ICPMS | Analytes nebulized followed by ion entrained in plasma gas and extracted by means of a differentially pumped vacuum interface into a mass spectrometer | 0.001 | Y | Y | Y | 30 |  |  | 10/30/2014 8:31:19 AM | 2001-09-18 | 9/18/2001 1:36:23 PM | NPEPPIN | NPEPPIN | Concentration | mg/L | Y | (fl. conc.) |  | Zinc Total (fl. conc.) | 7440-66-6 | Metals - Total; | Metals - Total; EMS code: ZN-T | Chemical | Total | LAB
ZN-T | ug/g | Zinc Total | 0196 | D&G: LTA/HCl: ICP | D&G: LTA/HCl: ICP | 1 | Y | Y | Y | 27 |  |  | 10/24/1996 2:08:13 PM | 1996-04-12 | 6/3/1998 6:51:41 PM | EMS |  | MassPerMass | mg/kg | Y | (mass conc.) |  | Zinc Total (mass conc.) | 7440-66-6 | Metals - Total; | Metals - Total; EMS code: ZN-T | Chemical | Total | LAB
ZN-T | mg/L | Zinc Total | X095 | FF,FA,HNO3,ICP | FF;FA;HNO3;ICP | 0.005 | Y | Y | Y | 19 |  |  | 10/24/1996 12:17:00 PM | 1996-04-12 | 6/3/1998 6:51:41 PM | EMS |  | Concentration | mg/L | Y | (fl. conc.) |  | Zinc Total (fl. conc.) | 7440-66-6 | Metals - Total; | Metals - Total; EMS code: ZN-T | Chemical | Total | LAB
ZN-T | ug/g wet | Zinc Total | X552 | Metals in Biota (wet), total, acid digestion, ICP-SFMS | One gram of wet, homogenized biota is weighed in a microwave TFM vessel, and 8mL of Nitric Acid and 1mL of Hydrogen Peroxide are added. The sample is digested at 200 degrees C for 15 minutes in a closed vessel high pressure microwave oven. The sample is cooled, the volume adjusted to 100mL and introduced to an inductively coupled argon plasma-sector field mass spectrometer (ICP-SFMS) by pneumatic nebulization. Each respective element is measured at a specific mass to charge ratio, m/z value expressed in amu (atomic mass units), and at a specified instrument resolution. LR: Low Resolution ~300, MR: Mid Resolution ~4000. Ref: 02-2704, National Laboratory for Environmental Testing, CCIW, Environment Canada, Burlington. Requested April 2003. | 0.01 | Y | Y | Y | 37 | 8/9/2016 | 10/16/2023 | 11/28/2023 2:21:33 PM | 2004-01-19 | 4/19/2021 1:22:00 PM | UMULLER | EMS | MassPerMass | µg/g | Y | (mass conc.) | (wet) | Zinc Total (wet) (mass conc.) | 7440-66-6 | Metals - Total; | Metals - Total; EMS code: ZN-T | Chemical | Total | LAB
ZN-T | ug/g | Zinc Total | FS82 | PBM,dry60C,10mesh,SALM-ICP-MS | PBM;dry60C;10mesh;SALM-ICP-MS | 0.5 | Y | Y | Y | 3476 | 5/21/2002 | 7/24/2024 | 7/30/2024 10:00:14 AM | 2002-11-12 |  | UMULLER |  | MassPerMass | mg/kg | Y | (mass conc.) |  | Zinc Total (mass conc.) | 7440-66-6 | Metals - Total; | Metals - Total; EMS code: ZN-T | Chemical | Total | LAB
ZN-T | mg/dm2/d | Zinc Total | CAL1 | Calculation (mg/dm2/d) | Calculation (mg/dm2/d) | 0.0006 | Y | Y | Y | 3046 |  |  | 7/3/2024 12:59:27 PM | 1996-04-12 | 6/3/1998 6:51:42 PM | EMS |  | MassFlux | mg/dm²/d | Y | (mass flux) |  | Zinc Total (mass flux) | 7440-66-6 | Metals - Total; | Metals - Total; EMS code: ZN-T | Chemical | Total | LAB
ZN-T | ug/m3 | Zinc Total | 5097 | N/P Dig(Filt+Wash),ICP | N/P Dig(Filt+Wash);ICP | 0.001 | Y | Y | Y | 10429 |  |  | 7/3/2024 12:59:27 PM | 1996-04-12 | 6/3/1998 6:51:41 PM | EMS |  | AirConcentration | µg/m³ | Y | (air conc.) |  | Zinc Total (air conc.) | 7440-66-6 | Metals - Total; | Metals - Total; EMS code: ZN-T | Chemical | Total | LAB
ZN-T | ug/g | Zinc Total | X551 | Metals in Biota (dry), total, acid digestion, ICP-SFMS | 0.1 gram of dried, homogenized biota is weighed in a microwave TFM vessel, and 8mL of Nitric Acid and 1mL of Hydrogen Peroxide are added. The sample is digested at 200 degrees C for 15 minutes in a closed vessel high pressure microwave oven. The sample is cooled, the volume adjusted to 100mL and introduced to an iductively coupled argon plasma-sector field mass spectrometer (ICP-SFMS) by pneumatic nebulization. Each respective element is measured at a specific mass to charge ratio, m/z value expressed in amu (atomic mass units), and at a specified instrument resolution. LR: Low Resolution ~300, MR: Mid Resolution ~4000 and HR: Hich Resolution ~10000. Ref: Method number 02-2705, National Laboratory for Environmental Testing, CCIW, Environment Canada, Burlington. Requested April 2003. | 0.1 | Y | Y | Y | 49 | 8/11/2015 | 10/30/2023 | 11/28/2023 2:21:33 PM | 2004-01-19 |  | UMULLER |  | MassPerMass | mg/kg | Y | (mass conc.) |  | Zinc Total (mass conc.) | 7440-66-6 | Metals - Total; | Metals - Total; EMS code: ZN-T | Chemical | Total | LAB
ZN-T | mg/L | Zinc Total | METN | MET-T-NP-CCMS-VA | Ultra trace metals in water are analyzed by CRC ICPMS, based on US EPA Method 6020B (July 2014). The detection limits provided can only be met for undigested samples. This procedure is intended for colorless, non-turbid, acid-preserved water samples (i.e. pristine water samples), having turbidity < 1 NTU and no odor. Where turbidity exceeds 1 NTU, and/or the sample is colored and has an odor, results may be biased low compared to true Total Metals concentrations. ALS recommends that turbidity analysis be requested on samples submitted for this test to aid with interpretation of results. Where turbidity is <1NTU, undigested metals are equivalent to total metals concentrations. | 0.0005 | Y | Y | Y | 1059 | 4/3/2012 | 4/4/2024 | 4/29/2024 4:53:06 PM | 2021-04-14 |  | RDGIBSON |  | Concentration | mg/L | Y | (fl. conc.) |  | Zinc Total (fl. conc.) | 7440-66-6 | Metals - Total; | Metals - Total; EMS code: ZN-T | Chemical | Total | LAB
ZN-T | ug/g wet | Zinc Total | BPSY | MET-WET-BPSY-CCMS-VA | Metals in Animal Tissue and Vegetation (Biota) - Prescriptive. EPA 200.3/6020B This method is conducted following British Columbia Lab Manual method "Metals in Animal Tissue and Vegetation (Biota) - Prescriptive". Tissue samples are digested with nitric and hydrochloric acids; in combination with addition of hydrogen peroxide. Instrumental analysis is by collision cell inductively coupled plasma - mass spectrometry (modified from EPA Method 6020B). | 0.2 | Y | Y | Y | 2 | 11/7/2019 | 11/7/2019 | 6/10/2021 1:57:10 PM | 2021-06-04 |  | EMS |  | MassPerMass | µg/g | Y | (mass conc.) | (wet) | Zinc Total (wet) (mass conc.) | 7440-66-6 | Metals - Total; | Metals - Total; EMS code: ZN-T | Chemical | Total | LAB
ZN-T | mg/L | Zinc Total | EQ01 | EQUIS (mg/L) | EQUIS (mg/L) | 0.005 | N | Y | Y | 7771 |  |  | 10/24/1996 3:16:18 PM | 1996-04-12 | 6/3/1998 6:51:41 PM | EMS |  | Concentration | mg/L | Y | (fl. conc.) |  | Zinc Total (fl. conc.) | 7440-66-6 | Metals - Total; | Metals - Total; EMS code: ZN-T | Chemical | Total | LAB
ZN-T | mg/L | Zinc Total | F108 | SPR,IDA Chelation,ICPMS | SPR;IDA Chelation;ICPMS |  | Y | Y | Y | 4 | 11/30/2000 10:32:00 AM | 11/30/2000 10:32:00 AM | 9/10/2002 9:29:49 AM | 2001-03-08 |  | NPEPPIN |  | Concentration | mg/L | Y | (fl. conc.) |  | Zinc Total (fl. conc.) | 7440-66-6 | Metals - Total; | Metals - Total; EMS code: ZN-T | Chemical | Total | LAB
ZN-T | t/d | Zinc Total | CALC | Calculation (t/d) | Calculation (t/d) | 0 | Y | Y | Y | 8387 |  |  | 7/3/2024 12:59:29 PM | 1996-04-12 | 6/3/1998 6:51:41 PM | EMS |  | MassFlow | t/d | Y | (mass flow) |  | Zinc Total (mass flow) | 7440-66-6 | Metals - Total; | Metals - Total; EMS code: ZN-T | Chemical | Total | LAB
ZN-T | mg/L | Zinc Total | X170 | ICP | This code is a catch-all for a large variety of ICP methods including water and also solids digested into solutions as an intermediate result. Different procedures will have greatly different detection limits. (a) For air particulates, preliminary results are reported in mg/L and loading results as mg/dm2/d. (b) For Hi Vol samples preliminiary results are same as reported in mg/L and loading results as ug/m3. (c) Note that for Rain samples there is no filtration nor digestion steps. | 0.005 | Y | Y | Y | 11949 | 3/28/1996 6:03:00 PM | 11/7/2018 | 7/25/2024 11:22:07 AM | 1996-04-12 | 8/15/2019 1:25:09 PM | EMS | KATFRASE | Concentration | mg/L | Y | (fl. conc.) |  | Zinc Total (fl. conc.) | 7440-66-6 | Metals - Total; | Metals - Total; EMS code: ZN-T | Chemical | Total | LAB
ZN-T | ug/m3 | Zinc Total | 5012 | N/P Dig:HGA Analysis | N/P Dig:HGA Analysis | 0.01 | N | Y | Y | 813 |  |  | 10/24/1996 6:54:32 AM | 1996-04-12 | 6/3/1998 6:51:41 PM | EMS |  | AirConcentration | µg/m³ | Y | (air conc.) |  | Zinc Total (air conc.) | 7440-66-6 | Metals - Total; | Metals - Total; EMS code: ZN-T | Chemical | Total | LAB
ZN-T | ug/m3 | Zinc Total | X180 | LVOL:Dig,AA | LVOL:Dig;AA | 0 | Y | Y | Y | 1832 |  |  | 3/9/1998 10:35:51 AM | 1996-04-12 | 6/3/1998 6:51:41 PM | EMS |  | AirConcentration | µg/m³ | Y | (air conc.) |  | Zinc Total (air conc.) | 7440-66-6 | Metals - Total; | Metals - Total; EMS code: ZN-T | Chemical | Total | LAB
ZN-T | ug/g | Zinc Total | 0192 | Hom, Nitric/Perchlor,ICP | Sample already in a dry state so does not require drying before homogenization and nitric perchlorid digestion followed by ICP analysis. | 1 | Y | Y | Y | 706 |  |  | 10/24/1996 12:50:25 PM | 1996-04-12 | 6/3/1998 6:51:41 PM | EMS |  | MassPerMass | mg/kg | Y | (mass conc.) |  | Zinc Total (mass conc.) | 7440-66-6 | Metals - Total; | Metals - Total; EMS code: ZN-T | Chemical | Total | LAB
ZN-T | mg/L | Zinc Total | X512 | HNO3,pHto7.5,MIK extr'n,AA | HNO3 Digestion;pH to 7.5; MIK solvent extraction;AA | 0.001 | Y | Y | Y | 2967 |  |  | 8/23/1999 4:52:44 PM | 1999-08-10 |  | NPEPPIN |  | Concentration | mg/L | Y | (fl. conc.) |  | Zinc Total (fl. conc.) | 7440-66-6 | Metals - Total; | Metals - Total; EMS code: ZN-T | Chemical | Total | LAB
ZN-T | mg/dm2/d | Zinc Total | CS15 | System Calculation | System Calculation (mg/dm2/d) |  | N | Y | N | 44 |  |  | 11/18/2000 1:30:25 AM | 1996-04-12 | 6/3/1998 6:51:44 PM | EMS |  | MassFlux | mg/dm²/d | Y | (mass flux) |  | Zinc Total (mass flux) | 7440-66-6 | Metals - Total; | Metals - Total; EMS code: ZN-T | Chemical | Total | LAB
ZR-D | mg/L | Zirconium Dissolved | X562 | Dissolved metals: FF, FA, HNO3, HR-ICPMS | Ultra trace metals in water are analyzed by high resolution inductively coupled plasma mass spectrometry (HR-ICPMS), based on US EPA Method 200.8, (Oct. 1999). This procedure is intended for pristine field-filtered acid-preserved water samples. The detection limits (LOR) for this test are based on lab instrumental analysis only, not including filtration. ALS-supplied field filtration equipment does not support these LOR. Therefore, because of the high probability of false positives due to filtration, it is strongly recommended that a filtration blank be analyzed to aid in data interpretation. | 5e-05 | Y | Y | Y | 22 | 10/20/2015 | 10/19/2023 | 11/9/2023 7:41:27 AM | 2016-03-29 |  | RDGIBSON |  | Concentration | mg/L | Y | (fl. conc.) |  | Zirconium Dissolved (fl. conc.) | 7440-67-7 | Metals - Dissolved; | Metals - Dissolved; EMS code: ZR-D | Chemical | Dissolved | LAB
ZR-D | mg/L | Zirconium Dissolved | F032 | Filt.,Acid,ICPMS | There is no acid digestion required for this analysis. A water sample is collected and preserved in 0.2 percent nitric acid. To analyse the extractable metals the sample is shaken, allowed to sit for at least twenty four hours, and a decanted portion is introduced to the spectrometer. When the dissolved metals are required, the samples must first be filtered through a 0.45 micron cellulose acetate filter paper and then followed by preservation in 0.2 percent nitric acid. It is important to note that different laboratories may have very different detection limits for some variables analyzed by ICPMS. | 0.001 | Y | Y | Y | 45933 | 5/7/1998 | 8/26/2024 | 8/26/2024 4:51:04 PM | 1998-05-14 | 6/3/1998 6:51:49 PM | NPEPPIN |  | Concentration | mg/L | Y | (fl. conc.) |  | Zirconium Dissolved (fl. conc.) | 7440-67-7 | Metals - Dissolved; | Metals - Dissolved; EMS code: ZR-D | Chemical | Dissolved | LAB
ZR-D | mg/L | Zirconium Dissolved | MWID | Metals in Water by ICPMS (digested) | Metals in Water by ICPMS (digested) Based on EPA 200.8 | 0.002 | Y | Y | Y | 452 | 10/26/2018 5:20:00 PM | 10/26/2018 5:20:00 PM | 1/27/2023 10:37:54 AM | 2009-12-24 |  | RDGIBSON |  | Concentration | mg/L | Y | (fl. conc.) |  | Zirconium Dissolved (fl. conc.) | 7440-67-7 | Metals - Dissolved; | Metals - Dissolved; EMS code: ZR-D | Chemical | Dissolved | LAB
ZR-D | mg/L | Zirconium Dissolved | X285 | HNO3 Dig:ICP Analysis | HNO3 Dig:ICP Analysis | 0.01 | Y | Y | Y | 647 | 1/30/2002 | 11/13/2002 | 2/19/2004 4:22:44 PM | 1998-05-22 | 6/3/1998 6:51:49 PM | NPEPPIN |  | Concentration | mg/L | Y | (fl. conc.) |  | Zirconium Dissolved (fl. conc.) | 7440-67-7 | Metals - Dissolved; | Metals - Dissolved; EMS code: ZR-D | Chemical | Dissolved | LAB
ZR-D | mg/L | Zirconium Dissolved | METC | MET-D-NP-CCMS-VA | Ultra trace metals in water are analyzed by CRC ICPMS, based on US EPA Method 6020B (July 2014). This procedure is intended for pristine field-filtered acid-preserved water samples. The detection limits (LOR) for this test are based on lab instrumental analysis only, not including filtration. ALS-supplied field filtration equipment does not support these LOR. Therefore, because of the high probability of false positives due to filtration, it is strongly recommended that a filtration blank be analysed to aid in data interpretation. | 0.0001 | Y | Y | Y | 1179 | 4/3/2012 | 4/9/2024 | 4/29/2024 4:53:06 PM | 2021-04-14 |  | RDGIBSON |  | Concentration | mg/L | Y | (fl. conc.) |  | Zirconium Dissolved (fl. conc.) | 7440-67-7 | Metals - Dissolved; | Metals - Dissolved; EMS code: ZR-D | Chemical | Dissolved | LAB
ZR-D | mg/L | Zirconium Dissolved | X170 | ICP | This code is a catch-all for a large variety of ICP methods including water and also solids digested into solutions as an intermediate result. Different procedures will have greatly different detection limits. (a) For air particulates, preliminary results are reported in mg/L and loading results as mg/dm2/d. (b) For Hi Vol samples preliminiary results are same as reported in mg/L and loading results as ug/m3. (c) Note that for Rain samples there is no filtration nor digestion steps. | 0.003 | Y | Y | Y | 2159 | 12/3/1996 | 6/24/2002 | 2/10/2004 1:00:41 AM | 1996-04-12 | 6/3/1998 6:51:41 PM | EMS |  | Concentration | mg/L | Y | (fl. conc.) |  | Zirconium Dissolved (fl. conc.) | 7440-67-7 | Metals - Dissolved; | Metals - Dissolved; EMS code: ZR-D | Chemical | Dissolved | LAB
ZR-D | mg/L | Zirconium Dissolved | F039 | Field filt + acid, ICP-AES | Dissolved and extractable water samples are ready for analysis once collected. Total samples are autoclaved with HNO3 and hydrochloric acid (HCI) in capped Teflon vials. Analysis is performed via Inductively Coupled Argon Plasma Atomic Emission Spectrometry (ICP-AES)(simultaneous multielement analysis), Graphite Furnace Atomic Absorption Spectrometry (GF-AAS)(low level analysis) and/or Flame Atomic Absorption Spectrometry (F-AAS). | 0.01 | Y | Y | Y | 2729 | 1/18/2002 | 7/5/2019 1:21:00 PM | 7/12/2019 6:49:16 PM | 1998-05-21 | 6/3/1998 6:51:49 PM | NPEPPIN |  | Concentration | mg/L | Y | (fl. conc.) |  | Zirconium Dissolved (fl. conc.) | 7440-67-7 | Metals - Dissolved; | Metals - Dissolved; EMS code: ZR-D | Chemical | Dissolved | LAB
ZR-D | mg/L | Zirconium Dissolved | MSEA | Metals in Seawater (undigested) by (CRC) ICP-MS | This method describes the multi-elemental determination of trace elements by Collision/Reaction Cell (CRC) ICP-MS. The method measures ions produced by a radio-frequency inductively coupled plasma. Analyte species originating in a liquid are pneumatically nebulized and the resulting aerosol transported by argon gas into the plasma torch. The ions produced are entrained in the plasma gas and extracted, by means of a differentially pumped vacuum interface, into a mass spectrometer. The ions flow through a Collision/Reaction Cell (CRC) where depending on the mode of analysis either the ions passes through the cell unhindered (no gas - classical ICP-MS mode), or collisions/reactions chemistry takes place between the ions and a neutral gas such as hydrogen or helium. Any polyatomic ions which give interference on the analytes are removed. The ions are then sorted according to their mass-to-charge ratios by a quadrupole mass spectrometer. The ions are then sorted according to their mass-to-charge ratios by a quadrupole mass spectrometer. The ions transmitted through the quadrupole are quantified by a channel electron multiplier and the ion information processed by a data handling system. Although most polyatomic interference is removed, any additional interference must be assessed and valid corrections applied or the data flagged to indicate problems. Interference correction must include compensation for background ions contributed by the plasma gas, reagents, and constituents of the sample matrix. Instrument drift as well as suppressions or enhancements of instrument response caused by the sample matrix must be corrected for by the use of internal standards. | 8e-05 | Y | Y | Y | 4 | 4/12/2023 | 4/29/2024 | 5/16/2024 2:00:14 PM | 2009-09-08 |  | RDGIBSON |  | Concentration | mg/L | Y | (fl. conc.) |  | Zirconium Dissolved (fl. conc.) | 7440-67-7 | Metals - Dissolved; | Metals - Dissolved; EMS code: ZR-D | Chemical | Dissolved | LAB
ZR-T | ppm (S) | Zirconium Total | 6041 | Aqua Regia by ICPMS | These were sediment/soils digested with Aqua-Regia (Nitric Acid & Hydrochloric Acids) and then analysis by ICPMS. Results would be reported in mg/kg and/or %. | 0.01 | Y | Y | Y | 21 | 10/13/2019 | 8/5/2020 | 8/10/2020 3:52:07 PM | 2019-09-06 | 10/21/2019 11:57:06 AM | KATFRASE | KATFRASE | DimensionlessRatio | ppm | N | (ratio) |  | Zirconium Total (ratio) | 7440-67-7 | Metals - Total; | Metals - Total; EMS code: ZR-T | Chemical | Total | LAB
ZR-T | ppm (S) | Zirconium Total | MS61 | ALS Minerals ME-MS61 | ALS Minerals ME-MS61. A prepared sample is digested with HNO3; HF; and HClO4 in a graphite heating block to near dryness. After cooling; the sample is boiled in diluted HCl. Bulk to volume with deionized water; mixed and analyzed by ICP-AES and ICP-MS for all elements. | 0.5 | Y | Y | Y | 0 |  |  |  | 2024-01-29 |  | EMS |  | DimensionlessRatio | ppm | N | (ratio) |  | Zirconium Total (ratio) | 7440-67-7 | Metals - Total; | Metals - Total; EMS code: ZR-T | Chemical | Total | LAB
ZR-T | mg/L | Zirconium Total | F082 | Aliquot,HNO3/HCL dig,ICPMS | Aliquot microwave digested with 3:1 HNO3/HCL. Analysis by ICPMS. For HiVol samples preliminary result are reported in mg/L and loading results are reported in ug/m3 | 0.0003 | Y | Y | Y | 14538 | 4/6/2015 | 8/23/2024 | 8/26/2024 3:51:07 PM | 2015-07-30 |  | RDGIBSON |  | Concentration | mg/L | Y | (fl. conc.) |  | Zirconium Total (fl. conc.) | 7440-67-7 | Metals - Total; | Metals - Total; EMS code: ZR-T | Chemical | Total | LAB
ZR-T | mg/L | Zirconium Total | X285 | HNO3 Dig:ICP Analysis | HNO3 Dig:ICP Analysis | 0.01 | Y | Y | Y | 529 | 8/10/1993 | 10/5/1993 | 4/30/2004 11:32:39 AM | 1998-05-22 | 6/3/1998 6:51:49 PM | NPEPPIN |  | Concentration | mg/L | Y | (fl. conc.) |  | Zirconium Total (fl. conc.) | 7440-67-7 | Metals - Total; | Metals - Total; EMS code: ZR-T | Chemical | Total | LAB
ZR-T | ug/g wet | Zirconium Total | X560 | Metals in Biota (wet), total, acid digestion, ICP-MS | Five grams of wet, homogenized biota is weighed in a centrifuge tube, and 2.5 mL of Nitric Acid and 2.5 mL of Hydrochloric acid are added. The sample is digested at room temperature for one hour and then digested at 90-95 degrees C in a water bath or heat block for one more hour. The sample is cooled and the volume is made to 50 mL with DI. The digestate is then diluted at ten fold and introduced to a collision/reaction cell inductively coupled argon plasma mass spectrometer (CRC ICPMS) by pneumatic nebulization. Each respective element is measured at a specific mass to charge ratio, m/z value expressed in amu (atomic mass units). - Maxxam Analytics (BC). Added to EMS Dec 10, 2008. | 0.6 | Y | Y | Y | 181 | 10/19/2009 | 10/30/2023 | 11/28/2023 2:21:33 PM | 2010-02-01 | 3/11/2015 4:17:10 PM | RDGIBSON | EMS | MassPerMass | µg/g | Y | (mass conc.) | (wet) | Zirconium Total (wet) (mass conc.) | 7440-67-7 | Metals - Total; | Metals - Total; EMS code: ZR-T | Chemical | Total | LAB
ZR-T | ug/g wet | Zirconium Total | BPSY | MET-WET-BPSY-CCMS-VA | Metals in Animal Tissue and Vegetation (Biota) - Prescriptive. EPA 200.3/6020B This method is conducted following British Columbia Lab Manual method "Metals in Animal Tissue and Vegetation (Biota) - Prescriptive". Tissue samples are digested with nitric and hydrochloric acids; in combination with addition of hydrogen peroxide. Instrumental analysis is by collision cell inductively coupled plasma - mass spectrometry (modified from EPA Method 6020B). | 0.04 | Y | Y | Y | 2 | 11/7/2019 | 11/7/2019 | 6/10/2021 1:57:10 PM | 2021-06-04 |  | EMS |  | MassPerMass | µg/g | Y | (mass conc.) | (wet) | Zirconium Total (wet) (mass conc.) | 7440-67-7 | Metals - Total; | Metals - Total; EMS code: ZR-T | Chemical | Total | LAB
ZR-T | mg/L | Zirconium Total | F033 | Acid Dig,Centrifuge,ICPMS | For the determination of total recoverable elements in water or waste water, take 100 ml aliquot from a well mixed, acid preserved, sample containing not more than 0.25% (w/v) total solids and transfer to a 250 ml Griffin beaker. (If total solids are greater than 0.25% reduce the size of the aliquot by a proportionate amount). Add 1 ml of conc. nitric acid and 0.5 ml conc. hydrochloric acid. Heat on a hot plate until the volume has been reduced to 15-20ml, ensuring that the sample does not boil. Allow to cool and transfer to a 50 ml volumetric flask. Dilute to volume with ASTM type I water. Centrifuge the sample or allow to stand overnight to separate insoluble material. Prior analysis, pipette 20 ml into a 50 ml volumetric flask and dilute to volume with ASTM type I water. Analysis should be performed within seventy two hours of the completed preparation. Note that different laboratories may have very different field preservation, digestion procedures and detection limits for this method, so these should be documented. | 0.001 | Y | Y | Y | 23501 | 5/7/1998 | 8/22/2024 3:24:00 AM | 8/25/2024 9:47:36 PM | 1998-05-14 | 6/3/1998 6:51:49 PM | NPEPPIN |  | Concentration | mg/L | Y | (fl. conc.) |  | Zirconium Total (fl. conc.) | 7440-67-7 | Metals - Total; | Metals - Total; EMS code: ZR-T | Chemical | Total | LAB
ZR-T | ppm (S) | Zirconium Total | MS81 | ALS Minerals: ME-MS81 | ALS Minerals: ME-MS81. A prepared sample (0.100 g) is added to lithium metaborate/lithium tetraborate flux; mixed well and fused in a furnace at 1025°C. The resulting melt is then cooled and dissolved in an acid mixture containing nitric; hydrochloric and hydrofluoric acids. This solution is then analyzed by inductively coupled plasma - mass spectrometry. | 2 | Y | Y | Y | 0 |  |  |  | 2024-01-29 |  | EMS |  | DimensionlessRatio | ppm | N | (ratio) |  | Zirconium Total (ratio) | 7440-67-7 | Metals - Total; | Metals - Total; EMS code: ZR-T | Chemical | Total | LAB
ZR-T | ug/g | Zirconium Total | FS82 | PBM,dry60C,10mesh,SALM-ICP-MS | PBM;dry60C;10mesh;SALM-ICP-MS | 0.5 | Y | Y | Y | 2030 | 1/31/2002 | 7/24/2024 | 7/29/2024 5:17:09 PM | 2002-11-12 |  | UMULLER |  | MassPerMass | mg/kg | Y | (mass conc.) |  | Zirconium Total (mass conc.) | 7440-67-7 | Metals - Total; | Metals - Total; EMS code: ZR-T | Chemical | Total | LAB
ZR-T | ug/g wet | Zirconium Total | MTIW | MET-WET-CCMS-MID-VA | Metals in Tissue by CRC ICPMS (WET) Methodology: This method is conducted following British Columbia Lab Manual method "Metals in Animal Tissue and Vegetation (Biota) - Prescriptive". Tissue samples are homogenized and sub-sampled prior to hotblock digestion with nitric and hydrochloric acids, in combination with addition of hydrogen peroxide. Instrumental analysis is by collision cell inductively coupled plasma - mass spectrometry (modified from EPA Method 6020A). Method Limitation: This method employs a strong acid/peroxide digestion, and is intended to provide a conservative estimate of bio-available metals. Near complete recoveries are achieved for most toxicologically important metals, but elements associated with recalcitrant minerals may be only partially recovered. | 0.04 | Y | Y | Y | 24 | 11/25/2018 | 12/6/2018 | 7/21/2021 2:55:08 PM | 2018-12-24 |  | EMS |  | MassPerMass | µg/g | Y | (mass conc.) | (wet) | Zirconium Total (wet) (mass conc.) | 7440-67-7 | Metals - Total; | Metals - Total; EMS code: ZR-T | Chemical | Total | LAB
ZR-T | mg/L | Zirconium Total | F121 | Acid,ICPMS | Acidified, no digestion, ICPMS |  | Y | Y | Y | 4852 | 3/12/2010 7:57:00 PM | 8/15/2024 9:10:00 PM | 8/26/2024 8:06:03 AM | 2002-07-22 |  | UMULLER |  | Concentration | mg/L | Y | (fl. conc.) |  | Zirconium Total (fl. conc.) | 7440-67-7 | Metals - Total; | Metals - Total; EMS code: ZR-T | Chemical | Total | LAB
ZR-T | mg/L | Zirconium Total | MWID | Metals in Water by ICPMS (digested) | Metals in Water by ICPMS (digested) Based on EPA 200.8 | 0.002 | Y | Y | Y | 350 |  |  | 1/19/2023 12:03:00 PM | 2009-12-24 |  | RDGIBSON |  | Concentration | mg/L | Y | (fl. conc.) |  | Zirconium Total (fl. conc.) | 7440-67-7 | Metals - Total; | Metals - Total; EMS code: ZR-T | Chemical | Total | LAB
ZR-T | ug/g | Zirconium Total | ICPM | Metals in Soil by CRC ICPMS | This analysis is carried out using procedures from CSR Analytical Method: "Strong Acid Leachable Metals (SALM) in Soil", BC Ministry of Environment, 26 June 2009, and procedures adapted from EPA Method 200.2. The sample is manually homogenized, dried at 60 degrees Celsius, sieved through a 63 um (230 mesh) sieve, and a representative subsample of the dry material is weighed. The sample is then digested at 95 degrees Celsius for 2 hours by block digester using concentrated nitric and hydrochloric acids. Instrumental analysis is by collision cell inductively coupled plasma - mass spectrometry (modifed from EPA Method 6020A). Method Limitation: This method is not a total digestion technique. It is a very strong acid digestion that is intended to dissolve those metals that may be environmentally available. By design, elements bound in silicate structures are not normally dissolved by this procedure as they are not usually mobile in the environment. | 1 | Y | Y | Y | 80 | 7/14/2022 | 9/24/2023 | 1/12/2024 11:16:53 AM | 2022-08-24 |  | EMS |  | MassPerMass | mg/kg | Y | (mass conc.) |  | Zirconium Total (mass conc.) | 7440-67-7 | Metals - Total; | Metals - Total; EMS code: ZR-T | Chemical | Total | LAB
ZR-T | mg/L | Zirconium Total | 0040 | HNO3 Dig: ICP Analysis | HNO3 Dig: ICP Analysis | 0.003 | Y | Y | Y | 1297 | 6/9/2003 12:01:00 PM | 12/1/2010 | 12/9/2010 3:55:27 PM | 1996-04-12 | 6/3/1998 6:51:41 PM | EMS |  | Concentration | mg/L | Y | (fl. conc.) |  | Zirconium Total (fl. conc.) | 7440-67-7 | Metals - Total; | Metals - Total; EMS code: ZR-T | Chemical | Total | LAB
ZR-T | mg/L | Zirconium Total | METN | MET-T-NP-CCMS-VA | Ultra trace metals in water are analyzed by CRC ICPMS, based on US EPA Method 6020B (July 2014). The detection limits provided can only be met for undigested samples. This procedure is intended for colorless, non-turbid, acid-preserved water samples (i.e. pristine water samples), having turbidity < 1 NTU and no odor. Where turbidity exceeds 1 NTU, and/or the sample is colored and has an odor, results may be biased low compared to true Total Metals concentrations. ALS recommends that turbidity analysis be requested on samples submitted for this test to aid with interpretation of results. Where turbidity is <1NTU, undigested metals are equivalent to total metals concentrations. | 0.0001 | Y | Y | Y | 1059 | 4/3/2012 | 4/4/2024 | 4/29/2024 4:53:06 PM | 2021-04-14 |  | RDGIBSON |  | Concentration | mg/L | Y | (fl. conc.) |  | Zirconium Total (fl. conc.) | 7440-67-7 | Metals - Total; | Metals - Total; EMS code: ZR-T | Chemical | Total | LAB
ZR-T | mg/L | Zirconium Total | TCLP | Toxicity Characteristic Leaching Procedure | The TCLP procedure is designed to simulate the climatic leaching action | 0.1 | Y | Y | Y | 22 | 12/8/2009 7:49:00 PM | 9/28/2011 5:53:00 PM | 11/3/2011 11:10:21 PM | 2009-03-06 |  | JCHUDAK |  | Concentration | mg/L | Y | (fl. conc.) |  | Zirconium Total (fl. conc.) | 7440-67-7 | Metals - Total; | Metals - Total; EMS code: ZR-T | Chemical | Total | LAB
ZR-T | mg/L | Zirconium Total | 0042 | GI HNO3 Dig: ICP | Sample aliquot taken from unfiltered unpreserved General Ions bottle; HNO3 digestion; ICP analysis. | 0.003 | Y | Y | Y | 8592 | 9/20/1996 | 7/5/2019 3:21:00 PM | 7/12/2019 6:49:13 PM | 1996-04-12 | 6/3/1998 6:51:41 PM | EMS |  | Concentration | mg/L | Y | (fl. conc.) |  | Zirconium Total (fl. conc.) | 7440-67-7 | Metals - Total; | Metals - Total; EMS code: ZR-T | Chemical | Total | LAB
ZR-T | ug/g wet | Zirconium Total | X552 | Metals in Biota (wet), total, acid digestion, ICP-SFMS | One gram of wet, homogenized biota is weighed in a microwave TFM vessel, and 8mL of Nitric Acid and 1mL of Hydrogen Peroxide are added. The sample is digested at 200 degrees C for 15 minutes in a closed vessel high pressure microwave oven. The sample is cooled, the volume adjusted to 100mL and introduced to an inductively coupled argon plasma-sector field mass spectrometer (ICP-SFMS) by pneumatic nebulization. Each respective element is measured at a specific mass to charge ratio, m/z value expressed in amu (atomic mass units), and at a specified instrument resolution. LR: Low Resolution ~300, MR: Mid Resolution ~4000. Ref: 02-2704, National Laboratory for Environmental Testing, CCIW, Environment Canada, Burlington. Requested April 2003. | 0.04 | Y | Y | Y | 37 | 8/9/2016 | 10/16/2023 | 11/28/2023 2:21:33 PM | 2015-03-11 |  | RDGIBSON |  | MassPerMass | µg/g | Y | (mass conc.) | (wet) | Zirconium Total (wet) (mass conc.) | 7440-67-7 | Metals - Total; | Metals - Total; EMS code: ZR-T | Chemical | Total | LAB
ZR-T | ug/g | Zirconium Total | MTID | MET-DRY-CCMS-VA | Metals in Tissue by CRC ICPMS (DRY) Methodology: This method is conducted following British Columbia Lab Manual method "Metals in Animal Tissue and Vegetation (Biota) - Prescriptive". Tissue samples are homogenized and sub-sampled prior to hotblock digestion with nitric and hydrochloric acids, in combination with addition of hydrogen peroxide. Instrumental analysis is by collision cell inductively coupled plasma - mass spectrometry (modified from EPA Method 6020A). Method Limitation: This method employs a strong acid/peroxide digestion, and is intended to provide a conservative estimate of bio-available metals. Near complete recoveries are achieved for most toxicologically important metals, but elements associated with recalcitrant minerals may be only partially recovered. | 0.2 | Y | Y | Y | 122 | 8/13/2014 | 9/8/2022 | 10/26/2022 7:14:15 AM | 2015-03-09 | 3/11/2015 4:12:05 PM | RDGIBSON | EMS | MassPerMass | mg/kg | Y | (mass conc.) |  | Zirconium Total (mass conc.) | 7440-67-7 | Metals - Total; | Metals - Total; EMS code: ZR-T | Chemical | Total | LAB
ZR-T | ug/g | Zirconium Total | FS01 | D&G,Nitric/HCl:ICP | Samples are dried at 60 C, then either sieved to -100mesh (150 um) and rolled to homogenize, or finely ground with a mortar and pestle or a shatter box laboratory mill and rolled (depending on the submitter's requirements). The samples are then weighed into Teflon digestion liners and digested with HNO3 and HCl in a microwave oven. Samples are cooled, volumized, and allowed to settle overnight before being decanted. Analysis is then performed via Inductively Coupled Argon Plasma Atomic Emission Spectrometry (ICP-AES) (Simultaneous multi element analysis). In ICP Spectrometry, the liquid sample is pumped to the nebulizer, where it is convertd into an aerosol. the nebulizer carrier gas flow than transports the wet aerosol past the impact bead, where the large droplets condense out and flow to the waste container. The remaining fine aerosol is carried into the plasma. There, the high temperature desolvates the aerosol and dissociates the sample material into individual atoms and ions, which are excited to emit light at wavelengths characteristic of the elements present. A diffraction grating in the spectrometer diffracts the various wavelengths of light through exit slits to the photo-multiplier tubes, which convert the light energy to electronic signals. | 0.5 | Y | Y | Y | 271 | 1/31/2002 | 3/22/2002 | 4/1/2004 1:19:36 AM | 2004-02-10 |  | UMULLER |  | MassPerMass | mg/kg | Y | (mass conc.) |  | Zirconium Total (mass conc.) | 7440-67-7 | Metals - Total; | Metals - Total; EMS code: ZR-T | Chemical | Total | LAB
ZR-T | ppm (S) | Zirconium Total | 6042 | Aqua Regia by ICPMS | These were sediment/soils digested with Aqua-Regia (Nitric Acid & Hydrochloric Acids) and then analysis by ICPMS. Results would be reported in ppm s. | 0.01 | Y | Y | Y | 71 | 5/14/2022 | 7/27/2024 | 7/29/2024 9:13:08 AM | 2022-06-28 |  | EMS |  | DimensionlessRatio | ppm | N | (ratio) |  | Zirconium Total (ratio) | 7440-67-7 | Metals - Total; | Metals - Total; EMS code: ZR-T | Chemical | Total | LAB
ZR-T | ug/g | Zirconium Total | X551 | Metals in Biota (dry), total, acid digestion, ICP-SFMS | 0.1 gram of dried, homogenized biota is weighed in a microwave TFM vessel, and 8mL of Nitric Acid and 1mL of Hydrogen Peroxide are added. The sample is digested at 200 degrees C for 15 minutes in a closed vessel high pressure microwave oven. The sample is cooled, the volume adjusted to 100mL and introduced to an iductively coupled argon plasma-sector field mass spectrometer (ICP-SFMS) by pneumatic nebulization. Each respective element is measured at a specific mass to charge ratio, m/z value expressed in amu (atomic mass units), and at a specified instrument resolution. LR: Low Resolution ~300, MR: Mid Resolution ~4000 and HR: Hich Resolution ~10000. Ref: Method number 02-2705, National Laboratory for Environmental Testing, CCIW, Environment Canada, Burlington. Requested April 2003. | 0.2 | Y | Y | Y | 43 | 8/11/2015 | 10/30/2023 | 11/28/2023 2:21:33 PM | 2015-03-11 |  | RDGIBSON |  | MassPerMass | mg/kg | Y | (mass conc.) |  | Zirconium Total (mass conc.) | 7440-67-7 | Metals - Total; | Metals - Total; EMS code: ZR-T | Chemical | Total | LAB
ZR-T | mg/L | Zirconium Total | F038 | HNO3/HCl dig,ICP-AES | Dissolved and extractable water samples are ready for analysis once collected. Total samples are autoclaved with HNO3 and hydrochloric acid (HCI) in capped Teflon vials. Analysis is performed via Inductively Coupled Argon Plasma Atomic Emission Spectrometry (ICP-AES)(simultaneous multielement analysis), Graphite Furnace Atomic Absorption Spectrometry (GF-AAS) (low level analysis) and/or Flame Atomic Absorption Spectrometry (F-AAS). | 0.01 | Y | Y | Y | 1 | 3/22/2002 | 3/22/2002 | 2/13/2004 1:01:35 AM | 1998-05-21 | 6/3/1998 6:51:49 PM | NPEPPIN |  | Concentration | mg/L | Y | (fl. conc.) |  | Zirconium Total (fl. conc.) | 7440-67-7 | Metals - Total; | Metals - Total; EMS code: ZR-T | Chemical | Total | LAB
ZR-T | mg/L | Zirconium Total | MSEA | Metals in Seawater (undigested) by (CRC) ICP-MS | This method describes the multi-elemental determination of trace elements by Collision/Reaction Cell (CRC) ICP-MS. The method measures ions produced by a radio-frequency inductively coupled plasma. Analyte species originating in a liquid are pneumatically nebulized and the resulting aerosol transported by argon gas into the plasma torch. The ions produced are entrained in the plasma gas and extracted, by means of a differentially pumped vacuum interface, into a mass spectrometer. The ions flow through a Collision/Reaction Cell (CRC) where depending on the mode of analysis either the ions passes through the cell unhindered (no gas - classical ICP-MS mode), or collisions/reactions chemistry takes place between the ions and a neutral gas such as hydrogen or helium. Any polyatomic ions which give interference on the analytes are removed. The ions are then sorted according to their mass-to-charge ratios by a quadrupole mass spectrometer. The ions are then sorted according to their mass-to-charge ratios by a quadrupole mass spectrometer. The ions transmitted through the quadrupole are quantified by a channel electron multiplier and the ion information processed by a data handling system. Although most polyatomic interference is removed, any additional interference must be assessed and valid corrections applied or the data flagged to indicate problems. Interference correction must include compensation for background ions contributed by the plasma gas, reagents, and constituents of the sample matrix. Instrument drift as well as suppressions or enhancements of instrument response caused by the sample matrix must be corrected for by the use of internal standards. | 8e-05 | Y | Y | Y | 35 | 2/15/2011 3:58:00 PM | 4/29/2024 | 6/7/2024 10:03:16 AM | 2009-09-08 |  | RDGIBSON |  | Concentration | mg/L | Y | (fl. conc.) |  | Zirconium Total (fl. conc.) | 7440-67-7 | Metals - Total; | Metals - Total; EMS code: ZR-T | Chemical | Total | LAB
ZR-T | ug/m3 | Zirconium Total | 5312 | Tef:N/P Dig:ICP Analysis | A nitric/perchloric acid digestion if used to bring the metals into solution prior to ICP analysis. For Hi-Vol samples preliminary results are reported as mg/L and 24 hour loading results as ug/m3. | 0.004 | Y | Y | Y | 1191 | 11/7/1997 | 5/15/2000 | 11/18/2000 1:29:41 AM | 1996-10-13 | 6/3/1998 6:51:46 PM | EMS |  | AirConcentration | µg/m³ | Y | (air conc.) |  | Zirconium Total (air conc.) | 7440-67-7 | Metals - Total; | Metals - Total; EMS code: ZR-T | Chemical | Total | LAB
ZR-T | mg/L | Zirconium Total | X561 | Total metals: FA, HNO3, HR-ICPMS (No Prep)Prep) | Ultra trace metals in water are analyzed by high resolution inductively coupled plasma mass spectrometry (HR-ICPMS), modified from US EPA 200.8, (Revision 5.5). The detection limits provided can only be met for undigested samples. This procedure is intended for colorless, non-turbid, acid-preserved water samples (i.e. pristine water samples), having turbidity < 1 NTU and no odor. Where turbidity exceeds 1 NTU, and/or the sample is colored and has an odor, results may be biased low compared to true Total Metals concentrations. ALS recommends that turbidity analysis be requested on samples submitted for this test to aid with interpretation of results. | 5e-05 | Y | Y | Y | 20 | 10/20/2015 | 10/5/2023 | 1/31/2024 2:37:41 PM | 2016-03-29 |  | RDGIBSON |  | Concentration | mg/L | Y | (fl. conc.) |  | Zirconium Total (fl. conc.) | 7440-67-7 | Metals - Total; | Metals - Total; EMS code: ZR-T | Chemical | Total | LAB
ZR-T | ppm (S) | Zirconium Total | MS41 | ALS Minerals ME-MS41 | ALS Minerals ME-MS41. A prepared sample is digested with aqua regia in a graphite heating block. After cooling; the resulting solution is diluted to with deionized water; mixed and analyzed by ICP-AES. Samples are then analyzed by ICP-MS for the remaining suite of elements. | 0.5 | Y | Y | Y | 0 |  |  |  | 2024-01-29 |  | EMS |  | DimensionlessRatio | ppm | N | (ratio) |  | Zirconium Total (ratio) | 7440-67-7 | Metals - Total; | Metals - Total; EMS code: ZR-T | Chemical | Total | LAB
ZR-T | mg/L | Zirconium Total | 1349 | Tot Metals in Seawater by CRC ICPMS (BC) | Tot Metals in Seawater by CRC ICPMS (BC) | 0.0005 | Y | Y | Y | 98 | 8/16/2019 | 12/5/2021 | 12/13/2021 12:34:17 PM | 2019-09-30 |  | KATFRASE |  | Concentration | mg/L | Y | (fl. conc.) |  | Zirconium Total (fl. conc.) | 7440-67-7 | Metals - Total; | Metals - Total; EMS code: ZR-T | Chemical | Total | LAB
ZR-T | mg/L | Zirconium Total | X170 | ICP | This code is a catch-all for a large variety of ICP methods including water and also solids digested into solutions as an intermediate result. Different procedures will have greatly different detection limits. (a) For air particulates, preliminary results are reported in mg/L and loading results as mg/dm2/d. (b) For Hi Vol samples preliminiary results are same as reported in mg/L and loading results as ug/m3. (c) Note that for Rain samples there is no filtration nor digestion steps. | 0.003 | Y | Y | Y | 170 |  |  | 10/24/2022 10:58:25 AM | 1996-04-12 | 6/3/1998 6:51:41 PM | EMS |  | Concentration | mg/L | Y | (fl. conc.) |  | Zirconium Total (fl. conc.) | 7440-67-7 | Metals - Total; | Metals - Total; EMS code: ZR-T | Chemical | Total | LAB
ZRTCLP | mg/L | Zirconium TCLP (Toxicity Characteristic Leaching Procedure) | FS82 | PBM,dry60C,10mesh,SALM-ICP-MS | PBM;dry60C;10mesh;SALM-ICP-MS | 10 | Y | Y | Y | 0 |  |  |  | 2024-01-24 |  | RDGIBSON |  | Concentration | mg/L | Y | (fl. conc.) |  | Zirconium TCLP (Toxicity Characteristic Leaching Procedure) (fl. conc.) |  | Metals - Extractable; | Metals - Extractable; EMS code: ZRTCLP | Chemical |  | LAB
Note on AE1: [Threaded comment] Your version of Excel allows you to read this threaded comment; however, any edits to it will get removed if the file is opened in a newer version of Excel. Learn more: https://go.microsoft.com/fwlink/?linkid=870924 Comment: Where data class = field result, device type = method short name