## Sheet1
Parm.Code | Unit | Parameter | Method.Code | Method | Method.Description | MDL | Dict.Active | Parameter.Active | Method.Active | Results | First.Used | Last.Used | Last.Upload | WHEN_CREATED | WHEN_UPDATED | WHO_CREATED | WHO_UPDATED | Sample.Unit.Group | Sample.Unit | Convertable.In.Samples | Sample.Unit.Group.Short.Name | Modifier | NewNameID | CAS | OP_Group | Description | Analysis.Type | Fraction | Classification | Device Type
0001 | Col. Unit | Color Apparent | EQ06 | EQUIS (Col units) | EQUIS (Col units) | 5 | N | Y | Y | 4164 |  |  | 10/22/1996 11:26:55 PM | 1996-04-12 | 6/3/1998 6:51:32 PM | EMS |  | Apperance | Col. Unit | N | (app.) |  | Color Apparent (app.) |  | Physical Properties; | Physical Properties; EMS code: 0001 | Physical |  | LAB
0001 | Col. Unit | Color Apparent | X152 | Visual Comparison to glass disks | Colour is determined by visual comparison of the sample with standard colour discs (Hellige colour Disc No.611-10, APHA Hazen platinum cobalt scale or equivalent). One unit of colour is that produced by 1.0 mg/L platinum in the form of the chloroplatinate ion. Sample centrifuged or filtered for True Color. | 5 | Y | Y | Y | 12769 | 1/2/1997 2:00:00 PM | 3/24/2012 3:10:00 AM | 3/28/2012 10:45:29 AM | 1996-04-12 | 6/3/1998 6:51:32 PM | EMS |  | Apperance | Col. Unit | N | (app.) |  | Color Apparent (app.) |  | Physical Properties; | Physical Properties; EMS code: 0001 | Physical |  | LAB
0001 | Col. Unit | Color Apparent | XM14 | Spectrophotometer | The pH of a filtered sample is adjusted to pH 7.6. Then the TAC is measured by integrating the absorbance of the sample between 400 and 700 nm. One unit of TAC is defined as the colour produced by 2 mg/L platinum in the form of the chloroplatinate ion. | 5 | Y | Y | Y | 2300 | 2/17/2005 | 10/5/2012 5:43:00 PM | 2/6/2019 3:01:07 PM | 1996-04-12 | 6/3/1998 6:51:32 PM | EMS |  | Apperance | Col. Unit | N | (app.) |  | Color Apparent (app.) |  | Physical Properties; | Physical Properties; EMS code: 0001 | Physical |  | LAB
0001 | Col. Unit | Color Apparent | X118 | Colour Comparator | Colour Comparator | 5 | Y | Y | Y | 355 | 3/22/2002 6:15:00 PM | 12/12/2011 12:02:00 PM | 2/9/2012 12:15:56 PM | 1996-04-12 | 6/3/1998 6:51:32 PM | EMS |  | Apperance | Col. Unit | N | (app.) |  | Color Apparent (app.) |  | Physical Properties; | Physical Properties; EMS code: 0001 | Physical |  | LAB
0002 | Col. Unit | Color True | X106 | Lab Cent,Visual Compariso | Lab Cent;Visual Comparison | 5 | Y | Y | Y | 23872 | 10/17/1996 7:41:00 AM | 5/21/2015 5:58:00 PM | 10/26/2017 9:58:33 AM | 1996-04-12 | 6/3/1998 6:51:32 PM | EMS |  | Apperance | Col. Unit | N | (app.) |  | Color True (app.) |  | Physical Properties; | Physical Properties; EMS code: 0002 | Physical |  | LAB
0002 | Col. Unit | Color True | 1301 | Filt,Colour Comparator | Filt;Colour Comparator | 5 | Y | Y | Y | 2144 | 5/5/1998 | 8/17/2024 4:56:00 PM | 8/23/2024 5:01:25 PM | 1996-04-12 | 6/3/1998 6:51:32 PM | EMS |  | Apperance | Col. Unit | N | (app.) |  | Color True (app.) |  | Physical Properties; | Physical Properties; EMS code: 0002 | Physical |  | LAB
0002 | Col. Unit | Color True | X118 | Colour Comparator | Colour Comparator | 5 | Y | Y | Y | 17 |  |  | 2/27/2001 12:50:43 PM | 1996-04-12 | 6/3/1998 6:51:32 PM | EMS |  | Apperance | Col. Unit | N | (app.) |  | Color True (app.) |  | Physical Properties; | Physical Properties; EMS code: 0002 | Physical |  | LAB
0002 | Col. Unit | Color True | X152 | Visual Comparison to glass disks | Colour is determined by visual comparison of the sample with standard colour discs (Hellige colour Disc No.611-10, APHA Hazen platinum cobalt scale or equivalent). One unit of colour is that produced by 1.0 mg/L platinum in the form of the chloroplatinate ion. Sample centrifuged or filtered for True Color. | 5 | Y | Y | Y | 13177 | 8/14/1996 4:28:00 PM | 8/4/2009 8:05:00 AM | 4/12/2010 4:51:56 PM | 1996-04-12 | 6/3/1998 6:51:32 PM | EMS |  | Apperance | Col. Unit | N | (app.) |  | Color True (app.) |  | Physical Properties; | Physical Properties; EMS code: 0002 | Physical |  | LAB
0002 | Col. Unit | Color True | X321 | LF, comparison to coloured solutions | The sample is centrifuged or filtered to remove turbidity and colour is determined by visual comparison of the sample with known concentrations of coloured solutions. | 5 | Y | Y | Y | 4895 | 6/4/1997 | 9/4/2009 3:50:00 PM | 10/22/2009 3:31:19 PM | 1996-04-12 | 6/3/1998 6:51:32 PM | EMS |  | Apperance | Col. Unit | N | (app.) |  | Color True (app.) |  | Physical Properties; | Physical Properties; EMS code: 0002 | Physical |  | LAB
0002 | Col. Unit | Color True | EQ06 | EQUIS (Col units) | EQUIS (Col units) | 5 | N | Y | Y | 7067 |  |  | 10/24/1996 3:16:58 PM | 1996-04-12 | 6/3/1998 6:51:32 PM | EMS |  | Apperance | Col. Unit | N | (app.) |  | Color True (app.) |  | Physical Properties; | Physical Properties; EMS code: 0002 | Physical |  | LAB
0002 | Col. Unit | Color True | XM14 | Spectrophotometer | The pH of a filtered sample is adjusted to pH 7.6. Then the TAC is measured by integrating the absorbance of the sample between 400 and 700 nm. One unit of TAC is defined as the colour produced by 2 mg/L platinum in the form of the chloroplatinate ion. | 5 | Y | Y | Y | 22839 | 9/4/1996 | 8/20/2024 | 8/26/2024 3:51:07 PM | 1996-04-12 | 6/3/1998 6:51:32 PM | EMS |  | Apperance | Col. Unit | N | (app.) |  | Color True (app.) |  | Physical Properties; | Physical Properties; EMS code: 0002 | Physical |  | LAB
0003 | ug/m3 | Oil & Grease | EQ40 | EQUIS (mg/m3) | EQUIS (mg/m3) | 0 | N | Y | Y | 38 |  |  | 10/22/1996 7:18:31 PM | 1996-04-12 | 6/3/1998 6:51:32 PM | EMS |  | AirConcentration | µg/m³ | Y | (air conc.) |  | Oil & Grease (air conc.) |  | Petroleum Hydrocarbons; | Petroleum Hydrocarbons; EMS code: 0003 | Chemical |  | LAB
0003 | mg/L | Oil & Grease | X055 | Dir. Hexane Extraction | The sample is placed in a separatory funnel and extracted with hexane into a tared flask. The hexane is evaporated from the extract and the residue is dried at 105C, cooled and weighed. | 1 | Y | Y | Y | 1234 | 11/24/1995 | 8/17/2024 | 8/19/2024 4:18:06 PM | 1996-04-12 | 6/3/1998 6:51:32 PM | EMS |  | Concentration | mg/L | Y | (fl. conc.) |  | Oil & Grease (fl. conc.) |  | Petroleum Hydrocarbons; | Petroleum Hydrocarbons; EMS code: 0003 | Chemical |  | LAB
0003 | ug/g | Oil & Grease | 0961 | Ultrasonic CH2Cl2 Extr:G | Ultrasonic CH2Cl2 Extr:G | 100 | Y | Y | Y | 350 | 10/18/2002 12:15:00 PM | 9/2/2014 | 7/19/2024 8:20:35 AM | 1996-04-12 | 6/3/1998 6:51:45 PM | EMS |  | MassPerMass | mg/kg | Y | (mass conc.) |  | Oil & Grease (mass conc.) |  | Petroleum Hydrocarbons; | Petroleum Hydrocarbons; EMS code: 0003 | Chemical |  | LAB
0003 | mg/L | Oil & Grease | X083 | Freon Extraction | Freon Extraction | 1 | Y | Y | Y | 58 |  |  | 10/24/1996 12:39:30 PM | 1996-04-12 | 6/3/1998 6:51:45 PM | EMS |  | Concentration | mg/L | Y | (fl. conc.) |  | Oil & Grease (fl. conc.) |  | Petroleum Hydrocarbons; | Petroleum Hydrocarbons; EMS code: 0003 | Chemical |  | LAB
0003 | ug/g | Oil & Grease | E567 | BC MOE Lab Manual (Oil & Grease) (mod) | BC MOE Lab Manual (Oil & Grease) (mod). A subsample of the sediment/soil is extracted with 1:1 hexane:acetone using a rotary extraction apparatus. The extract is analyzed gravimetrically. | 500 | Y | Y | Y | 0 |  |  |  | 2024-01-24 |  | RDGIBSON |  | MassPerMass | mg/kg | Y | (mass conc.) |  | Oil & Grease (mass conc.) |  | Petroleum Hydrocarbons; | Petroleum Hydrocarbons; EMS code: 0003 | Chemical |  | LAB
0003 | mg/L | Oil & Grease | X372 | Comb.Ext.Hexane | Particulates are separated from the sample by filtration and then extracted with hexane using a Soxhlet extractor. The liquid phase is then extracted with hexane and this extract is subsequently combined with the Soxhlet extract. The hexane is evaporated from the combined extract and the residue is weighed. | 1 | Y | Y | Y | 1403 | 4/1/1997 | 8/19/2024 11:19:00 AM | 8/22/2024 3:35:15 PM | 1996-04-12 | 6/3/1998 6:51:46 PM | EMS |  | Concentration | mg/L | Y | (fl. conc.) |  | Oil & Grease (fl. conc.) |  | Petroleum Hydrocarbons; | Petroleum Hydrocarbons; EMS code: 0003 | Chemical |  | LAB
0003 | mg/L | Oil & Grease | X046 | Pet. Ether Extraction | Pet. Ether Extraction | 1 | Y | Y | Y | 13 |  |  | 8/17/1999 10:12:39 PM | 1996-04-12 | 6/3/1998 6:51:32 PM | EMS |  | Concentration | mg/L | Y | (fl. conc.) |  | Oil & Grease (fl. conc.) |  | Petroleum Hydrocarbons; | Petroleum Hydrocarbons; EMS code: 0003 | Chemical |  | LAB
0003 | ug/g | Oil & Grease | FS07 | Freon 113 Extr,Infrared | Freon 113 soluble compounds are extracted from an acidified effluent sample and made up to a measured volume. Sediment samples are acidified then extracted using a wrist-action shaker. The "Oils and Greases" content of the Freon extracts is determined using an Infrared Spectrophotometer with commercial quantitation software. Quantitation is based on the comparison of the Absorbance difference between samples and a synthetic oil standard. As there is no sample Absorbance band at 4000 cm-1 it represents a useful reference value to compensate for minor differences between spectrophotometer quartz cells. Sample Absorbance at 2935 cm-1 represents the C-H stretching frequency of the aliphatic hydrocarbons of which typical environmental "Oils and Greases" are comprised. "Hydrocarbons" are quantitated in the same manner after treating the Freon 113 extract with Silica Gel. | 5 | Y | Y | Y | 179 | 6/30/1998 3:45:00 PM | 1/4/2002 8:07:00 AM | 5/26/2003 1:56:27 PM | 1996-07-16 | 6/3/1998 6:51:44 PM | MPAK |  | MassPerMass | mg/kg | Y | (mass conc.) |  | Oil & Grease (mass conc.) |  | Petroleum Hydrocarbons; | Petroleum Hydrocarbons; EMS code: 0003 | Chemical |  | LAB
0003 | mg/L | Oil & Grease | 1000 | Comb.Ext.Pet.Ether | Comb.Ext.Pet.Ether | 1 | Y | Y | Y | 9 |  |  | 10/24/1996 11:53:06 AM | 1996-04-12 | 6/3/1998 6:51:32 PM | EMS |  | Concentration | mg/L | Y | (fl. conc.) |  | Oil & Grease (fl. conc.) |  | Petroleum Hydrocarbons; | Petroleum Hydrocarbons; EMS code: 0003 | Chemical |  | LAB
0003 | mg/L | Oil & Grease | F020 | Freon 113 Extr,Infrared | Freon 113 soluble compounds are extracted from an acidified effluent sample and made up to a measured volume. Sediment samples are acidified then extracted using a wrist-action shaker. The "Oils and Greases" content of the Freon extracts is determined using an Infrared Spectrophotometer with commercial quantitation software. Quantitation is based on the comparison of the Absorbance difference between samples and a synthetic oil standard. As there is no sample Absorbance band at 4000 cm-1 it represents a useful reference value to compensate for minor differences between spectrophotometer quartz cells. Sample Absorbance at 2935 cm-1 represents the C-H stretching frequency of the aliphatic hydrcarbons of which typical environmental "Oils and Greases" are comprised. "Hydrocarbons" are quantitated in the same manner after treating the Freon 113 extract with Silica Gel. | 0.2 | Y | Y | Y | 161 | 9/10/1996 10:30:00 AM | 10/26/2001 1:55:00 PM | 5/20/2003 4:08:44 PM | 1996-07-16 | 6/3/1998 6:51:44 PM | MPAK |  | Concentration | mg/L | Y | (fl. conc.) |  | Oil & Grease (fl. conc.) |  | Petroleum Hydrocarbons; | Petroleum Hydrocarbons; EMS code: 0003 | Chemical |  | LAB
0003 | mg/L | Oil & Grease | EQ01 | EQUIS (mg/L) | EQUIS (mg/L) | 1 | N | Y | Y | 1457 |  |  | 10/23/1996 12:07:05 AM | 1996-04-12 | 6/3/1998 6:51:32 PM | EMS |  | Concentration | mg/L | Y | (fl. conc.) |  | Oil & Grease (fl. conc.) |  | Petroleum Hydrocarbons; | Petroleum Hydrocarbons; EMS code: 0003 | Chemical |  | LAB
0004 | pH units | pH | X330 | Meter | The conductance (= conductivity) of a sample is measured by use of a self-contained conductivity meter. pH is measured with a combined glass/reference (Ag/AgCl) electrode. For both conductivity and pH the measurements are made at 25 C, or the results are temperature-corrected to 25 C. Note that laboratory and field measurements are distinguished by the EMS Analytical Agency codes. Ex situ and in situ measurements are distinguished by the computer codes. | 2 | Y | Y | Y | 398621 | 9/29/1981 12:00:00 PM | 8/22/2024 12:44:00 PM | 8/26/2024 4:53:05 PM | 1996-04-17 | 12/20/2019 11:02:20 AM | EMS | RDGIBSON | Acidity | pH Units | Y | (acidity) |  | pH (acidity) |  | Physical Properties; | Physical Properties; EMS code: 0004 | Physical |  | FIELD_RESULT
0004 | pH units | pH | XM30 | Meter: Glass/Ref Regular Ionic Strength | A small polypropylene beaker containing about 40 mL of sample is placed in a constant-temperature water bath until it attains 25.0(+/-)0.5 C. The beaker is then loaded onto a sample tray and the pH is measured potentiometrically using a combined glass/reference electrode (Ag/AgCl as the reference half-cell) that has been calibrated against two certified reference buffer solutions; the calibration standards are chosen to bracket the sample pH as closely as possible. This method is not appropriate for low ionic strength solutions nor for solutions having pH below pH2 nor above pH 12. | 2 | Y | Y | Y | 24565 | 3/19/1996 3:47:00 PM | 6/26/2014 10:29:00 AM | 9/11/2019 8:47:49 AM | 1997-05-01 | 6/3/1998 6:51:46 PM | MPAK |  | Acidity | pH Units | Y | (acidity) |  | pH (acidity) |  | Physical Properties; | Physical Properties; EMS code: 0004 | Physical |  | LAB
0004 | pH units | pH | 5065 | Meter: Glass/Ref Low Ionic Strength | Low ionic strength glass electrode (Radiometer PHC 2701 or equivalent). Calibrate with low ionic strength buffers. Measurement accuracy and precision to be +- 0.01 pH unit or better. Measurements made at 25 C or results are temperature corrected to 25 C. Code "PH--" must be used for precipitation and code "0004" must be used for surface waters. |  | Y | Y | Y | 67341 | 7/17/1998 3:05:00 PM | 5/3/2025 12:21:00 AM | 8/26/2024 8:06:03 AM | 1996-04-17 | 6/3/1998 6:51:44 PM | EMS |  | Acidity | pH Units | Y | (acidity) |  | pH (acidity) |  | Physical Properties; | Physical Properties; EMS code: 0004 | Physical |  | LAB
0004 | pH units | pH | FS62 | pH@25C 1:2 Soil:H2O susp. | 1:2 soil: water mixture is placed in a constant temperature bath at 25 C and the pH is measured potentiometrically. | 2 | Y | Y | Y | 622 | 3/10/2009 | 7/28/2024 | 7/29/2024 5:17:09 PM | 1998-11-05 |  | NPEPPIN |  | Acidity | pH Units | Y | (acidity) |  | pH (acidity) |  | Physical Properties; | Physical Properties; EMS code: 0004 | Physical |  | LAB
0004 | pH units | pH | 1331 | pH by Meter (Automated) (seawater) | pH by Meter (Automated) (seawater) | 0.1 | Y | Y | Y | 755 | 8/14/2019 | 6/30/2024 | 7/11/2024 10:54:53 AM | 2019-09-30 |  | KATFRASE |  | Acidity | pH Units | Y | (acidity) |  | pH (acidity) |  | Physical Properties; | Physical Properties; EMS code: 0004 | Physical |  | LAB
0004 | pH units | pH | VOL11 | ALS Minerals OA-VOL11 | ALS Minerals OA-VOL11. Net Acid Generation (NAG) provides a quantitative estimation of the acid that can be generated by mine waste. Hydrogen peroxide is used to rapidly oxidize sulfides. NAG is reported in kg H2SO4/tonne at pH 4.5 and 7.0. | 0.1 | Y | Y | Y | 0 |  |  |  | 2024-01-29 |  | EMS |  | Acidity | pH Units | Y | (acidity) |  | pH (acidity) |  | Physical Properties; | Physical Properties; EMS code: 0004 | Physical |  | LAB
0004 | pH units | pH | ELE07 | ALS Minerals OA-ELE07 | ALS Minerals OA-ELE07. Sample mix with water in 1:1 ratio to make a paste; pH is then determined. | 0.1 | Y | Y | Y | 0 |  |  |  | 2024-01-29 |  | RDGIBSON |  | Acidity | pH Units | Y | (acidity) |  | pH (acidity) |  | Physical Properties; | Physical Properties; EMS code: 0004 | Physical |  | LAB
0004 | pH units | pH | FS18 | pH@25C 1:2Soil:0.01MCaCl2 susp. | 1:2 soil:0.01 M CaCl2 mixture is placed in a constant temperature bath at 25C and the pH is measured potentiometrically. Soil pH in soil:neutral salt suspension measures both active and exchangeable acidity. | 2 | Y | Y | Y | 37 | 10/29/1998 2:30:00 PM | 8/28/2002 1:40:00 PM | 5/26/2003 2:26:58 PM | 1998-05-01 | 6/3/1998 6:51:48 PM | NPEPPIN |  | Acidity | pH Units | Y | (acidity) |  | pH (acidity) |  | Physical Properties; | Physical Properties; EMS code: 0004 | Physical |  | LAB
0004 | pH units | pH | X003 | Paper | Paper | 1 | Y | Y | Y | 39 |  |  | 2/13/1997 2:21:33 PM | 1996-04-17 | 6/3/1998 6:51:44 PM | EMS |  | Acidity | pH Units | Y | (acidity) |  | pH (acidity) |  | Physical Properties; | Physical Properties; EMS code: 0004 | Physical |  | LAB
0004 | pH units | pH | FS17 | pH@25C 1:1Soil:H2O susp. | 1:1 soil:water mixture is placed in a constant temperature bath at 25C and the pH is measured potentiometrically. Soil pH in soil:water suspension measures active acidity. | 2 | Y | Y | Y | 54 | 12/7/2009 12:29:00 PM | 3/14/2016 1:58:00 PM | 3/24/2016 5:03:31 PM | 1998-05-01 | 6/3/1998 6:51:48 PM | NPEPPIN |  | Acidity | pH Units | Y | (acidity) |  | pH (acidity) |  | Physical Properties; | Physical Properties; EMS code: 0004 | Physical |  | LAB
0004 | pH units | pH | PH15 | pH @ 15C | APHA 4500-H pH is determined by potentiometric measurement with a pH electrode, and is conducted at 15 ? 1?C, and is used to calculate Un-Ionized Ammonia for the federal Wastewater Systems Effluent Regulation. | 0.1 | Y | Y | Y | 88 | 1/20/2022 | 7/2/2024 | 7/24/2024 4:32:03 PM | 2022-04-08 | 1/29/2024 10:26:44 AM | EMS | RDGIBSON | Acidity | pH Units | Y | (acidity) |  | pH (acidity) |  | Physical Properties; | Physical Properties; EMS code: 0004 | Physical |  | LAB
0004 | pH units | pH | PHKT | pH kit, LaMotte wide range (5.0-8.5) | pH kit; LaMotte wide range (5.0-8.5) | 5 | Y | Y | Y | 57 |  |  | 9/2/2011 3:37:09 PM | 2011-08-04 |  | RDGIBSON |  | Acidity | pH Units | Y | (acidity) |  | pH (acidity) |  | Physical Properties; | Physical Properties; EMS code: 0004 | Physical |  | LAB
0004 | pH units | pH | XH00 | Hydrolab | Hydrolab | 1 | Y | Y | Y | 16403 | 4/7/1997 1:25:00 PM | 12/16/1999 11:41:00 AM | 6/7/2024 5:30:18 PM | 1996-04-17 | 12/20/2019 11:02:20 AM | EMS | RDGIBSON | Acidity | pH Units | Y | (acidity) |  | pH (acidity) |  | Physical Properties; | Physical Properties; EMS code: 0004 | Physical |  | FIELD_RESULT | Hydrolab
0004 | pH units | pH | DBPFP | DBPFP Incubation | DBPFP Incubation | 0.1 | Y | Y | Y | 2 | 3/18/2023 4:37:00 PM | 5/16/2023 12:06:00 PM | 6/8/2023 12:14:56 PM | 2023-03-31 |  | EMS |  | Acidity | pH Units | Y | (acidity) |  | pH (acidity) |  | Physical Properties; | Physical Properties; EMS code: 0004 | Physical |  | LAB
0004 | pH units | pH | EQ05 | EQUIS (pH units) | EQUIS (pH units) | 1 | N | Y | Y | 47737 |  |  | 10/24/1996 3:17:07 PM | 1996-04-17 | 6/3/1998 6:51:44 PM | EMS |  | Acidity | pH Units | Y | (acidity) |  | pH (acidity) |  | Physical Properties; | Physical Properties; EMS code: 0004 | Physical |  | LAB
0004 | pH units | pH | F072 | 25C,Low Ionic Strength | 25C Low Ionic Strength Electrode; measurements made at 25 +- .5 C. | 2 | Y | Y | Y | 10 |  |  | 2/5/2001 4:22:00 PM | 1999-07-15 |  | NPEPPIN |  | Acidity | pH Units | Y | (acidity) |  | pH (acidity) |  | Physical Properties; | Physical Properties; EMS code: 0004 | Physical |  | LAB
0004 | pH units | pH | XM00 | Meter-in situ | Meter-in situ | 1 | Y | Y | Y | 90573 | 8/14/1998 12:59:00 PM | 8/7/2024 3:00:00 PM | 8/20/2024 7:35:41 AM | 1996-04-17 | 10/16/2012 4:12:05 PM | EMS | RDGIBSON | Acidity | pH Units | Y | (acidity) |  | pH (acidity) |  | Physical Properties; | Physical Properties; EMS code: 0004 | Physical |  | FIELD_RESULT | Meter-in situ
0004 | None | pH | AISP | Special Analysis | Special Analysis | 1 | N | Y | N | 2 |  |  | 10/24/1996 11:39:24 AM | 1996-04-17 | 6/3/1998 6:51:44 PM | EMS |  | None | None | N | (unitless) |  | pH (unitless) |  | Physical Properties; | Physical Properties; EMS code: 0004 | Physical |  | LAB
0004 | pH units | pH | XM15 | Continuous Rec.-in situ | Continuous Rec.-in situ | 1 | Y | Y | Y | 133548 | 4/30/1996 7:00:00 AM | 7/1/2024 | 8/22/2024 1:49:14 PM | 1996-04-17 | 12/20/2019 11:02:20 AM | EMS | RDGIBSON | Acidity | pH Units | Y | (acidity) |  | pH (acidity) |  | Physical Properties; | Physical Properties; EMS code: 0004 | Physical |  | FIELD_RESULT | Continuous Rec.-in situ
0005 | mg/L | Residue: Total (Total Solids) | X042 | Gravimetric 103C | A known volume of a well-mixed sample is transferred into an ignited and weighed porcelain dish, evaporated, dried and weighed. The difference in weights represents Total Residue. | 10 | Y | Y | Y | 1916 | 10/17/1997 6:32:00 PM | 12/22/2020 9:00:00 AM | 12/24/2020 2:37:05 PM | 1996-04-12 | 6/3/1998 6:51:45 PM | EMS |  | Concentration | mg/L | Y | (fl. conc.) |  | Residue: Total (Total Solids) (fl. conc.) |  | Physical Properties; | Physical Properties; EMS code: 0005 | Physical | Total | LAB
0005 | mg/L | Residue: Total (Total Solids) | X333 | Gravimetric, 180C | A well mixed sample is evaporated in a pre-weighed dish and dried to constant weight in the oven at 180 C. The increase in weight over the empty dish represents the total solids. | 1 | Y | Y | Y | 121 |  |  | 2/5/2001 4:22:01 PM | 1996-04-12 | 6/3/1998 6:51:32 PM | EMS |  | Concentration | mg/L | Y | (fl. conc.) |  | Residue: Total (Total Solids) (fl. conc.) |  | Physical Properties; | Physical Properties; EMS code: 0005 | Physical | Total | LAB
0005 | mg/L | Residue: Total (Total Solids) | CS01 | System Calculation | System Calculation (mg/L) |  | N | Y | N | 101028 |  |  | 8/26/2024 4:53:06 PM | 1996-04-12 | 6/3/1998 6:51:42 PM | EMS |  | Concentration | mg/L | Y | (fl. conc.) |  | Residue: Total (Total Solids) (fl. conc.) |  | Physical Properties; | Physical Properties; EMS code: 0005 | Physical | Total | LAB
0005 | mg/L | Residue: Total (Total Solids) | X048 | Gravimetric | Gravimetric | 0 | N | Y | Y | 151 |  |  | 10/24/1996 9:15:20 AM | 1996-10-13 | 6/3/1998 6:51:46 PM | EMS |  | Concentration | mg/L | Y | (fl. conc.) |  | Residue: Total (Total Solids) (fl. conc.) |  | Physical Properties; | Physical Properties; EMS code: 0005 | Physical | Total | LAB
0005 | mg/L | Residue: Total (Total Solids) | EQ01 | EQUIS (mg/L) | EQUIS (mg/L) | 0 | N | Y | Y | 12135 |  |  | 10/24/1996 3:16:57 PM | 1996-04-12 | 6/3/1998 6:51:32 PM | EMS |  | Concentration | mg/L | Y | (fl. conc.) |  | Residue: Total (Total Solids) (fl. conc.) |  | Physical Properties; | Physical Properties; EMS code: 0005 | Physical | Total | LAB
0005 | mg/L | Residue: Total (Total Solids) | X237 | Gravimetric | This code is a catch-all for a variety of gravimetric procedures. | 14 | Y | Y | Y | 3854 | 9/12/1996 1:04:00 PM | 8/1/1997 6:31:00 PM | 12/17/1997 5:20:24 PM | 1996-04-12 | 6/3/1998 6:51:32 PM | EMS |  | Concentration | mg/L | Y | (fl. conc.) |  | Residue: Total (Total Solids) (fl. conc.) |  | Physical Properties; | Physical Properties; EMS code: 0005 | Physical | Total | LAB
0005 | ug/g | Residue: Total (Total Solids) | CS02 | System Calculation | System Calculation (ug/g) |  | N | Y | N | 1051 |  |  | 8/13/2016 5:13:46 PM | 1996-04-12 | 6/3/1998 6:51:42 PM | EMS |  | MassPerMass | mg/kg | Y | (mass conc.) |  | Residue: Total (Total Solids) (mass conc.) |  | Physical Properties; | Physical Properties; EMS code: 0005 | Physical | Total | LAB
0006 | mg/L | Residue: Total Fixed (Total Solids; Fixed) | X237 | Gravimetric | This code is a catch-all for a variety of gravimetric procedures. | 1 | Y | Y | Y | 381 |  |  | 10/15/1997 1:04:34 PM | 1996-04-12 | 6/3/1998 6:51:32 PM | EMS |  | Concentration | mg/L | Y | (fl. conc.) |  | Residue: Total Fixed (Total Solids; Fixed) (fl. conc.) |  | Physical Properties; | Physical Properties; EMS code: 0006 | Physical | Total | LAB
0006 | mg/L | Residue: Total Fixed (Total Solids; Fixed) | EQ01 | EQUIS (mg/L) | EQUIS (mg/L) | 1 | N | Y | Y | 3330 |  |  | 10/24/1996 3:16:56 PM | 1996-04-12 | 6/3/1998 6:51:45 PM | EMS |  | Concentration | mg/L | Y | (fl. conc.) |  | Residue: Total Fixed (Total Solids; Fixed) (fl. conc.) |  | Physical Properties; | Physical Properties; EMS code: 0006 | Physical | Total | LAB
0006 | ug/g | Residue: Total Fixed (Total Solids; Fixed) | X479 | Ignition,550 C | Sediment Fixed Residue (SFR) represents the solids or ash remaining after ignition at 550 C. Sediment Volatile Residue (SVR) is the difference between the weight lost on ignition at 550 C and the weight of the sample dried at 105 C. SVR is a rough approximation of the amount of organic matter in the sediment. | 500 | Y | Y | Y | 2189 | 7/29/1997 6:31:00 PM | 8/29/2010 8:43:00 AM | 4/23/2012 2:34:28 PM | 1997-11-04 | 6/3/1998 6:51:47 PM | NPEPPIN |  | MassPerMass | mg/kg | Y | (mass conc.) |  | Residue: Total Fixed (Total Solids; Fixed) (mass conc.) |  | Physical Properties; | Physical Properties; EMS code: 0006 | Physical | Total | LAB
0006 | mg/L | Residue: Total Fixed (Total Solids; Fixed) | CS01 | System Calculation | System Calculation (mg/L) |  | N | Y | N | 125 |  |  | 9/10/2002 12:45:03 PM | 1996-04-12 | 6/3/1998 6:51:42 PM | EMS |  | Concentration | mg/L | Y | (fl. conc.) |  | Residue: Total Fixed (Total Solids; Fixed) (fl. conc.) |  | Physical Properties; | Physical Properties; EMS code: 0006 | Physical | Total | LAB
0007 | mg/L | Residue: Filterable 1.0u (TDS) | CS01 | System Calculation | System Calculation (mg/L) |  | N | Y | N | 222 |  |  | 12/11/2013 12:10:53 PM | 1996-04-12 | 6/3/1998 6:51:42 PM | EMS |  | Concentration | mg/L | Y | (fl. conc.) |  | Residue: Filterable 1.0u (TDS) (fl. conc.) |  | Physical Properties; | Physical Properties; EMS code: 0007 | Physical |  | LAB
0007 | mg/L | Residue: Filterable 1.0u (TDS) | X017 | Glass Fiber Filter | A well-mixed sample is filtered through a standard 0.1 um glass fibre filter. A measured portion of the filtrate is evaporated in a preweighed evaporating vessel and dried to constant weight at 180 C. The increase in dish weight represents the total dissolved solids. (The filtrate from residue, non-filterable may be used.) | 1 | Y | Y | Y | 21963 | 6/1/1998 | 3/25/2734 1:55:00 PM | 8/26/2024 2:41:32 PM | 1996-04-12 | 6/3/1998 6:51:32 PM | EMS |  | Concentration | mg/L | Y | (fl. conc.) |  | Residue: Filterable 1.0u (TDS) (fl. conc.) |  | Physical Properties; | Physical Properties; EMS code: 0007 | Physical |  | LAB
0007 | mg/L | Residue: Filterable 1.0u (TDS) | F120 | Grav,1.2u GF filter,180C | Gravimetric;1.2uGlass Fiber Filter; 180C | 1 | Y | Y | Y | 37892 |  |  | 8/20/2024 7:35:41 AM | 2002-03-08 |  | NPEPPIN |  | Concentration | mg/L | Y | (fl. conc.) |  | Residue: Filterable 1.0u (TDS) (fl. conc.) |  | Physical Properties; | Physical Properties; EMS code: 0007 | Physical |  | LAB
0007 | mg/L | Residue: Filterable 1.0u (TDS) | 1032 | Gravimetric 1.2um GF filter, 103-5 C | A well-mixed volume of the sample is filtered through a Whatman GF/C glass fibre filter (1.2um particle retention). The filtrate is collected in a flask and transferred to a clean, pre-weighed porcelain dish that has been muffled at 550C for 20 minutes; the dish containing the filtrate is then placed overnight in an oven set at 90-95C to permit the liquid to evaporate. The dish with the filterable residue is dried at 103C for one hour and weighed again to determine the total filterable residue(FR). | 5 | Y | Y | Y | 21731 | 4/11/1996 | 11/5/2038 | 8/23/2024 5:01:25 PM | 1996-04-12 | 6/3/1998 6:51:41 PM | EMS |  | Concentration | mg/L | Y | (fl. conc.) |  | Residue: Filterable 1.0u (TDS) (fl. conc.) |  | Physical Properties; | Physical Properties; EMS code: 0007 | Physical |  | LAB
0007 | mg/L | Residue: Filterable 1.0u (TDS) | X043 | Calc, TRes-NFRes | Calc; TRes-NFRes | 1 | Y | Y | Y | 218 |  |  | 1/28/1998 11:33:48 AM | 1996-04-12 | 6/3/1998 6:51:45 PM | EMS |  | Concentration | mg/L | Y | (fl. conc.) |  | Residue: Filterable 1.0u (TDS) (fl. conc.) |  | Physical Properties; | Physical Properties; EMS code: 0007 | Physical |  | LAB
0007 | mg/L | Residue: Filterable 1.0u (TDS) | XCAL | Calculation (mg/L) | Calculation (mg/L) | 1 | Y | Y | Y | 2290 | 3/3/2003 | 8/13/2024 | 8/19/2024 2:53:10 PM | 1996-04-12 | 6/3/1998 6:51:41 PM | EMS |  | Concentration | mg/L | Y | (fl. conc.) |  | Residue: Filterable 1.0u (TDS) (fl. conc.) |  | Physical Properties; | Physical Properties; EMS code: 0007 | Physical |  | LAB
0007 | mg/L | Residue: Filterable 1.0u (TDS) | EQ01 | EQUIS (mg/L) | EQUIS (mg/L) | 1 | N | Y | Y | 10472 |  |  | 10/24/1996 3:16:58 PM | 1996-04-12 | 6/3/1998 6:51:32 PM | EMS |  | Concentration | mg/L | Y | (fl. conc.) |  | Residue: Filterable 1.0u (TDS) (fl. conc.) |  | Physical Properties; | Physical Properties; EMS code: 0007 | Physical |  | LAB
0007 | mg/L | Residue: Filterable 1.0u (TDS) | X026 | Gravimetric 0.45u Filter | A well-mixed sample is filtered through a standard membrane filter. A measured portion of the filtrate is evaporated in a preweighed evaporating vessel and dried to constant weight at 180 C. The increase in dish weight represents the total dissolved solids. (The filtrate from residue, non-filterable may be used.) | 1 | Y | Y | Y | 40236 | 12/12/1996 | 8/24/2024 | 8/26/2024 4:53:05 PM | 1996-04-12 | 6/3/1998 6:51:32 PM | EMS |  | Concentration | mg/L | Y | (fl. conc.) |  | Residue: Filterable 1.0u (TDS) (fl. conc.) |  | Physical Properties; | Physical Properties; EMS code: 0007 | Physical |  | LAB
0007 | mg/L | Residue: Filterable 1.0u (TDS) | X518 | Homo,GFor mf,105,grav | Homogenize if D+G; GF/C or membrance filter;dry@105C;gravimetric |  | Y | Y | Y | 2196 |  |  | 10/15/1999 1:06:35 PM | 1999-04-22 |  | NPEPPIN |  | Concentration | mg/L | Y | (fl. conc.) |  | Residue: Filterable 1.0u (TDS) (fl. conc.) |  | Physical Properties; | Physical Properties; EMS code: 0007 | Physical |  | LAB
0008 | mg/L | Residue: Non-filterable (TSS) | CS01 | System Calculation | System Calculation (mg/L) |  | N | Y | N | 254 |  |  | 8/13/2016 5:14:00 PM | 1996-04-12 | 6/3/1998 6:51:42 PM | EMS |  | Concentration | mg/L | Y | (fl. conc.) |  | Residue: Non-filterable (TSS) (fl. conc.) |  | Physical Properties; | Physical Properties; EMS code: 0008 | Physical |  | LAB
0008 | mg/L | Residue: Non-filterable (TSS) | F042 | Whole Bottle, Grav., whole bottle, 1.2 um GF filter, 103-5C | The entire bottle sample is filtered through a Whatman GF/C glass fibre filter (1.2 u particle retention) that has been muffled at 550C for 20 minutes and pre-weighed. The filter with the residue is then dried at 103C for one hour and weighed once again to constant weight; the resulting weight differences gives the total non-filterable residue (NFR). | 5 | Y | Y | Y | 7166 | 4/23/1999 | 8/31/2005 | 1/30/2007 1:02:16 AM | 1997-05-08 | 6/3/1998 6:51:46 PM | MPAK |  | Concentration | mg/L | Y | (fl. conc.) |  | Residue: Non-filterable (TSS) (fl. conc.) |  | Physical Properties; | Physical Properties; EMS code: 0008 | Physical |  | LAB
0008 | mg/L | Residue: Non-filterable (TSS) | X049 | Muffle Furnace 600C | Muffle Furnace 600C | 1 | Y | Y | Y | 2302 |  |  | 10/24/1996 2:18:17 AM | 1996-04-12 | 6/3/1998 6:51:32 PM | EMS |  | Concentration | mg/L | Y | (fl. conc.) |  | Residue: Non-filterable (TSS) (fl. conc.) |  | Physical Properties; | Physical Properties; EMS code: 0008 | Physical |  | LAB
0008 | mg/L | Residue: Non-filterable (TSS) | X332 | Grav, Whole bottle 105C | The entire sample is filtered, with rinsing, through a pre-weighed glass fibre filter discs (Whatman 934-AH, or equivalent), and the residue on the filter is dried to constant weight at 103 - 105 C. The increase in weight of the filter is reported as nonfilterable residue. The filtrate may be used for residue, filterable. | 4 | Y | Y | Y | 37647 | 5/3/1995 | 7/1/2024 | 8/26/2024 2:41:32 PM | 1996-04-12 | 6/3/1998 6:51:32 PM | EMS |  | Concentration | mg/L | Y | (fl. conc.) |  | Residue: Non-filterable (TSS) (fl. conc.) |  | Physical Properties; | Physical Properties; EMS code: 0008 | Physical |  | LAB
0008 | mg/L | Residue: Non-filterable (TSS) | X035 | Cent,Glass Fiber Filt105 | Cent;Glass Fiber Filt105 | 1 | Y | Y | Y | 19691 | 1/26/2001 | 10/31/2021 6:59:00 AM | 7/29/2024 11:17:51 AM | 1996-04-12 | 6/3/1998 6:51:32 PM | EMS |  | Concentration | mg/L | Y | (fl. conc.) |  | Residue: Non-filterable (TSS) (fl. conc.) |  | Physical Properties; | Physical Properties; EMS code: 0008 | Physical |  | LAB
0008 | t/d | Residue: Non-filterable (TSS) | EQ30 | EQUIS (T/d) | EQUIS (T/d) | 0 | N | Y | Y | 579 |  |  | 10/23/1996 12:03:16 AM | 1996-04-12 | 6/3/1998 6:51:32 PM | EMS |  | MassFlow | t/d | Y | (mass flow) |  | Residue: Non-filterable (TSS) (mass flow) |  | Physical Properties; | Physical Properties; EMS code: 0008 | Physical |  | LAB
0008 | mg/L | Residue: Non-filterable (TSS) | X017 | Glass Fiber Filter | A well-mixed sample is filtered through a standard 0.1 um glass fibre filter. A measured portion of the filtrate is evaporated in a preweighed evaporating vessel and dried to constant weight at 180 C. The increase in dish weight represents the total dissolved solids. (The filtrate from residue, non-filterable may be used.) | 1 | Y | Y | Y | 154046 | 9/2/1996 | 7/1/2024 | 8/23/2024 8:10:36 AM | 1996-04-12 | 6/3/1998 6:51:32 PM | EMS |  | Concentration | mg/L | Y | (fl. conc.) |  | Residue: Non-filterable (TSS) (fl. conc.) |  | Physical Properties; | Physical Properties; EMS code: 0008 | Physical |  | LAB
0008 | mg/L | Residue: Non-filterable (TSS) | F043 | Grav, 1.2um GF filter, 103-5C | A homogeneous sample aliquot is filtered through a Whatman GF/C glass fibre filter (1.2 u particle retention) that has been muffled at 550C for 20 minutes and pre-weighed. The filter with the residue is then dried at 103C for one hour and weighed once again to constant weight; the resulting weight differences gives the total non-filterable residue (NFR). | 2 | Y | Y | Y | 58751 | 1/3/1997 8:07:00 PM | 8/22/2024 | 8/25/2024 9:47:36 PM | 1997-05-08 | 9/30/2019 3:33:12 PM | MPAK | KATFRASE | Concentration | mg/L | Y | (fl. conc.) |  | Residue: Non-filterable (TSS) (fl. conc.) |  | Physical Properties; | Physical Properties; EMS code: 0008 | Physical |  | LAB
0008 | mg/L | Residue: Non-filterable (TSS) | XCAL | Calculation (mg/L) | Calculation (mg/L) | 1 | Y | Y | Y | 25749 | 11/18/1996 9:30:00 AM | 6/30/2024 7:00:00 AM | 8/15/2024 9:21:46 AM | 1996-04-12 | 6/3/1998 6:51:41 PM | EMS |  | Concentration | mg/L | Y | (fl. conc.) |  | Residue: Non-filterable (TSS) (fl. conc.) |  | Physical Properties; | Physical Properties; EMS code: 0008 | Physical |  | LAB
0008 | t/d | Residue: Non-filterable (TSS) | CALC | Calculation (t/d) | Calculation (t/d) | 0 | Y | Y | Y | 35967 | 9/26/1996 7:00:00 AM | 3/1/2016 | 8/22/2024 1:49:14 PM | 1996-04-12 | 6/3/1998 6:51:32 PM | EMS |  | MassFlow | t/d | Y | (mass flow) |  | Residue: Non-filterable (TSS) (mass flow) |  | Physical Properties; | Physical Properties; EMS code: 0008 | Physical |  | LAB
0008 | mg/L | Residue: Non-filterable (TSS) | 1071 | Grav, Subsamp Buch 105C | Grav; Subsamp Buch 105C | 4 | Y | Y | Y | 21851 | 7/18/1996 8:02:00 PM | 10/17/2019 2:51:00 PM | 11/4/2019 9:26:29 AM | 1996-04-12 | 3/14/2000 2:37:57 PM | EMS | UMULLER | Concentration | mg/L | Y | (fl. conc.) |  | Residue: Non-filterable (TSS) (fl. conc.) |  | Physical Properties; | Physical Properties; EMS code: 0008 | Physical |  | LAB
0008 | mg/L | Residue: Non-filterable (TSS) | 1070 | Grav, 9 cm Buchner, 105C | Grav; 9 cm Buchner, 105C | 1 | Y | Y | Y | 31950 | 1/2/1997 8:00:00 AM | 8/23/2024 11:24:00 AM | 8/26/2024 9:40:02 AM | 1996-04-12 | 3/14/2000 2:37:53 PM | EMS | UMULLER | Concentration | mg/L | Y | (fl. conc.) |  | Residue: Non-filterable (TSS) (fl. conc.) |  | Physical Properties; | Physical Properties; EMS code: 0008 | Physical |  | LAB
0008 | mg/L | Residue: Non-filterable (TSS) | X199 | Buchner 5.5 cm GF | Buchner 5.5 cm GF | 1 | Y | Y | Y | 3207 | 9/13/1996 | 10/5/2020 8:00:00 PM | 11/27/2020 6:25:47 AM | 1996-04-12 | 6/3/1998 6:51:32 PM | EMS |  | Concentration | mg/L | Y | (fl. conc.) |  | Residue: Non-filterable (TSS) (fl. conc.) |  | Physical Properties; | Physical Properties; EMS code: 0008 | Physical |  | LAB
0008 | mg/L | Residue: Non-filterable (TSS) | X011 | Filter Paper,Dry 105C | Filter Paper;Dry 105C | 1 | Y | Y | Y | 24419 | 1/13/2010 | 6/30/2024 | 7/23/2024 11:41:51 AM | 1996-04-12 | 6/3/1998 6:51:32 PM | EMS |  | Concentration | mg/L | Y | (fl. conc.) |  | Residue: Non-filterable (TSS) (fl. conc.) |  | Physical Properties; | Physical Properties; EMS code: 0008 | Physical |  | LAB
0008 | mg/L | Residue: Non-filterable (TSS) | XM00 | Meter-in situ | Meter-in situ |  | Y | Y | Y | 10516 |  |  | 8/3/2024 3:33:09 PM | 2003-02-17 |  | UMULLER |  | Concentration | mg/L | Y | (fl. conc.) |  | Residue: Non-filterable (TSS) (fl. conc.) |  | Physical Properties; | Physical Properties; EMS code: 0008 | Physical |  | LAB
0008 | mg/L | Residue: Non-filterable (TSS) | X141 | Buchner 11 cm GF | Buchner 11 cm GF | 1 | Y | Y | Y | 13125 | 10/5/1996 | 10/5/1996 | 8/12/2024 12:02:12 PM | 1996-04-12 | 6/3/1998 6:51:32 PM | EMS |  | Concentration | mg/L | Y | (fl. conc.) |  | Residue: Non-filterable (TSS) (fl. conc.) |  | Physical Properties; | Physical Properties; EMS code: 0008 | Physical |  | LAB
0008 | mg/L | Residue: Non-filterable (TSS) | X101 | Filtration,GFC | Filtration;GFC | 1 | Y | Y | Y | 1328 |  |  | 10/31/2016 5:30:59 PM | 1996-04-12 | 6/3/1998 6:51:32 PM | EMS |  | Concentration | mg/L | Y | (fl. conc.) |  | Residue: Non-filterable (TSS) (fl. conc.) |  | Physical Properties; | Physical Properties; EMS code: 0008 | Physical |  | LAB
0008 | mg/L | Residue: Non-filterable (TSS) | X026 | Gravimetric 0.45u Filter | A well-mixed sample is filtered through a standard membrane filter. A measured portion of the filtrate is evaporated in a preweighed evaporating vessel and dried to constant weight at 180 C. The increase in dish weight represents the total dissolved solids. (The filtrate from residue, non-filterable may be used.) | 1 | Y | Y | Y | 85840 | 9/21/1994 | 8/24/2024 | 8/26/2024 4:53:05 PM | 1996-04-12 | 6/3/1998 6:51:32 PM | EMS |  | Concentration | mg/L | Y | (fl. conc.) |  | Residue: Non-filterable (TSS) (fl. conc.) |  | Physical Properties; | Physical Properties; EMS code: 0008 | Physical |  | LAB
0008 | mg/L | Residue: Non-filterable (TSS) | X024 | Grav. Gooch Cruc 105C | Grav. Gooch Cruc 105C | 1 | Y | Y | Y | 67708 | 1/14/1996 8:00:00 AM | 8/1/2024 | 8/23/2024 9:49:37 AM | 1996-04-12 | 6/3/1998 6:51:32 PM | EMS |  | Concentration | mg/L | Y | (fl. conc.) |  | Residue: Non-filterable (TSS) (fl. conc.) |  | Physical Properties; | Physical Properties; EMS code: 0008 | Physical |  | LAB
0008 | kg/adt | Residue: Non-filterable (TSS) | CAL5 | Calculation (kg/adt) | Calculation (kg/adt) | 0 | Y | Y | Y | 995 |  |  | 8/22/2001 9:59:25 AM | 1996-10-13 | 6/3/1998 6:51:46 PM | EMS |  | PerAirDriedTonne | kg/adt | Y | (/adt) |  | Residue: Non-filterable (TSS) (/adt) |  | Physical Properties; | Physical Properties; EMS code: 0008 | Physical |  | LAB
0008 | mg/L | Residue: Non-filterable (TSS) | EQ01 | EQUIS (mg/L) | EQUIS (mg/L) | 1 | N | Y | Y | 22054 |  |  | 10/24/1996 3:16:58 PM | 1996-04-12 | 6/3/1998 6:51:32 PM | EMS |  | Concentration | mg/L | Y | (fl. conc.) |  | Residue: Non-filterable (TSS) (fl. conc.) |  | Physical Properties; | Physical Properties; EMS code: 0008 | Physical |  | LAB
0009 | mg/L | Residue: Fixed Non-filterable (TSS; Fixed) | CS01 | System Calculation | System Calculation (mg/L) |  | N | Y | N | 6986 |  |  | 8/1/2024 5:01:10 PM | 1996-04-12 | 6/3/1998 6:51:42 PM | EMS |  | Concentration | mg/L | Y | (fl. conc.) |  | Residue: Fixed Non-filterable (TSS; Fixed) (fl. conc.) |  | Physical Properties; | Physical Properties; EMS code: 0009 | Physical |  | LAB
0009 | mg/L | Residue: Fixed Non-filterable (TSS; Fixed) | EQ01 | EQUIS (mg/L) | EQUIS (mg/L) | 1 | N | Y | Y | 1703 |  |  | 10/24/1996 3:16:20 PM | 1996-04-12 | 6/3/1998 6:51:32 PM | EMS |  | Concentration | mg/L | Y | (fl. conc.) |  | Residue: Fixed Non-filterable (TSS; Fixed) (fl. conc.) |  | Physical Properties; | Physical Properties; EMS code: 0009 | Physical |  | LAB
0009 | mg/L | Residue: Fixed Non-filterable (TSS; Fixed) | 1040 | Grav.Gooch Cruc.550C | Grav.Gooch Cruc.550C | 1 | Y | Y | Y | 1 |  |  | 10/23/1996 4:24:07 PM | 1996-04-12 | 6/3/1998 6:51:45 PM | EMS |  | Concentration | mg/L | Y | (fl. conc.) |  | Residue: Fixed Non-filterable (TSS; Fixed) (fl. conc.) |  | Physical Properties; | Physical Properties; EMS code: 0009 | Physical |  | LAB
0009 | mg/L | Residue: Fixed Non-filterable (TSS; Fixed) | 1050 | Grav, 9 cm Buchner, 550C | Residue from determination of residue, non-filterable, is ignited to constant weight at 550 C in a muffle furnace. Usually, a 15 to 20 minute ignition is required. The ignited residue is cooled in a desiccator and weighed. The cycle of igniting, cooling, desiccating and weighing is repeated until a constant weight is attained. | 1 | Y | Y | Y | 1323 |  |  | 3/21/1997 7:09:30 PM | 1996-04-12 | 6/3/1998 6:51:46 PM | EMS |  | Concentration | mg/L | Y | (fl. conc.) |  | Residue: Fixed Non-filterable (TSS; Fixed) (fl. conc.) |  | Physical Properties; | Physical Properties; EMS code: 0009 | Physical |  | LAB
0009 | mg/L | Residue: Fixed Non-filterable (TSS; Fixed) | F012 | 1.2 um GF filt:Grav. 550C | A homogeneous sample is filtered through a 5.5 cm Whatman GF/C glass fibre filter (1.2 um particle retention) that has been muffled at 550 C for 20 minutes and pre-weighed. For fixed non-filterable residue (FNFR) analysis, the filter with the residue is then dried at 103 C for one hour and weighed once again to constant weight. The filter with dried and weighed residue is then ignited in a muffle at 550 C to determine FNFR. For fixed filterable residue (FFR) analysis, evaporated overnight at 90-95C, then dried for 1 hour at 103C and weighed for total filterable residue (FR), then ignited for 1 hour at 550C to determine FFR. Volatile residue is the difference in weights after drying at 103 C and asking at 550 C. | 10 | Y | Y | Y | 2967 | 8/30/1996 6:31:00 PM | 7/31/2020 | 8/7/2020 3:05:13 PM | 1996-06-19 | 6/3/1998 6:51:44 PM | MPAK |  | Concentration | mg/L | Y | (fl. conc.) |  | Residue: Fixed Non-filterable (TSS; Fixed) (fl. conc.) |  | Physical Properties; | Physical Properties; EMS code: 0009 | Physical |  | LAB
0009 | t/d | Residue: Fixed Non-filterable (TSS; Fixed) | CS30 | System Calculation | System Calculation (t/d) |  | N | Y | N | 226 |  |  | 8/30/2001 2:50:05 PM | 1998-05-06 | 6/3/1998 6:51:49 PM | EMS |  | MassFlow | t/d | Y | (mass flow) |  | Residue: Fixed Non-filterable (TSS; Fixed) (mass flow) |  | Physical Properties; | Physical Properties; EMS code: 0009 | Physical |  | LAB
0009 | mg/L | Residue: Fixed Non-filterable (TSS; Fixed) | X079 | Gravimetric 20min | Gravimetric 20min | 1 | Y | Y | Y | 105 |  |  | 8/25/1999 5:38:49 PM | 1996-04-12 | 6/3/1998 6:51:32 PM | EMS |  | Concentration | mg/L | Y | (fl. conc.) |  | Residue: Fixed Non-filterable (TSS; Fixed) (fl. conc.) |  | Physical Properties; | Physical Properties; EMS code: 0009 | Physical |  | LAB
0010 | mg/L | Residue: Volatile Non-filterable (TSS; Volatile) | X036 | Cent,Glass Fiber Filt550 | Cent;Glass Fiber Filt550 | 1 | Y | Y | Y | 41 |  |  | 10/23/1996 11:05:53 PM | 1996-04-12 | 6/3/1998 6:51:32 PM | EMS |  | Concentration | mg/L | Y | (fl. conc.) |  | Residue: Volatile Non-filterable (TSS; Volatile) (fl. conc.) |  | Physical Properties; | Physical Properties; EMS code: 0010 | Physical |  | LAB
0010 | t/d | Residue: Volatile Non-filterable (TSS; Volatile) | CALC | Calculation (t/d) | Calculation (t/d) | 0 | Y | Y | Y | 256 |  |  | 8/30/2001 2:50:03 PM | 1996-04-12 | 6/3/1998 6:51:32 PM | EMS |  | MassFlow | t/d | Y | (mass flow) |  | Residue: Volatile Non-filterable (TSS; Volatile) (mass flow) |  | Physical Properties; | Physical Properties; EMS code: 0010 | Physical |  | LAB
0010 | mg/L | Residue: Volatile Non-filterable (TSS; Volatile) | X012 | Filter Paper,Dry 600C | Filter Paper;Dry 600C | 1 | Y | Y | Y | 248 | 1/3/1997 8:00:00 AM | 7/27/2001 8:00:00 AM | 8/31/2001 10:38:00 AM | 1996-04-12 | 6/3/1998 6:51:45 PM | EMS |  | Concentration | mg/L | Y | (fl. conc.) |  | Residue: Volatile Non-filterable (TSS; Volatile) (fl. conc.) |  | Physical Properties; | Physical Properties; EMS code: 0010 | Physical |  | LAB
0010 | mg/L | Residue: Volatile Non-filterable (TSS; Volatile) | 1050 | Grav, 9 cm Buchner, 550C | Residue from determination of residue, non-filterable, is ignited to constant weight at 550 C in a muffle furnace. Usually, a 15 to 20 minute ignition is required. The ignited residue is cooled in a desiccator and weighed. The cycle of igniting, cooling, desiccating and weighing is repeated until a constant weight is attained. | 1 | Y | Y | Y | 1349 | 4/4/2011 8:04:00 AM | 7/31/2024 | 8/1/2024 5:01:10 PM | 1996-04-12 | 6/3/1998 6:51:45 PM | EMS |  | Concentration | mg/L | Y | (fl. conc.) |  | Residue: Volatile Non-filterable (TSS; Volatile) (fl. conc.) |  | Physical Properties; | Physical Properties; EMS code: 0010 | Physical |  | LAB
0010 | t/d | Residue: Volatile Non-filterable (TSS; Volatile) | EQ30 | EQUIS (T/d) | EQUIS (T/d) | 0 | N | Y | Y | 13 |  |  | 10/22/1996 11:31:03 PM | 1996-04-12 | 6/3/1998 6:51:32 PM | EMS |  | MassFlow | t/d | Y | (mass flow) |  | Residue: Volatile Non-filterable (TSS; Volatile) (mass flow) |  | Physical Properties; | Physical Properties; EMS code: 0010 | Physical |  | LAB
0010 | mg/L | Residue: Volatile Non-filterable (TSS; Volatile) | CS01 | System Calculation | System Calculation (mg/L) |  | N | Y | N | 1671 |  |  | 8/13/2016 5:13:46 PM | 1996-04-12 | 6/3/1998 6:51:42 PM | EMS |  | Concentration | mg/L | Y | (fl. conc.) |  | Residue: Volatile Non-filterable (TSS; Volatile) (fl. conc.) |  | Physical Properties; | Physical Properties; EMS code: 0010 | Physical |  | LAB
0010 | mg/L | Residue: Volatile Non-filterable (TSS; Volatile) | X049 | Muffle Furnace 600C | Muffle Furnace 600C | 1 | Y | Y | Y | 2460 | 9/4/1996 | 4/8/2024 | 6/17/2024 3:26:08 PM | 1996-04-12 | 6/3/1998 6:51:32 PM | EMS |  | Concentration | mg/L | Y | (fl. conc.) |  | Residue: Volatile Non-filterable (TSS; Volatile) (fl. conc.) |  | Physical Properties; | Physical Properties; EMS code: 0010 | Physical |  | LAB
0010 | mg/L | Residue: Volatile Non-filterable (TSS; Volatile) | EQ01 | EQUIS (mg/L) | EQUIS (mg/L) | 1 | N | Y | Y | 820 |  |  | 10/23/1996 12:03:15 AM | 1996-04-12 | 6/3/1998 6:51:32 PM | EMS |  | Concentration | mg/L | Y | (fl. conc.) |  | Residue: Volatile Non-filterable (TSS; Volatile) (fl. conc.) |  | Physical Properties; | Physical Properties; EMS code: 0010 | Physical |  | LAB
0010 | mg/L | Residue: Volatile Non-filterable (TSS; Volatile) | XCAL | Calculation (mg/L) | Calculation (mg/L) | 5 | Y | Y | Y | 299 | 4/14/1998 6:31:00 PM | 8/11/2014 12:01:00 PM | 6/22/2015 2:30:55 PM | 1996-04-12 | 6/3/1998 6:51:41 PM | EMS |  | Concentration | mg/L | Y | (fl. conc.) |  | Residue: Volatile Non-filterable (TSS; Volatile) (fl. conc.) |  | Physical Properties; | Physical Properties; EMS code: 0010 | Physical |  | LAB
0010 | mg/L | Residue: Volatile Non-filterable (TSS; Volatile) | X175 | Gravimetric | Code X175 covers a variety of gravimetric procedures. | 1 | Y | Y | Y | 4870 |  |  | 7/29/2024 10:49:10 AM | 1996-04-12 | 6/3/1998 6:51:32 PM | EMS |  | Concentration | mg/L | Y | (fl. conc.) |  | Residue: Volatile Non-filterable (TSS; Volatile) (fl. conc.) |  | Physical Properties; | Physical Properties; EMS code: 0010 | Physical |  | LAB
0011 | uS/cm | Specific Conductance | COND | Conductivity APHA 2510B | Conductivity APHA 2510B. Conductance of a solution measured between two spatially fixed and chemically inert electrodes. | 1 | Y | Y | Y | 1082 | 7/8/2015 | 10/19/2023 | 7/31/2024 1:46:41 PM | 2016-06-30 |  | RDGIBSON |  | ElectricConductivity | µS/cm | Y | (elec. cond.) |  | Specific Conductance (elec. cond.) |  | Physical Properties; | Physical Properties; EMS code: 0011 | Physical |  | LAB
0011 | uS/cm | Specific Conductance | XH00 | Hydrolab | Hydrolab | 1 | Y | Y | Y | 12447 | 4/7/1997 1:25:00 PM | 12/16/1999 11:41:00 AM | 6/7/2024 5:30:18 PM | 1996-04-12 | 12/20/2019 11:02:20 AM | EMS | RDGIBSON | ElectricConductivity | µS/cm | Y | (elec. cond.) |  | Specific Conductance (elec. cond.) |  | Physical Properties; | Physical Properties; EMS code: 0011 | Physical |  | FIELD_RESULT | Hydrolab
0011 | uS/cm | Specific Conductance | X330 | Meter | The conductance (= conductivity) of a sample is measured by use of a self-contained conductivity meter. pH is measured with a combined glass/reference (Ag/AgCl) electrode. For both conductivity and pH the measurements are made at 25 C, or the results are temperature-corrected to 25 C. Note that laboratory and field measurements are distinguished by the EMS Analytical Agency codes. Ex situ and in situ measurements are distinguished by the computer codes. | 1 | Y | Y | Y | 257835 | 9/12/1994 | 8/24/2024 | 8/26/2024 4:53:05 PM | 1996-04-12 | 12/20/2019 11:02:20 AM | EMS | RDGIBSON | ElectricConductivity | µS/cm | Y | (elec. cond.) |  | Specific Conductance (elec. cond.) |  | Physical Properties; | Physical Properties; EMS code: 0011 | Physical |  | FIELD_RESULT | Meter
0011 | uS/cm | Specific Conductance | XM01 | DO Meter | Following the manufacturer's instructions, the probe with oxygen permeable membrane is calibrated against air or samples of known DO concentration (e.g. from Winkler titrations). The samples are then measured for DO, again following all precautions recommended by the manufacturer to insure acceptable results. | 1 | N | Y | Y | 603 |  |  | 3/26/1997 1:08:17 PM | 1996-04-12 | 6/3/1998 6:51:32 PM | EMS |  | ElectricConductivity | µS/cm | Y | (elec. cond.) |  | Specific Conductance (elec. cond.) |  | Physical Properties; | Physical Properties; EMS code: 0011 | Physical |  | LAB
0011 | uS/cm | Specific Conductance | XM10 | Meter | Meter (April 9/97 - changed description from 'Flow Meter' to 'Meter'; this method code is associated with parameter 0011 - Specific Conductance therefore 'Meter' is more appropriate. Note that X330 is also called 'Meter' and is the preferred method code.) | 2 | Y | Y | Y | 89916 | 5/7/1996 8:37:00 AM | 8/22/2024 | 8/23/2024 5:01:31 PM | 1996-09-17 | 6/3/1998 6:51:45 PM | MPAK |  | ElectricConductivity | µS/cm | Y | (elec. cond.) |  | Specific Conductance (elec. cond.) |  | Physical Properties; | Physical Properties; EMS code: 0011 | Physical |  | LAB
0011 | uS/cm | Specific Conductance | EQ04 | EQUIS (uS/cm) | EQUIS (uS/cm) | 1 | N | Y | Y | 35488 |  |  | 10/24/1996 3:17:07 PM | 1996-04-12 | 6/3/1998 6:51:32 PM | EMS |  | ElectricConductivity | µS/cm | Y | (elec. cond.) |  | Specific Conductance (elec. cond.) |  | Physical Properties; | Physical Properties; EMS code: 0011 | Physical |  | LAB
0011 | uS/cm | Specific Conductance | XM00 | Meter-in situ | Meter-in situ | 1 | Y | Y | Y | 138464 | 4/30/1996 7:00:00 AM | 8/1/2024 | 8/23/2024 9:49:37 AM | 1996-04-12 | 10/17/2012 10:34:38 AM | EMS | RDGIBSON | ElectricConductivity | µS/cm | Y | (elec. cond.) |  | Specific Conductance (elec. cond.) |  | Physical Properties; | Physical Properties; EMS code: 0011 | Physical |  | LAB
0013 | C | Temperature | XH00 | Hydrolab | Hydrolab | -20 | Y | Y | Y | 25717 | 1/2/1997 1:50:00 PM | 4/2/2008 11:50:00 AM | 5/31/2023 1:30:30 PM | 1996-04-12 | 12/20/2019 11:02:20 AM | EMS | RDGIBSON | Temperature | degC | Y | (temp.) |  | Temperature (temp.) |  | Physical Properties; | Physical Properties; EMS code: 0013 | Physical |  | FIELD_RESULT | Hydrolab
0013 | C | Temperature | XM15 | Continuous Rec.-in situ | Continuous Rec.-in situ | -20 | Y | Y | Y | 155286 | 9/2/1996 | 7/1/2024 | 8/22/2024 1:49:14 PM | 1996-04-12 | 6/3/1998 6:51:45 PM | EMS |  | Temperature | degC | Y | (temp.) |  | Temperature (temp.) |  | Physical Properties; | Physical Properties; EMS code: 0013 | Physical |  | FIELD_RESULT | Continuous Rec.-in situ
0013 | C | Temperature | EQ32 | EQUIS (C) | EQUIS (C) | -20 | N | Y | Y | 35935 |  |  | 10/24/1996 3:16:40 PM | 1996-04-12 | 6/3/1998 6:51:32 PM | EMS |  | Temperature | degC | Y | (temp.) |  | Temperature (temp.) |  | Physical Properties; | Physical Properties; EMS code: 0013 | Physical |  | LAB
0013 | C | Temperature | XM01 | DO Meter | Following the manufacturer's instructions, the probe with oxygen permeable membrane is calibrated against air or samples of known DO concentration (e.g. from Winkler titrations). The samples are then measured for DO, again following all precautions recommended by the manufacturer to insure acceptable results. | -20 | Y | Y | Y | 39336 | 1/7/1997 9:00:00 AM | 10/6/2015 9:48:00 AM | 7/25/2024 3:53:30 PM | 1996-04-12 | 12/20/2019 11:02:20 AM | EMS | RDGIBSON | Temperature | degC | Y | (temp.) |  | Temperature (temp.) |  | Physical Properties; | Physical Properties; EMS code: 0013 | Physical |  | FIELD_RESULT | DO Meter
0013 | C | Temperature | XM16 | Thermometer (Hg) | Thermometer (Hg) | -20 | Y | Y | Y | 55042 | 1/2/1996 7:59:00 AM | 8/1/2024 | 8/23/2024 9:49:37 AM | 1996-04-12 | 6/3/1998 6:51:32 PM | EMS |  | Temperature | degC | Y | (temp.) |  | Temperature (temp.) |  | Physical Properties; | Physical Properties; EMS code: 0013 | Physical |  | LAB
0013 | C | Temperature | XM02 | Thermometer | Thermometer | -20 | Y | Y | Y | 120035 | 1/2/1997 | 8/17/2024 4:09:00 PM | 8/25/2024 9:22:38 PM | 1996-04-12 | 10/16/2012 4:12:54 PM | EMS | RDGIBSON | Temperature | degC | Y | (temp.) |  | Temperature (temp.) |  | Physical Properties; | Physical Properties; EMS code: 0013 | Physical |  | FIELD_RESULT | Thermometer
0013 | C | Temperature | DBPFP | DBPFP Incubation | DBPFP Incubation | 1 | Y | Y | Y | 2 | 3/18/2023 4:37:00 PM | 5/23/2023 12:07:00 PM | 6/8/2023 12:14:56 PM | 2023-03-31 |  | EMS |  | Temperature | degC | Y | (temp.) |  | Temperature (temp.) |  | Physical Properties; | Physical Properties; EMS code: 0013 | Physical |  | LAB
0013 | C | Temperature | X330 | Meter | The conductance (= conductivity) of a sample is measured by use of a self-contained conductivity meter. pH is measured with a combined glass/reference (Ag/AgCl) electrode. For both conductivity and pH the measurements are made at 25 C, or the results are temperature-corrected to 25 C. Note that laboratory and field measurements are distinguished by the EMS Analytical Agency codes. Ex situ and in situ measurements are distinguished by the computer codes. | 0 | Y | Y | Y | 76562 | 3/1/2000 | 6/30/2024 | 8/23/2024 8:10:35 AM | 1997-04-09 | 12/20/2019 11:02:20 AM | MPAK | RDGIBSON | Temperature | degC | Y | (temp.) |  | Temperature (temp.) |  | Physical Properties; | Physical Properties; EMS code: 0013 | Physical |  | FIELD_RESULT | Meter
0013 | C | Temperature | EXOSO | EXO Sonde YSI | EXO Sonde YSI |  | Y | Y | Y | 382 |  |  | 7/15/2024 3:01:47 PM | 2023-06-12 |  | EMS |  | Temperature | degC | Y | (temp.) |  | Temperature (temp.) |  | Physical Properties; | Physical Properties; EMS code: 0013 | Physical |  | FIELD_RESULT | EXO Sonde YSI
0013 | C | Temperature | XT01 | Thermister | Thermister | -20 | Y | Y | Y | 231 |  |  | 10/24/1996 2:53:41 PM | 1996-04-12 | 6/3/1998 6:51:32 PM | EMS |  | Temperature | degC | Y | (temp.) |  | Temperature (temp.) |  | Physical Properties; | Physical Properties; EMS code: 0013 | Physical |  | LAB
0014 | mg/L | Oxygen Dissolved | XM01 | DO Meter | Following the manufacturer's instructions, the probe with oxygen permeable membrane is calibrated against air or samples of known DO concentration (e.g. from Winkler titrations). The samples are then measured for DO, again following all precautions recommended by the manufacturer to insure acceptable results. | 0 | Y | Y | Y | 138488 | 9/1/1996 | 7/1/2024 | 8/20/2024 8:39:43 AM | 1996-04-12 | 10/16/2012 4:12:27 PM | EMS | RDGIBSON | Concentration | mg/L | Y | (fl. conc.) |  | Oxygen Dissolved (fl. conc.) |  | Inorganics; | Inorganics; EMS code: 0014 | Chemical | Dissolved | LAB
0014 | mg/L | Oxygen Dissolved | EQ01 | EQUIS (mg/L) | EQUIS (mg/L) | 0 | N | Y | Y | 22367 |  |  | 10/24/1996 3:16:40 PM | 1996-04-12 | 6/3/1998 6:51:32 PM | EMS |  | Concentration | mg/L | Y | (fl. conc.) |  | Oxygen Dissolved (fl. conc.) |  | Inorganics; | Inorganics; EMS code: 0014 | Chemical | Dissolved | LAB
0014 | mg/L | Oxygen Dissolved | EXOSO | EXO Sonde YSI | EXO Sonde YSI |  | Y | Y | Y | 359 |  |  | 7/15/2024 3:01:47 PM | 2023-06-12 |  | EMS |  | Concentration | mg/L | Y | (fl. conc.) |  | Oxygen Dissolved (fl. conc.) |  | Inorganics; | Inorganics; EMS code: 0014 | Chemical | Dissolved | FIELD_RESULT | EXO Sonde YSI
0014 | mg/L | Oxygen Dissolved | X330 | Meter | The conductance (= conductivity) of a sample is measured by use of a self-contained conductivity meter. pH is measured with a combined glass/reference (Ag/AgCl) electrode. For both conductivity and pH the measurements are made at 25 C, or the results are temperature-corrected to 25 C. Note that laboratory and field measurements are distinguished by the EMS Analytical Agency codes. Ex situ and in situ measurements are distinguished by the computer codes. | 0.1 | Y | Y | Y | 10561 | 8/16/2011 9:00:00 AM | 11/7/2023 9:33:00 AM | 7/23/2024 2:26:52 PM | 2011-06-15 |  | RDGIBSON |  | Concentration | mg/L | Y | (fl. conc.) |  | Oxygen Dissolved (fl. conc.) |  | Inorganics; | Inorganics; EMS code: 0014 | Chemical | Dissolved | FIELD_RESULT | Meter
0014 | % | Oxygen Dissolved | CAL3 | Calculation (%) | Calculation (%) | 0 | Y | Y | Y | 54 |  |  | 1/29/2024 7:23:17 PM | 1996-04-12 | 6/3/1998 6:51:41 PM | EMS |  | DimensionlessRatio | % | N | (ratio) |  | Oxygen Dissolved (ratio) |  | Inorganics; | Inorganics; EMS code: 0014 | Chemical | Dissolved | LAB
0014 | mg/L | Oxygen Dissolved | XH00 | Hydrolab | Hydrolab | 0 | Y | Y | Y | 24113 | 4/2/2008 11:50:00 AM | 4/2/2008 11:50:00 AM | 11/21/2019 9:29:57 AM | 1996-04-12 | 12/20/2019 11:02:20 AM | EMS | RDGIBSON | Concentration | mg/L | Y | (fl. conc.) |  | Oxygen Dissolved (fl. conc.) |  | Inorganics; | Inorganics; EMS code: 0014 | Chemical | Dissolved | FIELD_RESULT | Hydrolab
0014 | mg/L | Oxygen Dissolved | X047 | Winkler-Azide Mod. | Winkler-Azide Mod. | 0 | Y | Y | Y | 128 |  |  | 3/3/2000 12:25:19 PM | 1996-04-12 | 6/3/1998 6:51:32 PM | EMS |  | Concentration | mg/L | Y | (fl. conc.) |  | Oxygen Dissolved (fl. conc.) |  | Inorganics; | Inorganics; EMS code: 0014 | Chemical | Dissolved | LAB
0014 | % | Oxygen Dissolved | EQ34 | EQUIS (%) | EQUIS (%) | 0 | N | Y | Y | 2065 |  |  | 10/22/1996 10:40:49 PM | 1996-04-12 | 6/3/1998 6:51:32 PM | EMS |  | DimensionlessRatio | % | N | (ratio) |  | Oxygen Dissolved (ratio) |  | Inorganics; | Inorganics; EMS code: 0014 | Chemical | Dissolved | LAB
0014 | mg/L | Oxygen Dissolved | X002 | Hach | Hach | 0 | Y | Y | Y | 602 | 3/21/2006 | 3/21/2006 | 3/29/2006 4:17:52 PM | 1996-04-12 | 6/3/1998 6:51:32 PM | EMS |  | Concentration | mg/L | Y | (fl. conc.) |  | Oxygen Dissolved (fl. conc.) |  | Inorganics; | Inorganics; EMS code: 0014 | Chemical | Dissolved | LAB
0014 | mg/L | Oxygen Dissolved | X001 | Winkler | Winkler | 0 | Y | Y | Y | 78 |  |  | 1/3/1997 11:11:33 AM | 1996-04-12 | 6/3/1998 6:51:32 PM | EMS |  | Concentration | mg/L | Y | (fl. conc.) |  | Oxygen Dissolved (fl. conc.) |  | Inorganics; | Inorganics; EMS code: 0014 | Chemical | Dissolved | LAB
0015 | NTU | Turbidity | 1150 | Nephel.Double Beam | Nephel.Double Beam | 0.1 | Y | Y | Y | 13 |  |  | 10/31/2016 5:30:59 PM | 1996-04-12 | 6/3/1998 6:51:32 PM | EMS |  | Turbidity | NTU | Y | (turb.) |  | Turbidity (turb.) |  | Physical Properties; | Physical Properties; EMS code: 0015 | Physical |  | LAB
0015 | NTU | Turbidity | XM12 | Turbidity Meter | Turbidity Meter | 0.1 | Y | Y | Y | 74215 | 4/24/1997 | 8/17/2024 4:31:00 PM | 8/23/2024 5:01:25 PM | 1996-04-12 | 10/16/2012 4:12:39 PM | EMS | RDGIBSON | Turbidity | NTU | Y | (turb.) |  | Turbidity (turb.) |  | Physical Properties; | Physical Properties; EMS code: 0015 | Physical |  | LAB
0015 | NTU | Turbidity | EQ07 | EQUIS (NTU) | EQUIS (NTU) | 0 | N | Y | Y | 20613 |  |  | 10/24/1996 3:17:06 PM | 1996-04-12 | 6/3/1998 6:51:32 PM | EMS |  | Turbidity | NTU | Y | (turb.) |  | Turbidity (turb.) |  | Physical Properties; | Physical Properties; EMS code: 0015 | Physical |  | LAB
0015 | mg/L | Turbidity | EQ01 | EQUIS (mg/L) | EQUIS (mg/L) | 0 | N | Y | Y | 159 |  |  | 10/24/1996 1:52:20 PM | 1996-04-12 | 6/3/1998 6:51:32 PM | EMS |  | Concentration | mg/L | Y | (fl. conc.) |  | Turbidity (fl. conc.) |  | Physical Properties; | Physical Properties; EMS code: 0015 | Physical |  | LAB
0015 | NTU | Turbidity | XM08 | Nephelometer | Nephelometer | 0.1 | Y | Y | Y | 96939 | 1/24/2001 | 8/24/2024 | 8/26/2024 4:53:05 PM | 1996-04-12 | 12/20/2019 11:02:20 AM | EMS | RDGIBSON | Turbidity | NTU | Y | (turb.) |  | Turbidity (turb.) |  | Physical Properties; | Physical Properties; EMS code: 0015 | Physical |  | LAB
0015 | NTU | Turbidity | 1140 | Nephel.Single Beam | Light is directed through a tube containing the sample, and the relative intensity of the scattered light, at 90 degrees to the light source, is measured by photocells. An instrument designed to do this is called a nephelometer, which is calibrated with a standard formazin reference suspension. The instrument readout is in terms of Nephelometric Turbidity Units (NTU), or the equivalent formazin turbidity units (FTU) or the equivalent Jackson Turbidity Units (JTU). | 0.05 | Y | Y | Y | 106176 | 7/18/1996 4:36:00 PM | 1/31/2015 | 8/26/2024 2:41:32 PM | 1996-04-12 | 6/3/1998 6:51:32 PM | EMS |  | Turbidity | NTU | Y | (turb.) |  | Turbidity (turb.) |  | Physical Properties; | Physical Properties; EMS code: 0015 | Physical |  | LAB
0015 | None | Turbidity | X189 | Visual- Wedge | Visual- Wedge | 0 | Y | Y | Y | 5408 |  |  | 4/25/2005 1:30:34 PM | 1996-04-12 | 6/3/1998 6:51:32 PM | EMS |  | None | None | N | (unitless) |  | Turbidity (unitless) |  | Physical Properties; | Physical Properties; EMS code: 0015 | Physical |  | LAB
0016 | mg/L | Chlorine Res:Free | X004 | Colorimetric Titr'n | Manual Colorimetric Titration with standard solutions to specified endpoint as identified by colorimetric indicator. The sample must not be filtered, diluted, concentrated, or altered in any way. | 0 | Y | Y | Y | 104 | 8/19/2011 10:35:00 AM | 9/27/2023 9:24:00 AM | 10/26/2023 5:01:16 PM | 1996-04-12 | 6/3/1998 6:51:32 PM | EMS |  | Concentration | mg/L | Y | (fl. conc.) |  | Chlorine Res:Free (fl. conc.) | 7782-50-5 | Inorganics; | Inorganics; EMS code: 0016 | Chemical |  | LAB
0016 | mg/L | Chlorine Res:Free | X103 | DPD Colorimetric | Liberated iodine reacts with DPD to produce a red solution which is measured spectrophotometrically at 515 nm. | 0 | Y | Y | Y | 201 | 4/3/2007 4:10:00 PM | 2/22/2018 | 3/1/2018 5:16:09 PM | 1996-04-12 | 6/3/1998 6:51:32 PM | EMS |  | Concentration | mg/L | Y | (fl. conc.) |  | Chlorine Res:Free (fl. conc.) | 7782-50-5 | Inorganics; | Inorganics; EMS code: 0016 | Chemical |  | LAB
0016 | mg/L | Chlorine Res:Free | X006 | Amperometric | Amperometric | 0 | Y | Y | Y | 30 |  |  | 10/24/1996 12:38:14 PM | 1996-04-12 | 3/18/2003 9:38:19 AM | EMS | UMULLER | Concentration | mg/L | Y | (fl. conc.) |  | Chlorine Res:Free (fl. conc.) | 7782-50-5 | Inorganics; | Inorganics; EMS code: 0016 | Chemical |  | LAB
0016 | mg/L | Chlorine Res:Free | EQ01 | EQUIS (mg/L) | EQUIS (mg/L) | 0 | N | Y | Y | 82 |  |  | 10/22/1996 9:25:42 PM | 1996-04-12 | 6/3/1998 6:51:45 PM | EMS |  | Concentration | mg/L | Y | (fl. conc.) |  | Chlorine Res:Free (fl. conc.) | 7782-50-5 | Inorganics; | Inorganics; EMS code: 0016 | Chemical |  | LAB
0016 | mg/L | Chlorine Res:Free | X002 | Hach | Hach | 0 | N | Y | Y | 257 |  |  | 3/13/1997 4:17:31 PM | 1996-04-12 | 6/3/1998 6:51:32 PM | EMS |  | Concentration | mg/L | Y | (fl. conc.) |  | Chlorine Res:Free (fl. conc.) | 7782-50-5 | Inorganics; | Inorganics; EMS code: 0016 | Chemical |  | LAB
0016 | mg/L | Chlorine Res:Free | FLD | Field sampling | Field sampling method | 0.01 | Y | Y | Y | 4 | 10/28/2022 | 2/22/2023 | 2/22/2023 12:07:20 PM | 2022-11-08 |  | EMS |  | Concentration | mg/L | Y | (fl. conc.) |  | Chlorine Res:Free (fl. conc.) | 7782-50-5 | Inorganics; | Inorganics; EMS code: 0016 | Chemical |  | FIELD_RESULT | Field sampling
0017 | kPa | Barometric Press. | XM03 | Barometer | Barometer | 90 | Y | Y | Y | 114 |  |  | 9/30/1998 1:47:34 PM | 1996-04-12 | 6/3/1998 6:51:32 PM | EMS |  | Pressure | kPa | Y | (pres.) |  | Barometric Press. (pres.) |  | Physical Properties; | Physical Properties; EMS code: 0017 | Physical |  | LAB
0017 | kPa | Barometric Press. | EQ28 | EQUIS (kPa) | EQUIS (kPa) | 90 | N | Y | Y | 3289 |  |  | 10/22/1996 10:56:57 PM | 1996-04-12 | 6/3/1998 6:51:32 PM | EMS |  | Pressure | kPa | Y | (pres.) |  | Barometric Press. (pres.) |  | Physical Properties; | Physical Properties; EMS code: 0017 | Physical |  | LAB
0018 | m3/d | Flow (Avg) | XM04 | Totalizer | Totalizer | 0 | Y | Y | Y | 77638 | 9/25/2022 | 8/10/2024 | 8/21/2024 11:21:28 AM | 1996-04-12 | 6/3/1998 6:51:32 PM | EMS |  | Flow | m³/d | Y | (flow) |  | Flow (Avg) (flow) |  | Physical Properties; | Physical Properties; EMS code: 0018 | Physical |  | LAB
0018 | m3/d | Flow (Avg) | XG01 | Parshall Flume | Parshall Flume | 0 | Y | Y | Y | 35758 | 4/1/1997 9:50:00 AM | 12/31/2016 7:15:00 AM | 7/3/2024 12:59:29 PM | 1996-04-12 | 6/3/1998 6:51:45 PM | EMS |  | Flow | m³/d | Y | (flow) |  | Flow (Avg) (flow) |  | Physical Properties; | Physical Properties; EMS code: 0018 | Physical |  | LAB
0018 | m3/d | Flow (Avg) | XM19 | Contin.-Electr.Sensors | Contin.-Electr.Sensors | 0 | Y | Y | Y | 1162 |  |  | 10/24/1996 2:37:36 PM | 1996-04-12 | 6/3/1998 6:51:32 PM | EMS |  | Flow | m³/d | Y | (flow) |  | Flow (Avg) (flow) |  | Physical Properties; | Physical Properties; EMS code: 0018 | Physical |  | LAB
0018 | m3/d | Flow (Avg) | XM10 | Meter | Meter (April 9/97 - changed description from 'Flow Meter' to 'Meter'; this method code is associated with parameter 0011 - Specific Conductance therefore 'Meter' is more appropriate. Note that X330 is also called 'Meter' and is the preferred method code.) | 0 | Y | Y | Y | 321010 | 1/2/1996 7:59:00 AM | 7/31/2024 11:59:00 PM | 8/22/2024 1:49:14 PM | 1996-04-12 | 6/3/1998 6:51:45 PM | EMS |  | Flow | m³/d | Y | (flow) |  | Flow (Avg) (flow) |  | Physical Properties; | Physical Properties; EMS code: 0018 | Physical |  | LAB
0018 | m3/d | Flow (Avg) | X462 | 90th percentile | 90th percentile | 0 | Y | Y | Y | 772 | 1/31/1996 7:00:00 AM | 2/4/1999 7:00:00 AM | 7/16/2024 3:11:23 PM | 1996-04-12 | 6/3/1998 6:51:42 PM | EMS |  | Flow | m³/d | Y | (flow) |  | Flow (Avg) (flow) |  | Physical Properties; | Physical Properties; EMS code: 0018 | Physical |  | LAB
0018 | m3/d | Flow (Avg) | EQ35 | EQUIS (m3/d) | EQUIS (m3/d) | 0 | N | Y | Y | 19415 |  |  | 10/24/1996 12:59:23 PM | 1996-04-12 | 6/3/1998 6:51:32 PM | EMS |  | Flow | m³/d | Y | (flow) |  | Flow (Avg) (flow) |  | Physical Properties; | Physical Properties; EMS code: 0018 | Physical |  | LAB
0018 | m3/d | Flow (Avg) | CAL4 | Calculation (m3/d) | Calculation (m3/d) | 0 | Y | Y | Y | 35830 | 10/2/1996 | 7/31/2024 11:59:00 PM | 8/21/2024 11:42:56 AM | 1996-04-12 | 6/3/1998 6:51:41 PM | EMS |  | Flow | m³/d | Y | (flow) |  | Flow (Avg) (flow) |  | Physical Properties; | Physical Properties; EMS code: 0018 | Physical |  | LAB
0018 | m3/d | Flow (Avg) | FLD | Field sampling | Field sampling method | 0 | Y | Y | Y | 33 | 12/30/2019 | 8/13/2024 | 8/13/2024 11:43:08 AM | 2019-07-10 |  | KATFRASE |  | Flow | m³/d | Y | (flow) |  | Flow (Avg) (flow) |  | Physical Properties; | Physical Properties; EMS code: 0018 | Physical |  | FIELD_RESULT | Field sampling
0018 | m3/d | Flow (Avg) | XG00 | Stream Gauge | Stream Gauge | 0 | Y | Y | Y | 2812 |  |  | 1/3/2013 2:59:08 PM | 1996-04-12 | 6/3/1998 6:51:32 PM | EMS |  | Flow | m³/d | Y | (flow) |  | Flow (Avg) (flow) |  | Physical Properties; | Physical Properties; EMS code: 0018 | Physical |  | LAB
0018 | m3/d | Flow (Avg) | XM00 | Meter-in situ | Meter-in situ | 0 | Y | Y | Y | 70434 | 9/2/1996 | 8/1/2024 | 8/23/2024 9:49:37 AM | 1996-04-12 | 6/3/1998 6:51:32 PM | EMS |  | Flow | m³/d | Y | (flow) |  | Flow (Avg) (flow) |  | Physical Properties; | Physical Properties; EMS code: 0018 | Physical |  | LAB
0018 | m3/d | Flow (Avg) | X144 | Volume/Time | Volume/Time | 0 | Y | Y | Y | 9255 | 1/9/2012 | 8/11/2015 9:00:00 AM | 7/29/2024 12:59:01 PM | 1996-04-12 | 6/3/1998 6:51:32 PM | EMS |  | Flow | m³/d | Y | (flow) |  | Flow (Avg) (flow) |  | Physical Properties; | Physical Properties; EMS code: 0018 | Physical |  | LAB
0018 | m3/d | Flow (Avg) | XG02 | V Notch Weir | V Notch Weir | 0 | Y | Y | Y | 32977 | 10/6/1998 10:30:00 AM | 1/1/2020 | 7/25/2024 11:22:07 AM | 1996-04-12 | 6/3/1998 6:51:32 PM | EMS |  | Flow | m³/d | Y | (flow) |  | Flow (Avg) (flow) |  | Physical Properties; | Physical Properties; EMS code: 0018 | Physical |  | LAB
0019 | m | Extinction Depth | XG04 | Secchi Disc + Box | Secchi Disc + Box | 0 | Y | Y | Y | 2030 | 8/19/1998 | 9/6/2017 | 12/9/2020 1:28:02 PM | 1996-04-12 | 6/3/1998 6:51:45 PM | EMS |  | Length | m | Y | (len.) |  | Extinction Depth (len.) |  | Field; | Field; EMS code: 0019 | Physical |  | FIELD_RESULT | Secchi Disc + Box
0019 | m | Extinction Depth | XG03 | Secchi Disc | Secchi Disc | 0 | Y | Y | Y | 7644 | 4/9/2002 9:00:00 AM | 5/24/2022 10:40:00 AM | 7/25/2024 2:56:05 PM | 1996-04-12 | 12/20/2019 11:02:20 AM | EMS | RDGIBSON | Length | m | Y | (len.) |  | Extinction Depth (len.) |  | Field; | Field; EMS code: 0019 | Physical |  | FIELD_RESULT | Secchi Disc
0019 | m | Extinction Depth | XM17 | Absorbometer | Absorbometer | 0 | Y | Y | Y | 618 | 2/17/2005 | 5/11/2015 | 5/15/2015 4:30:08 PM | 1996-04-12 | 6/3/1998 6:51:32 PM | EMS |  | Length | m | Y | (len.) |  | Extinction Depth (len.) |  | Field; | Field; EMS code: 0019 | Physical |  | FIELD_RESULT | Absorbometer
0019 | m | Extinction Depth | FLD | Field sampling | Field sampling method | 0 | Y | Y | Y | 1330 | 2/16/2016 3:00:00 PM | 8/21/2024 | 8/26/2024 1:03:06 PM | 2019-07-10 | 12/20/2019 10:20:07 AM | KATFRASE | RDGIBSON | Length | m | Y | (len.) |  | Extinction Depth (len.) |  | Field; | Field; EMS code: 0019 | Physical |  | FIELD_RESULT | Field sampling
0019 | m | Extinction Depth | EQ24 | EQUIS (m) | EQUIS (m) | 0 | N | Y | Y | 3918 |  |  | 10/23/1996 12:04:03 AM | 1996-04-12 | 6/3/1998 6:51:32 PM | EMS |  | Length | m | Y | (len.) |  | Extinction Depth (len.) |  | Field; | Field; EMS code: 0019 | Physical |  | FIELD_RESULT | EQUIS (m)
0020 | C | Temperature (Air) | XM02 | Thermometer | Thermometer | -20 | Y | Y | Y | 38239 | 6/4/1997 | 12/22/2015 1:10:00 PM | 9/11/2019 8:47:50 AM | 1996-04-12 | 12/20/2019 11:02:20 AM | EMS | RDGIBSON | Temperature | degC | Y | (temp.) |  | Temperature (Air) (temp.) |  | Field; | Field; EMS code: 0020 | Chemical |  | FIELD_RESULT | Thermometer
0020 | C | Temperature (Air) | XM01 | DO Meter | Following the manufacturer's instructions, the probe with oxygen permeable membrane is calibrated against air or samples of known DO concentration (e.g. from Winkler titrations). The samples are then measured for DO, again following all precautions recommended by the manufacturer to insure acceptable results. | -20 | Y | Y | Y | 184 |  |  | 6/12/1998 9:57:04 AM | 1996-04-12 | 12/20/2019 11:02:20 AM | EMS | RDGIBSON | Temperature | degC | Y | (temp.) |  | Temperature (Air) (temp.) |  | Field; | Field; EMS code: 0020 | Chemical |  | FIELD_RESULT | DO Meter
0020 | C | Temperature (Air) | EQ32 | EQUIS (C) | EQUIS (C) | -20 | N | Y | Y | 1212 |  |  | 10/22/1996 11:22:46 PM | 1996-04-12 | 6/3/1998 6:51:32 PM | EMS |  | Temperature | degC | Y | (temp.) |  | Temperature (Air) (temp.) |  | Field; | Field; EMS code: 0020 | Chemical |  | FIELD_RESULT | EQUIS (C)
0020 | C | Temperature (Air) | XH00 | Hydrolab | Hydrolab | -20 | Y | Y | Y | 132 |  |  | 12/5/2019 10:01:15 AM | 1996-04-12 | 12/20/2019 11:02:20 AM | EMS | RDGIBSON | Temperature | degC | Y | (temp.) |  | Temperature (Air) (temp.) |  | Field; | Field; EMS code: 0020 | Chemical |  | FIELD_RESULT | Hydrolab
0020 | C | Temperature (Air) | XT01 | Thermister | Thermister | -20 | Y | Y | Y | 71 |  |  | 1/8/1999 3:44:22 PM | 1996-04-12 | 12/20/2019 11:02:20 AM | EMS | RDGIBSON | Temperature | degC | Y | (temp.) |  | Temperature (Air) (temp.) |  | Field; | Field; EMS code: 0020 | Chemical |  | FIELD_RESULT | Thermister
0021 | m | Water Level | EQ24 | EQUIS (m) | EQUIS (m) | 0 | N | Y | Y | 494 |  |  | 10/23/1996 12:06:54 AM | 1996-04-12 | 6/3/1998 6:51:32 PM | EMS |  | Length | m | Y | (len.) |  | Water Level (len.) |  | Field; | Field; EMS code: 0021 | Chemical |  | FIELD_RESULT | EQUIS (m)
0021 | m | Water Level | XM18 | Wetted Tape | Wetted Tape | -30 | Y | Y | Y | 3038 | 9/18/1996 | 6/25/2024 10:40:00 AM | 8/13/2024 12:09:44 PM | 1996-04-12 | 6/3/1998 6:51:32 PM | EMS |  | Length | m | Y | (len.) |  | Water Level (len.) |  | Field; | Field; EMS code: 0021 | Chemical |  | FIELD_RESULT | Wetted Tape
0021 | m | Water Level | XG05 | Staff Gauge | Staff Gauge | 0 | Y | Y | Y | 11131 | 1/31/2001 | 12/1/2018 | 1/28/2022 1:14:28 PM | 1996-04-12 | 6/3/1998 6:51:32 PM | EMS |  | Length | m | Y | (len.) |  | Water Level (len.) |  | Field; | Field; EMS code: 0021 | Chemical |  | FIELD_RESULT | Staff Gauge
0022 | m | Groundwater Elevation- height of groundwater | XM18 | Wetted Tape | Wetted Tape |  | Y | Y | Y | 8107 | 1/21/1991 9:05:00 AM | 4/11/2024 | 4/26/2024 2:32:40 PM | 2019-08-29 |  | KATFRASE |  | Length | m | Y | (len.) |  | Groundwater Elevation- height of groundwater (len.) |  | Physical Properties; | Physical Properties; EMS code: 0022 | Physical |  | LAB
0023 | ml/L | Settleable Solids | 1010 | Imhoff Cone | A well mixed sample is introduced into a graduated Imhoff cone and allowed to stand for an hour (with gentle spinning of the cone at 45 minutes to minimize entrapped pockets of water). The volume of settled residue is recorded and reported as mL/L. | 0.5 | Y | Y | Y | 52 | 1/17/1997 8:00:00 AM | 8/12/2001 8:00:00 AM | 9/26/2001 2:46:16 PM | 1996-04-12 | 6/3/1998 6:51:32 PM | EMS |  | VolumePerVolume | mL/L | Y | (vol./vol.) |  | Settleable Solids (vol./vol.) |  | Physical Properties; | Physical Properties; EMS code: 0023 | Physical |  | LAB
0023 | mg/L | Settleable Solids | F036 | Gravimetric, GF/C glass Fiber Filt, 1-hr settling period | (This method is verified for parm=0023 Settleable Solids). The gravimetric analysis is performed by determining the amount of non-settleable solids that remains in suspension after one hour; this is measured by filtering a portion of the supernatent liquid through a GF/C glass fiber filter. The amount of non-settleable solids is then subtracted from the total non-filterable residue (NFR) to give the settleable residue. | 10 | Y | Y | Y | 27 |  |  | 7/4/2007 11:10:03 AM | 1997-02-12 | 6/3/1998 6:51:46 PM | MPAK |  | Concentration | mg/L | Y | (fl. conc.) |  | Settleable Solids (fl. conc.) |  | Physical Properties; | Physical Properties; EMS code: 0023 | Physical |  | LAB
0023 | ml/L | Settleable Solids | EQ10 | EQUIS (mL/L) | EQUIS (mL/L) | 0 | N | Y | Y | 4 |  |  | 10/22/1996 10:04:46 PM | 1996-04-12 | 6/3/1998 6:51:32 PM | EMS |  | VolumePerVolume | mL/L | Y | (vol./vol.) |  | Settleable Solids (vol./vol.) |  | Physical Properties; | Physical Properties; EMS code: 0023 | Physical |  | LAB
0024 | TAC | Color TAC | EQ08 | EQUIS (TAC) | EQUIS (TAC) | 1 | N | Y | Y | 3915 |  |  | 10/24/1996 3:15:30 PM | 1996-04-12 | 6/3/1998 6:51:32 PM | EMS |  | Apperance | TAC | N | (app.) |  | Color TAC (app.) |  | Physical Properties; | Physical Properties; EMS code: 0024 | Physical |  | LAB
0024 | TAC | Color TAC | XM14 | Spectrophotometer | The pH of a filtered sample is adjusted to pH 7.6. Then the TAC is measured by integrating the absorbance of the sample between 400 and 700 nm. One unit of TAC is defined as the colour produced by 2 mg/L platinum in the form of the chloroplatinate ion. | 1 | Y | Y | Y | 4340 | 9/24/1996 2:05:00 PM | 3/22/2001 1:36:00 PM | 11/1/2002 4:49:42 AM | 1996-04-12 | 6/3/1998 6:51:32 PM | EMS |  | Apperance | TAC | N | (app.) |  | Color TAC (app.) |  | Physical Properties; | Physical Properties; EMS code: 0024 | Physical |  | LAB
0025 | %(W/W) | Moisture | X233 | Hom, Gravimetric 103-105C | The sample is homogenized, and moisture is removed by heating at 105 C, and the residue is determined gravimetrically. (Oven drying at 103 - 105 C is taken as equivalent.) | 0.1 | Y | Y | Y | 2082 | 10/1/1996 11:15:00 AM | 3/15/2024 | 3/26/2024 8:53:57 AM | 1996-04-12 | 6/3/1998 6:51:32 PM | EMS |  | DimensionlessRatio | % | N | (ratio) | (wet) | Moisture (wet) (ratio) |  | Physical Properties; | Physical Properties; EMS code: 0025 | Physical |  | LAB
0025 | % | Moisture | X175 | Gravimetric | Code X175 covers a variety of gravimetric procedures. | 0.01 | Y | Y | Y | 2171 | 9/6/1996 4:43:00 PM | 6/17/2024 | 6/18/2024 11:02:03 PM | 1996-04-12 | 6/3/1998 6:51:32 PM | EMS |  | DimensionlessRatio | % | N | (ratio) |  | Moisture (ratio) |  | Physical Properties; | Physical Properties; EMS code: 0025 | Physical |  | LAB
0025 | kPa | Moisture | EQ28 | EQUIS (kPa) | EQUIS (kPa) | 0 | N | Y | Y | 39 |  |  | 10/22/1996 10:33:47 PM | 1996-04-12 | 6/3/1998 6:51:32 PM | EMS |  | Pressure | kPa | Y | (pres.) |  | Moisture (pres.) |  | Physical Properties; | Physical Properties; EMS code: 0025 | Physical |  | LAB
0025 | % | Moisture | EQ34 | EQUIS (%) | EQUIS (%) | 0 | N | Y | Y | 597 |  |  | 10/23/1996 12:08:25 AM | 1996-04-12 | 6/3/1998 6:51:45 PM | EMS |  | DimensionlessRatio | % | N | (ratio) |  | Moisture (ratio) |  | Physical Properties; | Physical Properties; EMS code: 0025 | Physical |  | LAB
0025 | % | Moisture | F054 | Dry @60C,Hom. | Dry @ 60 C in drying Oven; Homogenization after drying | 0.1 | Y | Y | Y | 147 | 2/4/2000 9:38:00 AM | 1/2/2002 12:22:00 PM | 9/10/2002 10:16:48 AM | 1998-03-10 | 6/3/1998 6:51:48 PM | NPEPPIN |  | DimensionlessRatio | % | N | (ratio) |  | Moisture (ratio) |  | Physical Properties; | Physical Properties; EMS code: 0025 | Physical |  | LAB
0025 | %(W/W) | Moisture | 0770 | Hom.:Grav.Freeze Dry | Sample moisture is determined by weighing samples before and after freeze-drying. | 0.1 | Y | Y | Y | 109 | 3/27/1998 2:06:00 PM | 2/1/2002 2:31:00 PM | 6/23/2003 1:15:26 PM | 1996-04-12 | 6/3/1998 6:51:32 PM | EMS |  | DimensionlessRatio | % | N | (ratio) | (wet) | Moisture (wet) (ratio) |  | Physical Properties; | Physical Properties; EMS code: 0025 | Physical |  | LAB
0025 | % | Moisture | MICM | Microdigestion: Hom.,Grav <60C | Microdigestion of tissues: Homog; dry <60C; gravimetric | 2 | Y | Y | Y | 36 | 2/13/2019 | 10/2/2023 | 11/28/2023 2:21:32 PM | 2019-04-24 |  | RPWILLIA |  | DimensionlessRatio | % | N | (ratio) |  | Moisture (ratio) |  | Physical Properties; | Physical Properties; EMS code: 0025 | Physical |  | LAB
0026 | L | Vol. of groundwater removed from well prior to sample coll. | XM20 | Extrapolation from pumping rate and pumping duration. | Extrapolation from pumping rate and pumping duration. |  | Y | Y | Y | 3831 | 4/15/2005 10:00:00 AM | 12/22/2015 1:10:00 PM | 9/11/2019 8:47:50 AM | 2019-08-29 |  | KATFRASE |  | Volume | L | Y | (vol.) |  | Vol. of groundwater removed from well prior to sample coll. (vol.) |  | Field; | Field; EMS code: 0026 | Chemical |  | FIELD_RESULT | Extrapolation from pumping rate and pumping duration.
0027 | m3/s W | Rate of groundwater extraction during purging of the well. | XM21 | Timed filling of measured bucket. | Timed filling of measured bucket. |  | Y | Y | Y | 3810 | 4/15/2005 10:00:00 AM | 12/22/2015 1:10:00 PM | 9/11/2019 8:47:50 AM | 2019-08-29 |  | KATFRASE |  | Flow | m³/s | Y | (flow) |  | Rate of groundwater extraction during purging of the well. (flow) |  | Field; | Field; EMS code: 0027 | Chemical |  | FIELD_RESULT | Timed filling of measured bucket.
0028 | mg/L | Total Solids | X237 | Gravimetric | This code is a catch-all for a variety of gravimetric procedures. | 1 | Y | Y | Y | 11224 | 4/9/1999 8:00:00 AM | 8/7/2024 | 8/21/2024 11:21:21 AM | 1996-04-12 | 6/3/1998 6:51:41 PM | EMS |  | Concentration | mg/L | Y | (fl. conc.) |  | Total Solids (fl. conc.) |  | Physical Properties; | Physical Properties; EMS code: 0028 | Physical | Total | LAB
0028 | %(W/W) | Total Solids | GRAV | Gravimetric - Percent | This method is a catch-all for gravimetric procedures with units in percent. | 0.1 | Y | Y | Y | 206 | 3/5/2010 9:05:00 AM | 3/15/2024 | 3/26/2024 8:53:57 AM | 2010-04-19 | 4/20/2010 9:27:33 AM | RDGIBSON | RDGIBSON | DimensionlessRatio | % | N | (ratio) | (wet) | Total Solids (wet) (ratio) |  | Physical Properties; | Physical Properties; EMS code: 0028 | Physical | Total | LAB
0028 | % | Total Solids | E157 | APHA 2540 B (mod) | Total Solids (TS) are determined by drying a pre-weighed sample in a pre-weighed dish to constant weight in an oven at 104 ? 1?C. The increase in weight over that of the empty dish represents the total solids. Results are reported as the percentage of the total sample. | 0.1 | Y | Y | Y | 19 | 4/27/2022 | 6/12/2024 | 6/18/2024 11:02:03 PM | 2021-08-17 |  | RDGIBSON |  | DimensionlessRatio | % | N | (ratio) |  | Total Solids (ratio) |  | Physical Properties; | Physical Properties; EMS code: 0028 | Physical | Total | LAB
0029 | mg/L | Floatable Solids | EQ01 | EQUIS (mg/L) | EQUIS (mg/L) | 0 | N | Y | Y | 1908 |  |  | 10/22/1996 3:24:09 PM | 1996-04-12 | 6/3/1998 6:51:32 PM | EMS |  | Concentration | mg/L | Y | (fl. conc.) |  | Floatable Solids (fl. conc.) |  | Physical Properties; | Physical Properties; EMS code: 0029 | Physical |  | LAB
0032 | % | Residue: Tot.Volatile (Total Solids; Volatile) | EQ34 | EQUIS (%) | EQUIS (%) | 0 | N | Y | Y | 198 |  |  | 10/23/1996 12:08:28 AM | 1996-04-12 | 6/3/1998 6:51:45 PM | EMS |  | DimensionlessRatio | % | N | (ratio) |  | Residue: Tot.Volatile (Total Solids; Volatile) (ratio) |  | Physical Properties; | Physical Properties; EMS code: 0032 | Physical | Total | LAB
0032 | % | Residue: Tot.Volatile (Total Solids; Volatile) | X048 | Gravimetric | Gravimetric | 0.1 | Y | Y | Y | 30 |  |  | 1/29/2024 9:21:02 PM | 1996-04-12 | 6/3/1998 6:51:42 PM | EMS |  | DimensionlessRatio | % | N | (ratio) |  | Residue: Tot.Volatile (Total Solids; Volatile) (ratio) |  | Physical Properties; | Physical Properties; EMS code: 0032 | Physical | Total | LAB
0032 | %(W/W) | Residue: Tot.Volatile (Total Solids; Volatile) | 0750 | D/G:/Grav. 550C | D/G:/Grav. 550C | 0.1 | Y | Y | Y | 708 | 3/5/2010 9:05:00 AM | 3/15/2024 | 3/26/2024 8:53:57 AM | 1996-04-12 | 6/3/1998 6:51:41 PM | EMS |  | DimensionlessRatio | % | N | (ratio) | (wet) | Residue: Tot.Volatile (Total Solids; Volatile) (wet) (ratio) |  | Physical Properties; | Physical Properties; EMS code: 0032 | Physical | Total | LAB
0032 | % | Residue: Tot.Volatile (Total Solids; Volatile) | X175 | Gravimetric | Code X175 covers a variety of gravimetric procedures. | 0 | Y | Y | Y | 9 |  |  | 5/15/1997 10:13:08 AM | 1996-04-12 | 6/3/1998 6:51:32 PM | EMS |  | DimensionlessRatio | % | N | (ratio) |  | Residue: Tot.Volatile (Total Solids; Volatile) (ratio) |  | Physical Properties; | Physical Properties; EMS code: 0032 | Physical | Total | LAB
0032 | ug/g | Residue: Tot.Volatile (Total Solids; Volatile) | X479 | Ignition,550 C | Sediment Fixed Residue (SFR) represents the solids or ash remaining after ignition at 550 C. Sediment Volatile Residue (SVR) is the difference between the weight lost on ignition at 550 C and the weight of the sample dried at 105 C. SVR is a rough approximation of the amount of organic matter in the sediment. | 500 | Y | Y | Y | 1052 | 3/14/1997 6:31:00 PM | 1/30/2002 7:14:00 AM | 5/28/2003 2:34:17 PM | 1997-11-04 | 6/3/1998 6:51:47 PM | NPEPPIN |  | MassPerMass | mg/kg | Y | (mass conc.) |  | Residue: Tot.Volatile (Total Solids; Volatile) (mass conc.) |  | Physical Properties; | Physical Properties; EMS code: 0032 | Physical | Total | LAB
0032 | % | Residue: Tot.Volatile (Total Solids; Volatile) | XCAL | Calculation (mg/L) | Calculation (mg/L) |  | Y | Y | Y | 19 | 5/2/2022 | 6/13/2024 | 6/18/2024 11:02:03 PM | 2022-05-30 |  | EMS |  | DimensionlessRatio | % | N | (ratio) |  | Residue: Tot.Volatile (Total Solids; Volatile) (ratio) |  | Physical Properties; | Physical Properties; EMS code: 0032 | Physical | Total | LAB
0032 | mg/L | Residue: Tot.Volatile (Total Solids; Volatile) | 1940 | Gravimetric 550C | For Total Volatile Residue (parm=0032): A known volume of a well-mixed sample is transferred into an pre-ignited porcelain dish, evaporated at 90 - 95 C, dried at 103 C, cooled and weighed. The porcelain dish with dried and weighed residue is then ignited in a muffle at 550 C and weighed again. The difference in weights represents Total Volatile Residue (TVR). The Total Fixed Residue (TFR) is the portion of sample left after ignited at 550 C. | 10 | Y | Y | Y | 152 | 11/6/1997 6:31:00 PM | 2/4/2010 12:15:00 PM | 3/29/2010 2:25:26 PM | 1997-02-27 | 6/3/1998 6:51:46 PM | MPAK |  | Concentration | mg/L | Y | (fl. conc.) |  | Residue: Tot.Volatile (Total Solids; Volatile) (fl. conc.) |  | Physical Properties; | Physical Properties; EMS code: 0032 | Physical | Total | LAB
0032 | %(W/W) | Residue: Tot.Volatile (Total Solids; Volatile) | X237 | Gravimetric | This code is a catch-all for a variety of gravimetric procedures. | 0.1 | Y | Y | Y | 579 | 9/12/1996 1:05:00 PM | 9/22/2004 8:50:00 AM | 2/24/2012 11:58:35 AM | 1996-04-12 | 6/3/1998 6:51:32 PM | EMS |  | DimensionlessRatio | % | N | (ratio) | (wet) | Residue: Tot.Volatile (Total Solids; Volatile) (wet) (ratio) |  | Physical Properties; | Physical Properties; EMS code: 0032 | Physical | Total | LAB
0032 | mg/L | Residue: Tot.Volatile (Total Solids; Volatile) | CS01 | System Calculation | System Calculation (mg/L) |  | N | Y | N | 2 |  |  | 10/15/1997 1:04:43 PM | 1996-04-12 | 6/3/1998 6:51:42 PM | EMS |  | Concentration | mg/L | Y | (fl. conc.) |  | Residue: Tot.Volatile (Total Solids; Volatile) (fl. conc.) |  | Physical Properties; | Physical Properties; EMS code: 0032 | Physical | Total | LAB
0033 | %(W/W) | Part. Size 16 Mesh | 0660 | 16 Mesh: 1.19 mm (dry) | 16 Mesh: 1.19 mm (dry) | 0.1 | Y | Y | Y | 336 | 2/19/1997 12:37:00 PM | 2/19/1997 3:40:00 PM | 2/24/1997 7:16:40 PM | 1996-04-12 | 6/3/1998 6:51:32 PM | EMS |  | DimensionlessRatio | % | N | (ratio) | (wet) | Part. Size 16 Mesh (wet) (ratio) |  | Sediment Size; | Sediment Size; EMS code: 0033 | Physical |  | LAB
0033 | %(W/W) | Part. Size 16 Mesh | 0661 | 16 Mesh: 1.19 mm (wet) | 16 Mesh: 1.19 mm (wet) | 0.1 | Y | Y | Y | 74 |  |  | 10/24/1996 7:19:03 AM | 1996-04-12 | 6/3/1998 6:51:32 PM | EMS |  | DimensionlessRatio | % | N | (ratio) | (wet) | Part. Size 16 Mesh (wet) (ratio) |  | Sediment Size; | Sediment Size; EMS code: 0033 | Physical |  | LAB
0033 | % | Part. Size 16 Mesh | EQ34 | EQUIS (%) | EQUIS (%) | 0 | N | Y | Y | 34 |  |  | 10/22/1996 10:25:21 PM | 1996-04-12 | 6/3/1998 6:51:45 PM | EMS |  | DimensionlessRatio | % | N | (ratio) |  | Part. Size 16 Mesh (ratio) |  | Sediment Size; | Sediment Size; EMS code: 0033 | Physical |  | LAB
0039 | mV | Oxid. Red. Potential | XH01 | Hydrolab Ag/AgCl | Hydrolab Ag/AgCl | -500 | Y | Y | Y | 11921 | 4/30/2010 9:22:00 AM | 12/7/2011 | 3/22/2019 9:52:29 AM | 1996-04-12 | 12/20/2019 11:02:20 AM | EMS | RDGIBSON | ElectricPotential | mV | Y | (elec. pot.) |  | Oxid. Red. Potential (elec. pot.) |  | Physical Properties; | Physical Properties; EMS code: 0039 | Physical |  | FIELD_RESULT | Hydrolab Ag/AgCl
0039 | mV | Oxid. Red. Potential | EQ33 | EQUIS (mV) | EQUIS (mV) | -500 | N | Y | Y | 1829 |  |  | 10/22/1996 11:28:37 PM | 1996-04-12 | 6/3/1998 6:51:32 PM | EMS |  | ElectricPotential | mV | Y | (elec. pot.) |  | Oxid. Red. Potential (elec. pot.) |  | Physical Properties; | Physical Properties; EMS code: 0039 | Physical |  | LAB
0039 | mV | Oxid. Red. Potential | CSFT | Client Supplied Field Test Result | Client Supplied Field Test Result - Generic | -1000 | Y | Y | Y | 1110 | 8/4/2015 | 7/8/2019 | 7/30/2024 9:20:20 AM | 2015-11-05 |  | RDGIBSON |  | ElectricPotential | mV | Y | (elec. pot.) |  | Oxid. Red. Potential (elec. pot.) |  | Physical Properties; | Physical Properties; EMS code: 0039 | Physical |  | FIELD_RESULT | Client Supplied Field Test Result
0039 | mV | Oxid. Red. Potential | ORP1 | Oxidation-Reduction Potential of Water | This analysis is carried out in accordance with the procedure described in the "ASTM" method D1498-00 "Oxidation-Reduction Potential of Water" published by the American Society for Testing and Materials (ASTM), May 2000. Results are reported as observed oxidation-reduction potential of the platinum metal-reference electrode employed, in mV. | -1000 | Y | Y | Y | 5529 | 7/14/2010 | 8/16/2024 | 8/20/2024 5:53:04 PM | 2016-02-15 |  | RDGIBSON |  | ElectricPotential | mV | Y | (elec. pot.) |  | Oxid. Red. Potential (elec. pot.) |  | Physical Properties; | Physical Properties; EMS code: 0039 | Physical |  | LAB
003A | %(W/W) | Part. Size 30 Mesh | 0662 | 30 Mesh: 0.59 mm (wet) | 30 Mesh: 0.59 mm (wet) | 0.1 | Y | Y | Y | 74 |  |  | 10/24/1996 7:19:03 AM | 1996-04-12 | 6/3/1998 6:51:45 PM | EMS |  | DimensionlessRatio | % | N | (ratio) | (wet) | Part. Size 30 Mesh (wet) (ratio) |  | Sediment Size; | Sediment Size; EMS code: 003A | Physical |  | LAB
003A | % | Part. Size 30 Mesh | EQ34 | EQUIS (%) | EQUIS (%) | 0 | N | Y | Y | 17 |  |  | 10/22/1996 10:25:21 PM | 1996-04-12 | 6/3/1998 6:51:32 PM | EMS |  | DimensionlessRatio | % | N | (ratio) |  | Part. Size 30 Mesh (ratio) |  | Sediment Size; | Sediment Size; EMS code: 003A | Physical |  | LAB
003A | %(W/W) | Part. Size 30 Mesh | 0670 | 30 Mesh: 0.59 mm (dry) | 30 Mesh: 0.59 mm (dry) | 0.1 | Y | Y | Y | 601 | 2/19/1997 12:37:00 PM | 7/6/2004 9:45:00 AM | 8/12/2004 1:08:24 AM | 1996-04-12 | 6/3/1998 6:51:32 PM | EMS |  | DimensionlessRatio | % | N | (ratio) | (wet) | Part. Size 30 Mesh (wet) (ratio) |  | Sediment Size; | Sediment Size; EMS code: 003A | Physical |  | LAB
003B | % | Part. Size 50 Mesh | EQ34 | EQUIS (%) | EQUIS (%) | 0 | N | Y | Y | 25 |  |  | 10/22/1996 10:25:21 PM | 1996-04-12 | 6/3/1998 6:51:32 PM | EMS |  | DimensionlessRatio | % | N | (ratio) |  | Part. Size 50 Mesh (ratio) |  | Sediment Size; | Sediment Size; EMS code: 003B | Physical |  | LAB
003B | %(W/W) | Part. Size 50 Mesh | 0680 | 50 Mesh: 0.297 mm (dry) | 50 Mesh: 0.297 mm (dry) | 0.1 | Y | Y | Y | 659 | 2/19/1997 12:37:00 PM | 7/6/2004 9:45:00 AM | 8/12/2004 1:08:25 AM | 1996-04-12 | 6/3/1998 6:51:45 PM | EMS |  | DimensionlessRatio | % | N | (ratio) | (wet) | Part. Size 50 Mesh (wet) (ratio) |  | Sediment Size; | Sediment Size; EMS code: 003B | Physical |  | LAB
003B | %(W/W) | Part. Size 50 Mesh | 0663 | 50 Mesh: 0.297 mm (wet) | 50 Mesh: 0.297 mm (wet) | 0.1 | Y | Y | Y | 78 |  |  | 10/24/1996 9:35:47 AM | 1996-04-12 | 6/3/1998 6:51:32 PM | EMS |  | DimensionlessRatio | % | N | (ratio) | (wet) | Part. Size 50 Mesh (wet) (ratio) |  | Sediment Size; | Sediment Size; EMS code: 003B | Physical |  | LAB
003B | %(W/W) | Part. Size 50 Mesh | X224 | 50 Mesh: <0.3 mm (dry) | 50 Mesh: <0.3 mm (dry) | 0.1 | Y | Y | Y | 80 |  |  | 10/24/1996 10:24:19 AM | 1996-04-12 | 6/3/1998 6:51:32 PM | EMS |  | DimensionlessRatio | % | N | (ratio) | (wet) | Part. Size 50 Mesh (wet) (ratio) |  | Sediment Size; | Sediment Size; EMS code: 003B | Physical |  | LAB
003C | %(W/W) | Part. Size 100 Mesh | 0690 | 100 Mesh: 0.149 mm (dry) | 100 Mesh: 0.149 mm (dry) | 0.1 | Y | Y | Y | 601 | 2/19/1997 12:38:00 PM | 7/6/2004 9:45:00 AM | 8/12/2004 1:08:25 AM | 1996-04-12 | 6/3/1998 6:51:45 PM | EMS |  | DimensionlessRatio | % | N | (ratio) | (wet) | Part. Size 100 Mesh (wet) (ratio) |  | Sediment Size; | Sediment Size; EMS code: 003C | Physical |  | LAB
003C | % | Part. Size 100 Mesh | EQ34 | EQUIS (%) | EQUIS (%) | 0 | N | Y | Y | 17 |  |  | 10/22/1996 10:25:21 PM | 1996-04-12 | 6/3/1998 6:51:32 PM | EMS |  | DimensionlessRatio | % | N | (ratio) |  | Part. Size 100 Mesh (ratio) |  | Sediment Size; | Sediment Size; EMS code: 003C | Physical |  | LAB
003C | %(W/W) | Part. Size 100 Mesh | 0664 | 100 Mesh: 0.149 mm (wet) | 100 Mesh: 0.149 mm (wet) | 0.1 | Y | Y | Y | 74 |  |  | 10/24/1996 7:19:03 AM | 1996-04-12 | 6/3/1998 6:51:32 PM | EMS |  | DimensionlessRatio | % | N | (ratio) | (wet) | Part. Size 100 Mesh (wet) (ratio) |  | Sediment Size; | Sediment Size; EMS code: 003C | Physical |  | LAB
003D | %(W/W) | Part. Size 140 Mesh | 0665 | 140 Mesh: 0.105 mm (wet) | 140 Mesh: 0.105 mm (wet) | 0.1 | Y | Y | Y | 74 |  |  | 10/24/1996 7:19:03 AM | 1996-04-12 | 6/3/1998 6:51:32 PM | EMS |  | DimensionlessRatio | % | N | (ratio) | (wet) | Part. Size 140 Mesh (wet) (ratio) |  | Sediment Size; | Sediment Size; EMS code: 003D | Physical |  | LAB
003D | %(W/W) | Part. Size 140 Mesh | 0700 | 140 Mesh: 0.105 mm (dry) | 140 Mesh: 0.105 mm (dry) | 0.1 | Y | Y | Y | 335 | 2/19/1997 12:38:00 PM | 2/19/1997 3:40:00 PM | 2/24/1997 7:16:41 PM | 1996-04-12 | 6/3/1998 6:51:32 PM | EMS |  | DimensionlessRatio | % | N | (ratio) | (wet) | Part. Size 140 Mesh (wet) (ratio) |  | Sediment Size; | Sediment Size; EMS code: 003D | Physical |  | LAB
003D | % | Part. Size 140 Mesh | EQ34 | EQUIS (%) | EQUIS (%) | 0 | N | Y | Y | 17 |  |  | 10/22/1996 10:25:21 PM | 1996-04-12 | 6/3/1998 6:51:45 PM | EMS |  | DimensionlessRatio | % | N | (ratio) |  | Part. Size 140 Mesh (ratio) |  | Sediment Size; | Sediment Size; EMS code: 003D | Physical |  | LAB
003E | %(W/W) | Part. Size 200 Mesh | 0666 | 200 Mesh: 0.074 mm (wet) | 200 Mesh: 0.074 mm (wet) | 0.1 | Y | Y | Y | 78 |  |  | 10/24/1996 9:35:47 AM | 1996-04-12 | 6/3/1998 6:51:32 PM | EMS |  | DimensionlessRatio | % | N | (ratio) | (wet) | Part. Size 200 Mesh (wet) (ratio) |  | Sediment Size; | Sediment Size; EMS code: 003E | Physical |  | LAB
003E | % | Part. Size 200 Mesh | EQ34 | EQUIS (%) | EQUIS (%) | 0 | N | Y | Y | 34 |  |  | 10/22/1996 10:25:21 PM | 1996-04-12 | 6/3/1998 6:51:32 PM | EMS |  | DimensionlessRatio | % | N | (ratio) |  | Part. Size 200 Mesh (ratio) |  | Sediment Size; | Sediment Size; EMS code: 003E | Physical |  | LAB
003E | %(W/W) | Part. Size 200 Mesh | 0710 | 200 Mesh: 0.074 mm (dry) | 200 Mesh: 0.074 mm (dry) | 0.1 | Y | Y | Y | 496 | 2/19/1997 12:38:00 PM | 4/3/2002 9:40:00 AM | 2/26/2004 1:05:25 AM | 1996-04-12 | 6/3/1998 6:51:32 PM | EMS |  | DimensionlessRatio | % | N | (ratio) | (wet) | Part. Size 200 Mesh (wet) (ratio) |  | Sediment Size; | Sediment Size; EMS code: 003E | Physical |  | LAB
003F | %(W/W) | Part. Size 270 Mesh | 0720 | 270 Mesh: 0.053 mm (dry) | 270 Mesh: 0.053 mm (dry) | 0.1 | Y | Y | Y | 601 | 2/19/1997 12:38:00 PM | 7/6/2004 9:45:00 AM | 8/12/2004 1:08:25 AM | 1996-04-12 | 6/3/1998 6:51:33 PM | EMS |  | DimensionlessRatio | % | N | (ratio) | (wet) | Part. Size 270 Mesh (wet) (ratio) |  | Sediment Size; | Sediment Size; EMS code: 003F | Physical |  | LAB
003F | %(W/W) | Part. Size 270 Mesh | 0667 | 270 Mesh: 0.053 mm (wet) | 270 Mesh: 0.053 mm (wet) | 0.1 | Y | Y | Y | 74 |  |  | 10/24/1996 7:19:03 AM | 1996-04-12 | 6/3/1998 6:51:32 PM | EMS |  | DimensionlessRatio | % | N | (ratio) | (wet) | Part. Size 270 Mesh (wet) (ratio) |  | Sediment Size; | Sediment Size; EMS code: 003F | Physical |  | LAB
003F | % | Part. Size 270 Mesh | EQ34 | EQUIS (%) | EQUIS (%) | 0 | N | Y | Y | 17 |  |  | 10/22/1996 10:25:21 PM | 1996-04-12 | 6/3/1998 6:51:33 PM | EMS |  | DimensionlessRatio | % | N | (ratio) |  | Part. Size 270 Mesh (ratio) |  | Sediment Size; | Sediment Size; EMS code: 003F | Physical |  | LAB
003G | % | Part. Size 400 Mesh | EQ34 | EQUIS (%) | EQUIS (%) | 0 | N | Y | Y | 26 |  |  | 10/22/1996 10:25:21 PM | 1996-04-12 | 6/3/1998 6:51:33 PM | EMS |  | DimensionlessRatio | % | N | (ratio) |  | Part. Size 400 Mesh (ratio) |  | Sediment Size; | Sediment Size; EMS code: 003G | Physical |  | LAB
003G | %(W/W) | Part. Size 400 Mesh | 0730 | 400 Mesh: 0.037 mm (dry) | 400 Mesh: 0.037 mm (dry) | 0.1 | Y | Y | Y | 601 | 2/19/1997 12:38:00 PM | 7/6/2004 9:45:00 AM | 8/12/2004 1:08:25 AM | 1996-04-12 | 6/3/1998 6:51:33 PM | EMS |  | DimensionlessRatio | % | N | (ratio) | (wet) | Part. Size 400 Mesh (wet) (ratio) |  | Sediment Size; | Sediment Size; EMS code: 003G | Physical |  | LAB
003G | %(W/W) | Part. Size 400 Mesh | 0668 | 400 Mesh: 0.037 mm (wet) | 400 Mesh: 0.037 mm (wet) | 0.1 | Y | Y | Y | 74 |  |  | 10/24/1996 7:19:03 AM | 1996-04-12 | 6/3/1998 6:51:33 PM | EMS |  | DimensionlessRatio | % | N | (ratio) | (wet) | Part. Size 400 Mesh (wet) (ratio) |  | Sediment Size; | Sediment Size; EMS code: 003G | Physical |  | LAB
003H | %(W/W) | Part. Size >400 Mesh | 0740 | Pan: <0.037 mm (dry) | Pan: <0.037 mm (dry) | 0.1 | Y | Y | Y | 659 | 2/19/1997 12:38:00 PM | 7/6/2004 9:45:00 AM | 8/12/2004 1:08:25 AM | 1996-04-12 | 6/3/1998 6:51:33 PM | EMS |  | DimensionlessRatio | % | N | (ratio) | (wet) | Part. Size >400 Mesh (wet) (ratio) |  | Sediment Size; | Sediment Size; EMS code: 003H | Physical |  | LAB
003H | % | Part. Size >400 Mesh | EQ34 | EQUIS (%) | EQUIS (%) | 0 | N | Y | Y | 34 |  |  | 10/22/1996 10:25:21 PM | 1996-04-12 | 6/3/1998 6:51:33 PM | EMS |  | DimensionlessRatio | % | N | (ratio) |  | Part. Size >400 Mesh (ratio) |  | Sediment Size; | Sediment Size; EMS code: 003H | Physical |  | LAB
003H | %(W/W) | Part. Size >400 Mesh | 0669 | Pan: <0.037 mm (wet) | Pan: <0.037 mm (wet) | 0.1 | Y | Y | Y | 74 |  |  | 10/24/1996 7:19:04 AM | 1996-04-12 | 6/3/1998 6:51:33 PM | EMS |  | DimensionlessRatio | % | N | (ratio) | (wet) | Part. Size >400 Mesh (wet) (ratio) |  | Sediment Size; | Sediment Size; EMS code: 003H | Physical |  | LAB
003K | %(W/W) | Part. Size 10 Mesh | X278 | 10 Mesh: > 2 mm (dry) | 10 Mesh: > 2 mm (dry) | 0.1 | Y | Y | Y | 308 | 5/27/2002 3:30:00 PM | 7/6/2004 9:45:00 AM | 8/12/2004 1:08:25 AM | 1996-04-12 | 6/3/1998 6:51:33 PM | EMS |  | DimensionlessRatio | % | N | (ratio) | (wet) | Part. Size 10 Mesh (wet) (ratio) |  | Sediment Size; | Sediment Size; EMS code: 003K | Physical |  | LAB
003L | %(W/W) | Part. Size 230 Mesh | X279 | 230 Mesh:2mm-.063mm(wet) | 230 Mesh:2mm-.063mm(wet) | 0.1 | Y | Y | Y | 64 | 10/27/1998 1:12:00 PM | 11/15/2001 8:02:00 AM | 11/26/2002 11:02:00 AM | 1996-04-12 | 6/3/1998 6:51:45 PM | EMS |  | DimensionlessRatio | % | N | (ratio) | (wet) | Part. Size 230 Mesh (wet) (ratio) |  | Sediment Size; | Sediment Size; EMS code: 003L | Physical |  | LAB
003L | %(W/W) | Part. Size 230 Mesh | X280 | 230 Mesh:2mm-.063mm(dry) | 230 Mesh:2mm-.063mm(dry) | 0.1 | Y | Y | Y | 42 |  |  | 10/24/1996 2:25:26 PM | 1996-04-12 | 6/3/1998 6:51:33 PM | EMS |  | DimensionlessRatio | % | N | (ratio) | (wet) | Part. Size 230 Mesh (wet) (ratio) |  | Sediment Size; | Sediment Size; EMS code: 003L | Physical |  | LAB
003M | %(W/W) | Part. Size 4000 Mesh | X282 | 4000 Mesh:.063mm-4um(dry | 4000 Mesh:.063mm-4um(dry | 0.1 | Y | Y | Y | 42 |  |  | 10/24/1996 2:25:26 PM | 1996-04-12 | 6/3/1998 6:51:33 PM | EMS |  | DimensionlessRatio | % | N | (ratio) | (wet) | Part. Size 4000 Mesh (wet) (ratio) |  | Sediment Size; | Sediment Size; EMS code: 003M | Physical |  | LAB
003N | %(W/W) | Part. Size>4000 Mesh | X284 | >4000 Mesh: < 4 um (dry) | >4000 Mesh: < 4 um (dry) | 0.1 | Y | Y | Y | 42 |  |  | 10/24/1996 2:25:26 PM | 1996-04-12 | 6/3/1998 6:51:33 PM | EMS |  | DimensionlessRatio | % | N | (ratio) | (wet) | Part. Size>4000 Mesh (wet) (ratio) |  | Sediment Size; | Sediment Size; EMS code: 003N | Physical |  | LAB
0041 | m/s | Water Velocity | XM13 | Velocity Meter | Velocity Meter | 0 | Y | Y | Y | 4 |  |  | 7/17/2015 11:47:21 AM | 1996-04-12 | 6/3/1998 6:51:33 PM | EMS |  | LinearVelocity | m/s | Y | (vel.) |  | Water Velocity (vel.) |  | Physical Properties; | Physical Properties; EMS code: 0041 | Physical |  | LAB
0041 | m/s | Water Velocity | EQ31 | EQUIS (m/s) | EQUIS (m/s) | 0 | N | Y | Y | 24 |  |  | 10/22/1996 7:31:35 PM | 1996-04-12 | 6/3/1998 6:51:33 PM | EMS |  | LinearVelocity | m/s | Y | (vel.) |  | Water Velocity (vel.) |  | Physical Properties; | Physical Properties; EMS code: 0041 | Physical |  | LAB
0041 | m/s | Water Velocity | XM06 | Clock Estimation | Clock Estimation | 0 | Y | Y | Y | 19 |  |  | 10/24/1996 11:36:04 AM | 1996-04-12 | 6/3/1998 6:51:33 PM | EMS |  | LinearVelocity | m/s | Y | (vel.) |  | Water Velocity (vel.) |  | Physical Properties; | Physical Properties; EMS code: 0041 | Physical |  | LAB
0042 | mV | Oxid. Red. Potential- Field | FLD | Field sampling | Field sampling method | -1000 | Y | Y | Y | 2627 | 3/20/2018 | 8/22/2024 | 8/26/2024 4:51:04 PM | 2019-08-14 | 12/20/2019 10:20:58 AM | KATFRASE | RDGIBSON | ElectricPotential | mV | Y | (elec. pot.) |  | Oxid. Red. Potential- Field (elec. pot.) |  | Field; | Field; EMS code: 0042 | Chemical |  | FIELD_RESULT | Field sampling
0042 | mV | Oxid. Red. Potential- Field | CSFT | Client Supplied Field Test Result | Client Supplied Field Test Result - Generic | -1000 | Y | Y | Y | 8529 | 9/10/2018 10:36:00 AM | 5/15/2024 3:35:00 PM | 8/20/2024 7:35:41 AM | 2019-07-10 | 7/10/2019 4:05:50 PM | KATFRASE | KATFRASE | ElectricPotential | mV | Y | (elec. pot.) |  | Oxid. Red. Potential- Field (elec. pot.) |  | Field; | Field; EMS code: 0042 | Chemical |  | FIELD_RESULT | Client Supplied Field Test Result
0101 | mg/L | Alkalinity; Phenolphthalein (as CaCO3) | X310 | Auto Potentiometric Titration | For alkalinity, the samples are (if pH >8.3) first titrated to a pH of 8.3, which corresponds to the equivalence point for the conversion of carbonate to bicarbonate ion (phenolphthalein alkalinity). The titration is continued to a pH of 4.5, which corresponds approx. to the equivalence point for the conversion of bicarbonate ion to carbonic acid (entire titration indicates total alkalinity). An extrapolation technique based on the near proportionality of concentration of H ions to excess of titrant beyond the equivalence pt. The amount of standard acid required to reduce pH exactly 0.30pH (ie. pH4.2) unit is measured carefully. Because this change in pH corresponds to an exact coupling of the H ion concentration, a simple extrapolation can be made to the equivalence point. The actual pH of the stoichiometric end point in alkalinity determinations is the pH at which the inflection in the titration curve occurs and is directly related to the amount of carbonate alkalinity originally present in the sample. A curve can be constructed by plotting pH vs mL of acid to find this point directly. *** For acidity, the samples are titrated to a pH of 8.3, which corresponds to stoichiometric neutralization of carbonic acid to bicarbonate. Because the colour change of phenolphthalein indicator is close to pH 8.3, this value generally is accepted as a standard end point for titration of total acidity, including CO2 and most weak acids an equivalence point for the conversion of carbonate to bicarbonate ion (phenolphthalein alkalinity). The actual pH of the stoichiometric end point in acidity determinations is the pH at which the inflection in the titration curve occurs and is directly related to the amount of titratable acids originally present in the sample. A curve can be constructed by plotting pH vs. mL of NaOH to find this point directly. (To obtain the actual substances contributing to acidity/alkalinity, a pH vs mL of acid/NaOH curve should be constructed.) | 0.5 | Y | Y | Y | 34740 | 8/2/1996 4:17:00 PM | 8/22/2024 12:47:00 PM | 8/26/2024 8:06:03 AM | 1996-06-19 | 6/3/1998 6:51:44 PM | MPAK |  | Concentration | mg/L | Y | (fl. conc.) |  | Alkalinity; Phenolphthalein (as CaCO3) (fl. conc.) |  | Major Ions; | Major Ions; EMS code: 0101 | Chemical |  | LAB
0101 | mg/L | Alkalinity; Phenolphthalein (as CaCO3) | 1200 | Manual Potentiometric Titration | Manual Potentiometric Titration with standard solution to specified pH. The sample must not be filtered, diluted, concentrated, or altered in any way. | 0.5 | Y | Y | Y | 7 | 9/29/2000 | 9/29/2000 | 10/5/2010 10:15:37 AM | 1996-04-12 | 6/3/1998 6:51:33 PM | EMS |  | Concentration | mg/L | Y | (fl. conc.) |  | Alkalinity; Phenolphthalein (as CaCO3) (fl. conc.) |  | Major Ions; | Major Ions; EMS code: 0101 | Chemical |  | LAB
0101 | mg/L | Alkalinity; Phenolphthalein (as CaCO3) | X330 | Meter | The conductance (= conductivity) of a sample is measured by use of a self-contained conductivity meter. pH is measured with a combined glass/reference (Ag/AgCl) electrode. For both conductivity and pH the measurements are made at 25 C, or the results are temperature-corrected to 25 C. Note that laboratory and field measurements are distinguished by the EMS Analytical Agency codes. Ex situ and in situ measurements are distinguished by the computer codes. | 0.5 | Y | Y | Y | 2755 |  |  | 2/13/1998 12:54:09 AM | 1996-04-12 | 6/3/1998 6:51:45 PM | EMS |  | Concentration | mg/L | Y | (fl. conc.) |  | Alkalinity; Phenolphthalein (as CaCO3) (fl. conc.) |  | Major Ions; | Major Ions; EMS code: 0101 | Chemical |  | LAB
0101 | mg/L | Alkalinity; Phenolphthalein (as CaCO3) | EQ01 | EQUIS (mg/L) | EQUIS (mg/L) | 0.5 | N | Y | Y | 3631 |  |  | 10/24/1996 3:16:20 PM | 1996-04-12 | 6/3/1998 6:51:33 PM | EMS |  | Concentration | mg/L | Y | (fl. conc.) |  | Alkalinity; Phenolphthalein (as CaCO3) (fl. conc.) |  | Major Ions; | Major Ions; EMS code: 0101 | Chemical |  | LAB
0101 | mg/L | Alkalinity; Phenolphthalein (as CaCO3) | X004 | Colorimetric Titr'n | Manual Colorimetric Titration with standard solutions to specified endpoint as identified by colorimetric indicator. The sample must not be filtered, diluted, concentrated, or altered in any way. | 0.5 | Y | Y | Y | 674 | 2/21/2003 | 1/12/2015 | 6/7/2024 5:30:18 PM | 1996-04-12 | 6/3/1998 6:51:33 PM | EMS |  | Concentration | mg/L | Y | (fl. conc.) |  | Alkalinity; Phenolphthalein (as CaCO3) (fl. conc.) |  | Major Ions; | Major Ions; EMS code: 0101 | Chemical |  | LAB
0102 | mg/L | Alkalinity Total 4.5 (as CaCO3) | X045 | Methyl Orange Titr'n | Methyl Orange Titr'n | 0.5 | Y | Y | Y | 1103 | 1/21/2011 | 9/5/2017 | 12/8/2023 4:30:26 PM | 1996-04-12 | 6/3/1998 6:51:33 PM | EMS |  | Concentration | mg/L | Y | (fl. conc.) |  | Alkalinity Total 4.5 (as CaCO3) (fl. conc.) |  | Major Ions; | Major Ions; EMS code: 0102 | Chemical | Total | LAB
0102 | mg/L | Alkalinity Total 4.5 (as CaCO3) | X330 | Meter | The conductance (= conductivity) of a sample is measured by use of a self-contained conductivity meter. pH is measured with a combined glass/reference (Ag/AgCl) electrode. For both conductivity and pH the measurements are made at 25 C, or the results are temperature-corrected to 25 C. Note that laboratory and field measurements are distinguished by the EMS Analytical Agency codes. Ex situ and in situ measurements are distinguished by the computer codes. | 0.5 | Y | Y | Y | 19503 | 8/1/2000 | 10/11/2012 10:04:00 AM | 8/26/2024 2:41:32 PM | 1996-04-12 | 6/3/1998 6:51:46 PM | EMS |  | Concentration | mg/L | Y | (fl. conc.) |  | Alkalinity Total 4.5 (as CaCO3) (fl. conc.) |  | Major Ions; | Major Ions; EMS code: 0102 | Chemical | Total | LAB
0102 | mg/L | Alkalinity Total 4.5 (as CaCO3) | X181 | Manual Titr'n,St. Acid | Manual Titr'n;St. Acid | 0.5 | Y | Y | Y | 2815 | 1/9/2007 3:18:00 PM | 4/3/2013 5:52:00 PM | 7/23/2024 11:41:51 AM | 1996-04-12 | 6/3/1998 6:51:33 PM | EMS |  | Concentration | mg/L | Y | (fl. conc.) |  | Alkalinity Total 4.5 (as CaCO3) (fl. conc.) |  | Major Ions; | Major Ions; EMS code: 0102 | Chemical | Total | LAB
0102 | mg/L | Alkalinity Total 4.5 (as CaCO3) | X310 | Auto Potentiometric Titration | For alkalinity, the samples are (if pH >8.3) first titrated to a pH of 8.3, which corresponds to the equivalence point for the conversion of carbonate to bicarbonate ion (phenolphthalein alkalinity). The titration is continued to a pH of 4.5, which corresponds approx. to the equivalence point for the conversion of bicarbonate ion to carbonic acid (entire titration indicates total alkalinity). An extrapolation technique based on the near proportionality of concentration of H ions to excess of titrant beyond the equivalence pt. The amount of standard acid required to reduce pH exactly 0.30pH (ie. pH4.2) unit is measured carefully. Because this change in pH corresponds to an exact coupling of the H ion concentration, a simple extrapolation can be made to the equivalence point. The actual pH of the stoichiometric end point in alkalinity determinations is the pH at which the inflection in the titration curve occurs and is directly related to the amount of carbonate alkalinity originally present in the sample. A curve can be constructed by plotting pH vs mL of acid to find this point directly. *** For acidity, the samples are titrated to a pH of 8.3, which corresponds to stoichiometric neutralization of carbonic acid to bicarbonate. Because the colour change of phenolphthalein indicator is close to pH 8.3, this value generally is accepted as a standard end point for titration of total acidity, including CO2 and most weak acids an equivalence point for the conversion of carbonate to bicarbonate ion (phenolphthalein alkalinity). The actual pH of the stoichiometric end point in acidity determinations is the pH at which the inflection in the titration curve occurs and is directly related to the amount of titratable acids originally present in the sample. A curve can be constructed by plotting pH vs. mL of NaOH to find this point directly. (To obtain the actual substances contributing to acidity/alkalinity, a pH vs mL of acid/NaOH curve should be constructed.) | 0.5 | Y | Y | Y | 103038 | 6/4/1997 | 8/22/2024 12:47:00 PM | 8/26/2024 4:53:05 PM | 1996-06-19 | 6/3/1998 6:51:44 PM | MPAK |  | Concentration | mg/L | Y | (fl. conc.) |  | Alkalinity Total 4.5 (as CaCO3) (fl. conc.) |  | Major Ions; | Major Ions; EMS code: 0102 | Chemical | Total | LAB
0102 | mg/L | Alkalinity Total 4.5 (as CaCO3) | X004 | Colorimetric Titr'n | Manual Colorimetric Titration with standard solutions to specified endpoint as identified by colorimetric indicator. The sample must not be filtered, diluted, concentrated, or altered in any way. | 0.5 | Y | Y | Y | 9094 | 2/21/2003 | 7/9/2015 | 8/28/2020 12:46:44 PM | 1996-04-12 | 6/3/1998 6:51:45 PM | EMS |  | Concentration | mg/L | Y | (fl. conc.) |  | Alkalinity Total 4.5 (as CaCO3) (fl. conc.) |  | Major Ions; | Major Ions; EMS code: 0102 | Chemical | Total | LAB
0102 | mg/L | Alkalinity Total 4.5 (as CaCO3) | 1200 | Manual Potentiometric Titration | Manual Potentiometric Titration with standard solution to specified pH. The sample must not be filtered, diluted, concentrated, or altered in any way. | 0.5 | Y | Y | Y | 951 | 2/28/2005 | 2/2/2015 | 2/9/2015 9:32:54 AM | 1996-04-12 | 6/3/1998 6:51:33 PM | EMS |  | Concentration | mg/L | Y | (fl. conc.) |  | Alkalinity Total 4.5 (as CaCO3) (fl. conc.) |  | Major Ions; | Major Ions; EMS code: 0102 | Chemical | Total | LAB
0102 | mg/L | Alkalinity Total 4.5 (as CaCO3) | EQ01 | EQUIS (mg/L) | EQUIS (mg/L) | 0.5 | N | Y | Y | 11620 |  |  | 10/24/1996 3:16:55 PM | 1996-04-12 | 6/3/1998 6:51:33 PM | EMS |  | Concentration | mg/L | Y | (fl. conc.) |  | Alkalinity Total 4.5 (as CaCO3) (fl. conc.) |  | Major Ions; | Major Ions; EMS code: 0102 | Chemical | Total | LAB
0103 | ug/g | Carbon Total Organic | EQ02 | EQUIS (ug/g) | EQUIS (ug/g) | 300 | N | Y | Y | 98 |  |  | 10/22/1996 11:18:42 PM | 1996-04-12 | 6/3/1998 6:51:33 PM | EMS |  | MassPerMass | mg/kg | Y | (mass conc.) |  | Carbon Total Organic (mass conc.) |  | Carbon; | Carbon; EMS code: 0103 | Chemical | Total | LAB
0103 | mg/L | Carbon Total Organic | EQ01 | EQUIS (mg/L) | EQUIS (mg/L) | 0.5 | N | Y | Y | 10666 |  |  | 10/24/1996 3:16:55 PM | 1996-04-12 | 6/3/1998 6:51:45 PM | EMS |  | Concentration | mg/L | Y | (fl. conc.) |  | Carbon Total Organic (fl. conc.) |  | Carbon; | Carbon; EMS code: 0103 | Chemical | Total | LAB
0103 | mg/L | Carbon Total Organic | X257 | Colorimetric | Colorimetric. General use for several parameters. E.g.Total Sulfide is determined by spectrophotometer using the methlyene blue colourimetric method. The H2S calculation converts Total Sulphide as (S2-) and reports it as Total Sulphide as (H2S). | 1 | Y | Y | Y | 21928 | 1/9/2002 | 6/26/2024 1:13:00 AM | 7/2/2024 4:37:03 PM | 1996-04-12 | 6/3/1998 6:51:33 PM | EMS |  | Concentration | mg/L | Y | (fl. conc.) |  | Carbon Total Organic (fl. conc.) |  | Carbon; | Carbon; EMS code: 0103 | Chemical | Total | LAB
0103 | % | Carbon Total Organic | TOCC | TOC Calculated | Total Organic Carbon (TOC) is calculated by the difference between total carbon (TC) and total inorganic carbon. (TIC). | 0.05 | Y | Y | Y | 101 | 7/25/2016 | 2/21/2024 | 2/26/2024 6:09:04 PM | 2018-01-10 |  | JKROGH |  | DimensionlessRatio | % | N | (ratio) |  | Carbon Total Organic (ratio) |  | Carbon; | Carbon; EMS code: 0103 | Chemical | Total | LAB
0103 | mg/L | Carbon Total Organic | X067 | TOC Analyzer | Organic carbon is converted to carbon dioxide (CO2) by catalytic combustion or wet chemistry oxidation. The CO2 is then measured by infrared detector. | 1 | Y | Y | Y | 60097 | 12/12/1996 | 8/21/2024 | 8/26/2024 4:53:05 PM | 1996-04-12 | 6/3/1998 6:51:33 PM | EMS |  | Concentration | mg/L | Y | (fl. conc.) |  | Carbon Total Organic (fl. conc.) |  | Carbon; | Carbon; EMS code: 0103 | Chemical | Total | LAB
0103 | % | Carbon Total Organic | IR17 | ALS Minerals C-IR17 | ALS Minerals C-IR17. Sample is heated at 105ºC with 50% HCl added dropwise to evolve inorganic carbonates as CO2. The process repeated several times; after which the residue is dried and washed with deionized water to remove the remaining acid and stop the reaction. The remainder of the sample; taken to be non-carbonate carbon, is analysed by induction furnace and infrared spectroscopy. | 0.02 | Y | Y | Y | 0 |  |  |  | 2024-01-29 |  | RDGIBSON |  | DimensionlessRatio | % | N | (ratio) |  | Carbon Total Organic (ratio) |  | Carbon; | Carbon; EMS code: 0103 | Chemical | Total | LAB
0103 | % | Carbon Total Organic | VELC | CNS Analysis in Soil, Vario El Cube | Elemental Analysis of Soil Samples by Elementar Vario Cube EL. DOCUMENT # CAL SOP-00263. Reference Method: Application Note, Elementar Form No. AN-A-030609-E-01, CNS analysis in soil with the vario EL Cube SECTION 1: Introduction The Vario El Cube is an elemental analyzer that uses a thermal conductivity detector to determine Carbon, Hydrogen, Nitrogen, and Sulfur in a soil sample. The thermal conductivity detector contains two chambers, one measuring cell for the analysis gas and a reference cell for the carrier gas. Both gases will flow through the cell during analysis. The gases will mix and cause a difference in electrical charge which is measured by a microcontroller. The microcontroller will convert and integrate the electrical signal into a digitized one. The signal is displayed as a measuring peak dependent on the time. The elemental content of the sample will be calculated from the integrals of the individual measuring peaks and the calibration factors of the elements. SECTION 2: Scope This procedure is applicable to the determination of Carbon, Organic Carbon, Hydrogen, Nitrogen and Sulphur in solids. SECTION 3: Reporting Detection Limits Total Nitrogen: 0.20% Total Carbon: 0.02% Total Organic Carbon: 0.02% Total Sulfur: 0.03% SECTION 4: Interferences 4.1 All samples are to be dried and ground as per AB SOP-00035: Grinding of Soil Samples to ensure all moisture is removed. During analysis the sample gas passes through reagent tubes containing anhydrone (magnesium perchlorate) to remove any water vapor produced during combustion. 4.2 Some samples may react strongly when the HCL is added which could cause a loss of sample. If this occurs a dilution must be done and the acid must be added drop by drop. SECTION 5: Sample Hold Time and Preservation No specific holding time is stated in the reference method. It is Maxxam?s policy to discard samples 60 days after analysis. No preservation required, store dry samples at ambient Temperature. | 0.02 | Y | Y | Y | 10 | 11/7/2014 11:48:00 AM | 8/31/2021 11:18:00 AM | 9/12/2021 3:13:22 PM | 2014-11-13 |  | RDGIBSON |  | DimensionlessRatio | % | N | (ratio) |  | Carbon Total Organic (ratio) |  | Carbon; | Carbon; EMS code: 0103 | Chemical | Total | LAB
0103 | % | Carbon Total Organic | IR06 | ALS Minerals C-IR06 | ALS Minerals C-IR06. Sample is treated with 50% HCl and heated for 5 mins at 125ºC to expel carbonates as CO2. The remaining solution is vacuum filtered through a filtering crucible, and the residue is analysed via induction furnace and infrared spectroscopy. | 0.01 | Y | Y | Y | 0 |  |  |  | 2024-01-29 |  | RDGIBSON |  | DimensionlessRatio | % | N | (ratio) |  | Carbon Total Organic (ratio) |  | Carbon; | Carbon; EMS code: 0103 | Chemical | Total | LAB
0103 | ug/g | Carbon Total Organic | CAL2 | Calculation (ug/g) | Calculation (ug/g) | 300 | Y | Y | Y | 282 | 10/1/1996 11:13:00 AM | 7/5/2001 7:30:00 AM | 8/21/2002 2:11:01 PM | 1996-04-12 | 6/3/1998 6:51:41 PM | EMS |  | MassPerMass | mg/kg | Y | (mass conc.) |  | Carbon Total Organic (mass conc.) |  | Carbon; | Carbon; EMS code: 0103 | Chemical | Total | LAB
0103 | ug/g | Carbon Total Organic | X231 | HCl Leach,Leco Furnace | The procedure consists of first air drying a subsample of sediment at room temperature. The dried sample is then digested with concetrated hydrochloric acid to remove inorganic carbon. The acid treated sample is dried at 60C. Iron tin fines are added to the sample prior to combustion in a Leco Induction Furnace. Carbon is determined volumetrically. Results are reported on a dry-weight basis. | 0.1 | Y | Y | Y | 629 | 2/25/2002 4:15:00 PM | 8/22/2006 | 7/23/2010 12:16:50 PM | 1996-04-12 | 6/3/1998 6:51:33 PM | EMS |  | MassPerMass | mg/kg | Y | (mass conc.) |  | Carbon Total Organic (mass conc.) |  | Carbon; | Carbon; EMS code: 0103 | Chemical | Total | LAB
0103 | % | Carbon Total Organic | TOC63 | TOC Calculation for 63um fraction | Total Organic Carbon (TOC) is calculated by the difference between total carbon (TC) and total inorganic carbon. (TIC). The TOC is determined on the fraction of the sample that passes a 63um screen. | 0.05 | Y | Y | Y | 19 | 9/28/2017 | 10/6/2018 | 10/10/2018 2:46:10 PM | 2018-01-10 |  | JKROGH |  | DimensionlessRatio | % | N | (ratio) |  | Carbon Total Organic (ratio) |  | Carbon; | Carbon; EMS code: 0103 | Chemical | Total | LAB
0103 | % | Carbon Total Organic | FS93 | Oxid to CO2,NDIR,Calc'n | Total Carbon is determined by high temperature oxidation of carbon to carbon diooxide which is then measured by means of a nondispersive infrared analyzer. Inorganic Carbon is determined by reaction with phosphoric acid to convert all carbonates to carbon dioxide which is also measured by means of a nondispersive infrared analyzer. Total Organic Carbon is determined as the difference between Total and Inorganic Carbons. | 0.05 | Y | Y | Y | 232 | 4/11/2000 10:02:00 AM | 9/30/2021 3:31:00 PM | 10/18/2021 9:46:09 AM | 2001-05-22 |  | NPEPPIN |  | DimensionlessRatio | % | N | (ratio) |  | Carbon Total Organic (ratio) |  | Carbon; | Carbon; EMS code: 0103 | Chemical | Total | LAB
0103 | mg/L | Carbon Total Organic | XCAL | Calculation (mg/L) | Calculation (mg/L) | 0.5 | Y | Y | Y | 13628 | 7/31/1996 12:51:00 PM | 8/21/2007 1:54:00 PM | 12/19/2022 12:39:07 PM | 1996-04-12 | 6/3/1998 6:51:41 PM | EMS |  | Concentration | mg/L | Y | (fl. conc.) |  | Carbon Total Organic (fl. conc.) |  | Carbon; | Carbon; EMS code: 0103 | Chemical | Total | LAB
0103 | mg/L | Carbon Total Organic | 0940 | Vap.bypass/FID Dohrmann | Vap.bypass/FID Dohrmann | 1 | Y | Y | Y | 6 |  |  | 10/24/1996 2:06:03 PM | 1996-04-12 | 6/3/1998 6:51:33 PM | EMS |  | Concentration | mg/L | Y | (fl. conc.) |  | Carbon Total Organic (fl. conc.) |  | Carbon; | Carbon; EMS code: 0103 | Chemical | Total | LAB
0104 | mg/L | Chloride Total | X044 | Ion Chromatography | A clear (Total) or filtered (Dissolved) water sample is injected into a stream of carbonate-bicarbonate eluent and passed through a series of ion exchanger. The anions of interested are separated on the basis of their relative affinities for a low capacity , strongly basic anion exchanger (guard and separate columns). The separated anions are directed through the anion self-regenerating suppressor (ASRS). Hydronium ions in the suppressor will neutralize the eluent to its weakly ionized form and convert the separated anions in their conductive acid form. The separated anions are measured by conductivity. They are identified on the basis of retention time as compared to standards. Quantitation is by measurement of peak area of peak height against a calibration curve generated from known standards. | 0.5 | Y | Y | Y | 8608 | 12/12/1996 | 9/18/2015 | 8/26/2024 2:41:32 PM | 1996-04-12 | 6/3/1998 6:51:33 PM | EMS |  | Concentration | mg/L | Y | (fl. conc.) |  | Chloride Total (fl. conc.) | 16887-00-6 | Major Ions; | Major Ions; EMS code: 0104 | Chemical | Total | LAB
0104 | mg/L | Chloride Total | X004 | Colorimetric Titr'n | Manual Colorimetric Titration with standard solutions to specified endpoint as identified by colorimetric indicator. The sample must not be filtered, diluted, concentrated, or altered in any way. | 0.5 | Y | Y | Y | 548 | 6/25/2010 3:15:00 PM | 6/25/2010 3:15:00 PM | 6/7/2024 5:30:18 PM | 1996-04-12 | 6/3/1998 6:51:33 PM | EMS |  | Concentration | mg/L | Y | (fl. conc.) |  | Chloride Total (fl. conc.) | 16887-00-6 | Major Ions; | Major Ions; EMS code: 0104 | Chemical | Total | LAB
0104 | mg/L | Chloride Total | X276 | Ferric Thiocyan:Col'm | Ferric Thiocyan:Col'm | 0.5 | Y | Y | Y | 734 | 1/15/2007 | 1/2/2015 | 1/29/2015 2:12:56 PM | 1996-04-12 | 6/3/1998 6:51:45 PM | EMS |  | Concentration | mg/L | Y | (fl. conc.) |  | Chloride Total (fl. conc.) | 16887-00-6 | Major Ions; | Major Ions; EMS code: 0104 | Chemical | Total | LAB
0104 | mg/L | Chloride Total | X460 | Mercuric Nitrate Titr'n (Manual) | Chloride is titrated (manual method) with standard mercuric nitrate in the presence of diphenylcarbazone at a pH of 2.5 +/- 0.1. At the end point of the titration, the excess mercuric ions are complexed by the diphenylcarbazone to form a purple color. | 0.5 | Y | Y | Y | 210 |  |  | 10/29/2001 11:34:39 AM | 1996-04-12 | 6/3/1998 6:51:33 PM | EMS |  | Concentration | mg/L | Y | (fl. conc.) |  | Chloride Total (fl. conc.) | 16887-00-6 | Major Ions; | Major Ions; EMS code: 0104 | Chemical | Total | LAB
0104 | mg/L | Chloride Total | EQ01 | EQUIS (mg/L) | EQUIS (mg/L) | 0.5 | N | Y | Y | 889 |  |  | 10/24/1996 3:16:17 PM | 1996-04-12 | 6/3/1998 6:51:33 PM | EMS |  | Concentration | mg/L | Y | (fl. conc.) |  | Chloride Total (fl. conc.) | 16887-00-6 | Major Ions; | Major Ions; EMS code: 0104 | Chemical | Total | LAB
0105 | ug/g | Cyanide S.A.D. | EQ02 | EQUIS (ug/g) | EQUIS (ug/g) | 0.02 | N | Y | Y | 2 |  |  | 10/22/1996 5:37:18 PM | 1996-04-12 | 6/3/1998 6:51:33 PM | EMS |  | MassPerMass | mg/kg | Y | (mass conc.) |  | Cyanide S.A.D. (mass conc.) | 57-12-5 | Inorganics; | Inorganics; EMS code: 0105 | Chemical |  | LAB
0105 | mg/L | Cyanide S.A.D. | 0520 | H2SO4 Dig:Pyr-Pyr Vi | H2SO4 Dig:Pyr-Pyr Vi | 0.005 | Y | Y | Y | 16 |  |  | 10/23/1996 11:21:46 PM | 1996-04-12 | 6/3/1998 6:51:45 PM | EMS |  | Concentration | mg/L | Y | (fl. conc.) |  | Cyanide S.A.D. (fl. conc.) | 57-12-5 | Inorganics; | Inorganics; EMS code: 0105 | Chemical |  | LAB
0105 | mg/L | Cyanide S.A.D. | CS01 | System Calculation | System Calculation (mg/L) |  | N | Y | N | 24 |  |  | 5/26/2003 2:05:58 PM | 1996-04-12 | 6/3/1998 6:51:44 PM | EMS |  | Concentration | mg/L | Y | (fl. conc.) |  | Cyanide S.A.D. (fl. conc.) | 57-12-5 | Inorganics; | Inorganics; EMS code: 0105 | Chemical |  | LAB
0105 | mg/L | Cyanide S.A.D. | X122 | NaOH,Pyr-Barb,Col'metric | NaOH;Pyr-Barb;Col'metric | 0.01 | Y | Y | Y | 742 | 7/27/2005 7:11:00 AM | 6/13/2016 4:43:00 PM | 12/15/2017 5:01:15 PM | 1996-04-12 | 6/3/1998 6:51:45 PM | EMS |  | Concentration | mg/L | Y | (fl. conc.) |  | Cyanide S.A.D. (fl. conc.) | 57-12-5 | Inorganics; | Inorganics; EMS code: 0105 | Chemical |  | LAB
0105 | mg/L | Cyanide S.A.D. | 0510 | HCl-NH2OH Dig:Isonic/Bar | HCl-NH2OH Dig:Isonic/Bar | 0.005 | Y | Y | Y | 30021 |  |  | 5/19/2015 3:10:36 PM | 1996-04-12 | 6/3/1998 6:51:33 PM | EMS |  | Concentration | mg/L | Y | (fl. conc.) |  | Cyanide S.A.D. (fl. conc.) | 57-12-5 | Inorganics; | Inorganics; EMS code: 0105 | Chemical |  | LAB
0105 | mg/L | Cyanide S.A.D. | X086 | NaOH,Bis-Pyr,Col'metric | NaOH;Bis-Pyr;Col'metric | 0.01 | Y | Y | Y | 98 |  |  | 10/24/1996 9:13:46 AM | 1996-04-12 | 6/3/1998 6:51:33 PM | EMS |  | Concentration | mg/L | Y | (fl. conc.) |  | Cyanide S.A.D. (fl. conc.) | 57-12-5 | Inorganics; | Inorganics; EMS code: 0105 | Chemical |  | LAB
0105 | mg/L | Cyanide S.A.D. | X092 | NaOH,Pyr-Pyr,Col'metric | NaOH;Pyr-Pyr;Col'metric | 0.01 | Y | Y | Y | 36 | 1/26/2001 | 4/16/2001 | 7/27/2001 11:21:20 AM | 1996-04-12 | 6/3/1998 6:51:33 PM | EMS |  | Concentration | mg/L | Y | (fl. conc.) |  | Cyanide S.A.D. (fl. conc.) | 57-12-5 | Inorganics; | Inorganics; EMS code: 0105 | Chemical |  | LAB
0105 | mg/L | Cyanide S.A.D. | X066 | CN Selective Electrode | CN Selective Electrode | 0.01 | Y | Y | Y | 2080 | 4/1/1997 9:50:00 AM | 6/22/2024 2:18:00 PM | 6/28/2024 5:01:19 PM | 1996-04-12 | 6/3/1998 6:51:45 PM | EMS |  | Concentration | mg/L | Y | (fl. conc.) |  | Cyanide S.A.D. (fl. conc.) | 57-12-5 | Inorganics; | Inorganics; EMS code: 0105 | Chemical |  | LAB
0105 | mg/L | Cyanide S.A.D. | F004 | HCL-HH Dig,Ni Complex,Col'm | The sample is refluxed for 1 hour in the presence of hydroxylamine hydrochloride/hydrochloric acid. This digestion medium minimizes the Fe-catalyzed decomposition of CNS. Simple cyanides and some complex metal cyanides eg. Cd, Cu, Ni, Zn react to produce HCN gas. The HCN generated reacts instantaneously with nickel ions in an ammoniacal solution to form the ionic tetracyanonickelate(II) anion complex, Ni(CN)4-2, which has a characteristic ultra-violet spectrum, and is measured at 267 and 284 nm. Note that thiocyanates do not react under these test conditions, but can cause positive bias if metal complexed cyanide is present. | 0.03 | Y | Y | Y | 83 | 12/6/1996 2:24:00 PM | 12/3/2001 12:54:00 PM | 5/26/2003 2:36:54 PM | 1996-06-19 | 6/3/1998 6:51:44 PM | MPAK |  | Concentration | mg/L | Y | (fl. conc.) |  | Cyanide S.A.D. (fl. conc.) | 57-12-5 | Inorganics; | Inorganics; EMS code: 0105 | Chemical |  | LAB
0105 | mg/L | Cyanide S.A.D. | X153 | UV:Auto Pyr-Barb Acid | UV:Auto Pyr-Barb Acid | 0.005 | Y | Y | Y | 2662 | 2/6/2002 | 7/25/2013 4:45:00 PM | 8/16/2013 10:46:01 AM | 1996-04-12 | 6/3/1998 6:51:33 PM | EMS |  | Concentration | mg/L | Y | (fl. conc.) |  | Cyanide S.A.D. (fl. conc.) | 57-12-5 | Inorganics; | Inorganics; EMS code: 0105 | Chemical |  | LAB
0105 | mg/L | Cyanide S.A.D. | X324 | HCL-HH dig, isonicotinic | The sample is subjected to a strong acid [hydrochloric acid - hydroxylamine hydrochloride] reflux digestion/distillation. HCN is liberated and trapped in a weak NaOH solution. An aliquot of this solution is then analyzed by a colorimetric technique, ref. Isonicotinic - Barbituric Acid method. | 0.05 | Y | Y | Y | 992 |  |  | 10/1/2003 9:40:21 AM | 1996-04-12 | 6/3/1998 6:51:45 PM | EMS |  | Concentration | mg/L | Y | (fl. conc.) |  | Cyanide S.A.D. (fl. conc.) | 57-12-5 | Inorganics; | Inorganics; EMS code: 0105 | Chemical |  | LAB
0105 | mg/L | Cyanide S.A.D. | AA09 | Man.Dist.,Auto-CN,Pyr-Barb-Low Level | A sample is placed in a sample tube then acidified and sealed to a collection tube containing a sodium hydroxide trapping solution. The tube is placed in a block heater and distilled through a hydrophobic membrane. The distillate is then analyzed for strong acid dissociable (SAD) cyanide by acid distillation / UV irradiation followed by an automated colourimetric determination | 0.0005 | Y | Y | Y | 951 | 5/15/2009 | 7/11/2024 10:47:00 AM | 8/21/2024 11:42:55 AM | 2009-07-20 | 7/22/2009 2:44:58 PM | RDGIBSON | RDGIBSON | Concentration | mg/L | Y | (fl. conc.) |  | Cyanide S.A.D. (fl. conc.) | 57-12-5 | Inorganics; | Inorganics; EMS code: 0105 | Chemical |  | LAB
0105 | mg/L | Cyanide S.A.D. | EQ01 | EQUIS (mg/L) | EQUIS (mg/L) | 0.01 | N | Y | Y | 2227 |  |  | 10/24/1996 3:14:41 PM | 1996-04-12 | 6/3/1998 6:51:33 PM | EMS |  | Concentration | mg/L | Y | (fl. conc.) |  | Cyanide S.A.D. (fl. conc.) | 57-12-5 | Inorganics; | Inorganics; EMS code: 0105 | Chemical |  | LAB
0105 | mg/L | Cyanide S.A.D. | 0511 | HCl/NH2OH Dig: NiCl2 | HCl/NH2OH Dig: NiCl2 | 0.005 | Y | Y | Y | 1532 |  |  | 10/24/1996 1:47:32 PM | 1996-04-12 | 6/3/1998 6:51:33 PM | EMS |  | Concentration | mg/L | Y | (fl. conc.) |  | Cyanide S.A.D. (fl. conc.) | 57-12-5 | Inorganics; | Inorganics; EMS code: 0105 | Chemical |  | LAB
0105 | mg/L | Cyanide S.A.D. | 0530 | H2SO4 Dig:UV | H2SO4 Dig:UV | 0.005 | Y | Y | Y | 499 |  |  | 12/14/2015 10:27:40 AM | 1996-04-12 | 6/3/1998 6:51:33 PM | EMS |  | Concentration | mg/L | Y | (fl. conc.) |  | Cyanide S.A.D. (fl. conc.) | 57-12-5 | Inorganics; | Inorganics; EMS code: 0105 | Chemical |  | LAB
0105 | ug/g wet | Cyanide S.A.D. | FS14 | Tetracyanonickelatate,UV Spec. | For weak acid dissociable cyanide, CN(WAD), the solid sample is added to type one water and is refluxed for 1 hour in the presence of Zinc acetate and sodium acetate. The HCN is liberated from this slightly acidified (pH 4.5 - 6.0) sample. The acetate buffer used contains Zinc salts to precipitate iron cyanide as a further assurance of the selectivity of the method. For strong acid dissociable cyanide, CN(SAD), the solid sample is added to Type 1 water and is refluxed for 1 hour in the presence of sulphuric acid and magnesium chloride. For both CN(WAD) and CN(SAD) the HCN generated reacts instantaneously with nickel ions in an ammoniacal solution to form the ionic tetracyanonickelate (II) anion complex, Ni (CN)4-2, which has a characteristic ultra-violet spectrum, and is measured at 267 and 284 nm. | 0.5 | Y | Y | Y | 8 | 8/24/2000 7:15:00 AM | 12/5/2001 3:11:00 PM | 5/26/2003 2:26:59 PM | 1997-06-23 | 6/3/1998 6:51:47 PM | NPEPPIN |  | MassPerMass | µg/g | Y | (mass conc.) | (wet) | Cyanide S.A.D. (wet) (mass conc.) | 57-12-5 | Inorganics; | Inorganics; EMS code: 0105 | Chemical |  | LAB
0106 | mg/L | Fluoride Total | EQ01 | EQUIS (mg/L) | EQUIS (mg/L) | 0.1 | N | Y | Y | 759 |  |  | 10/24/1996 3:16:17 PM | 1996-04-12 | 6/3/1998 6:51:45 PM | EMS |  | Concentration | mg/L | Y | (fl. conc.) |  | Fluoride Total (fl. conc.) | 16984-48-8 |  | EMS code: 0106 | Chemical | Total | LAB
0106 | ug/g | Fluoride Total | EQ02 | EQUIS (ug/g) | EQUIS (ug/g) | 20 | N | Y | Y | 23 |  |  | 10/22/1996 10:25:20 PM | 1996-04-12 | 6/3/1998 6:51:33 PM | EMS |  | MassPerMass | mg/kg | Y | (mass conc.) |  | Fluoride Total (mass conc.) | 16984-48-8 |  | EMS code: 0106 | Chemical | Total | LAB
0106 | ug/g | Fluoride Total | 0640 | Na2CO3/K2CO3 Fus. SIE SP | Na2CO3/K2CO3 Fus. SIE SP | 2 | Y | Y | Y | 2 |  |  | 10/23/1996 4:55:40 PM | 1996-04-12 | 6/3/1998 6:51:33 PM | EMS |  | MassPerMass | mg/kg | Y | (mass conc.) |  | Fluoride Total (mass conc.) | 16984-48-8 |  | EMS code: 0106 | Chemical | Total | LAB
0106 | ug/m3 | Fluoride Total | X188 | Impinger,Sel Ion Elect | Impinger;Sel Ion Elect | 0 | Y | Y | Y | 3239 |  |  | 7/3/2024 12:59:26 PM | 1996-04-12 | 6/3/1998 6:51:33 PM | EMS |  | AirConcentration | µg/m³ | Y | (air conc.) |  | Fluoride Total (air conc.) | 16984-48-8 |  | EMS code: 0106 | Chemical | Total | LAB
0106 | mg/L | Fluoride Total | X143 | Selective Ion Electrode | The analyte is determined potentiometrically by direct measurement using a selective ion electrode with standard reference electrode and either a meter with a direct concentration scale or a pH meter with an expanded mV scale. | 0.1 | Y | Y | Y | 569 | 1/24/2001 | 2/3/2021 5:13:00 PM | 2/10/2021 4:41:01 PM | 1996-04-12 | 6/3/1998 6:51:33 PM | EMS |  | Concentration | mg/L | Y | (fl. conc.) |  | Fluoride Total (fl. conc.) | 16984-48-8 |  | EMS code: 0106 | Chemical | Total | LAB
0106 | mg/L | Fluoride Total | 1120 | Manual S I E | Manual S I E | 0.1 | Y | Y | Y | 7 |  |  | 10/22/1997 11:51:12 AM | 1996-04-12 | 6/3/1998 6:51:33 PM | EMS |  | Concentration | mg/L | Y | (fl. conc.) |  | Fluoride Total (fl. conc.) | 16984-48-8 |  | EMS code: 0106 | Chemical | Total | LAB
0106 | mg/L | Fluoride Total | X044 | Ion Chromatography | A clear (Total) or filtered (Dissolved) water sample is injected into a stream of carbonate-bicarbonate eluent and passed through a series of ion exchanger. The anions of interested are separated on the basis of their relative affinities for a low capacity , strongly basic anion exchanger (guard and separate columns). The separated anions are directed through the anion self-regenerating suppressor (ASRS). Hydronium ions in the suppressor will neutralize the eluent to its weakly ionized form and convert the separated anions in their conductive acid form. The separated anions are measured by conductivity. They are identified on the basis of retention time as compared to standards. Quantitation is by measurement of peak area of peak height against a calibration curve generated from known standards. | 0.1 | Y | Y | Y | 14748 | 4/10/1996 6:56:00 AM | 8/17/2024 3:01:00 PM | 8/23/2024 5:01:25 PM | 1996-04-12 | 6/3/1998 6:51:33 PM | EMS |  | Concentration | mg/L | Y | (fl. conc.) |  | Fluoride Total (fl. conc.) | 16984-48-8 |  | EMS code: 0106 | Chemical | Total | LAB
0106 | mg/L | Fluoride Total | X110 | SPADNS | SPADNS | 0.1 | Y | Y | Y | 145 | 1/15/2007 3:15:00 PM | 1/15/2007 3:15:00 PM | 4/11/2007 4:21:18 PM | 1996-04-12 | 6/3/1998 6:51:33 PM | EMS |  | Concentration | mg/L | Y | (fl. conc.) |  | Fluoride Total (fl. conc.) | 16984-48-8 |  | EMS code: 0106 | Chemical | Total | LAB
0107 | mg/L | Hardness Total (Total) | MCAL | Calculation (Ca,Mg,Sr,Fe,Al,Zn,Mn,Cu,Ba,Pb) | Dissolved and extractable water samples are ready for analysis once collected. Total samples are autoclaved with HNO3 and hydrochloric acid (HCl) in capped Teflon vials. Analysis is performed via Inductively Coupled Argon Plasma Atomic Emission Spectrometry (ICP-AES) (simultaneous multielement analysis), Graphite Furnace Atomic Absorption Spectrometry (GF-AAS) (low level analysis) and/or Flame Atomic Absorption Spectrometry (F-AAS). Total Hardness is calculated from the dissolved (or extractable as a substitute) fractions of Ca, Mg, Sr, Fe, Al, Zn, Mn, Cu, Ba, Pb, according to the following formula: Hardness Total = 2.497248*Ca + 4.116885*Mg + 1.142316*Sr + 1.792213*Fe + 3.709568*Al + 1.531126*Zn + 1.821866*Mn + 1.575224*Cu + 0.728773*Ba + 0.483082*Pb (where Total Hardness is expressed as mg/L of CaCO3, and each element is expressed in mg/L. Check Media/Pres. codes for consistency with regard to acid preservation. | 1 | Y | Y | Y | 105 |  |  | 8/13/2024 12:09:44 PM | 1998-04-21 | 6/3/1998 6:51:48 PM | NPEPPIN |  | Concentration | mg/L | Y | (fl. conc.) |  | Hardness Total (Total) (fl. conc.) |  | Major Ions; | Major Ions; EMS code: 0107 | Chemical | Total | LAB
0107 | mg/L | Hardness Total (Total) | X040 | EDTA Titration | a) For Ca, the sample pH is adjusted to pH 12-13 so that the magnesium is precipitated as the hydroxide. Next, the sample is titrated with EDTA in the presence of an indicator. When sufficient EDTA has been added so as to complex all the calcium ions the indicator changes color. Results reported as mg/L as Ca. b) For hardness, a sample buffered to pH 10.0+/-0.1 is titrated with EDTA in the presence of an indicator. When sufficient EDTA has been added in order to complex all the calcium and magnesium ions the indicator changes color. Results reported as mg/L as CaCO3. | 0.1 | Y | Y | Y | 5787 | 12/1/1998 8:45:00 AM | 11/30/2012 12:01:00 AM | 12/20/2012 11:18:42 AM | 1996-04-12 | 6/3/1998 6:51:33 PM | EMS |  | Concentration | mg/L | Y | (fl. conc.) |  | Hardness Total (Total) (fl. conc.) |  | Major Ions; | Major Ions; EMS code: 0107 | Chemical | Total | LAB
0107 | ug/g wet | Hardness Total (Total) | CS44 | System Calculation | System Calculation (ug/g wet) Unit Code 44 |  | Y | Y | Y | 249 |  |  | 11/28/2023 2:21:33 PM | 2015-12-14 |  | RDGIBSON |  | MassPerMass | µg/g | Y | (mass conc.) | (wet) | Hardness Total (Total) (wet) (mass conc.) |  | Major Ions; | Major Ions; EMS code: 0107 | Chemical | Total | LAB
0107 | mg/L | Hardness Total (Total) | CCAL | Calculation (Ca, Mg) | Dissolved and extractable water samples are ready for analysis once collected. Total samples are autoclaved with HNO3 and hydrochloric acid (HCl) in capped Teflon vials. Analysis is performed via Inductively Coupled Argon Plasma Atomic Emission Spectrometry (ICP-AES) (simultaneous multielement analysis), Graphite Furnace Atomic Absorption Spectrometry (GF-AAS) (low level analysis) and/or Flame Atomic Absorption Spectrometry (F-AAS). Total Hardness is calculated from dissolved Calcium and dissolved Magnesium according to the following formula: Total Hardness=2.497248*Ca + 4.116885*Mg (where Total Hardness is expressed as mg/L of CaCO3, and Calcium and Magnesium are expressed as mg/L Ca and mg/L Mg respectively. Check Media/Pres. codes for consistency with regard to acid preservation. | 1 | Y | Y | Y | 32185 | 8/23/1996 6:04:00 PM | 8/16/2024 | 8/26/2024 8:06:03 AM | 1997-07-31 | 6/3/1998 6:51:47 PM | NPEPPIN |  | Concentration | mg/L | Y | (fl. conc.) |  | Hardness Total (Total) (fl. conc.) |  | Major Ions; | Major Ions; EMS code: 0107 | Chemical | Total | LAB
0107 | % | Hardness Total (Total) | CS34 | System Calculation | System Calculation (%) |  | N | Y | N | 92 |  |  | 7/29/2024 9:13:09 AM | 1996-04-12 | 6/3/1998 6:51:42 PM | EMS |  | DimensionlessRatio | % | N | (ratio) |  | Hardness Total (Total) (ratio) |  | Major Ions; | Major Ions; EMS code: 0107 | Chemical | Total | LAB
0107 | mg/L | Hardness Total (Total) | XCAL | Calculation (mg/L) | Calculation (mg/L) | 0.1 | Y | Y | Y | 34687 | 7/31/1996 6:03:00 PM | 8/23/2024 | 8/26/2024 4:53:05 PM | 1996-04-12 | 6/3/1998 6:51:42 PM | EMS |  | Concentration | mg/L | Y | (fl. conc.) |  | Hardness Total (Total) (fl. conc.) |  | Major Ions; | Major Ions; EMS code: 0107 | Chemical | Total | LAB
0107 | ug/g | Hardness Total (Total) | CS02 | System Calculation | System Calculation (ug/g) |  | N | Y | N | 6944 |  |  | 7/29/2024 5:17:09 PM | 1996-04-12 | 6/3/1998 6:51:42 PM | EMS |  | MassPerMass | mg/kg | Y | (mass conc.) |  | Hardness Total (Total) (mass conc.) |  | Major Ions; | Major Ions; EMS code: 0107 | Chemical | Total | LAB
0107 | mg/L | Hardness Total (Total) | EQ01 | EQUIS (mg/L) | EQUIS (mg/L) | 0.1 | N | Y | Y | 1875 |  |  | 10/24/1996 3:16:55 PM | 1996-04-12 | 6/3/1998 6:51:45 PM | EMS |  | Concentration | mg/L | Y | (fl. conc.) |  | Hardness Total (Total) (fl. conc.) |  | Major Ions; | Major Ions; EMS code: 0107 | Chemical | Total | LAB
0107 | mg/L | Hardness Total (Total) | CS01 | System Calculation | System Calculation (mg/L) |  | N | Y | N | 89246 |  |  | 8/26/2024 2:41:32 PM | 1996-04-12 | 6/3/1998 6:51:42 PM | EMS |  | Concentration | mg/L | Y | (fl. conc.) |  | Hardness Total (Total) (fl. conc.) |  | Major Ions; | Major Ions; EMS code: 0107 | Chemical | Total | LAB
0107 | mg/L | Hardness Total (Total) | X004 | Colorimetric Titr'n | Manual Colorimetric Titration with standard solutions to specified endpoint as identified by colorimetric indicator. The sample must not be filtered, diluted, concentrated, or altered in any way. | 0.1 | Y | Y | Y | 56 | 8/5/1997 6:06:00 PM | 12/1/2010 | 12/9/2010 3:55:25 PM | 1996-04-12 | 6/3/1998 6:51:33 PM | EMS |  | Concentration | mg/L | Y | (fl. conc.) |  | Hardness Total (Total) (fl. conc.) |  | Major Ions; | Major Ions; EMS code: 0107 | Chemical | Total | LAB
0107 | mg/L | Hardness Total (Total) | X102 | Col'metric Titr'n,Ca | Col'metric Titr'n;Ca | 0.1 | Y | Y | Y | 4 | 10/6/1998 10:30:00 AM | 10/27/1998 8:45:00 AM | 11/27/1998 7:55:35 AM | 1996-04-12 | 6/3/1998 6:51:45 PM | EMS |  | Concentration | mg/L | Y | (fl. conc.) |  | Hardness Total (Total) (fl. conc.) |  | Major Ions; | Major Ions; EMS code: 0107 | Chemical | Total | LAB
0107 | mg/dm2/d | Hardness Total (Total) | CS15 | System Calculation | System Calculation (mg/dm2/d) |  | N | Y | N | 269 |  |  | 7/31/2024 11:23:13 AM | 1996-04-12 | 6/3/1998 6:51:42 PM | EMS |  | MassFlux | mg/dm²/d | Y | (mass flux) |  | Hardness Total (Total) (mass flux) |  | Major Ions; | Major Ions; EMS code: 0107 | Chemical | Total | LAB
0108 | mg/L | Nitrogen Ammonia Total | CS01 | System Calculation | System Calculation (mg/L) |  | N | Y | N | 21142 |  |  | 8/23/2024 5:01:31 PM | 1996-04-12 | 6/3/1998 6:51:43 PM | EMS |  | Concentration | mg/L | Y | (fl. conc.) |  | Nitrogen Ammonia Total (fl. conc.) | 7664-41-7 | Nutrients; | Nutrients; EMS code: 0108 | Chemical | Total | LAB
0108 | mg/L | Nitrogen Ammonia Total | X054 | Phenate-Colorimetric | Phenate-Colorimetric | 0.01 | Y | Y | Y | 15588 | 5/17/1999 | 7/27/2024 8:00:00 AM | 8/13/2024 12:09:44 PM | 1996-04-12 | 6/3/1998 6:51:33 PM | EMS |  | Concentration | mg/L | Y | (fl. conc.) |  | Nitrogen Ammonia Total (fl. conc.) | 7664-41-7 | Nutrients; | Nutrients; EMS code: 0108 | Chemical | Total | LAB
0108 | mg/L | Nitrogen Ammonia Total | EQ01 | EQUIS (mg/L) | EQUIS (mg/L) | 0.01 | N | Y | Y | 4510 |  |  | 10/24/1996 3:17:07 PM | 1996-04-12 | 6/3/1998 6:51:33 PM | EMS |  | Concentration | mg/L | Y | (fl. conc.) |  | Nitrogen Ammonia Total (fl. conc.) | 7664-41-7 | Nutrients; | Nutrients; EMS code: 0108 | Chemical | Total | LAB
0108 | mg/L | Nitrogen Ammonia Total | OPAF | Ortho-phthaldialdehyde (OPA) reaction,flourometric detection | A fluorescent derivative is formed by reacting ortho-phthaldialdehyde (OPA) and sulphite with ammonium. A sample is injected in a matrix matched carrier which is then mixed with a sodium hydroxide solution to convert the free ammonium into ammonia gas. After mixing, this basic solution is pushed through a dialysis PTFE membrane to the receiver OPA solution which then reacts to form the fluorescent derivative. This reaction is further increased (Intensified) by passing through a 62 degree inline reaction loop which then is detected and quantified using an inline fluorescent detector. This method is applicable for the measurement of ammonium in drinking water, ground water, and can be acceptable for brackish waters. | 0.005 | Y | Y | Y | 37949 |  |  | 8/20/2024 7:35:41 AM | 2010-08-04 |  | RDGIBSON |  | Concentration | mg/L | Y | (fl. conc.) |  | Nitrogen Ammonia Total (fl. conc.) | 7664-41-7 | Nutrients; | Nutrients; EMS code: 0108 | Chemical | Total | LAB
0108 | mg/L | Nitrogen Ammonia Total | X488 | Salicylate Colorimetric | Salicylate Colorimetric | 0 | Y | Y | Y | 3750 |  |  | 8/23/2024 8:10:36 AM | 1998-07-10 |  | NPEPPIN |  | Concentration | mg/L | Y | (fl. conc.) |  | Nitrogen Ammonia Total (fl. conc.) | 7664-41-7 | Nutrients; | Nutrients; EMS code: 0108 | Chemical | Total | LAB
0108 | t/d | Nitrogen Ammonia Total | CALC | Calculation (t/d) | Calculation (t/d) | 0 | Y | Y | Y | 1124 |  |  | 4/11/2024 12:22:09 PM | 1996-04-12 | 2/26/2009 11:51:52 AM | EMS | RDGIBSON | MassFlow | t/d | Y | (mass flow) |  | Nitrogen Ammonia Total (mass flow) | 7664-41-7 | Nutrients; | Nutrients; EMS code: 0108 | Chemical | Total | LAB
0108 | mg/L | Nitrogen Ammonia Total | X143 | Selective Ion Electrode | The analyte is determined potentiometrically by direct measurement using a selective ion electrode with standard reference electrode and either a meter with a direct concentration scale or a pH meter with an expanded mV scale. | 0.01 | Y | Y | Y | 4003 | 1/8/1997 | 6/26/2024 | 7/30/2024 8:37:34 PM | 1996-04-12 | 6/3/1998 6:51:33 PM | EMS |  | Concentration | mg/L | Y | (fl. conc.) |  | Nitrogen Ammonia Total (fl. conc.) | 7664-41-7 | Nutrients; | Nutrients; EMS code: 0108 | Chemical | Total | LAB
0108 | mg/L | Nitrogen Ammonia Total | NH3F | NH3-F-VA | Ammonia in Water by Fluorescence | 0.005 | Y | Y | Y | 38555 | 1/23/2012 | 8/22/2024 | 8/26/2024 3:51:07 PM | 2018-04-03 |  | RDGIBSON |  | Concentration | mg/L | Y | (fl. conc.) |  | Nitrogen Ammonia Total (fl. conc.) | 7664-41-7 | Nutrients; | Nutrients; EMS code: 0108 | Chemical | Total | LAB
0108 | ug/m3 | Nitrogen Ammonia Total | X195 | Dist,@pH9.5,Nessler.-Air | Dist;@pH9.5;Nessler.-Air | 0 | Y | Y | Y | 3 |  |  | 10/24/1996 11:56:43 AM | 1996-04-12 | 6/3/1998 6:51:33 PM | EMS |  | AirConcentration | µg/m³ | Y | (air conc.) |  | Nitrogen Ammonia Total (air conc.) | 7664-41-7 | Nutrients; | Nutrients; EMS code: 0108 | Chemical | Total | LAB
0108 | mg/L | Nitrogen Ammonia Total | X478 | pH7.4 dist,Nessler or Titration | The unfiltered sample is buffered at pH 7.4 so as to inhibit conversion of organic nitrogen compounds to ammonia and yet provide an alkaline medium which allows for the recovery of free ammonia by distillation into boric acid solution. The ammonia in the distillate is then determined either by Nesslerization (low conc. 0.01 to 1.2 mg/L as N) or by titration (conc. 1. to 30. mg/L as N). Preferably fresh samples are analyzed without preservation. Where samples cannot be analyzed immediately they are acidified to pH 1.5 to 2 with H2SO4 and are analyzed within 36 hours of sample collection. | 0.01 | Y | Y | Y | 132 | 9/29/2000 | 12/1/2010 | 10/24/2022 10:58:17 AM | 1997-10-22 | 6/3/1998 6:51:47 PM | NPEPPIN |  | Concentration | mg/L | Y | (fl. conc.) |  | Nitrogen Ammonia Total (fl. conc.) | 7664-41-7 | Nutrients; | Nutrients; EMS code: 0108 | Chemical | Total | LAB
0108 | ug/m3 | Nitrogen Ammonia Total | X196 | Dist.,Acid Titration-Air | Dist.;Acid Titration-Air | 0 | Y | Y | Y | 34 |  |  | 10/24/1996 1:08:30 AM | 1996-04-12 | 6/3/1998 6:51:33 PM | EMS |  | AirConcentration | µg/m³ | Y | (air conc.) |  | Nitrogen Ammonia Total (air conc.) | 7664-41-7 | Nutrients; | Nutrients; EMS code: 0108 | Chemical | Total | LAB
0108 | mg/L | Nitrogen Ammonia Total | X004 | Colorimetric Titr'n | Manual Colorimetric Titration with standard solutions to specified endpoint as identified by colorimetric indicator. The sample must not be filtered, diluted, concentrated, or altered in any way. | 0.01 | Y | Y | Y | 7976 |  |  | 6/7/2024 5:30:18 PM | 1996-04-12 | 6/3/1998 6:51:33 PM | EMS |  | Concentration | mg/L | Y | (fl. conc.) |  | Nitrogen Ammonia Total (fl. conc.) | 7664-41-7 | Nutrients; | Nutrients; EMS code: 0108 | Chemical | Total | LAB
0108 | mg/L | Nitrogen Ammonia Total | X126 | Dist,Acid Titration | Dist;Acid Titration | 0.01 | Y | Y | Y | 2525 | 9/18/1996 | 9/18/2015 | 10/18/2023 8:55:31 AM | 1996-04-12 | 6/3/1998 6:51:33 PM | EMS |  | Concentration | mg/L | Y | (fl. conc.) |  | Nitrogen Ammonia Total (fl. conc.) | 7664-41-7 | Nutrients; | Nutrients; EMS code: 0108 | Chemical | Total | LAB
0108 | ug/m3 | Nitrogen Ammonia Total | EQ40 | EQUIS (mg/m3) | EQUIS (mg/m3) | 0 | N | Y | Y | 83 |  |  | 10/22/1996 11:53:53 PM | 1996-04-12 | 6/3/1998 6:51:33 PM | EMS |  | AirConcentration | µg/m³ | Y | (air conc.) |  | Nitrogen Ammonia Total (air conc.) | 7664-41-7 | Nutrients; | Nutrients; EMS code: 0108 | Chemical | Total | LAB
0108 | mg/L | Nitrogen Ammonia Total | X033 | FA,Dist,Titration | FA;Dist;Titration | 0.01 | Y | Y | Y | 534 |  |  | 3/8/2005 11:34:47 AM | 1996-04-12 | 6/3/1998 6:51:33 PM | EMS |  | Concentration | mg/L | Y | (fl. conc.) |  | Nitrogen Ammonia Total (fl. conc.) | 7664-41-7 | Nutrients; | Nutrients; EMS code: 0108 | Chemical | Total | LAB
0109 | mg/L | Total Nitrogen NO2 + NO3 | X044 | Ion Chromatography | A clear (Total) or filtered (Dissolved) water sample is injected into a stream of carbonate-bicarbonate eluent and passed through a series of ion exchanger. The anions of interested are separated on the basis of their relative affinities for a low capacity , strongly basic anion exchanger (guard and separate columns). The separated anions are directed through the anion self-regenerating suppressor (ASRS). Hydronium ions in the suppressor will neutralize the eluent to its weakly ionized form and convert the separated anions in their conductive acid form. The separated anions are measured by conductivity. They are identified on the basis of retention time as compared to standards. Quantitation is by measurement of peak area of peak height against a calibration curve generated from known standards. | 0.005 | Y | Y | Y | 5513 | 5/6/1981 12:00:00 PM | 3/24/2023 | 5/3/2023 11:44:08 AM | 1996-04-12 | 6/3/1998 6:51:45 PM | EMS |  | Concentration | mg/L | Y | (fl. conc.) |  | Total Nitrogen NO2 + NO3 (fl. conc.) | 14797-55-8, 14797-65-0 | Nutrients; | Nutrients; EMS code: 0109 | Chemical | Total | LAB
0109 | mg/L | Total Nitrogen NO2 + NO3 | X019 | Manual Diazotization | Manual Diazotization | 0.02 | Y | Y | Y | 316 | 9/29/2000 | 12/1/2010 | 12/9/2010 3:55:25 PM | 1996-04-12 | 6/3/1998 6:51:33 PM | EMS |  | Concentration | mg/L | Y | (fl. conc.) |  | Total Nitrogen NO2 + NO3 (fl. conc.) | 14797-55-8, 14797-65-0 | Nutrients; | Nutrients; EMS code: 0109 | Chemical | Total | LAB
0109 | mg/L | Total Nitrogen NO2 + NO3 | X519 | F,EDTA,Cd red'n,diaz,col'm | Filtered;EDTA etc.;Cd reduction;diazotization;colorimetric | 0.001 | Y | Y | Y | 54 | 1/6/2024 8:00:00 AM | 7/27/2024 8:00:00 AM | 8/8/2024 11:45:18 AM | 1999-04-22 |  | NPEPPIN |  | Concentration | mg/L | Y | (fl. conc.) |  | Total Nitrogen NO2 + NO3 (fl. conc.) | 14797-55-8, 14797-65-0 | Nutrients; | Nutrients; EMS code: 0109 | Chemical | Total | LAB
0109 | mg/L | Total Nitrogen NO2 + NO3 | CS01 | System Calculation | System Calculation (mg/L) |  | N | Y | N | 72488 |  |  | 8/26/2024 1:03:06 PM | 1996-04-12 | 6/3/1998 6:51:43 PM | EMS |  | Concentration | mg/L | Y | (fl. conc.) |  | Total Nitrogen NO2 + NO3 (fl. conc.) | 14797-55-8, 14797-65-0 | Nutrients; | Nutrients; EMS code: 0109 | Chemical | Total | LAB
0109 | ug/g | Total Nitrogen NO2 + NO3 | CS02 | System Calculation | System Calculation (ug/g) |  | N | Y | N | 14 |  |  | 11/23/2023 7:26:23 AM | 1996-04-12 | 6/3/1998 6:51:43 PM | EMS |  | MassPerMass | mg/kg | Y | (mass conc.) |  | Total Nitrogen NO2 + NO3 (mass conc.) | 14797-55-8, 14797-65-0 | Nutrients; | Nutrients; EMS code: 0109 | Chemical | Total | LAB
0109 | mg/L | Total Nitrogen NO2 + NO3 | EQ01 | EQUIS (mg/L) | EQUIS (mg/L) | 0.02 | N | Y | Y | 198 |  |  | 10/24/1996 3:17:05 PM | 1996-04-12 | 6/3/1998 6:51:33 PM | EMS |  | Concentration | mg/L | Y | (fl. conc.) |  | Total Nitrogen NO2 + NO3 (fl. conc.) | 14797-55-8, 14797-65-0 | Nutrients; | Nutrients; EMS code: 0109 | Chemical | Total | LAB
0110 | mg/L | Nitrogen NO3 Total | X143 | Selective Ion Electrode | The analyte is determined potentiometrically by direct measurement using a selective ion electrode with standard reference electrode and either a meter with a direct concentration scale or a pH meter with an expanded mV scale. | 0.02 | Y | Y | Y | 359 |  |  | 1/30/2002 4:36:40 PM | 1996-04-12 | 6/3/1998 6:51:45 PM | EMS |  | Concentration | mg/L | Y | (fl. conc.) |  | Nitrogen NO3 Total (fl. conc.) | 14797-55-8 | Nutrients; | Nutrients; EMS code: 0110 | Chemical | Total | LAB
0110 | mg/L | Nitrogen NO3 Total | CS01 | System Calculation | System Calculation (mg/L) |  | N | Y | N | 43 |  |  | 8/13/2016 5:14:18 PM | 1996-04-12 | 6/3/1998 6:51:43 PM | EMS |  | Concentration | mg/L | Y | (fl. conc.) |  | Nitrogen NO3 Total (fl. conc.) | 14797-55-8 | Nutrients; | Nutrients; EMS code: 0110 | Chemical | Total | LAB
0110 | mg/L | Nitrogen NO3 Total | XCAL | Calculation (mg/L) | Calculation (mg/L) | 0.002 | Y | Y | Y | 12023 | 4/23/1997 | 9/18/1997 | 6/7/2024 5:30:18 PM | 1996-04-12 | 6/3/1998 6:51:41 PM | EMS |  | Concentration | mg/L | Y | (fl. conc.) |  | Nitrogen NO3 Total (fl. conc.) | 14797-55-8 | Nutrients; | Nutrients; EMS code: 0110 | Chemical | Total | LAB
0110 | mg/L | Nitrogen NO3 Total | X044 | Ion Chromatography | A clear (Total) or filtered (Dissolved) water sample is injected into a stream of carbonate-bicarbonate eluent and passed through a series of ion exchanger. The anions of interested are separated on the basis of their relative affinities for a low capacity , strongly basic anion exchanger (guard and separate columns). The separated anions are directed through the anion self-regenerating suppressor (ASRS). Hydronium ions in the suppressor will neutralize the eluent to its weakly ionized form and convert the separated anions in their conductive acid form. The separated anions are measured by conductivity. They are identified on the basis of retention time as compared to standards. Quantitation is by measurement of peak area of peak height against a calibration curve generated from known standards. | 0.02 | Y | Y | Y | 5462 | 9/14/1994 | 4/4/2023 | 8/13/2024 12:09:44 PM | 1996-04-12 | 6/3/1998 6:51:33 PM | EMS |  | Concentration | mg/L | Y | (fl. conc.) |  | Nitrogen NO3 Total (fl. conc.) | 14797-55-8 | Nutrients; | Nutrients; EMS code: 0110 | Chemical | Total | LAB
0110 | mg/L | Nitrogen NO3 Total | EQ01 | EQUIS (mg/L) | EQUIS (mg/L) | 0.02 | N | Y | Y | 9886 |  |  | 10/24/1996 3:16:20 PM | 1996-04-12 | 6/3/1998 6:51:33 PM | EMS |  | Concentration | mg/L | Y | (fl. conc.) |  | Nitrogen NO3 Total (fl. conc.) | 14797-55-8 | Nutrients; | Nutrients; EMS code: 0110 | Chemical | Total | LAB
0110 | mg/L | Nitrogen NO3 Total | X056 | Chromotropic Acid Meth. | Chromotropic Acid Meth. | 0.02 | Y | Y | Y | 2876 |  |  | 10/24/1996 11:06:33 AM | 1996-04-12 | 6/3/1998 6:51:33 PM | EMS |  | Concentration | mg/L | Y | (fl. conc.) |  | Nitrogen NO3 Total (fl. conc.) | 14797-55-8 | Nutrients; | Nutrients; EMS code: 0110 | Chemical | Total | LAB
0111 | mg/L | Nitrogen NO2 Total | X020 | Manual Cadmium Reduction | The filtered sample is passed through a column containing granulated copper-cadmium which reduces Nitrate to Nitrite. Quantitative reduction is ensured by maintaining a stable H-ion concentration through addition of ammonium chloride (NH4Cl). The Nitrite that is originally present in the sample plus the reduced Nitrate then react with sulfanilamide under acidic conditions to form a diazo compound. This compound then complexes with N-1-naphthylethylenediamine dihydrochloride to form a reddish purple azo dye, the absorbance of which is measured at 550 nm. Nitrate (NO3) concentration can be calculated using the Nitrate & Nitrite result and subtracting the Nitrite result (NO23 - NO2 = NO3). Manual analysis. | 0.005 | Y | Y | Y | 5827 |  |  | 5/19/2015 3:10:10 PM | 1996-04-12 | 6/3/1998 6:51:45 PM | EMS |  | Concentration | mg/L | Y | (fl. conc.) |  | Nitrogen NO2 Total (fl. conc.) | 14797-65-0 | Nutrients; | Nutrients; EMS code: 0111 | Chemical | Total | LAB
0111 | mg/L | Nitrogen NO2 Total | X019 | Manual Diazotization | Manual Diazotization | 0.005 | Y | Y | Y | 11758 | 4/23/1997 | 11/25/2005 11:10:00 AM | 5/23/2013 1:08:49 PM | 1996-04-12 | 6/3/1998 6:51:33 PM | EMS |  | Concentration | mg/L | Y | (fl. conc.) |  | Nitrogen NO2 Total (fl. conc.) | 14797-65-0 | Nutrients; | Nutrients; EMS code: 0111 | Chemical | Total | LAB
0111 | mg/L | Nitrogen NO2 Total | CS01 | System Calculation | System Calculation (mg/L) |  | N | Y | N | 16 |  |  | 8/13/2016 5:09:01 PM | 1996-04-12 | 6/3/1998 6:51:43 PM | EMS |  | Concentration | mg/L | Y | (fl. conc.) |  | Nitrogen NO2 Total (fl. conc.) | 14797-65-0 | Nutrients; | Nutrients; EMS code: 0111 | Chemical | Total | LAB
0111 | mg/L | Nitrogen NO2 Total | EQ01 | EQUIS (mg/L) | EQUIS (mg/L) | 0.005 | N | Y | Y | 2909 |  |  | 10/24/1996 3:16:17 PM | 1996-04-12 | 6/3/1998 6:51:33 PM | EMS |  | Concentration | mg/L | Y | (fl. conc.) |  | Nitrogen NO2 Total (fl. conc.) | 14797-65-0 | Nutrients; | Nutrients; EMS code: 0111 | Chemical | Total | LAB
0111 | mg/L | Nitrogen NO2 Total | X143 | Selective Ion Electrode | The analyte is determined potentiometrically by direct measurement using a selective ion electrode with standard reference electrode and either a meter with a direct concentration scale or a pH meter with an expanded mV scale. | 0.02 | Y | Y | Y | 349 |  |  | 8/28/2020 12:46:44 PM | 1996-04-12 | 6/3/1998 6:51:33 PM | EMS |  | Concentration | mg/L | Y | (fl. conc.) |  | Nitrogen NO2 Total (fl. conc.) | 14797-65-0 | Nutrients; | Nutrients; EMS code: 0111 | Chemical | Total | LAB
0111 | mg/L | Nitrogen NO2 Total | X044 | Ion Chromatography | A clear (Total) or filtered (Dissolved) water sample is injected into a stream of carbonate-bicarbonate eluent and passed through a series of ion exchanger. The anions of interested are separated on the basis of their relative affinities for a low capacity , strongly basic anion exchanger (guard and separate columns). The separated anions are directed through the anion self-regenerating suppressor (ASRS). Hydronium ions in the suppressor will neutralize the eluent to its weakly ionized form and convert the separated anions in their conductive acid form. The separated anions are measured by conductivity. They are identified on the basis of retention time as compared to standards. Quantitation is by measurement of peak area of peak height against a calibration curve generated from known standards. | 0.005 | Y | Y | Y | 62680 | 9/14/1994 | 11/11/2023 | 8/26/2024 2:41:32 PM | 1996-04-12 | 6/3/1998 6:51:33 PM | EMS |  | Concentration | mg/L | Y | (fl. conc.) |  | Nitrogen NO2 Total (fl. conc.) | 14797-65-0 | Nutrients; | Nutrients; EMS code: 0111 | Chemical | Total | LAB
0112 | mg/L | Nitrogen Organic-Total | EQ01 | EQUIS (mg/L) | EQUIS (mg/L) | 0.02 | N | Y | Y | 13150 |  |  | 10/24/1996 3:17:07 PM | 1996-04-12 | 6/3/1998 6:51:33 PM | EMS |  | Concentration | mg/L | Y | (fl. conc.) |  | Nitrogen Organic-Total (fl. conc.) |  | Nutrients; | Nutrients; EMS code: 0112 | Chemical | Total | LAB
0112 | mg/L | Nitrogen Organic-Total | XCAL | Calculation (mg/L) | Calculation (mg/L) | 0.02 | Y | Y | Y | 36861 | 5/4/2002 | 8/23/2024 7:24:00 AM | 8/23/2024 5:01:31 PM | 1996-04-12 | 6/3/1998 6:51:41 PM | EMS |  | Concentration | mg/L | Y | (fl. conc.) |  | Nitrogen Organic-Total (fl. conc.) |  | Nutrients; | Nutrients; EMS code: 0112 | Chemical | Total | LAB
0112 | mg/L | Nitrogen Organic-Total | CS01 | System Calculation | System Calculation (mg/L) |  | N | Y | N | 45576 |  |  | 8/26/2024 3:51:07 PM | 1996-04-12 | 6/3/1998 6:51:43 PM | EMS |  | Concentration | mg/L | Y | (fl. conc.) |  | Nitrogen Organic-Total (fl. conc.) |  | Nutrients; | Nutrients; EMS code: 0112 | Chemical | Total | LAB
0113 | mg/L | Nitrogen Kjel.Tot(N) | X150 | H2O2/H2SO4Dig,Indophenol | H2O2/H2SO4Dig;Indophenol | 0.01 | Y | Y | Y | 459 |  |  | 7/25/2024 3:53:30 PM | 1996-04-12 | 8/25/2020 6:05:46 PM | EMS | RDGIBSON | Concentration | mg/L | Y | (fl. conc.) |  | Nitrogen Kjel.Tot(N) (fl. conc.) |  | Nutrients; | Nutrients; EMS code: 0113 | Chemical |  | LAB
0113 | mg/L | Nitrogen Kjel.Tot(N) | X306 | K2SO4/H2SO4 Dig:Auto Colo | K2SO4/H2SO4 Dig:Auto Colorimet | 0.01 | Y | Y | Y | 531 | 9/3/1996 | 4/4/2023 | 5/29/2023 7:32:19 AM | 1996-04-12 | 6/3/1998 6:51:33 PM | EMS |  | Concentration | mg/L | Y | (fl. conc.) |  | Nitrogen Kjel.Tot(N) (fl. conc.) |  | Nutrients; | Nutrients; EMS code: 0113 | Chemical |  | LAB
0113 | mg/L | Nitrogen Kjel.Tot(N) | X108 | LA,HNO3,Titration | LA;HNO3;Titration | 0.01 | Y | Y | Y | 23 |  |  | 10/24/1996 1:48:23 AM | 1996-04-12 | 6/3/1998 6:51:33 PM | EMS |  | Concentration | mg/L | Y | (fl. conc.) |  | Nitrogen Kjel.Tot(N) (fl. conc.) |  | Nutrients; | Nutrients; EMS code: 0113 | Chemical |  | LAB
0113 | mg/L | Nitrogen Kjel.Tot(N) | X098 | Dig,Titration | Dig;Titration | 0.01 | Y | Y | Y | 513 | 9/18/1996 | 9/18/2015 | 10/18/2023 8:55:31 AM | 1996-04-12 | 6/3/1998 6:51:45 PM | EMS |  | Concentration | mg/L | Y | (fl. conc.) |  | Nitrogen Kjel.Tot(N) (fl. conc.) |  | Nutrients; | Nutrients; EMS code: 0113 | Chemical |  | LAB
0113 | ug/g | Nitrogen Kjel.Tot(N) | EQ02 | EQUIS (ug/g) | EQUIS (ug/g) | 30 | N | Y | Y | 208 |  |  | 10/22/1996 11:18:42 PM | 1996-04-12 | 6/3/1998 6:51:33 PM | EMS |  | MassPerMass | mg/kg | Y | (mass conc.) |  | Nitrogen Kjel.Tot(N) (mass conc.) |  | Nutrients; | Nutrients; EMS code: 0113 | Chemical |  | LAB
0113 | mg/L | Nitrogen Kjel.Tot(N) | TKNF | TKN-F-VA | TKN in Water by Fluorescence | 0.05 | Y | Y | Y | 23094 | 10/28/2011 | 8/22/2024 | 8/26/2024 3:51:07 PM | 2015-03-09 |  | RDGIBSON |  | Concentration | mg/L | Y | (fl. conc.) |  | Nitrogen Kjel.Tot(N) (fl. conc.) |  | Nutrients; | Nutrients; EMS code: 0113 | Chemical |  | LAB
0113 | mg/L | Nitrogen Kjel.Tot(N) | TKN-F | TKN in Seawater by Fluorescence | TKN in Seawater by Fluorescence | 0.05 | Y | Y | Y | 102 | 8/16/2019 | 5/2/2024 | 6/7/2024 10:03:16 AM | 2019-09-30 |  | KATFRASE |  | Concentration | mg/L | Y | (fl. conc.) |  | Nitrogen Kjel.Tot(N) (fl. conc.) |  | Nutrients; | Nutrients; EMS code: 0113 | Chemical |  | LAB
0113 | ug/g | Nitrogen Kjel.Tot(N) | X245 | Block Digest/Auto.Color | Sample is ground and dried then subjected to hydrogen peroxide or sulphuric acid digestion in a block digestor which converts the Kjeldahl nitrogen containing compounds to ammonium sulphate and carbon dioxide. The digest is then taken up by an auto analyzer, neutralized and reacted with sodium salicylate and a chlorine source. Using sodium nitroprusside catalyst, a coloured indophenols complex is formed and absorbance at 660nm is measured. | 30 | Y | Y | Y | 652 | 5/22/2003 | 10/31/2023 | 11/23/2023 7:26:23 AM | 1996-04-12 | 6/3/1998 6:51:33 PM | EMS |  | MassPerMass | mg/kg | Y | (mass conc.) |  | Nitrogen Kjel.Tot(N) (mass conc.) |  | Nutrients; | Nutrients; EMS code: 0113 | Chemical |  | LAB
0113 | mg/L | Nitrogen Kjel.Tot(N) | 2660 | Froz.Dig.Aut.Color | Froz.Dig.Aut.Color | 0.01 | Y | Y | Y | 1489 |  |  | 8/14/1997 9:59:25 AM | 1996-04-12 | 6/3/1998 6:51:33 PM | EMS |  | Concentration | mg/L | Y | (fl. conc.) |  | Nitrogen Kjel.Tot(N) (fl. conc.) |  | Nutrients; | Nutrients; EMS code: 0113 | Chemical |  | LAB
0113 | mg/L | Nitrogen Kjel.Tot(N) | X126 | Dist,Acid Titration | Dist;Acid Titration | 0.01 | Y | Y | Y | 353 |  |  | 11/29/2021 1:31:17 PM | 1996-04-12 | 6/3/1998 6:51:45 PM | EMS |  | Concentration | mg/L | Y | (fl. conc.) |  | Nitrogen Kjel.Tot(N) (fl. conc.) |  | Nutrients; | Nutrients; EMS code: 0113 | Chemical |  | LAB
0113 | mg/L | Nitrogen Kjel.Tot(N) | X329 | Auto Digestion&Colorimet | The sample is automatically digested with a sulfuric acid solution containing a metal catalyst. Organic nitrogen is converted to ammonium sulfate. | 0.05 | Y | Y | Y | 120 |  |  | 8/13/2024 12:09:44 PM | 1996-04-12 | 6/3/1998 6:51:33 PM | EMS |  | Concentration | mg/L | Y | (fl. conc.) |  | Nitrogen Kjel.Tot(N) (fl. conc.) |  | Nutrients; | Nutrients; EMS code: 0113 | Chemical |  | LAB
0113 | % | Nitrogen Kjel.Tot(N) | STKN | CSSS (2008) 22.2.3 | See CSSS (2008) 22.2.3 for full method | 1.31 | Y | Y | Y | 20 | 5/3/2022 | 6/11/2024 | 6/18/2024 11:02:03 PM | 2022-07-05 |  | EMS |  | DimensionlessRatio | % | N | (ratio) |  | Nitrogen Kjel.Tot(N) (ratio) |  | Nutrients; | Nutrients; EMS code: 0113 | Chemical |  | LAB
0113 | mg/L | Nitrogen Kjel.Tot(N) | X107 | FA,HNO3,Titration | FA;HNO3;Titration | 0.01 | Y | Y | Y | 52 |  |  | 10/3/2005 9:51:09 AM | 1996-04-12 | 6/3/1998 6:51:33 PM | EMS |  | Concentration | mg/L | Y | (fl. conc.) |  | Nitrogen Kjel.Tot(N) (fl. conc.) |  | Nutrients; | Nutrients; EMS code: 0113 | Chemical |  | LAB
0113 | mg/L | Nitrogen Kjel.Tot(N) | X143 | Selective Ion Electrode | The analyte is determined potentiometrically by direct measurement using a selective ion electrode with standard reference electrode and either a meter with a direct concentration scale or a pH meter with an expanded mV scale. | 0.01 | Y | Y | Y | 25594 | 6/12/2001 | 10/3/2014 | 12/18/2023 2:41:39 PM | 1996-04-12 | 6/3/1998 6:51:33 PM | EMS |  | Concentration | mg/L | Y | (fl. conc.) |  | Nitrogen Kjel.Tot(N) (fl. conc.) |  | Nutrients; | Nutrients; EMS code: 0113 | Chemical |  | LAB
0113 | mg/L | Nitrogen Kjel.Tot(N) | X325 | HG SO4 Dig.AutoCol-Berth | The sample is heated in the presence of sulfuric acid, potassium sulfate and mercuric sulfate for 2.5 hours. The digest is cooled, diluted to 25 mL and placed on the autoanalyzer for NH3 determination. | 0.04 | Y | Y | Y | 9 |  |  | 3/19/2006 4:43:30 PM | 1996-04-12 | 6/3/1998 6:51:45 PM | EMS |  | Concentration | mg/L | Y | (fl. conc.) |  | Nitrogen Kjel.Tot(N) (fl. conc.) |  | Nutrients; | Nutrients; EMS code: 0113 | Chemical |  | LAB
0113 | mg/L | Nitrogen Kjel.Tot(N) | 136A | HgSO4 Dig.Auto Color | HgSO4 Dig.Auto Color | 0.04 | Y | Y | Y | 5515 | 1/9/2002 | 2/14/2002 11:45:00 AM | 11/16/2004 4:00:37 PM | 1996-04-12 | 6/3/1998 6:51:45 PM | EMS |  | Concentration | mg/L | Y | (fl. conc.) |  | Nitrogen Kjel.Tot(N) (fl. conc.) |  | Nutrients; | Nutrients; EMS code: 0113 | Chemical |  | LAB
0113 | mg/L | Nitrogen Kjel.Tot(N) | CS01 | System Calculation | System Calculation (mg/L) |  | N | Y | N | 2713 |  |  | 8/21/2024 8:46:13 PM | 1996-04-12 | 6/3/1998 6:51:43 PM | EMS |  | Concentration | mg/L | Y | (fl. conc.) |  | Nitrogen Kjel.Tot(N) (fl. conc.) |  | Nutrients; | Nutrients; EMS code: 0113 | Chemical |  | LAB
0113 | mg/L | Nitrogen Kjel.Tot(N) | XCAL | Calculation (mg/L) | Calculation (mg/L) | 0.02 | Y | Y | Y | 32763 | 6/15/2007 | 8/21/2024 | 8/26/2024 1:03:06 PM | 1997-04-24 | 6/3/1998 6:51:46 PM | MPAK |  | Concentration | mg/L | Y | (fl. conc.) |  | Nitrogen Kjel.Tot(N) (fl. conc.) |  | Nutrients; | Nutrients; EMS code: 0113 | Chemical |  | LAB
0113 | %(W/W) | Nitrogen Kjel.Tot(N) | BDAC | Block Dig.Auto.Color % | Block Dig.Auto.Color % | 0.01 | Y | Y | Y | 183 | 3/5/2010 10:25:00 AM | 3/13/2024 9:33:00 AM | 3/26/2024 8:53:57 AM | 2010-04-20 |  | RDGIBSON |  | DimensionlessRatio | % | N | (ratio) | (wet) | Nitrogen Kjel.Tot(N) (wet) (ratio) |  | Nutrients; | Nutrients; EMS code: 0113 | Chemical |  | LAB
0113 | mg/L | Nitrogen Kjel.Tot(N) | X075 | Nesslerization | Nesslerization | 0.01 | Y | Y | Y | 471 |  |  | 8/28/2020 12:46:44 PM | 1996-04-12 | 6/3/1998 6:51:33 PM | EMS |  | Concentration | mg/L | Y | (fl. conc.) |  | Nitrogen Kjel.Tot(N) (fl. conc.) |  | Nutrients; | Nutrients; EMS code: 0113 | Chemical |  | LAB
0113 | mg/L | Nitrogen Kjel.Tot(N) | 1360 | Block Dig.Auto.Color | Block Dig.Auto.Color | 0.01 | Y | Y | Y | 32834 | 1/26/2001 | 8/23/2024 7:24:00 AM | 8/25/2024 9:47:36 PM | 1996-04-12 | 6/3/1998 6:51:42 PM | EMS |  | Concentration | mg/L | Y | (fl. conc.) |  | Nitrogen Kjel.Tot(N) (fl. conc.) |  | Nutrients; | Nutrients; EMS code: 0113 | Chemical |  | LAB
0113 | mg/L | Nitrogen Kjel.Tot(N) | EQ01 | EQUIS (mg/L) | EQUIS (mg/L) | 0.01 | N | Y | Y | 15131 |  |  | 10/24/1996 3:16:58 PM | 1996-04-12 | 6/3/1998 6:51:45 PM | EMS |  | Concentration | mg/L | Y | (fl. conc.) |  | Nitrogen Kjel.Tot(N) (fl. conc.) |  | Nutrients; | Nutrients; EMS code: 0113 | Chemical |  | LAB
0114 | mg/L | Nitrogen Total | NCEM | Total Nitrogen Chemiluminescence | To determine the quantity of combustible Total Nitrogen in water, the aqueous sample is injected into a combustion furnace with a temperature of greater than 700oC (with a catalyst). The nitrogen is converted to NO. The carrier gas sweeps the NO to the nitrogen module where NO is mixed with ozone and then forms NO2 in its excited state. When NO2 in its excited state returns to ground state, extra energy is given off as light. This phenomenon is called chemiluminescence. This light is then quantified using a chemiluminescence detector that is attached to the Apollo 9000 HS instrument. | 0.05 | Y | Y | Y | 251 | 11/17/2010 1:15:00 PM | 5/30/2014 2:36:00 PM | 6/9/2014 3:03:49 PM | 2010-12-23 |  | RDGIBSON |  | Concentration | mg/L | Y | (fl. conc.) |  | Nitrogen Total (fl. conc.) |  | Nutrients; | Nutrients; EMS code: 0114 | Chemical | Total | LAB
0114 | mg/L | Nitrogen Total | F006 | K/Persulphate-UV Dig,Red'n,Col'm | Samples containing nitrogen compounds are digested and oxidized in-line to nitrate at 95 degrees Celsius by using an alkaline persulfate digestion solution and ultraviolet (UV) radiation. | 0.02 | Y | Y | Y | 16983 | 11/7/2007 | 6/28/2013 11:00:00 PM | 1/19/2023 12:38:36 PM | 2007-10-29 |  | RDGIBSON |  | Concentration | mg/L | Y | (fl. conc.) |  | Nitrogen Total (fl. conc.) |  | Nutrients; | Nutrients; EMS code: 0114 | Chemical | Total | LAB
0114 | % | Nitrogen Total | CAL3 | Calculation (%) | Calculation (%) | 0.01 | Y | Y | Y | 7 | 9/22/2022 11:33:00 AM | 9/22/2022 11:33:00 AM | 3/9/2023 4:55:51 PM | 2009-11-24 |  | RDGIBSON |  | DimensionlessRatio | % | N | (ratio) |  | Nitrogen Total (ratio) |  | Nutrients; | Nutrients; EMS code: 0114 | Chemical | Total | LAB
0114 | mg/L | Nitrogen Total | F005 | K/Persulphate Dig,Red'n,Col'm | The Total Nitrogen present in the sample is oxidized to nitrate by digesting the sample with potassium persulphate in an alkaline medium. The resulting Nitrate is reduced to Nitrite by passing it through a column containing granulated copper-cadmium. The nitrite then reacts with sulfanilamide under acidic conditions to form a diazo compound. This compound then complexes with N-1-naphthylethylenediamine dihydrochloride to form a reddish-purple azo dye, the absorbance of which is measured at 550 nm. | 0.02 | Y | Y | Y | 70351 | 8/12/1996 6:16:00 PM | 8/22/2024 | 8/26/2024 1:05:19 PM | 1996-06-20 | 6/3/1998 6:51:44 PM | MPAK |  | Concentration | mg/L | Y | (fl. conc.) |  | Nitrogen Total (fl. conc.) |  | Nutrients; | Nutrients; EMS code: 0114 | Chemical | Total | LAB
0114 | ug/g | Nitrogen Total | FS09 | K/Persulphate Dig,Red'n,Col'm | The Total Nitrogen present in the sample is oxidized to nitrate by digesting the sample with potassium persulphate in an alkaline medium. The resulting Nitrate is reduced to Nitrite by passing it through a column containing granulated copper-cadmium. The nitrite then reacts with sulfanilamide under acidic conditions to form a diazo compound. This compound then complexes with N-1-naphthylethylenediamine dihydrochloride to form a reddish-purple azo dye, the absorbance of which is measured at 550 nm. | 50 | Y | Y | Y | 523 | 8/7/1997 11:00:00 AM | 1/17/2002 2:54:00 PM | 5/28/2003 2:34:18 PM | 1997-01-13 | 6/3/1998 6:51:46 PM | MPAK |  | MassPerMass | mg/kg | Y | (mass conc.) |  | Nitrogen Total (mass conc.) |  | Nutrients; | Nutrients; EMS code: 0114 | Chemical | Total | LAB
0114 | mg/L | Nitrogen Total | CS01 | System Calculation | System Calculation (mg/L) |  | N | Y | N | 826 |  |  | 10/18/2023 8:55:31 AM | 1996-04-12 | 6/3/1998 6:51:43 PM | EMS |  | Concentration | mg/L | Y | (fl. conc.) |  | Nitrogen Total (fl. conc.) |  | Nutrients; | Nutrients; EMS code: 0114 | Chemical | Total | LAB
0114 | mg/L | Nitrogen Total | EQ01 | EQUIS (mg/L) | EQUIS (mg/L) | 0.02 | N | Y | Y | 11388 |  |  | 10/24/1996 3:16:58 PM | 1996-04-12 | 6/3/1998 6:51:33 PM | EMS |  | Concentration | mg/L | Y | (fl. conc.) |  | Nitrogen Total (fl. conc.) |  | Nutrients; | Nutrients; EMS code: 0114 | Chemical | Total | LAB
0114 | mg/L | Nitrogen Total | X156 | Auto UV Dig:Col'metric | Auto UV Dig:Col'metric | 0.02 | Y | Y | Y | 273 |  |  | 1/19/2023 12:38:40 PM | 1996-04-12 | 6/3/1998 6:51:33 PM | EMS |  | Concentration | mg/L | Y | (fl. conc.) |  | Nitrogen Total (fl. conc.) |  | Nutrients; | Nutrients; EMS code: 0114 | Chemical | Total | LAB
0114 | mg/L | Nitrogen Total | XCAL | Calculation (mg/L) | Calculation (mg/L) | 0.04 | Y | Y | Y | 30920 | 4/12/2002 | 8/23/2024 7:24:00 AM | 8/26/2024 2:41:32 PM | 1996-04-12 | 6/3/1998 6:51:41 PM | EMS |  | Concentration | mg/L | Y | (fl. conc.) |  | Nitrogen Total (fl. conc.) |  | Nutrients; | Nutrients; EMS code: 0114 | Chemical | Total | LAB
0114 | ug/g | Nitrogen Total | FB05 | K/Persulphate Dig,Red'n,Col'm | The Total Nitrogen present in the sample is oxidized to nitrate by digesting the sample with potassium persulphate in an alkaline medium. The resulting Nitrate is reduced to Nitrite by passing it through a column containing granulated copper-cadmium. The nitrite then reacts with sulfanilamide under acidic conditions to form a diazo compound. This compound then complexes with N-1-naphthylethylenediamine dihydrochloride to form a reddish-purple azo dye, the absorbance of which is measured at 550 nm. | 50 | Y | Y | Y | 287 | 2/14/2002 | 10/31/2023 | 11/23/2023 7:26:23 AM | 1997-01-13 | 6/3/1998 6:51:46 PM | MPAK |  | MassPerMass | mg/kg | Y | (mass conc.) |  | Nitrogen Total (mass conc.) |  | Nutrients; | Nutrients; EMS code: 0114 | Chemical | Total | LAB
0115 | mg/L | Biochemical Oxygen Demand | X140 | Unseeded Dil Nit Inh 5 Days | The biochemical oxygen demand is determined by measuring the amount of dissolved oxygen (DO) that is consumed in the sample under aerobic conditions during the incubation period (five days at 20 C). A polarographic membrane electrode that has been calibrated in saturated air is used to acquire all dissolved oxygen measurements. Nitrification inhibitor added. No seed is added. A modified Winkler procedure may be substituted for the DO probe procedure. Performance Base Method published circa Dec. 2000. | 5 | Y | Y | Y | 29106 | 9/2/1996 | 8/12/2024 9:53:00 AM | 8/20/2024 8:39:43 AM | 1996-04-12 | 6/3/1998 6:51:41 PM | EMS |  | Concentration | mg/L | Y | (fl. conc.) |  | Biochemical Oxygen Demand (fl. conc.) |  | Aggregate Organics; | Aggregate Organics; EMS code: 0115 | Chemical |  | LAB
0115 | mg/L | Biochemical Oxygen Demand | 1100 | Uns. Dil. Nit. Inh.5 Day | Uns. Dil. Nit. Inh.5 Day | 10 | N | Y | Y | 90 | 11/18/1996 9:30:00 AM | 11/18/1996 9:30:00 AM | 4/10/1997 4:09:38 PM | 1996-04-12 | 6/3/1998 6:51:33 PM | EMS |  | Concentration | mg/L | Y | (fl. conc.) |  | Biochemical Oxygen Demand (fl. conc.) |  | Aggregate Organics; | Aggregate Organics; EMS code: 0115 | Chemical |  | LAB
0115 | t/d | Biochemical Oxygen Demand | EQ30 | EQUIS (T/d) | EQUIS (T/d) | 0 | N | Y | Y | 379 |  |  | 10/23/1996 12:03:16 AM | 1996-04-12 | 6/3/1998 6:51:33 PM | EMS |  | MassFlow | t/d | Y | (mass flow) |  | Biochemical Oxygen Demand (mass flow) |  | Aggregate Organics; | Aggregate Organics; EMS code: 0115 | Chemical |  | LAB
0115 | mg/L | Biochemical Oxygen Demand | EQ01 | EQUIS (mg/L) | EQUIS (mg/L) | 10 | N | Y | Y | 3305 |  |  | 10/24/1996 3:16:44 PM | 1996-04-12 | 6/3/1998 6:51:33 PM | EMS |  | Concentration | mg/L | Y | (fl. conc.) |  | Biochemical Oxygen Demand (fl. conc.) |  | Aggregate Organics; | Aggregate Organics; EMS code: 0115 | Chemical |  | LAB
0115 | mg/L | Biochemical Oxygen Demand | X018 | Dil 5 Day | Dil 5 Day. Not known whether BOD or CBOD. | 10 | Y | Y | Y | 4278 | 5/12/1999 | 6/5/2000 | 8/13/2024 12:09:44 PM | 1996-04-12 | 6/3/1998 6:51:33 PM | EMS |  | Concentration | mg/L | Y | (fl. conc.) |  | Biochemical Oxygen Demand (fl. conc.) |  | Aggregate Organics; | Aggregate Organics; EMS code: 0115 | Chemical |  | LAB
0115 | kg/adt | Biochemical Oxygen Demand | CAL5 | Calculation (kg/adt) | Calculation (kg/adt) | 0 | Y | Y | Y | 418 |  |  | 8/22/2001 9:59:26 AM | 1996-10-13 | 6/3/1998 6:51:46 PM | EMS |  | PerAirDriedTonne | kg/adt | Y | (/adt) |  | Biochemical Oxygen Demand (/adt) |  | Aggregate Organics; | Aggregate Organics; EMS code: 0115 | Chemical |  | LAB
0115 | mg/L | Biochemical Oxygen Demand | XCAL | Calculation (mg/L) | Calculation (mg/L) | 10 | Y | Y | Y | 8875 | 9/9/1996 6:59:00 AM | 6/30/2024 7:00:00 AM | 8/15/2024 9:21:45 AM | 1996-04-12 | 8/25/2020 6:11:52 PM | EMS | RDGIBSON | Concentration | mg/L | Y | (fl. conc.) |  | Biochemical Oxygen Demand (fl. conc.) |  | Aggregate Organics; | Aggregate Organics; EMS code: 0115 | Chemical |  | LAB
0115 | mg/L | Biochemical Oxygen Demand | 1110 | Seeded Dil.Nit.Inh.5 Day | Seeded Dil.Nit.Inh.5 Day | 10 | Y | Y | Y | 4701 | 5/3/1995 | 5/17/1996 | 9/30/2009 2:50:30 PM | 1996-04-12 | 8/28/2008 4:12:18 PM | EMS | RDGIBSON | Concentration | mg/L | Y | (fl. conc.) |  | Biochemical Oxygen Demand (fl. conc.) |  | Aggregate Organics; | Aggregate Organics; EMS code: 0115 | Chemical |  | LAB
0115 | t/d | Biochemical Oxygen Demand | CALC | Calculation (t/d) | Calculation (t/d) | 0 | Y | Y | Y | 11619 | 10/7/1996 | 8/4/2024 | 8/22/2024 1:49:14 PM | 1996-04-12 | 6/3/1998 6:51:33 PM | EMS |  | MassFlow | t/d | Y | (mass flow) |  | Biochemical Oxygen Demand (mass flow) |  | Aggregate Organics; | Aggregate Organics; EMS code: 0115 | Chemical |  | LAB
0115 | mg/L | Biochemical Oxygen Demand | X482 | Unseeded Dil No Nit Inh 5 day | The biochemical oxygen demand is determined by measuring the amount of dissolved oxygen (DO) that is consumed in the sample under aerobic conditions during the incubation period (five days at 20 C). A polarographic membrane electrode that has been calibrated in saturated air is used to acquire all dissolved oxygen measurements. No nitrogen inhibitor added. No seed is added. A modified Winkler procedure may be substituted for the DO probe procedure. Performance Base Method published circa Dec. 2000. | 5 | Y | Y | Y | 1572 | 8/1/1997 | 6/28/2024 | 7/29/2024 3:17:44 PM | 1997-12-11 | 6/3/1998 6:51:47 PM | NPEPPIN |  | Concentration | mg/L | Y | (fl. conc.) |  | Biochemical Oxygen Demand (fl. conc.) |  | Aggregate Organics; | Aggregate Organics; EMS code: 0115 | Chemical |  | LAB
0115 | mg/L | Biochemical Oxygen Demand | X483 | Seeded Dil No Nit Inh | The biochemical oxygen demand is determined by measuring the amount of dissolved oxygen (DO) that is consumed in the sample under aerobic conditions during the incubation period (five days at 20 C) after seeding material has been added. A polarographic membrane electrode that has been calibrated in saturated air is used to acquire all dissolved oxygen measurements. No nitrogen inhibitor added. A modified Winkler procedure may be substituted for the DO probe procedure. Performance Base Method published circa Dec. 2000. | 5 | Y | Y | Y | 15919 | 1/22/1998 6:51:00 AM | 8/22/2024 8:00:00 AM | 8/25/2024 9:47:36 PM | 1997-12-11 | 6/3/1998 6:51:47 PM | NPEPPIN |  | Concentration | mg/L | Y | (fl. conc.) |  | Biochemical Oxygen Demand (fl. conc.) |  | Aggregate Organics; | Aggregate Organics; EMS code: 0115 | Chemical |  | LAB
0115 | mg/L | Biochemical Oxygen Demand | X013 | Seeded Dil Nit Inh 5 Day | The biochemical oxygen demand is determined by measuring the amount of dissolved oxygen (DO) that is consumed in the sample under aerobic conditions during the incubation period (five days at 20 C) after seeding material has been added. A polarographic membrane electrode that has been calibrated in saturated air is used to acquire all dissolved oxygen measurements. Nitrification inhibitor added. A modified Winkler procedure may be substituted for the DO probe procedure. Performance Base Method published circa Dec. 2000. | 5 | Y | Y | Y | 89511 | 9/8/1994 | 8/21/2024 2:06:00 PM | 8/23/2024 9:49:37 AM | 1996-04-12 | 6/3/1998 6:51:42 PM | EMS |  | Concentration | mg/L | Y | (fl. conc.) |  | Biochemical Oxygen Demand (fl. conc.) |  | Aggregate Organics; | Aggregate Organics; EMS code: 0115 | Chemical |  | LAB
0115 | ug/m3 | Biochemical Oxygen Demand | EQ40 | EQUIS (mg/m3) | EQUIS (mg/m3) | 0 | N | Y | Y | 16 |  |  | 10/22/1996 7:18:29 PM | 1996-04-12 | 6/3/1998 6:51:33 PM | EMS |  | AirConcentration | µg/m³ | Y | (air conc.) |  | Biochemical Oxygen Demand (air conc.) |  | Aggregate Organics; | Aggregate Organics; EMS code: 0115 | Chemical |  | LAB
0116 | mg/L | Chemical Oxygen Demand | EQ01 | EQUIS (mg/L) | EQUIS (mg/L) | 10 | N | Y | Y | 2255 |  |  | 10/23/1996 12:05:41 AM | 1996-04-12 | 6/3/1998 6:51:33 PM | EMS |  | Concentration | mg/L | Y | (fl. conc.) |  | Chemical Oxygen Demand (fl. conc.) |  | Aggregate Organics; | Aggregate Organics; EMS code: 0116 | Chemical |  | LAB
0116 | mg/L | Chemical Oxygen Demand | X065 | K2Cr2O7 Dig:Titr. | K2Cr2O7 Dig:Titr. | 10 | Y | Y | Y | 15953 | 12/12/1996 | 8/16/2024 | 8/19/2024 4:11:05 PM | 1996-04-12 | 6/3/1998 6:51:46 PM | EMS |  | Concentration | mg/L | Y | (fl. conc.) |  | Chemical Oxygen Demand (fl. conc.) |  | Aggregate Organics; | Aggregate Organics; EMS code: 0116 | Chemical |  | LAB
0116 | mg/L | Chemical Oxygen Demand | X559 | Haach,Mn(3)sol'nDig,KPHcol'm,510nm | Hach MnIII COD method uses the oxidant trivalent manganese. Sodium bismuthate is used to remove chloride interferences. Using a vacuum pump, an acidified homogenized sample is filtered through a glass fiber filter, then through sodium bismuthate into a vial containing the trivalent manganese solution. The glass fiber filter traps solids and is removed from the filter and placed into the MnIII vial. The capped vial is digested at 150 degrees C for 1 hour. As the MnIII oxidizes the organics present in the sample, the purple color of the MnIII solution decreases, becoming clear when all MnIII has been reduced. The COD concentration is proportional to the decrease in the purple color of the solution. COD is determined colorimetrically at 510 nm using a calibration curve from potassium hydrogen phthalate standards. | 4 | Y | Y | Y | 13019 | 1/1/2002 10:00:00 AM | 7/27/2024 8:00:00 AM | 8/23/2024 8:10:36 AM | 2004-07-20 |  | UMULLER |  | Concentration | mg/L | Y | (fl. conc.) |  | Chemical Oxygen Demand (fl. conc.) |  | Aggregate Organics; | Aggregate Organics; EMS code: 0116 | Chemical |  | LAB
0116 | mg/L | Chemical Oxygen Demand | X315 | Open reflux,K2Cr2O7/H2SO4 dig, FAS Titr. | Organic and oxidizable inorganic substances are oxidized by potassium dichromate in H2SO4 solution at reflux temperature for 2 hours. Excess dichromate is titrated with standard ferrous ammonium sulfate (0.1M) using orthophenanthroline ferrous complex (ferroin) as indicator. Open Reflux. | 5 | Y | Y | Y | 8681 | 1/2/1997 8:00:00 AM | 5/1/2014 1:46:00 PM | 2/11/2017 7:53:30 PM | 1996-04-12 | 6/3/1998 6:51:45 PM | EMS |  | Concentration | mg/L | Y | (fl. conc.) |  | Chemical Oxygen Demand (fl. conc.) |  | Aggregate Organics; | Aggregate Organics; EMS code: 0116 | Chemical |  | LAB
0116 | mg/L | Chemical Oxygen Demand | X504 | Closed refux,K2Cr2O7 dig,col'm,600nm | Organic and oxidizable inorganic substances are oxidized by potassium dichromate in H2SO4 solution at reflux temperature for 2 hours in a closed reflux. Oxygen consumption is measured colorimetrically against standards at 600nm. | 5 | Y | Y | Y | 8217 | 3/1/2000 | 8/16/2024 8:00:00 AM | 8/25/2024 9:47:36 PM | 1998-12-09 |  | NPEPPIN |  | Concentration | mg/L | Y | (fl. conc.) |  | Chemical Oxygen Demand (fl. conc.) |  | Aggregate Organics; | Aggregate Organics; EMS code: 0116 | Chemical |  | LAB
0116 | mg/L | Chemical Oxygen Demand | F031 | K2Cr2O7/H2SO4 Dig., Ag2SO4, FAS Titr. | Most types of organic matter, and any elements with a lower oxidation potential than chromium, are oxidized by boiling in a mixture with known excess of potassium dichromate and sulfuric acid and the excess dichromate is titrated with ferrous ammonium sulfate (FAS). The amount of oxidizable organic matter is proportional to the potassium dichromate consumed. Silver sulfate is added as a catalyst so that refractory organics such as acetic acid, which are biologically available to stream organisms, are included in the oxidation. With the catalyst, other compounds such as cellulose, which are not part of the immediate biochemical load, are also included. Any relationships among COD and BOD or any other tests are fortuitous, because condition of the test tend to give results indicating the susceptibility of a given sample to oxidation under specified conditions that are different for each test. (Note: This procedure is NOT appropriate for seawater.) | 10 | Y | Y | Y | 2330 | 7/30/1996 3:36:00 PM | 1/14/2015 | 8/28/2018 3:13:36 PM | 1997-01-14 | 6/3/1998 6:51:46 PM | MPAK |  | Concentration | mg/L | Y | (fl. conc.) |  | Chemical Oxygen Demand (fl. conc.) |  | Aggregate Organics; | Aggregate Organics; EMS code: 0116 | Chemical |  | LAB
0116 | mg/L | Chemical Oxygen Demand | X525 | Closed reflux,K2Cr2O7/H2SO4 dig,FAS,Titr. | Organic and oxidizable inorganic substances are oxidized by potassium dichromate in H2SO4 solution at reflux temperature for 2 hours. Excess dichromate is titrated with standard ferrous ammonium sulfate (0.1M) using orthophenanthroline ferrous complex (ferroin) as indicator. Closed Reflux. | 5 | Y | Y | Y | 4 |  |  | 6/7/2024 5:30:18 PM | 1999-07-15 |  | NPEPPIN |  | Concentration | mg/L | Y | (fl. conc.) |  | Chemical Oxygen Demand (fl. conc.) |  | Aggregate Organics; | Aggregate Organics; EMS code: 0116 | Chemical |  | LAB
0117 | mg/L | Phenols | 0550 | 4-Amino Anti Pyr:Vis | 4-Amino Anti Pyr:Vis | 0.002 | Y | Y | Y | 1159 | 9/26/2022 | 6/28/2024 | 6/28/2024 5:01:19 PM | 1996-04-12 | 6/3/1998 6:51:33 PM | EMS |  | Concentration | mg/L | Y | (fl. conc.) |  | Phenols (fl. conc.) |  | Organics; | Organics; EMS code: 0117 | Chemical |  | LAB
0117 | ug/m3 | Phenols | EQ40 | EQUIS (mg/m3) | EQUIS (mg/m3) | 0 | N | Y | Y | 73 |  |  | 10/22/1996 7:18:31 PM | 1996-04-12 | 6/3/1998 6:51:33 PM | EMS |  | AirConcentration | µg/m³ | Y | (air conc.) |  | Phenols (air conc.) |  | Organics; | Organics; EMS code: 0117 | Chemical |  | LAB
0117 | mg/L | Phenols | X094 | Dst,Amino Anti Pyr,KFeCN | Dst;Amino Anti Pyr;KFeCN | 0.002 | Y | Y | Y | 36 |  |  | 3/26/1997 2:27:34 PM | 1996-04-12 | 6/3/1998 6:51:45 PM | EMS |  | Concentration | mg/L | Y | (fl. conc.) |  | Phenols (fl. conc.) |  | Organics; | Organics; EMS code: 0117 | Chemical |  | LAB
0117 | mg/L | Phenols | AA02 | Automated 4-APP,dist | Automated 4-APP;dist | 0.001 | Y | Y | Y | 1681 | 1/6/2007 | 8/13/2024 | 8/16/2024 4:05:21 PM | 1996-04-12 | 6/3/1998 6:51:42 PM | EMS |  | Concentration | mg/L | Y | (fl. conc.) |  | Phenols (fl. conc.) |  | Organics; | Organics; EMS code: 0117 | Chemical |  | LAB
0117 | mg/L | Phenols | X257 | Colorimetric | Colorimetric. General use for several parameters. E.g.Total Sulfide is determined by spectrophotometer using the methlyene blue colourimetric method. The H2S calculation converts Total Sulphide as (S2-) and reports it as Total Sulphide as (H2S). | 0.01 | Y | Y | Y | 1897 | 6/4/1997 9:52:00 AM | 3/27/1998 9:34:00 AM | 8/23/2024 8:10:35 AM | 1996-04-12 | 6/3/1998 6:51:33 PM | EMS |  | Concentration | mg/L | Y | (fl. conc.) |  | Phenols (fl. conc.) |  | Organics; | Organics; EMS code: 0117 | Chemical |  | LAB
0117 | mg/L | Phenols | EQ01 | EQUIS (mg/L) | EQUIS (mg/L) | 0.002 | N | Y | Y | 1980 |  |  | 10/23/1996 12:06:54 AM | 1996-04-12 | 6/3/1998 6:51:33 PM | EMS |  | Concentration | mg/L | Y | (fl. conc.) |  | Phenols (fl. conc.) |  | Organics; | Organics; EMS code: 0117 | Chemical |  | LAB
0117 | mg/L | Phenols | AA01 | Automated 4-APP,dist | The phenolic compounds are removed from the sample interferences by means of on-line distillation under acidic conditions. The distillate is subsequently reacted with 4-aminoantipyrine and an alkaline ferricyanide to form a red complex which is measured colourimetrically at 505 nm. This method is applicable to the analysis of drinking, surface and saline waters, domestic and industrial wastes and acid rain (atmospheric deposition). | 0.002 | Y | Y | Y | 192 | 7/29/1998 1:46:00 PM | 9/17/2021 4:53:00 PM | 2/1/2022 2:38:14 PM | 1996-04-12 | 6/3/1998 6:51:33 PM | EMS |  | Concentration | mg/L | Y | (fl. conc.) |  | Phenols (fl. conc.) |  | Organics; | Organics; EMS code: 0117 | Chemical |  | LAB
0117 | mg/L | Phenols | X142 | Dst,Amino Anti Pyr,Col | The unfiltered preserved sample is distilled under acidic conditions to remove most interferences. The addition of 4-aminoantipyrine in the presence of potassium ferricyanide forms a red coloured antipyrine dye which remains in an aqueous solution and the absorbance is measured at 505 nm. This method determines phenol, ortho- and meta- substituted phenols as phenol. | 0.002 | Y | Y | Y | 112 |  |  | 10/24/1996 2:18:17 AM | 1996-04-12 | 6/3/1998 6:51:33 PM | EMS |  | Concentration | mg/L | Y | (fl. conc.) |  | Phenols (fl. conc.) |  | Organics; | Organics; EMS code: 0117 | Chemical |  | LAB
0118 | mg/L | Phosphorus Ortho | X125 | Dig,Stannous Cl2,Col'met | Dig;Stannous Cl2;Col'met | 0 | Y | Y | Y | 7140 | 8/9/1996 10:24:00 AM | 3/26/1998 6:30:00 PM | 8/26/2024 2:41:32 PM | 1996-04-12 | 6/3/1998 6:51:45 PM | EMS |  | Concentration | mg/L | Y | (fl. conc.) |  | Phosphorus Ortho (fl. conc.) | 14265-44-2 | Nutrients; | Nutrients; EMS code: 0118 | Chemical |  | LAB
0118 | mg/L | Phosphorus Ortho | EQ01 | EQUIS (mg/L) | EQUIS (mg/L) | 0 | N | Y | Y | 1124 |  |  | 10/24/1996 3:16:55 PM | 1996-04-12 | 6/3/1998 6:51:33 PM | EMS |  | Concentration | mg/L | Y | (fl. conc.) |  | Phosphorus Ortho (fl. conc.) | 14265-44-2 | Nutrients; | Nutrients; EMS code: 0118 | Chemical |  | LAB
0118 | mg/L | Phosphorus Ortho | X453 | Stannous Cl2, Col'metric | Stannous Cl2; Col'metric | 0.003 | Y | Y | Y | 2074 | 5/25/1999 | 7/3/2024 | 8/9/2024 3:11:45 PM | 1996-04-12 | 6/3/1998 6:51:33 PM | EMS |  | Concentration | mg/L | Y | (fl. conc.) |  | Phosphorus Ortho (fl. conc.) | 14265-44-2 | Nutrients; | Nutrients; EMS code: 0118 | Chemical |  | LAB
0119 | ug/g | Phenol | 1332 | Phenolics by Tumbler/GC-MS | Phenolics by Tumbler/GC-MS | 0.02 | Y | Y | Y | 15 | 1/17/2022 | 2/2/2022 | 1/11/2024 1:07:06 PM | 2022-07-07 |  | EMS |  | MassPerMass | mg/kg | Y | (mass conc.) |  | Phenol (mass conc.) | 108-95-2 | Organics; | Organics; EMS code: 0119 | Chemical |  | LAB
0119 | mg/L | Phenol | X268 | CH2Cl2 ext:GC/MS | CH2Cl2 ext:GC/MS | 0.001 | Y | Y | Y | 859 | 3/6/2007 | 8/14/2024 | 8/16/2024 4:05:21 PM | 2012-03-27 |  | RDGIBSON |  | Concentration | mg/L | Y | (fl. conc.) |  | Phenol (fl. conc.) | 108-95-2 | Organics; | Organics; EMS code: 0119 | Chemical |  | LAB
0119 | mg/L | Phenol | F028 | Acid,Extr.MeCl,HRGC/LRMS:TIM,SIM | A measured volume of sample is acidified and extracted with Methylene Chloride. For samples which have a high organic content that potentially produce chromatographic interferences an initial pre-extraction under basic conditions is performed. The sample extract is then evaporated to small volume and analysed by High Resolution Gas Chromatography / Low Resolution Mass Spectrometry (HRGC/LRMS). LRMS data is acquired in Total Ion Mode (TIM) and quantitation is performed using a single mass ion. Positive Identification of the various target compounds is made by comparison of absolute retention times to those of standards combined with computer based mass spectral mass/intensity library searching and matching routines. | 0.005 | Y | Y | Y | 114 | 4/10/1997 3:17:00 PM | 4/14/2015 12:26:00 AM | 4/16/2015 6:00:10 PM | 1997-01-14 | 6/3/1998 6:51:46 PM | MPAK |  | Concentration | mg/L | Y | (fl. conc.) |  | Phenol (fl. conc.) | 108-95-2 | Organics; | Organics; EMS code: 0119 | Chemical |  | LAB
0120 | mg/L | Silica Reactive | X338 | Auto Ascorb.red.het.blue | Ammonium molybdate at pH 1.2 reacts with silica to form yellow molybdosilicic acid which is reduced by ascorbic acid to produce an intensely blue heteropoly acid. The absorbance is read at 600 nm. | 0.5 | Y | Y | Y | 233 | 12/10/2019 9:05:00 AM | 2/26/2024 1:13:00 PM | 2/26/2024 5:01:21 PM | 1996-04-12 | 6/3/1998 6:51:33 PM | EMS |  | Concentration | mg/L | Y | (fl. conc.) |  | Silica Reactive (fl. conc.) | 7631-86-9 | Nutrients; | Nutrients; EMS code: 0120 | Chemical |  | LAB
0120 | mg/L | Silica Reactive | X339 | Auto molybdosilicate | Ammonium molybdate at pH 1.2 reacts with silica to form yellow molybdosilicic acid. The absorbance is read at 410 nm. | 1 | Y | Y | Y | 1 |  |  | 5/11/1998 2:17:56 PM | 1996-04-12 | 6/3/1998 6:51:45 PM | EMS |  | Concentration | mg/L | Y | (fl. conc.) |  | Silica Reactive (fl. conc.) | 7631-86-9 | Nutrients; | Nutrients; EMS code: 0120 | Chemical |  | LAB
0120 | mg/L | Silica Reactive | EQ01 | EQUIS (mg/L) | EQUIS (mg/L) | 0.5 | N | Y | Y | 585 |  |  | 10/24/1996 3:16:55 PM | 1996-04-12 | 6/3/1998 6:51:33 PM | EMS |  | Concentration | mg/L | Y | (fl. conc.) |  | Silica Reactive (fl. conc.) | 7631-86-9 | Nutrients; | Nutrients; EMS code: 0120 | Chemical |  | LAB
0121 | mg/L | Sulfate Total | X004 | Colorimetric Titr'n | Manual Colorimetric Titration with standard solutions to specified endpoint as identified by colorimetric indicator. The sample must not be filtered, diluted, concentrated, or altered in any way. | 1 | Y | Y | Y | 1598 | 9/29/2000 | 12/1/2010 | 6/7/2024 5:30:18 PM | 1996-04-12 | 6/3/1998 6:51:33 PM | EMS |  | Concentration | mg/L | Y | (fl. conc.) |  | Sulfate Total (fl. conc.) | 14808-79-8 | Major Ions; | Major Ions; EMS code: 0121 | Chemical | Total | LAB
0121 | mg/L | Sulfate Total | X061 | BaCl2,Gravimetric | Sulfate is precipitated as barium sulfate (BaSO4) in HCl medium by the addition of barium chloride. After the digestion period, the precipitate is filtered, washed with hot water until chloride free, ignited and weighed as BaSO4. | 1 | Y | Y | Y | 963 |  |  | 1/24/2004 1:05:00 PM | 1996-04-12 | 6/3/1998 6:51:45 PM | EMS |  | Concentration | mg/L | Y | (fl. conc.) |  | Sulfate Total (fl. conc.) | 14808-79-8 | Major Ions; | Major Ions; EMS code: 0121 | Chemical | Total | LAB
0121 | mg/L | Sulfate Total | X170 | ICP | This code is a catch-all for a large variety of ICP methods including water and also solids digested into solutions as an intermediate result. Different procedures will have greatly different detection limits. (a) For air particulates, preliminary results are reported in mg/L and loading results as mg/dm2/d. (b) For Hi Vol samples preliminiary results are same as reported in mg/L and loading results as ug/m3. (c) Note that for Rain samples there is no filtration nor digestion steps. | 1 | N | Y | Y | 223 |  |  | 3/4/1997 7:02:36 PM | 1996-04-12 | 6/3/1998 6:51:33 PM | EMS |  | Concentration | mg/L | Y | (fl. conc.) |  | Sulfate Total (fl. conc.) | 14808-79-8 | Major Ions; | Major Ions; EMS code: 0121 | Chemical | Total | LAB
0121 | mg/L | Sulfate Total | X044 | Ion Chromatography | A clear (Total) or filtered (Dissolved) water sample is injected into a stream of carbonate-bicarbonate eluent and passed through a series of ion exchanger. The anions of interested are separated on the basis of their relative affinities for a low capacity , strongly basic anion exchanger (guard and separate columns). The separated anions are directed through the anion self-regenerating suppressor (ASRS). Hydronium ions in the suppressor will neutralize the eluent to its weakly ionized form and convert the separated anions in their conductive acid form. The separated anions are measured by conductivity. They are identified on the basis of retention time as compared to standards. Quantitation is by measurement of peak area of peak height against a calibration curve generated from known standards. | 1 | Y | Y | Y | 60713 | 4/10/1996 6:56:00 AM | 11/27/2015 | 8/23/2024 8:10:35 AM | 1996-04-12 | 6/3/1998 6:51:33 PM | EMS |  | Concentration | mg/L | Y | (fl. conc.) |  | Sulfate Total (fl. conc.) | 14808-79-8 | Major Ions; | Major Ions; EMS code: 0121 | Chemical | Total | LAB
0121 | mg/L | Sulfate Total | X064 | BaCl2,Turbimetric | Sulfate ion is converted to a barium sulfate suspension under controlled conditions. The resulting turbidity is determined using a nephelometer or spectrophotometer and compared to a curve prepared from standard sulfate solutions. | 1 | Y | Y | Y | 7052 | 1/16/1998 | 2/11/2015 | 2/15/2018 11:14:41 AM | 1996-04-12 | 6/3/1998 6:51:33 PM | EMS |  | Concentration | mg/L | Y | (fl. conc.) |  | Sulfate Total (fl. conc.) | 14808-79-8 | Major Ions; | Major Ions; EMS code: 0121 | Chemical | Total | LAB
0121 | mg/L | Sulfate Total | EQ01 | EQUIS (mg/L) | EQUIS (mg/L) | 0 | N | Y | Y | 862 |  |  | 10/24/1996 3:16:17 PM | 1996-04-12 | 6/3/1998 6:51:33 PM | EMS |  | Concentration | mg/L | Y | (fl. conc.) |  | Sulfate Total (fl. conc.) | 14808-79-8 | Major Ions; | Major Ions; EMS code: 0121 | Chemical | Total | LAB
0122 | mg/L | Surfactants | X341 | Methylene blue,colorimetric(Anionic) | The sample, made just acid to phenolphthalein and treated with an excess of methylene blue, is extracted three times with chloroform (CHCl3). The combined chloroform extracts are washed with acidic buffer solution, dried and made to volume for colorimetric measurement at 652nm. The intensity of the resulting blue color in the organic phase is a measure of MBAS. Linear alkylbezene sulfonate (LAS) is the most widely used anionic surfactant and is used to standardize the MBAS method. This method is a measure of anionic surfactants. | 0.03 | Y | Y | Y | 297 | 2/27/2007 3:51:00 PM | 8/2/2024 | 8/2/2024 5:01:07 PM | 1996-04-12 | 6/3/1998 6:51:41 PM | EMS |  | Concentration | mg/L | Y | (fl. conc.) |  | Surfactants (fl. conc.) |  | Organics; | Organics; EMS code: 0122 | Chemical |  | LAB
0122 | mg/L | Surfactants | F017 | Azure A, Col'm(Anionic) | An Azure-A dye is added to an acidified sample forming a blue complex which is subsequently extracted into Chloroform. Standards, made using Sodium Lauryl Sulfate, are extracted similarly. The Absorbance of the Chloroform extracts is measured on a UV-VIS spectrophotometer at 623 nm. The concentration of LAS equivalents in the sample is determined by comparison of the sample Absorbance to a standard curve. Surfactants may be cationic, anionic, or non-ionic. This method is a measure of anionic surfactants. | 0.04 | Y | Y | Y | 6 | 5/19/2000 9:28:00 AM | 7/12/2001 2:32:00 PM | 8/22/2001 1:06:28 AM | 1996-06-19 | 6/3/1998 6:51:44 PM | MPAK |  | Concentration | mg/L | Y | (fl. conc.) |  | Surfactants (fl. conc.) |  | Organics; | Organics; EMS code: 0122 | Chemical |  | LAB
0122 | mg/L | Surfactants | F077 | Methyl Orange,col'm (cationic) | An excess of Methyl-Orange reagent is added to a buffered sample forming a yellow complex which is subsequently extracted into Chloroform. Standards, made using Cetyldimethylbenzylammonium Chloride (CDBAC), are extracted similarly. The Absorbance of the Chloroform extracts is measured on a UV-VIS spectrophotometer at 415nm. The concentration of CDBAC equivalents in the sample is determined by comparison of the sample Absorbance to that of the standards. This method is a measure of cationic surfactants. | 0.5 | Y | Y | Y | 6 | 5/19/2000 9:28:00 AM | 7/12/2001 2:32:00 PM | 8/22/2001 1:06:29 AM | 1999-10-04 |  | NPEPPIN |  | Concentration | mg/L | Y | (fl. conc.) |  | Surfactants (fl. conc.) |  | Organics; | Organics; EMS code: 0122 | Chemical |  | LAB
0122 | mg/L | Surfactants | EQ01 | EQUIS (mg/L) | EQUIS (mg/L) | 0 | N | Y | Y | 565 |  |  | 10/24/1996 3:16:17 PM | 1996-04-12 | 6/3/1998 6:51:33 PM | EMS |  | Concentration | mg/L | Y | (fl. conc.) |  | Surfactants (fl. conc.) |  | Organics; | Organics; EMS code: 0122 | Chemical |  | LAB
0123 | mg/L | Tannin &Lignin Total | 0951 | Heteropoly Acids - auto. | Aliquots of sample are reacted with Folin phenol reagent (a mixture of tungstic and molybdic acids) and, after time for reaction, with carbonate solution. The absorbance of the developed colour is measured at 700 nm using a spectro-photometer equipped with 1 cm cells. It should be emphasized that the reaction is not specific for lignin or tannin. Automated procedure. | 0.1 | Y | Y | Y | 40 | 1/7/1997 9:31:00 AM | 5/23/2001 1:40:00 PM | 2/14/2002 1:33:45 AM | 1996-04-12 | 6/3/1998 6:51:45 PM | EMS |  | Concentration | mg/L | Y | (fl. conc.) |  | Tannin &Lignin Total (fl. conc.) |  | Organics; | Organics; EMS code: 0123 | Chemical | Total | LAB
0123 | mg/L | Tannin &Lignin Total | 0950 | Heteropoly Acid | Heteropoly Acid | 0.1 | Y | Y | Y | 118 | 2/22/2002 12:40:00 PM | 10/1/2003 11:40:00 AM | 6/7/2024 5:30:18 PM | 1996-04-12 | 6/3/1998 6:51:33 PM | EMS |  | Concentration | mg/L | Y | (fl. conc.) |  | Tannin &Lignin Total (fl. conc.) |  | Organics; | Organics; EMS code: 0123 | Chemical | Total | LAB
0123 | mg/L | Tannin &Lignin Total | E563 | APHA 5550 B-Colorimetry | APHA 5550 B-Colorimetry This analysis is carried out using procedures adapted from APHA Method 5550 B. "Tannin & Lignin ". Both lignin and tannin contain aromatic hydroxyl groups that react with Folin phenol reagent (tungstophosphoric and molybdophosphoric acids) to form a blue color suitable for the estimation of tannin and lignin concentrations. However; the reaction is not specific for lignin or tannin; nor for compounds containing aromatic hydroxyl groups; in as much as many other reducing materials; both organic and inorganic; respond similarly. | 0.1 | Y | Y | Y | 14 | 11/29/2022 | 11/24/2023 | 6/7/2024 5:30:17 PM | 2021-12-30 |  | EMS |  | Concentration | mg/L | Y | (fl. conc.) |  | Tannin &Lignin Total (fl. conc.) |  | Organics; | Organics; EMS code: 0123 | Chemical | Total | LAB
0123 | mg/L | Tannin &Lignin Total | X120 | Heteropoly Acids,Col | Aliquots of sample are reacted with Folin phenol reagent (a mixture of tungstic and molybdic acids) and, after time for reaction, with carbonate solution. The absorbance of the developed colour is measured at 700 nm using a spectro-photometer equipped with 1 cm cells. It should be emphasized that the reaction is not specific for lignin or tannin. Manual procedure. | 0.1 | Y | Y | Y | 43 | 8/1/2006 | 10/3/2015 | 8/12/2016 5:10:57 PM | 1996-04-12 | 6/3/1998 6:51:33 PM | EMS |  | Concentration | mg/L | Y | (fl. conc.) |  | Tannin &Lignin Total (fl. conc.) |  | Organics; | Organics; EMS code: 0123 | Chemical | Total | LAB
0123 | mg/L | Tannin &Lignin Total | EQ01 | EQUIS (mg/L) | EQUIS (mg/L) | 0 | N | Y | Y | 3461 |  |  | 10/24/1996 3:16:55 PM | 1996-04-12 | 6/3/1998 6:51:33 PM | EMS |  | Concentration | mg/L | Y | (fl. conc.) |  | Tannin &Lignin Total (fl. conc.) |  | Organics; | Organics; EMS code: 0123 | Chemical | Total | LAB
0124 | % | Carbon Total Inorganic | FS86 | HCl/FeCl2, Gravimetric | HCl/FeCl2;Gravimetric |  | Y | Y | Y | 11 | 12/14/2000 2:12:00 PM | 12/14/2000 2:13:00 PM | 8/13/2002 1:05:58 PM | 2001-03-08 |  | NPEPPIN |  | DimensionlessRatio | % | N | (ratio) |  | Carbon Total Inorganic (ratio) |  | Carbon; | Carbon; EMS code: 0124 | Chemical | Total | LAB
0124 | % | Carbon Total Inorganic | CSSS20 | CSSS (2008) 20.2 | See Canadian Society of Soil Science (2008) section 20.2 | 0.05 | Y | Y | Y | 31 | 1/13/2022 | 2/20/2024 | 2/26/2024 6:09:04 PM | 2022-06-28 |  | EMS |  | DimensionlessRatio | % | N | (ratio) |  | Carbon Total Inorganic (ratio) |  | Carbon; | Carbon; EMS code: 0124 | Chemical | Total | LAB
0124 | ug/g | Carbon Total Inorganic | 0941 | Ash on Leco | Ash on Leco | 500 | Y | Y | Y | 1161 | 4/28/2000 7:31:00 PM | 8/29/2006 | 9/20/2006 1:54:26 AM | 1996-04-12 | 6/3/1998 6:51:33 PM | EMS |  | MassPerMass | mg/kg | Y | (mass conc.) |  | Carbon Total Inorganic (mass conc.) |  | Carbon; | Carbon; EMS code: 0124 | Chemical | Total | LAB
0124 | mg/L | Carbon Total Inorganic | XCAL | Calculation (mg/L) | Calculation (mg/L) | 0.5 | Y | Y | Y | 462 | 7/31/1996 12:51:00 PM | 10/6/1997 8:45:00 AM | 10/17/1997 6:11:01 PM | 1996-04-12 | 6/3/1998 6:51:41 PM | EMS |  | Concentration | mg/L | Y | (fl. conc.) |  | Carbon Total Inorganic (fl. conc.) |  | Carbon; | Carbon; EMS code: 0124 | Chemical | Total | LAB
0124 | mg/L | Carbon Total Inorganic | F127 | PersulDig:CO2Purge/Col'm,550nm | Total and Dissolved Inorganic Carbon by Automated Persulfate Colourimetric Method (TAACS 800 Industrial Method No. 860-87T, January 1989): The sample is mixed with phosphate buffer and persulphate in order to digest the inorganic compounds and convert them to CO2 Dialysis of the resultant CO2 through a gas-permeable silicone membrane to combine with a phenolphthalein solution causes a decrease on colour of the phenophthalein solution proportional to the original inorganic carbon concentration. This colour change is measured at a wavelength of 550nm. | 0.5 | Y | Y | Y | 2333 | 11/15/2004 | 11/24/2015 5:59:00 PM | 11/30/2015 4:44:35 PM | 2005-01-11 |  | UMULLER |  | Concentration | mg/L | Y | (fl. conc.) |  | Carbon Total Inorganic (fl. conc.) |  | Carbon; | Carbon; EMS code: 0124 | Chemical | Total | LAB
0124 | ug/g | Carbon Total Inorganic | CAL2 | Calculation (ug/g) | Calculation (ug/g) | 0.1 | Y | Y | Y | 84 | 5/25/2005 | 5/25/2005 | 10/15/2005 1:02:13 AM | 1996-04-12 | 6/3/1998 6:51:42 PM | EMS |  | MassPerMass | mg/kg | Y | (mass conc.) |  | Carbon Total Inorganic (mass conc.) |  | Carbon; | Carbon; EMS code: 0124 | Chemical | Total | LAB
0124 | ug/g | Carbon Total Inorganic | CS02 | System Calculation | System Calculation (ug/g) |  | N | Y | N | 100 |  |  | 4/26/2001 1:59:10 AM | 1996-04-12 | 6/3/1998 6:51:43 PM | EMS |  | MassPerMass | mg/kg | Y | (mass conc.) |  | Carbon Total Inorganic (mass conc.) |  | Carbon; | Carbon; EMS code: 0124 | Chemical | Total | LAB
0124 | mg/L | Carbon Total Inorganic | 0940 | Vap.bypass/FID Dohrmann | Vap.bypass/FID Dohrmann | 1 | Y | Y | Y | 3181 | 2/25/2005 | 8/17/2010 | 9/1/2010 12:17:21 PM | 1996-04-12 | 6/3/1998 6:51:33 PM | EMS |  | Concentration | mg/L | Y | (fl. conc.) |  | Carbon Total Inorganic (fl. conc.) |  | Carbon; | Carbon; EMS code: 0124 | Chemical | Total | LAB
0124 | mg/L | Carbon Total Inorganic | EQ01 | EQUIS (mg/L) | EQUIS (mg/L) | 0.5 | N | Y | Y | 4219 |  |  | 10/24/1996 3:16:34 PM | 1996-04-12 | 6/3/1998 6:51:33 PM | EMS |  | Concentration | mg/L | Y | (fl. conc.) |  | Carbon Total Inorganic (fl. conc.) |  | Carbon; | Carbon; EMS code: 0124 | Chemical | Total | LAB
0124 | % | Carbon Total Inorganic | IR0717 | ALS Minerals C-IR07/C-IR17 | ALS Minerals C-IR07/C-IR17. Total inorganic carbon is calculated by difference from total carbon (C-IR07) minus total organic carbon (C-IR17). | 0.01 | Y | Y | Y | 0 |  |  |  | 2024-01-29 |  | RDGIBSON |  | DimensionlessRatio | % | N | (ratio) |  | Carbon Total Inorganic (ratio) |  | Carbon; | Carbon; EMS code: 0124 | Chemical | Total | LAB
0124 | mg/L | Carbon Total Inorganic | 0930 | H3PO4/Quartz Tube/IR | H3PO4/Quartz Tube/IR | 2 | Y | Y | Y | 2082 | 6/4/1997 | 6/4/1997 | 2/7/2000 3:11:43 PM | 1996-04-12 | 6/3/1998 6:51:33 PM | EMS |  | Concentration | mg/L | Y | (fl. conc.) |  | Carbon Total Inorganic (fl. conc.) |  | Carbon; | Carbon; EMS code: 0124 | Chemical | Total | LAB
0124 | mg/L | Carbon Total Inorganic | X171 | Carbon Anal.,Infra Red | Analysis is comprised of 2 phases, the determination of Total Carbon (TC) or Dissolved Carbon (DC) and the determination of Total Inorganic Carbon (TIC) or Dissolved Inorganic Carbon (DIC). Total Organic Carbon (TOC) content, or Dissolved Organic Carbon (DOC) is then calculated by difference. A series of External Standard solutions, Potassium Biphthalate for TC (DC) and Sodium Carbonate/Sodium Bicarbonate for TIC (DIC) is used for instrument calibration. During TC (DC) analysis, a small portion of the sample (or standard) solution is introduced into a combustion tube, containing a Platinum catalyst at an elevated temperature with an Oxygen gas carrier, via an automated autosampler/injector. The sample's TC (DC) component is combusted/oxidized to form CO2. The CO2 flows through an IC reaction vessel, is cooled and dried by an electronic dehumidifier and passes through a halogen scrubber into a Non-dispersive Infrared Detector (NDIR). The NDIR produces an analog signal over a constant time base that a data processor converts into a peak, the area of which is proportional to the TC concentration of the sample. The TC (DC) concentration of a sample is determined electronically from a calibration curve expressing the relationship between the peak area of a standard and its concentration. Similarly, for the TIC (DIC) phase of the analysis, another small portion of sample (or standard) solution is introduced directly into the IC reaction vessel, containing a Phosphoric Acid solution through which carrier gas is bubbled. However, here only the IC component of the sample reacts to form CO2. The TOC (DOC) concentration is automatically determined by the data processor from the difference between TC (DC) and TIC (DOC) concentrations. | 0.5 | Y | Y | Y | 6243 | 1/9/1997 9:17:00 AM | 8/7/2024 | 8/13/2024 4:39:04 PM | 1996-04-12 | 6/3/1998 6:51:33 PM | EMS |  | Concentration | mg/L | Y | (fl. conc.) |  | Carbon Total Inorganic (fl. conc.) |  | Carbon; | Carbon; EMS code: 0124 | Chemical | Total | LAB
0124 | ug/g | Carbon Total Inorganic | EQ02 | EQUIS (ug/g) | EQUIS (ug/g) | 0.3 | N | Y | Y | 70 |  |  | 10/22/1996 11:18:43 PM | 1996-04-12 | 6/3/1998 6:51:33 PM | EMS |  | MassPerMass | mg/kg | Y | (mass conc.) |  | Carbon Total Inorganic (mass conc.) |  | Carbon; | Carbon; EMS code: 0124 | Chemical | Total | LAB
0125 | ug/m3 | Sulfide Total | EQ40 | EQUIS (mg/m3) | EQUIS (mg/m3) | 0 | N | Y | Y | 73 |  |  | 10/22/1996 7:18:31 PM | 1996-04-12 | 6/3/1998 6:51:33 PM | EMS |  | AirConcentration | µg/m³ | Y | (air conc.) |  | Sulfide Total (air conc.) | 18496-25-8 | Inorganics; | Inorganics;EMS code: 0125 | Chemical | Total | LAB
0125 | mg/L | Sulfide Total | EQ01 | EQUIS (mg/L) | EQUIS (mg/L) | 0 | N | Y | Y | 3543 |  |  | 10/23/1996 8:42:02 PM | 1996-04-12 | 6/3/1998 6:51:33 PM | EMS |  | Concentration | mg/L | Y | (fl. conc.) |  | Sulfide Total (fl. conc.) | 18496-25-8 | Volatile Organic Compounds (VOCs); | Volatile Organic Compounds (VOCs); EMS code: 0125 | Chemical | Total | LAB
0125 | mg/L | Sulfide Total | X031 | Iodometric Titration | The sample is acidified and sparged with an inert gas. Hydrogen sulfide, thus liberated, is trapped in a zinc acetate solution. Excess iodine is added and back-titrated with standard sodium thiosulfate solution. Iodine consumed is equivalent to the trapped sulfide. | 0.5 | Y | Y | Y | 741 |  |  | 3/22/2001 8:16:16 AM | 1996-04-12 | 6/3/1998 6:51:33 PM | EMS |  | Concentration | mg/L | Y | (fl. conc.) |  | Sulfide Total (fl. conc.) | 18496-25-8 | Volatile Organic Compounds (VOCs); | Volatile Organic Compounds (VOCs); EMS code: 0125 | Chemical | Total | LAB
0125 | mg/L | Sulfide Total | F022 | SAOB,Selective Ion Electrode | Sulfide ion is measured potentiometrically using a sulfide ion-selective electrode in conjuction with a double-junction, sleeve type reference electrode. Potentials are read using a pH metre having an expanded millivolt scale capable of being read to the nearest 0.1 mV. Samples are treated prior to analysis with sulfide anti-oxidant buffer (SAOB). This buffer fixes the solution pH at a highly alkaline level, and contains ascorbic acid to retard air oxidation of sulfide ion in solution. This ensures that sulfide present occurs chiefly as S= ion rather than as complexed HS- or H2S which are present at lower pH values. The sulfide ion concentration is measured indirectly by observing the silver ion activity in the sample. (Because of the extreme insolubility of silver sulfide, if either ion is present in a solution the other is absent.) A minute quantity of silver ions is released to the sample from the solid silver sulfide membrane of the electrode. These ions set up a potential with the silver ions in the internal reference solution. The silver ion activity in the sample is dependent on the sulfide ion activity in the sample. The relationships of the two ion activities is: E = Eb - 2.3* RT/nF* log As (where E=measured total potential of the system; Eb=the portion of the total potential due to the reference electrode and its filling solution; As=the activity of the sulfide ion in the sample; 2.3* RT/nF=Nernst factor) The potential (E) for each standard is plotted linearly against concentration on a log scale to produce a standard curve. | 0.04 | Y | Y | Y | 47 | 12/3/1996 9:29:00 AM | 12/27/2001 1:44:00 PM | 8/28/2020 12:46:44 PM | 1996-09-17 | 6/3/1998 6:51:45 PM | MPAK |  | Concentration | mg/L | Y | (fl. conc.) |  | Sulfide Total (fl. conc.) | 18496-25-8 | Volatile Organic Compounds (VOCs); | Volatile Organic Compounds (VOCs); EMS code: 0125 | Chemical | Total | LAB
0125 | % | Sulfide Total | SIR06A | ALS Minerals: S-IR06a | ALS Minerals: S-IR06a. Solid sample is boiled in dilute hydrochloric acid. The resulting solution is filtered and dried before being placed in a high temperature combustion analyzer for analysis. Results reported as the percentage sulfide sulfur. | 0.01 | Y | Y | Y | 0 |  |  |  | 2024-01-29 |  | EMS |  | DimensionlessRatio | % | N | (ratio) |  | Sulfide Total (ratio) | 18496-25-8 | Volatile Organic Compounds (VOCs); | Volatile Organic Compounds (VOCs); EMS code: 0125 | Chemical | Total | LAB
0125 | mg/L | Sulfide Total | 1410 | Titr.Iodine | Titr.Iodine | 0.5 | Y | Y | Y | 416 |  |  | 10/24/1996 1:26:31 PM | 1996-04-12 | 6/3/1998 6:51:45 PM | EMS |  | Concentration | mg/L | Y | (fl. conc.) |  | Sulfide Total (fl. conc.) | 18496-25-8 | Volatile Organic Compounds (VOCs); | Volatile Organic Compounds (VOCs); EMS code: 0125 | Chemical | Total | LAB
0125 | mg/L | Sulfide Total | X257 | Colorimetric | Colorimetric. General use for several parameters. E.g.Total Sulfide is determined by spectrophotometer using the methlyene blue colourimetric method. The H2S calculation converts Total Sulphide as (S2-) and reports it as Total Sulphide as (H2S). | 0.02 | Y | Y | Y | 9974 | 5/2/2002 | 6/30/2024 4:21:00 PM | 7/30/2024 12:35:04 PM | 1996-04-12 | 6/3/1998 6:51:33 PM | EMS |  | Concentration | mg/L | Y | (fl. conc.) |  | Sulfide Total (fl. conc.) | 18496-25-8 | Volatile Organic Compounds (VOCs); | Volatile Organic Compounds (VOCs); EMS code: 0125 | Chemical | Total | LAB
0125 | ug/g | Sulfide Total | FS84 | SAOB, Selective Ion Electrode | Sulfide ion is measured potentiometrically using a sulfide ion-selective electrode in conjuction with a double-junction, sleeve type reference electrode. Potentials are read using a pH metre having an expanded millivolt scale capable of being read to the nearest 0.1 mV. Samples are treated prior to analysis with sulfide anti-oxidant buffer (SAOB). This buffer fixes the solution pH at a highly alkaline level, and contains ascorbic acid to retard air oxidation of sulfide ion in solution. This ensures that sulfide present occurs chiefly as S= ion rather than as complexed HS- or H2S which are present at lower pH values. The sulfide ion concentration is measured indirectly by observing the silver ion activity in the sample. (Because of the extreme insolubility of silver sulfide, if either ion is present in a solution the other is absent.) A minute quantity of silver ions is released to the sample from the solid silver sulfide membrane of the electrode. These ions set up a potential with the silver ions in the internal reference solution. The silver ion activity in the sample is dependent on the sulfide ion activity in the sample. The relationships of the two ion activities is: E = Eb - 2.3* RT/nF* log As (where E=measured total potential of the system; Eb=the portion of the total potential due to the reference electrode and its filling solution; As=the activity of the sulfide ion in the sample; 2.3* RT/nF=Nernst factor) The potential (E) for each standard is plotted linearly against concentration on a log scale to produce a standard curve. | 0.04 | Y | Y | Y | 160 | 8/22/2000 11:26:00 AM | 9/8/2021 | 9/12/2021 3:13:22 PM | 2001-02-23 |  | NPEPPIN |  | MassPerMass | mg/kg | Y | (mass conc.) |  | Sulfide Total (mass conc.) | 18496-25-8 | Volatile Organic Compounds (VOCs); | Volatile Organic Compounds (VOCs); EMS code: 0125 | Chemical | Total | LAB
0125 | mg/L | Sulfide Total | 1075 | Colorimetric:Methyl.Blue | Colorimetric:Methyl.Blue | 0.5 | Y | Y | Y | 2959 | 7/25/2014 2:14:00 PM | 8/15/2024 | 8/20/2024 1:06:09 PM | 1996-04-12 | 6/3/1998 6:51:33 PM | EMS |  | Concentration | mg/L | Y | (fl. conc.) |  | Sulfide Total (fl. conc.) | 18496-25-8 | Volatile Organic Compounds (VOCs); | Volatile Organic Compounds (VOCs); EMS code: 0125 | Chemical | Total | LAB
0126 | % | Total Inorganic Carbon as CaCO3 | CSSS20 | CSSS (2008) 20.2 | See Canadian Society of Soil Science (2008) section 20.2 | 0.4 | Y | Y | Y | 31 | 1/13/2022 | 2/20/2024 | 2/26/2024 6:09:04 PM | 2022-06-28 |  | EMS |  | DimensionlessRatio | % | N | (ratio) |  | Total Inorganic Carbon as CaCO3 (ratio) |  | Carbon; | Carbon; EMS code: 0126 | Chemical | Total | LAB
0128 | mg/L | Resin Acids | X239 | Gas Chromatography | Gas Chromatography | 0.01 | Y | Y | Y | 658 |  |  | 6/24/2008 3:52:28 PM | 1996-04-12 | 6/3/1998 6:51:33 PM | EMS |  | Concentration | mg/L | Y | (fl. conc.) |  | Resin Acids (fl. conc.) |  | Organics; | Organics; EMS code: 0128 | Chemical |  | LAB
0128 | mg/L | Resin Acids | X257 | Colorimetric | Colorimetric. General use for several parameters. E.g.Total Sulfide is determined by spectrophotometer using the methlyene blue colourimetric method. The H2S calculation converts Total Sulphide as (S2-) and reports it as Total Sulphide as (H2S). | 1 | Y | Y | Y | 434 | 9/17/1996 | 3/7/2000 | 6/7/2024 5:30:15 PM | 1996-04-12 | 6/3/1998 6:51:33 PM | EMS |  | Concentration | mg/L | Y | (fl. conc.) |  | Resin Acids (fl. conc.) |  | Organics; | Organics; EMS code: 0128 | Chemical |  | LAB
0128 | mg/L | Resin Acids | X380 | DCM extr,diazo deriv,Florisil,GC/FID,sum | The water sample is acidified and extracted with dichloro-methane. The raw extract is concentrated and drivatized with diazomethane to produce the corresponding methyl ester derivatives. (Suitable substitute derivatives may be produced by other available techniques.) If required, the extracts are cleaned up by Florisil column chromatography. The derivatives are analyzed by gas chromatography with flame ionization detection. Resin acid components summed. | 0.001 | Y | Y | Y | 2745 | 7/19/1996 | 6/4/2024 | 7/29/2024 3:17:42 PM | 1996-04-12 | 6/3/1998 6:51:33 PM | EMS |  | Concentration | mg/L | Y | (fl. conc.) |  | Resin Acids (fl. conc.) |  | Organics; | Organics; EMS code: 0128 | Chemical |  | LAB
0130 | ppt | Salinity | EQ21 | EQUIS (ppT) | EQUIS (ppT) | 0 | N | Y | Y | 4104 |  |  | 10/22/1996 11:25:18 PM | 1996-04-12 | 6/3/1998 6:51:45 PM | EMS |  | DimensionlessRatio | ppt | N | (ratio) | (salinity) | Salinity (salinity) (ratio) |  | Major Ions; | Major Ions; EMS code: 0130 | Chemical |  | LAB
0130 | ppt | Salinity | FLD | Field sampling | Field sampling method | 0.01 | Y | Y | Y | 1 | 9/11/2023 | 9/11/2023 | 9/19/2023 2:38:05 PM | 2023-09-12 |  | EMS |  | DimensionlessRatio | ppt | N | (ratio) | (salinity) | Salinity (salinity) (ratio) |  | Major Ions; | Major Ions; EMS code: 0130 | Chemical |  | FIELD_RESULT | Field sampling
0130 | ppt | Salinity | F011 | Refractometer | A portion of the sample is placed on the window of the refractometer and the salinity is determined by reading a scale from one end of the refractometer, similar to viewing through a telescope. | 1 | Y | Y | Y | 9 | 8/30/2001 10:45:00 AM | 8/30/2001 10:46:00 AM | 3/12/2003 1:35:07 PM | 1996-06-19 | 6/3/1998 6:51:44 PM | MPAK |  | DimensionlessRatio | ppt | N | (ratio) | (salinity) | Salinity (salinity) (ratio) |  | Major Ions; | Major Ions; EMS code: 0130 | Chemical |  | LAB
0130 | ppt | Salinity | 1130 | Salinometer | The electrical conductance is measured using a conductivity bridge which has been calibrated with KCl solutions of known concentration. Salinity is determined by reference to the Practical Salinity Scale, 1978. | 0.5 | Y | Y | Y | 2592 | 6/5/2002 | 8/16/2024 1:33:00 PM | 8/22/2024 3:35:14 PM | 1996-04-12 | 6/3/1998 6:51:46 PM | EMS |  | DimensionlessRatio | ppt | N | (ratio) | (salinity) | Salinity (salinity) (ratio) |  | Major Ions; | Major Ions; EMS code: 0130 | Chemical |  | LAB
0130 | ppt | Salinity | XH00 | Hydrolab | Hydrolab | 0.01 | Y | Y | Y | 911 |  |  | 8/28/2019 11:56:29 AM | 2011-05-11 | 5/11/2011 2:12:34 PM | RDGIBSON | RDGIBSON | DimensionlessRatio | ppt | N | (ratio) | (salinity) | Salinity (salinity) (ratio) |  | Major Ions; | Major Ions; EMS code: 0130 | Chemical |  | FIELD_RESULT | Hydrolab
0130 | None | Salinity | SBCM | Salinity by Conductivity Method | Salinity in Water PURPOSE/PRINCIPLE OF METHOD: There are two definitions of the term salinity, absolute salinity and practical salinity. Absolute salinity is defined as the ratio of the mass of dissolved material in seawater to the mass of the seawater. This salinity cannot be measured directly and thus the practical salinity value has been adopted as the standard for reporting salinity values. Practical salinity is defined as the ratio of the electrical conductivity of a sample to that of a standard potassium chloride (KCl) solution at the same temperature and pressure. The salinity dependence of the conductivity ratio as a function of temperature of a given sample to a standard S=35 seawater sample is used to determine the salinity. In water, ions such as calcium, magnesium, sodium bicarbonate, nitrate, phosphate, iron, and sulphate all contribute to the salinity. Conductivity accounts for all ions in solution and provides no specific information regarding the type and concentration of any particular ion. Instead, a reading proportional to the combined effect of all ions is reported. SCOPE: Salinity is an important unitless property of industrial and natural waters. Because of its high sensitivity and ease of measurement, the conductivity method is most commonly used to determine salinity. DETECTION LIMITS: RDL =0.01 REFERENCE: Standard Methods for the Examination of Water and Wastewater, 22nd Edition, American Public Health Association, Washington, DC, 2011, Method 2520 B Salinity, Electrical Conductivity Method. | 0.01 | Y | Y | Y | 326 | 7/27/2012 6:40:00 PM | 8/2/2024 | 8/9/2024 9:09:08 AM | 2012-06-14 |  | RDGIBSON |  | None | None | N | (unitless) |  | Salinity (unitless) |  | Major Ions; | Major Ions; EMS code: 0130 | Chemical |  | LAB
0131 | mg/L | Acidity pH 8.3 (as CaCO3) | X062 | Manual Potentiometric Titration | The pH of the sample is determined and a measured amount of standard acid is added, as needed, to lower the pH to 4 or less. Hydrogen peroxide is added, the solution boiled for several minutes, cooled and titrated to pH 8.3 with standard base. The method measures the mineral acidity of the sample plus acidity from oxidation and hydrolysis of polyvalent cations, including salts of iron and aluminum. pH 8.3 end point determined by manual potentiometric titration. | 0.5 | N | Y | Y | 45 |  |  | 10/24/1996 1:51:39 PM | 1996-04-12 | 6/3/1998 6:51:33 PM | EMS |  | Concentration | mg/L | Y | (fl. conc.) |  | Acidity pH 8.3 (as CaCO3) (fl. conc.) |  | Major Ions; | Major Ions; EMS code: 0131 | Chemical |  | LAB
0131 | mg/L | Acidity pH 8.3 (as CaCO3) | F044 | Man. Potentiometric Titr'n | Manual potentiometric titration with solution of standard acid or of standard base to specified endpoints. For acidity determinations, the pH of the sample is determined and a measured amount of standard acid is added, as needed, to lower the pH to 4 or less. Hydrogen peroxide is added, the solution boiled for several minutes, cooled and titrated to pH 8.3 with standard base. The method measures the mineral acidity of the sample plus acidity from oxidation and hydrolysis of polyvalent cations, including salts of iron and aluminum. Because the color change of phenolphthalein indicator is close to pH 8.3, this value generally is accepted as a standard end point for titration of total acidity, including CO2 and most weak acids. | 0.5 | Y | Y | Y | 747 | 8/27/2001 | 5/13/2014 11:07:00 AM | 6/9/2014 3:03:08 PM | 1997-07-21 | 6/3/1998 6:51:47 PM | NPEPPIN |  | Concentration | mg/L | Y | (fl. conc.) |  | Acidity pH 8.3 (as CaCO3) (fl. conc.) |  | Major Ions; | Major Ions; EMS code: 0131 | Chemical |  | LAB
0131 | mg/L | Acidity pH 8.3 (as CaCO3) | EQ01 | EQUIS (mg/L) | EQUIS (mg/L) | 0.5 | N | Y | Y | 508 |  |  | 10/23/1996 12:04:06 AM | 1996-04-12 | 6/3/1998 6:51:33 PM | EMS |  | Concentration | mg/L | Y | (fl. conc.) |  | Acidity pH 8.3 (as CaCO3) (fl. conc.) |  | Major Ions; | Major Ions; EMS code: 0131 | Chemical |  | LAB
0131 | mg/L | Acidity pH 8.3 (as CaCO3) | X041 | Phenolphthalein Titr'n | The pH of the sample is determined and a measured amount of standard acid is added, as needed, to lower the pH to 4 or less. Hydrogen peroxide is added, the solution boiled for several minutes, cooled and titrated to pH 8.3 with standard base. The method measures the mineral acidity of the sample plus acidity from oxidation and hydrolysis of polyvalent cations, including salts of iron and aluminum. Because the color change of phenolphthalein indicator is close to pH 8.3, this value generally is accepted as a standard end point for titration of total acidity, including CO2 and most weak acids. | 0.5 | Y | Y | Y | 4631 | 4/27/2007 | 2/22/2020 | 6/28/2024 8:04:10 PM | 1996-04-12 | 6/3/1998 6:51:33 PM | EMS |  | Concentration | mg/L | Y | (fl. conc.) |  | Acidity pH 8.3 (as CaCO3) (fl. conc.) |  | Major Ions; | Major Ions; EMS code: 0131 | Chemical |  | LAB
0131 | mg/L | Acidity pH 8.3 (as CaCO3) | X310 | Auto Potentiometric Titration | For alkalinity, the samples are (if pH >8.3) first titrated to a pH of 8.3, which corresponds to the equivalence point for the conversion of carbonate to bicarbonate ion (phenolphthalein alkalinity). The titration is continued to a pH of 4.5, which corresponds approx. to the equivalence point for the conversion of bicarbonate ion to carbonic acid (entire titration indicates total alkalinity). An extrapolation technique based on the near proportionality of concentration of H ions to excess of titrant beyond the equivalence pt. The amount of standard acid required to reduce pH exactly 0.30pH (ie. pH4.2) unit is measured carefully. Because this change in pH corresponds to an exact coupling of the H ion concentration, a simple extrapolation can be made to the equivalence point. The actual pH of the stoichiometric end point in alkalinity determinations is the pH at which the inflection in the titration curve occurs and is directly related to the amount of carbonate alkalinity originally present in the sample. A curve can be constructed by plotting pH vs mL of acid to find this point directly. *** For acidity, the samples are titrated to a pH of 8.3, which corresponds to stoichiometric neutralization of carbonic acid to bicarbonate. Because the colour change of phenolphthalein indicator is close to pH 8.3, this value generally is accepted as a standard end point for titration of total acidity, including CO2 and most weak acids an equivalence point for the conversion of carbonate to bicarbonate ion (phenolphthalein alkalinity). The actual pH of the stoichiometric end point in acidity determinations is the pH at which the inflection in the titration curve occurs and is directly related to the amount of titratable acids originally present in the sample. A curve can be constructed by plotting pH vs. mL of NaOH to find this point directly. (To obtain the actual substances contributing to acidity/alkalinity, a pH vs mL of acid/NaOH curve should be constructed.) | 1 | Y | Y | Y | 17505 | 3/4/1996 | 8/21/2024 10:03:00 AM | 8/26/2024 4:53:05 PM | 1996-04-12 | 6/3/1998 6:51:41 PM | EMS |  | Concentration | mg/L | Y | (fl. conc.) |  | Acidity pH 8.3 (as CaCO3) (fl. conc.) |  | Major Ions; | Major Ions; EMS code: 0131 | Chemical |  | LAB
0132 | mg/L | Acidity pH 4.5 (as CaCO3) | X310 | Auto Potentiometric Titration | For alkalinity, the samples are (if pH >8.3) first titrated to a pH of 8.3, which corresponds to the equivalence point for the conversion of carbonate to bicarbonate ion (phenolphthalein alkalinity). The titration is continued to a pH of 4.5, which corresponds approx. to the equivalence point for the conversion of bicarbonate ion to carbonic acid (entire titration indicates total alkalinity). An extrapolation technique based on the near proportionality of concentration of H ions to excess of titrant beyond the equivalence pt. The amount of standard acid required to reduce pH exactly 0.30pH (ie. pH4.2) unit is measured carefully. Because this change in pH corresponds to an exact coupling of the H ion concentration, a simple extrapolation can be made to the equivalence point. The actual pH of the stoichiometric end point in alkalinity determinations is the pH at which the inflection in the titration curve occurs and is directly related to the amount of carbonate alkalinity originally present in the sample. A curve can be constructed by plotting pH vs mL of acid to find this point directly. *** For acidity, the samples are titrated to a pH of 8.3, which corresponds to stoichiometric neutralization of carbonic acid to bicarbonate. Because the colour change of phenolphthalein indicator is close to pH 8.3, this value generally is accepted as a standard end point for titration of total acidity, including CO2 and most weak acids an equivalence point for the conversion of carbonate to bicarbonate ion (phenolphthalein alkalinity). The actual pH of the stoichiometric end point in acidity determinations is the pH at which the inflection in the titration curve occurs and is directly related to the amount of titratable acids originally present in the sample. A curve can be constructed by plotting pH vs. mL of NaOH to find this point directly. (To obtain the actual substances contributing to acidity/alkalinity, a pH vs mL of acid/NaOH curve should be constructed.) | 1 | Y | Y | Y | 5085 | 1/21/1997 10:32:00 AM | 8/12/2024 9:19:00 PM | 8/15/2024 5:08:04 PM | 1996-04-12 | 6/3/1998 6:51:41 PM | EMS |  | Concentration | mg/L | Y | (fl. conc.) |  | Acidity pH 4.5 (as CaCO3) (fl. conc.) |  | Major Ions; | Major Ions; EMS code: 0132 | Chemical |  | LAB
0132 | mg/L | Acidity pH 4.5 (as CaCO3) | 1180 | Manual Potentiometric Titration | The acidity (pH 4.5) is determined by manual potentiometric titration of a sample aliquot with a standard solution of alkali to pH 4.5. | 0.5 | N | Y | Y | 286 |  |  | 10/24/1996 2:41:56 PM | 1996-04-12 | 6/3/1998 6:51:45 PM | EMS |  | Concentration | mg/L | Y | (fl. conc.) |  | Acidity pH 4.5 (as CaCO3) (fl. conc.) |  | Major Ions; | Major Ions; EMS code: 0132 | Chemical |  | LAB
0132 | mg/L | Acidity pH 4.5 (as CaCO3) | EQ01 | EQUIS (mg/L) | EQUIS (mg/L) | 0.5 | N | Y | Y | 14 |  |  | 10/22/1996 11:15:18 PM | 1996-04-12 | 6/3/1998 6:51:33 PM | EMS |  | Concentration | mg/L | Y | (fl. conc.) |  | Acidity pH 4.5 (as CaCO3) (fl. conc.) |  | Major Ions; | Major Ions; EMS code: 0132 | Chemical |  | LAB
0143 | mg/L | Chlorophyll A | EQ01 | EQUIS (mg/L) | EQUIS (mg/L) | 0.0005 | N | Y | Y | 3711 |  |  | 10/23/1996 12:04:08 AM | 1996-04-12 | 6/3/1998 6:51:33 PM | EMS |  | Concentration | mg/L | Y | (fl. conc.) |  | Chlorophyll A (fl. conc.) | 1406-65-1 | Biological Examination; | Biological Examination; EMS code: 0143 | Chemical |  | LAB
0143 | mg/L | Chlorophyll A | 1800 | Ext:SU:Dup:Vol:N.Acid. | This method describes the spectrophotometric analysis of chlorophyll a derived from material retained on a 0.45 um filter. The filter is ground and extracted in 90% acetone. Absorbances of the extract are read at 750, 664, 646 and 630 nm. The abbreviation "SU" stands for Scor-Unesco method of calculating chlorophyll. The method is the "Trichromatic Method" of determining chlorophyll, whereby, samples are not acidified and corrected for the presence of phaeophytin, but determines chlorophyll a, b, & c respectively by inserting the corrected optical densities in 3 separate equations. The "N. Acid" used to stand for extract "Not-acidified" when determining chlorophyll a. This method does not correct for the presence of phaeophytin. Acidification is used only when determining Phaeophytin. "Acid" means extract is acidified in order to determine chlorophyll a in the presence of phaeophytin, and the method is called the Lorenzon method. The optical densities read and equations used to calculate chlorophyll a are slightly different. This method corrects for the presence of phaeophytin. The terms SU and SP are synonymous, the method calculation is done by the "Scor-Unesco" (SU) written by Strickland & Parsons (SP). | 0.0005 | Y | Y | Y | 7024 | 9/18/1996 4:58:00 PM | 10/22/1997 3:45:00 PM | 5/28/2024 11:13:59 AM | 1996-04-12 | 12/9/1998 9:46:25 AM | EMS | NPEPPIN | Concentration | mg/L | Y | (fl. conc.) |  | Chlorophyll A (fl. conc.) | 1406-65-1 | Biological Examination; | Biological Examination; EMS code: 0143 | Chemical |  | LAB
0143 | g/m2 | Chlorophyll A | 1803 | Extr,SU,4 Filt,Substrate, N. Acid. | Spectrophotometric analysis of chlorophyll a derived from material scraped from substrate and retained on a 0.45 um filter. The filter is ground and extracted into 90% acetone. The absorbances of the extract are read at 750, 664, 646 and 630 nm. "N. Acid." (= not acidified) methods are not corrected for phaephytin; "Acid." (=acidified) methods are corrected for phaeophytin. | 0.0003 | Y | Y | Y | 76 |  |  | 10/24/1996 12:35:27 PM | 1996-04-12 | 12/9/1998 9:46:25 AM | EMS | NPEPPIN | MassPerArea | g/m² | Y | (mass/area) |  | Chlorophyll A (mass/area) | 1406-65-1 | Biological Examination; | Biological Examination; EMS code: 0143 | Chemical |  | LAB
0143 | mg/L | Chlorophyll A | X454 | Fluorimetric: Extraction | Fluorimetric: Extraction (was Colorimetric before Oct 2020) | 0.0005 | Y | Y | Y | 6904 | 6/12/2009 2:27:00 PM | 8/22/2024 | 8/26/2024 1:03:06 PM | 1996-04-12 | 12/9/1998 9:46:25 AM | EMS | NPEPPIN | Concentration | mg/L | Y | (fl. conc.) |  | Chlorophyll A (fl. conc.) | 1406-65-1 | Biological Examination; | Biological Examination; EMS code: 0143 | Chemical |  | LAB
0143 | mg/L | Chlorophyll A | XH00 | Hydrolab | Hydrolab | 3e-05 | Y | Y | Y | 325 | 2/26/2008 10:00:00 AM | 4/8/2008 9:40:00 AM | 6/6/2016 3:05:16 PM | 2011-08-04 | 8/4/2011 3:56:29 PM | RDGIBSON | RDGIBSON | Concentration | mg/L | Y | (fl. conc.) |  | Chlorophyll A (fl. conc.) | 1406-65-1 | Biological Examination; | Biological Examination; EMS code: 0143 | Chemical |  | FIELD_RESULT | Hydrolab
0143 | mg/L | Chlorophyll A | 1810 | Ext:SU:Sin:Vol:N.Acid | This method describes the spectrophotometric analysis of chlorophyll a derived from material retained on a 0.45 um filter. The filter is ground and extracted in 90% acetone. Absorbances of the extract are read at 750, 664, 646 and 630 nm. The abbreviation "SU" stands for Scor-Unesco method of calculating chlorophyll. The method is the Trichromatic Method of determining chlorophyll, whereby, samples are not acidified and corrected for the presence of phaeophytin, but determines chlorophyll a, b, & c respectively by inserting the corrected optical densities in 3 separate equations. The N. Acid used to stand for extract Not-acidified when determining chlorophyll a. This method does not correct for the presence of phaeophytin. Acidification is used only when determining Phaeophytin. "Acid" means extract is acidified in order to determine chlorophyll a in the presence of phaeophytin, and the method is called the Lorenzon method. The optical densities read and equations used to calculate chlorophyll a are slightly different. This method corrects for the presence of phaeophytin. The terms SU and SP are synonymous, the method calculation is done by the Scor-Unesco (SU) written by Strickland & Parsons (SP). | 0.0005 | Y | Y | Y | 10410 | 9/17/1996 1:56:00 PM | 8/15/2024 11:19:00 AM | 8/17/2024 4:36:03 PM | 1996-04-12 | 12/9/1998 9:46:25 AM | EMS | NPEPPIN | Concentration | mg/L | Y | (fl. conc.) |  | Chlorophyll A (fl. conc.) | 1406-65-1 | Biological Examination; | Biological Examination; EMS code: 0143 | Chemical |  | LAB
0143 | g/m2 | Chlorophyll A | 2630 | Sp.Sub.Styr.SU Sin Plug | Water samples are filtered through glass fibre filters and then frozen. Foam plugs are usually cut from a block with a cork borer of known diameter. The cells are disrupted by ultrasonification and the Chlorophyll is extracted with an aqueous magnesium Carbonate:Acetone mixture. By measuring the maximum Absorbance at about 663 nm before acidification and at about 665 nm after acidification the amount of Chlorophyll and Phaeophytin present in the sample can be calculated. The addition of acid to a Chlorophyll solution removes the bonded Magnesium of the Porphyrin ring thus converting it to Phaeophytin. This method uses a single plug. | 0.002 | Y | Y | Y | 1727 | 8/13/1997 1:08:00 PM | 7/30/2011 11:10:00 AM | 12/14/2015 10:27:35 AM | 1996-04-12 | 6/3/1998 6:51:33 PM | EMS |  | MassPerArea | g/m² | Y | (mass/area) |  | Chlorophyll A (mass/area) | 1406-65-1 | Biological Examination; | Biological Examination; EMS code: 0143 | Chemical |  | LAB
0143 | g/m2 | Chlorophyll A | 2631 | Sp.Sub.Styr.,SU Dup Plug | Water samples are filtered through glass fibre filters and then frozen. Foam plugs are usually cut from a block with a cork borer of known diameter. The cells are disrupted by ultrasonification and the Chlorophyll is extracted with an aqueous Magnesium Carbonate: Acetone mixture. By measuring the maximum Absorbance at about 663 nm before acidification and at about 665 nm after acidification the amount of Chlorophyll and Phaeophytin present in the sample can be calculated. The addition of acid to a Chlorophyll solution removes the bonded Magnesium of the Porphyrin ring thus converting it to Phaeophytin. This method uses a duplicate plug. | 0.002 | N | Y | Y | 171 | 8/13/1997 12:48:00 PM | 12/1/2001 11:31:00 AM | 5/20/2003 4:07:38 PM | 1996-04-12 | 6/24/2003 11:30:40 AM | EMS | UMULLER | MassPerArea | g/m² | Y | (mass/area) |  | Chlorophyll A (mass/area) | 1406-65-1 | Biological Examination; | Biological Examination; EMS code: 0143 | Chemical |  | LAB
0143 | mg/L | Chlorophyll A | 1830 | Ext:Lor:Sin:Sub:Acid. | Water samples are filtered through glass fibre filters and then frozen. Foam plugs are usually cut from a block with a cork borer of known diameter. The cells are disrupted by ultrasonification and the Chlorophyll is extracted with an aqueous Magnesium Caronate:Acetone mixture. Absorbance measured at about 750, 664, 646 and 630 nm before and after acidification with a few drops of HCl. "LOR." indicates Lorenzen calculation. The addition of acid to a Chlorophyll solution removes the bonded Magnesium of the Porphyrin ring thus converting it to Phaeophytin. This method is single filter. | 0.0002 | Y | Y | Y | 2402 | 9/19/1997 11:26:00 AM | 12/20/2001 9:53:00 AM | 5/20/2003 3:21:53 PM | 1996-04-12 | 6/3/1998 6:51:42 PM | EMS |  | Concentration | mg/L | Y | (fl. conc.) |  | Chlorophyll A (fl. conc.) | 1406-65-1 | Biological Examination; | Biological Examination; EMS code: 0143 | Chemical |  | LAB
0143 | g/m2 | Chlorophyll A | X455 | Ext:PS:Sin:Sub:Acid | Ext:PS:Sin:Sub:Acid | 0.0003 | Y | Y | Y | 294 |  |  | 7/11/2005 5:43:02 PM | 1996-04-12 | 12/9/1998 9:46:25 AM | EMS | NPEPPIN | MassPerArea | g/m² | Y | (mass/area) |  | Chlorophyll A (mass/area) | 1406-65-1 | Biological Examination; | Biological Examination; EMS code: 0143 | Chemical |  | LAB
0143 | g/m2 | Chlorophyll A | 1801 | Ext:SU:Dup:Subs:N.Acid.. | Spectrophotometric analysis of chlorophyll a derived from material scraped from substrate and retained on a 0.45 um filter. The filter is ground and extracted into 90% acetone. The absorbances of the extract are read at 750, 664, 646 and 630 nm. "N. Acid." (= not acidified) methods are not corrected for phaephytin; "Acid." (=acidified) methods are corrected for phaeophytin. | 0.0003 | Y | Y | Y | 289 |  |  | 10/30/2017 12:11:48 PM | 1996-04-12 | 12/9/1998 9:46:25 AM | EMS | NPEPPIN | MassPerArea | g/m² | Y | (mass/area) |  | Chlorophyll A (mass/area) | 1406-65-1 | Biological Examination; | Biological Examination; EMS code: 0143 | Chemical |  | LAB
0143 | g/m2 | Chlorophyll A | X445 | Chl-a Biota: Fluorometer | This analysis is done using procedures adapted from EPA Method 445.0. Chlorophyll-a is determined by a routine acetone extraction followed with analysis by fluorometry using the non-acidification procedure. This method is not subject to interferences from chlorophyll b. Note: Biota samples are typically submitted as scrapings on a filter. | 0.11 | Y | Y | Y | 50 | 8/19/2016 | 3/23/2020 | 3/26/2020 4:01:10 PM | 2016-10-05 |  | RRENTON |  | MassPerArea | g/m² | Y | (mass/area) |  | Chlorophyll A (mass/area) | 1406-65-1 | Biological Examination; | Biological Examination; EMS code: 0143 | Chemical |  | LAB
0143 | g/m2 | Chlorophyll A | 1811 | Ext:SU:Sin:Sub:N.Acid. | Spectrophotometric analysis of chlorophyll a derived from material scraped from substrate and retained on a 0.45 um filter. The filter is ground and extracted into 90% acetone. The absorbances of the extract are read at 750, 664, 646, and 630 nm. "N.Acid." (= not acidified) methods are not corrected for phaephytin; "Acid." (=acidified) methods are corrected for phaeophytin. | 0.0003 | Y | Y | Y | 3562 | 3/28/2002 3:55:00 PM | 10/1/2018 6:23:00 PM | 10/26/2020 2:41:03 PM | 1996-04-12 | 12/9/1998 9:46:25 AM | EMS | NPEPPIN | MassPerArea | g/m² | Y | (mass/area) |  | Chlorophyll A (mass/area) | 1406-65-1 | Biological Examination; | Biological Examination; EMS code: 0143 | Chemical |  | LAB
0143 | g/m2 | Chlorophyll A | 1804 | Extr,SU,6 Filt,Substrate, N. Acid. | Spectrophotometric analysis of chlorophyll a derived from material scraped from substrate and retained on a 0.45 um filter. The filter is ground and extracted into 90% acetone. The absorbances of the extract are read at 750, 664, 646 and 630 nm. "N. Acid." (= not acidified) methods are not corrected for phaephytin; "Acid." (=acidified) methods are corrected for phaeophytin. | 0.0003 | Y | Y | Y | 43 |  |  | 10/24/1996 11:18:40 AM | 1996-04-12 | 12/9/1998 9:46:25 AM | EMS | NPEPPIN | MassPerArea | g/m² | Y | (mass/area) |  | Chlorophyll A (mass/area) | 1406-65-1 | Biological Examination; | Biological Examination; EMS code: 0143 | Chemical |  | LAB
0143 | g/m2 | Chlorophyll A | TCAL | Calculation | Calculation |  | Y | Y | Y | 56 | 10/12/2023 | 8/21/2024 | 8/21/2024 11:02:08 PM | 2022-07-13 |  | EMS |  | MassPerArea | g/m² | Y | (mass/area) |  | Chlorophyll A (mass/area) | 1406-65-1 | Biological Examination; | Biological Examination; EMS code: 0143 | Chemical |  | LAB
0143 | mg | Chlorophyll A | XM40 | SM 10200 H (2017) Spectrophotometry | SM 10200 H (2017) Spectrophotometry | 0.0001 | Y | Y | Y | 3 | 11/1/2022 11:32:00 AM | 11/1/2022 11:32:00 AM | 11/10/2022 5:26:47 PM | 2022-11-08 |  | EMS |  | Mass | mg | Y | (mass) |  | Chlorophyll A (mass) | 1406-65-1 | Biological Examination; | Biological Examination; EMS code: 0143 | Chemical |  | LAB
0143 | mg | Chlorophyll A | E870 | EPA 445.0 (mod) | EPA 445.0 (mod) Chlorophyll a is determined by solvent extraction followed with analysis by fluorometry using the non-acidification procedure. This method is not subject to interferences from chlorophyll b. | 1e-06 | Y | Y | Y | 44 | 6/4/2024 | 8/20/2024 | 8/21/2024 11:02:08 PM | 2024-06-10 | 6/11/2024 8:38:45 AM | RDGIBSON | RDGIBSON | Mass | mg | Y | (mass) |  | Chlorophyll A (mass) | 1406-65-1 | Biological Examination; | Biological Examination; EMS code: 0143 | Chemical |  | LAB
0143 | mg/L | Chlorophyll A | EXOSO | EXO Sonde YSI | EXO Sonde YSI |  | Y | Y | Y | 155 |  |  | 7/15/2024 3:01:47 PM | 2023-06-16 |  | EMS |  | Concentration | mg/L | Y | (fl. conc.) |  | Chlorophyll A (fl. conc.) | 1406-65-1 | Biological Examination; | Biological Examination; EMS code: 0143 | Chemical |  | FIELD_RESULT | EXO Sonde YSI
0143 | g/m2 | Chlorophyll A | 1831 | Ext:Lor:Sin:Sub:Acid. | Spectrophotometric analysis of chlorophyll a derived from material scraped from substrate and retained on a 0.45 um filter. The filter is ground and extracted into 90% acetone. The absorbances of the extract are read at 750, 664, 646, and 630 nm. "N.Acid." (= not acidified) methods are not corrected for phaephytin; "Acid." (=acidified) methods are corrected for phaeophytin. | 0.0003 | Y | Y | Y | 567 | 9/23/1997 4:47:00 PM | 12/28/2001 4:00:00 PM | 1/17/2002 1:07:10 AM | 1996-04-12 | 12/9/1998 9:46:25 AM | EMS | NPEPPIN | MassPerArea | g/m² | Y | (mass/area) |  | Chlorophyll A (mass/area) | 1406-65-1 | Biological Examination; | Biological Examination; EMS code: 0143 | Chemical |  | LAB
0143 | g/m2 | Chlorophyll A | 1805 | Extr,SU,8 Filt,Substrate, N. Acid. | Spectrophotometric analysis of chlorophyll a derived from material scraped from substrate and retained on a 0.45 um filter. The filter is ground and extracted into 90% acetone. The absorbances of the extract are read at 750, 664, 646 and 630 nm. "N. Acid." (= not acidified) methods are not corrected for phaephytin; "Acid." (=acidified) methods are corrected for phaeophytin. | 0.0003 | Y | Y | Y | 18 |  |  | 10/24/1996 11:18:40 AM | 1996-04-12 | 12/9/1998 9:46:25 AM | EMS | NPEPPIN | MassPerArea | g/m² | Y | (mass/area) |  | Chlorophyll A (mass/area) | 1406-65-1 | Biological Examination; | Biological Examination; EMS code: 0143 | Chemical |  | LAB
0143 | g/m2 | Chlorophyll A | 1802 | Extr,SU, 3 Filters,Subst, N. Acid. | Spectrophotometric analysis of chlorophyll a derived from material scraped from substrate and retained on a 0.45 um filter. The filter is ground and extracted into 90% acetone. The absorbances of the extract are read at 750, 664, 646 and 630 nm. "N. Acid." (= not acidified) methods are not corrected for phaephytin; "Acid." (=acidified) methods are corrected for phaeophytin. | 0.0003 | Y | Y | Y | 147 |  |  | 10/24/1996 12:35:30 PM | 1996-04-12 | 12/9/1998 9:46:25 AM | EMS | NPEPPIN | MassPerArea | g/m² | Y | (mass/area) |  | Chlorophyll A (mass/area) | 1406-65-1 | Biological Examination; | Biological Examination; EMS code: 0143 | Chemical |  | LAB
0143 | g/m2 | Chlorophyll A | EQ13 | EQUIS (g/m2) | EQUIS (g/m2) | 0.0003 | N | Y | Y | 820 |  |  | 10/22/1996 11:30:23 PM | 1996-04-12 | 6/3/1998 6:51:33 PM | EMS |  | MassPerArea | g/m² | Y | (mass/area) |  | Chlorophyll A (mass/area) | 1406-65-1 | Biological Examination; | Biological Examination; EMS code: 0143 | Chemical |  | LAB
0143 | mg/L | Chlorophyll A | 1820 | Ext:Lor:Dup:Vol:Acid | Water samples are filtered through glass fibre filters and then frozen. Foam plugs are usually cut from a block with a cork borer of known diameter. The cells are disrupted by ultrasonification and the Chlorophyll is extracted with an aqueous Magnesium Caronate:Acetone mixture. Absorbance measured at about 750, 664, 646 and 630 nm before and after acidification with a few drops of HCl. "LOR." indicates Lorenzen calculation. The addition of acid to a Chlorophyll solution removes the bonded Magnesium of the Porphyrin ring thus converting it to Phaeophytin. This method is single filter. | 0.0002 | Y | Y | Y | 86 | 10/16/2000 3:03:00 PM | 12/20/2001 9:53:00 AM | 5/13/2003 2:37:47 PM | 1996-04-12 | 6/3/1998 6:51:33 PM | EMS |  | Concentration | mg/L | Y | (fl. conc.) |  | Chlorophyll A (fl. conc.) | 1406-65-1 | Biological Examination; | Biological Examination; EMS code: 0143 | Chemical |  | LAB
0143 | g/m2 | Chlorophyll A | 1821 | Ext:Lor:Dup:Sub:Acid. | Spectrophotometric analysis of chlorophyll a derived from material scraped from substrate and retained on a 0.45 um filter. The filter is ground and extracted into 90% acetone. The absorbances of the extract are read at 750, 664, 646, and 630 nm. "N.Acid." (= not acidified) methods are not corrected for phaephytin; "Acid." (=acidified) methods are corrected for phaeophytin. | 0.0003 | Y | Y | Y | 2 | 1/30/1997 9:12:00 AM | 1/30/1997 9:12:00 AM | 10/16/1997 6:03:32 PM | 1996-04-12 | 12/9/1998 9:46:25 AM | EMS | NPEPPIN | MassPerArea | g/m² | Y | (mass/area) |  | Chlorophyll A (mass/area) | 1406-65-1 | Biological Examination; | Biological Examination; EMS code: 0143 | Chemical |  | LAB
0144 | RFU | Chlorophyll A (RFU) | EXOSO | EXO Sonde YSI | EXO Sonde YSI |  | Y | Y | Y | 204 |  |  | 10/17/2023 2:53:37 PM | 2023-06-16 |  | EMS |  | Apperance | RFU | N | (app.) |  | Chlorophyll A (RFU) (app.) | 1406-65-1 | Biological Examination; | Biological Examination; EMS code: 0144 | Chemical |  | FIELD_RESULT | EXO Sonde YSI
0146 | mg/L | Phaeophytin A | EQ01 | EQUIS (mg/L) | EQUIS (mg/L) | 0.0005 | N | Y | Y | 1173 |  |  | 10/22/1996 11:30:33 PM | 1996-04-12 | 6/3/1998 6:51:33 PM | EMS |  | Concentration | mg/L | Y | (fl. conc.) |  | Phaeophytin A (fl. conc.) |  | Biological Examination; | Biological Examination; EMS code: 0146 | Chemical |  | LAB
0146 | g/m2 | Phaeophytin A | 2633 | Extr:Acid:SubsStyr-Sin Plug | Water samples are filtered through glass fibre filters and then frozen. Foam plugs are usually cut from a block with a cork borer of known diameter. The cells are disrupted by ultrasonification and the Chlorophyll is extracted with an aqueous Magnesium Carbonate:Acetone mixture. By measuring the maximum Absorbance at about 663 nm before acidification and at about 665 nm after acidification the amount of Chlorophyll and Phaeophytin present in the sample can be calculated. The addition of acid to a Chlorophyll solution removes the bonded Magnesium of the Porphyrin ring thus converting it to Phaeophytin. This method uses a single plug. | 0.002 | N | Y | Y | 36 | 8/13/1997 1:09:00 PM | 12/1/2001 11:42:00 AM | 5/20/2003 4:07:44 PM | 1996-04-12 | 6/24/2003 11:30:59 AM | EMS | UMULLER | MassPerArea | g/m² | Y | (mass/area) |  | Phaeophytin A (mass/area) |  | Biological Examination; | Biological Examination; EMS code: 0146 | Chemical |  | LAB
0146 | g/m2 | Phaeophytin A | 1851 | Ext:SU:Sin:Sub:Acidified | Spectrophotometric analysis of phaeophytin a derived from material retained on a 0.45 um filter. The filter is ground and extracted in 90% acetone, and the absorbances of the extract are read at 750 and 664 nm. The extract is then acidified by adding a few drops of HCl and the acidified absorbances are read at 750 and 664 nm. | 0.0003 | Y | Y | Y | 473 | 10/5/2004 | 10/5/2004 | 11/5/2004 1:44:04 AM | 1996-04-12 | 10/19/1998 2:10:53 PM | EMS | NPEPPIN | MassPerArea | g/m² | Y | (mass/area) |  | Phaeophytin A (mass/area) |  | Biological Examination; | Biological Examination; EMS code: 0146 | Chemical |  | LAB
0146 | mg/L | Phaeophytin A | 1850 | Ext:SU:Sin:Vol:Acidified | Water samples are filtered through glass fibre filters and then frozen. Foam plugs are usually cut from a block with a cork borer of known diameter. The cells are disrupted by ultrasonification and the Chlorophyll is extracted with an aqueous Magnesium Caronate:Acetone mixture. Absorbance measured at about 750, 664, 646 and 630 nm before and after acidification with a few drops of HCl. "LOR." indicates Lorenzen calculation. The addition of acid to a Chlorophyll solution removes the bonded Magnesium of the Porphyrin ring thus converting it to Phaeophytin. This method is single filter. | 0.0005 | Y | Y | Y | 2075 | 5/14/1997 3:47:00 PM | 8/1/2024 4:22:00 PM | 8/2/2024 5:04:09 AM | 1996-04-12 | 12/9/1998 9:43:51 AM | EMS | NPEPPIN | Concentration | mg/L | Y | (fl. conc.) |  | Phaeophytin A (fl. conc.) |  | Biological Examination; | Biological Examination; EMS code: 0146 | Chemical |  | LAB
0146 | g/m2 | Phaeophytin A | EQ13 | EQUIS (g/m2) | EQUIS (g/m2) | 0.0003 | N | Y | Y | 55 |  |  | 10/22/1996 9:26:47 PM | 1996-04-12 | 6/3/1998 6:51:33 PM | EMS |  | MassPerArea | g/m² | Y | (mass/area) |  | Phaeophytin A (mass/area) |  | Biological Examination; | Biological Examination; EMS code: 0146 | Chemical |  | LAB
0146 | g/m2 | Phaeophytin A | 1872 | Extr:Acid:Subst.Sty-Sgl | Extr:Acid:Subst.Sty-Sgl | 0.0003 | Y | Y | Y | 215 | 10/2/2002 | 9/14/2004 | 10/6/2004 11:24:27 AM | 1996-04-12 | 10/19/1998 2:10:53 PM | EMS | NPEPPIN | MassPerArea | g/m² | Y | (mass/area) |  | Phaeophytin A (mass/area) |  | Biological Examination; | Biological Examination; EMS code: 0146 | Chemical |  | LAB
0146 | g/m2 | Phaeophytin A | 1873 | Extr:Lorenz:Sin:Sub:Acid | Spectrophotometric analysis of phaeophytin a derived from material retained on a 0.45 um filter. The filter is ground and extracted in 90% acetone, and the absorbances of the extract are read at 750 and 664 nm. The extract is then acidified by adding a few drops of HCl and the acidified absorbances are read at 750 and 664 nm; Lorenzen calc. | 0.0003 | Y | Y | Y | 82 | 10/9/1997 4:55:00 PM | 10/16/1997 3:53:00 PM | 7/7/1999 7:28:00 PM | 1996-04-12 | 10/19/1998 2:10:53 PM | EMS | NPEPPIN | MassPerArea | g/m² | Y | (mass/area) |  | Phaeophytin A (mass/area) |  | Biological Examination; | Biological Examination; EMS code: 0146 | Chemical |  | LAB
0146 | g/m2 | Phaeophytin A | 2632 | Ext:Acid:Subs.Styr-Dup Plug | Water samples are filtered through glass fibre filters and then frozen. Foam plugs are usually cut from a block with a cork borer of known diameter. The cells are disrupted by ultrasonification and the Chlorophyll is extracted with an aqueous Magnesium Carbonate:Acetone mixture. By measuring the maximum Absorbance at about 663 nm before acidification and at about 665 nm after acidification the amount of Chlorophyll and Phaeophytin present in the sample can be calculated. The addition of acid to a Chlorophyll solution removes the bonded Magnesium of the Porphyrin ring thus converting it to Phaeophytin. This method uses a duplicate plug. | 0.002 | Y | Y | Y | 168 | 8/13/1997 1:02:00 PM | 12/1/2001 11:42:00 AM | 5/20/2003 4:07:41 PM | 1996-04-12 | 6/3/1998 6:51:42 PM | EMS |  | MassPerArea | g/m² | Y | (mass/area) |  | Phaeophytin A (mass/area) |  | Biological Examination; | Biological Examination; EMS code: 0146 | Chemical |  | LAB
0146 | mg/L | Phaeophytin A | 1830 | Ext:Lor:Sin:Sub:Acid. | Water samples are filtered through glass fibre filters and then frozen. Foam plugs are usually cut from a block with a cork borer of known diameter. The cells are disrupted by ultrasonification and the Chlorophyll is extracted with an aqueous Magnesium Caronate:Acetone mixture. Absorbance measured at about 750, 664, 646 and 630 nm before and after acidification with a few drops of HCl. "LOR." indicates Lorenzen calculation. The addition of acid to a Chlorophyll solution removes the bonded Magnesium of the Porphyrin ring thus converting it to Phaeophytin. This method is single filter. | 0.0002 | Y | Y | Y | 224 | 5/7/1998 12:27:00 PM | 12/19/2001 11:10:00 AM | 5/20/2003 3:21:55 PM | 1997-01-13 | 6/3/1998 6:51:46 PM | MPAK |  | Concentration | mg/L | Y | (fl. conc.) |  | Phaeophytin A (fl. conc.) |  | Biological Examination; | Biological Examination; EMS code: 0146 | Chemical |  | LAB
0146 | g/m2 | Phaeophytin A | 1841 | Ext:SU:Dup:Sub:Acidified | Spectrophotometric analysis of phaeophytin a derived from material retained on a 0.45 um filter. The filter is ground and extracted in 90% acetone, and the absorbances of the extract are read at 750 and 664 nm. The extract is then acidified by adding a few drops of HCl and the acidified absorbances are read at 750 and 664 nm. | 0.0003 | Y | Y | Y | 89 | 1/30/1997 9:13:00 AM | 1/30/1997 9:13:00 AM | 10/16/1997 6:03:32 PM | 1996-04-12 | 10/19/1998 2:10:53 PM | EMS | NPEPPIN | MassPerArea | g/m² | Y | (mass/area) |  | Phaeophytin A (mass/area) |  | Biological Examination; | Biological Examination; EMS code: 0146 | Chemical |  | LAB
0146 | mg | Phaeophytin A | XM40 | SM 10200 H (2017) Spectrophotometry | SM 10200 H (2017) Spectrophotometry | 0.0001 | Y | Y | Y | 3 | 11/1/2022 11:32:00 AM | 11/1/2022 11:32:00 AM | 11/10/2022 5:26:47 PM | 2022-11-08 |  | EMS |  | Mass | mg | Y | (mass) |  | Phaeophytin A (mass) |  | Biological Examination; | Biological Examination; EMS code: 0146 | Chemical |  | LAB
0146 | mg/L | Phaeophytin A | 1840 | Ext:SU:Dup:Vol:Acidified | Water samples are filtered through glass fibre filters and then frozen. Foam plugs are usually cut from a block with a cork borer of known diameter. The cells are disrupted by ultrasonification and the Chlorophyll is extracted with an aqueous Magnesium Caronate:Acetone mixture. Absorbance measured at about 750, 664, 646 and 630 nm before and after acidification with a few drops of HCl. "LOR." indicates Lorenzen calculation. The addition of acid to a Chlorophyll solution removes the bonded Magnesium of the Porphyrin ring thus converting it to Phaeophytin. This method is single filter. | 0.0005 | Y | Y | Y | 236 |  |  | 10/24/1996 1:41:21 PM | 1996-04-12 | 12/9/1998 9:43:51 AM | EMS | NPEPPIN | Concentration | mg/L | Y | (fl. conc.) |  | Phaeophytin A (fl. conc.) |  | Biological Examination; | Biological Examination; EMS code: 0146 | Chemical |  | LAB
0147 | MPN/100mL | E Coli | X015 | MPN Confirmed | Multiple-tube fermentation technique is the examination of replicate tubes and dilutions reported in terms of Most Probable Number (MPN) of organisms present. This number, based on certain probability formulas, is an estimate of the mean density of bacteria in the sample. MPN method is not influenced by turbidity. | 2 | Y | Y | Y | 1334 | 1/21/2005 | 1/19/2021 4:00:00 PM | 4/11/2023 5:01:29 PM | 1998-11-10 |  | NPEPPIN |  | MicrobialFormation | MPN/100mL | N | (microb.) |  | E Coli (microb.) |  | Bacteriological; | Bacteriological; EMS code: 0147 | Chemical |  | LAB
0147 | MPN/100mL | E Coli | X388 | Colilert | This method describes the selective isolation of total coliforms and E.coli from drinking water. Colilert medium is added to water samples and incubated for 18 - 24 hours at 35 degrees C for growth of total coliforms and E.coli. Colilert may be used as a presence/absence (P-A) or semi-quantitative 5 tube multiple tube fermentation (MPN) test. | 2 | Y | Y | Y | 2550 | 4/28/2000 | 8/15/2024 2:15:00 PM | 8/25/2024 9:47:36 PM | 1996-04-12 | 6/3/1998 6:51:45 PM | EMS |  | MicrobialFormation | MPN/100mL | N | (microb.) |  | E Coli (microb.) |  | Bacteriological; | Bacteriological; EMS code: 0147 | Chemical |  | LAB
0147 | CFU/100mL | E Coli | X022 | Membrane Filter Count | Non-turbid water samples are passed through a 0.45 um membrane filter which is placed on a selective medium for 18 - 24 hours incubation at 35 degrees C for growth of total coliform colonies, of fecal coliform colonies, fecal streptococci colonies, enterococci colonies, or of E. coli colonies. | 1 | Y | Y | Y | 52308 | 1/22/1998 | 8/15/2024 | 8/23/2024 9:50:35 AM | 1996-04-12 | 6/3/1998 6:51:42 PM | EMS |  | MicrobialFormation | CFU/100mL | N | (microb.) |  | E Coli (microb.) |  | Bacteriological; | Bacteriological; EMS code: 0147 | Chemical |  | LAB
0147 | MPN/g | E Coli | E014 | TMECC 07.00 (mod) | Samples are processed, inoculated into a series of fermentation tubes followed by incubation and subsequent inoculations dependent on the target orgniams. Fermentation tubes exhibiting positive reactions proceed to confirmation steps specific to the target organism. | 2 | Y | Y | Y | 44 | 1/17/2022 | 6/6/2024 | 6/18/2024 11:02:03 PM | 2021-08-17 |  | RDGIBSON |  | MicrobialFormation | MPN/g | N | (microb.) |  | E Coli (microb.) |  | Bacteriological; | Bacteriological; EMS code: 0147 | Chemical |  | LAB
0148 | MPN/100gm | Enterococcus | X390 | Confirmed MPN (wt) | This method describes a multiple tube fermentation technique which estimates the Most Probable Number (MPN) in solids, soil, and sludge or diluted samples of blended shellfish. The MPN method is not influenced by turbidity. | 2 | Y | Y | Y | 7 | 7/18/2002 | 7/18/2002 | 10/21/2002 3:11:33 PM | 1999-05-12 |  | NPEPPIN |  | MicrobialFormation | MPN/100g | N | (microb.) |  | Enterococcus (microb.) |  | Bacteriological; | Bacteriological; EMS code: 0148 | Chemical |  | LAB
0148 | CFU/100mL | Enterococcus | X022 | Membrane Filter Count | Non-turbid water samples are passed through a 0.45 um membrane filter which is placed on a selective medium for 18 - 24 hours incubation at 35 degrees C for growth of total coliform colonies, of fecal coliform colonies, fecal streptococci colonies, enterococci colonies, or of E. coli colonies. | 1 | Y | Y | Y | 17169 | 6/23/2000 | 8/15/2024 6:05:00 PM | 8/22/2024 3:35:15 PM | 1996-04-12 | 6/3/1998 6:51:33 PM | EMS |  | MicrobialFormation | CFU/100mL | N | (microb.) |  | Enterococcus (microb.) |  | Bacteriological; | Bacteriological; EMS code: 0148 | Chemical |  | LAB
0148 | MPN/g | Enterococcus | TMECC7 | TMECC 07.00 PATHOGENS | TMECC 07.00 PATHOGENS | 1 | Y | Y | Y | 24 | 3/10/2020 | 3/13/2020 | 2/22/2023 12:51:03 PM | 2023-01-31 |  | EMS |  | MicrobialFormation | MPN/g | N | (microb.) |  | Enterococcus (microb.) |  | Bacteriological; | Bacteriological; EMS code: 0148 | Chemical |  | LAB
0148 | MPN/100mL | Enterococcus | X015 | MPN Confirmed | Multiple-tube fermentation technique is the examination of replicate tubes and dilutions reported in terms of Most Probable Number (MPN) of organisms present. This number, based on certain probability formulas, is an estimate of the mean density of bacteria in the sample. MPN method is not influenced by turbidity. | 2 | Y | Y | Y | 126 | 9/7/2001 | 12/1/2022 1:00:00 PM | 10/31/2023 10:58:05 AM | 1999-10-15 |  | NPEPPIN |  | MicrobialFormation | MPN/100mL | N | (microb.) |  | Enterococcus (microb.) |  | Bacteriological; | Bacteriological; EMS code: 0148 | Chemical |  | LAB
0149 | mg/L | Carbonate Alkalinity (as CaCO3) | XCAL | Calculation (mg/L) | Calculation (mg/L) | 0 | Y | Y | Y | 14827 | 9/29/2000 | 8/14/2024 4:15:00 PM | 8/26/2024 8:06:03 AM | 1998-05-22 | 6/3/1998 6:51:49 PM | NPEPPIN |  | Concentration | mg/L | Y | (fl. conc.) |  | Carbonate Alkalinity (as CaCO3) (fl. conc.) |  | Major Ions; | Major Ions; EMS code: 0149 | Chemical |  | LAB
0149 | mg/L | Carbonate Alkalinity (as CaCO3) | X004 | Colorimetric Titr'n | Manual Colorimetric Titration with standard solutions to specified endpoint as identified by colorimetric indicator. The sample must not be filtered, diluted, concentrated, or altered in any way. | 0.85 | Y | Y | Y | 386 |  |  | 2/27/2023 9:10:15 AM | 2013-04-29 |  | RDGIBSON |  | Concentration | mg/L | Y | (fl. conc.) |  | Carbonate Alkalinity (as CaCO3) (fl. conc.) |  | Major Ions; | Major Ions; EMS code: 0149 | Chemical |  | LAB
0149 | mg/L | Carbonate Alkalinity (as CaCO3) | X310 | Auto Potentiometric Titration | For alkalinity, the samples are (if pH >8.3) first titrated to a pH of 8.3, which corresponds to the equivalence point for the conversion of carbonate to bicarbonate ion (phenolphthalein alkalinity). The titration is continued to a pH of 4.5, which corresponds approx. to the equivalence point for the conversion of bicarbonate ion to carbonic acid (entire titration indicates total alkalinity). An extrapolation technique based on the near proportionality of concentration of H ions to excess of titrant beyond the equivalence pt. The amount of standard acid required to reduce pH exactly 0.30pH (ie. pH4.2) unit is measured carefully. Because this change in pH corresponds to an exact coupling of the H ion concentration, a simple extrapolation can be made to the equivalence point. The actual pH of the stoichiometric end point in alkalinity determinations is the pH at which the inflection in the titration curve occurs and is directly related to the amount of carbonate alkalinity originally present in the sample. A curve can be constructed by plotting pH vs mL of acid to find this point directly. *** For acidity, the samples are titrated to a pH of 8.3, which corresponds to stoichiometric neutralization of carbonic acid to bicarbonate. Because the colour change of phenolphthalein indicator is close to pH 8.3, this value generally is accepted as a standard end point for titration of total acidity, including CO2 and most weak acids an equivalence point for the conversion of carbonate to bicarbonate ion (phenolphthalein alkalinity). The actual pH of the stoichiometric end point in acidity determinations is the pH at which the inflection in the titration curve occurs and is directly related to the amount of titratable acids originally present in the sample. A curve can be constructed by plotting pH vs. mL of NaOH to find this point directly. (To obtain the actual substances contributing to acidity/alkalinity, a pH vs mL of acid/NaOH curve should be constructed.) | 0.85 | Y | Y | Y | 20602 | 3/24/2011 | 8/22/2024 12:47:00 PM | 8/26/2024 4:53:05 PM | 2013-04-29 |  | RDGIBSON |  | Concentration | mg/L | Y | (fl. conc.) |  | Carbonate Alkalinity (as CaCO3) (fl. conc.) |  | Major Ions; | Major Ions; EMS code: 0149 | Chemical |  | LAB
0150 | mg/L | Bicarbonate Alkalinity (as CaCO3) | XCAL | Calculation (mg/L) | Calculation (mg/L) | 0 | Y | Y | Y | 15415 | 9/29/2000 | 8/14/2024 4:15:00 PM | 8/26/2024 8:06:03 AM | 1998-07-10 |  | NPEPPIN |  | Concentration | mg/L | Y | (fl. conc.) |  | Bicarbonate Alkalinity (as CaCO3) (fl. conc.) |  | Major Ions; | Major Ions; EMS code: 0150 | Chemical |  | LAB
0150 | mg/L | Bicarbonate Alkalinity (as CaCO3) | X310 | Auto Potentiometric Titration | For alkalinity, the samples are (if pH >8.3) first titrated to a pH of 8.3, which corresponds to the equivalence point for the conversion of carbonate to bicarbonate ion (phenolphthalein alkalinity). The titration is continued to a pH of 4.5, which corresponds approx. to the equivalence point for the conversion of bicarbonate ion to carbonic acid (entire titration indicates total alkalinity). An extrapolation technique based on the near proportionality of concentration of H ions to excess of titrant beyond the equivalence pt. The amount of standard acid required to reduce pH exactly 0.30pH (ie. pH4.2) unit is measured carefully. Because this change in pH corresponds to an exact coupling of the H ion concentration, a simple extrapolation can be made to the equivalence point. The actual pH of the stoichiometric end point in alkalinity determinations is the pH at which the inflection in the titration curve occurs and is directly related to the amount of carbonate alkalinity originally present in the sample. A curve can be constructed by plotting pH vs mL of acid to find this point directly. *** For acidity, the samples are titrated to a pH of 8.3, which corresponds to stoichiometric neutralization of carbonic acid to bicarbonate. Because the colour change of phenolphthalein indicator is close to pH 8.3, this value generally is accepted as a standard end point for titration of total acidity, including CO2 and most weak acids an equivalence point for the conversion of carbonate to bicarbonate ion (phenolphthalein alkalinity). The actual pH of the stoichiometric end point in acidity determinations is the pH at which the inflection in the titration curve occurs and is directly related to the amount of titratable acids originally present in the sample. A curve can be constructed by plotting pH vs. mL of NaOH to find this point directly. (To obtain the actual substances contributing to acidity/alkalinity, a pH vs mL of acid/NaOH curve should be constructed.) | 0.85 | Y | Y | Y | 20600 | 3/24/2011 | 8/22/2024 12:47:00 PM | 8/26/2024 4:53:05 PM | 2013-04-29 |  | RDGIBSON |  | Concentration | mg/L | Y | (fl. conc.) |  | Bicarbonate Alkalinity (as CaCO3) (fl. conc.) |  | Major Ions; | Major Ions; EMS code: 0150 | Chemical |  | LAB
0150 | mg/L | Bicarbonate Alkalinity (as CaCO3) | X004 | Colorimetric Titr'n | Manual Colorimetric Titration with standard solutions to specified endpoint as identified by colorimetric indicator. The sample must not be filtered, diluted, concentrated, or altered in any way. | 0.85 | Y | Y | Y | 386 |  |  | 2/27/2023 9:10:15 AM | 2013-04-29 |  | RDGIBSON |  | Concentration | mg/L | Y | (fl. conc.) |  | Bicarbonate Alkalinity (as CaCO3) (fl. conc.) |  | Major Ions; | Major Ions; EMS code: 0150 | Chemical |  | LAB
0157 | mg/L | Cyanide WAD | F013 | Tetracyanonickelate,UV Spec. | The sample is refluxed for 1 hour in the presence of Zinc Acetate and Sodium Acetate buffer of pH 4.5. Zinc Acetate is added to render Iron Cyanide insoluble. Simple cyanides and some of the complex metal cyanides (eg., Cd, Cu, Ni, Zn) react to produce HCN gas. The HCN produced reacts with the Ni ions in the tetranickelate solution to form the ionic tetracyanonickelate(11) anion complex, Ni(CN)4-2, which has a characteristic ultra-violet spectrum. | 0.03 | Y | Y | Y | 227 | 7/30/1996 1:53:00 PM | 12/5/2001 1:40:00 PM | 5/26/2003 2:36:54 PM | 1996-06-19 | 6/3/1998 6:51:44 PM | MPAK |  | Concentration | mg/L | Y | (fl. conc.) |  | Cyanide WAD (fl. conc.) | 57-12-5 | Inorganics; | Inorganics; EMS code: 0157 | Chemical |  | LAB
0157 | ug/g wet | Cyanide WAD | FS14 | Tetracyanonickelatate,UV Spec. | For weak acid dissociable cyanide, CN(WAD), the solid sample is added to type one water and is refluxed for 1 hour in the presence of Zinc acetate and sodium acetate. The HCN is liberated from this slightly acidified (pH 4.5 - 6.0) sample. The acetate buffer used contains Zinc salts to precipitate iron cyanide as a further assurance of the selectivity of the method. For strong acid dissociable cyanide, CN(SAD), the solid sample is added to Type 1 water and is refluxed for 1 hour in the presence of sulphuric acid and magnesium chloride. For both CN(WAD) and CN(SAD) the HCN generated reacts instantaneously with nickel ions in an ammoniacal solution to form the ionic tetracyanonickelate (II) anion complex, Ni (CN)4-2, which has a characteristic ultra-violet spectrum, and is measured at 267 and 284 nm. | 0.5 | Y | Y | Y | 5 | 12/5/2001 1:38:00 PM | 12/5/2001 3:11:00 PM | 5/26/2003 2:27:02 PM | 1997-06-23 | 6/3/1998 6:51:47 PM | NPEPPIN |  | MassPerMass | µg/g | Y | (mass conc.) | (wet) | Cyanide WAD (wet) (mass conc.) | 57-12-5 | Inorganics; | Inorganics; EMS code: 0157 | Chemical |  | LAB
0157 | mg/L | Cyanide WAD | X066 | CN Selective Electrode | CN Selective Electrode | 0.002 | Y | Y | Y | 1330 | 8/4/2015 10:36:00 AM | 1/7/2023 1:33:00 PM | 1/11/2023 5:01:27 PM | 2017-10-23 |  | JKROGH |  | Concentration | mg/L | Y | (fl. conc.) |  | Cyanide WAD (fl. conc.) | 57-12-5 | Inorganics; | Inorganics; EMS code: 0157 | Chemical |  | LAB
0157 | mg/L | Cyanide WAD | AA07 | Auto.Pyridine-Barbituric - low level | The sample is distilled from a phosphoric acid medium into a phosphate buffer solution (pH 5.2). Simple cyanides, eg,. those that distill from the phosphoric acid solution without any ultraviolet irradiation, are absorbed into the phosphate buffer solution after distillation. The cyanide in the buffer solution is converted to cyanogen chloride, (CNCI) by reaction with chloramine-T. The CNCI then forms a red coloured complex on addition of the pyridine-barbituric acid reagent and the absorbance is measured at 580 nm. The Konig (1,2) synthesis first described the specific and quantitative reaction of cyanogen halides to cleave pyridine to produce glutaconic aldehyde. This reaction has been applied to the determination of cyanide by oxidation of CN- to cyanogen halide, CNX, in which CN+ is the reactive species. The CN+ reacts with pyridine to produce an intermediate which hydrolyzes to a conjugated dialdehyde, glutaconic aldehyde. The glutaconic aldehyde is then coupled with a primary amine, of a compound containing reactive methylene hydrogens (RH2), barbituric acid in this case (3). | 0.0005 | Y | Y | Y | 5898 | 1/16/1997 4:42:00 PM | 7/11/2024 10:40:00 AM | 8/21/2024 11:42:55 AM | 1996-04-12 | 6/3/1998 6:51:33 PM | EMS |  | Concentration | mg/L | Y | (fl. conc.) |  | Cyanide WAD (fl. conc.) | 57-12-5 | Inorganics; | Inorganics; EMS code: 0157 | Chemical |  | LAB
0157 | mg/L | Cyanide WAD | X400 | Zinc Acetate,Isonicotin. | The sample is subjected to a weak acid [zinc acetate] reflux digestion to liberate hydrogen cyanide. The resulting HCN is distilled from the digestion mixture and trapped in a weak NaOH solution. This basic solution is then analyzed colorimetrically, ref. Isonicotinic - Barbituric Acid method. | 0.005 | Y | Y | Y | 1190 | 2/2/1999 11:00:00 AM | 11/6/2012 8:00:00 AM | 12/20/2012 11:18:32 AM | 1996-04-12 | 6/3/1998 6:51:45 PM | EMS |  | Concentration | mg/L | Y | (fl. conc.) |  | Cyanide WAD (fl. conc.) | 57-12-5 | Inorganics; | Inorganics; EMS code: 0157 | Chemical |  | LAB
0157 | mg/L | Cyanide WAD | EQ01 | EQUIS (mg/L) | EQUIS (mg/L) | 0.005 | N | Y | Y | 30 |  |  | 10/22/1996 10:26:53 PM | 1996-04-12 | 6/3/1998 6:51:45 PM | EMS |  | Concentration | mg/L | Y | (fl. conc.) |  | Cyanide WAD (fl. conc.) | 57-12-5 | Inorganics; | Inorganics; EMS code: 0157 | Chemical |  | LAB
0157 | mg/L | Cyanide WAD | X323 | Isonicotinic Barb Col'm | An aliquot of the alkaline distillate from the analyst's choice of digestion procedure is buffered to pH<8 and reacted with chloramine-T. Isonicotinic acid - barbituric acid solution is added and the absorption of the solution is measured at 578 nm. Soils and sediments can be analyzed by suspension in the digestion solution. | 0.01 | Y | Y | Y | 27668 |  |  | 5/19/2015 3:10:36 PM | 1996-04-12 | 6/3/1998 6:51:33 PM | EMS |  | Concentration | mg/L | Y | (fl. conc.) |  | Cyanide WAD (fl. conc.) | 57-12-5 | Inorganics; | Inorganics; EMS code: 0157 | Chemical |  | LAB
0157 | mg/L | Cyanide WAD | 0540 | Acet.Buff.Dig:UV | Acet.Buff.Dig:UV | 0.005 | Y | Y | Y | 2138 |  |  | 12/14/2015 10:27:40 AM | 1996-04-12 | 6/3/1998 6:51:33 PM | EMS |  | Concentration | mg/L | Y | (fl. conc.) |  | Cyanide WAD (fl. conc.) | 57-12-5 | Inorganics; | Inorganics; EMS code: 0157 | Chemical |  | LAB
0157 | mg/L | Cyanide WAD | X207 | Acet.Buff.Dig,Col'metric | Acet.Buff.Dig;Col'metric | 0.005 | Y | Y | Y | 7445 | 5/9/2013 | 8/22/2024 | 8/26/2024 3:51:07 PM | 1996-04-12 | 6/3/1998 6:51:33 PM | EMS |  | Concentration | mg/L | Y | (fl. conc.) |  | Cyanide WAD (fl. conc.) | 57-12-5 | Inorganics; | Inorganics; EMS code: 0157 | Chemical |  | LAB
0292 | Bq/L | Gross Alpha Total | EQ16 | EQUIS (Bq/L) | EQUIS (Bq/L) | 0.01 | N | Y | Y | 3 |  |  | 10/22/1996 8:04:40 PM | 1996-04-12 | 6/3/1998 6:51:33 PM | EMS |  | Radioactivity | Bq/L | Y | (rad./vol.) |  | Gross Alpha Total (rad./vol.) |  | Radionuclide; | Radionuclide; EMS code: 0292 | Chemical | Total | LAB
0450 | MPN/g | Coliform - Fecal | X390 | Confirmed MPN (wt) | This method describes a multiple tube fermentation technique which estimates the Most Probable Number (MPN) in solids, soil, and sludge or diluted samples of blended shellfish. The MPN method is not influenced by turbidity. | 2 | Y | Y | Y | 224 | 9/7/2001 | 6/6/2024 | 6/18/2024 11:02:03 PM | 1996-04-12 | 6/3/1998 6:51:33 PM | EMS |  | MicrobialFormation | MPN/g | N | (microb.) |  | Coliform - Fecal (microb.) |  | Bacteriological; | Bacteriological; EMS code: 0450 | Chemical |  | LAB
0450 | MPN/100mL | Coliform - Fecal | X015 | MPN Confirmed | Multiple-tube fermentation technique is the examination of replicate tubes and dilutions reported in terms of Most Probable Number (MPN) of organisms present. This number, based on certain probability formulas, is an estimate of the mean density of bacteria in the sample. MPN method is not influenced by turbidity. | 2 | Y | Y | Y | 13241 | 9/3/1996 | 8/15/2024 2:15:00 PM | 8/25/2024 9:47:36 PM | 1996-04-12 | 6/3/1998 6:51:33 PM | EMS |  | MicrobialFormation | MPN/100mL | N | (microb.) |  | Coliform - Fecal (microb.) |  | Bacteriological; | Bacteriological; EMS code: 0450 | Chemical |  | LAB
0450 | CFU/100mL | Coliform - Fecal | X022 | Membrane Filter Count | Non-turbid water samples are passed through a 0.45 um membrane filter which is placed on a selective medium for 18 - 24 hours incubation at 35 degrees C for growth of total coliform colonies, of fecal coliform colonies, fecal streptococci colonies, enterococci colonies, or of E. coli colonies. | 2 | Y | Y | Y | 93201 | 11/18/1996 9:30:00 AM | 8/15/2024 3:05:00 PM | 8/22/2024 5:01:12 PM | 1996-04-12 | 5/1/2002 1:54:20 PM | EMS | NPEPPIN | MicrobialFormation | CFU/100mL | N | (microb.) |  | Coliform - Fecal (microb.) |  | Bacteriological; | Bacteriological; EMS code: 0450 | Chemical |  | LAB
0450 | MPN/100mL | Coliform - Fecal | X388 | Colilert | This method describes the selective isolation of total coliforms and E.coli from drinking water. Colilert medium is added to water samples and incubated for 18 - 24 hours at 35 degrees C for growth of total coliforms and E.coli. Colilert may be used as a presence/absence (P-A) or semi-quantitative 5 tube multiple tube fermentation (MPN) test. | 2 | Y | Y | Y | 1352 | 4/2/2007 | 8/14/2024 | 8/21/2024 4:20:08 PM | 2001-08-22 |  | NPEPPIN |  | MicrobialFormation | MPN/100mL | N | (microb.) |  | Coliform - Fecal (microb.) |  | Bacteriological; | Bacteriological; EMS code: 0450 | Chemical |  | LAB
0450 | MPN/100mL | Coliform - Fecal | 2495 | MPN Completed | MPN Completed | 0 | Y | Y | Y | 1071 | 8/30/2007 | 1/26/2015 | 9/22/2021 3:20:50 PM | 1996-04-12 | 6/3/1998 6:51:33 PM | EMS |  | MicrobialFormation | MPN/100mL | N | (microb.) |  | Coliform - Fecal (microb.) |  | Bacteriological; | Bacteriological; EMS code: 0450 | Chemical |  | LAB
0450 | CFU/100mL | Coliform - Fecal | EQ22 | EQUIS (No/100ml) | EQUIS (No/100ml) | 0 | N | Y | Y | 361 |  |  | 10/24/1996 3:17:05 PM | 1996-04-12 | 6/3/1998 6:51:33 PM | EMS |  | MicrobialFormation | CFU/100mL | N | (microb.) |  | Coliform - Fecal (microb.) |  | Bacteriological; | Bacteriological; EMS code: 0450 | Chemical |  | LAB
0450 | None | Coliform - Fecal | BISP | Biology Special Analysis | Biology Special Analysis | 0 | N | Y | N | 1 |  |  | 10/24/1996 9:18:58 AM | 1996-04-12 | 6/3/1998 6:51:33 PM | EMS |  | None | None | N | (unitless) |  | Coliform - Fecal (unitless) |  | Bacteriological; | Bacteriological; EMS code: 0450 | Chemical |  | LAB
0450 | MPN/100mL | Coliform - Fecal | MTFC | Multiple Tube Fermentation | Multiple Tube Fermentation (concentration) |  | Y | Y | Y | 1501 |  |  | 8/23/2024 8:10:36 AM | 2022-09-15 |  | EMS |  | MicrobialFormation | MPN/100mL | N | (microb.) |  | Coliform - Fecal (microb.) |  | Bacteriological; | Bacteriological; EMS code: 0450 | Chemical |  | LAB
0450 | CFU/100mL | Coliform - Fecal | X550 | Enriched Membrane Filter Count | Place a sterile absorbent pad in the lid of a sterile culture dish and pipet at least 2.0 mL lauryl tryptose broth, prepared as directed in 9221B.1.a1), to saturate pad. Carefully remove any excess liquid from absorbent pad by decanting plate. Aseptically place filter through which the sample has been passed on pad. Incubate filter, without inverting dish, for 1.5 to 2h at 35 +/- 0.5 degree C in an atmosphere of at least 60% relative humidity. | 1 | Y | Y | Y | 2 |  |  | 10/22/2003 3:05:55 PM | 2003-08-12 | 8/12/2003 9:42:12 AM | UMULLER | UMULLER | MicrobialFormation | CFU/100mL | N | (microb.) |  | Coliform - Fecal (microb.) |  | Bacteriological; | Bacteriological; EMS code: 0450 | Chemical |  | LAB
0450 | MPN/100mL | Coliform - Fecal | EQ23 | EQUIS (MPN) | EQUIS (MPN) | 0 | N | Y | Y | 9340 |  |  | 10/24/1996 3:16:44 PM | 1996-04-12 | 6/3/1998 6:51:33 PM | EMS |  | MicrobialFormation | MPN/100mL | N | (microb.) |  | Coliform - Fecal (microb.) |  | Bacteriological; | Bacteriological; EMS code: 0450 | Chemical |  | LAB
0450 | CFU/100mL | Coliform - Fecal | 2481 | Filt Spl,MF Count -EL | Filt Spl;MF Count -EL | 0 | Y | Y | Y | 274 | 11/1/2022 | 10/19/2023 | 11/9/2023 7:41:27 AM | 1996-04-12 | 6/3/1998 6:51:33 PM | EMS |  | MicrobialFormation | CFU/100mL | N | (microb.) |  | Coliform - Fecal (microb.) |  | Bacteriological; | Bacteriological; EMS code: 0450 | Chemical |  | LAB
0451 | MPN/100mL | Coliform - Total | EQ23 | EQUIS (MPN) | EQUIS (MPN) | 0 | N | Y | Y | 7747 |  |  | 10/24/1996 3:16:44 PM | 1996-04-12 | 6/3/1998 6:51:33 PM | EMS |  | MicrobialFormation | MPN/100mL | N | (microb.) |  | Coliform - Total (microb.) |  | Bacteriological; | Bacteriological; EMS code: 0451 | Chemical | Total | LAB
0451 | MPN/100mL | Coliform - Total | X015 | MPN Confirmed | Multiple-tube fermentation technique is the examination of replicate tubes and dilutions reported in terms of Most Probable Number (MPN) of organisms present. This number, based on certain probability formulas, is an estimate of the mean density of bacteria in the sample. MPN method is not influenced by turbidity. | 2 | Y | Y | Y | 5507 | 9/3/1996 | 6/4/2024 | 7/29/2024 3:17:42 PM | 1996-04-12 | 6/3/1998 6:51:33 PM | EMS |  | MicrobialFormation | MPN/100mL | N | (microb.) |  | Coliform - Total (microb.) |  | Bacteriological; | Bacteriological; EMS code: 0451 | Chemical | Total | LAB
0451 | MPN/100mL | Coliform - Total | X388 | Colilert | This method describes the selective isolation of total coliforms and E.coli from drinking water. Colilert medium is added to water samples and incubated for 18 - 24 hours at 35 degrees C for growth of total coliforms and E.coli. Colilert may be used as a presence/absence (P-A) or semi-quantitative 5 tube multiple tube fermentation (MPN) test. | 2 | Y | Y | Y | 1660 | 1/22/2002 2:00:00 PM | 8/14/2024 3:55:00 PM | 8/25/2024 9:47:36 PM | 1996-04-12 | 6/3/1998 6:51:33 PM | EMS |  | MicrobialFormation | MPN/100mL | N | (microb.) |  | Coliform - Total (microb.) |  | Bacteriological; | Bacteriological; EMS code: 0451 | Chemical | Total | LAB
0451 | CFU/100mL | Coliform - Total | 2481 | Filt Spl,MF Count -EL | Filt Spl;MF Count -EL | 0 | Y | Y | Y | 71 |  |  | 2/27/2007 3:13:27 PM | 1996-04-12 | 5/1/2002 1:55:04 PM | EMS | NPEPPIN | MicrobialFormation | CFU/100mL | N | (microb.) |  | Coliform - Total (microb.) |  | Bacteriological; | Bacteriological; EMS code: 0451 | Chemical | Total | LAB
0451 | CFU/100mL | Coliform - Total | EQ22 | EQUIS (No/100ml) | EQUIS (No/100ml) | 0 | N | Y | Y | 180 |  |  | 10/24/1996 3:17:05 PM | 1996-04-12 | 5/1/2002 1:55:04 PM | EMS | NPEPPIN | MicrobialFormation | CFU/100mL | N | (microb.) |  | Coliform - Total (microb.) |  | Bacteriological; | Bacteriological; EMS code: 0451 | Chemical | Total | LAB
0451 | CFU/100mL | Coliform - Total | X022 | Membrane Filter Count | Non-turbid water samples are passed through a 0.45 um membrane filter which is placed on a selective medium for 18 - 24 hours incubation at 35 degrees C for growth of total coliform colonies, of fecal coliform colonies, fecal streptococci colonies, enterococci colonies, or of E. coli colonies. | 2 | Y | Y | Y | 13029 | 5/5/1998 | 7/17/2024 2:49:00 PM | 7/30/2024 5:52:05 PM | 1996-04-12 | 5/1/2002 1:55:04 PM | EMS | NPEPPIN | MicrobialFormation | CFU/100mL | N | (microb.) |  | Coliform - Total (microb.) |  | Bacteriological; | Bacteriological; EMS code: 0451 | Chemical | Total | LAB
0451 | MPN/g | Coliform - Total | X390 | Confirmed MPN (wt) | This method describes a multiple tube fermentation technique which estimates the Most Probable Number (MPN) in solids, soil, and sludge or diluted samples of blended shellfish. The MPN method is not influenced by turbidity. | 2 | Y | Y | Y | 25 | 1/17/2022 | 10/22/2022 5:05:00 PM | 1/11/2024 1:07:06 PM | 2022-07-27 |  | EMS |  | MicrobialFormation | MPN/g | N | (microb.) |  | Coliform - Total (microb.) |  | Bacteriological; | Bacteriological; EMS code: 0451 | Chemical | Total | LAB
0454 | MPN/g | Streptococcus -Fecal | X390 | Confirmed MPN (wt) | This method describes a multiple tube fermentation technique which estimates the Most Probable Number (MPN) in solids, soil, and sludge or diluted samples of blended shellfish. The MPN method is not influenced by turbidity. | 2 | Y | Y | Y | 7 | 2/26/2002 | 2/26/2002 | 3/9/2002 11:34:08 AM | 1996-04-12 | 6/3/1998 6:51:33 PM | EMS |  | MicrobialFormation | MPN/g | N | (microb.) |  | Streptococcus -Fecal (microb.) |  | Bacteriological; | Bacteriological; EMS code: 0454 | Chemical |  | LAB
0454 | MPN/100mL | Streptococcus -Fecal | X015 | MPN Confirmed | Multiple-tube fermentation technique is the examination of replicate tubes and dilutions reported in terms of Most Probable Number (MPN) of organisms present. This number, based on certain probability formulas, is an estimate of the mean density of bacteria in the sample. MPN method is not influenced by turbidity. | 0 | Y | Y | Y | 12 | 8/17/2014 1:00:00 PM | 8/18/2016 2:15:00 PM | 8/26/2016 5:00:15 PM | 1996-04-12 | 6/3/1998 6:51:41 PM | EMS |  | MicrobialFormation | MPN/100mL | N | (microb.) |  | Streptococcus -Fecal (microb.) |  | Bacteriological; | Bacteriological; EMS code: 0454 | Chemical |  | LAB
0454 | MPN/100mL | Streptococcus -Fecal | EQ23 | EQUIS (MPN) | EQUIS (MPN) | 0 | N | Y | Y | 135 |  |  | 10/24/1996 3:16:40 PM | 1996-04-12 | 6/3/1998 6:51:33 PM | EMS |  | MicrobialFormation | MPN/100mL | N | (microb.) |  | Streptococcus -Fecal (microb.) |  | Bacteriological; | Bacteriological; EMS code: 0454 | Chemical |  | LAB
0454 | CFU/100mL | Streptococcus -Fecal | X022 | Membrane Filter Count | Non-turbid water samples are passed through a 0.45 um membrane filter which is placed on a selective medium for 18 - 24 hours incubation at 35 degrees C for growth of total coliform colonies, of fecal coliform colonies, fecal streptococci colonies, enterococci colonies, or of E. coli colonies. | 0 | Y | Y | Y | 2957 | 4/6/2000 | 2/16/2010 | 2/19/2010 10:29:39 AM | 1996-04-12 | 6/3/1998 6:51:33 PM | EMS |  | MicrobialFormation | CFU/100mL | N | (microb.) |  | Streptococcus -Fecal (microb.) |  | Bacteriological; | Bacteriological; EMS code: 0454 | Chemical |  | LAB
0454 | CFU/100mL | Streptococcus -Fecal | EQ22 | EQUIS (No/100ml) | EQUIS (No/100ml) | 0 | N | Y | Y | 71 |  |  | 10/22/1996 11:02:53 PM | 1996-04-12 | 6/3/1998 6:51:33 PM | EMS |  | MicrobialFormation | CFU/100mL | N | (microb.) |  | Streptococcus -Fecal (microb.) |  | Bacteriological; | Bacteriological; EMS code: 0454 | Chemical |  | LAB
0457 | %(V/V) | Microtox IC50 5 Min | X393 | 1-IC50,5&15min M500 Micr | Living bioluminescent bacteria are exposed to a test sample, and the toxic effect of the sample on the organisms is measured. The Microtox test system measures the light output of rehydrated lyophilized bacteria after exposure to a specified dilution series of a sample, and compares it to the light output of a control blank (i.e./bacterial cell suspension in diluent only). It is assumed that the difference in light production is due to the effect of the sample. The degree of light loss (degree of metabolic inhibition in the bacteria) indicates the degree of toxicity of the sample and the IC50 is determined by a dose-response curve. The IC50 is defined as the statistically determined concentration of a sample causing a 50% decrease in the light emission of Microtox bacteria Vibrio fischeri (formerly Photobacterium phosphoreum) under defined conditions of exposure time (usually 5 minutes or 15 minutes) and test temperature. Note that fresh water samples are osmotically adjusted. However, there should be no adjustment for pH nor for colour. | 0 | Y | Y | Y | 88 | 4/3/2000 11:55:00 AM | 2/13/2002 1:05:00 PM | 2/17/2003 3:40:05 PM | 1996-04-12 | 6/3/1998 6:51:33 PM | EMS |  | DimensionlessRatio | % | N | (ratio) | (by vol.) | Microtox IC50 5 Min (by vol.) (ratio) |  | Toxicity; | Toxicity; EMS code: 0457 | Chemical |  | LAB
0457 | %(V/V) | Microtox IC50 5 Min | 1890 | Photometric by Vol. | Photometric by Vol. | 0.01 | Y | Y | Y | 109 | 9/29/2000 | 9/29/2000 | 10/4/2010 12:26:48 PM | 1996-04-12 | 6/3/1998 6:51:33 PM | EMS |  | DimensionlessRatio | % | N | (ratio) | (by vol.) | Microtox IC50 5 Min (by vol.) (ratio) |  | Toxicity; | Toxicity; EMS code: 0457 | Chemical |  | LAB
0457 | %(V/V) | Microtox IC50 5 Min | 1885 | Phot. Color Corr by Vol | Phot. Color Corr by Vol | 0.01 | Y | Y | Y | 66 |  |  | 10/24/1996 6:42:06 AM | 1996-04-12 | 6/3/1998 6:51:33 PM | EMS |  | DimensionlessRatio | % | N | (ratio) | (by vol.) | Microtox IC50 5 Min (by vol.) (ratio) |  | Toxicity; | Toxicity; EMS code: 0457 | Chemical |  | LAB
0457 | %(V/V) | Microtox IC50 5 Min | 1891 | Photo.Vol:pH Adjusted | Photo.Vol:pH Adjusted | 0.01 | Y | Y | Y | 6 |  |  | 10/24/1996 9:33:39 AM | 1996-04-12 | 6/3/1998 6:51:45 PM | EMS |  | DimensionlessRatio | % | N | (ratio) | (by vol.) | Microtox IC50 5 Min (by vol.) (ratio) |  | Toxicity; | Toxicity; EMS code: 0457 | Chemical |  | LAB
0457 | %(V/V) | Microtox IC50 5 Min | EQ11 | EQUIS (%V/V) | EQUIS (%V/V) | 0.01 | N | Y | Y | 98 |  |  | 10/22/1996 9:59:33 PM | 1996-04-12 | 6/3/1998 6:51:33 PM | EMS |  | DimensionlessRatio | % | N | (ratio) | (by vol.) | Microtox IC50 5 Min (by vol.) (ratio) |  | Toxicity; | Toxicity; EMS code: 0457 | Chemical |  | LAB
0457 | %(V/V) | Microtox IC50 5 Min | 1887 | Phot.C.Corr:pH Adj Vol. | Phot.C.Corr:pH Adj Vol. | 0.01 | Y | Y | Y | 3 |  |  | 10/24/1996 6:59:09 AM | 1996-04-12 | 6/3/1998 6:51:33 PM | EMS |  | DimensionlessRatio | % | N | (ratio) | (by vol.) | Microtox IC50 5 Min (by vol.) (ratio) |  | Toxicity; | Toxicity; EMS code: 0457 | Chemical |  | LAB
0458 | %(V/V) | Microtox IC50 15Min. | 1885 | Phot. Color Corr by Vol | Phot. Color Corr by Vol | 0.01 | Y | Y | Y | 4 |  |  | 8/26/1997 12:53:16 PM | 1996-04-12 | 6/3/1998 6:51:45 PM | EMS |  | DimensionlessRatio | % | N | (ratio) | (by vol.) | Microtox IC50 15Min. (by vol.) (ratio) |  | Toxicity; | Toxicity; EMS code: 0458 | Chemical |  | LAB
0458 | %(V/V) | Microtox IC50 15Min. | X393 | 1-IC50,5&15min M500 Micr | Living bioluminescent bacteria are exposed to a test sample, and the toxic effect of the sample on the organisms is measured. The Microtox test system measures the light output of rehydrated lyophilized bacteria after exposure to a specified dilution series of a sample, and compares it to the light output of a control blank (i.e./bacterial cell suspension in diluent only). It is assumed that the difference in light production is due to the effect of the sample. The degree of light loss (degree of metabolic inhibition in the bacteria) indicates the degree of toxicity of the sample and the IC50 is determined by a dose-response curve. The IC50 is defined as the statistically determined concentration of a sample causing a 50% decrease in the light emission of Microtox bacteria Vibrio fischeri (formerly Photobacterium phosphoreum) under defined conditions of exposure time (usually 5 minutes or 15 minutes) and test temperature. Note that fresh water samples are osmotically adjusted. However, there should be no adjustment for pH nor for colour. | 0 | Y | Y | Y | 90 | 4/3/2000 11:55:00 AM | 2/13/2002 1:05:00 PM | 2/17/2003 3:40:06 PM | 1996-04-12 | 6/3/1998 6:51:33 PM | EMS |  | DimensionlessRatio | % | N | (ratio) | (by vol.) | Microtox IC50 15Min. (by vol.) (ratio) |  | Toxicity; | Toxicity; EMS code: 0458 | Chemical |  | LAB
0458 | %(V/V) | Microtox IC50 15Min. | 1887 | Phot.C.Corr:pH Adj Vol. | Phot.C.Corr:pH Adj Vol. | 0.01 | Y | Y | Y | 4 |  |  | 10/24/1996 9:33:39 AM | 1996-04-12 | 6/3/1998 6:51:33 PM | EMS |  | DimensionlessRatio | % | N | (ratio) | (by vol.) | Microtox IC50 15Min. (by vol.) (ratio) |  | Toxicity; | Toxicity; EMS code: 0458 | Chemical |  | LAB
0458 | %(V/V) | Microtox IC50 15Min. | EQ11 | EQUIS (%V/V) | EQUIS (%V/V) | 0.01 | N | Y | Y | 64 |  |  | 10/22/1996 9:27:16 PM | 1996-04-12 | 6/3/1998 6:51:33 PM | EMS |  | DimensionlessRatio | % | N | (ratio) | (by vol.) | Microtox IC50 15Min. (by vol.) (ratio) |  | Toxicity; | Toxicity; EMS code: 0458 | Chemical |  | LAB
0458 | %(V/V) | Microtox IC50 15Min. | 1890 | Photometric by Vol. | Photometric by Vol. | 0.01 | Y | Y | Y | 43 | 9/29/2000 | 9/29/2000 | 10/4/2010 12:26:49 PM | 1996-04-12 | 6/3/1998 6:51:33 PM | EMS |  | DimensionlessRatio | % | N | (ratio) | (by vol.) | Microtox IC50 15Min. (by vol.) (ratio) |  | Toxicity; | Toxicity; EMS code: 0458 | Chemical |  | LAB
0459 | % | Bio. LT50:Fish | X542 | Rainbow/96hr/pH,hardness adj (%surv) | Rainbow trout, %survival after 96 hrs. Sample is adjusted for pH and hardness without dilution to within +/- 10% of the receiving water. | 0 | Y | Y | Y | 26 |  |  | 10/15/2018 12:58:48 PM | 2001-02-20 |  | NPEPPIN |  | DimensionlessRatio | % | N | (ratio) |  | Bio. LT50:Fish (ratio) |  | Toxicity; | Toxicity; EMS code: 0459 | Chemical |  | LAB
0459 | h | Bio. LT50:Fish | FA20 | Rainbow(%mortality)/100%conc.) | This method measures the time (in hours) for which 50% mortality occurs in undiluted (i.e. 100% conc.) effluent. | 0 | Y | Y | Y | 85 | 2/7/2000 8:34:00 AM | 10/16/2003 | 10/15/2018 12:58:48 PM | 2000-04-10 |  | NPEPPIN |  | Time | hr | Y | (time) |  | Bio. LT50:Fish (time) |  | Toxicity; | Toxicity; EMS code: 0459 | Chemical |  | LAB
0460 | mg | Biomass;dry wt @105C | X237 | Gravimetric | This code is a catch-all for a variety of gravimetric procedures. | 1 | Y | Y | Y | 1205 | 10/1/1996 11:41:00 AM | 10/7/1999 8:15:00 AM | 10/20/1999 1:36:58 AM | 1996-04-12 | 6/3/1998 6:51:45 PM | EMS |  | Mass | mg | Y | (mass) |  | Biomass;dry wt @105C (mass) |  | Biological Examination; | Biological Examination; EMS code: 0460 | Chemical |  | LAB
0460 | mg | Biomass;dry wt @105C | EQ03 | EQUIS (mg) | EQUIS (mg) | 1 | N | Y | Y | 366 |  |  | 10/22/1996 11:14:30 PM | 1996-04-12 | 6/3/1998 6:51:33 PM | EMS |  | Mass | mg | Y | (mass) |  | Biomass;dry wt @105C (mass) |  | Biological Examination; | Biological Examination; EMS code: 0460 | Chemical |  | LAB
0460 | mg/L | Biomass;dry wt @105C | X313 | Gravimetric, dry wt@105C | A measured volume of sample is filtered through a 0.45 um membrane filter or a pre-rinsed dried and pre-weighed glass fibre filter. The filter is rinsed, removed from the filtration apparatus and dried at 105 C for 24 hours than cooled in a desiccator and reweighed. Results may be reported in units of mg (per whole sample) but final results should be reported either in mg/L or mg/m2, whichever is appropriate. | 4 | Y | Y | Y | 19 | 10/30/1998 2:24:00 PM | 8/24/1999 4:36:00 PM | 8/27/1999 1:07:21 AM | 1996-04-12 | 6/3/1998 6:51:33 PM | EMS |  | Concentration | mg/L | Y | (fl. conc.) |  | Biomass;dry wt @105C (fl. conc.) |  | Biological Examination; | Biological Examination; EMS code: 0460 | Chemical |  | LAB
0461 | %(W/W) | Bio. LC50:Fish | FA19 | Rainbow(96hr)(%mortality/100%conc. | This method measures the concentration (expressed as % weight/weight) at which 50% mortality occurs. No pH adjustment usually is permitted. | 0 | Y | Y | Y | 27 | 2/4/2000 10:36:00 AM | 1/29/2002 2:12:00 PM | 1/5/2007 3:57:50 PM | 2000-04-10 |  | NPEPPIN |  | DimensionlessRatio | % | N | (ratio) | (wet) | Bio. LC50:Fish (wet) (ratio) |  | Toxicity; | Toxicity; EMS code: 0461 | Chemical |  | LAB
0461 | % (Mortality) | Bio. LC50:Fish | FA05 | Chinook/96hr(%mortality)/Salinity Adj./100% conc. | Chinook/96hr(%mortality)/Salinity Adjusted/100% concentration | 0 | Y | Y | Y | 1 |  |  | 4/22/1997 9:27:57 AM | 1997-02-26 | 6/3/1998 6:51:46 PM | MPAK |  | DimensionlessRatio | % | N | (ratio) | (mortality) | Bio. LC50:Fish (mortality) (ratio) |  | Toxicity; | Toxicity; EMS code: 0461 | Chemical |  | LAB
0461 | %(V/V) | Bio. LC50:Fish | FA15 | Rainbow/96hr/Vol/pH 6.5 adj. | Rainbow/96hr/Vol/pH 6.5 adjusted | 0 | Y | Y | Y | 1 |  |  | 10/12/2006 10:02:31 AM | 1997-03-04 | 6/3/1998 6:51:46 PM | MPAK |  | DimensionlessRatio | % | N | (ratio) | (by vol.) | Bio. LC50:Fish (by vol.) (ratio) |  | Toxicity; | Toxicity; EMS code: 0461 | Chemical |  | LAB
0461 | %(V/V) | Bio. LC50:Fish | X068 | Rainbow/96hr(Vol) (RM13) | (RM13) This method measures the concentration (expressed volume/volume) of sample at which 50% mortality (for LC50), at 20% mortality (for LC20), and at 10% mortality (for LC10) occurs. No pH adjustment is permitted. | 0 | Y | Y | Y | 5512 | 1/16/1996 | 8/2/2024 | 8/23/2024 9:49:36 AM | 1996-04-12 | 6/3/1998 6:51:33 PM | EMS |  | DimensionlessRatio | % | N | (ratio) | (by vol.) | Bio. LC50:Fish (by vol.) (ratio) |  | Toxicity; | Toxicity; EMS code: 0461 | Chemical |  | LAB
0461 | % (Mortality) | Bio. LC50:Fish | FA06 | Coho/96hr(%mortality)/Salinity Adj./100% conc. | Coho/96hr(%mortality)/Salinity Adjusted/100% concentration | 0 | Y | Y | Y | 43 |  |  | 4/22/1997 9:54:43 AM | 1997-02-26 | 6/3/1998 6:51:46 PM | MPAK |  | DimensionlessRatio | % | N | (ratio) | (mortality) | Bio. LC50:Fish (mortality) (ratio) |  | Toxicity; | Toxicity; EMS code: 0461 | Chemical |  | LAB
0461 | % (Mortality) | Bio. LC50:Fish | FA13 | RM13/Rainbow/96hr(%mortality)/100% conc. | Rainbow/96hr(%mortality)/100% concentration. (RM13) Biological Test Method: Reference Method for Determining Acute Lethality of Effluents to Rainbow Trout; EPS 1/RM/13 Second Edition December 2000 (with May 2007 and February 2016 amendments). Rainbow Trout are exposed in a single concentration of the sample and mortality is recorded over a period of 96 hours at 15C. The pass/fail single concentration test uses undiluted material at 100% concentration and determines the percent mortality after 96 hours. (Without pH Adjustment) | 0 | Y | Y | Y | 7750 | 1/4/1997 | 7/10/2024 | 8/23/2024 9:49:36 AM | 1997-02-28 | 6/3/1998 6:51:46 PM | MPAK |  | DimensionlessRatio | % | N | (ratio) | (mortality) | Bio. LC50:Fish (mortality) (ratio) |  | Toxicity; | Toxicity; EMS code: 0461 | Chemical |  | LAB
0461 | %(V/V) | Bio. LC50:Fish | X069 | Rainbow/96hr/Vol/pH 7 adj | Rainbow/96hr/Vol/pH 7 adj | 0 | Y | Y | Y | 1762 | 9/17/1996 | 10/6/2020 | 8/23/2024 8:10:36 AM | 1996-04-12 | 6/3/1998 6:51:33 PM | EMS |  | DimensionlessRatio | % | N | (ratio) | (by vol.) | Bio. LC50:Fish (by vol.) (ratio) |  | Toxicity; | Toxicity; EMS code: 0461 | Chemical |  | LAB
0461 | % (Mortality) | Bio. LC50:Fish | FA17 | Rainbow/96hr(%mortality)pH6.5adj. | Rainbow/96hr(%mortality/pH6.5adj. | 0 | Y | Y | Y | 50 |  |  | 7/20/2005 10:08:43 AM | 1998-07-10 |  | NPEPPIN |  | DimensionlessRatio | % | N | (ratio) | (mortality) | Bio. LC50:Fish (mortality) (ratio) |  | Toxicity; | Toxicity; EMS code: 0461 | Chemical |  | LAB
0462 | mg | Biomass;fixed @ 550C | X485 | Gravimetric (550C) | A measured volume of sample is filtered through a 0.45 um membrane filter or a pre-rinsed, dried and pre-weighed glass fibre filter. The filter is rinsed, removed from the filtration apparatus and dried at 105 C to constant weight, cooled in a desiccator and weighed. The residue is next ignited at 550 C for 1 hour, then cooled, rewetted to restore water of hydration of minerals, re-dried at 105 C to constant weight, cooled in a desiccator and reweighed. Volatile Biomass is the difference between the initial and fixed weights. Intermediate results should be reported in units of mg(i.e., mg for whole sample) but final results should be reported either in mg/L or mg/m2, whichever is appropriate. | 1 | Y | Y | Y | 38 | 12/3/1997 9:59:00 AM | 10/7/1999 8:20:00 AM | 10/20/1999 1:36:17 AM | 1998-01-08 | 6/3/1998 6:51:47 PM | NPEPPIN |  | Mass | mg | Y | (mass) |  | Biomass;fixed @ 550C (mass) |  | Biological Examination; | Biological Examination; EMS code: 0462 | Chemical |  | LAB
0462 | mg | Biomass;fixed @ 550C | EQ03 | EQUIS (mg) | EQUIS (mg) | 1 | N | Y | Y | 366 |  |  | 10/22/1996 11:14:30 PM | 1996-04-12 | 6/3/1998 6:51:33 PM | EMS |  | Mass | mg | Y | (mass) |  | Biomass;fixed @ 550C (mass) |  | Biological Examination; | Biological Examination; EMS code: 0462 | Chemical |  | LAB
0462 | mg | Biomass;fixed @ 550C | 1940 | Gravimetric 550C | For Total Volatile Residue (parm=0032): A known volume of a well-mixed sample is transferred into an pre-ignited porcelain dish, evaporated at 90 - 95 C, dried at 103 C, cooled and weighed. The porcelain dish with dried and weighed residue is then ignited in a muffle at 550 C and weighed again. The difference in weights represents Total Volatile Residue (TVR). The Total Fixed Residue (TFR) is the portion of sample left after ignited at 550 C. | 1 | N | Y | Y | 1077 |  |  | 10/24/1996 3:05:27 PM | 1996-04-12 | 6/3/1998 6:51:42 PM | EMS |  | Mass | mg | Y | (mass) |  | Biomass;fixed @ 550C (mass) |  | Biological Examination; | Biological Examination; EMS code: 0462 | Chemical |  | LAB
0462 | mg/L | Biomass;fixed @ 550C | X237 | Gravimetric | This code is a catch-all for a variety of gravimetric procedures. | 4 | Y | Y | Y | 86 | 10/1/1996 11:41:00 AM | 8/24/1999 4:37:00 PM | 10/20/1999 1:36:58 AM | 1996-04-12 | 6/3/1998 6:51:45 PM | EMS |  | Concentration | mg/L | Y | (fl. conc.) |  | Biomass;fixed @ 550C (fl. conc.) |  | Biological Examination; | Biological Examination; EMS code: 0462 | Chemical |  | LAB
0466 | %(V/V) | Bio. LC20:Fish | X068 | Rainbow/96hr(Vol) (RM13) | (RM13) This method measures the concentration (expressed volume/volume) of sample at which 50% mortality (for LC50), at 20% mortality (for LC20), and at 10% mortality (for LC10) occurs. No pH adjustment is permitted. | 0 | Y | Y | Y | 145 | 1/16/1996 | 1/9/2006 8:00:00 AM | 1/20/2022 4:45:11 PM | 1997-03-17 | 6/3/1998 6:51:46 PM | MPAK |  | DimensionlessRatio | % | N | (ratio) | (by vol.) | Bio. LC20:Fish (by vol.) (ratio) |  | Toxicity; | Toxicity; EMS code: 0466 | Chemical |  | LAB
0466 | % (Mortality) | Bio. LC20:Fish | FA13 | RM13/Rainbow/96hr(%mortality)/100% conc. | Rainbow/96hr(%mortality)/100% concentration. (RM13) Biological Test Method: Reference Method for Determining Acute Lethality of Effluents to Rainbow Trout; EPS 1/RM/13 Second Edition December 2000 (with May 2007 and February 2016 amendments). Rainbow Trout are exposed in a single concentration of the sample and mortality is recorded over a period of 96 hours at 15C. The pass/fail single concentration test uses undiluted material at 100% concentration and determines the percent mortality after 96 hours. (Without pH Adjustment) | 0 | Y | Y | Y | 749 | 9/22/1999 | 12/6/2022 | 7/25/2024 12:46:40 PM | 1997-03-14 | 6/3/1998 6:51:46 PM | MPAK |  | DimensionlessRatio | % | N | (ratio) | (mortality) | Bio. LC20:Fish (mortality) (ratio) |  | Toxicity; | Toxicity; EMS code: 0466 | Chemical |  | LAB
0470 | adt/d | Production Quant. | X464 | 90th%tile,Bleachd&Unblea | 90th%tile,Bleachd&Unblea | 0 | Y | Y | Y | 282 | 1/31/1996 7:00:00 AM | 2/4/1999 7:00:00 AM | 7/16/2024 3:11:23 PM | 1996-04-12 | 6/3/1998 6:51:42 PM | EMS |  | AirDriedProductionRate | adt/d | Y | (atd prod.) |  | Production Quant. (atd prod.) |  | Physical Properties; | Physical Properties; EMS code: 0470 | Physical |  | LAB
0470 | t | Production Quant. | XRPR | Fed,Annual Ref.Prod.Rate | Fed;Annual Ref.Prod.Rate | 0 | Y | Y | Y | 187 |  |  | 7/16/2024 3:11:23 PM | 1996-04-12 | 6/3/1998 6:51:42 PM | EMS |  | Mass | t | Y | (mass) |  | Production Quant. (mass) |  | Physical Properties; | Physical Properties; EMS code: 0470 | Physical |  | LAB
0470 | adt/d | Production Quant. | X014 | Pulp-Unbleach Air Dried | Pulp-Unbleach Air Dried | 0 | Y | Y | Y | 15448 | 5/7/1996 8:00:00 AM | 6/5/1996 8:00:00 AM | 7/16/2024 3:11:25 PM | 1996-04-12 | 6/3/1998 6:51:33 PM | EMS |  | AirDriedProductionRate | adt/d | Y | (atd prod.) |  | Production Quant. (atd prod.) |  | Physical Properties; | Physical Properties; EMS code: 0470 | Physical |  | LAB
0470 | m3/d | Production Quant. | X293 | Woodmill prod-beltweight | Woodmill prod-beltweight | 0 | Y | Y | Y | 142 |  |  | 10/23/1996 11:12:13 PM | 1996-04-12 | 6/3/1998 6:51:33 PM | EMS |  | Flow | m³/d | Y | (flow) |  | Production Quant. (flow) |  | Physical Properties; | Physical Properties; EMS code: 0470 | Physical |  | LAB
0470 | t/d | Production Quant. | X052 | Bleached Pulp-Air Dried | Bleached Pulp-Air Dried | 0 | Y | Y | Y | 23243 | 10/2/1996 | 3/1/2016 | 8/22/2024 1:49:14 PM | 1996-04-12 | 6/3/1998 6:51:33 PM | EMS |  | MassFlow | t/d | Y | (mass flow) |  | Production Quant. (mass flow) |  | Physical Properties; | Physical Properties; EMS code: 0470 | Physical |  | LAB
0470 | adt/d | Production Quant. | X037 | Pulp & Paper-Air Dried | Pulp & Paper-Air Dried | 0 | Y | Y | Y | 122623 | 9/2/1996 | 7/1/2024 | 8/20/2024 8:39:43 AM | 1996-04-12 | 6/3/1998 6:51:46 PM | EMS |  | AirDriedProductionRate | adt/d | Y | (atd prod.) |  | Production Quant. (atd prod.) |  | Physical Properties; | Physical Properties; EMS code: 0470 | Physical |  | LAB
0470 | t/d | Production Quant. | EQ30 | EQUIS (T/d) | EQUIS (T/d) | 0 | N | Y | Y | 2157 |  |  | 10/23/1996 11:12:44 PM | 1996-04-12 | 6/3/1998 6:51:33 PM | EMS |  | MassFlow | t/d | Y | (mass flow) |  | Production Quant. (mass flow) |  | Physical Properties; | Physical Properties; EMS code: 0470 | Physical |  | LAB
0470 | adt/d | Production Quant. | X463 | 90th%tile,Bleachd | 90th%tile,Bleachd | 0 | Y | Y | Y | 426 | 9/26/1996 7:00:00 AM | 2/4/1999 7:00:00 AM | 10/12/2023 11:44:29 AM | 1996-04-12 | 6/3/1998 6:51:42 PM | EMS |  | AirDriedProductionRate | adt/d | Y | (atd prod.) |  | Production Quant. (atd prod.) |  | Physical Properties; | Physical Properties; EMS code: 0470 | Physical |  | LAB
0470 | t | Production Quant. | EQ29 | EQUIS (T) | EQUIS (T) | 0 | N | Y | Y | 1 |  |  | 10/22/1996 7:13:51 PM | 1996-04-12 | 6/3/1998 6:51:33 PM | EMS |  | Mass | t | Y | (mass) |  | Production Quant. (mass) |  | Physical Properties; | Physical Properties; EMS code: 0470 | Physical |  | LAB
0470 | t/d | Production Quant. | X259 | Fish | Fish | 0 | Y | Y | Y | 8 |  |  | 10/24/1996 2:31:29 AM | 1996-04-12 | 6/3/1998 6:51:33 PM | EMS |  | MassFlow | t/d | Y | (mass flow) |  | Production Quant. (mass flow) |  | Physical Properties; | Physical Properties; EMS code: 0470 | Physical |  | LAB
0470 | adt/d | Production Quant. | X473 | Bleached Pulp | Bleached Pulp | 0 | Y | Y | Y | 61412 | 1/2/1996 8:00:00 AM | 2/25/2016 8:00:00 AM | 8/20/2024 8:39:43 AM | 1997-08-20 | 6/3/1998 6:51:47 PM | NPEPPIN |  | AirDriedProductionRate | adt/d | Y | (atd prod.) |  | Production Quant. (atd prod.) |  | Physical Properties; | Physical Properties; EMS code: 0470 | Physical |  | LAB
0470 | m3/min A | Production Quant. | EQ27 | EQUIS (m3/min A) | EQUIS (m3/min A) | 0 | N | Y | Y | 1127 |  |  | 10/22/1996 7:14:16 PM | 1996-04-12 | 6/3/1998 6:51:33 PM | EMS |  | Flow | m³/min | Y | (flow) |  | Production Quant. (flow) |  | Physical Properties; | Physical Properties; EMS code: 0470 | Physical |  | LAB
0471 | t/d | Waste Quantity | EQ30 | EQUIS (T/d) | EQUIS (T/d) | 0 | N | Y | Y | 39 |  |  | 10/23/1996 12:02:04 AM | 1996-04-12 | 6/3/1998 6:51:33 PM | EMS |  | MassFlow | t/d | Y | (mass flow) |  | Waste Quantity (mass flow) |  | Physical Properties; | Physical Properties; EMS code: 0471 | Physical |  | LAB
0471 | t/d | Waste Quantity | X292 | Weight/time | Weight/time | 0 | Y | Y | Y | 2460 | 12/8/2000 12:00:00 PM | 11/4/2003 | 7/18/2024 8:15:02 AM | 1996-04-12 | 6/3/1998 6:51:33 PM | EMS |  | MassFlow | t/d | Y | (mass flow) |  | Waste Quantity (mass flow) |  | Physical Properties; | Physical Properties; EMS code: 0471 | Physical |  | LAB
0474 | adt/d | 90th %tile; (Unbleached Production) | X489 | 90th%,Unbleach | 90th%tile, Unbleached | 0 | Y | Y | Y | 38 |  |  | 2/11/2020 9:48:37 AM | 1998-11-09 |  | NPEPPIN |  | AirDriedProductionRate | adt/d | Y | (atd prod.) |  | 90th %tile; (Unbleached Production) (atd prod.) |  | Physical Properties; | Physical Properties; EMS code: 0474 | Physical |  | LAB
0475 | adt/d | 90th %tile (Production) | X464 | 90th%tile,Bleachd&Unblea | 90th%tile,Bleachd&Unblea | 0 | Y | Y | Y | 550 | 1/31/1998 | 6/30/2024 7:00:00 AM | 7/29/2024 11:17:51 AM | 1998-10-07 |  | NPEPPIN |  | AirDriedProductionRate | adt/d | Y | (atd prod.) |  | 90th %tile (Production) (atd prod.) |  | Physical Properties; | Physical Properties; EMS code: 0475 | Physical |  | LAB
0476 | adt/d | 90th %tile (CB Prod) | X463 | 90th%tile,Bleachd | 90th%tile,Bleachd | 0 | Y | Y | Y | 771 | 1/31/1998 | 6/30/2024 7:00:00 AM | 7/29/2024 11:17:51 AM | 1998-10-07 |  | NPEPPIN |  | AirDriedProductionRate | adt/d | Y | (atd prod.) |  | 90th %tile (CB Prod) (atd prod.) |  | Physical Properties; | Physical Properties; EMS code: 0476 | Physical |  | LAB
0477 | m3/d | 90th %tile Effluent | X462 | 90th percentile | 90th percentile | 0 | Y | Y | Y | 739 | 1/31/1998 | 6/30/2024 7:00:00 AM | 7/29/2024 11:17:51 AM | 1998-10-07 | 11/2/1998 9:33:44 AM | NPEPPIN | NPEPPIN | Flow | m³/d | Y | (flow) |  | 90th %tile Effluent (flow) |  | Physical Properties; | Physical Properties; EMS code: 0477 | Physical |  | LAB
0660 | m3/min A | Discharge Rate | EQ27 | EQUIS (m3/min A) | EQUIS (m3/min A) | 0 | N | Y | Y | 5594 |  |  | 10/23/1996 12:01:37 AM | 1996-04-12 | 6/3/1998 6:51:33 PM | EMS |  | Flow | m³/min | Y | (flow) |  | Discharge Rate (flow) |  | Physical Properties; | Physical Properties; EMS code: 0660 | Physical |  | LAB
0660 | m3/d | Discharge Rate | EQ35 | EQUIS (m3/d) | EQUIS (m3/d) | 0 | N | Y | Y | 36 |  |  | 10/22/1996 9:14:49 PM | 1996-04-12 | 6/3/1998 6:51:33 PM | EMS |  | Flow | m³/d | Y | (flow) |  | Discharge Rate (flow) |  | Physical Properties; | Physical Properties; EMS code: 0660 | Physical |  | LAB
0X03 | ug/g | Oxychlordane | ZN03 | DCM+Ac.Extr:GC/MS | DCM+Ac.Extr:GC/MS EPA 8270E (mod) | 0.005 | Y | Y | Y | 2 | 11/14/2022 | 11/14/2022 | 12/15/2022 10:50:04 AM | 2022-06-23 |  | EMS |  | MassPerMass | mg/kg | Y | (mass conc.) |  | Oxychlordane (mass conc.) | 27304-13-8 | Pesticides; | Pesticides; EMS code: 0X03 | Chemical |  | LAB
0X03 | ug/g | Oxychlordane | E8081 | EPA 8081 | See methods EPA 3540, 3630, 3660, 8081 | 0.001 | Y | Y | Y | 9 | 3/18/2020 | 3/23/2020 | 2/22/2023 12:51:03 PM | 2022-07-29 |  | EMS |  | MassPerMass | mg/kg | Y | (mass conc.) |  | Oxychlordane (mass conc.) | 27304-13-8 | Pesticides; | Pesticides; EMS code: 0X03 | Chemical |  | LAB
0X03 | mg/L | Oxychlordane | X268 | CH2Cl2 ext:GC/MS | CH2Cl2 ext:GC/MS |  | Y | Y | Y | 16 | 12/29/2021 | 9/2/2022 | 11/10/2022 9:54:06 AM | 2022-04-21 |  | EMS |  | Concentration | mg/L | Y | (fl. conc.) |  | Oxychlordane (fl. conc.) | 27304-13-8 | Pesticides; | Pesticides; EMS code: 0X03 | Chemical |  | LAB
0X03 | mg/L | Oxychlordane | E660 | EPA 8270E (mod) | EPA 8270E (mod) Miscellaneous Pesticides by GC-MS | 2e-06 | Y | Y | Y | 2 | 9/14/2022 | 9/14/2022 | 12/8/2022 9:48:13 AM | 2022-12-06 |  | EMS |  | Concentration | mg/L | Y | (fl. conc.) |  | Oxychlordane (fl. conc.) | 27304-13-8 | Pesticides; | Pesticides; EMS code: 0X03 | Chemical |  | LAB
0X03 | mg/L | Oxychlordane | MW03 | OCP1-LL-SF-ECD-VA | This analysis is carried out using procedures adapted from "Test Methods for Evaluating Solid Waste" SW-846, Methods 3510, 3630, 3660 and 8081, published by the United States Environmental Protection Agency (EPA). The procedure involves extraction of the entire water sample with dichloromethane. The extract then undergoes one or more of the following clean-up procedures (if required): silica gel clean-up and/or sulphur clean-up. The final extract is analyzed by dual capillary column gas chromatography with electron capture detection (GC/ECD). | 5e-05 | Y | Y | Y | 10 | 9/20/2016 | 9/21/2016 | 8/15/2019 2:42:34 PM | 2019-07-08 | 8/15/2019 11:31:19 AM | KATFRASE | KATFRASE | Concentration | mg/L | Y | (fl. conc.) |  | Oxychlordane (fl. conc.) | 27304-13-8 | Pesticides; | Pesticides; EMS code: 0X03 | Chemical |  | LAB
1016 | mg/L | Chlorine Res:Total | X103 | DPD Colorimetric | Liberated iodine reacts with DPD to produce a red solution which is measured spectrophotometrically at 515 nm. | 0.2 | Y | Y | Y | 7472 | 1/7/1997 | 8/10/2024 | 8/21/2024 11:21:21 AM | 1996-04-12 | 7/15/1999 10:17:30 AM | EMS | NPEPPIN | Concentration | mg/L | Y | (fl. conc.) |  | Chlorine Res:Total (fl. conc.) | 7782-50-5 | Inorganics; | Inorganics; EMS code: 1016 | Chemical | Total | LAB
1016 | mg/L | Chlorine Res:Total | X006 | Amperometric | Amperometric | 0 | Y | Y | Y | 5943 | 1/1/2003 8:00:00 AM | 3/30/2018 8:00:00 AM | 4/27/2018 10:34:00 AM | 1996-04-12 | 3/18/2003 9:38:19 AM | EMS | UMULLER | Concentration | mg/L | Y | (fl. conc.) |  | Chlorine Res:Total (fl. conc.) | 7782-50-5 | Inorganics; | Inorganics; EMS code: 1016 | Chemical | Total | LAB
1016 | mg/L | Chlorine Res:Total | EQ01 | EQUIS (mg/L) | EQUIS (mg/L) | 0 | N | Y | Y | 13 |  |  | 10/22/1996 10:04:54 PM | 1996-04-12 | 7/15/1999 10:17:30 AM | EMS | NPEPPIN | Concentration | mg/L | Y | (fl. conc.) |  | Chlorine Res:Total (fl. conc.) | 7782-50-5 | Inorganics; | Inorganics; EMS code: 1016 | Chemical | Total | LAB
1016 | mg/L | Chlorine Res:Total | X004 | Colorimetric Titr'n | Manual Colorimetric Titration with standard solutions to specified endpoint as identified by colorimetric indicator. The sample must not be filtered, diluted, concentrated, or altered in any way. | 0 | Y | Y | Y | 611 | 6/28/2011 2:05:00 PM | 7/31/2024 9:38:00 AM | 8/6/2024 5:01:06 PM | 1996-04-12 | 6/3/1998 6:51:33 PM | EMS |  | Concentration | mg/L | Y | (fl. conc.) |  | Chlorine Res:Total (fl. conc.) | 7782-50-5 | Inorganics; | Inorganics; EMS code: 1016 | Chemical | Total | LAB
1016 | mg/L | Chlorine Res:Total | X002 | Hach | Hach | 0 | N | Y | Y | 1793 |  |  | 10/24/1996 12:35:58 PM | 1996-04-12 | 6/3/1998 6:51:33 PM | EMS |  | Concentration | mg/L | Y | (fl. conc.) |  | Chlorine Res:Total (fl. conc.) | 7782-50-5 | Inorganics; | Inorganics; EMS code: 1016 | Chemical | Total | LAB
1016 | mg/L | Chlorine Res:Total | X317 | Iodometric titr'n amperometric endpoint | Known amount of reducing agent (usually phenylarsine oxide) is added, and the excess reducing agent after reaction with chlorine residual is determined by iodometric titration, amperiometric endpoint. | 0 | Y | Y | Y | 3 |  |  | 2/2/1999 1:39:50 PM | 1996-04-12 | 6/3/1998 6:51:33 PM | EMS |  | Concentration | mg/L | Y | (fl. conc.) |  | Chlorine Res:Total (fl. conc.) | 7782-50-5 | Inorganics; | Inorganics; EMS code: 1016 | Chemical | Total | LAB
1018 | m3/s W | Flow (Rate) | X135 | Estimation | Estimation | 0 | Y | Y | Y | 27002 |  |  | 5/19/2015 3:10:37 PM | 1996-04-12 | 6/3/1998 6:51:45 PM | EMS |  | Flow | m³/s | Y | (flow) |  | Flow (Rate) (flow) |  | Physical Properties; | Physical Properties; EMS code: 1018 | Physical |  | LAB
1018 | m3/s W | Flow (Rate) | XG01 | Parshall Flume | Parshall Flume | 0 | Y | Y | Y | 457 |  |  | 10/23/1996 11:02:52 PM | 1996-04-12 | 6/3/1998 6:51:33 PM | EMS |  | Flow | m³/s | Y | (flow) |  | Flow (Rate) (flow) |  | Physical Properties; | Physical Properties; EMS code: 1018 | Physical |  | LAB
1018 | m3/s W | Flow (Rate) | XM00 | Meter-in situ | Meter-in situ | 0 | Y | Y | Y | 92 |  |  | 5/27/2024 9:55:15 AM | 1996-04-12 | 6/3/1998 6:51:45 PM | EMS |  | Flow | m³/s | Y | (flow) |  | Flow (Rate) (flow) |  | Physical Properties; | Physical Properties; EMS code: 1018 | Physical |  | LAB
1018 | m3/s W | Flow (Rate) | XG00 | Stream Gauge | Stream Gauge | 0 | Y | Y | Y | 1002 |  |  | 11/16/2004 4:00:39 PM | 1996-04-12 | 6/3/1998 6:51:33 PM | EMS |  | Flow | m³/s | Y | (flow) |  | Flow (Rate) (flow) |  | Physical Properties; | Physical Properties; EMS code: 1018 | Physical |  | LAB
1018 | m3/s W | Flow (Rate) | EQ26 | EQUIS (m3/s W) | EQUIS (m3/s W) | 0 | N | Y | Y | 6574 |  |  | 10/23/1996 12:07:11 AM | 1996-04-12 | 6/3/1998 6:51:33 PM | EMS |  | Flow | m³/s | Y | (flow) |  | Flow (Rate) (flow) |  | Physical Properties; | Physical Properties; EMS code: 1018 | Physical |  | LAB
1018 | m3/s W | Flow (Rate) | XG02 | V Notch Weir | V Notch Weir | 0 | Y | Y | Y | 25553 | 8/4/1998 9:50:00 AM | 12/7/1999 3:20:00 AM | 5/27/2024 9:55:15 AM | 1996-04-12 | 6/3/1998 6:51:33 PM | EMS |  | Flow | m³/s | Y | (flow) |  | Flow (Rate) (flow) |  | Physical Properties; | Physical Properties; EMS code: 1018 | Physical |  | LAB
1018 | m3/s W | Flow (Rate) | XM11 | Flow Meter | Flow Meter | 0 | Y | Y | Y | 45104 | 10/1/2022 | 6/30/2024 | 8/14/2024 11:49:28 AM | 1996-04-12 | 6/3/1998 6:51:33 PM | EMS |  | Flow | m³/s | Y | (flow) |  | Flow (Rate) (flow) |  | Physical Properties; | Physical Properties; EMS code: 1018 | Physical |  | LAB
1018 | m3/s W | Flow (Rate) | XM04 | Totalizer | Totalizer | 0 | Y | Y | Y | 18835 |  |  | 5/19/2015 3:10:30 PM | 1996-04-12 | 6/3/1998 6:51:33 PM | EMS |  | Flow | m³/s | Y | (flow) |  | Flow (Rate) (flow) |  | Physical Properties; | Physical Properties; EMS code: 1018 | Physical |  | LAB
1052 | SWU | Color S./W. Diss. | X010 | Filt,pH,Spec,465nm | A Single Wavelength Colour of a sample is determined by measuring the absorbance of the filtered sample at 465 nm after filtering and adjustment of pH to 7.6. One unit of SWC colour may be defined as the colour produced by 1 mg Pt/L in the form of choloroplatinate ion. Measurements at 466 nm are considered equivalent. | 5 | Y | Y | Y | 23743 | 12/10/1996 | 2/3/2016 | 8/22/2024 1:49:13 PM | 1996-04-12 | 6/3/1998 6:51:33 PM | EMS |  | Apperance | SWU | N | (app.) |  | Color S./W. Diss. (app.) |  | Physical Properties; | Physical Properties; EMS code: 1052 | Physical |  | LAB
1052 | SWU | Color S./W. Diss. | 1320 | Single wavelength,no pH adj | Colour is determined by spectrophotometric comparison of the sample, at a specific wavelength (400nm), with known concentrations of coloured solutions. Any turbidity must be removed. The pH of the sample must not be adjusted but should be determined and reported as well. | 5 | Y | Y | Y | 2351 | 9/24/1996 2:08:00 PM | 5/21/2015 5:28:00 PM | 5/27/2015 4:34:36 PM | 1996-04-12 | 6/3/1998 6:51:33 PM | EMS |  | Apperance | SWU | N | (app.) |  | Color S./W. Diss. (app.) |  | Physical Properties; | Physical Properties; EMS code: 1052 | Physical |  | LAB
1052 | SWU | Color S./W. Diss. | EQ09 | EQUIS (SWU) | EQUIS (SWU) | 5 | N | Y | Y | 2104 |  |  | 10/22/1996 11:42:08 PM | 1996-04-12 | 6/3/1998 6:51:33 PM | EMS |  | Apperance | SWU | N | (app.) |  | Color S./W. Diss. (app.) |  | Physical Properties; | Physical Properties; EMS code: 1052 | Physical |  | LAB
1101 | mg/L | Bromide Dissolved | XC01 | Ion Chromatography | Same as X044, but higher MDL. A clear (Total) or filtered (Dissolved) water sample is injected into a stream of carbonate-bicarbonate eluent and passed through a series of ion exchanger. The anions of interested are separated on the basis of their relative affinities for a low capacity , strongly basic anion exchanger (guard and separate columns). The separated anions are directed through the anion self-regenerating suppressor (ASRS). Hydronium ions in the suppressor will neutralize the eluent to its weakly ionized form and convert the separated anions in their conductive acid form. The separated anions are measured by conductivity. They are identified on the basis of retention time as compared to standards. Quantitation is by measurement of peak area of peak height against a calibration curve generated from known standards. | 0.4 | Y | Y | Y | 1044 | 7/14/2009 9:14:00 AM | 12/23/2011 2:18:00 PM | 1/24/2013 8:46:56 AM | 2009-07-22 | 7/22/2009 12:37:40 PM | RDGIBSON | RDGIBSON | Concentration | mg/L | Y | (fl. conc.) |  | Bromide Dissolved (fl. conc.) | 24959-67-9 | Metals - Dissolved; | Metals - Dissolved; EMS code: 1101 | Chemical | Dissolved | LAB
1101 | mg/L | Bromide Dissolved | 1342 | Bromide by IC (seawater) | Bromide by IC (seawater) | 5 | Y | Y | Y | 85 | 8/14/2019 | 4/25/2024 | 6/7/2024 10:03:16 AM | 2019-09-30 |  | KATFRASE |  | Concentration | mg/L | Y | (fl. conc.) |  | Bromide Dissolved (fl. conc.) | 24959-67-9 | Metals - Dissolved; | Metals - Dissolved; EMS code: 1101 | Chemical | Dissolved | LAB
1101 | mg/L | Bromide Dissolved | X044 | Ion Chromatography | A clear (Total) or filtered (Dissolved) water sample is injected into a stream of carbonate-bicarbonate eluent and passed through a series of ion exchanger. The anions of interested are separated on the basis of their relative affinities for a low capacity , strongly basic anion exchanger (guard and separate columns). The separated anions are directed through the anion self-regenerating suppressor (ASRS). Hydronium ions in the suppressor will neutralize the eluent to its weakly ionized form and convert the separated anions in their conductive acid form. The separated anions are measured by conductivity. They are identified on the basis of retention time as compared to standards. Quantitation is by measurement of peak area of peak height against a calibration curve generated from known standards. | 0.05 | Y | Y | Y | 91163 | 4/10/1996 6:56:00 AM | 8/21/2024 | 8/26/2024 4:53:05 PM | 1997-01-16 | 6/3/1998 6:51:46 PM | MPAK |  | Concentration | mg/L | Y | (fl. conc.) |  | Bromide Dissolved (fl. conc.) | 24959-67-9 | Metals - Dissolved; | Metals - Dissolved; EMS code: 1101 | Chemical | Dissolved | LAB
1101 | mg/L | Bromide Dissolved | 1340 | Auto.I.S.E. | Bromide is analyzed potentiometrically using a bromide Automated Ion Specific Electrode in conjunction with a standard reference electrode and a pH meter. The potential developed by the presence of bromide ions is measured by an expanded pH meter. | 0.01 | Y | Y | Y | 1013 | 10/29/2008 1:43:00 PM | 12/10/2010 1:52:00 PM | 7/8/2019 4:12:12 PM | 2008-11-12 | 11/12/2008 2:00:04 PM | JCHUDAK | JCHUDAK | Concentration | mg/L | Y | (fl. conc.) |  | Bromide Dissolved (fl. conc.) | 24959-67-9 | Metals - Dissolved; | Metals - Dissolved; EMS code: 1101 | Chemical | Dissolved | LAB
1102 | mg/L | Alkalinity Diss. 4.5 | EQ01 | EQUIS (mg/L) | EQUIS (mg/L) | 0.5 | N | Y | Y | 150 |  |  | 10/22/1996 8:39:06 PM | 1996-04-12 | 6/3/1998 6:51:33 PM | EMS |  | Concentration | mg/L | Y | (fl. conc.) |  | Alkalinity Diss. 4.5 (fl. conc.) |  | Major Ions; | Major Ions; EMS code: 1102 | Chemical |  | LAB
1102 | mg/L | Alkalinity Diss. 4.5 | X330 | Meter | The conductance (= conductivity) of a sample is measured by use of a self-contained conductivity meter. pH is measured with a combined glass/reference (Ag/AgCl) electrode. For both conductivity and pH the measurements are made at 25 C, or the results are temperature-corrected to 25 C. Note that laboratory and field measurements are distinguished by the EMS Analytical Agency codes. Ex situ and in situ measurements are distinguished by the computer codes. | 0.5 | Y | Y | Y | 10 |  |  | 5/26/1999 9:41:29 AM | 1996-04-12 | 6/3/1998 6:51:33 PM | EMS |  | Concentration | mg/L | Y | (fl. conc.) |  | Alkalinity Diss. 4.5 (fl. conc.) |  | Major Ions; | Major Ions; EMS code: 1102 | Chemical |  | LAB
1103 | mg/L | Carbon Dissolved Organic | X067 | TOC Analyzer | Organic carbon is converted to carbon dioxide (CO2) by catalytic combustion or wet chemistry oxidation. The CO2 is then measured by infrared detector. | 1 | Y | Y | Y | 66436 | 4/29/2002 | 8/21/2024 | 8/26/2024 4:53:05 PM | 2003-03-18 | 3/18/2003 2:40:28 PM | UMULLER | UMULLER | Concentration | mg/L | Y | (fl. conc.) |  | Carbon Dissolved Organic (fl. conc.) |  | Carbon; | Carbon; EMS code: 1103 | Chemical | Dissolved | LAB
1103 | mg/L | Carbon Dissolved Organic | X314 | Dig:FID TOC Dohrmann | Organic carbon is converted to carbon dioxide (CO2) by catalytic combustion or wet chemistry oxidation. The CO2 is then converted to methane (CH4) and measured by flame ionization detector. | 1 | Y | Y | Y | 4964 |  |  | 8/26/2024 2:41:32 PM | 1996-04-12 | 6/3/1998 6:51:45 PM | EMS |  | Concentration | mg/L | Y | (fl. conc.) |  | Carbon Dissolved Organic (fl. conc.) |  | Carbon; | Carbon; EMS code: 1103 | Chemical | Dissolved | LAB
1103 | mg/L | Carbon Dissolved Organic | 0914 | Colorimetric (DOC/TOC) | Colorimetric (DOC/TOC) | 1 | Y | Y | Y | 22852 | 3/2/2002 | 8/13/2024 8:48:00 PM | 8/20/2024 11:38:02 AM | 1996-04-12 | 6/3/1998 6:51:45 PM | EMS |  | Concentration | mg/L | Y | (fl. conc.) |  | Carbon Dissolved Organic (fl. conc.) |  | Carbon; | Carbon; EMS code: 1103 | Chemical | Dissolved | LAB
1103 | mg/L | Carbon Dissolved Organic | XCAL | Calculation (mg/L) | Calculation (mg/L) | 0.5 | Y | Y | Y | 9426 | 7/31/1996 12:36:00 PM | 8/21/2007 1:54:00 PM | 12/19/2022 12:39:07 PM | 1996-04-12 | 6/3/1998 6:51:41 PM | EMS |  | Concentration | mg/L | Y | (fl. conc.) |  | Carbon Dissolved Organic (fl. conc.) |  | Carbon; | Carbon; EMS code: 1103 | Chemical | Dissolved | LAB
1103 | mg/L | Carbon Dissolved Organic | EQ01 | EQUIS (mg/L) | EQUIS (mg/L) | 1 | N | Y | Y | 8 |  |  | 10/22/1996 9:28:27 PM | 1996-04-12 | 6/3/1998 6:51:33 PM | EMS |  | Concentration | mg/L | Y | (fl. conc.) |  | Carbon Dissolved Organic (fl. conc.) |  | Carbon; | Carbon; EMS code: 1103 | Chemical | Dissolved | LAB
1104 | mg/L | Chloride Dissolved | EQ01 | EQUIS (mg/L) | EQUIS (mg/L) | 0.5 | N | Y | Y | 8801 |  |  | 10/24/1996 3:16:54 PM | 1996-04-12 | 6/3/1998 6:51:33 PM | EMS |  | Concentration | mg/L | Y | (fl. conc.) |  | Chloride Dissolved (fl. conc.) | 16887-00-6 | Major Ions; | Major Ions; EMS code: 1104 | Chemical | Dissolved | LAB
1104 | mg/L | Chloride Dissolved | X316 | Auto colorimetric, HgSCN | Thiocyanate ion (SCN) is liberated from mercuric thiocyanate through sequestration of mercury by chloride ion to form un-ionized mercuric chloride. Ferric nitrate reagent provides ferric ion which, with SCN, forms highly coloured ferric thiocyanate in a concentration proportional to the original chloride concentration. The instrument range of 0.5 to 50 mg/L may be extended with sample dilution. Where particulate matter is present, the sample must be filtered or centrifuged prior to the determination. | 0.5 | Y | Y | Y | 34650 | 6/22/2006 | 8/14/2024 2:09:00 PM | 8/26/2024 8:06:03 AM | 1996-04-12 | 6/3/1998 6:51:33 PM | EMS |  | Concentration | mg/L | Y | (fl. conc.) |  | Chloride Dissolved (fl. conc.) | 16887-00-6 | Major Ions; | Major Ions; EMS code: 1104 | Chemical | Dissolved | LAB
1104 | mg/L | Chloride Dissolved | 1343 | Chloride by IC (seawater) | Chloride by IC (seawater) | 50 | Y | Y | Y | 90 | 8/14/2019 | 4/27/2024 | 6/7/2024 10:03:16 AM | 2019-09-30 |  | KATFRASE |  | Concentration | mg/L | Y | (fl. conc.) |  | Chloride Dissolved (fl. conc.) | 16887-00-6 | Major Ions; | Major Ions; EMS code: 1104 | Chemical | Dissolved | LAB
1104 | mg/L | Chloride Dissolved | 1330 | Auto. mercuric thiocyan. | Auto. mercuric thiocyan. | 0.5 | Y | Y | Y | 12367 | 9/20/1996 | 12/4/1996 | 4/16/2005 7:35:13 PM | 1996-04-12 | 6/3/1998 6:51:46 PM | EMS |  | Concentration | mg/L | Y | (fl. conc.) |  | Chloride Dissolved (fl. conc.) | 16887-00-6 | Major Ions; | Major Ions; EMS code: 1104 | Chemical | Dissolved | LAB
1104 | ug/g | Chloride Dissolved | CLPST | Chloride in soil (paste) by IC | A soil extract produced by the saturated paste extraction procedure is analyzed for chloride by Ion Chromatography with conductivity detection. | 12 | Y | Y | Y | 14 | 10/31/2022 | 10/31/2023 | 11/23/2023 7:26:23 AM | 2018-01-09 |  | JKROGH |  | MassPerMass | mg/kg | Y | (mass conc.) |  | Chloride Dissolved (mass conc.) | 16887-00-6 | Major Ions; | Major Ions; EMS code: 1104 | Chemical | Dissolved | LAB
1104 | mg/L | Chloride Dissolved | X119 | AgNO3 Titration | AgNO3 Titration | 0.5 | Y | Y | Y | 1143 | 3/1/2000 | 12/1/2010 | 2/10/2017 9:37:40 AM | 1996-04-12 | 6/3/1998 6:51:45 PM | EMS |  | Concentration | mg/L | Y | (fl. conc.) |  | Chloride Dissolved (fl. conc.) | 16887-00-6 | Major Ions; | Major Ions; EMS code: 1104 | Chemical | Dissolved | LAB
1104 | mg/L | Chloride Dissolved | X044 | Ion Chromatography | A clear (Total) or filtered (Dissolved) water sample is injected into a stream of carbonate-bicarbonate eluent and passed through a series of ion exchanger. The anions of interested are separated on the basis of their relative affinities for a low capacity , strongly basic anion exchanger (guard and separate columns). The separated anions are directed through the anion self-regenerating suppressor (ASRS). Hydronium ions in the suppressor will neutralize the eluent to its weakly ionized form and convert the separated anions in their conductive acid form. The separated anions are measured by conductivity. They are identified on the basis of retention time as compared to standards. Quantitation is by measurement of peak area of peak height against a calibration curve generated from known standards. | 0.05 | Y | Y | Y | 132152 | 9/13/1994 | 8/21/2024 | 8/26/2024 4:53:05 PM | 1996-10-13 | 8/15/2019 11:48:40 AM | EMS | KATFRASE | Concentration | mg/L | Y | (fl. conc.) |  | Chloride Dissolved (fl. conc.) | 16887-00-6 | Major Ions; | Major Ions; EMS code: 1104 | Chemical | Dissolved | LAB
1104 | ug/g | Chloride Dissolved | 15SW | Colourimetric or IC | Colourimetric or IC | 2.5 | Y | Y | Y | 2 | 10/18/1996 2:17:00 PM | 10/18/1996 2:21:00 PM | 8/14/1997 4:59:37 PM | 1996-04-12 | 6/3/1998 6:51:33 PM | EMS |  | MassPerMass | mg/kg | Y | (mass conc.) |  | Chloride Dissolved (mass conc.) | 16887-00-6 | Major Ions; | Major Ions; EMS code: 1104 | Chemical | Dissolved | LAB
1104 | mg/L | Chloride Dissolved | X276 | Ferric Thiocyan:Col'm | Ferric Thiocyan:Col'm | 1 | Y | Y | Y | 29 | 1/8/2015 | 8/6/2024 | 8/21/2024 11:21:21 AM | 2015-04-15 |  | RDGIBSON |  | Concentration | mg/L | Y | (fl. conc.) |  | Chloride Dissolved (fl. conc.) | 16887-00-6 | Major Ions; | Major Ions; EMS code: 1104 | Chemical | Dissolved | LAB
1104 | mg/L | Chloride Dissolved | X131 | SCN,Colorimetric | SCN;Colorimetric | 0.5 | Y | Y | Y | 1270 | 1/11/2007 2:20:00 PM | 4/18/2007 9:55:00 AM | 4/8/2012 3:42:07 PM | 1996-04-12 | 6/3/1998 6:51:33 PM | EMS |  | Concentration | mg/L | Y | (fl. conc.) |  | Chloride Dissolved (fl. conc.) | 16887-00-6 | Major Ions; | Major Ions; EMS code: 1104 | Chemical | Dissolved | LAB
1105 | mg/L | Cyanide Dissolved | X091 | Decant,NaOH,Pyr-Pyr,Col | Decant;NaOH;Pyr-Pyr;Col | 0.01 | Y | Y | Y | 4 |  |  | 10/23/1996 9:56:55 PM | 1996-04-12 | 6/3/1998 6:51:33 PM | EMS |  | Concentration | mg/L | Y | (fl. conc.) |  | Cyanide Dissolved (fl. conc.) | 57-12-5 | Inorganics; | Inorganics; EMS code: 1105 | Chemical | Dissolved | LAB
1105 | mg/L | Cyanide Dissolved | EQ01 | EQUIS (mg/L) | EQUIS (mg/L) | 0.01 | N | Y | Y | 1004 |  |  | 10/24/1996 11:35:58 AM | 1996-04-12 | 6/3/1998 6:51:33 PM | EMS |  | Concentration | mg/L | Y | (fl. conc.) |  | Cyanide Dissolved (fl. conc.) | 57-12-5 | Inorganics; | Inorganics; EMS code: 1105 | Chemical | Dissolved | LAB
1106 | mg/L | Fluoride Dissolved | 1344 | Fluoride by IC (seawater) | Fluoride by IC (seawater) | 1 | Y | Y | Y | 85 | 8/14/2019 | 4/25/2024 | 6/7/2024 10:03:16 AM | 2019-09-30 |  | KATFRASE |  | Concentration | mg/L | Y | (fl. conc.) |  | Fluoride Dissolved (fl. conc.) | 16984-48-8 | Major Ions; | Major Ions; EMS code: 1106 | Chemical | Dissolved | LAB
1106 | mg/L | Fluoride Dissolved | X044 | Ion Chromatography | A clear (Total) or filtered (Dissolved) water sample is injected into a stream of carbonate-bicarbonate eluent and passed through a series of ion exchanger. The anions of interested are separated on the basis of their relative affinities for a low capacity , strongly basic anion exchanger (guard and separate columns). The separated anions are directed through the anion self-regenerating suppressor (ASRS). Hydronium ions in the suppressor will neutralize the eluent to its weakly ionized form and convert the separated anions in their conductive acid form. The separated anions are measured by conductivity. They are identified on the basis of retention time as compared to standards. Quantitation is by measurement of peak area of peak height against a calibration curve generated from known standards. | 0.02 | Y | Y | Y | 85248 | 9/24/1996 10:07:00 AM | 8/21/2024 | 8/26/2024 4:53:05 PM | 1997-02-12 | 8/15/2019 11:49:07 AM | MPAK | KATFRASE | Concentration | mg/L | Y | (fl. conc.) |  | Fluoride Dissolved (fl. conc.) | 16984-48-8 | Major Ions; | Major Ions; EMS code: 1106 | Chemical | Dissolved | LAB
1106 | mg/L | Fluoride Dissolved | 1341 | I.S.E. | I.S.E. | 0.1 | Y | Y | Y | 11037 | 9/29/2000 | 10/5/2006 | 7/23/2024 11:41:49 AM | 1996-04-12 | 6/3/1998 6:51:33 PM | EMS |  | Concentration | mg/L | Y | (fl. conc.) |  | Fluoride Dissolved (fl. conc.) | 16984-48-8 | Major Ions; | Major Ions; EMS code: 1106 | Chemical | Dissolved | LAB
1106 | mg/L | Fluoride Dissolved | 1340 | Auto.I.S.E. | Bromide is analyzed potentiometrically using a bromide Automated Ion Specific Electrode in conjunction with a standard reference electrode and a pH meter. The potential developed by the presence of bromide ions is measured by an expanded pH meter. | 0.1 | Y | Y | Y | 553 |  |  | 3/2/2005 9:12:56 PM | 1996-04-12 | 6/3/1998 6:51:33 PM | EMS |  | Concentration | mg/L | Y | (fl. conc.) |  | Fluoride Dissolved (fl. conc.) | 16984-48-8 | Major Ions; | Major Ions; EMS code: 1106 | Chemical | Dissolved | LAB
1106 | mg/L | Fluoride Dissolved | X517 | pH5.8,Ionic Str.adj,SIE | pH 5.8 adj;Ionic Strength Adj; Selective Ion Electrode | 0.05 | Y | Y | Y | 7423 |  |  | 8/23/2002 10:11:07 AM | 1999-04-22 |  | NPEPPIN |  | Concentration | mg/L | Y | (fl. conc.) |  | Fluoride Dissolved (fl. conc.) | 16984-48-8 | Major Ions; | Major Ions; EMS code: 1106 | Chemical | Dissolved | LAB
1106 | mg/L | Fluoride Dissolved | EQ01 | EQUIS (mg/L) | EQUIS (mg/L) | 0.1 | N | Y | Y | 3374 |  |  | 10/24/1996 3:16:55 PM | 1996-04-12 | 6/3/1998 6:51:33 PM | EMS |  | Concentration | mg/L | Y | (fl. conc.) |  | Fluoride Dissolved (fl. conc.) | 16984-48-8 | Major Ions; | Major Ions; EMS code: 1106 | Chemical | Dissolved | LAB
1106 | t/d | Fluoride Dissolved | CALC | Calculation (t/d) | Calculation (t/d) | 0 | Y | Y | Y | 693 |  |  | 12/14/2006 2:21:55 PM | 1996-04-12 | 6/3/1998 6:51:45 PM | EMS |  | MassFlow | t/d | Y | (mass flow) |  | Fluoride Dissolved (mass flow) | 16984-48-8 | Major Ions; | Major Ions; EMS code: 1106 | Chemical | Dissolved | LAB
1106 | mg/L | Fluoride Dissolved | X114 | F Electr,known add meth | F Electr;known add meth | 0.1 | Y | Y | Y | 1155 |  |  | 3/9/1998 11:07:48 AM | 1996-04-12 | 6/3/1998 6:51:33 PM | EMS |  | Concentration | mg/L | Y | (fl. conc.) |  | Fluoride Dissolved (fl. conc.) | 16984-48-8 | Major Ions; | Major Ions; EMS code: 1106 | Chemical | Dissolved | LAB
1106 | mg/L | Fluoride Dissolved | X143 | Selective Ion Electrode | The analyte is determined potentiometrically by direct measurement using a selective ion electrode with standard reference electrode and either a meter with a direct concentration scale or a pH meter with an expanded mV scale. | 0.05 | Y | Y | Y | 21303 | 8/10/1993 | 8/18/2024 2:44:00 AM | 8/26/2024 8:06:03 AM | 1996-04-12 | 6/3/1998 6:51:42 PM | EMS |  | Concentration | mg/L | Y | (fl. conc.) |  | Fluoride Dissolved (fl. conc.) | 16984-48-8 | Major Ions; | Major Ions; EMS code: 1106 | Chemical | Dissolved | LAB
1107 | mg/L | Hardness (Dissolved) (as CaCO3) | CCAL | Calculation (Ca, Mg) | Dissolved and extractable water samples are ready for analysis once collected. Total samples are autoclaved with HNO3 and hydrochloric acid (HCl) in capped Teflon vials. Analysis is performed via Inductively Coupled Argon Plasma Atomic Emission Spectrometry (ICP-AES) (simultaneous multielement analysis), Graphite Furnace Atomic Absorption Spectrometry (GF-AAS) (low level analysis) and/or Flame Atomic Absorption Spectrometry (F-AAS). Total Hardness is calculated from dissolved Calcium and dissolved Magnesium according to the following formula: Total Hardness=2.497248*Ca + 4.116885*Mg (where Total Hardness is expressed as mg/L of CaCO3, and Calcium and Magnesium are expressed as mg/L Ca and mg/L Mg respectively. Check Media/Pres. codes for consistency with regard to acid preservation. | 1 | Y | Y | Y | 94370 | 4/11/1996 6:01:00 PM | 8/23/2024 | 8/26/2024 2:41:32 PM | 1997-01-14 | 6/3/1998 6:51:46 PM | MPAK |  | Concentration | mg/L | Y | (fl. conc.) |  | Hardness (Dissolved) (as CaCO3) (fl. conc.) |  | Major Ions; | Major Ions; EMS code: 1107 | Chemical | Dissolved | LAB
1107 | mg/L | Hardness (Dissolved) (as CaCO3) | EQ01 | EQUIS (mg/L) | EQUIS (mg/L) | 0.1 | N | Y | Y | 8493 |  |  | 10/24/1996 11:36:32 AM | 1996-04-12 | 6/3/1998 6:51:33 PM | EMS |  | Concentration | mg/L | Y | (fl. conc.) |  | Hardness (Dissolved) (as CaCO3) (fl. conc.) |  | Major Ions; | Major Ions; EMS code: 1107 | Chemical | Dissolved | LAB
1107 | mg/L | Hardness (Dissolved) (as CaCO3) | MCAL | Calculation (Ca,Mg,Sr,Fe,Al,Zn,Mn,Cu,Ba,Pb) | Dissolved and extractable water samples are ready for analysis once collected. Total samples are autoclaved with HNO3 and hydrochloric acid (HCl) in capped Teflon vials. Analysis is performed via Inductively Coupled Argon Plasma Atomic Emission Spectrometry (ICP-AES) (simultaneous multielement analysis), Graphite Furnace Atomic Absorption Spectrometry (GF-AAS) (low level analysis) and/or Flame Atomic Absorption Spectrometry (F-AAS). Total Hardness is calculated from the dissolved (or extractable as a substitute) fractions of Ca, Mg, Sr, Fe, Al, Zn, Mn, Cu, Ba, Pb, according to the following formula: Hardness Total = 2.497248*Ca + 4.116885*Mg + 1.142316*Sr + 1.792213*Fe + 3.709568*Al + 1.531126*Zn + 1.821866*Mn + 1.575224*Cu + 0.728773*Ba + 0.483082*Pb (where Total Hardness is expressed as mg/L of CaCO3, and each element is expressed in mg/L. Check Media/Pres. codes for consistency with regard to acid preservation. | 1 | Y | Y | Y | 5530 | 4/11/1996 6:01:00 PM | 4/16/2015 4:57:00 PM | 9/11/2019 8:47:50 AM | 1997-01-14 | 6/3/1998 6:51:46 PM | MPAK |  | Concentration | mg/L | Y | (fl. conc.) |  | Hardness (Dissolved) (as CaCO3) (fl. conc.) |  | Major Ions; | Major Ions; EMS code: 1107 | Chemical | Dissolved | LAB
1107 | mg/L | Hardness (Dissolved) (as CaCO3) | CS01 | System Calculation | System Calculation (mg/L) |  | N | Y | N | 16781 |  |  | 7/31/2024 1:30:04 PM | 1996-04-12 | 6/3/1998 6:51:43 PM | EMS |  | Concentration | mg/L | Y | (fl. conc.) |  | Hardness (Dissolved) (as CaCO3) (fl. conc.) |  | Major Ions; | Major Ions; EMS code: 1107 | Chemical | Dissolved | LAB
1107 | mg/L | Hardness (Dissolved) (as CaCO3) | X040 | EDTA Titration | a) For Ca, the sample pH is adjusted to pH 12-13 so that the magnesium is precipitated as the hydroxide. Next, the sample is titrated with EDTA in the presence of an indicator. When sufficient EDTA has been added so as to complex all the calcium ions the indicator changes color. Results reported as mg/L as Ca. b) For hardness, a sample buffered to pH 10.0+/-0.1 is titrated with EDTA in the presence of an indicator. When sufficient EDTA has been added in order to complex all the calcium and magnesium ions the indicator changes color. Results reported as mg/L as CaCO3. | 0.1 | Y | Y | Y | 2605 |  |  | 8/23/1999 4:10:47 PM | 1996-04-12 | 6/3/1998 6:51:45 PM | EMS |  | Concentration | mg/L | Y | (fl. conc.) |  | Hardness (Dissolved) (as CaCO3) (fl. conc.) |  | Major Ions; | Major Ions; EMS code: 1107 | Chemical | Dissolved | LAB
1107 | mg/L | Hardness (Dissolved) (as CaCO3) | XCAL | Calculation (mg/L) | Calculation (mg/L) | 0.5 | Y | Y | Y | 67856 | 8/12/1996 6:06:00 PM | 8/26/2024 | 8/26/2024 4:53:05 PM | 1996-04-12 | 8/15/2019 1:27:08 PM | EMS | KATFRASE | Concentration | mg/L | Y | (fl. conc.) |  | Hardness (Dissolved) (as CaCO3) (fl. conc.) |  | Major Ions; | Major Ions; EMS code: 1107 | Chemical | Dissolved | LAB
1108 | mg/L | Nitrogen Ammonia Dissolved | 2680 | Froz.Aut.Bertholot Meth. | Froz.Aut.Bertholot Meth. | 0.005 | Y | Y | Y | 1915 |  |  | 1/5/2000 7:22:22 AM | 1996-04-12 | 6/3/1998 6:51:33 PM | EMS |  | Concentration | mg/L | Y | (fl. conc.) |  | Nitrogen Ammonia Dissolved (fl. conc.) | 7664-41-7 | Nutrients; | Nutrients; EMS code: 1108 | Chemical | Dissolved | LAB
1108 | mg/L | Nitrogen Ammonia Dissolved | OPAF | Ortho-phthaldialdehyde (OPA) reaction,flourometric detection | A fluorescent derivative is formed by reacting ortho-phthaldialdehyde (OPA) and sulphite with ammonium. A sample is injected in a matrix matched carrier which is then mixed with a sodium hydroxide solution to convert the free ammonium into ammonia gas. After mixing, this basic solution is pushed through a dialysis PTFE membrane to the receiver OPA solution which then reacts to form the fluorescent derivative. This reaction is further increased (Intensified) by passing through a 62 degree inline reaction loop which then is detected and quantified using an inline fluorescent detector. This method is applicable for the measurement of ammonium in drinking water, ground water, and can be acceptable for brackish waters. | 0.005 | Y | Y | Y | 12 | 11/1/2022 | 10/19/2023 | 11/9/2023 7:41:27 AM | 2010-08-04 |  | RDGIBSON |  | Concentration | mg/L | Y | (fl. conc.) |  | Nitrogen Ammonia Dissolved (fl. conc.) | 7664-41-7 | Nutrients; | Nutrients; EMS code: 1108 | Chemical | Dissolved | LAB
1108 | t/d | Nitrogen Ammonia Dissolved | CALC | Calculation (t/d) | Calculation (t/d) | 0 | Y | Y | Y | 6528 |  |  | 7/25/2024 3:53:30 PM | 1996-04-12 | 6/3/1998 6:51:33 PM | EMS |  | MassFlow | t/d | Y | (mass flow) |  | Nitrogen Ammonia Dissolved (mass flow) | 7664-41-7 | Nutrients; | Nutrients; EMS code: 1108 | Chemical | Dissolved | LAB
1108 | mg/L | Nitrogen Ammonia Dissolved | X054 | Phenate-Colorimetric | Phenate-Colorimetric | 0.005 | Y | Y | Y | 27011 | 4/23/1997 | 8/20/2024 | 8/25/2024 9:47:36 PM | 1996-04-12 | 6/3/1998 6:51:33 PM | EMS |  | Concentration | mg/L | Y | (fl. conc.) |  | Nitrogen Ammonia Dissolved (fl. conc.) | 7664-41-7 | Nutrients; | Nutrients; EMS code: 1108 | Chemical | Dissolved | LAB
1108 | mg/L | Nitrogen Ammonia Dissolved | F001 | Phenol/NaHypochlorite Oxid,Indopheno:Col'm | Phenol, sodium nitroprusside, and an oxidizing solution composed of sodium hypochlorite in a basic citrate solution are sequentially added to the sample. Under basic conditions, ammonia reacts with phenol and is oxidized by sodium hypochlorite to form an organic complex, indophenol blue. The blue colour formed is intensified with sodium nitroprusside, and the absorbance of the resulting complex is measured at 630 nm. Sodium citrate complexes calcium and magnesium ions and prevents precipitation of their respective hydroxides. A heating coil is used to accelerate the colour development. | 0.005 | Y | Y | Y | 21082 | 8/5/1993 | 12/12/2020 | 4/6/2023 5:10:03 PM | 1996-06-20 | 6/3/1998 6:51:44 PM | MPAK |  | Concentration | mg/L | Y | (fl. conc.) |  | Nitrogen Ammonia Dissolved (fl. conc.) | 7664-41-7 | Nutrients; | Nutrients; EMS code: 1108 | Chemical | Dissolved | LAB
1108 | mg/L | Nitrogen Ammonia Dissolved | EQ01 | EQUIS (mg/L) | EQUIS (mg/L) | 0.005 | N | Y | Y | 14236 |  |  | 10/24/1996 3:16:44 PM | 1996-04-12 | 6/3/1998 6:51:33 PM | EMS |  | Concentration | mg/L | Y | (fl. conc.) |  | Nitrogen Ammonia Dissolved (fl. conc.) | 7664-41-7 | Nutrients; | Nutrients; EMS code: 1108 | Chemical | Dissolved | LAB
1108 | mg/L | Nitrogen Ammonia Dissolved | X143 | Selective Ion Electrode | The analyte is determined potentiometrically by direct measurement using a selective ion electrode with standard reference electrode and either a meter with a direct concentration scale or a pH meter with an expanded mV scale. | 0.01 | Y | Y | Y | 1628 | 1/22/1998 | 11/29/2022 9:45:00 AM | 12/5/2022 5:01:14 PM | 1996-04-12 | 6/3/1998 6:51:33 PM | EMS |  | Concentration | mg/L | Y | (fl. conc.) |  | Nitrogen Ammonia Dissolved (fl. conc.) | 7664-41-7 | Nutrients; | Nutrients; EMS code: 1108 | Chemical | Dissolved | LAB
1108 | mg/L | Nitrogen Ammonia Dissolved | X154 | Auto Berthelot (Indo) | Auto Berthelot (Indo) | 0.002 | Y | Y | Y | 29362 | 7/19/1996 | 4/4/2023 | 8/22/2024 1:49:13 PM | 1996-04-12 | 6/3/1998 6:51:45 PM | EMS |  | Concentration | mg/L | Y | (fl. conc.) |  | Nitrogen Ammonia Dissolved (fl. conc.) | 7664-41-7 | Nutrients; | Nutrients; EMS code: 1108 | Chemical | Dissolved | LAB
1108 | mg/L | Nitrogen Ammonia Dissolved | X112 | Dist@pH9.5,Nessleriz. | Dist@pH9.5;Nessleriz. | 0.005 | Y | Y | Y | 2844 |  |  | 12/14/2006 2:22:12 PM | 1996-04-12 | 6/3/1998 6:51:45 PM | EMS |  | Concentration | mg/L | Y | (fl. conc.) |  | Nitrogen Ammonia Dissolved (fl. conc.) | 7664-41-7 | Nutrients; | Nutrients; EMS code: 1108 | Chemical | Dissolved | LAB
1108 | mg/L | Nitrogen Ammonia Dissolved | NH3F | NH3-F-VA | Ammonia in Water by Fluorescence | 0.005 | Y | Y | Y | 9293 | 1/23/2011 | 8/16/2024 | 8/26/2024 3:51:07 PM | 2015-03-09 |  | RDGIBSON |  | Concentration | mg/L | Y | (fl. conc.) |  | Nitrogen Ammonia Dissolved (fl. conc.) | 7664-41-7 | Nutrients; | Nutrients; EMS code: 1108 | Chemical | Dissolved | LAB
1108 | mg/L | Nitrogen Ammonia Dissolved | X326 | Berthelot Colorimetric | Alkaline phenol and hypochlorite react with ammonia to form indophenol blue that is proportional to ammonia concentration. Sodium nitroprusside intensifies the blue colour thus formed. Analyses are performed with an autoanalyzer with spectrometer and 630-660 nm filters and 15 mm or 50 mm tubular flow cell. A manual version of this method may also be employed. Note that calcium and magnesium ions may be present in concentrations sufficient to cause precipitation problems during analysis. Also, sample turbidity and colour may interfere with this method. | 0.005 | Y | Y | Y | 64865 | 9/17/1996 | 8/20/2024 1:55:00 PM | 8/26/2024 2:41:32 PM | 1996-04-12 | 6/3/1998 6:51:33 PM | EMS |  | Concentration | mg/L | Y | (fl. conc.) |  | Nitrogen Ammonia Dissolved (fl. conc.) | 7664-41-7 | Nutrients; | Nutrients; EMS code: 1108 | Chemical | Dissolved | LAB
1108 | mg/L | Nitrogen Ammonia Dissolved | X126 | Dist,Acid Titration | Dist;Acid Titration | 0.005 | Y | Y | Y | 70 |  |  | 1/29/2002 8:48:19 AM | 1996-04-12 | 6/3/1998 6:51:33 PM | EMS |  | Concentration | mg/L | Y | (fl. conc.) |  | Nitrogen Ammonia Dissolved (fl. conc.) | 7664-41-7 | Nutrients; | Nutrients; EMS code: 1108 | Chemical | Dissolved | LAB
1108 | ug/g | Nitrogen Ammonia Dissolved | 15SW | Colourimetric or IC | Colourimetric or IC | 0.025 | Y | Y | Y | 29716 | 11/1/2002 | 10/31/2023 | 11/23/2023 7:26:23 AM | 1996-04-12 | 6/3/1998 6:51:33 PM | EMS |  | MassPerMass | mg/kg | Y | (mass conc.) |  | Nitrogen Ammonia Dissolved (mass conc.) | 7664-41-7 | Nutrients; | Nutrients; EMS code: 1108 | Chemical | Dissolved | LAB
1108-E | mg/L | Ammonia Extractable | CSSS15 | CSSS Ch. 15 | See Canadian Society of Soil Science (2008) Chapter 15 and Journal of Environmental Monitoring 2005, 7, 37-42 | 1 | Y | Y | Y | 15 | 1/11/2022 | 2/6/2022 | 1/11/2024 1:07:06 PM | 2022-06-29 |  | EMS |  | Concentration | mg/L | Y | (fl. conc.) |  | Ammonia Extractable (fl. conc.) | 7664-41-7 | Nutrients; | Nutrients; EMS code: 1108-E | Chemical | Extractable | LAB
1108-S | mg/L | Ammonia; soluble ion content (as N) | CSSS15 | CSSS Ch. 15 | See Canadian Society of Soil Science (2008) Chapter 15 and Journal of Environmental Monitoring 2005, 7, 37-42 | 1 | Y | Y | Y | 0 |  |  |  | 2024-02-02 |  | EMS |  | Concentration | mg/L | Y | (fl. conc.) |  | Ammonia; soluble ion content (as N) (fl. conc.) |  | Nutrients; | Nutrients; EMS code: 1108-S | Chemical |  | LAB
1109 | mg/L | Nitrate(NO3) + Nitrite(NO2) Dissolved | X328 | Auto Cadmium Red'n,Diazo | The filtered sample is passed through a column containing granulated copper-cadmium which reduces Nitrate to Nitrite. Quantitative reduction is ensured by maintaining a stable H-ion concentration through addition of ammonium chloride (NH4Cl). The Nitrite that is originally present in the sample plus the reduced Nitrate then react with sulfanilamide under acidic conditions to form a diazo compound. This compound then complexes with N-1-naphthylethylenediamine dihydrochloride to form a reddish purple azo dye, the absorbance of which is measured at 550 nm. Nitrate (NO3) concentration can be calculated using the Nitrate & Nitrite result and subtracting the Nitrite result (NO23 - NO2 = NO3). Automated analysis. | 0.002 | Y | Y | Y | 28982 | 4/25/1995 6:16:00 PM | 12/12/2020 | 8/26/2024 2:41:32 PM | 1996-04-12 | 6/3/1998 6:51:33 PM | EMS |  | Concentration | mg/L | Y | (fl. conc.) |  | Nitrate(NO3) + Nitrite(NO2) Dissolved (fl. conc.) | 14797-55-8, 14797-65-0 | Nutrients; | Nutrients; EMS code: 1109 | Chemical | Dissolved | LAB
1109 | mg/L | Nitrate(NO3) + Nitrite(NO2) Dissolved | X155 | Auto Cd:Diazotization | Auto Cd:Diazotization | 0.002 | Y | Y | Y | 4608 | 1/4/2007 4:08:00 PM | 8/17/2024 2:09:00 PM | 8/23/2024 5:01:25 PM | 1996-04-12 | 6/3/1998 6:51:33 PM | EMS |  | Concentration | mg/L | Y | (fl. conc.) |  | Nitrate(NO3) + Nitrite(NO2) Dissolved (fl. conc.) | 14797-55-8, 14797-65-0 | Nutrients; | Nutrients; EMS code: 1109 | Chemical | Dissolved | LAB
1109 | mg/L | Nitrate(NO3) + Nitrite(NO2) Dissolved | X020 | Manual Cadmium Reduction | The filtered sample is passed through a column containing granulated copper-cadmium which reduces Nitrate to Nitrite. Quantitative reduction is ensured by maintaining a stable H-ion concentration through addition of ammonium chloride (NH4Cl). The Nitrite that is originally present in the sample plus the reduced Nitrate then react with sulfanilamide under acidic conditions to form a diazo compound. This compound then complexes with N-1-naphthylethylenediamine dihydrochloride to form a reddish purple azo dye, the absorbance of which is measured at 550 nm. Nitrate (NO3) concentration can be calculated using the Nitrate & Nitrite result and subtracting the Nitrite result (NO23 - NO2 = NO3). Manual analysis. | 0.02 | Y | Y | Y | 3168 |  |  | 3/7/2003 12:11:11 PM | 1996-04-12 | 6/3/1998 6:51:33 PM | EMS |  | Concentration | mg/L | Y | (fl. conc.) |  | Nitrate(NO3) + Nitrite(NO2) Dissolved (fl. conc.) | 14797-55-8, 14797-65-0 | Nutrients; | Nutrients; EMS code: 1109 | Chemical | Dissolved | LAB
1109 | mg/L | Nitrate(NO3) + Nitrite(NO2) Dissolved | CS01 | System Calculation | System Calculation (mg/L) |  | N | Y | N | 94658 |  |  | 8/26/2024 4:53:06 PM | 1996-04-12 | 6/3/1998 6:51:43 PM | EMS |  | Concentration | mg/L | Y | (fl. conc.) |  | Nitrate(NO3) + Nitrite(NO2) Dissolved (fl. conc.) | 14797-55-8, 14797-65-0 | Nutrients; | Nutrients; EMS code: 1109 | Chemical | Dissolved | LAB
1109 | mg/L | Nitrate(NO3) + Nitrite(NO2) Dissolved | X044 | Ion Chromatography | A clear (Total) or filtered (Dissolved) water sample is injected into a stream of carbonate-bicarbonate eluent and passed through a series of ion exchanger. The anions of interested are separated on the basis of their relative affinities for a low capacity , strongly basic anion exchanger (guard and separate columns). The separated anions are directed through the anion self-regenerating suppressor (ASRS). Hydronium ions in the suppressor will neutralize the eluent to its weakly ionized form and convert the separated anions in their conductive acid form. The separated anions are measured by conductivity. They are identified on the basis of retention time as compared to standards. Quantitation is by measurement of peak area of peak height against a calibration curve generated from known standards. | 0.0051 | Y | Y | Y | 14379 | 8/12/2014 | 9/29/2020 | 12/13/2023 2:16:30 PM | 2015-03-11 |  | RDGIBSON |  | Concentration | mg/L | Y | (fl. conc.) |  | Nitrate(NO3) + Nitrite(NO2) Dissolved (fl. conc.) | 14797-55-8, 14797-65-0 | Nutrients; | Nutrients; EMS code: 1109 | Chemical | Dissolved | LAB
1109 | mg/L | Nitrate(NO3) + Nitrite(NO2) Dissolved | EQ01 | EQUIS (mg/L) | EQUIS (mg/L) | 0.02 | N | Y | Y | 15213 |  |  | 10/24/1996 3:16:44 PM | 1996-04-12 | 6/3/1998 6:51:45 PM | EMS |  | Concentration | mg/L | Y | (fl. conc.) |  | Nitrate(NO3) + Nitrite(NO2) Dissolved (fl. conc.) | 14797-55-8, 14797-65-0 | Nutrients; | Nutrients; EMS code: 1109 | Chemical | Dissolved | LAB
1109 | mg/L | Nitrate(NO3) + Nitrite(NO2) Dissolved | XCAL | Calculation (mg/L) | Calculation (mg/L) | 0.002 | Y | Y | Y | 21231 | 1/8/2010 1:51:00 PM | 8/22/2024 | 8/26/2024 4:51:04 PM | 2011-02-28 | 2/28/2011 1:57:11 PM | RDGIBSON | RDGIBSON | Concentration | mg/L | Y | (fl. conc.) |  | Nitrate(NO3) + Nitrite(NO2) Dissolved (fl. conc.) | 14797-55-8, 14797-65-0 | Nutrients; | Nutrients; EMS code: 1109 | Chemical | Dissolved | LAB
1109 | mg/L | Nitrate(NO3) + Nitrite(NO2) Dissolved | X133 | Cd Red'n,Diazotiz,Col | Cd Red'n;Diazotiz;Col | 0.02 | Y | Y | Y | 987 | 11/28/2012 11:09:00 AM | 5/28/2014 5:04:00 AM | 2/3/2015 3:06:19 PM | 1996-04-12 | 6/3/1998 6:51:46 PM | EMS |  | Concentration | mg/L | Y | (fl. conc.) |  | Nitrate(NO3) + Nitrite(NO2) Dissolved (fl. conc.) | 14797-55-8, 14797-65-0 | Nutrients; | Nutrients; EMS code: 1109 | Chemical | Dissolved | LAB
1109 | mg/L | Nitrate(NO3) + Nitrite(NO2) Dissolved | 1350 | Nitrite & Nitrate in Seawater (Calc) | Nitrite & Nitrate in Seawater (Calc) | 0.51 | Y | Y | Y | 67 | 8/16/2019 | 9/15/2019 | 10/7/2019 5:29:52 PM | 2019-09-30 |  | KATFRASE |  | Concentration | mg/L | Y | (fl. conc.) |  | Nitrate(NO3) + Nitrite(NO2) Dissolved (fl. conc.) | 14797-55-8, 14797-65-0 | Nutrients; | Nutrients; EMS code: 1109 | Chemical | Dissolved | LAB
1109 | mg/L | Nitrate(NO3) + Nitrite(NO2) Dissolved | X082 | Auto. Cadmium Reduction | Auto. Cadmium Reduction | 0.02 | Y | Y | Y | 91226 | 8/5/1993 | 8/14/2024 5:01:00 PM | 8/26/2024 8:06:03 AM | 1996-04-12 | 6/3/1998 6:51:33 PM | EMS |  | Concentration | mg/L | Y | (fl. conc.) |  | Nitrate(NO3) + Nitrite(NO2) Dissolved (fl. conc.) | 14797-55-8, 14797-65-0 | Nutrients; | Nutrients; EMS code: 1109 | Chemical | Dissolved | LAB
1109 | mg/L | Nitrate(NO3) + Nitrite(NO2) Dissolved | 2670 | Froz.Aut.Cadmium Red'n | Froz.Aut.Cadmium Red'n | 0.02 | Y | Y | Y | 1582 |  |  | 10/24/1996 3:06:02 PM | 1996-04-12 | 6/3/1998 6:51:33 PM | EMS |  | Concentration | mg/L | Y | (fl. conc.) |  | Nitrate(NO3) + Nitrite(NO2) Dissolved (fl. conc.) | 14797-55-8, 14797-65-0 | Nutrients; | Nutrients; EMS code: 1109 | Chemical | Dissolved | LAB
1109 | mg/L | Nitrate(NO3) + Nitrite(NO2) Dissolved | 1353 | Auto. Cadmium Reduc. L L | Auto. Cadmium Reduc. L L | 0.005 | N | Y | Y | 1141 |  |  | 10/24/1996 2:45:43 PM | 1996-04-12 | 6/3/1998 6:51:42 PM | EMS |  | Concentration | mg/L | Y | (fl. conc.) |  | Nitrate(NO3) + Nitrite(NO2) Dissolved (fl. conc.) | 14797-55-8, 14797-65-0 | Nutrients; | Nutrients; EMS code: 1109 | Chemical | Dissolved | LAB
1109 | mg/L | Nitrate(NO3) + Nitrite(NO2) Dissolved | F046 | Low Level,Auto Cd Red'n,Diazo | Low level Nitrate/Nitrite samples must be field filtered and collected in glass bottles with no dead space. The filtered sample is passed through a column containing granulated copper-cadmium which reduces Nitrate to Nitrite. Quantitative reduction is ensured by maintaining a stable H-ion concentration through addition of ammonium chloride (NH4Cl). The Nitrite that is originally present in the sample plus the reduced Nitrate then react with sulphanilamide under acidic conditions to form a diazo compound. This compound then complexes with N-1-naphthylethylenediamine dihydrochloride to form a reddish-purple azo dye, the absorbance of which is measured at 550 nm. Nitrate (NO3) concentration can be calculated using the Nitrate & Nitrite result and subtracting the Nitrite result (NO23 - NO2 = NO3) | 0.002 | Y | Y | Y | 1304 | 5/27/1997 11:15:00 AM | 12/5/2020 | 5/26/2003 11:42:43 AM | 1997-10-20 | 6/3/1998 6:51:47 PM | NPEPPIN |  | Concentration | mg/L | Y | (fl. conc.) |  | Nitrate(NO3) + Nitrite(NO2) Dissolved (fl. conc.) | 14797-55-8, 14797-65-0 | Nutrients; | Nutrients; EMS code: 1109 | Chemical | Dissolved | LAB
1109 | mg/L | Nitrate(NO3) + Nitrite(NO2) Dissolved | X519 | F,EDTA,Cd red'n,diaz,col'm | Filtered;EDTA etc.;Cd reduction;diazotization;colorimetric | 0.001 | Y | Y | Y | 20457 |  |  | 10/15/1999 1:06:52 PM | 1999-04-22 |  | NPEPPIN |  | Concentration | mg/L | Y | (fl. conc.) |  | Nitrate(NO3) + Nitrite(NO2) Dissolved (fl. conc.) | 14797-55-8, 14797-65-0 | Nutrients; | Nutrients; EMS code: 1109 | Chemical | Dissolved | LAB
1109 | ug/g | Nitrate(NO3) + Nitrite(NO2) Dissolved | 15SW | Colourimetric or IC | Colourimetric or IC | 0.1 | Y | Y | Y | 21 | 8/30/2001 | 11/9/2009 5:50:00 PM | 12/23/2009 11:11:37 AM | 1996-04-12 | 6/3/1998 6:51:33 PM | EMS |  | MassPerMass | mg/kg | Y | (mass conc.) |  | Nitrate(NO3) + Nitrite(NO2) Dissolved (mass conc.) | 14797-55-8, 14797-65-0 | Nutrients; | Nutrients; EMS code: 1109 | Chemical | Dissolved | LAB
1109-A | ug/g | Nitrate and Nitrite (as N) available | A4500 | APHA 4500-NO3 | See APHA 4500-NO3 for full method details. | 1 | Y | Y | Y | 18 | 9/2/2021 | 2/13/2023 | 1/11/2024 1:07:06 PM | 2022-06-29 |  | EMS |  | MassPerMass | mg/kg | Y | (mass conc.) |  | Nitrate and Nitrite (as N) available (mass conc.) | 14797-55-8, 14797-65-0 | Nutrients; | Nutrients; EMS code: 1109-A | Chemical |  | LAB
1110 | mg/L | Nitrate (NO3) Dissolved | X133 | Cd Red'n,Diazotiz,Col | Cd Red'n;Diazotiz;Col | 0.02 | Y | Y | Y | 6773 | 7/21/2000 | 2/17/2009 3:05:00 PM | 8/28/2018 3:14:03 PM | 1996-04-12 | 6/3/1998 6:51:33 PM | EMS |  | Concentration | mg/L | Y | (fl. conc.) |  | Nitrate (NO3) Dissolved (fl. conc.) | 14797-55-8 | Nutrients; | Nutrients; EMS code: 1110 | Chemical | Dissolved | LAB
1110 | mg/L | Nitrate (NO3) Dissolved | X044 | Ion Chromatography | A clear (Total) or filtered (Dissolved) water sample is injected into a stream of carbonate-bicarbonate eluent and passed through a series of ion exchanger. The anions of interested are separated on the basis of their relative affinities for a low capacity , strongly basic anion exchanger (guard and separate columns). The separated anions are directed through the anion self-regenerating suppressor (ASRS). Hydronium ions in the suppressor will neutralize the eluent to its weakly ionized form and convert the separated anions in their conductive acid form. The separated anions are measured by conductivity. They are identified on the basis of retention time as compared to standards. Quantitation is by measurement of peak area of peak height against a calibration curve generated from known standards. | 0.005 | Y | Y | Y | 189967 | 4/10/1996 6:56:00 AM | 8/21/2024 | 8/26/2024 4:53:05 PM | 1996-04-12 | 8/15/2019 11:49:48 AM | EMS | KATFRASE | Concentration | mg/L | Y | (fl. conc.) |  | Nitrate (NO3) Dissolved (fl. conc.) | 14797-55-8 | Nutrients; | Nutrients; EMS code: 1110 | Chemical | Dissolved | LAB
1110 | ug/g | Nitrate (NO3) Dissolved | CS02 | System Calculation | System Calculation (ug/g) |  | N | Y | N | 7 |  |  | 12/23/2009 11:11:40 AM | 1996-04-12 | 6/3/1998 6:51:43 PM | EMS |  | MassPerMass | mg/kg | Y | (mass conc.) |  | Nitrate (NO3) Dissolved (mass conc.) | 14797-55-8 | Nutrients; | Nutrients; EMS code: 1110 | Chemical | Dissolved | LAB
1110 | mg/L | Nitrate (NO3) Dissolved | X056 | Chromotropic Acid Meth. | Chromotropic Acid Meth. | 0.02 | Y | Y | Y | 293 |  |  | 10/24/1996 12:10:42 PM | 1996-04-12 | 6/3/1998 6:51:33 PM | EMS |  | Concentration | mg/L | Y | (fl. conc.) |  | Nitrate (NO3) Dissolved (fl. conc.) | 14797-55-8 | Nutrients; | Nutrients; EMS code: 1110 | Chemical | Dissolved | LAB
1110 | mg/L | Nitrate (NO3) Dissolved | 1346 | Nitrate in Seawater by IC | Nitrate in Seawater by IC | 0.5 | Y | Y | Y | 108 | 8/14/2019 | 4/25/2024 | 6/7/2024 10:03:16 AM | 2019-09-30 |  | KATFRASE |  | Concentration | mg/L | Y | (fl. conc.) |  | Nitrate (NO3) Dissolved (fl. conc.) | 14797-55-8 | Nutrients; | Nutrients; EMS code: 1110 | Chemical | Dissolved | LAB
1110 | mg/L | Nitrate (NO3) Dissolved | F124 | UV Spectrophotometric Screen | Measurement of Nitrate is accomplished by the preparation of Nitrate standards in the range of 0-7mg/L-NO3-N. Determination of sample nitrate concentration is accomplished by reading the UV absorbance of the sample at a wavelength at 220nm and at 275nm to determine interference. A standard curve is constructed by plotting the absorbances due to NO3 conc. of standards. Using the corrected sample absorbances, sample concentrations are obtained directly from the curve. |  | Y | Y | Y | 15181 |  |  | 7/29/2024 2:00:44 PM | 2004-02-10 |  | UMULLER |  | Concentration | mg/L | Y | (fl. conc.) |  | Nitrate (NO3) Dissolved (fl. conc.) | 14797-55-8 | Nutrients; | Nutrients; EMS code: 1110 | Chemical | Dissolved | LAB
1110 | mg/L | Nitrate (NO3) Dissolved | CS01 | System Calculation | System Calculation (mg/L) |  | N | Y | N | 14516 |  |  | 4/23/2021 12:06:22 PM | 1996-04-12 | 6/3/1998 6:51:43 PM | EMS |  | Concentration | mg/L | Y | (fl. conc.) |  | Nitrate (NO3) Dissolved (fl. conc.) | 14797-55-8 | Nutrients; | Nutrients; EMS code: 1110 | Chemical | Dissolved | LAB
1110 | mg/L | Nitrate (NO3) Dissolved | X257 | Colorimetric | Colorimetric. General use for several parameters. E.g.Total Sulfide is determined by spectrophotometer using the methlyene blue colourimetric method. The H2S calculation converts Total Sulphide as (S2-) and reports it as Total Sulphide as (H2S). | 0.02 | Y | Y | Y | 6313 | 10/5/2022 | 8/6/2024 | 8/21/2024 11:21:21 AM | 1996-04-12 | 6/3/1998 6:51:33 PM | EMS |  | Concentration | mg/L | Y | (fl. conc.) |  | Nitrate (NO3) Dissolved (fl. conc.) | 14797-55-8 | Nutrients; | Nutrients; EMS code: 1110 | Chemical | Dissolved | LAB
1110 | mg/L | Nitrate (NO3) Dissolved | XCAL | Calculation (mg/L) | Calculation (mg/L) | 0.002 | Y | Y | Y | 37011 | 8/5/1993 | 8/17/2024 2:09:00 PM | 8/26/2024 8:06:03 AM | 1996-04-12 | 6/3/1998 6:51:41 PM | EMS |  | Concentration | mg/L | Y | (fl. conc.) |  | Nitrate (NO3) Dissolved (fl. conc.) | 14797-55-8 | Nutrients; | Nutrients; EMS code: 1110 | Chemical | Dissolved | LAB
1110 | mg/L | Nitrate (NO3) Dissolved | EQ01 | EQUIS (mg/L) | EQUIS (mg/L) | 0.02 | N | Y | Y | 4647 |  |  | 10/24/1996 3:16:57 PM | 1996-04-12 | 6/3/1998 6:51:33 PM | EMS |  | Concentration | mg/L | Y | (fl. conc.) |  | Nitrate (NO3) Dissolved (fl. conc.) | 14797-55-8 | Nutrients; | Nutrients; EMS code: 1110 | Chemical | Dissolved | LAB
1110 | ug/g | Nitrate (NO3) Dissolved | X082 | Auto. Cadmium Reduction | Auto. Cadmium Reduction | 1 | Y | Y | Y | 354 |  |  | 6/22/2020 3:01:47 PM | 2020-05-28 | 5/28/2020 2:58:28 PM | RDGIBSON | RDGIBSON | MassPerMass | mg/kg | Y | (mass conc.) |  | Nitrate (NO3) Dissolved (mass conc.) | 14797-55-8 | Nutrients; | Nutrients; EMS code: 1110 | Chemical | Dissolved | LAB
1110 | mg/L | Nitrate (NO3) Dissolved | X155 | Auto Cd:Diazotization | Auto Cd:Diazotization | 0.001 | Y | Y | Y | 15 | 4/11/2002 | 5/28/2003 | 2/24/2004 2:45:17 PM | 2003-03-18 | 3/18/2003 2:51:45 PM | UMULLER | UMULLER | Concentration | mg/L | Y | (fl. conc.) |  | Nitrate (NO3) Dissolved (fl. conc.) | 14797-55-8 | Nutrients; | Nutrients; EMS code: 1110 | Chemical | Dissolved | LAB
1110-E | mg/L | Nitrate (as N) Extractable | E300 | EPA 300.1/CSSS 15 | See EPA method 300.1 and Canadian Society of Soil Sciences (2008) Chapter 15 | 0.5 | Y | Y | Y | 15 | 1/11/2022 | 2/5/2022 | 1/11/2024 1:07:06 PM | 2022-06-29 |  | EMS |  | Concentration | mg/L | Y | (fl. conc.) |  | Nitrate (as N) Extractable (fl. conc.) | 14797-55-8 | Nutrients; | Nutrients; EMS code: 1110-E | Chemical | Extractable | LAB
1110-E | ug/g | Nitrate (as N) Extractable | CAL2 | Calculation (ug/g) | Calculation (ug/g) | 0.5 | Y | Y | Y | 15 | 1/13/2022 | 2/8/2022 | 1/11/2024 1:07:06 PM | 2022-07-06 |  | EMS |  | MassPerMass | mg/kg | Y | (mass conc.) |  | Nitrate (as N) Extractable (mass conc.) | 14797-55-8 | Nutrients; | Nutrients; EMS code: 1110-E | Chemical | Extractable | LAB
1110-L | mg/L | Nitrate (as N) Leachable | MEND | MEND (mod) EPA 300.1 (mod) | MEND (mod) EPA 300.1 (mod). Inorganic anions are analyzed by obtaining an extract from a soil sample that has been added in a 3:1 ratio of deionized water to soil; then leached and the extract filtered prior to analysis; which is then analyzed by Ion Chromatography with conductivity and/or UV detection. | 0.005 | Y | Y | Y | 0 |  |  |  | 2024-01-24 |  | RDGIBSON |  | Concentration | mg/L | Y | (fl. conc.) |  | Nitrate (as N) Leachable (fl. conc.) |  | Nutrients; | Nutrients; EMS code: 1110-L | Chemical |  | LAB
1110-N | ug/g | Nitrate, Water-Soluble (as N) | ICAN | Ion Chromatography | Ion Chromatography | 0.05 | Y | Y | Y | 7 | 9/22/2022 9:13:00 AM | 9/22/2022 9:13:00 AM | 3/9/2023 4:55:51 PM | 2022-09-29 |  | EMS |  | MassPerMass | mg/kg | Y | (mass conc.) |  | Nitrate, Water-Soluble (as N) (mass conc.) | 14797-55-8 | Nutrients; | Nutrients; EMS code: 1110-N | Chemical | Soluble | LAB
1110-S | ug/g | Nitrate (as N); soluble ion content | CAL2 | Calculation (ug/g) | Calculation (ug/g) | 0.5 | Y | Y | Y | 0 |  |  |  | 2024-02-02 |  | EMS |  | MassPerMass | mg/kg | Y | (mass conc.) |  | Nitrate (as N); soluble ion content (mass conc.) |  | Nutrients; | Nutrients; EMS code: 1110-S | Chemical |  | LAB
1110-S | mg/L | Nitrate (as N); soluble ion content | E300 | EPA 300.1/CSSS 15 | See EPA method 300.1 and Canadian Society of Soil Sciences (2008) Chapter 15 | 0.5 | Y | Y | Y | 0 |  |  |  | 2024-02-02 |  | EMS |  | Concentration | mg/L | Y | (fl. conc.) |  | Nitrate (as N); soluble ion content (fl. conc.) |  | Nutrients; | Nutrients; EMS code: 1110-S | Chemical |  | LAB
1111 | mg/L | Nitrogen - Nitrite Dissolved (NO2) | CS01 | System Calculation | System Calculation (mg/L) |  | N | Y | N | 5111 |  |  | 8/26/2024 2:41:32 PM | 1996-04-12 | 6/3/1998 6:51:43 PM | EMS |  | Concentration | mg/L | Y | (fl. conc.) |  | Nitrogen - Nitrite Dissolved (NO2) (fl. conc.) | 14797-65-0 | Nutrients; | Nutrients; EMS code: 1111 | Chemical | Dissolved | LAB
1111 | mg/L | Nitrogen - Nitrite Dissolved (NO2) | X155 | Auto Cd:Diazotization | Auto Cd:Diazotization | 0.001 | Y | Y | Y | 5524 | 10/22/1997 | 5/6/2004 | 8/28/2018 3:14:03 PM | 1996-04-12 | 6/3/1998 6:51:41 PM | EMS |  | Concentration | mg/L | Y | (fl. conc.) |  | Nitrogen - Nitrite Dissolved (NO2) (fl. conc.) | 14797-65-0 | Nutrients; | Nutrients; EMS code: 1111 | Chemical | Dissolved | LAB
1111 | ug/g | Nitrogen - Nitrite Dissolved (NO2) | 15SW | Colourimetric or IC | Colourimetric or IC | 0.025 | Y | Y | Y | 387 | 6/7/2000 | 10/31/2023 | 11/23/2023 7:26:23 AM | 1996-04-12 | 6/3/1998 6:51:45 PM | EMS |  | MassPerMass | mg/kg | Y | (mass conc.) |  | Nitrogen - Nitrite Dissolved (NO2) (mass conc.) | 14797-65-0 | Nutrients; | Nutrients; EMS code: 1111 | Chemical | Dissolved | LAB
1111 | mg/L | Nitrogen - Nitrite Dissolved (NO2) | X327 | Auto Colorimet,diazotiz. | Nitrites react with sulphanilamide under acidic conditions to form a diazo compound. This compound compexes with N-1-naphthylethylene diamine dihydrochloride to form a reddish-purple azo dye, the absorbance of which is measured at 540 nm. | 0.005 | Y | Y | Y | 4483 | 9/17/1996 | 8/14/2024 3:35:00 PM | 8/16/2024 5:07:04 PM | 1996-04-12 | 6/3/1998 6:51:33 PM | EMS |  | Concentration | mg/L | Y | (fl. conc.) |  | Nitrogen - Nitrite Dissolved (NO2) (fl. conc.) | 14797-65-0 | Nutrients; | Nutrients; EMS code: 1111 | Chemical | Dissolved | LAB
1111 | mg/L | Nitrogen - Nitrite Dissolved (NO2) | 1345 | Nitrite in Seawater by IC | Nitrite in Seawater by IC | 0.1 | Y | Y | Y | 90 | 8/14/2019 | 4/27/2024 | 6/7/2024 10:03:16 AM | 2019-09-30 |  | KATFRASE |  | Concentration | mg/L | Y | (fl. conc.) |  | Nitrogen - Nitrite Dissolved (NO2) (fl. conc.) | 14797-65-0 | Nutrients; | Nutrients; EMS code: 1111 | Chemical | Dissolved | LAB
1111 | mg/L | Nitrogen - Nitrite Dissolved (NO2) | 2690 | Froz.Auto.Diazotiz'n | Froz.Auto.Diazotiz'n | 0.005 | Y | Y | Y | 858 |  |  | 3/31/1998 11:11:38 AM | 1996-04-12 | 6/3/1998 6:51:33 PM | EMS |  | Concentration | mg/L | Y | (fl. conc.) |  | Nitrogen - Nitrite Dissolved (NO2) (fl. conc.) | 14797-65-0 | Nutrients; | Nutrients; EMS code: 1111 | Chemical | Dissolved | LAB
1111 | mg/L | Nitrogen - Nitrite Dissolved (NO2) | X044 | Ion Chromatography | A clear (Total) or filtered (Dissolved) water sample is injected into a stream of carbonate-bicarbonate eluent and passed through a series of ion exchanger. The anions of interested are separated on the basis of their relative affinities for a low capacity , strongly basic anion exchanger (guard and separate columns). The separated anions are directed through the anion self-regenerating suppressor (ASRS). Hydronium ions in the suppressor will neutralize the eluent to its weakly ionized form and convert the separated anions in their conductive acid form. The separated anions are measured by conductivity. They are identified on the basis of retention time as compared to standards. Quantitation is by measurement of peak area of peak height against a calibration curve generated from known standards. | 0.001 | Y | Y | Y | 120767 | 9/24/1996 10:07:00 AM | 8/21/2024 | 8/26/2024 4:53:05 PM | 1997-02-12 | 8/15/2019 11:50:16 AM | MPAK | KATFRASE | Concentration | mg/L | Y | (fl. conc.) |  | Nitrogen - Nitrite Dissolved (NO2) (fl. conc.) | 14797-65-0 | Nutrients; | Nutrients; EMS code: 1111 | Chemical | Dissolved | LAB
1111 | mg/L | Nitrogen - Nitrite Dissolved (NO2) | X081 | Auto. Diazotization | The Nitrite that is present in the filtered sample react with sulfanilamide under acidic conditions to form a diazo compound. This compound then complexes with N-1-naphthylethylenediamine dihydrochloride to form a reddish-purple azo dye, the absorbance of which is measured at 550 nm. | 0.002 | Y | Y | Y | 59224 | 8/5/1993 | 8/14/2024 3:12:00 PM | 8/26/2024 8:06:03 AM | 1996-04-12 | 6/3/1998 6:51:33 PM | EMS |  | Concentration | mg/L | Y | (fl. conc.) |  | Nitrogen - Nitrite Dissolved (NO2) (fl. conc.) | 14797-65-0 | Nutrients; | Nutrients; EMS code: 1111 | Chemical | Dissolved | LAB
1111 | mg/L | Nitrogen - Nitrite Dissolved (NO2) | X132 | Diazotization,Col'metric | Diazotization;Col'metric | 0.005 | Y | Y | Y | 6001 | 9/3/1996 | 8/17/2024 10:57:00 AM | 8/23/2024 5:01:25 PM | 1996-04-12 | 6/3/1998 6:51:33 PM | EMS |  | Concentration | mg/L | Y | (fl. conc.) |  | Nitrogen - Nitrite Dissolved (NO2) (fl. conc.) | 14797-65-0 | Nutrients; | Nutrients; EMS code: 1111 | Chemical | Dissolved | LAB
1111 | mg/L | Nitrogen - Nitrite Dissolved (NO2) | EQ01 | EQUIS (mg/L) | EQUIS (mg/L) | 0.005 | N | Y | Y | 10311 |  |  | 10/24/1996 3:16:57 PM | 1996-04-12 | 6/3/1998 6:51:33 PM | EMS |  | Concentration | mg/L | Y | (fl. conc.) |  | Nitrogen - Nitrite Dissolved (NO2) (fl. conc.) | 14797-65-0 | Nutrients; | Nutrients; EMS code: 1111 | Chemical | Dissolved | LAB
1111 | mg/L | Nitrogen - Nitrite Dissolved (NO2) | X019 | Manual Diazotization | Manual Diazotization | 0.005 | Y | Y | Y | 2465 | 1/4/2007 12:16:00 PM | 2/17/2009 3:05:00 PM | 5/8/2013 11:58:53 AM | 1996-04-12 | 6/3/1998 6:51:33 PM | EMS |  | Concentration | mg/L | Y | (fl. conc.) |  | Nitrogen - Nitrite Dissolved (NO2) (fl. conc.) | 14797-65-0 | Nutrients; | Nutrients; EMS code: 1111 | Chemical | Dissolved | LAB
1111-A | ug/g | Nitrite available (as N) | A4500 | APHA 4500-NO3 | See APHA 4500-NO3 for full method details. | 0.4 | Y | Y | Y | 2 | 9/2/2021 | 9/2/2021 | 10/19/2022 1:02:09 PM | 2022-10-19 |  | EMS |  | MassPerMass | mg/kg | Y | (mass conc.) |  | Nitrite available (as N) (mass conc.) | 14797-65-0 | Nutrients; | Nutrients; EMS code: 1111-A | Chemical |  | LAB
1111-L | mg/L | Nitrite (as N) Leachable | MEND | MEND (mod) EPA 300.1 (mod) | MEND (mod) EPA 300.1 (mod). Inorganic anions are analyzed by obtaining an extract from a soil sample that has been added in a 3:1 ratio of deionized water to soil; then leached and the extract filtered prior to analysis; which is then analyzed by Ion Chromatography with conductivity and/or UV detection. | 0.001 | Y | Y | Y | 0 |  |  |  | 2024-01-24 |  | RDGIBSON |  | Concentration | mg/L | Y | (fl. conc.) |  | Nitrite (as N) Leachable (fl. conc.) |  | Nutrients; | Nutrients; EMS code: 1111-L | Chemical |  | LAB
1111-N | ug/g | Nitrite - Water-Soluble (as N) | ICAN | Ion Chromatography | Ion Chromatography | 0.05 | Y | Y | Y | 7 | 9/22/2022 9:13:00 AM | 9/22/2022 9:13:00 AM | 3/9/2023 4:55:51 PM | 2022-09-29 |  | EMS |  | MassPerMass | mg/kg | Y | (mass conc.) |  | Nitrite - Water-Soluble (as N) (mass conc.) | 14797-65-0 | Nutrients; | Nutrients; EMS code: 1111-N | Chemical | Soluble | LAB
1112 | mg/L | Nitrogen Organic Total Dissolved | XCAL | Calculation (mg/L) | Calculation (mg/L) | 0.02 | Y | Y | Y | 6653 | 9/26/2005 | 8/19/2024 | 8/21/2024 8:46:12 PM | 1997-05-08 | 6/3/1998 6:51:46 PM | MPAK |  | Concentration | mg/L | Y | (fl. conc.) |  | Nitrogen Organic Total Dissolved (fl. conc.) |  | Nutrients; | Nutrients; EMS code: 1112 | Chemical | Dissolved | LAB
1113 | mg/L | Nitrogen (Kjeldahl) Total Dissolved | 1360 | Block Dig.Auto.Color | Block Dig.Auto.Color | 0.01 | Y | Y | Y | 1362 | 6/4/1997 | 3/27/2024 | 4/10/2024 5:01:15 PM | 1996-04-12 | 6/3/1998 6:51:33 PM | EMS |  | Concentration | mg/L | Y | (fl. conc.) |  | Nitrogen (Kjeldahl) Total Dissolved (fl. conc.) |  | Nutrients; | Nutrients; EMS code: 1113 | Chemical | Dissolved | LAB
1113 | mg/L | Nitrogen (Kjeldahl) Total Dissolved | DKNF | DKN,Filt,Block Dig.,Fluor. | This analysis is carried out using procedures adapted from APHA Method 4500-Norg D. "Block Digestion and Flow Injection Analysis". Dissolved Kjeldahl Nitrogen is determined using block digestion followed by Flow-injection analysis with fluorescence detection. | 0.05 | Y | Y | Y | 278 | 8/6/2015 | 7/18/2024 | 7/24/2024 7:36:07 PM | 2015-12-15 |  | RDGIBSON |  | Concentration | mg/L | Y | (fl. conc.) |  | Nitrogen (Kjeldahl) Total Dissolved (fl. conc.) |  | Nutrients; | Nutrients; EMS code: 1113 | Chemical | Dissolved | LAB
1113 | mg/L | Nitrogen (Kjeldahl) Total Dissolved | 2660 | Froz.Dig.Aut.Color | Froz.Dig.Aut.Color | 0.01 | Y | Y | Y | 3 |  |  | 10/23/1996 10:46:49 PM | 1996-04-12 | 6/3/1998 6:51:33 PM | EMS |  | Concentration | mg/L | Y | (fl. conc.) |  | Nitrogen (Kjeldahl) Total Dissolved (fl. conc.) |  | Nutrients; | Nutrients; EMS code: 1113 | Chemical | Dissolved | LAB
1113 | mg/L | Nitrogen (Kjeldahl) Total Dissolved | CS01 | System Calculation | System Calculation (mg/L) |  | N | Y | N | 31277 |  |  | 8/21/2024 8:46:13 PM | 1996-04-12 | 6/3/1998 6:51:43 PM | EMS |  | Concentration | mg/L | Y | (fl. conc.) |  | Nitrogen (Kjeldahl) Total Dissolved (fl. conc.) |  | Nutrients; | Nutrients; EMS code: 1113 | Chemical | Dissolved | LAB
1113 | mg/L | Nitrogen (Kjeldahl) Total Dissolved | EQ01 | EQUIS (mg/L) | EQUIS (mg/L) | 0.005 | N | Y | Y | 261 |  |  | 10/24/1996 3:16:44 PM | 1996-04-12 | 6/3/1998 6:51:45 PM | EMS |  | Concentration | mg/L | Y | (fl. conc.) |  | Nitrogen (Kjeldahl) Total Dissolved (fl. conc.) |  | Nutrients; | Nutrients; EMS code: 1113 | Chemical | Dissolved | LAB
1113 | mg/L | Nitrogen (Kjeldahl) Total Dissolved | 136A | HgSO4 Dig.Auto Color | HgSO4 Dig.Auto Color | 0.04 | Y | Y | Y | 272 |  |  | 5/6/1999 2:18:55 PM | 1996-04-12 | 6/3/1998 6:51:33 PM | EMS |  | Concentration | mg/L | Y | (fl. conc.) |  | Nitrogen (Kjeldahl) Total Dissolved (fl. conc.) |  | Nutrients; | Nutrients; EMS code: 1113 | Chemical | Dissolved | LAB
1113 | mg/L | Nitrogen (Kjeldahl) Total Dissolved | XCAL | Calculation (mg/L) | Calculation (mg/L) | 0.02 | Y | Y | Y | 17152 | 2/18/2005 | 5/4/2021 | 5/11/2021 4:36:09 PM | 1996-09-17 | 6/3/1998 6:51:45 PM | MPAK |  | Concentration | mg/L | Y | (fl. conc.) |  | Nitrogen (Kjeldahl) Total Dissolved (fl. conc.) |  | Nutrients; | Nutrients; EMS code: 1113 | Chemical | Dissolved | LAB
1114 | mg/L | Nitrogen Total Dissolved | EQ01 | EQUIS (mg/L) | EQUIS (mg/L) | 0.02 | N | Y | Y | 472 |  |  | 10/22/1996 9:23:30 PM | 1996-04-12 | 6/3/1998 6:51:33 PM | EMS |  | Concentration | mg/L | Y | (fl. conc.) |  | Nitrogen Total Dissolved (fl. conc.) |  | Nutrients; | Nutrients; EMS code: 1114 | Chemical | Dissolved | LAB
1114 | mg/L | Nitrogen Total Dissolved | F006 | K/Persulphate-UV Dig,Red'n,Col'm | Samples containing nitrogen compounds are digested and oxidized in-line to nitrate at 95 degrees Celsius by using an alkaline persulfate digestion solution and ultraviolet (UV) radiation. | 0.02 | Y | Y | Y | 7530 | 11/7/2007 | 12/1/2017 5:11:00 PM | 12/4/2017 5:01:12 PM | 2007-10-29 |  | RDGIBSON |  | Concentration | mg/L | Y | (fl. conc.) |  | Nitrogen Total Dissolved (fl. conc.) |  | Nutrients; | Nutrients; EMS code: 1114 | Chemical | Dissolved | LAB
1114 | mg/L | Nitrogen Total Dissolved | X156 | Auto UV Dig:Col'metric | Auto UV Dig:Col'metric | 0.02 | Y | Y | Y | 4 | 12/1/2010 | 12/1/2010 | 12/9/2010 3:55:26 PM | 1996-04-12 | 6/3/1998 6:51:33 PM | EMS |  | Concentration | mg/L | Y | (fl. conc.) |  | Nitrogen Total Dissolved (fl. conc.) |  | Nutrients; | Nutrients; EMS code: 1114 | Chemical | Dissolved | LAB
1114 | mg/L | Nitrogen Total Dissolved | XCAL | Calculation (mg/L) | Calculation (mg/L) | 0.04 | Y | Y | Y | 349 | 12/9/1992 4:16:00 PM | 4/10/2024 | 4/16/2024 2:12:02 PM | 1996-04-12 | 6/3/1998 6:51:42 PM | EMS |  | Concentration | mg/L | Y | (fl. conc.) |  | Nitrogen Total Dissolved (fl. conc.) |  | Nutrients; | Nutrients; EMS code: 1114 | Chemical | Dissolved | LAB
1114 | mg/L | Nitrogen Total Dissolved | X520 | F,persul,hydz-sulf,diaz,col | Filtered;persulfate oxidation;hydz-sul. reduction;diazotization;colorimetric |  | Y | Y | Y | 349 |  |  | 9/9/1999 12:47:44 PM | 1999-04-22 |  | NPEPPIN |  | Concentration | mg/L | Y | (fl. conc.) |  | Nitrogen Total Dissolved (fl. conc.) |  | Nutrients; | Nutrients; EMS code: 1114 | Chemical | Dissolved | LAB
1114 | mg/L | Nitrogen Total Dissolved | CS01 | System Calculation | System Calculation (mg/L) |  | N | Y | N | 9087 |  |  | 4/10/2024 5:01:15 PM | 1996-04-12 | 6/3/1998 6:51:43 PM | EMS |  | Concentration | mg/L | Y | (fl. conc.) |  | Nitrogen Total Dissolved (fl. conc.) |  | Nutrients; | Nutrients; EMS code: 1114 | Chemical | Dissolved | LAB
1114 | mg/L | Nitrogen Total Dissolved | F003 | FF,K/Persulphate Dig,Red'n,Col'm | The Total Nitrogen present in the sample is oxidized to nitrate by digesting the sample with potassium persulphate in an alkaline medium. The resulting Nitrate is reduced to Nitrite by passing it through a column containing granulated copper-cadmium. The nitrite then reacts with sulfanilamide under acidic conditions to form a diazo compound. This compound then complexes with N-1-naphthylethylenediamine dihydrochloride to form a reddish-purple azo dye, the absorbance of which is measured at 550nm. | 0.02 | Y | Y | Y | 18186 | 10/27/1998 6:17:00 PM | 8/18/2024 | 8/22/2024 4:39:06 PM | 1996-06-20 | 6/3/1998 6:51:44 PM | MPAK |  | Concentration | mg/L | Y | (fl. conc.) |  | Nitrogen Total Dissolved (fl. conc.) |  | Nutrients; | Nutrients; EMS code: 1114 | Chemical | Dissolved | LAB
1114 | mg/L | Nitrogen Total Dissolved | X183 | UV Oxid'n, Cd red'n, diazn, colorimetric | Organic nitrogenous compounds, including ammonia, are oxidized by ultraviolet irradiation to a mixture of nitrate and nitrite. This mixture then is all converted to nitrite in the cadmium reduction column. These nitrites react with sulphanilamide under acidic conditions to form a diazo compound. This compound compexes with N-1-naphthylethylenediamine dihydrochloride to form a reddish-purple azo dye, the absorbance of which is measured at 550 nm. | 0.01 | Y | Y | Y | 21681 |  |  | 6/18/2018 10:00:07 AM | 1996-04-12 | 8/9/1999 1:04:30 PM | EMS | NPEPPIN | Concentration | mg/L | Y | (fl. conc.) |  | Nitrogen Total Dissolved (fl. conc.) |  | Nutrients; | Nutrients; EMS code: 1114 | Chemical | Dissolved | LAB
1118 | mg/L | Phosphorus Ort.Dis-P | X157 | Auto Ascorbic Acid Red. | The filtered sample is reacted with a mixture of ammonium molybdate and potassium antimonyl tartrate in acid solution. Ascorbic acid is then added to produce a blue coloured product with an absorbance maximum at 880 nm. The absorbance of the solution is measured and the phosphorus concentration is determined by comparison with standards treated in the same manner. Automated procedure. | 0.002 | Y | Y | Y | 4207 | 9/18/1996 | 11/27/2015 | 7/30/2024 9:20:20 AM | 1996-04-12 | 6/3/1998 6:51:33 PM | EMS |  | Concentration | mg/L | Y | (fl. conc.) |  | Phosphorus Ort.Dis-P (fl. conc.) | 7723-14-0 | Nutrients; | Nutrients; EMS code: 1118 | Chemical |  | LAB
1118 | mg/L | Phosphorus Ort.Dis-P | EQ01 | EQUIS (mg/L) | EQUIS (mg/L) | 0.003 | N | Y | Y | 10866 |  |  | 10/24/1996 3:16:40 PM | 1996-04-12 | 6/3/1998 6:51:45 PM | EMS |  | Concentration | mg/L | Y | (fl. conc.) |  | Phosphorus Ort.Dis-P (fl. conc.) | 7723-14-0 | Nutrients; | Nutrients; EMS code: 1118 | Chemical |  | LAB
1118 | mg/L | Phosphorus Ort.Dis-P | X023 | Dig Manual Ascorbic Acid | Dig Manual Ascorbic Acid | 0.003 | Y | Y | Y | 39737 | 11/18/1996 9:30:00 AM | 11/18/1996 9:30:00 AM | 8/20/2024 7:35:41 AM | 1996-04-12 | 6/3/1998 6:51:33 PM | EMS |  | Concentration | mg/L | Y | (fl. conc.) |  | Phosphorus Ort.Dis-P (fl. conc.) | 7723-14-0 | Nutrients; | Nutrients; EMS code: 1118 | Chemical |  | LAB
1118 | mg/L | Phosphorus Ort.Dis-P | X453 | Stannous Cl2, Col'metric | Stannous Cl2; Col'metric | 0.003 | Y | Y | Y | 1013 | 5/5/1998 | 5/24/2019 5:03:00 AM | 5/30/2019 9:17:09 AM | 1996-04-12 | 6/3/1998 6:51:45 PM | EMS |  | Concentration | mg/L | Y | (fl. conc.) |  | Phosphorus Ort.Dis-P (fl. conc.) | 7723-14-0 | Nutrients; | Nutrients; EMS code: 1118 | Chemical |  | LAB
1118 | t/d | Phosphorus Ort.Dis-P | CALC | Calculation (t/d) | Calculation (t/d) | 0 | Y | Y | Y | 372 |  |  | 2/24/2012 12:45:47 PM | 1996-04-12 | 6/3/1998 6:51:33 PM | EMS |  | MassFlow | t/d | Y | (mass flow) |  | Phosphorus Ort.Dis-P (mass flow) | 7723-14-0 | Nutrients; | Nutrients; EMS code: 1118 | Chemical |  | LAB
1118 | mg/L | Phosphorus Ort.Dis-P | 1420 | Auto. Ascorbic - L Level | The ortho-phosphate reacts with ammonium molybdate and antimony tartrate in an acidic medium to form an antimony-phosphomolybdate complex, which is reduced to an intensely blue coloured complex by ascorbic acid. The absorbance of this complex is measured at 880nm. Note that for low level analyses are performed on the entire bottle of sample, not on an aliquot from the bottle. | 0.001 | Y | Y | Y | 24697 | 8/20/1997 6:19:00 PM | 8/22/2024 | 8/26/2024 3:51:07 PM | 1996-04-12 | 6/3/1998 6:51:41 PM | EMS |  | Concentration | mg/L | Y | (fl. conc.) |  | Phosphorus Ort.Dis-P (fl. conc.) | 7723-14-0 | Nutrients; | Nutrients; EMS code: 1118 | Chemical |  | LAB
1118 | ug/g | Phosphorus Ort.Dis-P | DI-P | 1:5 Soil Leach acids col | 1:5 Soil Leach acids col | 0.015 | Y | Y | Y | 3 | 12/7/2004 | 12/7/2004 | 1/5/2005 1:01:24 AM | 1996-10-13 | 6/3/1998 6:51:46 PM | EMS |  | MassPerMass | mg/kg | Y | (mass conc.) |  | Phosphorus Ort.Dis-P (mass conc.) | 7723-14-0 | Nutrients; | Nutrients; EMS code: 1118 | Chemical |  | LAB
1118 | mg/L | Phosphorus Ort.Dis-P | 1380 | Auto.Ascorbic Acid | The ortho-phosphate reacts with ammonium molybdate and antimony tartrate in an acidic medium to form an antimony-phosphomolybdate complex, which is reduced to an intensely blue coloured complex by ascorbic acid. The absorbance of this complex is measured at 880 nm. | 0.002 | Y | Y | Y | 99377 | 8/9/1996 6:16:00 PM | 8/14/2024 2:06:00 PM | 8/26/2024 8:06:03 AM | 1996-04-12 | 6/3/1998 6:51:33 PM | EMS |  | Concentration | mg/L | Y | (fl. conc.) |  | Phosphorus Ort.Dis-P (fl. conc.) | 7723-14-0 | Nutrients; | Nutrients; EMS code: 1118 | Chemical |  | LAB
1118 | mg/L | Phosphorus Ort.Dis-P | X096 | Ascorbic Acid | Ascorbic Acid | 0.003 | Y | Y | Y | 7420 | 9/30/1994 | 8/17/2024 3:52:00 PM | 8/25/2024 9:22:38 PM | 1996-04-12 | 6/3/1998 6:51:33 PM | EMS |  | Concentration | mg/L | Y | (fl. conc.) |  | Phosphorus Ort.Dis-P (fl. conc.) | 7723-14-0 | Nutrients; | Nutrients; EMS code: 1118 | Chemical |  | LAB
1118 | mg/L | Phosphorus Ort.Dis-P | X044 | Ion Chromatography | A clear (Total) or filtered (Dissolved) water sample is injected into a stream of carbonate-bicarbonate eluent and passed through a series of ion exchanger. The anions of interested are separated on the basis of their relative affinities for a low capacity , strongly basic anion exchanger (guard and separate columns). The separated anions are directed through the anion self-regenerating suppressor (ASRS). Hydronium ions in the suppressor will neutralize the eluent to its weakly ionized form and convert the separated anions in their conductive acid form. The separated anions are measured by conductivity. They are identified on the basis of retention time as compared to standards. Quantitation is by measurement of peak area of peak height against a calibration curve generated from known standards. | 0.05 | Y | Y | Y | 21072 | 12/20/1996 7:12:00 PM | 5/17/2024 5:10:00 PM | 8/25/2024 9:47:36 PM | 1997-01-16 | 6/3/1998 6:51:46 PM | MPAK |  | Concentration | mg/L | Y | (fl. conc.) |  | Phosphorus Ort.Dis-P (fl. conc.) | 7723-14-0 | Nutrients; | Nutrients; EMS code: 1118 | Chemical |  | LAB
1118 | mg/L | Phosphorus Ort.Dis-P | 2700 | Froz.Auto.Ascorbic Acid | Froz.Auto.Ascorbic Acid | 0.003 | Y | Y | Y | 1422 |  |  | 10/24/1996 3:06:02 PM | 1996-04-12 | 6/3/1998 6:51:33 PM | EMS |  | Concentration | mg/L | Y | (fl. conc.) |  | Phosphorus Ort.Dis-P (fl. conc.) | 7723-14-0 | Nutrients; | Nutrients; EMS code: 1118 | Chemical |  | LAB
1120 | mg/L | Silica Reactive Diss | X535 | Auto moly-blue, Col'm,(not ascorbic) | Colourimetry using heteropoly blue method on an autoanalyzer. If turbid the sample is passed through a 0.45 u membrane filter. If the sample is sea water all standard SIO2 solutions are prepared with synthetic sea water instead of distilled water. A filtrate aliquot is mixed with a solution of (NH4)6MO7O24 in dil. H2SO4. This mixture is then mixed with oxalic acid solution, followed by a 1-Amino-2-Naphthol-4-Sul-Phonic acid solution, containing NA2SO3 and NAHSO3. The resulting heteropoly blue colour is measured spectrophotometrically at 660 mu and compared with identically prepared standard SIO2 solutions. Similar to ascorbic acid method, but using 1-Amino-2-Naphthol-4-Sulphuric acid as the reducing agent. | 0.02 | Y | Y | Y | 4539 |  |  | 9/9/1999 12:47:48 PM | 1999-08-06 |  | NPEPPIN |  | Concentration | mg/L | Y | (fl. conc.) |  | Silica Reactive Diss (fl. conc.) | 7631-86-9 | Nutrients; | Nutrients; EMS code: 1120 | Chemical |  | LAB
1120 | mg/L | Silica Reactive Diss | EQ01 | EQUIS (mg/L) | EQUIS (mg/L) | 0.2 | N | Y | Y | 8477 |  |  | 10/24/1996 3:16:02 PM | 1996-04-12 | 6/3/1998 6:51:33 PM | EMS |  | Concentration | mg/L | Y | (fl. conc.) |  | Silica Reactive Diss (fl. conc.) | 7631-86-9 | Nutrients; | Nutrients; EMS code: 1120 | Chemical |  | LAB
1120 | mg/L | Silica Reactive Diss | X160 | Auto Molybdate-Blue Col.(Ascorbic) | Not all forms of silica in solution will react to give the silicomolybdate complex. Silicic acid readily polymerizes at the pH of seawater (pH=8.0), and only straight chain polymers of relatively short length will react with molybdate at any appreciable speed. Therefore, the term "Reactive Silica" is best used. This automated procedure for the determination of reactive silica is based on the reduction by ascorbic acid of silicomolybdate in acidic solution to heteropoly blue. The colour is measured at 660 nm in a 1.5 cm cell. Oxalic acid is introduced to the sample stream before the addition of ascorbic acid to eliminate interference from phosphates. The sample should be collected in a clean linear polyethylene bottle. The samples are not preserved. The samples should be free from turbidity. Allow the samples to settle before decanting an aliquot, or filter the samples through 0.45 m filter paper. Samples are stored at 4C prior to analysis. The samples should be analysed within 28 days. Note: Do not use glass containers. | 0.1 | Y | Y | Y | 21527 | 8/7/1996 6:17:00 PM | 8/25/2024 | 8/26/2024 1:03:06 PM | 1996-04-12 | 6/3/1998 6:51:33 PM | EMS |  | Concentration | mg/L | Y | (fl. conc.) |  | Silica Reactive Diss (fl. conc.) | 7631-86-9 | Nutrients; | Nutrients; EMS code: 1120 | Chemical |  | LAB
1121 | mg/L | Sulfate Dissolved | EQ01 | EQUIS (mg/L) | EQUIS (mg/L) | 0 | N | Y | Y | 10925 |  |  | 10/24/1996 3:16:55 PM | 1996-04-12 | 6/3/1998 6:51:33 PM | EMS |  | Concentration | mg/L | Y | (fl. conc.) |  | Sulfate Dissolved (fl. conc.) | 14808-79-8 | Major Ions; | Major Ions; EMS code: 1121 | Chemical | Dissolved | LAB
1121 | mg/L | Sulfate Dissolved | X058 | Turbimetric | Turbimetric | 1 | Y | Y | Y | 40626 | 6/29/2006 | 8/22/2024 3:22:00 PM | 8/26/2024 9:40:02 AM | 1996-10-11 | 6/3/1998 6:51:46 PM | NPEPPIN |  | Concentration | mg/L | Y | (fl. conc.) |  | Sulfate Dissolved (fl. conc.) | 14808-79-8 | Major Ions; | Major Ions; EMS code: 1121 | Chemical | Dissolved | LAB
1121 | mg/L | Sulfate Dissolved | X061 | BaCl2,Gravimetric | Sulfate is precipitated as barium sulfate (BaSO4) in HCl medium by the addition of barium chloride. After the digestion period, the precipitate is filtered, washed with hot water until chloride free, ignited and weighed as BaSO4. | 1 | Y | Y | Y | 605 |  |  | 3/9/1998 11:07:49 AM | 1996-04-12 | 6/3/1998 6:51:33 PM | EMS |  | Concentration | mg/L | Y | (fl. conc.) |  | Sulfate Dissolved (fl. conc.) | 14808-79-8 | Major Ions; | Major Ions; EMS code: 1121 | Chemical | Dissolved | LAB
1121 | t/d | Sulfate Dissolved | CALC | Calculation (t/d) | Calculation (t/d) | 0 | Y | Y | Y | 661 |  |  | 12/14/2006 2:22:12 PM | 1996-04-12 | 6/3/1998 6:51:45 PM | EMS |  | MassFlow | t/d | Y | (mass flow) |  | Sulfate Dissolved (mass flow) | 14808-79-8 | Major Ions; | Major Ions; EMS code: 1121 | Chemical | Dissolved | LAB
1121 | mg/L | Sulfate Dissolved | 1347 | Sulfate by IC (seawater) | Sulfate by IC (seawater) | 30 | Y | Y | Y | 85 | 8/14/2019 | 4/25/2024 | 6/7/2024 10:03:16 AM | 2019-09-30 |  | KATFRASE |  | Concentration | mg/L | Y | (fl. conc.) |  | Sulfate Dissolved (fl. conc.) | 14808-79-8 | Major Ions; | Major Ions; EMS code: 1121 | Chemical | Dissolved | LAB
1121 | mg/L | Sulfate Dissolved | X044 | Ion Chromatography | A clear (Total) or filtered (Dissolved) water sample is injected into a stream of carbonate-bicarbonate eluent and passed through a series of ion exchanger. The anions of interested are separated on the basis of their relative affinities for a low capacity , strongly basic anion exchanger (guard and separate columns). The separated anions are directed through the anion self-regenerating suppressor (ASRS). Hydronium ions in the suppressor will neutralize the eluent to its weakly ionized form and convert the separated anions in their conductive acid form. The separated anions are measured by conductivity. They are identified on the basis of retention time as compared to standards. Quantitation is by measurement of peak area of peak height against a calibration curve generated from known standards. | 0.3 | Y | Y | Y | 132239 | 9/24/1996 10:08:00 AM | 8/24/2024 | 8/26/2024 4:53:05 PM | 1996-04-12 | 8/15/2019 11:50:43 AM | EMS | KATFRASE | Concentration | mg/L | Y | (fl. conc.) |  | Sulfate Dissolved (fl. conc.) | 14808-79-8 | Major Ions; | Major Ions; EMS code: 1121 | Chemical | Dissolved | LAB
1121 | mg/L | Sulfate Dissolved | X064 | BaCl2,Turbimetric | Sulfate ion is converted to a barium sulfate suspension under controlled conditions. The resulting turbidity is determined using a nephelometer or spectrophotometer and compared to a curve prepared from standard sulfate solutions. | 1 | Y | Y | Y | 16745 | 1/15/2007 8:52:00 AM | 7/6/2016 | 8/26/2024 2:41:32 PM | 1996-04-12 | 6/3/1998 6:51:45 PM | EMS |  | Concentration | mg/L | Y | (fl. conc.) |  | Sulfate Dissolved (fl. conc.) | 14808-79-8 | Major Ions; | Major Ions; EMS code: 1121 | Chemical | Dissolved | LAB
1121 | mg/L | Sulfate Dissolved | 1400 | Auto Methyl Thymol Blue | After being passed through a cation-exchange column, the sample is reacted with an alcohol solution of barium chloride and methylthymol blue (MTB) at pH 2.5-3.0 to form barium sulfate. The pH of this solution is raised to 12.5-13.0 so that the excess barium reacts with MTB. The uncomplexed MTB is equivalent to the amount of sulfate present. | 1 | Y | Y | Y | 18618 | 8/10/1993 | 10/22/2003 | 4/30/2004 1:59:49 PM | 1996-04-12 | 6/3/1998 6:51:33 PM | EMS |  | Concentration | mg/L | Y | (fl. conc.) |  | Sulfate Dissolved (fl. conc.) | 14808-79-8 | Major Ions; | Major Ions; EMS code: 1121 | Chemical | Dissolved | LAB
1121 | ug/g | Sulfate Dissolved | DI-S | 1:5 Soil Leach, IC | 1:5 Soil Leach; IC | 5 | Y | Y | Y | 17 | 12/6/2004 | 10/31/2023 | 11/23/2023 7:26:23 AM | 1996-10-13 | 6/3/1998 6:51:46 PM | EMS |  | MassPerMass | mg/kg | Y | (mass conc.) |  | Sulfate Dissolved (mass conc.) | 14808-79-8 | Major Ions; | Major Ions; EMS code: 1121 | Chemical | Dissolved | LAB
1122 | mg/L | Surfactants Diss. | EQ01 | EQUIS (mg/L) | EQUIS (mg/L) | 0 | N | Y | Y | 203 |  |  | 10/22/1996 9:38:47 PM | 1996-04-12 | 6/3/1998 6:51:33 PM | EMS |  | Concentration | mg/L | Y | (fl. conc.) |  | Surfactants Diss. (fl. conc.) |  | Organics; | Organics; EMS code: 1122 | Chemical |  | LAB
1123 | mg/L | Tannin &Lignin Diss. | 0950 | Heteropoly Acid | Heteropoly Acid | 0.1 | Y | Y | Y | 39 |  |  | 2/13/1998 1:21:03 AM | 1996-04-12 | 6/3/1998 6:51:46 PM | EMS |  | Concentration | mg/L | Y | (fl. conc.) |  | Tannin &Lignin Diss. (fl. conc.) |  | Organics; | Organics; EMS code: 1123 | Chemical |  | LAB
1123 | mg/L | Tannin &Lignin Diss. | EQ01 | EQUIS (mg/L) | EQUIS (mg/L) | 0 | N | Y | Y | 22 |  |  | 10/22/1996 9:46:44 PM | 1996-04-12 | 6/3/1998 6:51:45 PM | EMS |  | Concentration | mg/L | Y | (fl. conc.) |  | Tannin &Lignin Diss. (fl. conc.) |  | Organics; | Organics; EMS code: 1123 | Chemical |  | LAB
1124 | mg/L | Carbon Dissolved Inorganic | CS01 | System Calculation | System Calculation (mg/L) |  | N | Y | N | 132 |  |  | 6/18/2018 10:00:07 AM | 1996-04-12 | 6/3/1998 6:51:43 PM | EMS |  | Concentration | mg/L | Y | (fl. conc.) |  | Carbon Dissolved Inorganic (fl. conc.) |  | Carbon; | Carbon; EMS code: 1124 | Chemical | Dissolved | LAB
1124 | mg/L | Carbon Dissolved Inorganic | EQ01 | EQUIS (mg/L) | EQUIS (mg/L) | 1 | N | Y | Y | 2 |  |  | 10/22/1996 11:00:26 PM | 1996-04-12 | 6/3/1998 6:51:33 PM | EMS |  | Concentration | mg/L | Y | (fl. conc.) |  | Carbon Dissolved Inorganic (fl. conc.) |  | Carbon; | Carbon; EMS code: 1124 | Chemical | Dissolved | LAB
1124 | mg/L | Carbon Dissolved Inorganic | X171 | Carbon Anal.,Infra Red | Analysis is comprised of 2 phases, the determination of Total Carbon (TC) or Dissolved Carbon (DC) and the determination of Total Inorganic Carbon (TIC) or Dissolved Inorganic Carbon (DIC). Total Organic Carbon (TOC) content, or Dissolved Organic Carbon (DOC) is then calculated by difference. A series of External Standard solutions, Potassium Biphthalate for TC (DC) and Sodium Carbonate/Sodium Bicarbonate for TIC (DIC) is used for instrument calibration. During TC (DC) analysis, a small portion of the sample (or standard) solution is introduced into a combustion tube, containing a Platinum catalyst at an elevated temperature with an Oxygen gas carrier, via an automated autosampler/injector. The sample's TC (DC) component is combusted/oxidized to form CO2. The CO2 flows through an IC reaction vessel, is cooled and dried by an electronic dehumidifier and passes through a halogen scrubber into a Non-dispersive Infrared Detector (NDIR). The NDIR produces an analog signal over a constant time base that a data processor converts into a peak, the area of which is proportional to the TC concentration of the sample. The TC (DC) concentration of a sample is determined electronically from a calibration curve expressing the relationship between the peak area of a standard and its concentration. Similarly, for the TIC (DIC) phase of the analysis, another small portion of sample (or standard) solution is introduced directly into the IC reaction vessel, containing a Phosphoric Acid solution through which carrier gas is bubbled. However, here only the IC component of the sample reacts to form CO2. The TOC (DOC) concentration is automatically determined by the data processor from the difference between TC (DC) and TIC (DOC) concentrations. | 0.5 | Y | Y | Y | 4053 | 7/30/1997 9:28:00 AM | 8/21/2007 1:54:00 PM | 9/11/2019 8:46:21 AM | 1997-02-13 | 6/3/1998 6:51:46 PM | MPAK |  | Concentration | mg/L | Y | (fl. conc.) |  | Carbon Dissolved Inorganic (fl. conc.) |  | Carbon; | Carbon; EMS code: 1124 | Chemical | Dissolved | LAB
1124 | mg/L | Carbon Dissolved Inorganic | F127 | PersulDig:CO2Purge/Col'm,550nm | Total and Dissolved Inorganic Carbon by Automated Persulfate Colourimetric Method (TAACS 800 Industrial Method No. 860-87T, January 1989): The sample is mixed with phosphate buffer and persulphate in order to digest the inorganic compounds and convert them to CO2 Dialysis of the resultant CO2 through a gas-permeable silicone membrane to combine with a phenolphthalein solution causes a decrease on colour of the phenophthalein solution proportional to the original inorganic carbon concentration. This colour change is measured at a wavelength of 550nm. | 0.5 | Y | Y | Y | 930 | 4/9/2010 12:00:00 PM | 8/14/2024 | 8/19/2024 11:34:07 AM | 2010-08-25 |  | RDGIBSON |  | Concentration | mg/L | Y | (fl. conc.) |  | Carbon Dissolved Inorganic (fl. conc.) |  | Carbon; | Carbon; EMS code: 1124 | Chemical | Dissolved | LAB
1124 | mg/L | Carbon Dissolved Inorganic | XCAL | Calculation (mg/L) | Calculation (mg/L) | 0.5 | Y | Y | Y | 625 | 7/31/1996 12:36:00 PM | 7/6/2016 5:10:00 PM | 7/8/2016 10:37:08 AM | 1996-04-12 | 6/3/1998 6:51:41 PM | EMS |  | Concentration | mg/L | Y | (fl. conc.) |  | Carbon Dissolved Inorganic (fl. conc.) |  | Carbon; | Carbon; EMS code: 1124 | Chemical | Dissolved | LAB
1125 | mg/L | Sulfide Dissolved | X257 | Colorimetric | Colorimetric. General use for several parameters. E.g.Total Sulfide is determined by spectrophotometer using the methlyene blue colourimetric method. The H2S calculation converts Total Sulphide as (S2-) and reports it as Total Sulphide as (H2S). | 0.02 | Y | Y | Y | 1609 | 1/13/2015 4:00:00 PM | 4/24/2019 3:30:00 PM | 7/25/2024 2:48:45 PM | 2012-03-07 |  | RDGIBSON |  | Concentration | mg/L | Y | (fl. conc.) |  | Sulfide Dissolved (fl. conc.) | 18496-25-8 | Volatile Organic Compounds (VOCs); | Volatile Organic Compounds (VOCs); EMS code: 1125 | Chemical | Dissolved | LAB
1126 | mg/L | Carbon Total Dissolved | CS01 | System Calculation | System Calculation (mg/L) |  | N | Y | N | 5598 |  |  | 8/19/2024 11:34:07 AM | 1996-04-12 | 6/3/1998 6:51:43 PM | EMS |  | Concentration | mg/L | Y | (fl. conc.) |  | Carbon Total Dissolved (fl. conc.) |  | Carbon; | Carbon; EMS code: 1126 | Chemical | Dissolved | LAB
1126 | mg/L | Carbon Total Dissolved | X171 | Carbon Anal.,Infra Red | Analysis is comprised of 2 phases, the determination of Total Carbon (TC) or Dissolved Carbon (DC) and the determination of Total Inorganic Carbon (TIC) or Dissolved Inorganic Carbon (DIC). Total Organic Carbon (TOC) content, or Dissolved Organic Carbon (DOC) is then calculated by difference. A series of External Standard solutions, Potassium Biphthalate for TC (DC) and Sodium Carbonate/Sodium Bicarbonate for TIC (DIC) is used for instrument calibration. During TC (DC) analysis, a small portion of the sample (or standard) solution is introduced into a combustion tube, containing a Platinum catalyst at an elevated temperature with an Oxygen gas carrier, via an automated autosampler/injector. The sample's TC (DC) component is combusted/oxidized to form CO2. The CO2 flows through an IC reaction vessel, is cooled and dried by an electronic dehumidifier and passes through a halogen scrubber into a Non-dispersive Infrared Detector (NDIR). The NDIR produces an analog signal over a constant time base that a data processor converts into a peak, the area of which is proportional to the TC concentration of the sample. The TC (DC) concentration of a sample is determined electronically from a calibration curve expressing the relationship between the peak area of a standard and its concentration. Similarly, for the TIC (DIC) phase of the analysis, another small portion of sample (or standard) solution is introduced directly into the IC reaction vessel, containing a Phosphoric Acid solution through which carrier gas is bubbled. However, here only the IC component of the sample reacts to form CO2. The TOC (DOC) concentration is automatically determined by the data processor from the difference between TC (DC) and TIC (DOC) concentrations. | 0.5 | Y | Y | Y | 135 |  |  | 6/18/2018 10:00:07 AM | 1997-01-14 | 6/3/1998 6:51:46 PM | MPAK |  | Concentration | mg/L | Y | (fl. conc.) |  | Carbon Total Dissolved (fl. conc.) |  | Carbon; | Carbon; EMS code: 1126 | Chemical | Dissolved | LAB
1126 | mg/L | Carbon Total Dissolved | 0911 | Beckman/IR Det-cellulose | Beckman/IR Det-cellulose | 3 | Y | Y | Y | 23 |  |  | 10/24/1996 1:54:13 PM | 1996-04-12 | 6/3/1998 6:51:33 PM | EMS |  | Concentration | mg/L | Y | (fl. conc.) |  | Carbon Total Dissolved (fl. conc.) |  | Carbon; | Carbon; EMS code: 1126 | Chemical | Dissolved | LAB
1127 | % | Fixed Solids | E169 | APHA 2540 G (mod) | Fixed Solids (FS) are determined by drying a pre-weighed sample in a pre-weighed dish to constant weight in an oven at 104 ? 1?C. This residue is ignited to constant weight at 550?C. The remaining solids represent the Fixed Solids (FS). Results are reported as the percentage of the total solids. | 0.1 | Y | Y | Y | 19 | 4/27/2022 | 6/12/2024 | 6/18/2024 11:02:03 PM | 2021-08-17 |  | RDGIBSON |  | DimensionlessRatio | % | N | (ratio) |  | Fixed Solids (ratio) |  | Physical Properties; | Physical Properties; EMS code: 1127 | Physical |  | LAB
1127 | mg/L | Fixed Solids | X322 | APHA Method 2540 | This analysis is carried out using procedures adapted from APHA Method 2540 "Solids". Solids are determined gravimetrically. Fixed Solids is determined by ashing (at 550C) the residue of a sample evaporated to dryness. | 10 | Y | Y | Y | 1 | 12/18/2019 | 12/18/2019 | 12/24/2019 9:21:13 AM | 2017-10-24 |  | JKROGH |  | Concentration | mg/L | Y | (fl. conc.) |  | Fixed Solids (fl. conc.) |  | Physical Properties; | Physical Properties; EMS code: 1127 | Physical |  | LAB
1268 | mg/L | Arochlor 1268 (PCB) CHLOROBIPHENYL (68% CL) | F018 | Hexane extr,GC/ECD | Effluent and swab samples are extracted with hexane, the extract concentrated, and then treated using Florisil, concentrated Sulfuric Acid, activated Copper and Mercury (as required) to remove interferences. The samples are then analysed using Megabore Column Gas Chromatography with Electron Capture Detection. Oil samples are first diluted with Hexane then treated in the manner outlined for effluent and swab samples. The samples are then analysed using Megabore Column Gas Chromatography with Electron Capture Detection. |  | Y | Y | Y | 1 | 9/5/2023 6:46:00 PM | 9/5/2023 6:46:00 PM | 9/11/2023 5:22:03 PM | 2009-01-09 |  | RDGIBSON |  | Concentration | mg/L | Y | (fl. conc.) |  | Arochlor 1268 (PCB) CHLOROBIPHENYL (68% CL) (fl. conc.) | 11100-14-4 | Polychlorinated biphenylds (PCBs); | Polychlorinated biphenylds (PCBs); EMS code: 1268 | Chemical |  | LAB
1268 | mg/L | Arochlor 1268 (PCB) CHLOROBIPHENYL (68% CL) | E660 | EPA 8270E (mod) | EPA 8270E (mod) Miscellaneous Pesticides by GC-MS | 2e-05 | Y | Y | Y | 8 | 5/24/2023 | 11/1/2023 | 11/16/2023 3:38:09 PM | 2023-08-16 |  | EMS |  | Concentration | mg/L | Y | (fl. conc.) |  | Arochlor 1268 (PCB) CHLOROBIPHENYL (68% CL) (fl. conc.) | 11100-14-4 | Polychlorinated biphenylds (PCBs); | Polychlorinated biphenylds (PCBs); EMS code: 1268 | Chemical |  | LAB
1268 | ug/g | Arochlor 1268 (PCB) CHLOROBIPHENYL (68% CL) | X383 | Extr, GC/MS or GC/ECD | Determination of Polychloriated Biphenyls (PCBs) In Soil by GC/ECD. PURPOSE/PRINCIPLE OF METHOD: This document details the requirements and procedures for the analysis of polychlorinated biphenyls (PCBs) in soil by gas chromatography electron capture detector (GC/ECD). A homogenized soil sample is extracted using dichloromethane (DCM). The extract may be cleaned up using column chromatography with 2% Deactivated Florisil. When contamination with sulphur is suspected, further cleanup using copper is required. Analysis is performed using a GC/ECD. Similar methods have been developed for other matrices including waste and transformer oils, water, swabs, florisil tubes, biota. Generally, the sample extraction procedures differ while the GC analysis methods are similar or identical for all matrices. SCOPE: This procedure describes the determination of PCBs in soil and sediments. Other materials such as wood waste may also be analyzed by this method. DETECTION LIMITS: Target Compound Reported Detection Limit (mg/g) Arochlor 1242 0.03, Arochlor 1248 0.03, Arochlor 1254 0.03, Arochlor 1260 0.03, INTERFERENCES: Interferences (e.g. co-eluting peaks on the chromatogram) may be due to sulphur, phthalate esters, or from chlorinated compounds. Suspect samples must be reanalyzed using appropriate cleanup steps. SAMPLE HANDLING & PRESERVATION: Sample Requirements: Whenever possible, samples must completely fill the bottle such that little or no headspace is present to eliminate the loss of volatile hydrocarbon components. (~125g of sample) Sample Containers: Glass jars that have been solvent rinsed with acetone then hexane, then dried. Shipment and Storage: Samples are to be refrigerated (at 4±2°C) during shipping and prior to analysis. Samples suspected to contain high levels of analyte are to be segregated from other samples. Holding Time: Holding time prior to extraction is 14 days. REFERENCE: U.S. EPA Method 8080:Organochlorine Pesticides and PCBs Analysis. | 0.02 | Y | Y | Y | 46 | 3/16/2020 | 9/7/2022 | 1/11/2024 1:07:06 PM | 2020-08-28 |  | RDGIBSON |  | MassPerMass | mg/kg | Y | (mass conc.) |  | Arochlor 1268 (PCB) CHLOROBIPHENYL (68% CL) (mass conc.) | 11100-14-4 | Polychlorinated biphenylds (PCBs); | Polychlorinated biphenylds (PCBs); EMS code: 1268 | Chemical |  | LAB
1268 | mg/L | Arochlor 1268 (PCB) CHLOROBIPHENYL (68% CL) | E-8082 | EPA-8082A | See reference for EPA method 8082A POLYCHLORINATED BIPHENYLS (PCBs) BY GAS CHROMATOGRAPHY |  | Y | Y | Y | 14 | 12/28/2021 | 1/11/2022 | 9/2/2022 6:21:52 PM | 2022-04-20 |  | EMS |  | Concentration | mg/L | Y | (fl. conc.) |  | Arochlor 1268 (PCB) CHLOROBIPHENYL (68% CL) (fl. conc.) | 11100-14-4 | Polychlorinated biphenylds (PCBs); | Polychlorinated biphenylds (PCBs); EMS code: 1268 | Chemical |  | LAB
1461 | %(V/V) | Bio.LC10:Fish | X068 | Rainbow/96hr(Vol) (RM13) | (RM13) This method measures the concentration (expressed volume/volume) of sample at which 50% mortality (for LC50), at 20% mortality (for LC20), and at 10% mortality (for LC10) occurs. No pH adjustment is permitted. | 0 | Y | Y | Y | 193 | 7/20/1998 12:00:00 PM | 5/21/2001 12:00:00 PM | 7/25/2024 12:46:40 PM | 1996-04-12 | 6/3/1998 6:51:45 PM | EMS |  | DimensionlessRatio | % | N | (ratio) | (by vol.) | Bio.LC10:Fish (by vol.) (ratio) |  | Toxicity; | Toxicity; EMS code: 1461 | Chemical |  | LAB
1461 | % (Mortality) | Bio.LC10:Fish | FA13 | RM13/Rainbow/96hr(%mortality)/100% conc. | Rainbow/96hr(%mortality)/100% concentration. (RM13) Biological Test Method: Reference Method for Determining Acute Lethality of Effluents to Rainbow Trout; EPS 1/RM/13 Second Edition December 2000 (with May 2007 and February 2016 amendments). Rainbow Trout are exposed in a single concentration of the sample and mortality is recorded over a period of 96 hours at 15C. The pass/fail single concentration test uses undiluted material at 100% concentration and determines the percent mortality after 96 hours. (Without pH Adjustment) | 0 | Y | Y | Y | 88 |  |  | 3/27/2009 4:32:37 PM | 1998-11-04 |  | NPEPPIN |  | DimensionlessRatio | % | N | (ratio) | (mortality) | Bio.LC10:Fish (mortality) (ratio) |  | Toxicity; | Toxicity; EMS code: 1461 | Chemical |  | LAB
1461 | mg/L | Bio.LC10:Fish | FA01 | Coho/96 hr (wt)/Salinity adj. | Effluents having salinity >10 ppt must be run with either Coho or Chinook suitably aclimated to appropriate salinity. If this bioassay is performed on lower salinity samples, all test tanks must be adjusted to similar salinity. | 0 | Y | Y | Y | 16 |  |  | 2/25/2009 4:42:49 PM | 1997-02-14 | 6/3/1998 6:51:46 PM | MPAK |  | Concentration | mg/L | Y | (fl. conc.) |  | Bio.LC10:Fish (fl. conc.) |  | Toxicity; | Toxicity; EMS code: 1461 | Chemical |  | LAB
2016 | mg/L | Chlorine Dem. Free | THMHAA | THM/HAA CALC | THM/HAA CALC | 0.1 | Y | Y | Y | 2 | 3/18/2023 4:37:00 PM | 5/23/2023 12:06:00 PM | 6/8/2023 12:14:56 PM | 2023-03-31 |  | EMS |  | Concentration | mg/L | Y | (fl. conc.) |  | Chlorine Dem. Free (fl. conc.) | 7782-50-5 |  | EMS code: 2016 | Chemical |  | LAB
2101 | mg/L | Bromate Dissolved | F078 | Oxy-halide,ICP,Cond. detector | An approximately 225 uL of sample is introduced into an ion chromatograph (IC). The anions of interest are separated and measured, using a system comprised of a guard column, analytical column, suppressor device, conductivity detector, a postcolumn reagent delivery system (pneumatically controlled), a heated postcolumn reaction coil, and an ultraviolet/visible (UV/VIS) absorbance detector. They are identified on the basis of retention time as compared to standards. Quantitation is by measurement of peak area against a calibration curve generated from known standards. | 10 | Y | Y | Y | 1400 |  |  | 7/29/2024 2:02:49 PM | 2000-03-01 |  | NPEPPIN |  | Concentration | mg/L | Y | (fl. conc.) |  | Bromate Dissolved (fl. conc.) | 15541-45-4 | Inorganics; | Inorganics; EMS code: 2101 | Chemical | Dissolved | LAB
2105 | mg/L | Cyanide(SAD) + SCN | AA04 | UV,Auto-CN,Pyr-Barb - low level | Metallocyanide compounds are converted to simple cyanides by means of on-line UV irradiation and distillation under acidic conditions. The hydrogen cyanide is subsequently reacted with chloramine-T to form cyanogen chloride which reacts with pyridine and barbituric acid to form a red coloured complex which is measured colorimetrically at 570nm. Interferences are eliminated or reduced to a minimum by the on-line distillation procedure. Sulfide interferences may be eliminated by precipitating samples suspected of high H2S levels with the addition of PbCO3. This procedure does yield positive CN result for thiocyanate. | 0.0005 | Y | Y | Y | 3053 | 1/21/1997 12:54:00 PM | 1/11/2002 7:51:00 AM | 5/26/2003 2:05:08 PM | 1996-04-12 | 6/3/1998 6:51:41 PM | EMS |  | Concentration | mg/L | Y | (fl. conc.) |  | Cyanide(SAD) + SCN (fl. conc.) | 57-12-5, 302-04-5 | Inorganics; | Inorganics; EMS code: 2105 | Chemical |  | LAB
2105 | mg/L | Cyanide(SAD) + SCN | CS01 | System Calculation | System Calculation (mg/L) |  | N | Y | N | 28233 |  |  | 1/11/2023 5:01:27 PM | 1996-04-12 | 6/3/1998 6:51:44 PM | EMS |  | Concentration | mg/L | Y | (fl. conc.) |  | Cyanide(SAD) + SCN (fl. conc.) | 57-12-5, 302-04-5 | Inorganics; | Inorganics; EMS code: 2105 | Chemical |  | LAB
2107 | mg/L | Hardness Total (Extractable) | CCAL | Calculation (Ca, Mg) | Dissolved and extractable water samples are ready for analysis once collected. Total samples are autoclaved with HNO3 and hydrochloric acid (HCl) in capped Teflon vials. Analysis is performed via Inductively Coupled Argon Plasma Atomic Emission Spectrometry (ICP-AES) (simultaneous multielement analysis), Graphite Furnace Atomic Absorption Spectrometry (GF-AAS) (low level analysis) and/or Flame Atomic Absorption Spectrometry (F-AAS). Total Hardness is calculated from dissolved Calcium and dissolved Magnesium according to the following formula: Total Hardness=2.497248*Ca + 4.116885*Mg (where Total Hardness is expressed as mg/L of CaCO3, and Calcium and Magnesium are expressed as mg/L Ca and mg/L Mg respectively. Check Media/Pres. codes for consistency with regard to acid preservation. | 1 | Y | Y | Y | 3811 | 8/23/1996 6:04:00 PM | 12/10/2011 12:01:00 PM | 9/11/2019 8:47:41 AM | 1998-04-21 | 6/3/1998 6:51:48 PM | NPEPPIN |  | Concentration | mg/L | Y | (fl. conc.) |  | Hardness Total (Extractable) (fl. conc.) |  | Major Ions; | Major Ions; EMS code: 2107 | Chemical | Extractable | LAB
2107 | mg/L | Hardness Total (Extractable) | MCAL | Calculation (Ca,Mg,Sr,Fe,Al,Zn,Mn,Cu,Ba,Pb) | Dissolved and extractable water samples are ready for analysis once collected. Total samples are autoclaved with HNO3 and hydrochloric acid (HCl) in capped Teflon vials. Analysis is performed via Inductively Coupled Argon Plasma Atomic Emission Spectrometry (ICP-AES) (simultaneous multielement analysis), Graphite Furnace Atomic Absorption Spectrometry (GF-AAS) (low level analysis) and/or Flame Atomic Absorption Spectrometry (F-AAS). Total Hardness is calculated from the dissolved (or extractable as a substitute) fractions of Ca, Mg, Sr, Fe, Al, Zn, Mn, Cu, Ba, Pb, according to the following formula: Hardness Total = 2.497248*Ca + 4.116885*Mg + 1.142316*Sr + 1.792213*Fe + 3.709568*Al + 1.531126*Zn + 1.821866*Mn + 1.575224*Cu + 0.728773*Ba + 0.483082*Pb (where Total Hardness is expressed as mg/L of CaCO3, and each element is expressed in mg/L. Check Media/Pres. codes for consistency with regard to acid preservation. | 1 | Y | Y | Y | 2242 | 1/12/2000 10:50:00 AM | 3/19/2008 11:25:00 AM | 9/11/2019 8:47:41 AM | 1998-04-21 | 6/3/1998 6:51:48 PM | NPEPPIN |  | Concentration | mg/L | Y | (fl. conc.) |  | Hardness Total (Extractable) (fl. conc.) |  | Major Ions; | Major Ions; EMS code: 2107 | Chemical | Extractable | LAB
2107 | mg/L | Hardness Total (Extractable) | HCAL | Unfilt,Calculation (Ca,Mg) | 'Shortcut Metals': no acid preservation, no filtration, no digestion for following extractable elements: Ca, K, Mg, Na and Si as analyzed by ICP. | 1 | Y | Y | Y | 496 | 1/9/1997 6:05:00 PM | 3/23/2012 3:11:00 AM | 3/28/2012 10:45:29 AM | 1998-04-21 | 6/3/1998 6:51:48 PM | NPEPPIN |  | Concentration | mg/L | Y | (fl. conc.) |  | Hardness Total (Extractable) (fl. conc.) |  | Major Ions; | Major Ions; EMS code: 2107 | Chemical | Extractable | LAB
2125 | ug/g | Sulfide Volatile | X257 | Colorimetric | Colorimetric. General use for several parameters. E.g.Total Sulfide is determined by spectrophotometer using the methlyene blue colourimetric method. The H2S calculation converts Total Sulphide as (S2-) and reports it as Total Sulphide as (H2S). | 5 | Y | Y | Y | 70 |  |  | 10/24/1996 2:25:27 PM | 1996-04-12 | 6/3/1998 6:51:33 PM | EMS |  | MassPerMass | mg/kg | Y | (mass conc.) |  | Sulfide Volatile (mass conc.) |  | Volatile Organic Compounds (VOCs); | Volatile Organic Compounds (VOCs); EMS code: 2125 | Chemical |  | LAB
2128 | umol/g | Acid VolatileSulfide | FS66 | 6NHCl,NaOH,Col'm | Hydrochloric acid, 6.0 N HCI, is added to the air-dried sediment samples. The hydrogen sulfide (H2S) produced is carried into two sodium hydroxide (NaOH) sulfide traps by purified nitrogen gas. The acid volatile sulfide absorbed in the NaOH is then determined colourimetrically. The detection limit is 0.2 umol/g. | 0.2 | Y | Y | Y | 19 | 3/13/2020 | 3/24/2020 | 2/22/2023 12:51:03 PM | 1999-07-09 |  | NPEPPIN |  | ParticlesPerMass | µmol/g | Y | (part. /m) |  | Acid VolatileSulfide (part. /m) |  |  | EMS code: 2128 | Chemical |  | LAB
2461 | None | Fish Bioassay(Pass/Fail) | EQ00 | EQUIS (None) | EQUIS (None) | 0 | N | Y | Y | 2 |  |  | 10/22/1996 10:05:30 PM | 1996-04-12 | 6/3/1998 6:51:33 PM | EMS |  | None | None | N | (unitless) |  | Fish Bioassay(Pass/Fail) (unitless) |  | Toxicity; | Toxicity; EMS code: 2461 | Chemical |  | LAB
2461 | % (Mortality) | Fish Bioassay(Pass/Fail) | FA13 | RM13/Rainbow/96hr(%mortality)/100% conc. | Rainbow/96hr(%mortality)/100% concentration. (RM13) Biological Test Method: Reference Method for Determining Acute Lethality of Effluents to Rainbow Trout; EPS 1/RM/13 Second Edition December 2000 (with May 2007 and February 2016 amendments). Rainbow Trout are exposed in a single concentration of the sample and mortality is recorded over a period of 96 hours at 15C. The pass/fail single concentration test uses undiluted material at 100% concentration and determines the percent mortality after 96 hours. (Without pH Adjustment) |  | Y | Y | Y | 4140 | 11/1/2022 | 11/1/2022 | 7/29/2024 4:35:58 PM | 2015-06-24 |  | RDGIBSON |  | DimensionlessRatio | % | N | (ratio) | (mortality) | Fish Bioassay(Pass/Fail) (mortality) (ratio) |  | Toxicity; | Toxicity; EMS code: 2461 | Chemical |  | LAB
2461 | % | Fish Bioassay(Pass/Fail) | X104 | Rainbow/96hr(%survival) | Rainbow/96hr(%survival) | 0 | Y | Y | Y | 569 |  |  | 3/24/2020 4:08:02 PM | 1996-04-12 | 6/3/1998 6:51:45 PM | EMS |  | DimensionlessRatio | % | N | (ratio) |  | Fish Bioassay(Pass/Fail) (ratio) |  | Toxicity; | Toxicity; EMS code: 2461 | Chemical |  | LAB
2461 | % | Fish Bioassay(Pass/Fail) | EQ34 | EQUIS (%) | EQUIS (%) | 0 | N | Y | Y | 180 |  |  | 10/23/1996 12:03:06 AM | 1996-04-12 | 6/3/1998 6:51:33 PM | EMS |  | DimensionlessRatio | % | N | (ratio) |  | Fish Bioassay(Pass/Fail) (ratio) |  | Toxicity; | Toxicity; EMS code: 2461 | Chemical |  | LAB
2462 | % (Mortality) | Rainbow Trout Bioassay (Pass/Fail) | RM59 | RM59/Rainbow/96hr/(% mortality)/pH stabilization-CO2 inject | EC Test Method Biological Test Method: Procedure for pH Stabilization During the Testing of Acute Lethality of Wastewater Effluent to Rainbow Trout; EPS 1/RM/50 - March 2008. A revised version of this method has been developed specifically for pulp and paper effluents; EPS 1/RM/59; March 2018. This method is specific for the pH controller technique described in EPS 1/RM/50 as an add on to the acute fish toxicity test; EPS 1/RM/13. Method Summary Rainbow Trout are exposed in single or multiple concentration(s) of the sample and mortality is recorded over a period of 96 hours at 15?1?C. The pass/fail single concentration test uses undiluted material at 100% concentration and determines the percent mortality after 96 hours. The LC50 and LC20 tests use multiple concentrations of material diluted in laboratory water and determines the concentration of that is lethal to 50% and 20% of the test fish respectively. Parallel testing with and without pH stabilization is recommended to demonstrate the presence of ammonia toxicity. Upward pH drift in the effluent test solution causes an increase in the concentration of un-ionized ammonia which is the most acutely toxic form present. The loss of CO2 is caused by the standard test procedure of aerating the effluent sample with lab air. The pH stabilization is intended to replace the CO2 lost during aeration in order to maintain the pH throughout the test at the same level found in the original sample. The pH stabilization is accomplished by using either the CO2 injection or the pH controller techniques. The procedure is only applicable when used explicitly with conjunction to the reference method EPS 1/RM/13 ?Biological Test Method: Reference Method for Determining Acute Lethality of Effluents to Rainbow Trout?. The pH stabilization applies to both single and multiple concentration tests. Applications For use only with wastewater effluents. The LC50 test is required for permitted discharges. |  | Y | Y | Y | 77 |  |  | 7/25/2024 3:53:30 PM | 2020-05-14 | 5/14/2020 4:16:58 PM | RDGIBSON | RDGIBSON | DimensionlessRatio | % | N | (ratio) | (mortality) | Rainbow Trout Bioassay (Pass/Fail) (mortality) (ratio) |  | Toxicity; | Toxicity; EMS code: 2462 | Chemical |  | LAB
24DB | ug/g | Dichlorophenoxy(2;4-) Butyric Acid; 4- [2;4-DB] | ZN03 | DCM+Ac.Extr:GC/MS | DCM+Ac.Extr:GC/MS EPA 8270E (mod) | 0.0005 | Y | Y | Y | 15 | 1/14/2022 | 2/2/2022 | 1/11/2024 1:07:06 PM | 2022-07-27 |  | EMS |  | MassPerMass | mg/kg | Y | (mass conc.) |  | Dichlorophenoxy(2;4-) Butyric Acid; 4- [2;4-DB] (mass conc.) |  | Pesticides; | Pesticides; EMS code: 24DB | Chemical |  | LAB
24DB | mg/L | Dichlorophenoxy(2;4-) Butyric Acid; 4- [2;4-DB] | P015 | Phenoxyacid Herbicides by LC-MS/MS | Phenoxyacid Herbicides by LC-MS/MS | 5e-05 | Y | Y | Y | 68 | 12/6/2021 | 9/9/2022 | 12/8/2022 12:57:05 PM | 2021-12-09 | 12/9/2021 3:22:12 PM | EMS | EMS | Concentration | mg/L | Y | (fl. conc.) |  | Dichlorophenoxy(2;4-) Butyric Acid; 4- [2;4-DB] (fl. conc.) |  | Pesticides; | Pesticides; EMS code: 24DB | Chemical |  | LAB
2BUOH | mg/L | 2-Butanol CAS: 78-92-2 | E-8015 | EPA 8015D | EPA Method 8015D (SW-846): Nonhalogenated Organics Using GC/FID |  | Y | Y | Y | 14 | 12/29/2021 | 1/12/2022 | 9/2/2022 6:21:52 PM | 2022-04-04 |  | EMS |  | Concentration | mg/L | Y | (fl. conc.) |  | 2-Butanol (fl. conc.) | 78-92-2 | Organics; | Organics; EMS code: 2BUOH | Chemical |  | LAB
2BUOH | ug/g | 2-Butanol CAS: 78-92-2 | E8015S | EPA 8015 D in soil | See EPA 8015D | 10 | Y | Y | Y | 15 | 1/12/2022 | 2/1/2022 | 1/11/2024 1:07:06 PM | 2022-06-23 |  | EMS |  | MassPerMass | mg/kg | Y | (mass conc.) |  | 2-Butanol (mass conc.) | 78-92-2 | Organics; | Organics; EMS code: 2BUOH | Chemical |  | LAB
2HOP | mg/L | 2-Hydroxyphenol (catechol) CAS: 120-80-9 | X268 | CH2Cl2 ext:GC/MS | CH2Cl2 ext:GC/MS |  | Y | Y | Y | 33 | 12/29/2021 | 6/7/2024 | 6/26/2024 4:28:10 PM | 2022-03-17 |  | EMS |  | Concentration | mg/L | Y | (fl. conc.) |  | 2-Hydroxyphenol (catechol) (fl. conc.) | 120-80-9 | Organics; | Organics; EMS code: 2HOP | Chemical |  | LAB
3016 | mg/L | Chlorine Dem.Total | X006 | Amperometric | Amperometric | 0 | Y | Y | Y | 5079 |  |  | 8/23/2024 8:10:36 AM | 1996-04-12 | 3/18/2003 9:38:19 AM | EMS | UMULLER | Concentration | mg/L | Y | (fl. conc.) |  | Chlorine Dem.Total (fl. conc.) |  | Inorganics; | Inorganics; EMS code: 3016 | Chemical | Total | LAB
3HOP | mg/L | Resorcinol CAS: 108-46-3 | X268 | CH2Cl2 ext:GC/MS | CH2Cl2 ext:GC/MS |  | Y | Y | Y | 37 | 12/29/2021 | 6/7/2024 | 6/26/2024 4:28:10 PM | 2022-03-21 |  | EMS |  | Concentration | mg/L | Y | (fl. conc.) |  | Resorcinol (fl. conc.) | 108-46-3 | Organics; | Organics; EMS code: 3HOP | Chemical |  | LAB
4HOP | mg/L | Hydroquinone CAS: 123-31-9 | X268 | CH2Cl2 ext:GC/MS | CH2Cl2 ext:GC/MS |  | Y | Y | Y | 37 | 12/29/2021 | 6/7/2024 | 6/26/2024 4:28:10 PM | 2022-03-21 |  | EMS |  | Concentration | mg/L | Y | (fl. conc.) |  | Hydroquinone (fl. conc.) | 123-31-9 | Organics; | Organics; EMS code: 4HOP | Chemical |  | LAB
A001 | mg/L | AMPA CAS: 77521-29-0 | X365 | Derivat,extr,HPLC/UV-VIS | An aliquot of the water sample is reacted with dinitrofluoro-benzene (DNFB) to form derivatives of glyphosate and AMPA. Interferences are removed from the reaction mixture by extraction with dichloromethane. A subsequent extraction with iso-butanol removes the derivatives. The derivatives are then back-extracted into a borax solution in preparation for analysis by high performance liquid chromatography. | 0.005 | Y | Y | Y | 19 | 1/3/2001 8:19:00 AM | 7/29/2005 7:01:00 AM | 6/2/2009 12:00:55 PM | 1996-04-12 | 6/3/1998 6:51:45 PM | EMS |  | Concentration | mg/L | Y | (fl. conc.) |  | AMPA (fl. conc.) | 77521-29-0 | Pesticides; | Pesticides; EMS code: A001 | Chemical |  | LAB
A001 | ug/g | AMPA CAS: 77521-29-0 | P013 | Gly and AMPA | Gly and AMPA | 0.05 | Y | Y | Y | 24 | 10/31/1997 9:51:00 AM | 11/5/1997 1:24:00 PM | 11/7/1997 6:08:28 PM | 1996-04-12 | 6/3/1998 6:51:33 PM | EMS |  | MassPerMass | mg/kg | Y | (mass conc.) |  | AMPA (mass conc.) | 77521-29-0 | Pesticides; | Pesticides; EMS code: A001 | Chemical |  | LAB
A001 | mg/L | AMPA CAS: 77521-29-0 | P012 | Glyphosate & AMPA in H2O | Glyphosate & AMPA in H2O | 0.005 | Y | Y | Y | 8 | 9/24/2002 10:55:00 AM | 9/24/2002 10:55:00 AM | 3/24/2004 1:04:11 AM | 1996-04-12 | 6/3/1998 6:51:33 PM | EMS |  | Concentration | mg/L | Y | (fl. conc.) |  | AMPA (fl. conc.) | 77521-29-0 | Pesticides; | Pesticides; EMS code: A001 | Chemical |  | LAB
A002 | mg/L | Aldrin | E660 | EPA 8270E (mod) | EPA 8270E (mod) Miscellaneous Pesticides by GC-MS | 2e-06 | Y | Y | Y | 2 | 9/14/2022 | 9/14/2022 | 12/8/2022 9:48:13 AM | 2022-12-06 |  | EMS |  | Concentration | mg/L | Y | (fl. conc.) |  | Aldrin (fl. conc.) | 309-00-2 | Pesticides; | Pesticides; EMS code: A002 | Chemical |  | LAB
A002 | mg/L | Aldrin | F027 | Hexane Extr,GC/ECD | Effluent samples are solvent extracted with Hexane and the extract dried over Sodium Sulfate, concentrated and cleaned up, as required, using Florisil, activated Copper metal, concentrated Sulfuric Acid and Mercury, to remove interferences. The extracts are then analysed using megabore column Gas Chromatography with Electron Capture detection. | 5e-05 | Y | Y | Y | 20 | 3/3/2000 7:36:00 AM | 1/22/2002 10:58:00 AM | 6/18/2003 2:53:25 PM | 2002-06-05 | 6/5/2002 3:51:38 PM | UMULLER | UMULLER | Concentration | mg/L | Y | (fl. conc.) |  | Aldrin (fl. conc.) | 309-00-2 | Pesticides; | Pesticides; EMS code: A002 | Chemical |  | LAB
A002 | ug/g | Aldrin | E8081 | EPA 8081 | See methods EPA 3540, 3630, 3660, 8081 | 0.001 | Y | Y | Y | 9 | 3/18/2020 | 3/23/2020 | 2/22/2023 12:51:03 PM | 2022-07-29 |  | EMS |  | MassPerMass | mg/kg | Y | (mass conc.) |  | Aldrin (mass conc.) | 309-00-2 | Pesticides; | Pesticides; EMS code: A002 | Chemical |  | LAB
A002 | ug/g | Aldrin | ZN03 | DCM+Ac.Extr:GC/MS | DCM+Ac.Extr:GC/MS EPA 8270E (mod) | 0.03 | Y | Y | Y | 19 | 1/17/2022 | 11/14/2022 | 1/11/2024 1:07:06 PM | 1996-04-12 | 11/16/1999 2:27:40 PM | EMS | NPEPPIN | MassPerMass | mg/kg | Y | (mass conc.) |  | Aldrin (mass conc.) | 309-00-2 | Pesticides; | Pesticides; EMS code: A002 | Chemical |  | LAB
A002 | mg/L | Aldrin | X268 | CH2Cl2 ext:GC/MS | CH2Cl2 ext:GC/MS |  | Y | Y | Y | 16 | 12/29/2021 | 9/2/2022 | 11/10/2022 9:54:06 AM | 2022-04-08 |  | EMS |  | Concentration | mg/L | Y | (fl. conc.) |  | Aldrin (fl. conc.) | 309-00-2 | Pesticides; | Pesticides; EMS code: A002 | Chemical |  | LAB
A002 | mg/L | Aldrin | ZN01 | DCM+Ac.Extr:GC/MS | DCM+Ac.Extr:GC/MS | 0.0003 | Y | Y | Y | 52 | 10/3/2002 | 3/29/2012 10:15:00 PM | 5/3/2012 9:27:49 AM | 1996-04-12 | 11/16/1999 2:56:50 PM | EMS | NPEPPIN | Concentration | mg/L | Y | (fl. conc.) |  | Aldrin (fl. conc.) | 309-00-2 | Pesticides; | Pesticides; EMS code: A002 | Chemical |  | LAB
A002 | ug/g wet | Aldrin | X220 | Hex/AcExt,Part,Chrom,GC | Hex/AcExt;Part;Chrom;GC | 0.0002 | Y | Y | Y | 72 |  |  | 10/24/1996 10:01:24 AM | 1996-04-12 | 6/3/1998 6:51:41 PM | EMS |  | MassPerMass | µg/g | Y | (mass conc.) | (wet) | Aldrin (wet) (mass conc.) | 309-00-2 | Pesticides; | Pesticides; EMS code: A002 | Chemical |  | LAB
A002 | ug/g | Aldrin | X213 | Hex/AcExt,Part,Chrom,GC | Hex/AcExt;Part;Chrom;GC | 0.001 | Y | Y | Y | 187 | 4/20/2002 10:10:00 AM | 10/13/2002 10:30:00 AM | 2/26/2004 1:05:26 AM | 1996-04-12 | 6/3/1998 6:51:42 PM | EMS |  | MassPerMass | mg/kg | Y | (mass conc.) |  | Aldrin (mass conc.) | 309-00-2 | Pesticides; | Pesticides; EMS code: A002 | Chemical |  | LAB
A002 | mg/L | Aldrin | EQ01 | EQUIS (mg/L) | EQUIS (mg/L) | 0 | N | Y | Y | 17 |  |  | 10/22/1996 11:26:53 PM | 1996-04-12 | 6/3/1998 6:51:33 PM | EMS |  | Concentration | mg/L | Y | (fl. conc.) |  | Aldrin (fl. conc.) | 309-00-2 | Pesticides; | Pesticides; EMS code: A002 | Chemical |  | LAB
A002 | mg/L | Aldrin | MW03 | OCP1-LL-SF-ECD-VA | This analysis is carried out using procedures adapted from "Test Methods for Evaluating Solid Waste" SW-846, Methods 3510, 3630, 3660 and 8081, published by the United States Environmental Protection Agency (EPA). The procedure involves extraction of the entire water sample with dichloromethane. The extract then undergoes one or more of the following clean-up procedures (if required): silica gel clean-up and/or sulphur clean-up. The final extract is analyzed by dual capillary column gas chromatography with electron capture detection (GC/ECD). | 5e-05 | Y | Y | Y | 10 | 9/20/2016 | 9/21/2016 | 8/15/2019 2:42:34 PM | 2019-08-15 |  | KATFRASE |  | Concentration | mg/L | Y | (fl. conc.) |  | Aldrin (fl. conc.) | 309-00-2 | Pesticides; | Pesticides; EMS code: A002 | Chemical |  | LAB
A003 | mg/L | Aldicarb | X304 | CH2Cl2 Extr:LC/MS | CH2Cl2 Extr:LC/MS | 1 | Y | Y | Y | 3 | 8/11/2004 8:04:00 AM | 7/29/2005 10:46:00 AM | 6/2/2009 12:00:56 PM | 1996-04-12 | 6/3/1998 6:51:33 PM | EMS |  | Concentration | mg/L | Y | (fl. conc.) |  | Aldicarb (fl. conc.) | 116-06-3 | Pesticides; | Pesticides; EMS code: A003 | Chemical |  | LAB
A004 | mg/L | Alachlor | E660 | EPA 8270E (mod) | EPA 8270E (mod) Miscellaneous Pesticides by GC-MS | 0.0001 | Y | Y | Y | 68 | 12/7/2021 | 9/12/2022 | 12/8/2022 12:57:05 PM | 2021-12-09 | 12/9/2021 3:22:12 PM | EMS | EMS | Concentration | mg/L | Y | (fl. conc.) |  | Alachlor (fl. conc.) | 15972-60-8 | Pesticides; | Pesticides; EMS code: A004 | Chemical |  | LAB
A004 | mg/L | Alachlor | ZN01 | DCM+Ac.Extr:GC/MS | DCM+Ac.Extr:GC/MS | 0.0005 | Y | Y | Y | 28 | 7/28/2005 7:03:00 AM | 3/29/2012 10:15:00 PM | 5/3/2012 9:27:49 AM | 1996-04-12 | 11/16/1999 2:11:55 PM | EMS | NPEPPIN | Concentration | mg/L | Y | (fl. conc.) |  | Alachlor (fl. conc.) | 15972-60-8 | Pesticides; | Pesticides; EMS code: A004 | Chemical |  | LAB
A004 | ug/g | Alachlor | ZN03 | DCM+Ac.Extr:GC/MS | DCM+Ac.Extr:GC/MS EPA 8270E (mod) | 0.05 | Y | Y | Y | 16 | 1/25/2022 | 11/2/2022 4:41:00 AM | 1/7/2023 3:12:10 PM | 1996-04-12 | 11/16/1999 2:11:55 PM | EMS | NPEPPIN | MassPerMass | mg/kg | Y | (mass conc.) |  | Alachlor (mass conc.) | 15972-60-8 | Pesticides; | Pesticides; EMS code: A004 | Chemical |  | LAB
A007 | ug/g | Acephate | ZN03 | DCM+Ac.Extr:GC/MS | DCM+Ac.Extr:GC/MS EPA 8270E (mod) | 0.1 | Y | Y | Y | 19 | 4/20/2002 10:25:00 AM | 10/16/2002 3:20:00 PM | 2/25/2004 1:09:43 AM | 1996-04-12 | 11/16/1999 2:11:31 PM | EMS | NPEPPIN | MassPerMass | mg/kg | Y | (mass conc.) |  | Acephate (mass conc.) | 30560-19-1 | Pesticides; | Pesticides; EMS code: A007 | Chemical |  | LAB
A007 | mg/L | Acephate | ZN01 | DCM+Ac.Extr:GC/MS | DCM+Ac.Extr:GC/MS | 0.001 | Y | Y | Y | 4 | 10/3/2002 | 7/29/2005 7:10:00 AM | 6/2/2009 12:00:56 PM | 1996-04-12 | 11/16/1999 2:56:50 PM | EMS | NPEPPIN | Concentration | mg/L | Y | (fl. conc.) |  | Acephate (fl. conc.) | 30560-19-1 | Pesticides; | Pesticides; EMS code: A007 | Chemical |  | LAB
A008 | mg/L | Azinphos Methyl | E660 | EPA 8270E (mod) | EPA 8270E (mod) Miscellaneous Pesticides by GC-MS | 0.0001 | Y | Y | Y | 68 | 12/7/2021 | 9/12/2022 | 12/8/2022 12:57:05 PM | 2021-12-09 | 12/9/2021 3:22:12 PM | EMS | EMS | Concentration | mg/L | Y | (fl. conc.) |  | Azinphos Methyl (fl. conc.) | 86-50-0 | Pesticides; | Pesticides; EMS code: A008 | Chemical |  | LAB
A008 | ug/g | Azinphos Methyl | X260 | Sox-Ac/CH2Cl2 ext:GC/NPD | Sox-Ac/CH2Cl2 ext:GC/NPD | 0.05 | Y | Y | Y | 117 | 4/20/2002 10:25:00 AM | 10/16/2002 3:20:00 PM | 2/26/2004 1:05:26 AM | 1996-04-12 | 6/3/1998 6:51:33 PM | EMS |  | MassPerMass | mg/kg | Y | (mass conc.) |  | Azinphos Methyl (mass conc.) | 86-50-0 | Pesticides; | Pesticides; EMS code: A008 | Chemical |  | LAB
A008 | mg/L | Azinphos Methyl | ZN01 | DCM+Ac.Extr:GC/MS | DCM+Ac.Extr:GC/MS | 0.001 | Y | Y | Y | 1 | 7/29/2005 7:10:00 AM | 7/29/2005 7:10:00 AM | 6/2/2009 12:00:56 PM | 1996-04-12 | 11/16/1999 4:06:27 PM | EMS | NPEPPIN | Concentration | mg/L | Y | (fl. conc.) |  | Azinphos Methyl (fl. conc.) | 86-50-0 | Pesticides; | Pesticides; EMS code: A008 | Chemical |  | LAB
A008 | ug/g | Azinphos Methyl | ZN03 | DCM+Ac.Extr:GC/MS | DCM+Ac.Extr:GC/MS EPA 8270E (mod) | 0.1 | Y | Y | Y | 16 | 1/25/2022 | 11/2/2022 4:41:00 AM | 1/7/2023 3:12:10 PM | 1999-11-16 |  | NPEPPIN |  | MassPerMass | mg/kg | Y | (mass conc.) |  | Azinphos Methyl (mass conc.) | 86-50-0 | Pesticides; | Pesticides; EMS code: A008 | Chemical |  | LAB
A008 | mg/L | Azinphos Methyl | F053 | DCM extr'n,conc'n,HGRC-PFPD/N,P det'n | Effluent samples are extracted with Dichloromethane. The extract is then concentrated and analysed using High Resolution Gas Chromatography with Pulsed Flame Photometric Dectection. Confirmation, if required, is by High Resolution Gas Chromatography with Nitrogen Phosphorus Detection. | 0.0001 | Y | Y | Y | 31 | 3/3/2000 7:36:00 AM | 1/22/2002 10:26:00 AM | 6/18/2003 2:53:25 PM | 1997-10-20 | 6/3/1998 6:51:47 PM | NPEPPIN |  | Concentration | mg/L | Y | (fl. conc.) |  | Azinphos Methyl (fl. conc.) | 86-50-0 | Pesticides; | Pesticides; EMS code: A008 | Chemical |  | LAB
A009 | mg/L | Azinphos Ethyl | ZN01 | DCM+Ac.Extr:GC/MS | DCM+Ac.Extr:GC/MS | 0.0005 | Y | Y | Y | 1 | 7/29/2005 7:10:00 AM | 7/29/2005 7:10:00 AM | 6/2/2009 12:00:56 PM | 1996-04-12 | 11/16/1999 4:06:27 PM | EMS | NPEPPIN | Concentration | mg/L | Y | (fl. conc.) |  | Azinphos Ethyl (fl. conc.) | 2642-71-9 | Pesticides; | Pesticides; EMS code: A009 | Chemical |  | LAB
A010 | mg/L | Aspon | ZN01 | DCM+Ac.Extr:GC/MS | DCM+Ac.Extr:GC/MS | 0.0002 | Y | Y | Y | 52 | 10/3/2002 | 3/29/2012 10:15:00 PM | 5/3/2012 9:27:49 AM | 1996-04-12 | 11/16/1999 2:56:50 PM | EMS | NPEPPIN | Concentration | mg/L | Y | (fl. conc.) |  | Aspon (fl. conc.) |  | Pesticides; | Pesticides; EMS code: A010 | Chemical |  | LAB
A013 | mg/L | Aldicarb Sulfone | P045 | DCM Extr:HPLC | DCM Extr:HPLC | 0.0005 | Y | Y | Y | 1 | 7/29/2005 7:10:00 AM | 7/29/2005 7:10:00 AM | 6/2/2009 12:00:56 PM | 1996-04-12 | 6/3/1998 6:51:41 PM | EMS |  | Concentration | mg/L | Y | (fl. conc.) |  | Aldicarb Sulfone (fl. conc.) | 1646-88-4 | Pesticides; | Pesticides; EMS code: A013 | Chemical |  | LAB
A014 | mg/L | Aldicarb Sulfoxide | P045 | DCM Extr:HPLC | DCM Extr:HPLC | 0.0005 | Y | Y | Y | 1 | 7/29/2005 7:10:00 AM | 7/29/2005 7:10:00 AM | 6/2/2009 12:00:56 PM | 1996-04-12 | 6/3/1998 6:51:42 PM | EMS |  | Concentration | mg/L | Y | (fl. conc.) |  | Aldicarb Sulfoxide (fl. conc.) |  | Pesticides; | Pesticides; EMS code: A014 | Chemical |  | LAB
A015 | mg/L | De-ethyl Atrazine | E751 | E751D Pesticides in Water by LC-MS/MS (POS1 List) | E751D Pesticides in Water by LC-MS/MS (POS1 List) (MOECC E3501) | 5e-05 | Y | Y | Y | 39 | 12/9/2021 | 12/31/2021 | 2/15/2022 9:50:04 AM | 2021-12-14 |  | EMS |  | Concentration | mg/L | Y | (fl. conc.) |  | De-ethyl Atrazine (fl. conc.) | 6190-65-4 | Pesticides; | Pesticides; EMS code: A015 | Chemical |  | LAB
A015 | mg/L | De-ethyl Atrazine | ZN01 | DCM+Ac.Extr:GC/MS | DCM+Ac.Extr:GC/MS | 0.0003 | Y | Y | Y | 28 | 7/28/2005 7:03:00 AM | 3/29/2012 10:15:00 PM | 5/3/2012 9:27:49 AM | 1996-04-12 | 11/16/1999 4:06:27 PM | EMS | NPEPPIN | Concentration | mg/L | Y | (fl. conc.) |  | De-ethyl Atrazine (fl. conc.) | 6190-65-4 | Pesticides; | Pesticides; EMS code: A015 | Chemical |  | LAB
A015 | ug/g | De-ethyl Atrazine | ZN03 | DCM+Ac.Extr:GC/MS | DCM+Ac.Extr:GC/MS EPA 8270E (mod) | 0.03 | Y | Y | Y | 2 | 9/8/2022 | 9/8/2022 | 12/15/2022 10:50:07 AM | 1996-04-12 | 11/16/1999 2:27:40 PM | EMS | NPEPPIN | MassPerMass | mg/kg | Y | (mass conc.) |  | De-ethyl Atrazine (mass conc.) | 6190-65-4 | Pesticides; | Pesticides; EMS code: A015 | Chemical |  | LAB
A015 | mg/L | De-ethyl Atrazine | E660 | EPA 8270E (mod) | EPA 8270E (mod) Miscellaneous Pesticides by GC-MS | 0.0001 | Y | Y | Y | 68 | 12/7/2021 | 9/12/2022 | 12/8/2022 12:57:05 PM | 2021-12-09 | 12/9/2021 3:22:12 PM | EMS | EMS | Concentration | mg/L | Y | (fl. conc.) |  | De-ethyl Atrazine (fl. conc.) | 6190-65-4 | Pesticides; | Pesticides; EMS code: A015 | Chemical |  | LAB
A016 | mg/L | Atrazine | ZN01 | DCM+Ac.Extr:GC/MS | DCM+Ac.Extr:GC/MS | 0.0002 | Y | Y | Y | 52 | 10/3/2002 | 3/29/2012 10:15:00 PM | 5/3/2012 9:27:49 AM | 1996-04-12 | 11/16/1999 2:56:50 PM | EMS | NPEPPIN | Concentration | mg/L | Y | (fl. conc.) |  | Atrazine (fl. conc.) | 1912-24-9 | Pesticides; | Pesticides; EMS code: A016 | Chemical |  | LAB
A016 | mg/L | Atrazine | E751 | E751D Pesticides in Water by LC-MS/MS (POS1 List) | E751D Pesticides in Water by LC-MS/MS (POS1 List) (MOECC E3501) | 5e-05 | Y | Y | Y | 39 | 12/9/2021 | 12/31/2021 | 2/15/2022 9:50:04 AM | 2021-12-14 |  | EMS |  | Concentration | mg/L | Y | (fl. conc.) |  | Atrazine (fl. conc.) | 1912-24-9 | Pesticides; | Pesticides; EMS code: A016 | Chemical |  | LAB
A016 | ug/g | Atrazine | ZN03 | DCM+Ac.Extr:GC/MS | DCM+Ac.Extr:GC/MS EPA 8270E (mod) | 0.02 | Y | Y | Y | 14 | 1/25/2022 | 9/8/2022 | 12/15/2022 10:50:07 AM | 1996-04-12 | 11/16/1999 2:27:40 PM | EMS | NPEPPIN | MassPerMass | mg/kg | Y | (mass conc.) |  | Atrazine (mass conc.) | 1912-24-9 | Pesticides; | Pesticides; EMS code: A016 | Chemical |  | LAB
A016 | mg/L | Atrazine | F056 | DCM extr,HRGC-NPD | Effluent samples are extracted with Dichloromethane. The extract is concentrated and analysed using High Resolution Gas Chromatography with a Thermionic Specific Detector (TSD), alternately referred to as a Nitrogen Phosphorus Detector (NPD). | 0.0005 | Y | Y | Y | 12 | 4/3/2000 10:51:00 AM | 2/16/2001 3:36:00 PM | 9/11/2002 9:34:55 AM | 1998-10-02 | 10/2/1998 10:50:33 AM | NPEPPIN | NPEPPIN | Concentration | mg/L | Y | (fl. conc.) |  | Atrazine (fl. conc.) | 1912-24-9 | Pesticides; | Pesticides; EMS code: A016 | Chemical |  | LAB
A016 | mg/L | Atrazine | E660 | EPA 8270E (mod) | EPA 8270E (mod) Miscellaneous Pesticides by GC-MS | 0.0001 | Y | Y | Y | 68 | 12/7/2021 | 9/12/2022 | 12/8/2022 12:57:05 PM | 2021-12-09 | 12/9/2021 3:22:12 PM | EMS | EMS | Concentration | mg/L | Y | (fl. conc.) |  | Atrazine (fl. conc.) | 1912-24-9 | Pesticides; | Pesticides; EMS code: A016 | Chemical |  | LAB
A016 | ug/g | Atrazine | FS61 | DCM +MeOH extr,HRGC-NPD | Sediment samples are shaken with a mixture of Dichloromethane and Methanol, filtered then re-extracted with Dichloromethane. The extract is concentrated and analysed using High Resolution Gas Chromatography with a Thermionic Specific Detector (TSD), alternately referred to as a Nitrogen Phosphorus Detector (NPD). | 0.01 | Y | Y | Y | 34 | 4/20/2002 3:40:00 PM | 7/16/2002 3:20:00 PM | 2/26/2004 1:05:26 AM | 1998-10-02 |  | NPEPPIN |  | MassPerMass | mg/kg | Y | (mass conc.) |  | Atrazine (mass conc.) | 1912-24-9 | Pesticides; | Pesticides; EMS code: A016 | Chemical |  | LAB
A017 | mg/L | Arochlor 1242 (PCB) | E660 | EPA 8270E (mod) | EPA 8270E (mod) Miscellaneous Pesticides by GC-MS | 2e-05 | Y | Y | Y | 8 | 5/24/2023 | 11/1/2023 | 11/16/2023 3:38:09 PM | 2023-08-16 |  | EMS |  | Concentration | mg/L | Y | (fl. conc.) |  | Arochlor 1242 (PCB) (fl. conc.) | 53469-21-9 | Polychlorinated biphenylds (PCBs); | Polychlorinated biphenylds (PCBs); EMS code: A017 | Chemical |  | LAB
A017 | ug/g wet | Arochlor 1242 (PCB) | FB03 | Frozen,Acetone Extr,GC/ECD | Sediment and homogenized biota samples are Acetone extracted using a wrist-action shaker, partitioned with 2% NaCl, then treated using Florisil, concentrated Sulfuric Acid, activated Copper and Mercury (as required) to remove interferences. The samples are then analysed using Megabore Column Gas Chromatography with Electron Capture Detection. Note that biota results are reported on a "wet weight" basis. | 0.01 | Y | Y | Y | 2 | 4/12/2001 11:51:00 AM | 4/12/2001 11:51:00 AM | 9/11/2002 3:23:55 PM | 1997-05-15 | 6/3/1998 6:51:47 PM | MPAK |  | MassPerMass | µg/g | Y | (mass conc.) | (wet) | Arochlor 1242 (PCB) (wet) (mass conc.) | 53469-21-9 | Polychlorinated biphenylds (PCBs); | Polychlorinated biphenylds (PCBs); EMS code: A017 | Chemical |  | LAB
A017 | ug/g | Arochlor 1242 (PCB) | FS06 | Frozen,Acetone extr,GC/ECD | Sediment and homogenized biota samples are Acetone extracted using a wrist-action shaker, partitioned with 2% NaCl, then treated using Florisil, concentrated Sulfuric Acid, activated Copper and Mercury (as required) to remove interferences. The samples are then analysed using Megabore Column Gas Chromatography with Electron Capture Detection. | 0.01 | Y | Y | Y | 62 | 6/29/1999 11:10:00 AM | 11/9/2001 1:34:00 PM | 8/21/2002 2:09:53 PM | 1997-05-15 | 6/3/1998 6:51:47 PM | MPAK |  | MassPerMass | mg/kg | Y | (mass conc.) |  | Arochlor 1242 (PCB) (mass conc.) | 53469-21-9 | Polychlorinated biphenylds (PCBs); | Polychlorinated biphenylds (PCBs); EMS code: A017 | Chemical |  | LAB
A017 | mg/L | Arochlor 1242 (PCB) | X303 | CH2Cl2 Extr:GC/ECD | Polychlorinated Biphenyls (PCBs) In Water by Gas Chromatography-Electron Capture Detector (GC/ECD). A water sample, usually 1000 mL, is extracted using dichloromethane. Sulphuric acid clean-up step is performed in every sample. Dirty samples with dark colored extracts are put through a clean-up step using 2% Deactivated Florisil. If sulphur is present in the water samples, especially for the river or stream water samples, a copper metal clean-up step may be used. A final analysis is performed using a GC equipped with an ECD. SCOPE: This procedure describes the determination of PCBs in water. If Organochlorine Pesticides (OC's) and PCB's are to be analyzed on the same water sample, a separate test method (Method for the Determination of Pesticides in Water by GC/MS and LC/MS or HPLC Fluorescence Detection DCN 62-C-043) is used. DETECTION LIMITS: Target Compound Reported Detection Limit (mg/L) Arochlor 1221 0.1 Arochlor 1242 0.1 Arochlor 1248 0.1 Arochlor 1254 0.1 Arochlor 1260 0.1 INTERFERENCES: Interferences (e.g. co-eluting peaks on the chromatogram) may be due to sulfur,phthalate esters, or from chlorinated compounds. Suspect samples must be reanalyzed using appropriate cleanup steps. SAMPLE HANDLING & PRESERVATION: Sample Containers: For sample collection, use 1 litre amber glass bottles that have been pre-rinsed with acetone and hexane then dried. Bottle lids must be teflon-lined, or alternatively lined with solvent-rinsed aluminium foil. Shipment and Storage: Samples must be refrigerated (at 2-8°C) and stored in the dark. Holding Time: Samples must be extracted within 7 days of sampling; REFERENCE: U.S. EPA Method 8080: Organochlorine Pesticides and PCB's Analysis. SW-846 3rd Ed.,September, 1986. | 0.001 | Y | Y | Y | 2 | 12/21/2012 10:31:00 AM | 4/2/2013 3:22:00 PM | 4/12/2013 3:34:31 PM | 2013-01-30 |  | RDGIBSON |  | Concentration | mg/L | Y | (fl. conc.) |  | Arochlor 1242 (PCB) (fl. conc.) | 53469-21-9 | Polychlorinated biphenylds (PCBs); | Polychlorinated biphenylds (PCBs); EMS code: A017 | Chemical |  | LAB
A017 | ug/g | Arochlor 1242 (PCB) | X383 | Extr, GC/MS or GC/ECD | Determination of Polychloriated Biphenyls (PCBs) In Soil by GC/ECD. PURPOSE/PRINCIPLE OF METHOD: This document details the requirements and procedures for the analysis of polychlorinated biphenyls (PCBs) in soil by gas chromatography electron capture detector (GC/ECD). A homogenized soil sample is extracted using dichloromethane (DCM). The extract may be cleaned up using column chromatography with 2% Deactivated Florisil. When contamination with sulphur is suspected, further cleanup using copper is required. Analysis is performed using a GC/ECD. Similar methods have been developed for other matrices including waste and transformer oils, water, swabs, florisil tubes, biota. Generally, the sample extraction procedures differ while the GC analysis methods are similar or identical for all matrices. SCOPE: This procedure describes the determination of PCBs in soil and sediments. Other materials such as wood waste may also be analyzed by this method. DETECTION LIMITS: Target Compound Reported Detection Limit (mg/g) Arochlor 1242 0.03, Arochlor 1248 0.03, Arochlor 1254 0.03, Arochlor 1260 0.03, INTERFERENCES: Interferences (e.g. co-eluting peaks on the chromatogram) may be due to sulphur, phthalate esters, or from chlorinated compounds. Suspect samples must be reanalyzed using appropriate cleanup steps. SAMPLE HANDLING & PRESERVATION: Sample Requirements: Whenever possible, samples must completely fill the bottle such that little or no headspace is present to eliminate the loss of volatile hydrocarbon components. (~125g of sample) Sample Containers: Glass jars that have been solvent rinsed with acetone then hexane, then dried. Shipment and Storage: Samples are to be refrigerated (at 4±2°C) during shipping and prior to analysis. Samples suspected to contain high levels of analyte are to be segregated from other samples. Holding Time: Holding time prior to extraction is 14 days. REFERENCE: U.S. EPA Method 8080:Organochlorine Pesticides and PCBs Analysis. | 0.03 | Y | Y | Y | 46 | 3/16/2020 | 9/7/2022 | 1/11/2024 1:07:06 PM | 2013-01-30 |  | RDGIBSON |  | MassPerMass | mg/kg | Y | (mass conc.) |  | Arochlor 1242 (PCB) (mass conc.) | 53469-21-9 | Polychlorinated biphenylds (PCBs); | Polychlorinated biphenylds (PCBs); EMS code: A017 | Chemical |  | LAB
A017 | ug/g | Arochlor 1242 (PCB) | F115 | Hex dil,Hex extr'n,GC/ECD | Effluent and swab samples are extracted with hexane, the extract concentrated, and then treated using Florisil, concentrated Sulfuric Acid, activated Copper and Mercury (as required) to remove interferences. The samples are then analysed using Megabore Column Gas Chromatography with Electron Capture Detection. Oil samples are first diluted with Hexane then treated in the manner outlined for effluent and swab samples. The samples are then analysed using Megabore Column Gas Chromatography with Electron Capture Detection. | 0.1 | Y | Y | Y | 3 | 4/19/2001 9:04:00 AM | 4/19/2001 9:06:00 AM | 6/1/2001 3:54:26 AM | 2001-05-24 |  | NPEPPIN |  | MassPerMass | mg/kg | Y | (mass conc.) |  | Arochlor 1242 (PCB) (mass conc.) | 53469-21-9 | Polychlorinated biphenylds (PCBs); | Polychlorinated biphenylds (PCBs); EMS code: A017 | Chemical |  | LAB
A017 | mg/L | Arochlor 1242 (PCB) | EQ01 | EQUIS (mg/L) | EQUIS (mg/L) | 0.0004 | N | Y | Y | 3 |  |  | 10/22/1996 3:25:28 PM | 1996-04-12 | 6/3/1998 6:51:33 PM | EMS |  | Concentration | mg/L | Y | (fl. conc.) |  | Arochlor 1242 (PCB) (fl. conc.) | 53469-21-9 | Polychlorinated biphenylds (PCBs); | Polychlorinated biphenylds (PCBs); EMS code: A017 | Chemical |  | LAB
A017 | mg/L | Arochlor 1242 (PCB) | E-8082 | EPA-8082A | See reference for EPA method 8082A POLYCHLORINATED BIPHENYLS (PCBs) BY GAS CHROMATOGRAPHY |  | Y | Y | Y | 14 | 12/28/2021 | 1/11/2022 | 9/2/2022 6:21:52 PM | 2022-04-06 |  | EMS |  | Concentration | mg/L | Y | (fl. conc.) |  | Arochlor 1242 (PCB) (fl. conc.) | 53469-21-9 | Polychlorinated biphenylds (PCBs); | Polychlorinated biphenylds (PCBs); EMS code: A017 | Chemical |  | LAB
A017 | ug/g wet | Arochlor 1242 (PCB) | EQ44 | EQUIS (ug/g wet) | EQUIS (ug/g wet) | 0.02 | N | Y | Y | 2 |  |  | 10/22/1996 9:03:11 PM | 1996-04-12 | 6/3/1998 6:51:33 PM | EMS |  | MassPerMass | µg/g | Y | (mass conc.) | (wet) | Arochlor 1242 (PCB) (wet) (mass conc.) | 53469-21-9 | Polychlorinated biphenylds (PCBs); | Polychlorinated biphenylds (PCBs); EMS code: A017 | Chemical |  | LAB
A017 | mg/L | Arochlor 1242 (PCB) | F018 | Hexane extr,GC/ECD | Effluent and swab samples are extracted with hexane, the extract concentrated, and then treated using Florisil, concentrated Sulfuric Acid, activated Copper and Mercury (as required) to remove interferences. The samples are then analysed using Megabore Column Gas Chromatography with Electron Capture Detection. Oil samples are first diluted with Hexane then treated in the manner outlined for effluent and swab samples. The samples are then analysed using Megabore Column Gas Chromatography with Electron Capture Detection. | 0.0001 | Y | Y | Y | 18 | 2/28/2000 8:40:00 AM | 9/5/2023 6:46:00 PM | 9/11/2023 5:22:03 PM | 1997-05-15 | 6/3/1998 6:51:47 PM | MPAK |  | Concentration | mg/L | Y | (fl. conc.) |  | Arochlor 1242 (PCB) (fl. conc.) | 53469-21-9 | Polychlorinated biphenylds (PCBs); | Polychlorinated biphenylds (PCBs); EMS code: A017 | Chemical |  | LAB
A018 | ug/g wet | Arochlor 1221 (PCB) Chlorodiphenyl (21% Cl) | FB03 | Frozen,Acetone Extr,GC/ECD | Sediment and homogenized biota samples are Acetone extracted using a wrist-action shaker, partitioned with 2% NaCl, then treated using Florisil, concentrated Sulfuric Acid, activated Copper and Mercury (as required) to remove interferences. The samples are then analysed using Megabore Column Gas Chromatography with Electron Capture Detection. Note that biota results are reported on a "wet weight" basis. | 0.01 | Y | Y | Y | 2 | 4/12/2001 11:51:00 AM | 4/12/2001 11:51:00 AM | 9/11/2002 3:23:55 PM | 1997-05-15 | 6/3/1998 6:51:47 PM | MPAK |  | MassPerMass | µg/g | Y | (mass conc.) | (wet) | Arochlor 1221 (PCB) Chlorodiphenyl (21% Cl) (wet) (mass conc.) | 11104-28-2 | Polychlorinated biphenylds (PCBs); | Polychlorinated biphenylds (PCBs); EMS code: A018 | Chemical |  | LAB
A018 | mg/L | Arochlor 1221 (PCB) Chlorodiphenyl (21% Cl) | F018 | Hexane extr,GC/ECD | Effluent and swab samples are extracted with hexane, the extract concentrated, and then treated using Florisil, concentrated Sulfuric Acid, activated Copper and Mercury (as required) to remove interferences. The samples are then analysed using Megabore Column Gas Chromatography with Electron Capture Detection. Oil samples are first diluted with Hexane then treated in the manner outlined for effluent and swab samples. The samples are then analysed using Megabore Column Gas Chromatography with Electron Capture Detection. | 0.0001 | Y | Y | Y | 18 | 2/28/2000 8:40:00 AM | 9/5/2023 6:46:00 PM | 9/11/2023 5:22:03 PM | 1997-05-15 | 6/3/1998 6:51:47 PM | MPAK |  | Concentration | mg/L | Y | (fl. conc.) |  | Arochlor 1221 (PCB) Chlorodiphenyl (21% Cl) (fl. conc.) | 11104-28-2 | Polychlorinated biphenylds (PCBs); | Polychlorinated biphenylds (PCBs); EMS code: A018 | Chemical |  | LAB
A018 | ug/g | Arochlor 1221 (PCB) Chlorodiphenyl (21% Cl) | X383 | Extr, GC/MS or GC/ECD | Determination of Polychloriated Biphenyls (PCBs) In Soil by GC/ECD. PURPOSE/PRINCIPLE OF METHOD: This document details the requirements and procedures for the analysis of polychlorinated biphenyls (PCBs) in soil by gas chromatography electron capture detector (GC/ECD). A homogenized soil sample is extracted using dichloromethane (DCM). The extract may be cleaned up using column chromatography with 2% Deactivated Florisil. When contamination with sulphur is suspected, further cleanup using copper is required. Analysis is performed using a GC/ECD. Similar methods have been developed for other matrices including waste and transformer oils, water, swabs, florisil tubes, biota. Generally, the sample extraction procedures differ while the GC analysis methods are similar or identical for all matrices. SCOPE: This procedure describes the determination of PCBs in soil and sediments. Other materials such as wood waste may also be analyzed by this method. DETECTION LIMITS: Target Compound Reported Detection Limit (mg/g) Arochlor 1242 0.03, Arochlor 1248 0.03, Arochlor 1254 0.03, Arochlor 1260 0.03, INTERFERENCES: Interferences (e.g. co-eluting peaks on the chromatogram) may be due to sulphur, phthalate esters, or from chlorinated compounds. Suspect samples must be reanalyzed using appropriate cleanup steps. SAMPLE HANDLING & PRESERVATION: Sample Requirements: Whenever possible, samples must completely fill the bottle such that little or no headspace is present to eliminate the loss of volatile hydrocarbon components. (~125g of sample) Sample Containers: Glass jars that have been solvent rinsed with acetone then hexane, then dried. Shipment and Storage: Samples are to be refrigerated (at 4±2°C) during shipping and prior to analysis. Samples suspected to contain high levels of analyte are to be segregated from other samples. Holding Time: Holding time prior to extraction is 14 days. REFERENCE: U.S. EPA Method 8080:Organochlorine Pesticides and PCBs Analysis. | 0.03 | Y | Y | Y | 46 | 3/16/2020 | 9/7/2022 | 1/11/2024 1:07:06 PM | 2013-01-30 |  | RDGIBSON |  | MassPerMass | mg/kg | Y | (mass conc.) |  | Arochlor 1221 (PCB) Chlorodiphenyl (21% Cl) (mass conc.) | 11104-28-2 | Polychlorinated biphenylds (PCBs); | Polychlorinated biphenylds (PCBs); EMS code: A018 | Chemical |  | LAB
A018 | ug/g | Arochlor 1221 (PCB) Chlorodiphenyl (21% Cl) | FS06 | Frozen,Acetone extr,GC/ECD | Sediment and homogenized biota samples are Acetone extracted using a wrist-action shaker, partitioned with 2% NaCl, then treated using Florisil, concentrated Sulfuric Acid, activated Copper and Mercury (as required) to remove interferences. The samples are then analysed using Megabore Column Gas Chromatography with Electron Capture Detection. | 0.01 | Y | Y | Y | 60 | 2/18/2000 5:06:00 PM | 11/9/2001 1:34:00 PM | 8/21/2002 2:09:53 PM | 1997-05-15 | 6/3/1998 6:51:47 PM | MPAK |  | MassPerMass | mg/kg | Y | (mass conc.) |  | Arochlor 1221 (PCB) Chlorodiphenyl (21% Cl) (mass conc.) | 11104-28-2 | Polychlorinated biphenylds (PCBs); | Polychlorinated biphenylds (PCBs); EMS code: A018 | Chemical |  | LAB
A018 | mg/L | Arochlor 1221 (PCB) Chlorodiphenyl (21% Cl) | E-8082 | EPA-8082A | See reference for EPA method 8082A POLYCHLORINATED BIPHENYLS (PCBs) BY GAS CHROMATOGRAPHY |  | Y | Y | Y | 14 | 12/28/2021 | 1/11/2022 | 9/2/2022 6:21:52 PM | 2022-04-06 |  | EMS |  | Concentration | mg/L | Y | (fl. conc.) |  | Arochlor 1221 (PCB) Chlorodiphenyl (21% Cl) (fl. conc.) | 11104-28-2 | Polychlorinated biphenylds (PCBs); | Polychlorinated biphenylds (PCBs); EMS code: A018 | Chemical |  | LAB
A018 | mg/L | Arochlor 1221 (PCB) Chlorodiphenyl (21% Cl) | E660 | EPA 8270E (mod) | EPA 8270E (mod) Miscellaneous Pesticides by GC-MS | 2e-05 | Y | Y | Y | 8 | 5/24/2023 | 11/1/2023 | 11/16/2023 3:38:09 PM | 2023-08-16 |  | EMS |  | Concentration | mg/L | Y | (fl. conc.) |  | Arochlor 1221 (PCB) Chlorodiphenyl (21% Cl) (fl. conc.) | 11104-28-2 | Polychlorinated biphenylds (PCBs); | Polychlorinated biphenylds (PCBs); EMS code: A018 | Chemical |  | LAB
A018 | ug/g | Arochlor 1221 (PCB) Chlorodiphenyl (21% Cl) | F115 | Hex dil,Hex extr'n,GC/ECD | Effluent and swab samples are extracted with hexane, the extract concentrated, and then treated using Florisil, concentrated Sulfuric Acid, activated Copper and Mercury (as required) to remove interferences. The samples are then analysed using Megabore Column Gas Chromatography with Electron Capture Detection. Oil samples are first diluted with Hexane then treated in the manner outlined for effluent and swab samples. The samples are then analysed using Megabore Column Gas Chromatography with Electron Capture Detection. | 0.1 | Y | Y | Y | 3 | 4/19/2001 9:04:00 AM | 4/19/2001 9:06:00 AM | 6/1/2001 3:54:27 AM | 2001-05-24 |  | NPEPPIN |  | MassPerMass | mg/kg | Y | (mass conc.) |  | Arochlor 1221 (PCB) Chlorodiphenyl (21% Cl) (mass conc.) | 11104-28-2 | Polychlorinated biphenylds (PCBs); | Polychlorinated biphenylds (PCBs); EMS code: A018 | Chemical |  | LAB
A019 | mg/L | Arochlor 1232 (PCB) | E660 | EPA 8270E (mod) | EPA 8270E (mod) Miscellaneous Pesticides by GC-MS | 2e-05 | Y | Y | Y | 8 | 5/24/2023 | 11/1/2023 | 11/16/2023 3:38:09 PM | 2023-08-16 |  | EMS |  | Concentration | mg/L | Y | (fl. conc.) |  | Arochlor 1232 (PCB) (fl. conc.) | 11141-16-5 | Polychlorinated biphenylds (PCBs); | Polychlorinated biphenylds (PCBs); EMS code: A019 | Chemical |  | LAB
A019 | ug/g | Arochlor 1232 (PCB) | FS06 | Frozen,Acetone extr,GC/ECD | Sediment and homogenized biota samples are Acetone extracted using a wrist-action shaker, partitioned with 2% NaCl, then treated using Florisil, concentrated Sulfuric Acid, activated Copper and Mercury (as required) to remove interferences. The samples are then analysed using Megabore Column Gas Chromatography with Electron Capture Detection. | 0.01 | Y | Y | Y | 60 | 2/18/2000 5:06:00 PM | 11/9/2001 1:34:00 PM | 8/21/2002 2:09:54 PM | 1997-05-15 | 6/3/1998 6:51:47 PM | MPAK |  | MassPerMass | mg/kg | Y | (mass conc.) |  | Arochlor 1232 (PCB) (mass conc.) | 11141-16-5 | Polychlorinated biphenylds (PCBs); | Polychlorinated biphenylds (PCBs); EMS code: A019 | Chemical |  | LAB
A019 | ug/g wet | Arochlor 1232 (PCB) | FB03 | Frozen,Acetone Extr,GC/ECD | Sediment and homogenized biota samples are Acetone extracted using a wrist-action shaker, partitioned with 2% NaCl, then treated using Florisil, concentrated Sulfuric Acid, activated Copper and Mercury (as required) to remove interferences. The samples are then analysed using Megabore Column Gas Chromatography with Electron Capture Detection. Note that biota results are reported on a "wet weight" basis. | 0.01 | Y | Y | Y | 2 | 4/12/2001 11:51:00 AM | 4/12/2001 11:51:00 AM | 9/11/2002 3:23:55 PM | 1997-05-15 | 6/3/1998 6:51:47 PM | MPAK |  | MassPerMass | µg/g | Y | (mass conc.) | (wet) | Arochlor 1232 (PCB) (wet) (mass conc.) | 11141-16-5 | Polychlorinated biphenylds (PCBs); | Polychlorinated biphenylds (PCBs); EMS code: A019 | Chemical |  | LAB
A019 | mg/L | Arochlor 1232 (PCB) | F018 | Hexane extr,GC/ECD | Effluent and swab samples are extracted with hexane, the extract concentrated, and then treated using Florisil, concentrated Sulfuric Acid, activated Copper and Mercury (as required) to remove interferences. The samples are then analysed using Megabore Column Gas Chromatography with Electron Capture Detection. Oil samples are first diluted with Hexane then treated in the manner outlined for effluent and swab samples. The samples are then analysed using Megabore Column Gas Chromatography with Electron Capture Detection. | 0.0001 | Y | Y | Y | 18 | 2/28/2000 8:40:00 AM | 9/5/2023 6:46:00 PM | 9/11/2023 5:22:04 PM | 1997-05-15 | 6/3/1998 6:51:47 PM | MPAK |  | Concentration | mg/L | Y | (fl. conc.) |  | Arochlor 1232 (PCB) (fl. conc.) | 11141-16-5 | Polychlorinated biphenylds (PCBs); | Polychlorinated biphenylds (PCBs); EMS code: A019 | Chemical |  | LAB
A019 | ug/g | Arochlor 1232 (PCB) | F115 | Hex dil,Hex extr'n,GC/ECD | Effluent and swab samples are extracted with hexane, the extract concentrated, and then treated using Florisil, concentrated Sulfuric Acid, activated Copper and Mercury (as required) to remove interferences. The samples are then analysed using Megabore Column Gas Chromatography with Electron Capture Detection. Oil samples are first diluted with Hexane then treated in the manner outlined for effluent and swab samples. The samples are then analysed using Megabore Column Gas Chromatography with Electron Capture Detection. | 0.1 | Y | Y | Y | 3 | 4/19/2001 9:04:00 AM | 4/19/2001 9:06:00 AM | 6/1/2001 3:54:27 AM | 2001-05-24 |  | NPEPPIN |  | MassPerMass | mg/kg | Y | (mass conc.) |  | Arochlor 1232 (PCB) (mass conc.) | 11141-16-5 | Polychlorinated biphenylds (PCBs); | Polychlorinated biphenylds (PCBs); EMS code: A019 | Chemical |  | LAB
A019 | mg/L | Arochlor 1232 (PCB) | E-8082 | EPA-8082A | See reference for EPA method 8082A POLYCHLORINATED BIPHENYLS (PCBs) BY GAS CHROMATOGRAPHY |  | Y | Y | Y | 14 | 12/28/2021 | 1/11/2022 | 9/2/2022 6:21:52 PM | 2022-04-06 |  | EMS |  | Concentration | mg/L | Y | (fl. conc.) |  | Arochlor 1232 (PCB) (fl. conc.) | 11141-16-5 | Polychlorinated biphenylds (PCBs); | Polychlorinated biphenylds (PCBs); EMS code: A019 | Chemical |  | LAB
A019 | ug/g | Arochlor 1232 (PCB) | X383 | Extr, GC/MS or GC/ECD | Determination of Polychloriated Biphenyls (PCBs) In Soil by GC/ECD. PURPOSE/PRINCIPLE OF METHOD: This document details the requirements and procedures for the analysis of polychlorinated biphenyls (PCBs) in soil by gas chromatography electron capture detector (GC/ECD). A homogenized soil sample is extracted using dichloromethane (DCM). The extract may be cleaned up using column chromatography with 2% Deactivated Florisil. When contamination with sulphur is suspected, further cleanup using copper is required. Analysis is performed using a GC/ECD. Similar methods have been developed for other matrices including waste and transformer oils, water, swabs, florisil tubes, biota. Generally, the sample extraction procedures differ while the GC analysis methods are similar or identical for all matrices. SCOPE: This procedure describes the determination of PCBs in soil and sediments. Other materials such as wood waste may also be analyzed by this method. DETECTION LIMITS: Target Compound Reported Detection Limit (mg/g) Arochlor 1242 0.03, Arochlor 1248 0.03, Arochlor 1254 0.03, Arochlor 1260 0.03, INTERFERENCES: Interferences (e.g. co-eluting peaks on the chromatogram) may be due to sulphur, phthalate esters, or from chlorinated compounds. Suspect samples must be reanalyzed using appropriate cleanup steps. SAMPLE HANDLING & PRESERVATION: Sample Requirements: Whenever possible, samples must completely fill the bottle such that little or no headspace is present to eliminate the loss of volatile hydrocarbon components. (~125g of sample) Sample Containers: Glass jars that have been solvent rinsed with acetone then hexane, then dried. Shipment and Storage: Samples are to be refrigerated (at 4±2°C) during shipping and prior to analysis. Samples suspected to contain high levels of analyte are to be segregated from other samples. Holding Time: Holding time prior to extraction is 14 days. REFERENCE: U.S. EPA Method 8080:Organochlorine Pesticides and PCBs Analysis. | 0.03 | Y | Y | Y | 46 | 3/16/2020 | 9/7/2022 | 1/11/2024 1:07:06 PM | 2013-01-30 |  | RDGIBSON |  | MassPerMass | mg/kg | Y | (mass conc.) |  | Arochlor 1232 (PCB) (mass conc.) | 11141-16-5 | Polychlorinated biphenylds (PCBs); | Polychlorinated biphenylds (PCBs); EMS code: A019 | Chemical |  | LAB
A020 | mg/L | Arochlor 1016 (PCB) | E660 | EPA 8270E (mod) | EPA 8270E (mod) Miscellaneous Pesticides by GC-MS | 2e-05 | Y | Y | Y | 8 | 5/24/2023 | 11/1/2023 | 11/16/2023 3:38:09 PM | 2023-08-16 |  | EMS |  | Concentration | mg/L | Y | (fl. conc.) |  | Arochlor 1016 (PCB) (fl. conc.) | 12674-11-2 | Polychlorinated biphenylds (PCBs); | Polychlorinated biphenylds (PCBs); EMS code: A020 | Chemical |  | LAB
A020 | ug/g wet | Arochlor 1016 (PCB) | FB03 | Frozen,Acetone Extr,GC/ECD | Sediment and homogenized biota samples are Acetone extracted using a wrist-action shaker, partitioned with 2% NaCl, then treated using Florisil, concentrated Sulfuric Acid, activated Copper and Mercury (as required) to remove interferences. The samples are then analysed using Megabore Column Gas Chromatography with Electron Capture Detection. Note that biota results are reported on a "wet weight" basis. | 0.01 | Y | Y | Y | 2 | 4/12/2001 11:51:00 AM | 4/12/2001 11:51:00 AM | 9/11/2002 3:23:55 PM | 1997-05-15 | 6/3/1998 6:51:47 PM | MPAK |  | MassPerMass | µg/g | Y | (mass conc.) | (wet) | Arochlor 1016 (PCB) (wet) (mass conc.) | 12674-11-2 | Polychlorinated biphenylds (PCBs); | Polychlorinated biphenylds (PCBs); EMS code: A020 | Chemical |  | LAB
A020 | ug/g | Arochlor 1016 (PCB) | X383 | Extr, GC/MS or GC/ECD | Determination of Polychloriated Biphenyls (PCBs) In Soil by GC/ECD. PURPOSE/PRINCIPLE OF METHOD: This document details the requirements and procedures for the analysis of polychlorinated biphenyls (PCBs) in soil by gas chromatography electron capture detector (GC/ECD). A homogenized soil sample is extracted using dichloromethane (DCM). The extract may be cleaned up using column chromatography with 2% Deactivated Florisil. When contamination with sulphur is suspected, further cleanup using copper is required. Analysis is performed using a GC/ECD. Similar methods have been developed for other matrices including waste and transformer oils, water, swabs, florisil tubes, biota. Generally, the sample extraction procedures differ while the GC analysis methods are similar or identical for all matrices. SCOPE: This procedure describes the determination of PCBs in soil and sediments. Other materials such as wood waste may also be analyzed by this method. DETECTION LIMITS: Target Compound Reported Detection Limit (mg/g) Arochlor 1242 0.03, Arochlor 1248 0.03, Arochlor 1254 0.03, Arochlor 1260 0.03, INTERFERENCES: Interferences (e.g. co-eluting peaks on the chromatogram) may be due to sulphur, phthalate esters, or from chlorinated compounds. Suspect samples must be reanalyzed using appropriate cleanup steps. SAMPLE HANDLING & PRESERVATION: Sample Requirements: Whenever possible, samples must completely fill the bottle such that little or no headspace is present to eliminate the loss of volatile hydrocarbon components. (~125g of sample) Sample Containers: Glass jars that have been solvent rinsed with acetone then hexane, then dried. Shipment and Storage: Samples are to be refrigerated (at 4±2°C) during shipping and prior to analysis. Samples suspected to contain high levels of analyte are to be segregated from other samples. Holding Time: Holding time prior to extraction is 14 days. REFERENCE: U.S. EPA Method 8080:Organochlorine Pesticides and PCBs Analysis. | 0.03 | Y | Y | Y | 46 | 3/16/2020 | 9/7/2022 | 1/11/2024 1:07:06 PM | 2013-01-30 |  | RDGIBSON |  | MassPerMass | mg/kg | Y | (mass conc.) |  | Arochlor 1016 (PCB) (mass conc.) | 12674-11-2 | Polychlorinated biphenylds (PCBs); | Polychlorinated biphenylds (PCBs); EMS code: A020 | Chemical |  | LAB
A020 | ug/g | Arochlor 1016 (PCB) | FS06 | Frozen,Acetone extr,GC/ECD | Sediment and homogenized biota samples are Acetone extracted using a wrist-action shaker, partitioned with 2% NaCl, then treated using Florisil, concentrated Sulfuric Acid, activated Copper and Mercury (as required) to remove interferences. The samples are then analysed using Megabore Column Gas Chromatography with Electron Capture Detection. | 0.01 | Y | Y | Y | 60 | 2/18/2000 5:06:00 PM | 11/9/2001 1:34:00 PM | 8/21/2002 2:09:54 PM | 1997-05-15 | 6/3/1998 6:51:47 PM | MPAK |  | MassPerMass | mg/kg | Y | (mass conc.) |  | Arochlor 1016 (PCB) (mass conc.) | 12674-11-2 | Polychlorinated biphenylds (PCBs); | Polychlorinated biphenylds (PCBs); EMS code: A020 | Chemical |  | LAB
A020 | ug/g | Arochlor 1016 (PCB) | F115 | Hex dil,Hex extr'n,GC/ECD | Effluent and swab samples are extracted with hexane, the extract concentrated, and then treated using Florisil, concentrated Sulfuric Acid, activated Copper and Mercury (as required) to remove interferences. The samples are then analysed using Megabore Column Gas Chromatography with Electron Capture Detection. Oil samples are first diluted with Hexane then treated in the manner outlined for effluent and swab samples. The samples are then analysed using Megabore Column Gas Chromatography with Electron Capture Detection. | 0.1 | Y | Y | Y | 3 | 4/19/2001 9:04:00 AM | 4/19/2001 9:06:00 AM | 6/1/2001 3:54:28 AM | 2001-05-24 |  | NPEPPIN |  | MassPerMass | mg/kg | Y | (mass conc.) |  | Arochlor 1016 (PCB) (mass conc.) | 12674-11-2 | Polychlorinated biphenylds (PCBs); | Polychlorinated biphenylds (PCBs); EMS code: A020 | Chemical |  | LAB
A020 | mg/L | Arochlor 1016 (PCB) | F018 | Hexane extr,GC/ECD | Effluent and swab samples are extracted with hexane, the extract concentrated, and then treated using Florisil, concentrated Sulfuric Acid, activated Copper and Mercury (as required) to remove interferences. The samples are then analysed using Megabore Column Gas Chromatography with Electron Capture Detection. Oil samples are first diluted with Hexane then treated in the manner outlined for effluent and swab samples. The samples are then analysed using Megabore Column Gas Chromatography with Electron Capture Detection. | 0.0001 | Y | Y | Y | 18 | 2/28/2000 8:40:00 AM | 9/5/2023 6:46:00 PM | 9/11/2023 5:22:04 PM | 1997-05-15 | 6/3/1998 6:51:47 PM | MPAK |  | Concentration | mg/L | Y | (fl. conc.) |  | Arochlor 1016 (PCB) (fl. conc.) | 12674-11-2 | Polychlorinated biphenylds (PCBs); | Polychlorinated biphenylds (PCBs); EMS code: A020 | Chemical |  | LAB
A020 | mg/L | Arochlor 1016 (PCB) | E-8082 | EPA-8082A | See reference for EPA method 8082A POLYCHLORINATED BIPHENYLS (PCBs) BY GAS CHROMATOGRAPHY |  | Y | Y | Y | 14 | 12/28/2021 | 1/11/2022 | 9/2/2022 6:21:52 PM | 2022-04-06 |  | EMS |  | Concentration | mg/L | Y | (fl. conc.) |  | Arochlor 1016 (PCB) (fl. conc.) | 12674-11-2 | Polychlorinated biphenylds (PCBs); | Polychlorinated biphenylds (PCBs); EMS code: A020 | Chemical |  | LAB
A021 | mg/L | Arochlor 1254 (PCB) | E660 | EPA 8270E (mod) | EPA 8270E (mod) Miscellaneous Pesticides by GC-MS | 2e-05 | Y | Y | Y | 8 | 5/24/2023 | 11/1/2023 | 11/16/2023 3:38:09 PM | 2023-08-16 |  | EMS |  | Concentration | mg/L | Y | (fl. conc.) |  | Arochlor 1254 (PCB) (fl. conc.) | 11097-69-1 | Polychlorinated biphenylds (PCBs); | Polychlorinated biphenylds (PCBs); EMS code: A021 | Chemical |  | LAB
A021 | ug/g wet | Arochlor 1254 (PCB) | FB03 | Frozen,Acetone Extr,GC/ECD | Sediment and homogenized biota samples are Acetone extracted using a wrist-action shaker, partitioned with 2% NaCl, then treated using Florisil, concentrated Sulfuric Acid, activated Copper and Mercury (as required) to remove interferences. The samples are then analysed using Megabore Column Gas Chromatography with Electron Capture Detection. Note that biota results are reported on a "wet weight" basis. | 0.01 | Y | Y | Y | 2 | 4/12/2001 11:51:00 AM | 4/12/2001 11:51:00 AM | 9/11/2002 3:23:56 PM | 1997-05-15 | 6/3/1998 6:51:47 PM | MPAK |  | MassPerMass | µg/g | Y | (mass conc.) | (wet) | Arochlor 1254 (PCB) (wet) (mass conc.) | 11097-69-1 | Polychlorinated biphenylds (PCBs); | Polychlorinated biphenylds (PCBs); EMS code: A021 | Chemical |  | LAB
A021 | ug/g | Arochlor 1254 (PCB) | X383 | Extr, GC/MS or GC/ECD | Determination of Polychloriated Biphenyls (PCBs) In Soil by GC/ECD. PURPOSE/PRINCIPLE OF METHOD: This document details the requirements and procedures for the analysis of polychlorinated biphenyls (PCBs) in soil by gas chromatography electron capture detector (GC/ECD). A homogenized soil sample is extracted using dichloromethane (DCM). The extract may be cleaned up using column chromatography with 2% Deactivated Florisil. When contamination with sulphur is suspected, further cleanup using copper is required. Analysis is performed using a GC/ECD. Similar methods have been developed for other matrices including waste and transformer oils, water, swabs, florisil tubes, biota. Generally, the sample extraction procedures differ while the GC analysis methods are similar or identical for all matrices. SCOPE: This procedure describes the determination of PCBs in soil and sediments. Other materials such as wood waste may also be analyzed by this method. DETECTION LIMITS: Target Compound Reported Detection Limit (mg/g) Arochlor 1242 0.03, Arochlor 1248 0.03, Arochlor 1254 0.03, Arochlor 1260 0.03, INTERFERENCES: Interferences (e.g. co-eluting peaks on the chromatogram) may be due to sulphur, phthalate esters, or from chlorinated compounds. Suspect samples must be reanalyzed using appropriate cleanup steps. SAMPLE HANDLING & PRESERVATION: Sample Requirements: Whenever possible, samples must completely fill the bottle such that little or no headspace is present to eliminate the loss of volatile hydrocarbon components. (~125g of sample) Sample Containers: Glass jars that have been solvent rinsed with acetone then hexane, then dried. Shipment and Storage: Samples are to be refrigerated (at 4±2°C) during shipping and prior to analysis. Samples suspected to contain high levels of analyte are to be segregated from other samples. Holding Time: Holding time prior to extraction is 14 days. REFERENCE: U.S. EPA Method 8080:Organochlorine Pesticides and PCBs Analysis. | 0.03 | Y | Y | Y | 46 | 3/16/2020 | 9/7/2022 | 1/11/2024 1:07:06 PM | 2013-01-30 |  | RDGIBSON |  | MassPerMass | mg/kg | Y | (mass conc.) |  | Arochlor 1254 (PCB) (mass conc.) | 11097-69-1 | Polychlorinated biphenylds (PCBs); | Polychlorinated biphenylds (PCBs); EMS code: A021 | Chemical |  | LAB
A021 | ug/g wet | Arochlor 1254 (PCB) | EQ44 | EQUIS (ug/g wet) | EQUIS (ug/g wet) | 0.02 | N | Y | Y | 2 |  |  | 10/22/1996 9:03:11 PM | 1996-04-12 | 6/3/1998 6:51:33 PM | EMS |  | MassPerMass | µg/g | Y | (mass conc.) | (wet) | Arochlor 1254 (PCB) (wet) (mass conc.) | 11097-69-1 | Polychlorinated biphenylds (PCBs); | Polychlorinated biphenylds (PCBs); EMS code: A021 | Chemical |  | LAB
A021 | mg/L | Arochlor 1254 (PCB) | EQ01 | EQUIS (mg/L) | EQUIS (mg/L) | 0.0004 | N | Y | Y | 19 |  |  | 10/22/1996 11:26:53 PM | 1996-04-12 | 6/3/1998 6:51:33 PM | EMS |  | Concentration | mg/L | Y | (fl. conc.) |  | Arochlor 1254 (PCB) (fl. conc.) | 11097-69-1 | Polychlorinated biphenylds (PCBs); | Polychlorinated biphenylds (PCBs); EMS code: A021 | Chemical |  | LAB
A021 | ug/g | Arochlor 1254 (PCB) | F115 | Hex dil,Hex extr'n,GC/ECD | Effluent and swab samples are extracted with hexane, the extract concentrated, and then treated using Florisil, concentrated Sulfuric Acid, activated Copper and Mercury (as required) to remove interferences. The samples are then analysed using Megabore Column Gas Chromatography with Electron Capture Detection. Oil samples are first diluted with Hexane then treated in the manner outlined for effluent and swab samples. The samples are then analysed using Megabore Column Gas Chromatography with Electron Capture Detection. | 0.1 | Y | Y | Y | 3 | 4/19/2001 9:04:00 AM | 4/19/2001 9:06:00 AM | 6/1/2001 3:54:28 AM | 2001-05-24 |  | NPEPPIN |  | MassPerMass | mg/kg | Y | (mass conc.) |  | Arochlor 1254 (PCB) (mass conc.) | 11097-69-1 | Polychlorinated biphenylds (PCBs); | Polychlorinated biphenylds (PCBs); EMS code: A021 | Chemical |  | LAB
A021 | ug/g | Arochlor 1254 (PCB) | FS06 | Frozen,Acetone extr,GC/ECD | Sediment and homogenized biota samples are Acetone extracted using a wrist-action shaker, partitioned with 2% NaCl, then treated using Florisil, concentrated Sulfuric Acid, activated Copper and Mercury (as required) to remove interferences. The samples are then analysed using Megabore Column Gas Chromatography with Electron Capture Detection. | 0.01 | Y | Y | Y | 60 | 2/18/2000 5:06:00 PM | 11/9/2001 1:34:00 PM | 8/21/2002 2:09:54 PM | 1997-05-15 | 6/3/1998 6:51:47 PM | MPAK |  | MassPerMass | mg/kg | Y | (mass conc.) |  | Arochlor 1254 (PCB) (mass conc.) | 11097-69-1 | Polychlorinated biphenylds (PCBs); | Polychlorinated biphenylds (PCBs); EMS code: A021 | Chemical |  | LAB
A021 | mg/L | Arochlor 1254 (PCB) | E-8082 | EPA-8082A | See reference for EPA method 8082A POLYCHLORINATED BIPHENYLS (PCBs) BY GAS CHROMATOGRAPHY |  | Y | Y | Y | 14 | 12/28/2021 | 1/11/2022 | 9/2/2022 6:21:52 PM | 2022-04-06 |  | EMS |  | Concentration | mg/L | Y | (fl. conc.) |  | Arochlor 1254 (PCB) (fl. conc.) | 11097-69-1 | Polychlorinated biphenylds (PCBs); | Polychlorinated biphenylds (PCBs); EMS code: A021 | Chemical |  | LAB
A021 | mg/L | Arochlor 1254 (PCB) | X303 | CH2Cl2 Extr:GC/ECD | Polychlorinated Biphenyls (PCBs) In Water by Gas Chromatography-Electron Capture Detector (GC/ECD). A water sample, usually 1000 mL, is extracted using dichloromethane. Sulphuric acid clean-up step is performed in every sample. Dirty samples with dark colored extracts are put through a clean-up step using 2% Deactivated Florisil. If sulphur is present in the water samples, especially for the river or stream water samples, a copper metal clean-up step may be used. A final analysis is performed using a GC equipped with an ECD. SCOPE: This procedure describes the determination of PCBs in water. If Organochlorine Pesticides (OC's) and PCB's are to be analyzed on the same water sample, a separate test method (Method for the Determination of Pesticides in Water by GC/MS and LC/MS or HPLC Fluorescence Detection DCN 62-C-043) is used. DETECTION LIMITS: Target Compound Reported Detection Limit (mg/L) Arochlor 1221 0.1 Arochlor 1242 0.1 Arochlor 1248 0.1 Arochlor 1254 0.1 Arochlor 1260 0.1 INTERFERENCES: Interferences (e.g. co-eluting peaks on the chromatogram) may be due to sulfur,phthalate esters, or from chlorinated compounds. Suspect samples must be reanalyzed using appropriate cleanup steps. SAMPLE HANDLING & PRESERVATION: Sample Containers: For sample collection, use 1 litre amber glass bottles that have been pre-rinsed with acetone and hexane then dried. Bottle lids must be teflon-lined, or alternatively lined with solvent-rinsed aluminium foil. Shipment and Storage: Samples must be refrigerated (at 2-8°C) and stored in the dark. Holding Time: Samples must be extracted within 7 days of sampling; REFERENCE: U.S. EPA Method 8080: Organochlorine Pesticides and PCB's Analysis. SW-846 3rd Ed.,September, 1986. | 0.001 | Y | Y | Y | 2 | 12/21/2012 10:31:00 AM | 4/2/2013 3:22:00 PM | 4/12/2013 3:34:31 PM | 2013-01-30 |  | RDGIBSON |  | Concentration | mg/L | Y | (fl. conc.) |  | Arochlor 1254 (PCB) (fl. conc.) | 11097-69-1 | Polychlorinated biphenylds (PCBs); | Polychlorinated biphenylds (PCBs); EMS code: A021 | Chemical |  | LAB
A021 | mg/L | Arochlor 1254 (PCB) | F018 | Hexane extr,GC/ECD | Effluent and swab samples are extracted with hexane, the extract concentrated, and then treated using Florisil, concentrated Sulfuric Acid, activated Copper and Mercury (as required) to remove interferences. The samples are then analysed using Megabore Column Gas Chromatography with Electron Capture Detection. Oil samples are first diluted with Hexane then treated in the manner outlined for effluent and swab samples. The samples are then analysed using Megabore Column Gas Chromatography with Electron Capture Detection. | 0.0001 | Y | Y | Y | 18 | 2/28/2000 8:40:00 AM | 9/5/2023 6:46:00 PM | 9/11/2023 5:22:04 PM | 1997-05-15 | 6/3/1998 6:51:47 PM | MPAK |  | Concentration | mg/L | Y | (fl. conc.) |  | Arochlor 1254 (PCB) (fl. conc.) | 11097-69-1 | Polychlorinated biphenylds (PCBs); | Polychlorinated biphenylds (PCBs); EMS code: A021 | Chemical |  | LAB
A022 | mg/L | Arochlor 1260 (PCB) | F018 | Hexane extr,GC/ECD | Effluent and swab samples are extracted with hexane, the extract concentrated, and then treated using Florisil, concentrated Sulfuric Acid, activated Copper and Mercury (as required) to remove interferences. The samples are then analysed using Megabore Column Gas Chromatography with Electron Capture Detection. Oil samples are first diluted with Hexane then treated in the manner outlined for effluent and swab samples. The samples are then analysed using Megabore Column Gas Chromatography with Electron Capture Detection. | 0.0001 | Y | Y | Y | 18 | 2/28/2000 8:40:00 AM | 9/5/2023 6:46:00 PM | 9/11/2023 5:22:04 PM | 1997-05-15 | 6/3/1998 6:51:47 PM | MPAK |  | Concentration | mg/L | Y | (fl. conc.) |  | Arochlor 1260 (PCB) (fl. conc.) | 11096-82-5 | Polychlorinated biphenylds (PCBs); | Polychlorinated biphenylds (PCBs); EMS code: A022 | Chemical |  | LAB
A022 | ug/g wet | Arochlor 1260 (PCB) | FB03 | Frozen,Acetone Extr,GC/ECD | Sediment and homogenized biota samples are Acetone extracted using a wrist-action shaker, partitioned with 2% NaCl, then treated using Florisil, concentrated Sulfuric Acid, activated Copper and Mercury (as required) to remove interferences. The samples are then analysed using Megabore Column Gas Chromatography with Electron Capture Detection. Note that biota results are reported on a "wet weight" basis. | 0.01 | Y | Y | Y | 2 | 4/12/2001 11:51:00 AM | 4/12/2001 11:51:00 AM | 9/11/2002 3:23:56 PM | 1997-05-15 | 6/3/1998 6:51:47 PM | MPAK |  | MassPerMass | µg/g | Y | (mass conc.) | (wet) | Arochlor 1260 (PCB) (wet) (mass conc.) | 11096-82-5 | Polychlorinated biphenylds (PCBs); | Polychlorinated biphenylds (PCBs); EMS code: A022 | Chemical |  | LAB
A022 | ug/g | Arochlor 1260 (PCB) | FS06 | Frozen,Acetone extr,GC/ECD | Sediment and homogenized biota samples are Acetone extracted using a wrist-action shaker, partitioned with 2% NaCl, then treated using Florisil, concentrated Sulfuric Acid, activated Copper and Mercury (as required) to remove interferences. The samples are then analysed using Megabore Column Gas Chromatography with Electron Capture Detection. | 0.01 | Y | Y | Y | 60 | 2/18/2000 5:06:00 PM | 11/9/2001 1:34:00 PM | 8/21/2002 2:09:55 PM | 1997-05-15 | 6/3/1998 6:51:47 PM | MPAK |  | MassPerMass | mg/kg | Y | (mass conc.) |  | Arochlor 1260 (PCB) (mass conc.) | 11096-82-5 | Polychlorinated biphenylds (PCBs); | Polychlorinated biphenylds (PCBs); EMS code: A022 | Chemical |  | LAB
A022 | mg/L | Arochlor 1260 (PCB) | E-8082 | EPA-8082A | See reference for EPA method 8082A POLYCHLORINATED BIPHENYLS (PCBs) BY GAS CHROMATOGRAPHY |  | Y | Y | Y | 14 | 12/28/2021 | 1/11/2022 | 9/2/2022 6:21:52 PM | 2022-04-06 |  | EMS |  | Concentration | mg/L | Y | (fl. conc.) |  | Arochlor 1260 (PCB) (fl. conc.) | 11096-82-5 | Polychlorinated biphenylds (PCBs); | Polychlorinated biphenylds (PCBs); EMS code: A022 | Chemical |  | LAB
A022 | ug/g | Arochlor 1260 (PCB) | X383 | Extr, GC/MS or GC/ECD | Determination of Polychloriated Biphenyls (PCBs) In Soil by GC/ECD. PURPOSE/PRINCIPLE OF METHOD: This document details the requirements and procedures for the analysis of polychlorinated biphenyls (PCBs) in soil by gas chromatography electron capture detector (GC/ECD). A homogenized soil sample is extracted using dichloromethane (DCM). The extract may be cleaned up using column chromatography with 2% Deactivated Florisil. When contamination with sulphur is suspected, further cleanup using copper is required. Analysis is performed using a GC/ECD. Similar methods have been developed for other matrices including waste and transformer oils, water, swabs, florisil tubes, biota. Generally, the sample extraction procedures differ while the GC analysis methods are similar or identical for all matrices. SCOPE: This procedure describes the determination of PCBs in soil and sediments. Other materials such as wood waste may also be analyzed by this method. DETECTION LIMITS: Target Compound Reported Detection Limit (mg/g) Arochlor 1242 0.03, Arochlor 1248 0.03, Arochlor 1254 0.03, Arochlor 1260 0.03, INTERFERENCES: Interferences (e.g. co-eluting peaks on the chromatogram) may be due to sulphur, phthalate esters, or from chlorinated compounds. Suspect samples must be reanalyzed using appropriate cleanup steps. SAMPLE HANDLING & PRESERVATION: Sample Requirements: Whenever possible, samples must completely fill the bottle such that little or no headspace is present to eliminate the loss of volatile hydrocarbon components. (~125g of sample) Sample Containers: Glass jars that have been solvent rinsed with acetone then hexane, then dried. Shipment and Storage: Samples are to be refrigerated (at 4±2°C) during shipping and prior to analysis. Samples suspected to contain high levels of analyte are to be segregated from other samples. Holding Time: Holding time prior to extraction is 14 days. REFERENCE: U.S. EPA Method 8080:Organochlorine Pesticides and PCBs Analysis. | 0.03 | Y | Y | Y | 46 | 3/16/2020 | 9/7/2022 | 1/11/2024 1:07:06 PM | 2013-01-30 |  | RDGIBSON |  | MassPerMass | mg/kg | Y | (mass conc.) |  | Arochlor 1260 (PCB) (mass conc.) | 11096-82-5 | Polychlorinated biphenylds (PCBs); | Polychlorinated biphenylds (PCBs); EMS code: A022 | Chemical |  | LAB
A022 | ug/g wet | Arochlor 1260 (PCB) | EQ44 | EQUIS (ug/g wet) | EQUIS (ug/g wet) | 0.02 | N | Y | Y | 2 |  |  | 10/22/1996 9:03:11 PM | 1996-04-12 | 6/3/1998 6:51:33 PM | EMS |  | MassPerMass | µg/g | Y | (mass conc.) | (wet) | Arochlor 1260 (PCB) (wet) (mass conc.) | 11096-82-5 | Polychlorinated biphenylds (PCBs); | Polychlorinated biphenylds (PCBs); EMS code: A022 | Chemical |  | LAB
A022 | mg/L | Arochlor 1260 (PCB) | EQ01 | EQUIS (mg/L) | EQUIS (mg/L) | 0.0004 | N | Y | Y | 19 |  |  | 10/22/1996 11:26:53 PM | 1996-04-12 | 6/3/1998 6:51:45 PM | EMS |  | Concentration | mg/L | Y | (fl. conc.) |  | Arochlor 1260 (PCB) (fl. conc.) | 11096-82-5 | Polychlorinated biphenylds (PCBs); | Polychlorinated biphenylds (PCBs); EMS code: A022 | Chemical |  | LAB
A022 | mg/L | Arochlor 1260 (PCB) | E660 | EPA 8270E (mod) | EPA 8270E (mod) Miscellaneous Pesticides by GC-MS | 2e-05 | Y | Y | Y | 8 | 5/24/2023 | 11/1/2023 | 11/16/2023 3:38:09 PM | 2023-08-16 |  | EMS |  | Concentration | mg/L | Y | (fl. conc.) |  | Arochlor 1260 (PCB) (fl. conc.) | 11096-82-5 | Polychlorinated biphenylds (PCBs); | Polychlorinated biphenylds (PCBs); EMS code: A022 | Chemical |  | LAB
A022 | ug/g | Arochlor 1260 (PCB) | F115 | Hex dil,Hex extr'n,GC/ECD | Effluent and swab samples are extracted with hexane, the extract concentrated, and then treated using Florisil, concentrated Sulfuric Acid, activated Copper and Mercury (as required) to remove interferences. The samples are then analysed using Megabore Column Gas Chromatography with Electron Capture Detection. Oil samples are first diluted with Hexane then treated in the manner outlined for effluent and swab samples. The samples are then analysed using Megabore Column Gas Chromatography with Electron Capture Detection. | 0.1 | Y | Y | Y | 3 | 4/19/2001 9:04:00 AM | 4/19/2001 9:06:00 AM | 6/1/2001 3:54:28 AM | 2001-05-24 |  | NPEPPIN |  | MassPerMass | mg/kg | Y | (mass conc.) |  | Arochlor 1260 (PCB) (mass conc.) | 11096-82-5 | Polychlorinated biphenylds (PCBs); | Polychlorinated biphenylds (PCBs); EMS code: A022 | Chemical |  | LAB
A022 | mg/L | Arochlor 1260 (PCB) | X303 | CH2Cl2 Extr:GC/ECD | Polychlorinated Biphenyls (PCBs) In Water by Gas Chromatography-Electron Capture Detector (GC/ECD). A water sample, usually 1000 mL, is extracted using dichloromethane. Sulphuric acid clean-up step is performed in every sample. Dirty samples with dark colored extracts are put through a clean-up step using 2% Deactivated Florisil. If sulphur is present in the water samples, especially for the river or stream water samples, a copper metal clean-up step may be used. A final analysis is performed using a GC equipped with an ECD. SCOPE: This procedure describes the determination of PCBs in water. If Organochlorine Pesticides (OC's) and PCB's are to be analyzed on the same water sample, a separate test method (Method for the Determination of Pesticides in Water by GC/MS and LC/MS or HPLC Fluorescence Detection DCN 62-C-043) is used. DETECTION LIMITS: Target Compound Reported Detection Limit (mg/L) Arochlor 1221 0.1 Arochlor 1242 0.1 Arochlor 1248 0.1 Arochlor 1254 0.1 Arochlor 1260 0.1 INTERFERENCES: Interferences (e.g. co-eluting peaks on the chromatogram) may be due to sulfur,phthalate esters, or from chlorinated compounds. Suspect samples must be reanalyzed using appropriate cleanup steps. SAMPLE HANDLING & PRESERVATION: Sample Containers: For sample collection, use 1 litre amber glass bottles that have been pre-rinsed with acetone and hexane then dried. Bottle lids must be teflon-lined, or alternatively lined with solvent-rinsed aluminium foil. Shipment and Storage: Samples must be refrigerated (at 2-8°C) and stored in the dark. Holding Time: Samples must be extracted within 7 days of sampling; REFERENCE: U.S. EPA Method 8080: Organochlorine Pesticides and PCB's Analysis. SW-846 3rd Ed.,September, 1986. | 0.001 | Y | Y | Y | 2 | 12/21/2012 10:31:00 AM | 4/2/2013 3:22:00 PM | 4/12/2013 3:34:31 PM | 2013-01-30 |  | RDGIBSON |  | Concentration | mg/L | Y | (fl. conc.) |  | Arochlor 1260 (PCB) (fl. conc.) | 11096-82-5 | Polychlorinated biphenylds (PCBs); | Polychlorinated biphenylds (PCBs); EMS code: A022 | Chemical |  | LAB
A023 | mg/L | Arochlor 1248 (PCB) | F018 | Hexane extr,GC/ECD | Effluent and swab samples are extracted with hexane, the extract concentrated, and then treated using Florisil, concentrated Sulfuric Acid, activated Copper and Mercury (as required) to remove interferences. The samples are then analysed using Megabore Column Gas Chromatography with Electron Capture Detection. Oil samples are first diluted with Hexane then treated in the manner outlined for effluent and swab samples. The samples are then analysed using Megabore Column Gas Chromatography with Electron Capture Detection. | 0.0001 | Y | Y | Y | 18 | 2/28/2000 8:40:00 AM | 9/5/2023 6:46:00 PM | 9/11/2023 5:22:04 PM | 1997-05-15 | 6/3/1998 6:51:47 PM | MPAK |  | Concentration | mg/L | Y | (fl. conc.) |  | Arochlor 1248 (PCB) (fl. conc.) | 12672-29-6 | Polychlorinated biphenylds (PCBs); | Polychlorinated biphenylds (PCBs); EMS code: A023 | Chemical |  | LAB
A023 | mg/L | Arochlor 1248 (PCB) | E660 | EPA 8270E (mod) | EPA 8270E (mod) Miscellaneous Pesticides by GC-MS | 2e-05 | Y | Y | Y | 8 | 5/24/2023 | 11/1/2023 | 11/16/2023 3:38:09 PM | 2023-08-16 |  | EMS |  | Concentration | mg/L | Y | (fl. conc.) |  | Arochlor 1248 (PCB) (fl. conc.) | 12672-29-6 | Polychlorinated biphenylds (PCBs); | Polychlorinated biphenylds (PCBs); EMS code: A023 | Chemical |  | LAB
A023 | ug/g | Arochlor 1248 (PCB) | F115 | Hex dil,Hex extr'n,GC/ECD | Effluent and swab samples are extracted with hexane, the extract concentrated, and then treated using Florisil, concentrated Sulfuric Acid, activated Copper and Mercury (as required) to remove interferences. The samples are then analysed using Megabore Column Gas Chromatography with Electron Capture Detection. Oil samples are first diluted with Hexane then treated in the manner outlined for effluent and swab samples. The samples are then analysed using Megabore Column Gas Chromatography with Electron Capture Detection. | 0.1 | Y | Y | Y | 3 | 4/19/2001 9:04:00 AM | 4/19/2001 9:06:00 AM | 6/1/2001 3:54:29 AM | 2001-05-24 |  | NPEPPIN |  | MassPerMass | mg/kg | Y | (mass conc.) |  | Arochlor 1248 (PCB) (mass conc.) | 12672-29-6 | Polychlorinated biphenylds (PCBs); | Polychlorinated biphenylds (PCBs); EMS code: A023 | Chemical |  | LAB
A023 | ug/g wet | Arochlor 1248 (PCB) | FB03 | Frozen,Acetone Extr,GC/ECD | Sediment and homogenized biota samples are Acetone extracted using a wrist-action shaker, partitioned with 2% NaCl, then treated using Florisil, concentrated Sulfuric Acid, activated Copper and Mercury (as required) to remove interferences. The samples are then analysed using Megabore Column Gas Chromatography with Electron Capture Detection. Note that biota results are reported on a "wet weight" basis. | 0.01 | Y | Y | Y | 2 | 4/12/2001 11:51:00 AM | 4/12/2001 11:51:00 AM | 9/11/2002 3:23:56 PM | 1997-05-15 | 6/3/1998 6:51:47 PM | MPAK |  | MassPerMass | µg/g | Y | (mass conc.) | (wet) | Arochlor 1248 (PCB) (wet) (mass conc.) | 12672-29-6 | Polychlorinated biphenylds (PCBs); | Polychlorinated biphenylds (PCBs); EMS code: A023 | Chemical |  | LAB
A023 | ug/g | Arochlor 1248 (PCB) | X383 | Extr, GC/MS or GC/ECD | Determination of Polychloriated Biphenyls (PCBs) In Soil by GC/ECD. PURPOSE/PRINCIPLE OF METHOD: This document details the requirements and procedures for the analysis of polychlorinated biphenyls (PCBs) in soil by gas chromatography electron capture detector (GC/ECD). A homogenized soil sample is extracted using dichloromethane (DCM). The extract may be cleaned up using column chromatography with 2% Deactivated Florisil. When contamination with sulphur is suspected, further cleanup using copper is required. Analysis is performed using a GC/ECD. Similar methods have been developed for other matrices including waste and transformer oils, water, swabs, florisil tubes, biota. Generally, the sample extraction procedures differ while the GC analysis methods are similar or identical for all matrices. SCOPE: This procedure describes the determination of PCBs in soil and sediments. Other materials such as wood waste may also be analyzed by this method. DETECTION LIMITS: Target Compound Reported Detection Limit (mg/g) Arochlor 1242 0.03, Arochlor 1248 0.03, Arochlor 1254 0.03, Arochlor 1260 0.03, INTERFERENCES: Interferences (e.g. co-eluting peaks on the chromatogram) may be due to sulphur, phthalate esters, or from chlorinated compounds. Suspect samples must be reanalyzed using appropriate cleanup steps. SAMPLE HANDLING & PRESERVATION: Sample Requirements: Whenever possible, samples must completely fill the bottle such that little or no headspace is present to eliminate the loss of volatile hydrocarbon components. (~125g of sample) Sample Containers: Glass jars that have been solvent rinsed with acetone then hexane, then dried. Shipment and Storage: Samples are to be refrigerated (at 4±2°C) during shipping and prior to analysis. Samples suspected to contain high levels of analyte are to be segregated from other samples. Holding Time: Holding time prior to extraction is 14 days. REFERENCE: U.S. EPA Method 8080:Organochlorine Pesticides and PCBs Analysis. | 0.03 | Y | Y | Y | 46 | 3/16/2020 | 9/7/2022 | 1/11/2024 1:07:06 PM | 2013-01-30 |  | RDGIBSON |  | MassPerMass | mg/kg | Y | (mass conc.) |  | Arochlor 1248 (PCB) (mass conc.) | 12672-29-6 | Polychlorinated biphenylds (PCBs); | Polychlorinated biphenylds (PCBs); EMS code: A023 | Chemical |  | LAB
A023 | mg/L | Arochlor 1248 (PCB) | E-8082 | EPA-8082A | See reference for EPA method 8082A POLYCHLORINATED BIPHENYLS (PCBs) BY GAS CHROMATOGRAPHY |  | Y | Y | Y | 14 | 12/28/2021 | 1/11/2022 | 9/2/2022 6:21:52 PM | 2022-04-06 |  | EMS |  | Concentration | mg/L | Y | (fl. conc.) |  | Arochlor 1248 (PCB) (fl. conc.) | 12672-29-6 | Polychlorinated biphenylds (PCBs); | Polychlorinated biphenylds (PCBs); EMS code: A023 | Chemical |  | LAB
A023 | ug/g | Arochlor 1248 (PCB) | FS06 | Frozen,Acetone extr,GC/ECD | Sediment and homogenized biota samples are Acetone extracted using a wrist-action shaker, partitioned with 2% NaCl, then treated using Florisil, concentrated Sulfuric Acid, activated Copper and Mercury (as required) to remove interferences. The samples are then analysed using Megabore Column Gas Chromatography with Electron Capture Detection. | 0.01 | Y | Y | Y | 60 | 2/18/2000 5:06:00 PM | 11/9/2001 1:34:00 PM | 8/21/2002 2:09:55 PM | 1997-05-15 | 6/3/1998 6:51:47 PM | MPAK |  | MassPerMass | mg/kg | Y | (mass conc.) |  | Arochlor 1248 (PCB) (mass conc.) | 12672-29-6 | Polychlorinated biphenylds (PCBs); | Polychlorinated biphenylds (PCBs); EMS code: A023 | Chemical |  | LAB
A023 | mg/L | Arochlor 1248 (PCB) | X303 | CH2Cl2 Extr:GC/ECD | Polychlorinated Biphenyls (PCBs) In Water by Gas Chromatography-Electron Capture Detector (GC/ECD). A water sample, usually 1000 mL, is extracted using dichloromethane. Sulphuric acid clean-up step is performed in every sample. Dirty samples with dark colored extracts are put through a clean-up step using 2% Deactivated Florisil. If sulphur is present in the water samples, especially for the river or stream water samples, a copper metal clean-up step may be used. A final analysis is performed using a GC equipped with an ECD. SCOPE: This procedure describes the determination of PCBs in water. If Organochlorine Pesticides (OC's) and PCB's are to be analyzed on the same water sample, a separate test method (Method for the Determination of Pesticides in Water by GC/MS and LC/MS or HPLC Fluorescence Detection DCN 62-C-043) is used. DETECTION LIMITS: Target Compound Reported Detection Limit (mg/L) Arochlor 1221 0.1 Arochlor 1242 0.1 Arochlor 1248 0.1 Arochlor 1254 0.1 Arochlor 1260 0.1 INTERFERENCES: Interferences (e.g. co-eluting peaks on the chromatogram) may be due to sulfur,phthalate esters, or from chlorinated compounds. Suspect samples must be reanalyzed using appropriate cleanup steps. SAMPLE HANDLING & PRESERVATION: Sample Containers: For sample collection, use 1 litre amber glass bottles that have been pre-rinsed with acetone and hexane then dried. Bottle lids must be teflon-lined, or alternatively lined with solvent-rinsed aluminium foil. Shipment and Storage: Samples must be refrigerated (at 2-8°C) and stored in the dark. Holding Time: Samples must be extracted within 7 days of sampling; REFERENCE: U.S. EPA Method 8080: Organochlorine Pesticides and PCB's Analysis. SW-846 3rd Ed.,September, 1986. | 0.001 | Y | Y | Y | 2 | 12/21/2012 10:31:00 AM | 4/2/2013 3:22:00 PM | 4/12/2013 3:34:31 PM | 2013-01-30 |  | RDGIBSON |  | Concentration | mg/L | Y | (fl. conc.) |  | Arochlor 1248 (PCB) (fl. conc.) | 12672-29-6 | Polychlorinated biphenylds (PCBs); | Polychlorinated biphenylds (PCBs); EMS code: A023 | Chemical |  | LAB
A024 | mg/L | Arochlor 1262 (PCB) | E-8082 | EPA-8082A | See reference for EPA method 8082A POLYCHLORINATED BIPHENYLS (PCBs) BY GAS CHROMATOGRAPHY |  | Y | Y | Y | 14 | 12/28/2021 | 1/11/2022 | 9/2/2022 6:21:52 PM | 2022-04-06 |  | EMS |  | Concentration | mg/L | Y | (fl. conc.) |  | Arochlor 1262 (PCB) (fl. conc.) | 37324-23-5 | Polychlorinated biphenylds (PCBs); | Polychlorinated biphenylds (PCBs); EMS code: A024 | Chemical |  | LAB
A024 | ug/g | Arochlor 1262 (PCB) | FS06 | Frozen,Acetone extr,GC/ECD | Sediment and homogenized biota samples are Acetone extracted using a wrist-action shaker, partitioned with 2% NaCl, then treated using Florisil, concentrated Sulfuric Acid, activated Copper and Mercury (as required) to remove interferences. The samples are then analysed using Megabore Column Gas Chromatography with Electron Capture Detection. | 0.01 | Y | Y | Y | 60 | 2/18/2000 5:06:00 PM | 11/9/2001 1:34:00 PM | 8/21/2002 2:09:55 PM | 1997-05-26 | 6/3/1998 6:51:47 PM | MPAK |  | MassPerMass | mg/kg | Y | (mass conc.) |  | Arochlor 1262 (PCB) (mass conc.) | 37324-23-5 | Polychlorinated biphenylds (PCBs); | Polychlorinated biphenylds (PCBs); EMS code: A024 | Chemical |  | LAB
A024 | ug/g | Arochlor 1262 (PCB) | X383 | Extr, GC/MS or GC/ECD | Determination of Polychloriated Biphenyls (PCBs) In Soil by GC/ECD. PURPOSE/PRINCIPLE OF METHOD: This document details the requirements and procedures for the analysis of polychlorinated biphenyls (PCBs) in soil by gas chromatography electron capture detector (GC/ECD). A homogenized soil sample is extracted using dichloromethane (DCM). The extract may be cleaned up using column chromatography with 2% Deactivated Florisil. When contamination with sulphur is suspected, further cleanup using copper is required. Analysis is performed using a GC/ECD. Similar methods have been developed for other matrices including waste and transformer oils, water, swabs, florisil tubes, biota. Generally, the sample extraction procedures differ while the GC analysis methods are similar or identical for all matrices. SCOPE: This procedure describes the determination of PCBs in soil and sediments. Other materials such as wood waste may also be analyzed by this method. DETECTION LIMITS: Target Compound Reported Detection Limit (mg/g) Arochlor 1242 0.03, Arochlor 1248 0.03, Arochlor 1254 0.03, Arochlor 1260 0.03, INTERFERENCES: Interferences (e.g. co-eluting peaks on the chromatogram) may be due to sulphur, phthalate esters, or from chlorinated compounds. Suspect samples must be reanalyzed using appropriate cleanup steps. SAMPLE HANDLING & PRESERVATION: Sample Requirements: Whenever possible, samples must completely fill the bottle such that little or no headspace is present to eliminate the loss of volatile hydrocarbon components. (~125g of sample) Sample Containers: Glass jars that have been solvent rinsed with acetone then hexane, then dried. Shipment and Storage: Samples are to be refrigerated (at 4±2°C) during shipping and prior to analysis. Samples suspected to contain high levels of analyte are to be segregated from other samples. Holding Time: Holding time prior to extraction is 14 days. REFERENCE: U.S. EPA Method 8080:Organochlorine Pesticides and PCBs Analysis. | 0.02 | Y | Y | Y | 46 | 3/16/2020 | 9/7/2022 | 1/11/2024 1:07:06 PM | 2020-08-28 |  | RDGIBSON |  | MassPerMass | mg/kg | Y | (mass conc.) |  | Arochlor 1262 (PCB) (mass conc.) | 37324-23-5 | Polychlorinated biphenylds (PCBs); | Polychlorinated biphenylds (PCBs); EMS code: A024 | Chemical |  | LAB
A024 | mg/L | Arochlor 1262 (PCB) | E660 | EPA 8270E (mod) | EPA 8270E (mod) Miscellaneous Pesticides by GC-MS | 2e-05 | Y | Y | Y | 8 | 5/24/2023 | 11/1/2023 | 11/16/2023 3:38:09 PM | 2023-08-16 |  | EMS |  | Concentration | mg/L | Y | (fl. conc.) |  | Arochlor 1262 (PCB) (fl. conc.) | 37324-23-5 | Polychlorinated biphenylds (PCBs); | Polychlorinated biphenylds (PCBs); EMS code: A024 | Chemical |  | LAB
A024 | ug/g wet | Arochlor 1262 (PCB) | FB03 | Frozen,Acetone Extr,GC/ECD | Sediment and homogenized biota samples are Acetone extracted using a wrist-action shaker, partitioned with 2% NaCl, then treated using Florisil, concentrated Sulfuric Acid, activated Copper and Mercury (as required) to remove interferences. The samples are then analysed using Megabore Column Gas Chromatography with Electron Capture Detection. Note that biota results are reported on a "wet weight" basis. | 0.01 | Y | Y | Y | 2 | 4/12/2001 11:51:00 AM | 4/12/2001 11:51:00 AM | 9/11/2002 3:23:57 PM | 1997-05-26 | 6/3/1998 6:51:47 PM | MPAK |  | MassPerMass | µg/g | Y | (mass conc.) | (wet) | Arochlor 1262 (PCB) (wet) (mass conc.) | 37324-23-5 | Polychlorinated biphenylds (PCBs); | Polychlorinated biphenylds (PCBs); EMS code: A024 | Chemical |  | LAB
A024 | ug/g | Arochlor 1262 (PCB) | F115 | Hex dil,Hex extr'n,GC/ECD | Effluent and swab samples are extracted with hexane, the extract concentrated, and then treated using Florisil, concentrated Sulfuric Acid, activated Copper and Mercury (as required) to remove interferences. The samples are then analysed using Megabore Column Gas Chromatography with Electron Capture Detection. Oil samples are first diluted with Hexane then treated in the manner outlined for effluent and swab samples. The samples are then analysed using Megabore Column Gas Chromatography with Electron Capture Detection. | 0.1 | Y | Y | Y | 3 | 4/19/2001 9:04:00 AM | 4/19/2001 9:06:00 AM | 6/1/2001 3:54:29 AM | 2001-05-24 |  | NPEPPIN |  | MassPerMass | mg/kg | Y | (mass conc.) |  | Arochlor 1262 (PCB) (mass conc.) | 37324-23-5 | Polychlorinated biphenylds (PCBs); | Polychlorinated biphenylds (PCBs); EMS code: A024 | Chemical |  | LAB
A024 | mg/L | Arochlor 1262 (PCB) | F018 | Hexane extr,GC/ECD | Effluent and swab samples are extracted with hexane, the extract concentrated, and then treated using Florisil, concentrated Sulfuric Acid, activated Copper and Mercury (as required) to remove interferences. The samples are then analysed using Megabore Column Gas Chromatography with Electron Capture Detection. Oil samples are first diluted with Hexane then treated in the manner outlined for effluent and swab samples. The samples are then analysed using Megabore Column Gas Chromatography with Electron Capture Detection. | 0.0001 | Y | Y | Y | 18 | 2/28/2000 8:40:00 AM | 9/5/2023 6:46:00 PM | 9/11/2023 5:22:04 PM | 1997-05-26 | 6/3/1998 6:51:47 PM | MPAK |  | Concentration | mg/L | Y | (fl. conc.) |  | Arochlor 1262 (PCB) (fl. conc.) | 37324-23-5 | Polychlorinated biphenylds (PCBs); | Polychlorinated biphenylds (PCBs); EMS code: A024 | Chemical |  | LAB
A026 | mg/L | Ametryn | E660 | EPA 8270E (mod) | EPA 8270E (mod) Miscellaneous Pesticides by GC-MS | 0.0001 | Y | Y | Y | 68 | 12/7/2021 | 9/12/2022 | 12/8/2022 12:57:05 PM | 2021-12-09 | 12/9/2021 3:22:12 PM | EMS | EMS | Concentration | mg/L | Y | (fl. conc.) |  | Ametryn (fl. conc.) | 834-12-8 | Pesticides; | Pesticides; EMS code: A026 | Chemical |  | LAB
A026 | ug/g | Ametryn | ZN03 | DCM+Ac.Extr:GC/MS | DCM+Ac.Extr:GC/MS EPA 8270E (mod) | 0.05 | Y | Y | Y | 14 | 1/25/2022 | 9/8/2022 | 12/15/2022 10:50:07 AM | 2003-06-16 |  | UMULLER |  | MassPerMass | mg/kg | Y | (mass conc.) |  | Ametryn (mass conc.) | 834-12-8 | Pesticides; | Pesticides; EMS code: A026 | Chemical |  | LAB
A029 | mg/L | Amyl OH | P050 | MR's | MR's | 0.001 | Y | Y | Y | 5 |  |  | 10/23/1996 8:47:21 PM | 1996-04-12 | 6/3/1998 6:51:33 PM | EMS |  | Concentration | mg/L | Y | (fl. conc.) |  | Amyl OH (fl. conc.) |  | Organics; | Organics; EMS code: A029 | Chemical |  | LAB
A030 | mg/L | Abietic Acid | F009 | pH 9,Diethyl Ether:HRGC/LRMS | Effluent samples are first adjusted to pH 9 and then extracted with Diethyl Ether. The extracts are rotary evaporated and the resulting residues, containing the Resin Acids and other solvent soluble organics are then redissolved in Methanol, transferred to small glass vials, methylated with Diazomethane then evaporated to dryness under Nitrogen. [Methylation Reaction R-COOH(Resin/Fatty Acids) + CH2N2(Diazomethane) = R-COOCH3(Methyl ester) + N2) The residue is re-dissolved in Methyl t-Butyl Ether (MTBE) and injected on a High Resolution Gas Chromatograph/Low Resolution Mass Spectrometer (HRGC/LRMS) system for analysis. LRMS data is acquired in Total Ion Mode (TIM) and quantitation is performed using selected mass ions. Positive identification of the various Resin Acids is made on the basis of a comparison of absolute retention times to those of the external standards and computer based mass spectral mass/intensity library matching routines. | 0.005 | Y | Y | Y | 160 | 12/27/1996 4:15:00 PM | 8/20/2002 | 4/1/2004 1:08:45 AM | 1996-09-16 | 6/3/1998 6:51:44 PM | MPAK |  | Concentration | mg/L | Y | (fl. conc.) |  | Abietic Acid (fl. conc.) |  | Organics; | Organics; EMS code: A030 | Chemical |  | LAB
A030 | mg/L | Abietic Acid | P033 | Resin&Fatty Acids,GC/MS | Resin&Fatty Acids;GC/MS | 0.001 | Y | Y | Y | 34 |  |  | 6/7/2024 5:30:18 PM | 1996-04-12 | 6/3/1998 6:51:45 PM | EMS |  | Concentration | mg/L | Y | (fl. conc.) |  | Abietic Acid (fl. conc.) |  | Organics; | Organics; EMS code: A030 | Chemical |  | LAB
A030 | mg/L | Abietic Acid | X239 | Gas Chromatography | Gas Chromatography | 0.001 | Y | Y | Y | 12 |  |  | 10/23/1996 11:56:36 PM | 1996-04-12 | 6/3/1998 6:51:33 PM | EMS |  | Concentration | mg/L | Y | (fl. conc.) |  | Abietic Acid (fl. conc.) |  | Organics; | Organics; EMS code: A030 | Chemical |  | LAB
A030 | mg/L | Abietic Acid | P030 | Resin Acids | The water sample is acidified and extracted with dichloro-methane. The raw extract is concentrated and derivatized with diazomethane to produce the corresponding methyl ester derivatives. (Suitable substitute derivatives may be produced by other available techniques.) If required, the extracts are cleaned up by Florisil column chromatography. The derivatives are analyzed by gas chromatography with flame ionization detection. | 0.001 | Y | Y | Y | 477 | 8/26/2014 | 8/15/2017 | 8/15/2017 4:37:09 PM | 1996-04-12 | 6/3/1998 6:51:33 PM | EMS |  | Concentration | mg/L | Y | (fl. conc.) |  | Abietic Acid (fl. conc.) |  | Organics; | Organics; EMS code: A030 | Chemical |  | LAB
A030 | mg/L | Abietic Acid | P035 | DCM-DIAZO-FLOR | DCM-DIAZO-FLOR | 0.0005 | Y | Y | Y | 22 |  |  | 10/24/1996 12:44:23 PM | 1996-04-12 | 6/3/1998 6:51:33 PM | EMS |  | Concentration | mg/L | Y | (fl. conc.) |  | Abietic Acid (fl. conc.) |  | Organics; | Organics; EMS code: A030 | Chemical |  | LAB
A030 | mg/L | Abietic Acid | E677 | Resins & Fatty Acids are methylated then analyzed by GCMS | EPA 8270E; AB Env. 129; NCASI RA/FA-85.02 (mod) Resins & Fatty Acids are methylated then analyzed by GCMS. Total fatty acids is the sum of arachidic acid; behenic acid; lauric acid; lignoceric acid; linoleic acid; linolenic acid; myristic acid; oleic acid; palmitic acid; and stearic acid. Total resin acids is the sum of abietic acid; 12-chlorodehydroabietic acid; 14-chlorodehydroabietic acid; dehydroabietic acid; 12;14-dichlorodehydroabietic acid; isopimaric acid; levopimaric acid; neoabietic acid; palustric acid; pimaric acid and sandaracopimaric acid. Total resin acids (BC WWQG) is the sum of abietic acid; neoabietic acid; pimaric acid; isopimaric acid; and sandaracopimaric acid. | 0.001 | Y | Y | Y | 12 | 11/28/2022 | 6/15/2024 | 7/8/2024 11:10:02 AM | 2021-12-30 |  | EMS |  | Concentration | mg/L | Y | (fl. conc.) |  | Abietic Acid (fl. conc.) |  | Organics; | Organics; EMS code: A030 | Chemical |  | LAB
A030 | mg/L | Abietic Acid | RFAS | Separation of resin and fatty acids (RFAs) in water by GC/MS | Scientific Principle The separation of resin and fatty acids (RFAs) in water samples is performed using gas chromatography/mass spectrometry (GC/MS) in the Selected Ion Monitoring (SIM) mode. The separation is based on compound size; boiling point and polarity. Analyte concentration is calculated using the internal standard method. RFA compounds are extracted from water samples using a liquid-liquid solvent extraction with DCM. Extraction occurs after the sample is brought to a pH of ~ 2 and placed in a rotary tumbler (at approx. 40 rpm) for 1 hour. The resulting extract is transferred to a vial; dried with acidified baked sodium sulfate; methylated with diazomethane and concentrated to 1mL under nitrogen prior to analysis. The resulting extracts are vialed and fortified with labeled internal standard compounds prior to analysis by gas chromatography with a (50%-Phenyl)-methylpolysiloxane (DB-17) column and a mass spectrometer detector in SIM Mode. | 0.006 | Y | Y | Y | 1 | 9/17/2021 7:49:00 PM | 9/17/2021 7:49:00 PM | 2/1/2022 2:38:14 PM | 2022-01-19 |  | EMS |  | Concentration | mg/L | Y | (fl. conc.) |  | Abietic Acid (fl. conc.) |  | Organics; | Organics; EMS code: A030 | Chemical |  | LAB
A030 | ug/g | Abietic Acid | FS05 | MeOH,DCM extr:HRGC/LRMS | Sediment samples are extracted by shaking with Methanol in a screw cap centrifuge bottle. The extracts are rotary evaporated and the resulting residues, containing the Resin Acids and other solvent soluble organics, are then redissolved in Methanol, transferred to small glass vials, methylated with Diazomethane then evaporated to dryness under Nitrogen. [Methylation Reaction R-COOH(Resin/Fatty Acids) + CH2N2 (Diazomethane) = R-COOCH3 (Methyl ester) + N2] The residue is re-dissolved in Methyl t-Butyl Ether (MTBE) and injected on a High Resolution Gas Chromatograph / Low Resolution Mass Spectrometer (HRGC/LRMS) system for analysis. LRMS data is acquired in Total Ion Mode (TIM) and quantitation is performed using selected mass ions. Positive identification of the various Resin Acids is made on the basis of a comparison of absolute retention times to those of the external standards and computer based mass spectral mass/intensity library matching routines. | 0.2 | Y | Y | Y | 1 | 9/14/2001 3:29:00 PM | 9/14/2001 3:29:00 PM | 2/17/2003 3:39:54 PM | 1996-09-16 | 6/3/1998 6:51:44 PM | MPAK |  | MassPerMass | mg/kg | Y | (mass conc.) |  | Abietic Acid (mass conc.) |  | Organics; | Organics; EMS code: A030 | Chemical |  | LAB
A031 | ug/g | 8(14)-Abietenic Acid | FS05 | MeOH,DCM extr:HRGC/LRMS | Sediment samples are extracted by shaking with Methanol in a screw cap centrifuge bottle. The extracts are rotary evaporated and the resulting residues, containing the Resin Acids and other solvent soluble organics, are then redissolved in Methanol, transferred to small glass vials, methylated with Diazomethane then evaporated to dryness under Nitrogen. [Methylation Reaction R-COOH(Resin/Fatty Acids) + CH2N2 (Diazomethane) = R-COOCH3 (Methyl ester) + N2] The residue is re-dissolved in Methyl t-Butyl Ether (MTBE) and injected on a High Resolution Gas Chromatograph / Low Resolution Mass Spectrometer (HRGC/LRMS) system for analysis. LRMS data is acquired in Total Ion Mode (TIM) and quantitation is performed using selected mass ions. Positive identification of the various Resin Acids is made on the basis of a comparison of absolute retention times to those of the external standards and computer based mass spectral mass/intensity library matching routines. | 0.2 | Y | Y | Y | 1 | 9/14/2001 3:29:00 PM | 9/14/2001 3:29:00 PM | 2/17/2003 3:39:55 PM | 1997-01-13 | 6/3/1998 6:51:46 PM | MPAK |  | MassPerMass | mg/kg | Y | (mass conc.) |  | 8(14)-Abietenic Acid (mass conc.) | 19407-37-5 | Organics; | Organics; EMS code: A031 | Chemical |  | LAB
A031 | mg/L | 8(14)-Abietenic Acid | F009 | pH 9,Diethyl Ether:HRGC/LRMS | Effluent samples are first adjusted to pH 9 and then extracted with Diethyl Ether. The extracts are rotary evaporated and the resulting residues, containing the Resin Acids and other solvent soluble organics are then redissolved in Methanol, transferred to small glass vials, methylated with Diazomethane then evaporated to dryness under Nitrogen. [Methylation Reaction R-COOH(Resin/Fatty Acids) + CH2N2(Diazomethane) = R-COOCH3(Methyl ester) + N2) The residue is re-dissolved in Methyl t-Butyl Ether (MTBE) and injected on a High Resolution Gas Chromatograph/Low Resolution Mass Spectrometer (HRGC/LRMS) system for analysis. LRMS data is acquired in Total Ion Mode (TIM) and quantitation is performed using selected mass ions. Positive identification of the various Resin Acids is made on the basis of a comparison of absolute retention times to those of the external standards and computer based mass spectral mass/intensity library matching routines. | 0.005 | Y | Y | Y | 131 | 12/27/1996 4:16:00 PM | 1/8/2002 12:49:00 PM | 5/28/2003 11:38:54 AM | 1997-01-13 | 6/3/1998 6:51:46 PM | MPAK |  | Concentration | mg/L | Y | (fl. conc.) |  | 8(14)-Abietenic Acid (fl. conc.) | 19407-37-5 | Organics; | Organics; EMS code: A031 | Chemical |  | LAB
A032 | mg/L | Anilazine | F056 | DCM extr,HRGC-NPD | Effluent samples are extracted with Dichloromethane. The extract is concentrated and analysed using High Resolution Gas Chromatography with a Thermionic Specific Detector (TSD), alternately referred to as a Nitrogen Phosphorus Detector (NPD). | 0.005 | Y | Y | Y | 12 | 4/3/2000 10:51:00 AM | 2/16/2001 3:36:00 PM | 9/11/2002 9:34:56 AM | 1998-10-02 | 10/2/1998 10:50:33 AM | NPEPPIN | NPEPPIN | Concentration | mg/L | Y | (fl. conc.) |  | Anilazine (fl. conc.) |  | Pesticides; | Pesticides; EMS code: A032 | Chemical |  | LAB
AC-F | ueq/L | Acidity:Free | 5063 | Calculation-Electrometer | Calculation-Electrometer | 0.1 | Y | Y | Y | 6399 | 4/30/1997 3:50:00 PM | 4/10/2008 | 5/28/2008 2:08:44 AM | 1996-04-12 | 6/3/1998 6:51:33 PM | EMS |  | Equivalents | µEq/L | Y | (equiv./vol.) |  | Acidity:Free (equiv./vol.) |  | Major Ions; | Major Ions; EMS code: AC-F | Chemical |  | LAB
AC-F | ueq/L | Acidity:Free | EQ18 | EQUIS (ueq/L) | EQUIS (ueq/L) | 0.1 | N | Y | Y | 126 |  |  | 10/22/1996 11:12:48 PM | 1996-04-12 | 6/3/1998 6:51:33 PM | EMS |  | Equivalents | µEq/L | Y | (equiv./vol.) |  | Acidity:Free (equiv./vol.) |  | Major Ions; | Major Ions; EMS code: AC-F | Chemical |  | LAB
AC-T | ueq/L | Acidity:Total | 5062 | Electrom/Grans Plot | A glass electrode, calibrated with a pH 4.10 H2SO4 buffer, is used with a digital pH/mV meter to measure sample pH. Acidity is then determined on the same aliquot by titration with uL incrments of 0.01N NaOH. The volume of NaOH added to the sample is plotted against the Gran's function, calculated from the readings obtained during the titration. Equivalence point for strong acidity is obtained by extrapolating the Gran's functions to the volume axis. Total alkalinity is determined in a similar manner using 0.01N H2SO4. | 15 | Y | Y | Y | 464 |  |  | 10/24/1996 5:20:14 AM | 1996-04-12 | 6/3/1998 6:51:33 PM | EMS |  | Equivalents | µEq/L | Y | (equiv./vol.) |  | Acidity:Total (equiv./vol.) |  | Major Ions; | Major Ions; EMS code: AC-T | Chemical | Total | LAB
AC-T | ueq/L | Acidity:Total | EQ18 | EQUIS (ueq/L) | EQUIS (ueq/L) | 15 | N | Y | Y | 118 |  |  | 10/22/1996 11:12:48 PM | 1996-04-12 | 6/3/1998 6:51:33 PM | EMS |  | Equivalents | µEq/L | Y | (equiv./vol.) |  | Acidity:Total (equiv./vol.) |  | Major Ions; | Major Ions; EMS code: AC-T | Chemical | Total | LAB
AC83 | ueq/L | Acidity to pH 8.3 | 5062 | Electrom/Grans Plot | A glass electrode, calibrated with a pH 4.10 H2SO4 buffer, is used with a digital pH/mV meter to measure sample pH. Acidity is then determined on the same aliquot by titration with uL incrments of 0.01N NaOH. The volume of NaOH added to the sample is plotted against the Gran's function, calculated from the readings obtained during the titration. Equivalence point for strong acidity is obtained by extrapolating the Gran's functions to the volume axis. Total alkalinity is determined in a similar manner using 0.01N H2SO4. | 0.1 | Y | Y | Y | 6052 | 4/30/1997 3:50:00 PM | 4/10/2008 | 5/28/2008 2:08:44 AM | 1996-04-12 | 6/3/1998 6:51:33 PM | EMS |  | Equivalents | µEq/L | Y | (equiv./vol.) |  | Acidity to pH 8.3 (equiv./vol.) |  | Major Ions; | Major Ions; EMS code: AC83 | Chemical |  | LAB
ACD9 | % | Acridine d9 | X262 | CH2Cl2 Ext:Silica:GC/MS | The sample is placed in a separatory funnel, the pH of the sample is adjusted to alkaline and then acidic conditions, and extracted each time with dichloromethane to selectively remove compounds. The final extracts (basic and acidic) can be combined or run separately. The extract is cleaned by silica gel column chromatography, then analyzed by GC/MS. | 1 | Y | Y | Y | 1 | 8/28/2014 | 8/28/2014 | 3/11/2015 3:57:18 PM | 2015-03-10 |  | RDGIBSON |  | DimensionlessRatio | % | N | (ratio) |  | Acridine d9 (ratio) | 260-94-6 | Organics; | Organics; EMS code: ACD9 | Chemical |  | SURROGATE_RESULT
ACD9 | % | Acridine d9 | SURR | Surrogate Recovery | Surrogate Recovery | 1 | Y | Y | Y | 705 | 1/13/2015 | 9/9/2022 | 1/11/2024 1:07:06 PM | 2015-07-07 |  | RDGIBSON |  | DimensionlessRatio | % | N | (ratio) |  | Acridine d9 (ratio) | 260-94-6 | Organics; | Organics; EMS code: ACD9 | Chemical |  | SURROGATE_RESULT
ACEN | % | Acenaphthene d10 | FS87 | DCM,Ac/Hex extr,PAM,HRGC/LRMS | DCM;Ac/Hex extr;PAM;HRGC/LRMS | 1 | Y | Y | Y | 43 | 5/20/2015 | 5/23/2015 | 7/3/2015 3:19:03 PM | 2015-05-27 | 5/27/2015 3:49:01 PM | RDGIBSON | RDGIBSON | DimensionlessRatio | % | N | (ratio) |  | Acenaphthene d10 (ratio) | 15067-26-2 | Polycyclic Aromatic Hydrocarbons (PAHs); | Polycyclic Aromatic Hydrocarbons (PAHs); EMS code: ACEN | Chemical |  | SURROGATE_RESULT
ACEN | mg/L | Acenaphthene d10 | MW02 | DCM extrn,GC/MS-SIM (PAH) | Water samples are liquid - liquid extracted with dichloromethane. The extract is then cleaned by silica gel column chromatography and analyzed by gas chromatography - mass spectroscopy. | 1e-05 | Y | Y | Y | 71 | 2/5/2002 | 8/16/2004 | 9/21/2004 1:03:53 AM | 1996-04-12 | 6/3/1998 6:51:41 PM | EMS |  | Concentration | mg/L | Y | (fl. conc.) |  | Acenaphthene d10 (fl. conc.) | 15067-26-2 | Polycyclic Aromatic Hydrocarbons (PAHs); | Polycyclic Aromatic Hydrocarbons (PAHs); EMS code: ACEN | Chemical |  | SURROGATE_RESULT
ACEN | % | Acenaphthene d10 | X262 | CH2Cl2 Ext:Silica:GC/MS | The sample is placed in a separatory funnel, the pH of the sample is adjusted to alkaline and then acidic conditions, and extracted each time with dichloromethane to selectively remove compounds. The final extracts (basic and acidic) can be combined or run separately. The extract is cleaned by silica gel column chromatography, then analyzed by GC/MS. | 1 | Y | Y | Y | 1 | 8/28/2014 | 8/28/2014 | 3/11/2015 3:57:18 PM | 2015-03-09 |  | RDGIBSON |  | DimensionlessRatio | % | N | (ratio) |  | Acenaphthene d10 (ratio) | 15067-26-2 | Polycyclic Aromatic Hydrocarbons (PAHs); | Polycyclic Aromatic Hydrocarbons (PAHs); EMS code: ACEN | Chemical |  | SURROGATE_RESULT
ACEN | % | Acenaphthene d10 | SURR | Surrogate Recovery | Surrogate Recovery | 30 | Y | Y | Y | 537 | 8/18/2003 | 6/5/2019 | 6/5/2019 5:06:05 PM | 1996-04-12 | 6/3/1998 6:51:33 PM | EMS |  | DimensionlessRatio | % | N | (ratio) |  | Acenaphthene d10 (ratio) | 15067-26-2 | Polycyclic Aromatic Hydrocarbons (PAHs); | Polycyclic Aromatic Hydrocarbons (PAHs); EMS code: ACEN | Chemical |  | SURROGATE_RESULT
ACEN | % | Acenaphthene d10 | PAH1 | PUF:GC/MS | To understand the extent of human exposure to B[a]P and other PAHs, a reliable sampling and analytical method has been established which involves a combination quartz fibre filter/absorbent cartridge with subsequent extraction and analysis by gas chromatography (GC) with mass spectrometry (MS) detection (GC/MS). The analytical methods are a modification of EPA Test Method 625, Methods for Organic Chemical Analysis of Municipal and Industrial Wastewater, and Methods 8270, Test Methods for Evaluation of Solid Waste. PAH1 utilizes PUF as an adsorbant; PAH2 utilizes XAD-2 as an adsorbant. Note that isotopic surrogates may be added and are reported as % recovery. | 0 | Y | Y | Y | 100 | 10/14/1997 2:50:00 PM | 4/22/1998 5:40:00 PM | 1/19/1999 6:45:09 PM | 1996-04-12 | 6/3/1998 6:51:33 PM | EMS |  | DimensionlessRatio | % | N | (ratio) |  | Acenaphthene d10 (ratio) | 15067-26-2 | Polycyclic Aromatic Hydrocarbons (PAHs); | Polycyclic Aromatic Hydrocarbons (PAHs); EMS code: ACEN | Chemical |  | SURROGATE_RESULT
ACEN | ug/g | Acenaphthene d10 | MP02 | GC/MS | GC/MS | 0.001 | Y | Y | Y | 101 | 3/28/2002 | 3/28/2011 | 5/9/2011 5:20:45 PM | 1996-04-12 | 6/3/1998 6:51:33 PM | EMS |  | MassPerMass | mg/kg | Y | (mass conc.) |  | Acenaphthene d10 (mass conc.) | 15067-26-2 | Polycyclic Aromatic Hydrocarbons (PAHs); | Polycyclic Aromatic Hydrocarbons (PAHs); EMS code: ACEN | Chemical |  | SURROGATE_RESULT
ACT | mg/L | Acetaminophen CAS: 103-90-2 | XMSP | HPLC and Tri. Quadpole MS/MS | EPA 1694 An aliquot of filtered sample containing 5% organic is injected directly. Each compound is separated by reversed phase HPLC. Analyzed by Electro Spray Ionization MS/MS detection using Triple Quadrupole MS/MS detector. Quantiitation with internal or external standard. |  | Y | Y | Y | 12 | 5/8/2020 | 6/4/2024 | 8/21/2024 3:32:03 PM | 2022-05-03 |  | EMS |  | Concentration | mg/L | Y | (fl. conc.) |  | Acetaminophen (fl. conc.) | 103-90-2 | Pharmaceuticals and Personal Care Products (PPCP); | Pharmaceuticals and Personal Care Products (PPCP); EMS code: ACT | Chemical |  | LAB
ACTA | mg/L | Acetic Acid | VFAD | Volatile Fatty Acids by GC/FID | Volatile Fatty Acids by GC/FID; Standard Methods for the Examination of Water and Wastewater; 21st Edition, Method 5560 D; pages 5-59 to 5-61. 5560 D. Gas Chromatographic Method Principle: This gas chromatographic procedure may be usedto determine the individual concentrations of the following fattyacids: acetic; propionic; butyric; isobutyric; valeric; andisovaleric acids in digester sludge samples. Caproic and heptanoicacids also may be determined; but with less certaintybecause of their reduced solubility in water. Volatile fatty acidsare generally soluble in digester sludge and exist in either theprotonated or unprotonated form; depending on the pH of thesludge sample. Acetic and propionic acids usually are in higherconcentrations than the other fatty acids in digester sludge. Allacids must be converted to the protonated (volatile) form beforeinjection into the gas chromatograph (GC) so vaporization mayoccur. The fatty acids must also be separated from the solidsmaterial in the sample before introduction into the GC to minimizedegradation of the GC column. Digester sludge samples areprepared by acidification; centrifugation; and filtration. The sampleis analyzed by direct injection into a gas chromatographequipped with a flame ionization detector after both centrifugationand filtration.b. Interference: The blank amount may be high relative tosample analyte. High blanks may be caused by a buildup ofcontaminants in the injector and guard column or by samplecarry-over. The analyst must be aware of the blank levels andcorrect unacceptably high blank levels that are above the reportinglimits.c. Minimum detectable concentration: The method detectionlevel has been determined by the USEPA method1 in reagentwater to be about 3 mg/L for acetic acid and about 1 mg/L for allother target compounds. | 3 | Y | Y | Y | 5 | 3/14/2017 10:30:00 AM | 11/19/2017 5:10:00 PM | 6/20/2018 9:02:11 AM | 2018-06-19 |  | RDGIBSON |  | Concentration | mg/L | Y | (fl. conc.) |  | Acetic Acid (fl. conc.) |  | Organics; | Organics; EMS code: ACTA | Chemical |  | LAB
AD | ug/g | Aldrin + Dieldrin CAS: 309-00-2 and CAS: 60-57-1 | ZN03 | DCM+Ac.Extr:GC/MS | DCM+Ac.Extr:GC/MS EPA 8270E (mod) | 0.03 | Y | Y | Y | 2 | 11/14/2022 | 11/14/2022 | 12/15/2022 10:50:04 AM | 2022-06-21 |  | EMS |  | MassPerMass | mg/kg | Y | (mass conc.) |  | Aldrin + Dieldrin (mass conc.) | 309-00-2, 60-57-1 | Pesticides; | Pesticides; EMS code: AD | Chemical |  | LAB
AD | mg/L | Aldrin + Dieldrin CAS: 309-00-2 and CAS: 60-57-1 | E660 | EPA 8270E (mod) | EPA 8270E (mod) Miscellaneous Pesticides by GC-MS | 3e-06 | Y | Y | Y | 2 | 9/14/2022 | 9/14/2022 | 12/8/2022 9:48:13 AM | 2022-12-06 |  | EMS |  | Concentration | mg/L | Y | (fl. conc.) |  | Aldrin + Dieldrin (fl. conc.) | 309-00-2, 60-57-1 | Pesticides; | Pesticides; EMS code: AD | Chemical |  | LAB
AD | mg/L | Aldrin + Dieldrin CAS: 309-00-2 and CAS: 60-57-1 | X268 | CH2Cl2 ext:GC/MS | CH2Cl2 ext:GC/MS |  | Y | Y | Y | 16 | 12/29/2021 | 9/2/2022 | 11/10/2022 9:54:06 AM | 2022-03-21 |  | EMS |  | Concentration | mg/L | Y | (fl. conc.) |  | Aldrin + Dieldrin (fl. conc.) | 309-00-2, 60-57-1 | Pesticides; | Pesticides; EMS code: AD | Chemical |  | LAB
ADDP | mg/L | Atrazine-desthyl-desisopropyl | E751 | E751D Pesticides in Water by LC-MS/MS (POS1 List) | E751D Pesticides in Water by LC-MS/MS (POS1 List) (MOECC E3501) | 0.0005 | Y | Y | Y | 39 | 12/9/2021 | 12/31/2021 | 2/15/2022 9:50:04 AM | 2021-12-14 |  | EMS |  | Concentration | mg/L | Y | (fl. conc.) |  | Atrazine-desthyl-desisopropyl (fl. conc.) | 3397-62-4 | Pesticides; | Pesticides; EMS code: ADDP | Chemical |  | LAB
ADONA | mg/L | Perfluorononanoic acid, 4,8-dioxa-3H- (ADONA)CAS:919005-14-4 | E3533 | MECP E3533 | PFAs in water by direct injection LC/MS/MS | 2e-05 | Y | Y | Y | 2 | 6/3/2024 | 6/3/2024 | 8/21/2024 3:32:03 PM | 2023-11-15 | 6/27/2024 5:06:31 PM | EMS | EMS | Concentration | mg/L | Y | (fl. conc.) |  | Perfluorononanoic acid, 4,8-dioxa-3H- (ADONA) (fl. conc.) | 919005-14-4 | Polychlorinated biphenylds (PCBs); | Polychlorinated biphenylds (PCBs); EMS code: ADONA | Chemical |  | LAB
AEHS | None | Acid Extrctble Herbs | SCC1 | Acid Extrctble Herb Scan | Acid Extrctble Herb Scan | 0 | N | Y | Y | 4 |  |  | 10/24/1996 12:45:23 PM | 1996-04-12 | 6/3/1998 6:51:33 PM | EMS |  | None | None | N | (unitless) |  | Acid Extrctble Herbs (unitless) |  |  | EMS code: AEHS | Chemical |  | LAB
AG-D | mg/L | Silver Dissolved | F023 | HNO3,GFAA | Dissolved and extractable water samples are ready for analysis once collected. Total samples are autoclaved with HNO3 and hydrochloric acid (HCl) in capped Teflon vials. Analysis is performed via Inductively Coupled Argon Plasma Atomic Emission Spectrometry (ICP-AES) (simultaneous multielement analysis), Graphite Furnace Atomic Absorption Spectrometry (GF-AAS) (low level analysis) and/or Flame Atomic Absorption Spectrometry (F-AAS). Note that extractable samples are acidified with 0.5 mL HNO3 per 100 mL of sample at time of sample collection. | 0.0005 | Y | Y | Y | 18 | 10/18/1996 6:11:00 PM | 6/1/2000 8:33:00 AM | 10/18/2000 1:03:49 AM | 1997-04-02 | 6/3/1998 6:51:46 PM | MPAK |  | Concentration | mg/L | Y | (fl. conc.) |  | Silver Dissolved (fl. conc.) | 7440-22-4 | Metals - Dissolved; | Metals - Dissolved; EMS code: AG-D | Chemical | Dissolved | LAB
AG-D | mg/L | Silver Dissolved | F068 | HNO3dig,TeflonTube/Autoclave,ICP-MS | This method is applicable to determiniation of dissolved, extractable, and total recoverable metals in ground waters, surface waters, drinking waters, and wastewaters. Dissolved, extractable, and total recoverable sample fractions are acidified to pH <1.5 with analytical grade nitric acid (HNO3) to prevent the metals from adsorbing to the container material or precipitating. Extractable and total recoverable water samples are collected in 250 ml acid washed polyethylene bottles and are preserved with Nitric acid (0.5 ml acid for 100 mLs of sample). All preserved samples are stored at room temperature and can be held for up to 6 months. Unfiltered, un-preserved samples are stored at 4C for no more than 30 days. For dissolved and extractable water fractions, no sample preparation is necessary. Samples are simply poured in autosampler tubes (i.e. 15 ml centrifuge tubes), and placed on the autosampler and analyzed. Samples for total recoverable metals are digested with Nitric acid in special Teflon tubes in an autoclave prior to analyses. | 2e-05 | Y | Y | Y | 1513 | 6/1/1999 1:19:00 PM | 4/17/2015 1:06:00 PM | 8/11/2021 12:04:11 PM | 1999-06-28 |  | NPEPPIN |  | Concentration | mg/L | Y | (fl. conc.) |  | Silver Dissolved (fl. conc.) | 7440-22-4 | Metals - Dissolved; | Metals - Dissolved; EMS code: AG-D | Chemical | Dissolved | LAB
AG-D | mg/L | Silver Dissolved | 5138 | AA Analysis | AA Analysis | 0.005 | Y | Y | Y | 70 |  |  | 10/24/1996 6:09:27 AM | 1996-04-12 | 6/3/1998 6:51:33 PM | EMS |  | Concentration | mg/L | Y | (fl. conc.) |  | Silver Dissolved (fl. conc.) | 7440-22-4 | Metals - Dissolved; | Metals - Dissolved; EMS code: AG-D | Chemical | Dissolved | LAB
AG-D | mg/L | Silver Dissolved | METC | MET-D-NP-CCMS-VA | Ultra trace metals in water are analyzed by CRC ICPMS, based on US EPA Method 6020B (July 2014). This procedure is intended for pristine field-filtered acid-preserved water samples. The detection limits (LOR) for this test are based on lab instrumental analysis only, not including filtration. ALS-supplied field filtration equipment does not support these LOR. Therefore, because of the high probability of false positives due to filtration, it is strongly recommended that a filtration blank be analysed to aid in data interpretation. | 5e-06 | Y | Y | Y | 1776 | 4/3/2012 | 8/15/2024 | 8/23/2024 9:50:35 AM | 2019-08-20 |  | KATFRASE |  | Concentration | mg/L | Y | (fl. conc.) |  | Silver Dissolved (fl. conc.) | 7440-22-4 | Metals - Dissolved; | Metals - Dissolved; EMS code: AG-D | Chemical | Dissolved | LAB
AG-D | None | Silver Dissolved | AISP | Special Analysis | Special Analysis | 0 | N | Y | N | 3 |  |  | 10/24/1996 11:57:22 AM | 1996-04-12 | 6/3/1998 6:51:33 PM | EMS |  | None | None | N | (unitless) |  | Silver Dissolved (unitless) | 7440-22-4 | Metals - Dissolved; | Metals - Dissolved; EMS code: AG-D | Chemical | Dissolved | LAB
AG-D | mg/L | Silver Dissolved | X357 | LF,LA:HNO3:GFAA | Samples are lab-filtered (0.45 um) and lab-acidified with HNO3; analysis by GFAA. | 0.0002 | Y | Y | Y | 8 |  |  | 10/23/1996 9:10:03 PM | 1996-04-12 | 6/3/1998 6:51:41 PM | EMS |  | Concentration | mg/L | Y | (fl. conc.) |  | Silver Dissolved (fl. conc.) | 7440-22-4 | Metals - Dissolved; | Metals - Dissolved; EMS code: AG-D | Chemical | Dissolved | LAB
AG-D | mg/L | Silver Dissolved | X085 | LF,LA,HNO3,AA | Samples are lab-filtered (0.45 um), and lab-acidified (HNO3); analysis by AA. | 0.005 | Y | Y | Y | 37 | 6/12/2000 | 12/1/2010 | 12/9/2010 3:55:26 PM | 1996-04-12 | 6/3/1998 6:51:41 PM | EMS |  | Concentration | mg/L | Y | (fl. conc.) |  | Silver Dissolved (fl. conc.) | 7440-22-4 | Metals - Dissolved; | Metals - Dissolved; EMS code: AG-D | Chemical | Dissolved | LAB
AG-D | mg/L | Silver Dissolved | F049 | Filt,HNO3 dig,conc'n,nebul'ICP-OES | This method provides an automated procedure for the Inductively Coupled Argon Plasma Optical Emission Spectrophotometric (ICP-OES) determination and quantitation of trace amounts of metals in surface, ground and precipitation waters after a manual acid digestion and concentration procedure and ultrasonic nebulisation. Organically-bound metals and metal complexes are converted to metal ions by digestion with nitric acid. To increase the signal detected by the ICP, the samples are concentrated 10 times. The liquid sample is pumped onto the piezoelectric transducer, where the ultrasonic action efficiently converts it to a fine, dense aerosol. The nebulizer carrier gas flow then transports the wet aerosol through a heated U-tube, where the solvent is vaporized. Subsequently, most of the solvent vapour is removed from the aerosol stream by passing the stream through a chilled condenser. The result is a dry, analyte-laden aerosol which is introduced to the plasma. There, the high temperature desolvates the aerosol and dissociates the sample material into individual atoms and ions, which are excited to emit light at wave lengths characteristic of the elements present. Dissolved samples are field filtered through 0.45 micron cellulose acetate filter prior to acid preservation. Samples for extractable analysis are acidified with 2 mL concentrated HNO3 per litre and allowed to stand a minimum of 24 hours before analysis. Clients providing these samples are advised to preserve samples at the collection site. | 0.01 | Y | Y | Y | 1421 |  |  | 5/19/2015 3:10:33 PM | 1998-11-17 |  | NPEPPIN |  | Concentration | mg/L | Y | (fl. conc.) |  | Silver Dissolved (fl. conc.) | 7440-22-4 | Metals - Dissolved; | Metals - Dissolved; EMS code: AG-D | Chemical | Dissolved | LAB
AG-D | mg/L | Silver Dissolved | MWID | Metals in Water by ICPMS (digested) | Metals in Water by ICPMS (digested) Based on EPA 200.8 | 0.0001 | Y | Y | Y | 31768 | 10/26/2018 5:20:00 PM | 10/26/2018 5:20:00 PM | 7/31/2024 1:46:41 PM | 2009-12-24 |  | RDGIBSON |  | Concentration | mg/L | Y | (fl. conc.) |  | Silver Dissolved (fl. conc.) | 7440-22-4 | Metals - Dissolved; | Metals - Dissolved; EMS code: AG-D | Chemical | Dissolved | LAB
AG-D | mg/L | Silver Dissolved | X285 | HNO3 Dig:ICP Analysis | HNO3 Dig:ICP Analysis | 0.015 | Y | Y | Y | 673 | 1/30/2002 | 11/13/2002 | 3/18/2004 10:28:33 AM | 1998-05-22 | 6/3/1998 6:51:49 PM | NPEPPIN |  | Concentration | mg/L | Y | (fl. conc.) |  | Silver Dissolved (fl. conc.) | 7440-22-4 | Metals - Dissolved; | Metals - Dissolved; EMS code: AG-D | Chemical | Dissolved | LAB
AG-D | mg/L | Silver Dissolved | EQ01 | EQUIS (mg/L) | EQUIS (mg/L) | 0 | N | Y | Y | 7 |  |  | 10/22/1996 7:21:47 PM | 1996-04-12 | 6/3/1998 6:51:33 PM | EMS |  | Concentration | mg/L | Y | (fl. conc.) |  | Silver Dissolved (fl. conc.) | 7440-22-4 | Metals - Dissolved; | Metals - Dissolved; EMS code: AG-D | Chemical | Dissolved | LAB
AG-D | mg/L | Silver Dissolved | X170 | ICP | This code is a catch-all for a large variety of ICP methods including water and also solids digested into solutions as an intermediate result. Different procedures will have greatly different detection limits. (a) For air particulates, preliminary results are reported in mg/L and loading results as mg/dm2/d. (b) For Hi Vol samples preliminiary results are same as reported in mg/L and loading results as ug/m3. (c) Note that for Rain samples there is no filtration nor digestion steps. | 0.01 | Y | Y | Y | 7144 | 4/11/1996 6:01:00 PM | 11/6/2018 | 6/12/2020 1:21:11 PM | 1996-04-12 | 6/3/1998 6:51:33 PM | EMS |  | Concentration | mg/L | Y | (fl. conc.) |  | Silver Dissolved (fl. conc.) | 7440-22-4 | Metals - Dissolved; | Metals - Dissolved; EMS code: AG-D | Chemical | Dissolved | LAB
AG-D | mg/L | Silver Dissolved | F032 | Filt.,Acid,ICPMS | There is no acid digestion required for this analysis. A water sample is collected and preserved in 0.2 percent nitric acid. To analyse the extractable metals the sample is shaken, allowed to sit for at least twenty four hours, and a decanted portion is introduced to the spectrometer. When the dissolved metals are required, the samples must first be filtered through a 0.45 micron cellulose acetate filter paper and then followed by preservation in 0.2 percent nitric acid. It is important to note that different laboratories may have very different detection limits for some variables analyzed by ICPMS. | 2e-06 | Y | Y | Y | 74657 | 3/10/1998 2:07:00 PM | 8/26/2024 | 8/26/2024 4:53:05 PM | 1997-10-20 | 8/15/2019 10:46:55 AM | NPEPPIN | KATFRASE | Concentration | mg/L | Y | (fl. conc.) |  | Silver Dissolved (fl. conc.) | 7440-22-4 | Metals - Dissolved; | Metals - Dissolved; EMS code: AG-D | Chemical | Dissolved | LAB
AG-D | mg/L | Silver Dissolved | F039 | Field filt + acid, ICP-AES | Dissolved and extractable water samples are ready for analysis once collected. Total samples are autoclaved with HNO3 and hydrochloric acid (HCI) in capped Teflon vials. Analysis is performed via Inductively Coupled Argon Plasma Atomic Emission Spectrometry (ICP-AES)(simultaneous multielement analysis), Graphite Furnace Atomic Absorption Spectrometry (GF-AAS)(low level analysis) and/or Flame Atomic Absorption Spectrometry (F-AAS). | 0.01 | Y | Y | Y | 4813 | 7/29/1996 6:03:00 PM | 7/5/2019 1:21:00 PM | 9/11/2019 8:47:46 AM | 1997-05-08 | 6/3/1998 6:51:46 PM | MPAK |  | Concentration | mg/L | Y | (fl. conc.) |  | Silver Dissolved (fl. conc.) | 7440-22-4 | Metals - Dissolved; | Metals - Dissolved; EMS code: AG-D | Chemical | Dissolved | LAB
AG-D | mg/L | Silver Dissolved | 0050 | Graphite Furnace AA | Graphite Furnace AA | 0.0005 | Y | Y | Y | 118 | 1/3/2003 | 1/3/2003 | 1/13/2003 3:03:12 PM | 1996-04-12 | 6/3/1998 6:51:33 PM | EMS |  | Concentration | mg/L | Y | (fl. conc.) |  | Silver Dissolved (fl. conc.) | 7440-22-4 | Metals - Dissolved; | Metals - Dissolved; EMS code: AG-D | Chemical | Dissolved | LAB
AG-D | mg/L | Silver Dissolved | X544 | FF,FA,HNO3,ICPMS | Analytes nebulized followed by ion entrained in plasma gas and extracted by means of a differentially pumped vacuum interface into a mass spectrometer | 0.0001 | Y | Y | Y | 30 |  |  | 10/30/2014 8:31:16 AM | 2001-09-18 | 9/18/2001 1:36:23 PM | NPEPPIN | NPEPPIN | Concentration | mg/L | Y | (fl. conc.) |  | Silver Dissolved (fl. conc.) | 7440-22-4 | Metals - Dissolved; | Metals - Dissolved; EMS code: AG-D | Chemical | Dissolved | LAB
AG-D | mg/L | Silver Dissolved | MSEA | Metals in Seawater (undigested) by (CRC) ICP-MS | This method describes the multi-elemental determination of trace elements by Collision/Reaction Cell (CRC) ICP-MS. The method measures ions produced by a radio-frequency inductively coupled plasma. Analyte species originating in a liquid are pneumatically nebulized and the resulting aerosol transported by argon gas into the plasma torch. The ions produced are entrained in the plasma gas and extracted, by means of a differentially pumped vacuum interface, into a mass spectrometer. The ions flow through a Collision/Reaction Cell (CRC) where depending on the mode of analysis either the ions passes through the cell unhindered (no gas - classical ICP-MS mode), or collisions/reactions chemistry takes place between the ions and a neutral gas such as hydrogen or helium. Any polyatomic ions which give interference on the analytes are removed. The ions are then sorted according to their mass-to-charge ratios by a quadrupole mass spectrometer. The ions are then sorted according to their mass-to-charge ratios by a quadrupole mass spectrometer. The ions transmitted through the quadrupole are quantified by a channel electron multiplier and the ion information processed by a data handling system. Although most polyatomic interference is removed, any additional interference must be assessed and valid corrections applied or the data flagged to indicate problems. Interference correction must include compensation for background ions contributed by the plasma gas, reagents, and constituents of the sample matrix. Instrument drift as well as suppressions or enhancements of instrument response caused by the sample matrix must be corrected for by the use of internal standards. | 2e-05 | Y | Y | Y | 4 | 4/12/2023 | 4/29/2024 | 5/16/2024 2:00:14 PM | 2009-09-08 |  | RDGIBSON |  | Concentration | mg/L | Y | (fl. conc.) |  | Silver Dissolved (fl. conc.) | 7440-22-4 | Metals - Dissolved; | Metals - Dissolved; EMS code: AG-D | Chemical | Dissolved | LAB
AG-D | mg/L | Silver Dissolved | X562 | Dissolved metals: FF, FA, HNO3, HR-ICPMS | Ultra trace metals in water are analyzed by high resolution inductively coupled plasma mass spectrometry (HR-ICPMS), based on US EPA Method 200.8, (Oct. 1999). This procedure is intended for pristine field-filtered acid-preserved water samples. The detection limits (LOR) for this test are based on lab instrumental analysis only, not including filtration. ALS-supplied field filtration equipment does not support these LOR. Therefore, because of the high probability of false positives due to filtration, it is strongly recommended that a filtration blank be analyzed to aid in data interpretation. | 5e-06 | Y | Y | Y | 1264 | 6/8/2015 | 10/19/2023 | 11/9/2023 7:41:27 AM | 2015-10-22 |  | RDGIBSON |  | Concentration | mg/L | Y | (fl. conc.) |  | Silver Dissolved (fl. conc.) | 7440-22-4 | Metals - Dissolved; | Metals - Dissolved; EMS code: AG-D | Chemical | Dissolved | LAB
AG-D | mg/L | Silver Dissolved | X116 | FF,FA,HNO3,Graph.Furn,AA | Samples are field-filtered (0.45 um), field-acidified with HNO3; analysis by GFAA. | 0.0001 | Y | Y | Y | 6454 |  |  | 8/26/2024 2:41:32 PM | 1996-04-12 | 6/3/1998 6:51:33 PM | EMS |  | Concentration | mg/L | Y | (fl. conc.) |  | Silver Dissolved (fl. conc.) | 7440-22-4 | Metals - Dissolved; | Metals - Dissolved; EMS code: AG-D | Chemical | Dissolved | LAB
AG-D | mg/L | Silver Dissolved | X556 | Metals in water, dissolved, "in-bottle" digestion, ICP-SFMS | The whole water sample is filtered through a 0.45 micron cellulose acetate filter paper and the filtrate is stored in a HDPE/Teflon bottle. The container is de-capped in a Class 100 workstation, acidified to 2% (v/v) HNO3 and 1% (v/v) HCl and recapped. The sample is "in-bottle" digested in a convection oven, set at 60 degrees C for 16 hours. The sample is cooled and introduced to an inductively coupled argon plasma-sector field mass spectrometer (ICP-SFMS) by pneumatic nebulization. Each respective element is measured at a specific mass to charge ratio, m/z value expressed in amu (atomic mass units), and at a specified intrument resolution. LR: Low Resolution ~300, MR: Mid Resolution ~4000 and HR: High Resolution ~10000. Ref: Method number 02-2003, National Laboratory for Environmental Testing, CCIW, Environment Canada, Burlington, Requested April 2003. | 1e-06 | Y | Y | Y | 4232 | 6/11/2003 | 12/6/2023 2:04:00 AM | 2/14/2024 9:05:32 AM | 2004-01-19 |  | UMULLER |  | Concentration | mg/L | Y | (fl. conc.) |  | Silver Dissolved (fl. conc.) | 7440-22-4 | Metals - Dissolved; | Metals - Dissolved; EMS code: AG-D | Chemical | Dissolved | LAB
AG-D | mg/L | Silver Dissolved | 0051 | Graphite Furnace AA | Graphite Furnace | 0.0001 | Y | Y | Y | 25 |  |  | 10/24/1996 3:06:39 PM | 1996-04-12 | 6/3/1998 6:51:33 PM | EMS |  | Concentration | mg/L | Y | (fl. conc.) |  | Silver Dissolved (fl. conc.) | 7440-22-4 | Metals - Dissolved; | Metals - Dissolved; EMS code: AG-D | Chemical | Dissolved | LAB
AG-E | mg/L | Silver Extractable | F069 | SEM extr,6N HCl,ICP/OES | 6.0 N Hydrochloric acid (HCI) is added to an aliquot of the sediment sample. The extract produced from the addition of the acid is then analyzed for simultaneously extracted metals (SEM) using inductively coupled plasma - optical emission spectrophotometry. The detection limits depend on the metal parameters in question and the instrumentation used. | 0.01 | Y | Y | Y | 1 |  |  | 11/25/2021 1:50:15 PM | 1999-07-09 |  | NPEPPIN |  | Concentration | mg/L | Y | (fl. conc.) |  | Silver Extractable (fl. conc.) | 7440-22-4 | Metals - Extractable; | Metals - Extractable; EMS code: AG-E | Chemical | Extractable | LAB
AG-E | mg/L | Silver Extractable | X526 | EDTA pres,pHadj.3.5,APDC/MIBK extr,AA | The sample is preserved with EDTA. pH is adjusted to 3.5 and silver is extracted by the complexing agent APDC in MIBK. The solvent is aspirated in an atomic absoption instrument the absorbance is measured SP. |  | Y | Y | Y | 27 |  |  | 8/24/1999 10:22:57 AM | 1999-07-21 |  | NPEPPIN |  | Concentration | mg/L | Y | (fl. conc.) |  | Silver Extractable (fl. conc.) | 7440-22-4 | Metals - Extractable; | Metals - Extractable; EMS code: AG-E | Chemical | Extractable | LAB
AG-E | mg/L | Silver Extractable | F047 | 24hr HNO3 dil'n,ICP-MS | There is no acid digestion required for this analysis. A water sample is collected and preserved in 0.2 percent nitric acid. To analyse the extractable metals the sample is shaken, allowed to sit for at least twenty four hours, and a decanted portion is introduced to the spectrometer. When the dissolved metals are required, the samples must first be filtered through a 0.45 micron cellulose acetate filter paper and then followed by preservation in 0.2 percent nitric acid. It is important to note that different laboratories may have very different detection limits for some variables analyzed by ICP/MS. | 5e-06 | Y | Y | Y | 2676 | 4/16/1998 10:21:00 AM | 6/26/2003 | 9/11/2019 8:46:17 AM | 1997-10-20 | 6/3/1998 6:51:47 PM | NPEPPIN |  | Concentration | mg/L | Y | (fl. conc.) |  | Silver Extractable (fl. conc.) | 7440-22-4 | Metals - Extractable; | Metals - Extractable; EMS code: AG-E | Chemical | Extractable | LAB
AG-E | mg/L | Silver Extractable | X553 | Metals in water, extractable, ICP-SFMS | The whole water sample is collected in a HDPE/Teflon bottle. In a Class 100 workstation it is preserved with concentrated Nitric Acid at concentration of 0.2 percent (v/v). The acid preservative extracts the metals from material in the sample and a decanted portion is introduced directly to an inductively coupled argon plasma-sector field mass spectrometer (ICP-SFMS) by pneumatic nebulization. Each respective element is measured at a specific mass to charge ratio, m/z value expressed in amu (atomic mass units), and at a specified instrument resolution. LR: Low Resolution ~300, MR: Mid Resolution ~4000 and HR: High Resolution ~10000. Ref: Method number 02-2061, National Laboratory for Environmental Testing, CCIW, Environment Canada, Burlington. Requested April 2003. | 1e-06 | Y | Y | Y | 2596 | 5/22/2003 | 9/21/2007 | 11/28/2023 2:11:53 PM | 2004-01-19 |  | UMULLER |  | Concentration | mg/L | Y | (fl. conc.) |  | Silver Extractable (fl. conc.) | 7440-22-4 | Metals - Extractable; | Metals - Extractable; EMS code: AG-E | Chemical | Extractable | LAB
AG-E | mg/L | Silver Extractable | F040 | Field Acid, lab filt, ICP-AES | Dissolved and extractable water samples are ready for analysis once collected. Total samples are autoclaved with HNO3 and hydrochloric acid (HCI) in capped Teflon vials. Analysis is performed via Inductively Coupled Argon Plasma Atomic Emission Spectrometry (ICP-AES)(simultaneous multielement analysis), Graphite Furnace Atomic Absorption Spectrometry (GF-AAS)(low level analysis) and/or Flame Atomic Absorption Spectrometry (F-AAS). | 0.01 | Y | Y | Y | 2361 | 8/5/1997 6:06:00 PM | 3/14/2008 | 9/11/2019 8:47:41 AM | 1997-05-08 | 6/3/1998 6:51:46 PM | MPAK |  | Concentration | mg/L | Y | (fl. conc.) |  | Silver Extractable (fl. conc.) | 7440-22-4 | Metals - Extractable; | Metals - Extractable; EMS code: AG-E | Chemical | Extractable | LAB
AG-E | mg/L | Silver Extractable | X527 | EDTA pres,pHadj.3.5-6.5EP/DPC extr,AA | The sample, preserved with EDTA, is digested with HNO3. The pH of an aliqout of this solution is adjusted to 3.5 - 6.5. The aliquot is extd. With ethyl propionate contg. diPhenylthiocarbazone. |  | Y | Y | Y | 95 |  |  | 9/9/1999 11:09:42 AM | 1999-07-21 |  | NPEPPIN |  | Concentration | mg/L | Y | (fl. conc.) |  | Silver Extractable (fl. conc.) | 7440-22-4 | Metals - Extractable; | Metals - Extractable; EMS code: AG-E | Chemical | Extractable | LAB
AG-T | ppm (S) | Silver Total | MS61 | ALS Minerals ME-MS61 | ALS Minerals ME-MS61. A prepared sample is digested with HNO3; HF; and HClO4 in a graphite heating block to near dryness. After cooling; the sample is boiled in diluted HCl. Bulk to volume with deionized water; mixed and analyzed by ICP-AES and ICP-MS for all elements. | 0.01 | Y | Y | Y | 0 |  |  |  | 2024-01-29 |  | EMS |  | DimensionlessRatio | ppm | N | (ratio) |  | Silver Total (ratio) | 7440-22-4 | Metals - Total; | Metals - Total; EMS code: AG-T | Chemical | Total | LAB
AG-T | ppm (S) | Silver Total | 6042 | Aqua Regia by ICPMS | These were sediment/soils digested with Aqua-Regia (Nitric Acid & Hydrochloric Acids) and then analysis by ICPMS. Results would be reported in ppm s. | 0.001 | Y | Y | Y | 71 | 5/14/2022 | 7/27/2024 | 7/29/2024 9:13:08 AM | 2022-06-28 |  | EMS |  | DimensionlessRatio | ppm | N | (ratio) |  | Silver Total (ratio) | 7440-22-4 | Metals - Total; | Metals - Total; EMS code: AG-T | Chemical | Total | LAB
AG-T | mg/L | Silver Total | F121 | Acid,ICPMS | Acidified, no digestion, ICPMS | 1e-05 | Y | Y | Y | 8440 | 3/12/2010 7:57:00 PM | 8/15/2024 9:10:00 PM | 8/26/2024 8:06:03 AM | 2002-07-17 | 7/16/2012 3:30:50 PM | UMULLER | RDGIBSON | Concentration | mg/L | Y | (fl. conc.) |  | Silver Total (fl. conc.) | 7440-22-4 | Metals - Total; | Metals - Total; EMS code: AG-T | Chemical | Total | LAB
AG-T | ug/m3 | Silver Total | 6040 | Aqua Regia Dig:ICP | Aqua regia digestion is used to bring the metals into solution prior to ICP analysis. For Hi-Vol samples, preliminary results are reported as mg/L and24 hr loadings as ug/m3 | 0.002 | Y | Y | Y | 437 | 2/13/2002 | 2/9/2005 | 2/23/2005 1:13:46 AM | 1998-05-05 | 6/3/1998 6:51:48 PM | NPEPPIN |  | AirConcentration | µg/m³ | Y | (air conc.) |  | Silver Total (air conc.) | 7440-22-4 | Metals - Total; | Metals - Total; EMS code: AG-T | Chemical | Total | LAB
AG-T | mg/L | Silver Total | 0051 | Graphite Furnace AA | Graphite Furnace | 0.0001 | Y | Y | Y | 292 | 11/29/1996 | 11/29/1996 | 3/8/2001 9:35:11 AM | 1996-04-12 | 6/3/1998 6:51:41 PM | EMS |  | Concentration | mg/L | Y | (fl. conc.) |  | Silver Total (fl. conc.) | 7440-22-4 | Metals - Total; | Metals - Total; EMS code: AG-T | Chemical | Total | LAB
AG-T | ug/g | Silver Total | X230 | Hom,HNO3/HCl,ICP | The sample is dried, ground to pass 100 mesh, and digested with a nitric-hydrochloric acid mixture to solubilize the solid matter and to remove the organic material by oxidation and volatilization. The completed digests are submitted to metals instrumentation personnel for metals analysis by Inductively Coupled Atomic Emission Spectrophotometry (ICP-AES), or hydride analysis by atomic absorption. | 1 | Y | Y | Y | 45 | 2/12/2002 | 12/30/2005 | 2/3/2006 1:13:12 AM | 1998-11-05 |  | NPEPPIN |  | MassPerMass | mg/kg | Y | (mass conc.) |  | Silver Total (mass conc.) | 7440-22-4 | Metals - Total; | Metals - Total; EMS code: AG-T | Chemical | Total | LAB
AG-T | mg/L | Silver Total | F068 | HNO3dig,TeflonTube/Autoclave,ICP-MS | This method is applicable to determiniation of dissolved, extractable, and total recoverable metals in ground waters, surface waters, drinking waters, and wastewaters. Dissolved, extractable, and total recoverable sample fractions are acidified to pH <1.5 with analytical grade nitric acid (HNO3) to prevent the metals from adsorbing to the container material or precipitating. Extractable and total recoverable water samples are collected in 250 ml acid washed polyethylene bottles and are preserved with Nitric acid (0.5 ml acid for 100 mLs of sample). All preserved samples are stored at room temperature and can be held for up to 6 months. Unfiltered, un-preserved samples are stored at 4C for no more than 30 days. For dissolved and extractable water fractions, no sample preparation is necessary. Samples are simply poured in autosampler tubes (i.e. 15 ml centrifuge tubes), and placed on the autosampler and analyzed. Samples for total recoverable metals are digested with Nitric acid in special Teflon tubes in an autoclave prior to analyses. | 2e-05 | Y | Y | Y | 1894 | 6/1/1999 1:23:00 PM | 7/4/2008 9:48:00 AM | 7/30/2024 6:32:12 AM | 1999-06-28 |  | NPEPPIN |  | Concentration | mg/L | Y | (fl. conc.) |  | Silver Total (fl. conc.) | 7440-22-4 | Metals - Total; | Metals - Total; EMS code: AG-T | Chemical | Total | LAB
AG-T | ug/m3 | Silver Total | 5312 | Tef:N/P Dig:ICP Analysis | A nitric/perchloric acid digestion if used to bring the metals into solution prior to ICP analysis. For Hi-Vol samples preliminary results are reported as mg/L and 24 hour loading results as ug/m3. | 0.002 | Y | Y | Y | 1185 | 11/7/1997 | 3/29/2000 | 9/28/2000 1:34:39 PM | 1996-10-13 | 6/3/1998 6:51:46 PM | EMS |  | AirConcentration | µg/m³ | Y | (air conc.) |  | Silver Total (air conc.) | 7440-22-4 | Metals - Total; | Metals - Total; EMS code: AG-T | Chemical | Total | LAB
AG-T | ppm (S) | Silver Total | MS41 | ALS Minerals ME-MS41 | ALS Minerals ME-MS41. A prepared sample is digested with aqua regia in a graphite heating block. After cooling; the resulting solution is diluted to with deionized water; mixed and analyzed by ICP-AES. Samples are then analyzed by ICP-MS for the remaining suite of elements. | 0.01 | Y | Y | Y | 0 |  |  |  | 2024-01-29 |  | EMS |  | DimensionlessRatio | ppm | N | (ratio) |  | Silver Total (ratio) | 7440-22-4 | Metals - Total; | Metals - Total; EMS code: AG-T | Chemical | Total | LAB
AG-T | ug/g | Silver Total | ICPM | Metals in Soil by CRC ICPMS | This analysis is carried out using procedures from CSR Analytical Method: "Strong Acid Leachable Metals (SALM) in Soil", BC Ministry of Environment, 26 June 2009, and procedures adapted from EPA Method 200.2. The sample is manually homogenized, dried at 60 degrees Celsius, sieved through a 63 um (230 mesh) sieve, and a representative subsample of the dry material is weighed. The sample is then digested at 95 degrees Celsius for 2 hours by block digester using concentrated nitric and hydrochloric acids. Instrumental analysis is by collision cell inductively coupled plasma - mass spectrometry (modifed from EPA Method 6020A). Method Limitation: This method is not a total digestion technique. It is a very strong acid digestion that is intended to dissolve those metals that may be environmentally available. By design, elements bound in silicate structures are not normally dissolved by this procedure as they are not usually mobile in the environment. | 0.05 | Y | Y | Y | 103 | 9/19/2017 | 9/24/2023 | 1/12/2024 11:16:52 AM | 2018-01-09 |  | JKROGH |  | MassPerMass | mg/kg | Y | (mass conc.) |  | Silver Total (mass conc.) | 7440-22-4 | Metals - Total; | Metals - Total; EMS code: AG-T | Chemical | Total | LAB
AG-T | ug/g wet | Silver Total | AGIW | AG-WET-CCMS-MID-VA | Silver in Tissue by CRC ICPMS (WET) Methodology: This method is conducted following British Columbia Lab Manual method "Metals in Animal Tissue and Vegetation (Biota) - Prescriptive". Tissue samples are homogenized and sub-sampled prior to hotblock digestion with nitric and hydrochloric acids, in combination with addition of hydrogen peroxide. Instrumental analysis is by collision cell inductively coupled plasma - mass spectrometry (modified from EPA Method 6020A). Method Limitation: This method employs a strong acid/peroxide digestion, and is intended to provide a conservative estimate of bio-available metals. Near complete recoveries are achieved for most toxicologically important metals, but elements associated with recalcitrant minerals may be only partially recovered. | 0.001 | Y | Y | Y | 24 | 11/25/2018 | 12/6/2018 | 7/21/2021 2:55:07 PM | 2018-12-24 |  | RDGIBSON |  | MassPerMass | µg/g | Y | (mass conc.) | (wet) | Silver Total (wet) (mass conc.) | 7440-22-4 | Metals - Total; | Metals - Total; EMS code: AG-T | Chemical | Total | LAB
AG-T | mg/L | Silver Total | X285 | HNO3 Dig:ICP Analysis | HNO3 Dig:ICP Analysis | 0.015 | Y | Y | Y | 645 | 8/10/1993 | 10/5/1993 | 4/30/2004 1:54:38 PM | 1996-04-12 | 6/3/1998 6:51:33 PM | EMS |  | Concentration | mg/L | Y | (fl. conc.) |  | Silver Total (fl. conc.) | 7440-22-4 | Metals - Total; | Metals - Total; EMS code: AG-T | Chemical | Total | LAB
AG-T | ppm (S) | Silver Total | 6041 | Aqua Regia by ICPMS | These were sediment/soils digested with Aqua-Regia (Nitric Acid & Hydrochloric Acids) and then analysis by ICPMS. Results would be reported in mg/kg and/or %. | 0.001 | Y | Y | Y | 21 | 10/13/2019 | 8/5/2020 | 8/10/2020 3:52:06 PM | 2019-09-06 | 10/21/2019 11:53:54 AM | KATFRASE | KATFRASE | DimensionlessRatio | ppm | N | (ratio) |  | Silver Total (ratio) | 7440-22-4 | Metals - Total; | Metals - Total; EMS code: AG-T | Chemical | Total | LAB
AG-T | ug/g wet | Silver Total | X552 | Metals in Biota (wet), total, acid digestion, ICP-SFMS | One gram of wet, homogenized biota is weighed in a microwave TFM vessel, and 8mL of Nitric Acid and 1mL of Hydrogen Peroxide are added. The sample is digested at 200 degrees C for 15 minutes in a closed vessel high pressure microwave oven. The sample is cooled, the volume adjusted to 100mL and introduced to an inductively coupled argon plasma-sector field mass spectrometer (ICP-SFMS) by pneumatic nebulization. Each respective element is measured at a specific mass to charge ratio, m/z value expressed in amu (atomic mass units), and at a specified instrument resolution. LR: Low Resolution ~300, MR: Mid Resolution ~4000. Ref: 02-2704, National Laboratory for Environmental Testing, CCIW, Environment Canada, Burlington. Requested April 2003. | 0.0001 | Y | Y | Y | 14 | 2/21/2019 | 2/21/2019 | 4/28/2021 11:01:11 AM | 2004-01-19 | 4/19/2021 1:22:00 PM | UMULLER | EMS | MassPerMass | µg/g | Y | (mass conc.) | (wet) | Silver Total (wet) (mass conc.) | 7440-22-4 | Metals - Total; | Metals - Total; EMS code: AG-T | Chemical | Total | LAB
AG-T | ug/m3 | Silver Total | F088 | Aliquot,HNO3/HCl dig, ICPMS,loading | Aliquot microwave digested with 3:1 HNO3/HCL. Analysis by ICPMS. For HiVol samples preliminary result are reported in mg/L and loading results are reported in ug/m3 |  | Y | Y | Y | 567 | 8/17/2000 10:46:00 AM | 12/14/2006 | 12/30/2006 1:01:35 AM | 2000-08-16 |  | NPEPPIN |  | AirConcentration | µg/m³ | Y | (air conc.) |  | Silver Total (air conc.) | 7440-22-4 | Metals - Total; | Metals - Total; EMS code: AG-T | Chemical | Total | LAB
AG-T | mg/L | Silver Total | TCLP | Toxicity Characteristic Leaching Procedure | The TCLP procedure is designed to simulate the climatic leaching action | 0.1 | Y | Y | Y | 38 | 12/8/2009 7:49:00 PM | 9/28/2011 5:53:00 PM | 8/24/2016 2:59:13 PM | 2009-03-06 |  | JCHUDAK |  | Concentration | mg/L | Y | (fl. conc.) |  | Silver Total (fl. conc.) | 7440-22-4 | Metals - Total; | Metals - Total; EMS code: AG-T | Chemical | Total | LAB
AG-T | mg/L | Silver Total | X545 | FA,HNO3,ICPMS | Analytes nebulized followed by ion entrained in plasma gas and extracted by means of a differentially pumped vacuum interface into a mass spectrometer | 0.0001 | Y | Y | Y | 30 |  |  | 10/30/2014 8:31:16 AM | 2001-09-18 | 9/18/2001 1:36:23 PM | NPEPPIN | NPEPPIN | Concentration | mg/L | Y | (fl. conc.) |  | Silver Total (fl. conc.) | 7440-22-4 | Metals - Total; | Metals - Total; EMS code: AG-T | Chemical | Total | LAB
AG-T | ug/g | Silver Total | PB02 | PBM,HNO3/HCl dig,dry60C,100mesh, | Performance Based Method. Samples are dried at 60 C and sieved using a .149 mm (100 mesh) sieve. Solid samples are digested with the combination of nitric and hydrochloric acids with a 1:1 ratio of acids. For environmental purposes, strong acid leaches are used routinely since they have the advantage of speed, low cost, and superio reproducibility. All metals trapped within the refractory matrix are unlikely to have environmental consequences. Instrumental analysis of extracts produced using this digestion method can be performed by a variety of techniques. |  | Y | Y | Y | 52 | 5/24/2001 3:42:00 PM | 11/26/2001 10:23:00 AM | 5/26/2003 2:27:08 PM | 2001-07-18 |  | NPEPPIN |  | MassPerMass | mg/kg | Y | (mass conc.) |  | Silver Total (mass conc.) | 7440-22-4 | Metals - Total; | Metals - Total; EMS code: AG-T | Chemical | Total | LAB
AG-T | mg/L | Silver Total | METN | MET-T-NP-CCMS-VA | Ultra trace metals in water are analyzed by CRC ICPMS, based on US EPA Method 6020B (July 2014). The detection limits provided can only be met for undigested samples. This procedure is intended for colorless, non-turbid, acid-preserved water samples (i.e. pristine water samples), having turbidity < 1 NTU and no odor. Where turbidity exceeds 1 NTU, and/or the sample is colored and has an odor, results may be biased low compared to true Total Metals concentrations. ALS recommends that turbidity analysis be requested on samples submitted for this test to aid with interpretation of results. Where turbidity is <1NTU, undigested metals are equivalent to total metals concentrations. | 5e-06 | Y | Y | Y | 3648 | 4/3/2012 | 8/19/2024 | 8/26/2024 1:03:06 PM | 2019-08-20 |  | KATFRASE |  | Concentration | mg/L | Y | (fl. conc.) |  | Silver Total (fl. conc.) | 7440-22-4 | Metals - Total; | Metals - Total; EMS code: AG-T | Chemical | Total | LAB
AG-T | mg/L | Silver Total | 0050 | Graphite Furnace AA | Graphite Furnace AA | 0.0005 | Y | Y | Y | 203 |  |  | 10/24/1996 1:43:19 PM | 1996-04-12 | 6/3/1998 6:51:33 PM | EMS |  | Concentration | mg/L | Y | (fl. conc.) |  | Silver Total (fl. conc.) | 7440-22-4 | Metals - Total; | Metals - Total; EMS code: AG-T | Chemical | Total | LAB
AG-T | ug/g wet | Silver Total | BPSY | MET-WET-BPSY-CCMS-VA | Metals in Animal Tissue and Vegetation (Biota) - Prescriptive. EPA 200.3/6020B This method is conducted following British Columbia Lab Manual method "Metals in Animal Tissue and Vegetation (Biota) - Prescriptive". Tissue samples are digested with nitric and hydrochloric acids; in combination with addition of hydrogen peroxide. Instrumental analysis is by collision cell inductively coupled plasma - mass spectrometry (modified from EPA Method 6020B). | 0.001 | Y | Y | Y | 2 | 11/7/2019 | 11/7/2019 | 6/10/2021 1:57:10 PM | 2021-06-09 | 6/9/2021 4:18:59 PM | RDGIBSON | RDGIBSON | MassPerMass | µg/g | Y | (mass conc.) | (wet) | Silver Total (wet) (mass conc.) | 7440-22-4 | Metals - Total; | Metals - Total; EMS code: AG-T | Chemical | Total | LAB
AG-T | mg/L | Silver Total | F033 | Acid Dig,Centrifuge,ICPMS | For the determination of total recoverable elements in water or waste water, take 100 ml aliquot from a well mixed, acid preserved, sample containing not more than 0.25% (w/v) total solids and transfer to a 250 ml Griffin beaker. (If total solids are greater than 0.25% reduce the size of the aliquot by a proportionate amount). Add 1 ml of conc. nitric acid and 0.5 ml conc. hydrochloric acid. Heat on a hot plate until the volume has been reduced to 15-20ml, ensuring that the sample does not boil. Allow to cool and transfer to a 50 ml volumetric flask. Dilute to volume with ASTM type I water. Centrifuge the sample or allow to stand overnight to separate insoluble material. Prior analysis, pipette 20 ml into a 50 ml volumetric flask and dilute to volume with ASTM type I water. Analysis should be performed within seventy two hours of the completed preparation. Note that different laboratories may have very different field preservation, digestion procedures and detection limits for this method, so these should be documented. | 1e-05 | Y | Y | Y | 36965 | 11/20/1997 9:32:00 AM | 8/22/2024 3:24:00 AM | 8/25/2024 9:47:36 PM | 1998-05-14 | 6/3/1998 6:51:49 PM | NPEPPIN |  | Concentration | mg/L | Y | (fl. conc.) |  | Silver Total (fl. conc.) | 7440-22-4 | Metals - Total; | Metals - Total; EMS code: AG-T | Chemical | Total | LAB
AG-T | mg/L | Silver Total | MWID | Metals in Water by ICPMS (digested) | Metals in Water by ICPMS (digested) Based on EPA 200.8 | 0.0001 | Y | Y | Y | 36830 |  |  | 7/31/2024 1:46:41 PM | 2009-12-24 |  | RDGIBSON |  | Concentration | mg/L | Y | (fl. conc.) |  | Silver Total (fl. conc.) | 7440-22-4 | Metals - Total; | Metals - Total; EMS code: AG-T | Chemical | Total | LAB
AG-T | mg/L | Silver Total | X561 | Total metals: FA, HNO3, HR-ICPMS (No Prep)Prep) | Ultra trace metals in water are analyzed by high resolution inductively coupled plasma mass spectrometry (HR-ICPMS), modified from US EPA 200.8, (Revision 5.5). The detection limits provided can only be met for undigested samples. This procedure is intended for colorless, non-turbid, acid-preserved water samples (i.e. pristine water samples), having turbidity < 1 NTU and no odor. Where turbidity exceeds 1 NTU, and/or the sample is colored and has an odor, results may be biased low compared to true Total Metals concentrations. ALS recommends that turbidity analysis be requested on samples submitted for this test to aid with interpretation of results. | 5e-06 | Y | Y | Y | 2756 | 6/8/2015 | 10/5/2023 | 1/31/2024 2:37:40 PM | 2015-10-22 |  | RDGIBSON |  | Concentration | mg/L | Y | (fl. conc.) |  | Silver Total (fl. conc.) | 7440-22-4 | Metals - Total; | Metals - Total; EMS code: AG-T | Chemical | Total | LAB
AG-T | ug/g | Silver Total | FS01 | D&G,Nitric/HCl:ICP | Samples are dried at 60 C, then either sieved to -100mesh (150 um) and rolled to homogenize, or finely ground with a mortar and pestle or a shatter box laboratory mill and rolled (depending on the submitter's requirements). The samples are then weighed into Teflon digestion liners and digested with HNO3 and HCl in a microwave oven. Samples are cooled, volumized, and allowed to settle overnight before being decanted. Analysis is then performed via Inductively Coupled Argon Plasma Atomic Emission Spectrometry (ICP-AES) (Simultaneous multi element analysis). In ICP Spectrometry, the liquid sample is pumped to the nebulizer, where it is convertd into an aerosol. the nebulizer carrier gas flow than transports the wet aerosol past the impact bead, where the large droplets condense out and flow to the waste container. The remaining fine aerosol is carried into the plasma. There, the high temperature desolvates the aerosol and dissociates the sample material into individual atoms and ions, which are excited to emit light at wavelengths characteristic of the elements present. A diffraction grating in the spectrometer diffracts the various wavelengths of light through exit slits to the photo-multiplier tubes, which convert the light energy to electronic signals. | 2 | Y | Y | Y | 2642 | 8/30/1996 1:52:00 PM | 10/10/2007 | 5/24/2008 1:15:02 AM | 1996-06-19 | 6/3/1998 6:51:44 PM | MPAK |  | MassPerMass | mg/kg | Y | (mass conc.) |  | Silver Total (mass conc.) | 7440-22-4 | Metals - Total; | Metals - Total; EMS code: AG-T | Chemical | Total | LAB
AG-T | mg/L | Silver Total | F048 | Acid dig,conc'n,nebul.,ICP-OES | This method provides an automated procedure for the Inductively Coupled Argon Plasma Optical Emission Spectrophotometric (ICP-OES) determination and quantitation of trace amounts of metals in surface, ground and precipitation waters after a manual acid digestion and concentration procedure and ultrasonic nebulisation. Organically-bound metals and metal complexes are converted to metal ions by digestion with nitric acid [method details says HCl]. To increase the signal detected by the ICP, the samples are concentrated 10 times. The liquid sample is pumped onto the piezoelectric transducer, where the ultrasonic action efficiently converts it to a fine, dense aerosol. The nebulizer carrier gas flow then transports the wet aerosol through a heated U-tube, where the solvent is vaporized. Subsequently, most of the solvent vapour is removed from the aerosol stream by passing the stream through a chilled condenser. The result is a dry, analyte-laden aerosol which is introduced to the plasma. There, the high temperature desolvates the aerosol and dissociates the sample material into individual atoms and ions, which are excited to emit light at wave lengths characteristic of the elements present. Dissolved samples are field filtered through 0.45 micron cellulose acetate filter prior to acid preservation. Samples for extractable analysis are acidified with 2 mL concentrated HNO3 per litre and allowed to stand a minimum of 24 hours before analysis. Clients providing these samples are advised to preserve samples at the collection site. | 0.01 | Y | Y | Y | 5118 | 7/4/1997 2:36:00 PM | 4/23/2003 | 9/11/2019 8:46:38 AM | 1998-11-13 |  | NPEPPIN |  | Concentration | mg/L | Y | (fl. conc.) |  | Silver Total (fl. conc.) | 7440-22-4 | Metals - Total; | Metals - Total; EMS code: AG-T | Chemical | Total | LAB
AG-T | mg/L | Silver Total | X072 | FA,HNO3,Graph.Furnace,AA | Samples are field-acidified with HNO3; HNO3 digestion; analysis by GFAA. | 0.0001 | Y | Y | Y | 5407 |  |  | 8/26/2024 2:41:32 PM | 1996-04-12 | 6/3/1998 6:51:33 PM | EMS |  | Concentration | mg/L | Y | (fl. conc.) |  | Silver Total (fl. conc.) | 7440-22-4 | Metals - Total; | Metals - Total; EMS code: AG-T | Chemical | Total | LAB
AG-T | mg/L | Silver Total | 0042 | GI HNO3 Dig: ICP | Sample aliquot taken from unfiltered unpreserved General Ions bottle; HNO3 digestion; ICP analysis. | 0.03 | Y | Y | Y | 5941 | 9/20/1996 | 7/5/2019 3:21:00 PM | 7/12/2019 6:49:13 PM | 1996-04-12 | 6/3/1998 6:51:33 PM | EMS |  | Concentration | mg/L | Y | (fl. conc.) |  | Silver Total (fl. conc.) | 7440-22-4 | Metals - Total; | Metals - Total; EMS code: AG-T | Chemical | Total | LAB
AG-T | ug/g wet | Silver Total | X560 | Metals in Biota (wet), total, acid digestion, ICP-MS | Five grams of wet, homogenized biota is weighed in a centrifuge tube, and 2.5 mL of Nitric Acid and 2.5 mL of Hydrochloric acid are added. The sample is digested at room temperature for one hour and then digested at 90-95 degrees C in a water bath or heat block for one more hour. The sample is cooled and the volume is made to 50 mL with DI. The digestate is then diluted at ten fold and introduced to a collision/reaction cell inductively coupled argon plasma mass spectrometer (CRC ICPMS) by pneumatic nebulization. Each respective element is measured at a specific mass to charge ratio, m/z value expressed in amu (atomic mass units). - Maxxam Analytics (BC). Added to EMS Dec 10, 2008. | 0.05 | Y | Y | Y | 432 | 11/15/2008 4:00:00 AM | 2/7/2022 | 3/29/2022 5:04:16 PM | 2008-12-11 | 3/11/2015 4:17:10 PM | RDGIBSON | EMS | MassPerMass | µg/g | Y | (mass conc.) | (wet) | Silver Total (wet) (mass conc.) | 7440-22-4 | Metals - Total; | Metals - Total; EMS code: AG-T | Chemical | Total | LAB
AG-T | ug/g | Silver Total | MTID | MET-DRY-CCMS-VA | Metals in Tissue by CRC ICPMS (DRY) Methodology: This method is conducted following British Columbia Lab Manual method "Metals in Animal Tissue and Vegetation (Biota) - Prescriptive". Tissue samples are homogenized and sub-sampled prior to hotblock digestion with nitric and hydrochloric acids, in combination with addition of hydrogen peroxide. Instrumental analysis is by collision cell inductively coupled plasma - mass spectrometry (modified from EPA Method 6020A). Method Limitation: This method employs a strong acid/peroxide digestion, and is intended to provide a conservative estimate of bio-available metals. Near complete recoveries are achieved for most toxicologically important metals, but elements associated with recalcitrant minerals may be only partially recovered. |  | Y | Y | Y | 6 | 2/7/2022 | 2/7/2022 | 3/29/2022 5:04:16 PM | 2022-03-25 |  | EMS |  | MassPerMass | mg/kg | Y | (mass conc.) |  | Silver Total (mass conc.) | 7440-22-4 | Metals - Total; | Metals - Total; EMS code: AG-T | Chemical | Total | LAB
AG-T | ug/g | Silver Total | X551 | Metals in Biota (dry), total, acid digestion, ICP-SFMS | 0.1 gram of dried, homogenized biota is weighed in a microwave TFM vessel, and 8mL of Nitric Acid and 1mL of Hydrogen Peroxide are added. The sample is digested at 200 degrees C for 15 minutes in a closed vessel high pressure microwave oven. The sample is cooled, the volume adjusted to 100mL and introduced to an iductively coupled argon plasma-sector field mass spectrometer (ICP-SFMS) by pneumatic nebulization. Each respective element is measured at a specific mass to charge ratio, m/z value expressed in amu (atomic mass units), and at a specified instrument resolution. LR: Low Resolution ~300, MR: Mid Resolution ~4000 and HR: Hich Resolution ~10000. Ref: Method number 02-2705, National Laboratory for Environmental Testing, CCIW, Environment Canada, Burlington. Requested April 2003. | 0.001 | Y | Y | Y | 9 | 12/22/2015 | 11/2/2022 11:44:00 AM | 11/10/2022 4:22:46 PM | 2004-01-19 |  | UMULLER |  | MassPerMass | mg/kg | Y | (mass conc.) |  | Silver Total (mass conc.) | 7440-22-4 | Metals - Total; | Metals - Total; EMS code: AG-T | Chemical | Total | LAB
AG-T | ug/g | Silver Total | FS82 | PBM,dry60C,10mesh,SALM-ICP-MS | PBM;dry60C;10mesh;SALM-ICP-MS | 0.05 | Y | Y | Y | 3648 | 5/21/2002 | 7/24/2024 | 7/29/2024 5:17:09 PM | 2002-11-12 |  | UMULLER |  | MassPerMass | mg/kg | Y | (mass conc.) |  | Silver Total (mass conc.) | 7440-22-4 | Metals - Total; | Metals - Total; EMS code: AG-T | Chemical | Total | LAB
AG-T | mg/L | Silver Total | 0040 | HNO3 Dig: ICP Analysis | HNO3 Dig: ICP Analysis | 0.01 | Y | Y | Y | 1600 |  |  | 3/23/2010 5:08:07 PM | 1996-04-12 | 6/3/1998 6:51:33 PM | EMS |  | Concentration | mg/L | Y | (fl. conc.) |  | Silver Total (fl. conc.) | 7440-22-4 | Metals - Total; | Metals - Total; EMS code: AG-T | Chemical | Total | LAB
AG-T | mg/L | Silver Total | F082 | Aliquot,HNO3/HCL dig,ICPMS | Aliquot microwave digested with 3:1 HNO3/HCL. Analysis by ICPMS. For HiVol samples preliminary result are reported in mg/L and loading results are reported in ug/m3 | 2e-05 | Y | Y | Y | 27437 | 7/26/2005 7:04:00 AM | 8/23/2024 | 8/26/2024 4:53:05 PM | 2000-06-13 | 8/15/2019 11:02:43 AM | NPEPPIN | KATFRASE | Concentration | mg/L | Y | (fl. conc.) |  | Silver Total (fl. conc.) | 7440-22-4 | Metals - Total; | Metals - Total; EMS code: AG-T | Chemical | Total | LAB
AG-T | mg/L | Silver Total | X555 | Metals in water, total, "in-bottle" digestion, ICP-SFMS | The whole water sample collected in a HDPE/Teflon bottle, is de-capped in a Class 100 workstation, acidified to 2% (v/v) HNO3 and 1% (v/v) HCl and recapped. The sample is "in-bottle" digested in a convection oven, set at 60 degrees C for 16 hours. The sample is cooled and introduced to an inductively coupled argon plasma-sector field mass spectrometer (ICP-SFMS) by pneumatic nebulization. Each respective element is measured at a specific mass to charge ratio, m/z value expressed in amu (atomic mass units), and at resolution. LR: Low Resolution ~300, MR: Mid Resolution ~4000 and HR: High a specified instrument Resolution ~10000. Ref: Method number 02-2003, National Laboratory for Environmental Testing, CCIW, Environment Canada, Burlington. Requested April 2003. | 1e-06 | Y | Y | Y | 17333 | 5/20/2003 | 12/6/2023 2:08:00 AM | 2/14/2024 9:05:32 AM | 2004-01-14 |  | UMULLER |  | Concentration | mg/L | Y | (fl. conc.) |  | Silver Total (fl. conc.) | 7440-22-4 | Metals - Total; | Metals - Total; EMS code: AG-T | Chemical | Total | LAB
AG-T | mg/L | Silver Total | F024 | HNO3/HCl,GFAA | Dissolved and extractable water samples are ready for analysis once collected. Total samples are autoclaved with HNO3 and hydrochloric acid (HCl) in capped Teflon vials. Analysis is performed via Inductively Coupled Argon Plasma Atomic Emission Spectrometry (ICP-AES) (simultaneous multielement analysis), Graphite Furnace Atomic Absorption Spectrometry (GF-AAS) (low level analysis) and/or Flame Atomic Absorption Spectrometry (F-AAS). | 0.0006 | Y | Y | Y | 31 | 10/18/1996 6:11:00 PM | 12/1/2010 | 12/9/2010 3:55:26 PM | 1997-04-02 | 6/3/1998 6:51:46 PM | MPAK |  | Concentration | mg/L | Y | (fl. conc.) |  | Silver Total (fl. conc.) | 7440-22-4 | Metals - Total; | Metals - Total; EMS code: AG-T | Chemical | Total | LAB
AG-T | mg/L | Silver Total | MSEA | Metals in Seawater (undigested) by (CRC) ICP-MS | This method describes the multi-elemental determination of trace elements by Collision/Reaction Cell (CRC) ICP-MS. The method measures ions produced by a radio-frequency inductively coupled plasma. Analyte species originating in a liquid are pneumatically nebulized and the resulting aerosol transported by argon gas into the plasma torch. The ions produced are entrained in the plasma gas and extracted, by means of a differentially pumped vacuum interface, into a mass spectrometer. The ions flow through a Collision/Reaction Cell (CRC) where depending on the mode of analysis either the ions passes through the cell unhindered (no gas - classical ICP-MS mode), or collisions/reactions chemistry takes place between the ions and a neutral gas such as hydrogen or helium. Any polyatomic ions which give interference on the analytes are removed. The ions are then sorted according to their mass-to-charge ratios by a quadrupole mass spectrometer. The ions are then sorted according to their mass-to-charge ratios by a quadrupole mass spectrometer. The ions transmitted through the quadrupole are quantified by a channel electron multiplier and the ion information processed by a data handling system. Although most polyatomic interference is removed, any additional interference must be assessed and valid corrections applied or the data flagged to indicate problems. Interference correction must include compensation for background ions contributed by the plasma gas, reagents, and constituents of the sample matrix. Instrument drift as well as suppressions or enhancements of instrument response caused by the sample matrix must be corrected for by the use of internal standards. | 2e-05 | Y | Y | Y | 441 | 8/12/2009 12:08:00 AM | 4/29/2024 | 6/7/2024 10:03:16 AM | 2009-09-08 |  | RDGIBSON |  | Concentration | mg/L | Y | (fl. conc.) |  | Silver Total (fl. conc.) | 7440-22-4 | Metals - Total; | Metals - Total; EMS code: AG-T | Chemical | Total | LAB
AG-T | mg/L | Silver Total | X073 | FA,HNO3,Direct Flame,AA | Samples are field-acidified with HNO3; HNO3 digestion; analysis by AA. | 0.005 | Y | Y | Y | 72 |  |  | 6/25/2003 9:22:37 AM | 1996-04-12 | 6/3/1998 6:51:33 PM | EMS |  | Concentration | mg/L | Y | (fl. conc.) |  | Silver Total (fl. conc.) | 7440-22-4 | Metals - Total; | Metals - Total; EMS code: AG-T | Chemical | Total | LAB
AG-T | mg/L | Silver Total | X170 | ICP | This code is a catch-all for a large variety of ICP methods including water and also solids digested into solutions as an intermediate result. Different procedures will have greatly different detection limits. (a) For air particulates, preliminary results are reported in mg/L and loading results as mg/dm2/d. (b) For Hi Vol samples preliminiary results are same as reported in mg/L and loading results as ug/m3. (c) Note that for Rain samples there is no filtration nor digestion steps. | 0.01 | Y | Y | Y | 6418 | 3/28/1996 6:03:00 PM | 11/7/2018 | 10/24/2022 10:58:25 AM | 1996-04-12 | 6/3/1998 6:51:41 PM | EMS |  | Concentration | mg/L | Y | (fl. conc.) |  | Silver Total (fl. conc.) | 7440-22-4 | Metals - Total; | Metals - Total; EMS code: AG-T | Chemical | Total | LAB
AG-T | mg/L | Silver Total | 0200 | 20%A.R.Dig:AA Anal. | 20%A.R.Dig:AA Anal. | 0.005 | Y | Y | Y | 82 |  |  | 10/24/1996 6:09:27 AM | 1996-04-12 | 6/3/1998 6:51:33 PM | EMS |  | Concentration | mg/L | Y | (fl. conc.) |  | Silver Total (fl. conc.) | 7440-22-4 | Metals - Total; | Metals - Total; EMS code: AG-T | Chemical | Total | LAB
AG-T | mg/L | Silver Total | EQ01 | EQUIS (mg/L) | EQUIS (mg/L) | 0 | N | Y | Y | 47 |  |  | 10/22/1996 11:26:55 PM | 1996-04-12 | 6/3/1998 6:51:33 PM | EMS |  | Concentration | mg/L | Y | (fl. conc.) |  | Silver Total (fl. conc.) | 7440-22-4 | Metals - Total; | Metals - Total; EMS code: AG-T | Chemical | Total | LAB
AG-T | mg/L | Silver Total | 1349 | Tot Metals in Seawater by CRC ICPMS (BC) | Tot Metals in Seawater by CRC ICPMS (BC) | 0.0001 | Y | Y | Y | 98 | 8/16/2019 | 12/5/2021 | 12/13/2021 12:34:16 PM | 2019-09-30 |  | KATFRASE |  | Concentration | mg/L | Y | (fl. conc.) |  | Silver Total (fl. conc.) | 7440-22-4 | Metals - Total; | Metals - Total; EMS code: AG-T | Chemical | Total | LAB
AG-T | mg/L | Silver Total | F038 | HNO3/HCl dig,ICP-AES | Dissolved and extractable water samples are ready for analysis once collected. Total samples are autoclaved with HNO3 and hydrochloric acid (HCI) in capped Teflon vials. Analysis is performed via Inductively Coupled Argon Plasma Atomic Emission Spectrometry (ICP-AES)(simultaneous multielement analysis), Graphite Furnace Atomic Absorption Spectrometry (GF-AAS) (low level analysis) and/or Flame Atomic Absorption Spectrometry (F-AAS). | 0.01 | Y | Y | Y | 5417 | 8/1/1996 6:04:00 PM | 6/25/2010 3:15:00 PM | 9/11/2019 8:47:42 AM | 1997-05-08 | 6/3/1998 6:51:46 PM | MPAK |  | Concentration | mg/L | Y | (fl. conc.) |  | Silver Total (fl. conc.) | 7440-22-4 | Metals - Total; | Metals - Total; EMS code: AG-T | Chemical | Total | LAB
AG-T | mg/dm2/d | Silver Total | DUST | HNO3 dig.(microwave or hot block),ICPMS | HNO3 dig.(microwave or hot block);ICPMS |  | Y | Y | Y | 420 | 7/19/2016 1:11:00 PM | 7/31/2024 | 7/31/2024 11:23:12 AM | 2008-02-07 |  | RDGIBSON |  | MassFlux | mg/dm²/d | Y | (mass flux) |  | Silver Total (mass flux) | 7440-22-4 | Metals - Total; | Metals - Total; EMS code: AG-T | Chemical | Total | LAB
AG-T | ug/g | Silver Total | FB04 | Freeze Dry,Hom,Acid Dig,ICP | Sample moisture is determined by weighing samples before and after freeze-drying. Freeze-dried samples are ground in a Braun blender before being weighed into Teflon PFA liners and digested with concentrated HNO3, concentrated HCl, and 50% H2O2 in a microwave oven; which results in a complete sample dissolution. Samples are collected and volumized with Type I water. Analysis is performed via Inductively Coupled Argon Plasma Atomic Emission Spectrometry (ICP-AES) (simultaneous multielement analysis), Graphite Furnace Atomic Absorption Spectrometry (GF-AAS) (low level Cd and Pb analysis), Cold Vapor Atomic Fluorescence Spectrometry (CV-AFS) (Mercury analysis) and/or Flame Atomic Absorption Spectrometry (F-AAS) (K analysis). | 0.8 | Y | Y | Y | 163 | 5/8/2000 8:36:00 AM | 2/1/2002 2:30:00 PM | 6/23/2003 1:15:26 PM | 1996-06-19 | 6/3/1998 6:51:44 PM | MPAK |  | MassPerMass | mg/kg | Y | (mass conc.) |  | Silver Total (mass conc.) | 7440-22-4 | Metals - Total; | Metals - Total; EMS code: AG-T | Chemical | Total | LAB
AG-T | ug/g | Silver Total | X492 | Homo,HCl-HNO3 dig,ICP-OES,dry calc | Homogenize samples and take sub-samples. For Total Recoverable Metals Digest with Aqua Regia. Analyses by ICPOES (Inductively Coupled Plasma Optical Emission Spectroscopy). Perform moisture analysis on separate sub-sample and calculate dry value for trace elements. | 2 | Y | Y | Y | 1 |  |  | 8/24/2016 2:59:13 PM | 1998-11-13 |  | NPEPPIN |  | MassPerMass | mg/kg | Y | (mass conc.) |  | Silver Total (mass conc.) | 7440-22-4 | Metals - Total; | Metals - Total; EMS code: AG-T | Chemical | Total | LAB
AK-G | meq/L | Gran Alkalinity (as -H+) | 5151 | Gran Titrating | This analysis is carried out using procedures adapted from Alberta Environment Method No. 51515. Gran Alkalinity is determined by titrating with strong acid to pH 4.0, purging with nitrogen gas to remove CO2, and back-titrating with strong alkali to pH 5.6. Gran Alkalinity is expressed as negative H+ milli-equivalent/L (-H+ meq/L). | 0.1 | Y | Y | Y | 159 | 6/9/2015 | 5/24/2016 | 12/9/2016 1:27:42 PM | 2016-11-28 |  | RRENTON |  | Equivalents | mEq/L | Y | (equiv./vol.) |  | Gran Alkalinity (as -H+) (equiv./vol.) |  | Major Ions; | Major Ions; EMS code: AK-G | Chemical |  | LAB
AK-H | mg/L | Hydroxide Alkalinity (as CaCO3) | XCAL | Calculation (mg/L) | Calculation (mg/L) | 0.5 | Y | Y | Y | 11721 | 10/18/2006 | 8/14/2024 4:15:00 PM | 8/26/2024 8:06:03 AM | 2008-12-10 |  | RDGIBSON |  | Concentration | mg/L | Y | (fl. conc.) |  | Hydroxide Alkalinity (as CaCO3) (fl. conc.) |  | Major Ions; | Major Ions; EMS code: AK-H | Chemical |  | LAB
AK-H | mg/L | Hydroxide Alkalinity (as CaCO3) | X004 | Colorimetric Titr'n | Manual Colorimetric Titration with standard solutions to specified endpoint as identified by colorimetric indicator. The sample must not be filtered, diluted, concentrated, or altered in any way. | 0.05 | Y | Y | Y | 397 | 8/13/2008 | 1/12/2015 | 2/27/2023 9:10:15 AM | 2009-03-31 | 3/31/2009 8:38:36 AM | JCHUDAK | JCHUDAK | Concentration | mg/L | Y | (fl. conc.) |  | Hydroxide Alkalinity (as CaCO3) (fl. conc.) |  | Major Ions; | Major Ions; EMS code: AK-H | Chemical |  | LAB
AK-H | mg/L | Hydroxide Alkalinity (as CaCO3) | X310 | Auto Potentiometric Titration | For alkalinity, the samples are (if pH >8.3) first titrated to a pH of 8.3, which corresponds to the equivalence point for the conversion of carbonate to bicarbonate ion (phenolphthalein alkalinity). The titration is continued to a pH of 4.5, which corresponds approx. to the equivalence point for the conversion of bicarbonate ion to carbonic acid (entire titration indicates total alkalinity). An extrapolation technique based on the near proportionality of concentration of H ions to excess of titrant beyond the equivalence pt. The amount of standard acid required to reduce pH exactly 0.30pH (ie. pH4.2) unit is measured carefully. Because this change in pH corresponds to an exact coupling of the H ion concentration, a simple extrapolation can be made to the equivalence point. The actual pH of the stoichiometric end point in alkalinity determinations is the pH at which the inflection in the titration curve occurs and is directly related to the amount of carbonate alkalinity originally present in the sample. A curve can be constructed by plotting pH vs mL of acid to find this point directly. *** For acidity, the samples are titrated to a pH of 8.3, which corresponds to stoichiometric neutralization of carbonic acid to bicarbonate. Because the colour change of phenolphthalein indicator is close to pH 8.3, this value generally is accepted as a standard end point for titration of total acidity, including CO2 and most weak acids an equivalence point for the conversion of carbonate to bicarbonate ion (phenolphthalein alkalinity). The actual pH of the stoichiometric end point in acidity determinations is the pH at which the inflection in the titration curve occurs and is directly related to the amount of titratable acids originally present in the sample. A curve can be constructed by plotting pH vs. mL of NaOH to find this point directly. (To obtain the actual substances contributing to acidity/alkalinity, a pH vs mL of acid/NaOH curve should be constructed.) | 0.05 | Y | Y | Y | 20430 | 3/24/2011 | 8/22/2024 12:47:00 PM | 8/26/2024 4:53:05 PM | 2009-03-31 | 3/31/2009 8:37:29 AM | JCHUDAK | JCHUDAK | Concentration | mg/L | Y | (fl. conc.) |  | Hydroxide Alkalinity (as CaCO3) (fl. conc.) |  | Major Ions; | Major Ions; EMS code: AK-H | Chemical |  | LAB
AK-L | mg/L | Alkalinity Total (as CaCO3) Leachable | E294 | Alkalinity by titration (Shakeflask, 3:1 Ratio with Water) | BC MINISTRY OF ENERGY AND MINES A sample is extracted with water at a 3:1 liquid to solid ratio for 24 hours; then filtered through a 0.45 micron membrane filter. Analysis is by autotitration. | 2 | Y | Y | Y | 0 |  |  |  | 2024-01-24 |  | RDGIBSON |  | Concentration | mg/L | Y | (fl. conc.) |  | Alkalinity Total (as CaCO3) Leachable (fl. conc.) |  |  | EMS code: AK-L | Physical |  | LAB
AK-T | mg/L | Alkalinity:Total | X310 | Auto Potentiometric Titration | For alkalinity, the samples are (if pH >8.3) first titrated to a pH of 8.3, which corresponds to the equivalence point for the conversion of carbonate to bicarbonate ion (phenolphthalein alkalinity). The titration is continued to a pH of 4.5, which corresponds approx. to the equivalence point for the conversion of bicarbonate ion to carbonic acid (entire titration indicates total alkalinity). An extrapolation technique based on the near proportionality of concentration of H ions to excess of titrant beyond the equivalence pt. The amount of standard acid required to reduce pH exactly 0.30pH (ie. pH4.2) unit is measured carefully. Because this change in pH corresponds to an exact coupling of the H ion concentration, a simple extrapolation can be made to the equivalence point. The actual pH of the stoichiometric end point in alkalinity determinations is the pH at which the inflection in the titration curve occurs and is directly related to the amount of carbonate alkalinity originally present in the sample. A curve can be constructed by plotting pH vs mL of acid to find this point directly. *** For acidity, the samples are titrated to a pH of 8.3, which corresponds to stoichiometric neutralization of carbonic acid to bicarbonate. Because the colour change of phenolphthalein indicator is close to pH 8.3, this value generally is accepted as a standard end point for titration of total acidity, including CO2 and most weak acids an equivalence point for the conversion of carbonate to bicarbonate ion (phenolphthalein alkalinity). The actual pH of the stoichiometric end point in acidity determinations is the pH at which the inflection in the titration curve occurs and is directly related to the amount of titratable acids originally present in the sample. A curve can be constructed by plotting pH vs. mL of NaOH to find this point directly. (To obtain the actual substances contributing to acidity/alkalinity, a pH vs mL of acid/NaOH curve should be constructed.) | 1 | Y | Y | Y | 0 |  |  |  | 2023-10-10 | 10/10/2023 3:55:42 PM | EMS | EMS | Concentration | mg/L | Y | (fl. conc.) |  | Alkalinity:Total (fl. conc.) |  | Major Ions; | Major Ions; EMS code: AK-T | Chemical | Total | LAB
AK-T | mg/L | Alkalinity:Total | X004 | Colorimetric Titr'n | Manual Colorimetric Titration with standard solutions to specified endpoint as identified by colorimetric indicator. The sample must not be filtered, diluted, concentrated, or altered in any way. | 1 | Y | Y | Y | 44022 | 11/18/1997 4:00:00 PM | 8/6/2024 | 8/21/2024 11:21:21 AM | 1996-04-12 | 6/3/1998 6:51:41 PM | EMS |  | Concentration | mg/L | Y | (fl. conc.) |  | Alkalinity:Total (fl. conc.) |  | Major Ions; | Major Ions; EMS code: AK-T | Chemical | Total | LAB
AK-T | ueq/L | Alkalinity:Total | EQ18 | EQUIS (ueq/L) | EQUIS (ueq/L) | 0 | N | Y | Y | 118 |  |  | 10/22/1996 11:12:48 PM | 1996-04-12 | 6/3/1998 6:51:33 PM | EMS |  | Equivalents | µEq/L | Y | (equiv./vol.) |  | Alkalinity:Total (equiv./vol.) |  | Major Ions; | Major Ions; EMS code: AK-T | Chemical | Total | LAB
AK-T | ueq/L | Alkalinity:Total | 5062 | Electrom/Grans Plot | A glass electrode, calibrated with a pH 4.10 H2SO4 buffer, is used with a digital pH/mV meter to measure sample pH. Acidity is then determined on the same aliquot by titration with uL incrments of 0.01N NaOH. The volume of NaOH added to the sample is plotted against the Gran's function, calculated from the readings obtained during the titration. Equivalence point for strong acidity is obtained by extrapolating the Gran's functions to the volume axis. Total alkalinity is determined in a similar manner using 0.01N H2SO4. | 0.1 | Y | Y | Y | 8248 | 4/30/1997 3:50:00 PM | 5/29/2015 12:43:00 PM | 6/8/2015 4:04:22 PM | 1996-04-12 | 6/3/1998 6:51:33 PM | EMS |  | Equivalents | µEq/L | Y | (equiv./vol.) |  | Alkalinity:Total (equiv./vol.) |  | Major Ions; | Major Ions; EMS code: AK-T | Chemical | Total | LAB
AL-D | mg/L | Aluminum Dissolved | F039 | Field filt + acid, ICP-AES | Dissolved and extractable water samples are ready for analysis once collected. Total samples are autoclaved with HNO3 and hydrochloric acid (HCI) in capped Teflon vials. Analysis is performed via Inductively Coupled Argon Plasma Atomic Emission Spectrometry (ICP-AES)(simultaneous multielement analysis), Graphite Furnace Atomic Absorption Spectrometry (GF-AAS)(low level analysis) and/or Flame Atomic Absorption Spectrometry (F-AAS). | 0.05 | Y | Y | Y | 6515 | 7/29/1996 6:03:00 PM | 7/5/2019 1:21:00 PM | 9/11/2019 8:47:50 AM | 1997-05-08 | 6/3/1998 6:51:46 PM | MPAK |  | Concentration | mg/L | Y | (fl. conc.) |  | Aluminum Dissolved (fl. conc.) | 7429-90-5 | Metals - Dissolved; | Metals - Dissolved; EMS code: AL-D | Chemical | Dissolved | LAB
AL-D | mg/L | Aluminum Dissolved | MWID | Metals in Water by ICPMS (digested) | Metals in Water by ICPMS (digested) Based on EPA 200.8 | 0.01 | Y | Y | Y | 31789 | 10/26/2018 5:20:00 PM | 10/26/2018 5:20:00 PM | 7/31/2024 1:46:41 PM | 2009-12-24 |  | RDGIBSON |  | Concentration | mg/L | Y | (fl. conc.) |  | Aluminum Dissolved (fl. conc.) | 7429-90-5 | Metals - Dissolved; | Metals - Dissolved; EMS code: AL-D | Chemical | Dissolved | LAB
AL-D | mg/L | Aluminum Dissolved | X556 | Metals in water, dissolved, "in-bottle" digestion, ICP-SFMS | The whole water sample is filtered through a 0.45 micron cellulose acetate filter paper and the filtrate is stored in a HDPE/Teflon bottle. The container is de-capped in a Class 100 workstation, acidified to 2% (v/v) HNO3 and 1% (v/v) HCl and recapped. The sample is "in-bottle" digested in a convection oven, set at 60 degrees C for 16 hours. The sample is cooled and introduced to an inductively coupled argon plasma-sector field mass spectrometer (ICP-SFMS) by pneumatic nebulization. Each respective element is measured at a specific mass to charge ratio, m/z value expressed in amu (atomic mass units), and at a specified intrument resolution. LR: Low Resolution ~300, MR: Mid Resolution ~4000 and HR: High Resolution ~10000. Ref: Method number 02-2003, National Laboratory for Environmental Testing, CCIW, Environment Canada, Burlington, Requested April 2003. | 0.0002 | Y | Y | Y | 4226 | 6/11/2003 | 12/6/2023 2:04:00 AM | 2/14/2024 9:05:32 AM | 2004-01-19 |  | UMULLER |  | Concentration | mg/L | Y | (fl. conc.) |  | Aluminum Dissolved (fl. conc.) | 7429-90-5 | Metals - Dissolved; | Metals - Dissolved; EMS code: AL-D | Chemical | Dissolved | LAB
AL-D | mg/L | Aluminum Dissolved | F032 | Filt.,Acid,ICPMS | There is no acid digestion required for this analysis. A water sample is collected and preserved in 0.2 percent nitric acid. To analyse the extractable metals the sample is shaken, allowed to sit for at least twenty four hours, and a decanted portion is introduced to the spectrometer. When the dissolved metals are required, the samples must first be filtered through a 0.45 micron cellulose acetate filter paper and then followed by preservation in 0.2 percent nitric acid. It is important to note that different laboratories may have very different detection limits for some variables analyzed by ICPMS. | 0.005 | Y | Y | Y | 78695 | 8/16/1996 2:35:00 PM | 8/26/2024 | 8/26/2024 4:53:05 PM | 1997-01-16 | 8/15/2019 10:51:15 AM | MPAK | KATFRASE | Concentration | mg/L | Y | (fl. conc.) |  | Aluminum Dissolved (fl. conc.) | 7429-90-5 | Metals - Dissolved; | Metals - Dissolved; EMS code: AL-D | Chemical | Dissolved | LAB
AL-D | mg/L | Aluminum Dissolved | 0200 | 20%A.R.Dig:AA Anal. | 20%A.R.Dig:AA Anal. | 0.5 | Y | Y | Y | 1 |  |  | 10/24/1996 2:24:33 AM | 1996-04-12 | 6/3/1998 6:51:41 PM | EMS |  | Concentration | mg/L | Y | (fl. conc.) |  | Aluminum Dissolved (fl. conc.) | 7429-90-5 | Metals - Dissolved; | Metals - Dissolved; EMS code: AL-D | Chemical | Dissolved | LAB
AL-D | mg/L | Aluminum Dissolved | EQ01 | EQUIS (mg/L) | EQUIS (mg/L) | 0.01 | N | Y | Y | 1414 |  |  | 10/23/1996 12:07:03 AM | 1996-04-12 | 6/3/1998 6:51:33 PM | EMS |  | Concentration | mg/L | Y | (fl. conc.) |  | Aluminum Dissolved (fl. conc.) | 7429-90-5 | Metals - Dissolved; | Metals - Dissolved; EMS code: AL-D | Chemical | Dissolved | LAB
AL-D | mg/L | Aluminum Dissolved | 0050 | Graphite Furnace AA | Graphite Furnace AA | 0.01 | Y | Y | Y | 1152 | 9/29/2000 | 12/1/2010 | 12/9/2010 3:55:26 PM | 1996-04-12 | 6/3/1998 6:51:41 PM | EMS |  | Concentration | mg/L | Y | (fl. conc.) |  | Aluminum Dissolved (fl. conc.) | 7429-90-5 | Metals - Dissolved; | Metals - Dissolved; EMS code: AL-D | Chemical | Dissolved | LAB
AL-D | mg/L | Aluminum Dissolved | X028 | FA,HNO3,AA Analysis | FA;HNO3;AA Analysis | 0.5 | Y | Y | Y | 39 |  |  | 7/15/1998 11:52:31 AM | 1996-04-12 | 6/3/1998 6:51:33 PM | EMS |  | Concentration | mg/L | Y | (fl. conc.) |  | Aluminum Dissolved (fl. conc.) | 7429-90-5 | Metals - Dissolved; | Metals - Dissolved; EMS code: AL-D | Chemical | Dissolved | LAB
AL-D | mg/L | Aluminum Dissolved | X170 | ICP | This code is a catch-all for a large variety of ICP methods including water and also solids digested into solutions as an intermediate result. Different procedures will have greatly different detection limits. (a) For air particulates, preliminary results are reported in mg/L and loading results as mg/dm2/d. (b) For Hi Vol samples preliminiary results are same as reported in mg/L and loading results as ug/m3. (c) Note that for Rain samples there is no filtration nor digestion steps. | 0.2 | Y | Y | Y | 13512 | 4/11/1996 6:01:00 PM | 11/6/2018 | 6/12/2020 1:21:11 PM | 1996-04-12 | 6/3/1998 6:51:33 PM | EMS |  | Concentration | mg/L | Y | (fl. conc.) |  | Aluminum Dissolved (fl. conc.) | 7429-90-5 | Metals - Dissolved; | Metals - Dissolved; EMS code: AL-D | Chemical | Dissolved | LAB
AL-D | mg/L | Aluminum Dissolved | 0060 | HNO3 Dig: HGA Analysis | HNO3 Dig: HGA Analysis | 0.01 | Y | Y | Y | 66 |  |  | 3/18/2004 10:28:49 AM | 1996-04-12 | 6/3/1998 6:51:33 PM | EMS |  | Concentration | mg/L | Y | (fl. conc.) |  | Aluminum Dissolved (fl. conc.) | 7429-90-5 | Metals - Dissolved; | Metals - Dissolved; EMS code: AL-D | Chemical | Dissolved | LAB
AL-D | mg/L | Aluminum Dissolved | METB | MET-D-NP-BCMOE-MS-VA | Ultra trace metals in water are analyzed by CRC ICPMS, based on US EPA Method 6020B (July 2014). This procedure is intended for pristine field-filtered acid-preserved water samples. The detection limits (LOR) for this test are based on lab instrumental analysis only, not including filtration. ALS-supplied field filtration equipment does not support these LOR. Therefore, because of the high probability of false positives due to filtration, it is strongly recommended that a filtration blank be analysed to aid in data interpretation. BC Ministry of Environment version. | 0.0005 | Y | Y | Y | 596 | 6/11/2019 | 8/15/2024 | 8/23/2024 9:50:35 AM | 2019-08-20 |  | KATFRASE |  | Concentration | mg/L | Y | (fl. conc.) |  | Aluminum Dissolved (fl. conc.) | 7429-90-5 | Metals - Dissolved; | Metals - Dissolved; EMS code: AL-D | Chemical | Dissolved | LAB
AL-D | mg/L | Aluminum Dissolved | METC | MET-D-NP-CCMS-VA | Ultra trace metals in water are analyzed by CRC ICPMS, based on US EPA Method 6020B (July 2014). This procedure is intended for pristine field-filtered acid-preserved water samples. The detection limits (LOR) for this test are based on lab instrumental analysis only, not including filtration. ALS-supplied field filtration equipment does not support these LOR. Therefore, because of the high probability of false positives due to filtration, it is strongly recommended that a filtration blank be analysed to aid in data interpretation. | 0.0005 | Y | Y | Y | 1179 | 4/3/2012 | 4/9/2024 | 4/29/2024 4:53:06 PM | 2021-04-14 |  | RDGIBSON |  | Concentration | mg/L | Y | (fl. conc.) |  | Aluminum Dissolved (fl. conc.) | 7429-90-5 | Metals - Dissolved; | Metals - Dissolved; EMS code: AL-D | Chemical | Dissolved | LAB
AL-D | mg/L | Aluminum Dissolved | F049 | Filt,HNO3 dig,conc'n,nebul'ICP-OES | This method provides an automated procedure for the Inductively Coupled Argon Plasma Optical Emission Spectrophotometric (ICP-OES) determination and quantitation of trace amounts of metals in surface, ground and precipitation waters after a manual acid digestion and concentration procedure and ultrasonic nebulisation. Organically-bound metals and metal complexes are converted to metal ions by digestion with nitric acid. To increase the signal detected by the ICP, the samples are concentrated 10 times. The liquid sample is pumped onto the piezoelectric transducer, where the ultrasonic action efficiently converts it to a fine, dense aerosol. The nebulizer carrier gas flow then transports the wet aerosol through a heated U-tube, where the solvent is vaporized. Subsequently, most of the solvent vapour is removed from the aerosol stream by passing the stream through a chilled condenser. The result is a dry, analyte-laden aerosol which is introduced to the plasma. There, the high temperature desolvates the aerosol and dissociates the sample material into individual atoms and ions, which are excited to emit light at wave lengths characteristic of the elements present. Dissolved samples are field filtered through 0.45 micron cellulose acetate filter prior to acid preservation. Samples for extractable analysis are acidified with 2 mL concentrated HNO3 per litre and allowed to stand a minimum of 24 hours before analysis. Clients providing these samples are advised to preserve samples at the collection site. | 0.002 | Y | Y | Y | 1441 |  |  | 5/19/2015 3:10:33 PM | 1997-10-20 | 6/3/1998 6:51:47 PM | NPEPPIN |  | Concentration | mg/L | Y | (fl. conc.) |  | Aluminum Dissolved (fl. conc.) | 7429-90-5 | Metals - Dissolved; | Metals - Dissolved; EMS code: AL-D | Chemical | Dissolved | LAB
AL-D | mg/L | Aluminum Dissolved | F023 | HNO3,GFAA | Dissolved and extractable water samples are ready for analysis once collected. Total samples are autoclaved with HNO3 and hydrochloric acid (HCl) in capped Teflon vials. Analysis is performed via Inductively Coupled Argon Plasma Atomic Emission Spectrometry (ICP-AES) (simultaneous multielement analysis), Graphite Furnace Atomic Absorption Spectrometry (GF-AAS) (low level analysis) and/or Flame Atomic Absorption Spectrometry (F-AAS). Note that extractable samples are acidified with 0.5 mL HNO3 per 100 mL of sample at time of sample collection. | 0.002 | Y | Y | Y | 10 | 6/12/1997 6:16:00 PM | 11/5/1998 6:04:00 PM | 12/4/1998 6:04:42 PM | 1997-04-02 | 6/3/1998 6:51:46 PM | MPAK |  | Concentration | mg/L | Y | (fl. conc.) |  | Aluminum Dissolved (fl. conc.) | 7429-90-5 | Metals - Dissolved; | Metals - Dissolved; EMS code: AL-D | Chemical | Dissolved | LAB
AL-D | mg/L | Aluminum Dissolved | X085 | LF,LA,HNO3,AA | Samples are lab-filtered (0.45 um), and lab-acidified (HNO3); analysis by AA. | 0.5 | Y | Y | Y | 6533 | 12/12/1996 | 9/18/2015 | 10/28/2015 9:54:16 AM | 1996-04-12 | 6/3/1998 6:51:33 PM | EMS |  | Concentration | mg/L | Y | (fl. conc.) |  | Aluminum Dissolved (fl. conc.) | 7429-90-5 | Metals - Dissolved; | Metals - Dissolved; EMS code: AL-D | Chemical | Dissolved | LAB
AL-D | mg/L | Aluminum Dissolved | X116 | FF,FA,HNO3,Graph.Furn,AA | Samples are field-filtered (0.45 um), field-acidified with HNO3; analysis by GFAA. | 0.0001 | Y | Y | Y | 16046 |  |  | 8/26/2024 2:41:32 PM | 1996-04-12 | 6/3/1998 6:51:33 PM | EMS |  | Concentration | mg/L | Y | (fl. conc.) |  | Aluminum Dissolved (fl. conc.) | 7429-90-5 | Metals - Dissolved; | Metals - Dissolved; EMS code: AL-D | Chemical | Dissolved | LAB
AL-D | mg/L | Aluminum Dissolved | MSEA | Metals in Seawater (undigested) by (CRC) ICP-MS | This method describes the multi-elemental determination of trace elements by Collision/Reaction Cell (CRC) ICP-MS. The method measures ions produced by a radio-frequency inductively coupled plasma. Analyte species originating in a liquid are pneumatically nebulized and the resulting aerosol transported by argon gas into the plasma torch. The ions produced are entrained in the plasma gas and extracted, by means of a differentially pumped vacuum interface, into a mass spectrometer. The ions flow through a Collision/Reaction Cell (CRC) where depending on the mode of analysis either the ions passes through the cell unhindered (no gas - classical ICP-MS mode), or collisions/reactions chemistry takes place between the ions and a neutral gas such as hydrogen or helium. Any polyatomic ions which give interference on the analytes are removed. The ions are then sorted according to their mass-to-charge ratios by a quadrupole mass spectrometer. The ions are then sorted according to their mass-to-charge ratios by a quadrupole mass spectrometer. The ions transmitted through the quadrupole are quantified by a channel electron multiplier and the ion information processed by a data handling system. Although most polyatomic interference is removed, any additional interference must be assessed and valid corrections applied or the data flagged to indicate problems. Interference correction must include compensation for background ions contributed by the plasma gas, reagents, and constituents of the sample matrix. Instrument drift as well as suppressions or enhancements of instrument response caused by the sample matrix must be corrected for by the use of internal standards. | 0.0003 | Y | Y | Y | 4 | 4/12/2023 | 4/29/2024 | 5/16/2024 2:00:14 PM | 2009-09-08 |  | RDGIBSON |  | Concentration | mg/L | Y | (fl. conc.) |  | Aluminum Dissolved (fl. conc.) | 7429-90-5 | Metals - Dissolved; | Metals - Dissolved; EMS code: AL-D | Chemical | Dissolved | LAB
AL-D | mg/L | Aluminum Dissolved | X562 | Dissolved metals: FF, FA, HNO3, HR-ICPMS | Ultra trace metals in water are analyzed by high resolution inductively coupled plasma mass spectrometry (HR-ICPMS), based on US EPA Method 200.8, (Oct. 1999). This procedure is intended for pristine field-filtered acid-preserved water samples. The detection limits (LOR) for this test are based on lab instrumental analysis only, not including filtration. ALS-supplied field filtration equipment does not support these LOR. Therefore, because of the high probability of false positives due to filtration, it is strongly recommended that a filtration blank be analyzed to aid in data interpretation. | 0.0002 | Y | Y | Y | 1264 | 6/8/2015 | 10/19/2023 | 11/9/2023 7:41:27 AM | 2015-10-22 |  | RDGIBSON |  | Concentration | mg/L | Y | (fl. conc.) |  | Aluminum Dissolved (fl. conc.) | 7429-90-5 | Metals - Dissolved; | Metals - Dissolved; EMS code: AL-D | Chemical | Dissolved | LAB
AL-D | mg/L | Aluminum Dissolved | 5138 | AA Analysis | AA Analysis | 0.5 | Y | Y | Y | 12 |  |  | 10/24/1996 9:37:08 AM | 1996-04-12 | 6/3/1998 6:51:33 PM | EMS |  | Concentration | mg/L | Y | (fl. conc.) |  | Aluminum Dissolved (fl. conc.) | 7429-90-5 | Metals - Dissolved; | Metals - Dissolved; EMS code: AL-D | Chemical | Dissolved | LAB
AL-D | mg/L | Aluminum Dissolved | X235 | HCl Dig,AA Analysis | HCl Dig;AA Analysis | 0.005 | Y | Y | Y | 353 |  |  | 11/29/2001 3:36:32 PM | 1996-04-12 | 6/3/1998 6:51:33 PM | EMS |  | Concentration | mg/L | Y | (fl. conc.) |  | Aluminum Dissolved (fl. conc.) | 7429-90-5 | Metals - Dissolved; | Metals - Dissolved; EMS code: AL-D | Chemical | Dissolved | LAB
AL-D | mg/L | Aluminum Dissolved | F068 | HNO3dig,TeflonTube/Autoclave,ICP-MS | This method is applicable to determiniation of dissolved, extractable, and total recoverable metals in ground waters, surface waters, drinking waters, and wastewaters. Dissolved, extractable, and total recoverable sample fractions are acidified to pH <1.5 with analytical grade nitric acid (HNO3) to prevent the metals from adsorbing to the container material or precipitating. Extractable and total recoverable water samples are collected in 250 ml acid washed polyethylene bottles and are preserved with Nitric acid (0.5 ml acid for 100 mLs of sample). All preserved samples are stored at room temperature and can be held for up to 6 months. Unfiltered, un-preserved samples are stored at 4C for no more than 30 days. For dissolved and extractable water fractions, no sample preparation is necessary. Samples are simply poured in autosampler tubes (i.e. 15 ml centrifuge tubes), and placed on the autosampler and analyzed. Samples for total recoverable metals are digested with Nitric acid in special Teflon tubes in an autoclave prior to analyses. | 0.0002 | Y | Y | Y | 1398 | 6/1/1999 1:19:00 PM | 4/17/2015 1:06:00 PM | 8/11/2021 12:04:11 PM | 1999-06-28 |  | NPEPPIN |  | Concentration | mg/L | Y | (fl. conc.) |  | Aluminum Dissolved (fl. conc.) | 7429-90-5 | Metals - Dissolved; | Metals - Dissolved; EMS code: AL-D | Chemical | Dissolved | LAB
AL-D | mg/L | Aluminum Dissolved | X234 | Aqua Reg. Dig,AA | Aqua Reg. Dig;AA | 0.05 | Y | Y | Y | 1 |  |  | 10/24/1996 9:16:04 AM | 1996-04-12 | 6/3/1998 6:51:33 PM | EMS |  | Concentration | mg/L | Y | (fl. conc.) |  | Aluminum Dissolved (fl. conc.) | 7429-90-5 | Metals - Dissolved; | Metals - Dissolved; EMS code: AL-D | Chemical | Dissolved | LAB
AL-D | mg/L | Aluminum Dissolved | X285 | HNO3 Dig:ICP Analysis | HNO3 Dig:ICP Analysis | 0.2 | Y | Y | Y | 660 | 1/30/2002 | 11/13/2002 | 3/17/2004 12:36:33 PM | 1998-05-22 | 6/3/1998 6:51:49 PM | NPEPPIN |  | Concentration | mg/L | Y | (fl. conc.) |  | Aluminum Dissolved (fl. conc.) | 7429-90-5 | Metals - Dissolved; | Metals - Dissolved; EMS code: AL-D | Chemical | Dissolved | LAB
AL-E | mg/L | Aluminum Extractable | X553 | Metals in water, extractable, ICP-SFMS | The whole water sample is collected in a HDPE/Teflon bottle. In a Class 100 workstation it is preserved with concentrated Nitric Acid at concentration of 0.2 percent (v/v). The acid preservative extracts the metals from material in the sample and a decanted portion is introduced directly to an inductively coupled argon plasma-sector field mass spectrometer (ICP-SFMS) by pneumatic nebulization. Each respective element is measured at a specific mass to charge ratio, m/z value expressed in amu (atomic mass units), and at a specified instrument resolution. LR: Low Resolution ~300, MR: Mid Resolution ~4000 and HR: High Resolution ~10000. Ref: Method number 02-2061, National Laboratory for Environmental Testing, CCIW, Environment Canada, Burlington. Requested April 2003. | 1e-06 | Y | Y | Y | 2584 | 5/22/2003 | 9/21/2007 | 4/18/2018 9:28:02 AM | 2004-01-19 |  | UMULLER |  | Concentration | mg/L | Y | (fl. conc.) |  | Aluminum Extractable (fl. conc.) | 7429-90-5 | Metals - Extractable; | Metals - Extractable; EMS code: AL-E | Chemical | Extractable | LAB
AL-E | ug/g | Aluminum Extractable | X230 | Hom,HNO3/HCl,ICP | The sample is dried, ground to pass 100 mesh, and digested with a nitric-hydrochloric acid mixture to solubilize the solid matter and to remove the organic material by oxidation and volatilization. The completed digests are submitted to metals instrumentation personnel for metals analysis by Inductively Coupled Atomic Emission Spectrophotometry (ICP-AES), or hydride analysis by atomic absorption. | 5 | Y | Y | Y | 74 | 12/13/2011 12:02:00 PM | 12/13/2011 12:02:00 PM | 2/9/2012 12:15:56 PM | 1996-04-12 | 6/3/1998 6:51:41 PM | EMS |  | MassPerMass | mg/kg | Y | (mass conc.) |  | Aluminum Extractable (mass conc.) | 7429-90-5 | Metals - Extractable; | Metals - Extractable; EMS code: AL-E | Chemical | Extractable | LAB
AL-E | mg/L | Aluminum Extractable | F040 | Field Acid, lab filt, ICP-AES | Dissolved and extractable water samples are ready for analysis once collected. Total samples are autoclaved with HNO3 and hydrochloric acid (HCI) in capped Teflon vials. Analysis is performed via Inductively Coupled Argon Plasma Atomic Emission Spectrometry (ICP-AES)(simultaneous multielement analysis), Graphite Furnace Atomic Absorption Spectrometry (GF-AAS)(low level analysis) and/or Flame Atomic Absorption Spectrometry (F-AAS). | 0.05 | Y | Y | Y | 2350 | 8/5/1997 6:06:00 PM | 3/14/2008 | 9/11/2019 8:47:41 AM | 1997-05-08 | 6/3/1998 6:51:46 PM | MPAK |  | Concentration | mg/L | Y | (fl. conc.) |  | Aluminum Extractable (fl. conc.) | 7429-90-5 | Metals - Extractable; | Metals - Extractable; EMS code: AL-E | Chemical | Extractable | LAB
AL-E | mg/L | Aluminum Extractable | F069 | SEM extr,6N HCl,ICP/OES | 6.0 N Hydrochloric acid (HCI) is added to an aliquot of the sediment sample. The extract produced from the addition of the acid is then analyzed for simultaneously extracted metals (SEM) using inductively coupled plasma - optical emission spectrophotometry. The detection limits depend on the metal parameters in question and the instrumentation used. | 0.2 | Y | Y | Y | 9 | 3/26/2012 3:01:00 AM | 3/26/2012 3:01:00 AM | 3/28/2012 10:45:29 AM | 1999-07-09 |  | NPEPPIN |  | Concentration | mg/L | Y | (fl. conc.) |  | Aluminum Extractable (fl. conc.) | 7429-90-5 | Metals - Extractable; | Metals - Extractable; EMS code: AL-E | Chemical | Extractable | LAB
AL-L | mg/L | Aluminum Leachable | MEND | MEND (mod) EPA 300.1 (mod) | MEND (mod) EPA 300.1 (mod). Inorganic anions are analyzed by obtaining an extract from a soil sample that has been added in a 3:1 ratio of deionized water to soil; then leached and the extract filtered prior to analysis; which is then analyzed by Ion Chromatography with conductivity and/or UV detection. | 0.005 | Y | Y | Y | 0 |  |  |  | 2024-01-29 |  | RDGIBSON |  | Concentration | mg/L | Y | (fl. conc.) |  | Aluminum Leachable (fl. conc.) | 7429-90-5 | Metals - Extractable; | Metals - Extractable; EMS code: AL-L | Chemical |  | LAB
AL-S | mg/L | Aluminum-Soluble | X170 | ICP | This code is a catch-all for a large variety of ICP methods including water and also solids digested into solutions as an intermediate result. Different procedures will have greatly different detection limits. (a) For air particulates, preliminary results are reported in mg/L and loading results as mg/dm2/d. (b) For Hi Vol samples preliminiary results are same as reported in mg/L and loading results as ug/m3. (c) Note that for Rain samples there is no filtration nor digestion steps. | 0.02 | Y | Y | Y | 405 |  |  | 10/24/1996 4:26:40 AM | 1996-04-12 | 6/3/1998 6:51:41 PM | EMS |  | Concentration | mg/L | Y | (fl. conc.) |  | Aluminum-Soluble (fl. conc.) | 7429-90-5 | Metals - Soluble; | Metals - Soluble; EMS code: AL-S | Chemical | Soluble | LAB
AL-S | % | Aluminum-Soluble | EQ34 | EQUIS (%) | EQUIS (%) | 0 | N | Y | Y | 13 |  |  | 10/22/1996 7:35:43 PM | 1996-04-12 | 6/3/1998 6:51:33 PM | EMS |  | DimensionlessRatio | % | N | (ratio) |  | Aluminum-Soluble (ratio) | 7429-90-5 | Metals - Soluble; | Metals - Soluble; EMS code: AL-S | Chemical | Soluble | LAB
AL-S | mg/L | Aluminum-Soluble | EQ01 | EQUIS (mg/L) | EQUIS (mg/L) | 0.02 | N | Y | Y | 38 |  |  | 10/22/1996 10:26:39 PM | 1996-04-12 | 6/3/1998 6:51:33 PM | EMS |  | Concentration | mg/L | Y | (fl. conc.) |  | Aluminum-Soluble (fl. conc.) | 7429-90-5 | Metals - Soluble; | Metals - Soluble; EMS code: AL-S | Chemical | Soluble | LAB
AL-S | mg/L | Aluminum-Soluble | 5233 | Precip Prep/HGA Analysis | Precip Prep/HGA Analysis | 0.01 | Y | Y | Y | 23 |  |  | 10/23/1996 10:23:23 PM | 1996-04-12 | 6/3/1998 6:51:33 PM | EMS |  | Concentration | mg/L | Y | (fl. conc.) |  | Aluminum-Soluble (fl. conc.) | 7429-90-5 | Metals - Soluble; | Metals - Soluble; EMS code: AL-S | Chemical | Soluble | LAB
AL-T | mg/L | Aluminum Total | 0061 | HNO3 Dig: GFAA Analysis | HNO3 Dig: GFAA Analysis | 0.02 | Y | Y | Y | 88 |  |  | 8/26/2004 11:46:08 AM | 1996-04-12 | 6/3/1998 6:51:34 PM | EMS |  | Concentration | mg/L | Y | (fl. conc.) |  | Aluminum Total (fl. conc.) | 7429-90-5 | Metals - Total; | Metals - Total; EMS code: AL-T | Chemical | Total | LAB
AL-T | mg/L | Aluminum Total | X561 | Total metals: FA, HNO3, HR-ICPMS (No Prep)Prep) | Ultra trace metals in water are analyzed by high resolution inductively coupled plasma mass spectrometry (HR-ICPMS), modified from US EPA 200.8, (Revision 5.5). The detection limits provided can only be met for undigested samples. This procedure is intended for colorless, non-turbid, acid-preserved water samples (i.e. pristine water samples), having turbidity < 1 NTU and no odor. Where turbidity exceeds 1 NTU, and/or the sample is colored and has an odor, results may be biased low compared to true Total Metals concentrations. ALS recommends that turbidity analysis be requested on samples submitted for this test to aid with interpretation of results. | 0.0002 | Y | Y | Y | 2755 | 6/8/2015 | 10/5/2023 | 1/31/2024 2:37:41 PM | 2015-10-22 |  | RDGIBSON |  | Concentration | mg/L | Y | (fl. conc.) |  | Aluminum Total (fl. conc.) | 7429-90-5 | Metals - Total; | Metals - Total; EMS code: AL-T | Chemical | Total | LAB
AL-T | mg/L | Aluminum Total | 0020 | HNO3 Dig:AA Analysis | HNO3 Dig:AA Analysis | 0.02 | Y | Y | Y | 13 | 6/25/2010 3:15:00 PM | 12/1/2010 | 12/9/2010 3:55:26 PM | 1996-04-12 | 6/3/1998 6:51:33 PM | EMS |  | Concentration | mg/L | Y | (fl. conc.) |  | Aluminum Total (fl. conc.) | 7429-90-5 | Metals - Total; | Metals - Total; EMS code: AL-T | Chemical | Total | LAB
AL-T | ug/g | Aluminum Total | FS01 | D&G,Nitric/HCl:ICP | Samples are dried at 60 C, then either sieved to -100mesh (150 um) and rolled to homogenize, or finely ground with a mortar and pestle or a shatter box laboratory mill and rolled (depending on the submitter's requirements). The samples are then weighed into Teflon digestion liners and digested with HNO3 and HCl in a microwave oven. Samples are cooled, volumized, and allowed to settle overnight before being decanted. Analysis is then performed via Inductively Coupled Argon Plasma Atomic Emission Spectrometry (ICP-AES) (Simultaneous multi element analysis). In ICP Spectrometry, the liquid sample is pumped to the nebulizer, where it is convertd into an aerosol. the nebulizer carrier gas flow than transports the wet aerosol past the impact bead, where the large droplets condense out and flow to the waste container. The remaining fine aerosol is carried into the plasma. There, the high temperature desolvates the aerosol and dissociates the sample material into individual atoms and ions, which are excited to emit light at wavelengths characteristic of the elements present. A diffraction grating in the spectrometer diffracts the various wavelengths of light through exit slits to the photo-multiplier tubes, which convert the light energy to electronic signals. | 8 | Y | Y | Y | 2194 | 8/30/1996 1:52:00 PM | 3/22/2002 | 4/1/2004 1:19:34 AM | 1996-11-20 | 6/3/1998 6:51:46 PM | MPAK |  | MassPerMass | mg/kg | Y | (mass conc.) |  | Aluminum Total (mass conc.) | 7429-90-5 | Metals - Total; | Metals - Total; EMS code: AL-T | Chemical | Total | LAB
AL-T | mg/L | Aluminum Total | X349 | FA:HNO3:ICP | Samples are field-acidified with HNO3; HNO3 digestion; analysis by ICP/AES. | 0.2 | Y | Y | Y | 30 |  |  | 10/30/2014 8:31:17 AM | 1998-05-22 | 6/3/1998 6:51:49 PM | NPEPPIN |  | Concentration | mg/L | Y | (fl. conc.) |  | Aluminum Total (fl. conc.) | 7429-90-5 | Metals - Total; | Metals - Total; EMS code: AL-T | Chemical | Total | LAB
AL-T | ug/m3 | Aluminum Total | F088 | Aliquot,HNO3/HCl dig, ICPMS,loading | Aliquot microwave digested with 3:1 HNO3/HCL. Analysis by ICPMS. For HiVol samples preliminary result are reported in mg/L and loading results are reported in ug/m3 |  | Y | Y | Y | 565 | 8/17/2000 10:46:00 AM | 12/14/2006 | 12/30/2006 1:01:35 AM | 2000-08-16 |  | NPEPPIN |  | AirConcentration | µg/m³ | Y | (air conc.) |  | Aluminum Total (air conc.) | 7429-90-5 | Metals - Total; | Metals - Total; EMS code: AL-T | Chemical | Total | LAB
AL-T | ug/g | Aluminum Total | FB04 | Freeze Dry,Hom,Acid Dig,ICP | Sample moisture is determined by weighing samples before and after freeze-drying. Freeze-dried samples are ground in a Braun blender before being weighed into Teflon PFA liners and digested with concentrated HNO3, concentrated HCl, and 50% H2O2 in a microwave oven; which results in a complete sample dissolution. Samples are collected and volumized with Type I water. Analysis is performed via Inductively Coupled Argon Plasma Atomic Emission Spectrometry (ICP-AES) (simultaneous multielement analysis), Graphite Furnace Atomic Absorption Spectrometry (GF-AAS) (low level Cd and Pb analysis), Cold Vapor Atomic Fluorescence Spectrometry (CV-AFS) (Mercury analysis) and/or Flame Atomic Absorption Spectrometry (F-AAS) (K analysis). | 4 | Y | Y | Y | 241 | 8/22/1996 6:11:00 PM | 2/1/2002 2:30:00 PM | 6/23/2003 1:15:27 PM | 1997-02-27 | 6/3/1998 6:51:46 PM | MPAK |  | MassPerMass | mg/kg | Y | (mass conc.) |  | Aluminum Total (mass conc.) | 7429-90-5 | Metals - Total; | Metals - Total; EMS code: AL-T | Chemical | Total | LAB
AL-T | ug/g | Aluminum Total | X551 | Metals in Biota (dry), total, acid digestion, ICP-SFMS | 0.1 gram of dried, homogenized biota is weighed in a microwave TFM vessel, and 8mL of Nitric Acid and 1mL of Hydrogen Peroxide are added. The sample is digested at 200 degrees C for 15 minutes in a closed vessel high pressure microwave oven. The sample is cooled, the volume adjusted to 100mL and introduced to an iductively coupled argon plasma-sector field mass spectrometer (ICP-SFMS) by pneumatic nebulization. Each respective element is measured at a specific mass to charge ratio, m/z value expressed in amu (atomic mass units), and at a specified instrument resolution. LR: Low Resolution ~300, MR: Mid Resolution ~4000 and HR: Hich Resolution ~10000. Ref: Method number 02-2705, National Laboratory for Environmental Testing, CCIW, Environment Canada, Burlington. Requested April 2003. | 0.2 | Y | Y | Y | 49 | 8/11/2015 | 10/30/2023 | 11/28/2023 2:21:33 PM | 2004-01-19 |  | UMULLER |  | MassPerMass | mg/kg | Y | (mass conc.) |  | Aluminum Total (mass conc.) | 7429-90-5 | Metals - Total; | Metals - Total; EMS code: AL-T | Chemical | Total | LAB
AL-T | mg/L | Aluminum Total | X072 | FA,HNO3,Graph.Furnace,AA | Samples are field-acidified with HNO3; HNO3 digestion; analysis by GFAA. | 0.001 | Y | Y | Y | 11415 |  |  | 8/26/2024 2:41:32 PM | 1996-04-12 | 6/3/1998 6:51:34 PM | EMS |  | Concentration | mg/L | Y | (fl. conc.) |  | Aluminum Total (fl. conc.) | 7429-90-5 | Metals - Total; | Metals - Total; EMS code: AL-T | Chemical | Total | LAB
AL-T | mg/L | Aluminum Total | 1349 | Tot Metals in Seawater by CRC ICPMS (BC) | Tot Metals in Seawater by CRC ICPMS (BC) | 0.005 | Y | Y | Y | 98 | 8/16/2019 | 12/5/2021 | 12/13/2021 12:34:16 PM | 2019-09-30 |  | KATFRASE |  | Concentration | mg/L | Y | (fl. conc.) |  | Aluminum Total (fl. conc.) | 7429-90-5 | Metals - Total; | Metals - Total; EMS code: AL-T | Chemical | Total | LAB
AL-T | mg/L | Aluminum Total | F024 | HNO3/HCl,GFAA | Dissolved and extractable water samples are ready for analysis once collected. Total samples are autoclaved with HNO3 and hydrochloric acid (HCl) in capped Teflon vials. Analysis is performed via Inductively Coupled Argon Plasma Atomic Emission Spectrometry (ICP-AES) (simultaneous multielement analysis), Graphite Furnace Atomic Absorption Spectrometry (GF-AAS) (low level analysis) and/or Flame Atomic Absorption Spectrometry (F-AAS). | 0.002 | Y | Y | Y | 10 | 8/9/1996 8:53:00 AM | 7/27/1999 2:56:00 PM | 8/6/1999 6:06:22 PM | 1997-04-02 | 6/3/1998 6:51:46 PM | MPAK |  | Concentration | mg/L | Y | (fl. conc.) |  | Aluminum Total (fl. conc.) | 7429-90-5 | Metals - Total; | Metals - Total; EMS code: AL-T | Chemical | Total | LAB
AL-T | % | Aluminum Total | EQ34 | EQUIS (%) | EQUIS (%) | 0.01 | N | Y | Y | 50 |  |  | 10/22/1996 7:35:00 PM | 1996-04-12 | 6/3/1998 6:51:34 PM | EMS |  | DimensionlessRatio | % | N | (ratio) |  | Aluminum Total (ratio) | 7429-90-5 | Metals - Total; | Metals - Total; EMS code: AL-T | Chemical | Total | LAB
AL-T | mg/dm2/d | Aluminum Total | DUST | HNO3 dig.(microwave or hot block),ICPMS | HNO3 dig.(microwave or hot block);ICPMS |  | Y | Y | Y | 267 | 7/19/2016 1:11:00 PM | 7/31/2024 | 7/31/2024 11:23:12 AM | 2008-02-07 |  | RDGIBSON |  | MassFlux | mg/dm²/d | Y | (mass flux) |  | Aluminum Total (mass flux) | 7429-90-5 | Metals - Total; | Metals - Total; EMS code: AL-T | Chemical | Total | LAB
AL-T | mg/L | Aluminum Total | MWID | Metals in Water by ICPMS (digested) | Metals in Water by ICPMS (digested) Based on EPA 200.8 | 0.01 | Y | Y | Y | 36830 |  |  | 7/31/2024 1:46:41 PM | 2009-12-24 |  | RDGIBSON |  | Concentration | mg/L | Y | (fl. conc.) |  | Aluminum Total (fl. conc.) | 7429-90-5 | Metals - Total; | Metals - Total; EMS code: AL-T | Chemical | Total | LAB
AL-T | mg/L | Aluminum Total | METN | MET-T-NP-CCMS-VA | Ultra trace metals in water are analyzed by CRC ICPMS, based on US EPA Method 6020B (July 2014). The detection limits provided can only be met for undigested samples. This procedure is intended for colorless, non-turbid, acid-preserved water samples (i.e. pristine water samples), having turbidity < 1 NTU and no odor. Where turbidity exceeds 1 NTU, and/or the sample is colored and has an odor, results may be biased low compared to true Total Metals concentrations. ALS recommends that turbidity analysis be requested on samples submitted for this test to aid with interpretation of results. Where turbidity is <1NTU, undigested metals are equivalent to total metals concentrations. | 0.0005 | Y | Y | Y | 1059 | 4/3/2012 | 4/4/2024 | 4/29/2024 4:53:06 PM | 2021-05-03 |  | RDGIBSON |  | Concentration | mg/L | Y | (fl. conc.) |  | Aluminum Total (fl. conc.) | 7429-90-5 | Metals - Total; | Metals - Total; EMS code: AL-T | Chemical | Total | LAB
AL-T | ug/g | Aluminum Total | PB02 | PBM,HNO3/HCl dig,dry60C,100mesh, | Performance Based Method. Samples are dried at 60 C and sieved using a .149 mm (100 mesh) sieve. Solid samples are digested with the combination of nitric and hydrochloric acids with a 1:1 ratio of acids. For environmental purposes, strong acid leaches are used routinely since they have the advantage of speed, low cost, and superio reproducibility. All metals trapped within the refractory matrix are unlikely to have environmental consequences. Instrumental analysis of extracts produced using this digestion method can be performed by a variety of techniques. |  | Y | Y | Y | 10 | 9/13/2001 | 9/13/2001 | 11/6/2002 3:52:32 PM | 2001-07-18 |  | NPEPPIN |  | MassPerMass | mg/kg | Y | (mass conc.) |  | Aluminum Total (mass conc.) | 7429-90-5 | Metals - Total; | Metals - Total; EMS code: AL-T | Chemical | Total | LAB
AL-T | mg/L | Aluminum Total | X555 | Metals in water, total, "in-bottle" digestion, ICP-SFMS | The whole water sample collected in a HDPE/Teflon bottle, is de-capped in a Class 100 workstation, acidified to 2% (v/v) HNO3 and 1% (v/v) HCl and recapped. The sample is "in-bottle" digested in a convection oven, set at 60 degrees C for 16 hours. The sample is cooled and introduced to an inductively coupled argon plasma-sector field mass spectrometer (ICP-SFMS) by pneumatic nebulization. Each respective element is measured at a specific mass to charge ratio, m/z value expressed in amu (atomic mass units), and at resolution. LR: Low Resolution ~300, MR: Mid Resolution ~4000 and HR: High a specified instrument Resolution ~10000. Ref: Method number 02-2003, National Laboratory for Environmental Testing, CCIW, Environment Canada, Burlington. Requested April 2003. | 0.0002 | Y | Y | Y | 16889 | 5/20/2003 | 12/6/2023 2:08:00 AM | 2/14/2024 9:05:32 AM | 2004-01-14 |  | UMULLER |  | Concentration | mg/L | Y | (fl. conc.) |  | Aluminum Total (fl. conc.) | 7429-90-5 | Metals - Total; | Metals - Total; EMS code: AL-T | Chemical | Total | LAB
AL-T | ug/g | Aluminum Total | MTID | MET-DRY-CCMS-VA | Metals in Tissue by CRC ICPMS (DRY) Methodology: This method is conducted following British Columbia Lab Manual method "Metals in Animal Tissue and Vegetation (Biota) - Prescriptive". Tissue samples are homogenized and sub-sampled prior to hotblock digestion with nitric and hydrochloric acids, in combination with addition of hydrogen peroxide. Instrumental analysis is by collision cell inductively coupled plasma - mass spectrometry (modified from EPA Method 6020A). Method Limitation: This method employs a strong acid/peroxide digestion, and is intended to provide a conservative estimate of bio-available metals. Near complete recoveries are achieved for most toxicologically important metals, but elements associated with recalcitrant minerals may be only partially recovered. | 2 | Y | Y | Y | 122 | 8/13/2014 | 9/8/2022 | 10/26/2022 7:14:14 AM | 2015-03-09 | 3/11/2015 4:12:05 PM | RDGIBSON | EMS | MassPerMass | mg/kg | Y | (mass conc.) |  | Aluminum Total (mass conc.) | 7429-90-5 | Metals - Total; | Metals - Total; EMS code: AL-T | Chemical | Total | LAB
AL-T | mg/L | Aluminum Total | X170 | ICP | This code is a catch-all for a large variety of ICP methods including water and also solids digested into solutions as an intermediate result. Different procedures will have greatly different detection limits. (a) For air particulates, preliminary results are reported in mg/L and loading results as mg/dm2/d. (b) For Hi Vol samples preliminiary results are same as reported in mg/L and loading results as ug/m3. (c) Note that for Rain samples there is no filtration nor digestion steps. | 0.02 | Y | Y | Y | 6585 | 3/28/1996 6:03:00 PM | 11/7/2018 | 10/24/2022 10:58:25 AM | 1996-04-12 | 6/3/1998 6:51:34 PM | EMS |  | Concentration | mg/L | Y | (fl. conc.) |  | Aluminum Total (fl. conc.) | 7429-90-5 | Metals - Total; | Metals - Total; EMS code: AL-T | Chemical | Total | LAB
AL-T | ug/g | Aluminum Total | FS82 | PBM,dry60C,10mesh,SALM-ICP-MS | PBM;dry60C;10mesh;SALM-ICP-MS | 100 | Y | Y | Y | 3574 | 5/21/2002 | 7/24/2024 | 7/29/2024 5:17:09 PM | 2002-11-12 | 11/12/2002 10:30:23 AM | UMULLER | UMULLER | MassPerMass | mg/kg | Y | (mass conc.) |  | Aluminum Total (mass conc.) | 7429-90-5 | Metals - Total; | Metals - Total; EMS code: AL-T | Chemical | Total | LAB
AL-T | mg/L | Aluminum Total | 0040 | HNO3 Dig: ICP Analysis | HNO3 Dig: ICP Analysis | 0.02 | Y | Y | Y | 11768 | 5/8/2003 9:40:00 AM | 9/12/2003 3:16:00 PM | 3/23/2010 5:08:08 PM | 1996-04-12 | 6/3/1998 6:51:33 PM | EMS |  | Concentration | mg/L | Y | (fl. conc.) |  | Aluminum Total (fl. conc.) | 7429-90-5 | Metals - Total; | Metals - Total; EMS code: AL-T | Chemical | Total | LAB
AL-T | ug/g wet | Aluminum Total | MTIW | MET-WET-CCMS-MID-VA | Metals in Tissue by CRC ICPMS (WET) Methodology: This method is conducted following British Columbia Lab Manual method "Metals in Animal Tissue and Vegetation (Biota) - Prescriptive". Tissue samples are homogenized and sub-sampled prior to hotblock digestion with nitric and hydrochloric acids, in combination with addition of hydrogen peroxide. Instrumental analysis is by collision cell inductively coupled plasma - mass spectrometry (modified from EPA Method 6020A). Method Limitation: This method employs a strong acid/peroxide digestion, and is intended to provide a conservative estimate of bio-available metals. Near complete recoveries are achieved for most toxicologically important metals, but elements associated with recalcitrant minerals may be only partially recovered. | 1 | Y | Y | Y | 24 | 11/25/2018 | 12/6/2018 | 7/21/2021 2:55:07 PM | 2018-12-24 |  | RDGIBSON |  | MassPerMass | µg/g | Y | (mass conc.) | (wet) | Aluminum Total (wet) (mass conc.) | 7429-90-5 | Metals - Total; | Metals - Total; EMS code: AL-T | Chemical | Total | LAB
AL-T | mg/L | Aluminum Total | MSEA | Metals in Seawater (undigested) by (CRC) ICP-MS | This method describes the multi-elemental determination of trace elements by Collision/Reaction Cell (CRC) ICP-MS. The method measures ions produced by a radio-frequency inductively coupled plasma. Analyte species originating in a liquid are pneumatically nebulized and the resulting aerosol transported by argon gas into the plasma torch. The ions produced are entrained in the plasma gas and extracted, by means of a differentially pumped vacuum interface, into a mass spectrometer. The ions flow through a Collision/Reaction Cell (CRC) where depending on the mode of analysis either the ions passes through the cell unhindered (no gas - classical ICP-MS mode), or collisions/reactions chemistry takes place between the ions and a neutral gas such as hydrogen or helium. Any polyatomic ions which give interference on the analytes are removed. The ions are then sorted according to their mass-to-charge ratios by a quadrupole mass spectrometer. The ions are then sorted according to their mass-to-charge ratios by a quadrupole mass spectrometer. The ions transmitted through the quadrupole are quantified by a channel electron multiplier and the ion information processed by a data handling system. Although most polyatomic interference is removed, any additional interference must be assessed and valid corrections applied or the data flagged to indicate problems. Interference correction must include compensation for background ions contributed by the plasma gas, reagents, and constituents of the sample matrix. Instrument drift as well as suppressions or enhancements of instrument response caused by the sample matrix must be corrected for by the use of internal standards. | 0.0003 | Y | Y | Y | 441 | 8/12/2009 12:08:00 AM | 4/29/2024 | 6/7/2024 10:03:16 AM | 2009-09-08 |  | RDGIBSON |  | Concentration | mg/L | Y | (fl. conc.) |  | Aluminum Total (fl. conc.) | 7429-90-5 | Metals - Total; | Metals - Total; EMS code: AL-T | Chemical | Total | LAB
AL-T | mg/L | Aluminum Total | X028 | FA,HNO3,AA Analysis | FA;HNO3;AA Analysis | 0.5 | Y | Y | Y | 47 |  |  | 10/11/2001 10:17:22 AM | 1996-04-12 | 6/3/1998 6:51:34 PM | EMS |  | Concentration | mg/L | Y | (fl. conc.) |  | Aluminum Total (fl. conc.) | 7429-90-5 | Metals - Total; | Metals - Total; EMS code: AL-T | Chemical | Total | LAB
AL-T | mg/L | Aluminum Total | F033 | Acid Dig,Centrifuge,ICPMS | For the determination of total recoverable elements in water or waste water, take 100 ml aliquot from a well mixed, acid preserved, sample containing not more than 0.25% (w/v) total solids and transfer to a 250 ml Griffin beaker. (If total solids are greater than 0.25% reduce the size of the aliquot by a proportionate amount). Add 1 ml of conc. nitric acid and 0.5 ml conc. hydrochloric acid. Heat on a hot plate until the volume has been reduced to 15-20ml, ensuring that the sample does not boil. Allow to cool and transfer to a 50 ml volumetric flask. Dilute to volume with ASTM type I water. Centrifuge the sample or allow to stand overnight to separate insoluble material. Prior analysis, pipette 20 ml into a 50 ml volumetric flask and dilute to volume with ASTM type I water. Analysis should be performed within seventy two hours of the completed preparation. Note that different laboratories may have very different field preservation, digestion procedures and detection limits for this method, so these should be documented. | 0.001 | Y | Y | Y | 37098 | 10/15/1996 9:26:00 AM | 8/22/2024 3:24:00 AM | 8/25/2024 9:47:36 PM | 1997-01-16 | 6/3/1998 6:51:46 PM | MPAK |  | Concentration | mg/L | Y | (fl. conc.) |  | Aluminum Total (fl. conc.) | 7429-90-5 | Metals - Total; | Metals - Total; EMS code: AL-T | Chemical | Total | LAB
AL-T | mg/L | Aluminum Total | 0200 | 20%A.R.Dig:AA Anal. | 20%A.R.Dig:AA Anal. | 0.5 | Y | Y | Y | 39 |  |  | 10/24/1996 9:37:08 AM | 1996-04-12 | 6/3/1998 6:51:41 PM | EMS |  | Concentration | mg/L | Y | (fl. conc.) |  | Aluminum Total (fl. conc.) | 7429-90-5 | Metals - Total; | Metals - Total; EMS code: AL-T | Chemical | Total | LAB
AL-T | mg/L | Aluminum Total | X145 | Dig, ICP | Dig; ICP | 0.2 | Y | Y | Y | 320 |  |  | 11/29/2001 3:36:33 PM | 1996-04-12 | 6/3/1998 6:51:41 PM | EMS |  | Concentration | mg/L | Y | (fl. conc.) |  | Aluminum Total (fl. conc.) | 7429-90-5 | Metals - Total; | Metals - Total; EMS code: AL-T | Chemical | Total | LAB
AL-T | % | Aluminum Total | MS61 | ALS Minerals ME-MS61 | ALS Minerals ME-MS61. A prepared sample is digested with HNO3; HF; and HClO4 in a graphite heating block to near dryness. After cooling; the sample is boiled in diluted HCl. Bulk to volume with deionized water; mixed and analyzed by ICP-AES and ICP-MS for all elements. | 0.01 | Y | Y | Y | 0 |  |  |  | 2024-01-29 |  | EMS |  | DimensionlessRatio | % | N | (ratio) |  | Aluminum Total (ratio) | 7429-90-5 | Metals - Total; | Metals - Total; EMS code: AL-T | Chemical | Total | LAB
AL-T | ug/g | Aluminum Total | ICPM | Metals in Soil by CRC ICPMS | This analysis is carried out using procedures from CSR Analytical Method: "Strong Acid Leachable Metals (SALM) in Soil", BC Ministry of Environment, 26 June 2009, and procedures adapted from EPA Method 200.2. The sample is manually homogenized, dried at 60 degrees Celsius, sieved through a 63 um (230 mesh) sieve, and a representative subsample of the dry material is weighed. The sample is then digested at 95 degrees Celsius for 2 hours by block digester using concentrated nitric and hydrochloric acids. Instrumental analysis is by collision cell inductively coupled plasma - mass spectrometry (modifed from EPA Method 6020A). Method Limitation: This method is not a total digestion technique. It is a very strong acid digestion that is intended to dissolve those metals that may be environmentally available. By design, elements bound in silicate structures are not normally dissolved by this procedure as they are not usually mobile in the environment. | 50 | Y | Y | Y | 103 | 9/19/2017 | 9/24/2023 | 1/12/2024 11:16:52 AM | 2018-01-09 |  | JKROGH |  | MassPerMass | mg/kg | Y | (mass conc.) |  | Aluminum Total (mass conc.) | 7429-90-5 | Metals - Total; | Metals - Total; EMS code: AL-T | Chemical | Total | LAB
AL-T | mg/L | Aluminum Total | 0042 | GI HNO3 Dig: ICP | Sample aliquot taken from unfiltered unpreserved General Ions bottle; HNO3 digestion; ICP analysis. | 0.06 | Y | Y | Y | 6807 | 9/20/1996 | 7/5/2019 3:21:00 PM | 7/12/2019 6:49:13 PM | 1996-04-12 | 6/3/1998 6:51:33 PM | EMS |  | Concentration | mg/L | Y | (fl. conc.) |  | Aluminum Total (fl. conc.) | 7429-90-5 | Metals - Total; | Metals - Total; EMS code: AL-T | Chemical | Total | LAB
AL-T | mg/L | Aluminum Total | F038 | HNO3/HCl dig,ICP-AES | Dissolved and extractable water samples are ready for analysis once collected. Total samples are autoclaved with HNO3 and hydrochloric acid (HCI) in capped Teflon vials. Analysis is performed via Inductively Coupled Argon Plasma Atomic Emission Spectrometry (ICP-AES)(simultaneous multielement analysis), Graphite Furnace Atomic Absorption Spectrometry (GF-AAS) (low level analysis) and/or Flame Atomic Absorption Spectrometry (F-AAS). | 0.06 | Y | Y | Y | 5401 | 8/1/1996 6:04:00 PM | 5/6/2008 | 9/11/2019 8:47:42 AM | 1997-05-08 | 6/3/1998 6:51:46 PM | MPAK |  | Concentration | mg/L | Y | (fl. conc.) |  | Aluminum Total (fl. conc.) | 7429-90-5 | Metals - Total; | Metals - Total; EMS code: AL-T | Chemical | Total | LAB
AL-T | mg/L | Aluminum Total | 0060 | HNO3 Dig: HGA Analysis | HNO3 Dig: HGA Analysis | 0.01 | Y | Y | Y | 1458 |  |  | 10/24/1996 12:33:07 PM | 1996-04-12 | 6/3/1998 6:51:33 PM | EMS |  | Concentration | mg/L | Y | (fl. conc.) |  | Aluminum Total (fl. conc.) | 7429-90-5 | Metals - Total; | Metals - Total; EMS code: AL-T | Chemical | Total | LAB
AL-T | mg/L | Aluminum Total | METT | MET-T-NP-BCMOE-MS-VA | Ultra trace metals in water are analyzed by CRC ICPMS, based on US EPA Method 6020B (July 2014). The detection limits provided can only be met for undigested samples. This procedure is intended for colorless, non-turbid, acid-preserved water samples (i.e. pristine water samples), having turbidity < 1 NTU and no odor. Where turbidity exceeds 1 NTU, and/or the sample is colored and has an odor, results may be biased low compared to true Total Metals concentrations. ALS recommends that turbidity analysis be requested on samples submitted for this test to aid with interpretation of results. Where turbidity is <1NTU, undigested metals are equivalent to total metals concentrations. BC Ministry of Environment version. | 0.0005 | Y | Y | Y | 2589 | 6/5/2019 | 8/18/2024 | 8/26/2024 1:03:06 PM | 2019-08-20 |  | KATFRASE |  | Concentration | mg/L | Y | (fl. conc.) |  | Aluminum Total (fl. conc.) | 7429-90-5 | Metals - Total; | Metals - Total; EMS code: AL-T | Chemical | Total | LAB
AL-T | mg/L | Aluminum Total | EQ01 | EQUIS (mg/L) | EQUIS (mg/L) | 0.01 | N | Y | Y | 2143 |  |  | 10/23/1996 12:08:27 AM | 1996-04-12 | 6/3/1998 6:51:34 PM | EMS |  | Concentration | mg/L | Y | (fl. conc.) |  | Aluminum Total (fl. conc.) | 7429-90-5 | Metals - Total; | Metals - Total; EMS code: AL-T | Chemical | Total | LAB
AL-T | ug/g | Aluminum Total | 0192 | Hom, Nitric/Perchlor,ICP | Sample already in a dry state so does not require drying before homogenization and nitric perchlorid digestion followed by ICP analysis. | 2 | Y | Y | Y | 708 |  |  | 10/24/1996 12:50:25 PM | 1996-04-12 | 6/3/1998 6:51:34 PM | EMS |  | MassPerMass | mg/kg | Y | (mass conc.) |  | Aluminum Total (mass conc.) | 7429-90-5 | Metals - Total; | Metals - Total; EMS code: AL-T | Chemical | Total | LAB
AL-T | ug/g wet | Aluminum Total | X560 | Metals in Biota (wet), total, acid digestion, ICP-MS | Five grams of wet, homogenized biota is weighed in a centrifuge tube, and 2.5 mL of Nitric Acid and 2.5 mL of Hydrochloric acid are added. The sample is digested at room temperature for one hour and then digested at 90-95 degrees C in a water bath or heat block for one more hour. The sample is cooled and the volume is made to 50 mL with DI. The digestate is then diluted at ten fold and introduced to a collision/reaction cell inductively coupled argon plasma mass spectrometer (CRC ICPMS) by pneumatic nebulization. Each respective element is measured at a specific mass to charge ratio, m/z value expressed in amu (atomic mass units). - Maxxam Analytics (BC). Added to EMS Dec 10, 2008. | 1 | Y | Y | Y | 541 | 11/15/2008 4:00:00 AM | 10/30/2023 | 11/28/2023 2:21:33 PM | 2008-12-11 | 3/11/2015 4:17:10 PM | RDGIBSON | EMS | MassPerMass | µg/g | Y | (mass conc.) | (wet) | Aluminum Total (wet) (mass conc.) | 7429-90-5 | Metals - Total; | Metals - Total; EMS code: AL-T | Chemical | Total | LAB
AL-T | mg/L | Aluminum Total | X234 | Aqua Reg. Dig,AA | Aqua Reg. Dig;AA | 0.005 | Y | Y | Y | 254 |  |  | 10/24/1996 12:12:10 PM | 1996-04-12 | 6/3/1998 6:51:41 PM | EMS |  | Concentration | mg/L | Y | (fl. conc.) |  | Aluminum Total (fl. conc.) | 7429-90-5 | Metals - Total; | Metals - Total; EMS code: AL-T | Chemical | Total | LAB
AL-T | mg/L | Aluminum Total | 0050 | Graphite Furnace AA | Graphite Furnace AA | 0.01 | Y | Y | Y | 23 |  |  | 10/24/1996 5:49:21 AM | 1996-04-12 | 6/3/1998 6:51:41 PM | EMS |  | Concentration | mg/L | Y | (fl. conc.) |  | Aluminum Total (fl. conc.) | 7429-90-5 | Metals - Total; | Metals - Total; EMS code: AL-T | Chemical | Total | LAB
AL-T | mg/L | Aluminum Total | F082 | Aliquot,HNO3/HCL dig,ICPMS | Aliquot microwave digested with 3:1 HNO3/HCL. Analysis by ICPMS. For HiVol samples preliminary result are reported in mg/L and loading results are reported in ug/m3 | 0.005 | Y | Y | Y | 28072 | 7/26/2005 7:04:00 AM | 8/23/2024 | 8/26/2024 4:53:05 PM | 2000-06-13 | 8/15/2019 11:04:01 AM | NPEPPIN | KATFRASE | Concentration | mg/L | Y | (fl. conc.) |  | Aluminum Total (fl. conc.) | 7429-90-5 | Metals - Total; | Metals - Total; EMS code: AL-T | Chemical | Total | LAB
AL-T | % | Aluminum Total | MS41 | ALS Minerals ME-MS41 | ALS Minerals ME-MS41. A prepared sample is digested with aqua regia in a graphite heating block. After cooling; the resulting solution is diluted to with deionized water; mixed and analyzed by ICP-AES. Samples are then analyzed by ICP-MS for the remaining suite of elements. | 0.01 | Y | Y | Y | 0 |  |  |  | 2024-01-29 |  | EMS |  | DimensionlessRatio | % | N | (ratio) |  | Aluminum Total (ratio) | 7429-90-5 | Metals - Total; | Metals - Total; EMS code: AL-T | Chemical | Total | LAB
AL-T | ug/g | Aluminum Total | 0190 | D&G:N/P Dig:ICP Anal. | Sample is dried and grind-mixed before nitric perchloric digestion followed by ICP analysis. | 2 | Y | Y | Y | 949 |  |  | 10/24/1996 3:06:16 PM | 1996-04-12 | 6/3/1998 6:51:41 PM | EMS |  | MassPerMass | mg/kg | Y | (mass conc.) |  | Aluminum Total (mass conc.) | 7429-90-5 | Metals - Total; | Metals - Total; EMS code: AL-T | Chemical | Total | LAB
AL-T | ug/m3 | Aluminum Total | 6040 | Aqua Regia Dig:ICP | Aqua regia digestion is used to bring the metals into solution prior to ICP analysis. For Hi-Vol samples, preliminary results are reported as mg/L and24 hr loadings as ug/m3 | 0.005 | Y | Y | Y | 437 | 2/13/2002 | 2/9/2005 | 2/23/2005 1:13:46 AM | 2004-02-10 |  | UMULLER |  | AirConcentration | µg/m³ | Y | (air conc.) |  | Aluminum Total (air conc.) | 7429-90-5 | Metals - Total; | Metals - Total; EMS code: AL-T | Chemical | Total | LAB
AL-T | mg/L | Aluminum Total | F068 | HNO3dig,TeflonTube/Autoclave,ICP-MS | This method is applicable to determiniation of dissolved, extractable, and total recoverable metals in ground waters, surface waters, drinking waters, and wastewaters. Dissolved, extractable, and total recoverable sample fractions are acidified to pH <1.5 with analytical grade nitric acid (HNO3) to prevent the metals from adsorbing to the container material or precipitating. Extractable and total recoverable water samples are collected in 250 ml acid washed polyethylene bottles and are preserved with Nitric acid (0.5 ml acid for 100 mLs of sample). All preserved samples are stored at room temperature and can be held for up to 6 months. Unfiltered, un-preserved samples are stored at 4C for no more than 30 days. For dissolved and extractable water fractions, no sample preparation is necessary. Samples are simply poured in autosampler tubes (i.e. 15 ml centrifuge tubes), and placed on the autosampler and analyzed. Samples for total recoverable metals are digested with Nitric acid in special Teflon tubes in an autoclave prior to analyses. | 0.0002 | Y | Y | Y | 3719 | 6/1/1999 1:23:00 PM | 7/4/2008 9:48:00 AM | 7/30/2024 6:32:12 AM | 1999-06-28 |  | NPEPPIN |  | Concentration | mg/L | Y | (fl. conc.) |  | Aluminum Total (fl. conc.) | 7429-90-5 | Metals - Total; | Metals - Total; EMS code: AL-T | Chemical | Total | LAB
AL-T | mg/L | Aluminum Total | F121 | Acid,ICPMS | Acidified, no digestion, ICPMS | 0.0002 | Y | Y | Y | 8443 | 3/12/2010 7:57:00 PM | 8/15/2024 9:10:00 PM | 8/26/2024 8:06:03 AM | 2002-07-17 | 11/27/2012 4:31:21 PM | UMULLER | RDGIBSON | Concentration | mg/L | Y | (fl. conc.) |  | Aluminum Total (fl. conc.) | 7429-90-5 | Metals - Total; | Metals - Total; EMS code: AL-T | Chemical | Total | LAB
AL-T | ug/g wet | Aluminum Total | X552 | Metals in Biota (wet), total, acid digestion, ICP-SFMS | One gram of wet, homogenized biota is weighed in a microwave TFM vessel, and 8mL of Nitric Acid and 1mL of Hydrogen Peroxide are added. The sample is digested at 200 degrees C for 15 minutes in a closed vessel high pressure microwave oven. The sample is cooled, the volume adjusted to 100mL and introduced to an inductively coupled argon plasma-sector field mass spectrometer (ICP-SFMS) by pneumatic nebulization. Each respective element is measured at a specific mass to charge ratio, m/z value expressed in amu (atomic mass units), and at a specified instrument resolution. LR: Low Resolution ~300, MR: Mid Resolution ~4000. Ref: 02-2704, National Laboratory for Environmental Testing, CCIW, Environment Canada, Burlington. Requested April 2003. | 0.02 | Y | Y | Y | 37 | 8/9/2016 | 10/16/2023 | 11/28/2023 2:21:32 PM | 2004-01-19 | 4/19/2021 1:22:00 PM | UMULLER | EMS | MassPerMass | µg/g | Y | (mass conc.) | (wet) | Aluminum Total (wet) (mass conc.) | 7429-90-5 | Metals - Total; | Metals - Total; EMS code: AL-T | Chemical | Total | LAB
AL-T | mg/L | Aluminum Total | X285 | HNO3 Dig:ICP Analysis | HNO3 Dig:ICP Analysis | 0.2 | Y | Y | Y | 726 | 8/10/1993 | 10/5/1993 | 4/30/2004 1:54:38 PM | 1996-04-12 | 6/3/1998 6:51:34 PM | EMS |  | Concentration | mg/L | Y | (fl. conc.) |  | Aluminum Total (fl. conc.) | 7429-90-5 | Metals - Total; | Metals - Total; EMS code: AL-T | Chemical | Total | LAB
AL-T | ug/g | Aluminum Total | FS89 | Homo,HNO3 dig,ICP-AES | Homo;HNO3 dig;ICP-AES | 8 | Y | Y | Y | 452 | 1/5/2001 | 10/10/2007 | 5/24/2008 1:15:02 AM | 2001-03-08 |  | NPEPPIN |  | MassPerMass | mg/kg | Y | (mass conc.) |  | Aluminum Total (mass conc.) | 7429-90-5 | Metals - Total; | Metals - Total; EMS code: AL-T | Chemical | Total | LAB
AL-T | mg/L | Aluminum Total | F048 | Acid dig,conc'n,nebul.,ICP-OES | This method provides an automated procedure for the Inductively Coupled Argon Plasma Optical Emission Spectrophotometric (ICP-OES) determination and quantitation of trace amounts of metals in surface, ground and precipitation waters after a manual acid digestion and concentration procedure and ultrasonic nebulisation. Organically-bound metals and metal complexes are converted to metal ions by digestion with nitric acid [method details says HCl]. To increase the signal detected by the ICP, the samples are concentrated 10 times. The liquid sample is pumped onto the piezoelectric transducer, where the ultrasonic action efficiently converts it to a fine, dense aerosol. The nebulizer carrier gas flow then transports the wet aerosol through a heated U-tube, where the solvent is vaporized. Subsequently, most of the solvent vapour is removed from the aerosol stream by passing the stream through a chilled condenser. The result is a dry, analyte-laden aerosol which is introduced to the plasma. There, the high temperature desolvates the aerosol and dissociates the sample material into individual atoms and ions, which are excited to emit light at wave lengths characteristic of the elements present. Dissolved samples are field filtered through 0.45 micron cellulose acetate filter prior to acid preservation. Samples for extractable analysis are acidified with 2 mL concentrated HNO3 per litre and allowed to stand a minimum of 24 hours before analysis. Clients providing these samples are advised to preserve samples at the collection site. | 0.002 | Y | Y | Y | 6161 | 7/4/1997 2:35:00 PM | 7/17/2012 7:03:00 AM | 9/11/2019 8:46:38 AM | 1997-10-20 | 6/3/1998 6:51:47 PM | NPEPPIN |  | Concentration | mg/L | Y | (fl. conc.) |  | Aluminum Total (fl. conc.) | 7429-90-5 | Metals - Total; | Metals - Total; EMS code: AL-T | Chemical | Total | LAB
AL-T | mg/dm2/d | Aluminum Total | EQ15 | EQUIS (mg/dm2/d) | EQUIS (mg/dm2/d) | 0 | N | Y | Y | 14 |  |  | 10/22/1996 7:35:14 PM | 1996-04-12 | 6/3/1998 6:51:41 PM | EMS |  | MassFlux | mg/dm²/d | Y | (mass flux) |  | Aluminum Total (mass flux) | 7429-90-5 | Metals - Total; | Metals - Total; EMS code: AL-T | Chemical | Total | LAB
AL-T | ug/g wet | Aluminum Total | BPSY | MET-WET-BPSY-CCMS-VA | Metals in Animal Tissue and Vegetation (Biota) - Prescriptive. EPA 200.3/6020B This method is conducted following British Columbia Lab Manual method "Metals in Animal Tissue and Vegetation (Biota) - Prescriptive". Tissue samples are digested with nitric and hydrochloric acids; in combination with addition of hydrogen peroxide. Instrumental analysis is by collision cell inductively coupled plasma - mass spectrometry (modified from EPA Method 6020B). | 1 | Y | Y | Y | 2 | 11/7/2019 | 11/7/2019 | 6/10/2021 1:57:10 PM | 2021-06-04 |  | EMS |  | MassPerMass | µg/g | Y | (mass conc.) | (wet) | Aluminum Total (wet) (mass conc.) | 7429-90-5 | Metals - Total; | Metals - Total; EMS code: AL-T | Chemical | Total | LAB
AL-T | ug/g | Aluminum Total | EQ02 | EQUIS (ug/g) | EQUIS (ug/g) | 2 | N | Y | Y | 760 |  |  | 10/23/1996 12:08:28 AM | 1996-04-12 | 6/3/1998 6:51:34 PM | EMS |  | MassPerMass | mg/kg | Y | (mass conc.) |  | Aluminum Total (mass conc.) | 7429-90-5 | Metals - Total; | Metals - Total; EMS code: AL-T | Chemical | Total | LAB
AL-T | % | Aluminum Total | 6041 | Aqua Regia by ICPMS | These were sediment/soils digested with Aqua-Regia (Nitric Acid & Hydrochloric Acids) and then analysis by ICPMS. Results would be reported in mg/kg and/or %. | 0.01 | Y | Y | Y | 92 | 10/13/2019 | 7/27/2024 | 7/29/2024 9:13:08 AM | 2019-09-06 |  | KATFRASE |  | DimensionlessRatio | % | N | (ratio) |  | Aluminum Total (ratio) | 7429-90-5 | Metals - Total; | Metals - Total; EMS code: AL-T | Chemical | Total | LAB
AL-T | ug/g | Aluminum Total | X230 | Hom,HNO3/HCl,ICP | The sample is dried, ground to pass 100 mesh, and digested with a nitric-hydrochloric acid mixture to solubilize the solid matter and to remove the organic material by oxidation and volatilization. The completed digests are submitted to metals instrumentation personnel for metals analysis by Inductively Coupled Atomic Emission Spectrophotometry (ICP-AES), or hydride analysis by atomic absorption. | 2 | Y | Y | Y | 181 | 2/12/2002 | 12/30/2005 | 2/3/2006 1:13:12 AM | 1996-04-12 | 6/3/1998 6:51:34 PM | EMS |  | MassPerMass | mg/kg | Y | (mass conc.) |  | Aluminum Total (mass conc.) | 7429-90-5 | Metals - Total; | Metals - Total; EMS code: AL-T | Chemical | Total | LAB
AL-T | ug/g | Aluminum Total | 0196 | D&G: LTA/HCl: ICP | D&G: LTA/HCl: ICP | 2 | Y | Y | Y | 27 |  |  | 10/24/1996 2:08:12 PM | 1996-04-12 | 6/3/1998 6:51:34 PM | EMS |  | MassPerMass | mg/kg | Y | (mass conc.) |  | Aluminum Total (mass conc.) | 7429-90-5 | Metals - Total; | Metals - Total; EMS code: AL-T | Chemical | Total | LAB
AL2-O3 | % | Aluminum (as Al2O3) | ICP06 | ALS Minerals ME-ICP06 | ALS Minerals ME-ICP06. A prepared sample is fused with a lithium tetraborate/lithium metaborate flux and poured into a graphite crucible and fused in a furnace at 1000°C. The resulting melt is then cooled and dissolved in 100 mL of 4% HNO3/2% HCl followed by ICP-AES. | 0.01 | Y | Y | Y | 0 |  |  |  | 2024-01-29 | 1/29/2024 11:34:32 AM | RDGIBSON | RDGIBSON | DimensionlessRatio | % | N | (ratio) |  | Aluminum (as Al2O3) (ratio) |  | Metals; | Metals; EMS code: AL2-O3 | Chemical |  | LAB
ALBU | mg/L | Albuterol / Salbutamol CAS: 18559-94-9 | XMSP | HPLC and Tri. Quadpole MS/MS | EPA 1694 An aliquot of filtered sample containing 5% organic is injected directly. Each compound is separated by reversed phase HPLC. Analyzed by Electro Spray Ionization MS/MS detection using Triple Quadrupole MS/MS detector. Quantiitation with internal or external standard. |  | Y | Y | Y | 12 | 5/8/2020 | 6/4/2024 | 8/21/2024 3:32:03 PM | 2022-05-04 |  | EMS |  | Concentration | mg/L | Y | (fl. conc.) |  | Albuterol / Salbutamol (fl. conc.) | 18559-94-9 | Pharmaceuticals and Personal Care Products (PPCP); | Pharmaceuticals and Personal Care Products (PPCP); EMS code: ALBU | Chemical |  | LAB
AMFG | mg | Amphipod;FW;Growth | F117 | H.Azteca,sed,14dy growth | H. Azteca;sed;14 day growth bioassay |  | Y | Y | Y | 1 | 10/27/2000 11:58:00 AM | 10/27/2000 11:58:00 AM | 6/25/2002 11:46:48 AM | 2001-06-29 |  | NPEPPIN |  | Mass | mg | Y | (mass) |  | Amphipod;FW;Growth (mass) |  |  | EMS code: AMFG | Chemical |  | LAB
AMFS | % | Amphipod;FW;Survival | F118 | H.Azteca,sed,14dy surv. | H/. Azteca;sed;14 day survival bioassay |  | Y | Y | Y | 1 | 10/27/2000 11:58:00 AM | 10/27/2000 11:58:00 AM | 6/25/2002 11:46:49 AM | 2001-06-29 |  | NPEPPIN |  | DimensionlessRatio | % | N | (ratio) |  | Amphipod;FW;Survival (ratio) |  |  | EMS code: AMFS | Chemical |  | LAB
AMIT | mg/L | Amitriptyline CAS: 50-48-6 | XMSP | HPLC and Tri. Quadpole MS/MS | EPA 1694 An aliquot of filtered sample containing 5% organic is injected directly. Each compound is separated by reversed phase HPLC. Analyzed by Electro Spray Ionization MS/MS detection using Triple Quadrupole MS/MS detector. Quantiitation with internal or external standard. |  | Y | Y | Y | 12 | 5/8/2020 | 6/4/2024 | 8/21/2024 3:32:03 PM | 2022-05-04 |  | EMS |  | Concentration | mg/L | Y | (fl. conc.) |  | Amitriptyline (fl. conc.) | 50-48-6 | Pharmaceuticals and Personal Care Products (PPCP); | Pharmaceuticals and Personal Care Products (PPCP); EMS code: AMIT | Chemical |  | LAB
AML | mg/L | Amlodipine CAS: 88150-42-9 | XMSP | HPLC and Tri. Quadpole MS/MS | EPA 1694 An aliquot of filtered sample containing 5% organic is injected directly. Each compound is separated by reversed phase HPLC. Analyzed by Electro Spray Ionization MS/MS detection using Triple Quadrupole MS/MS detector. Quantiitation with internal or external standard. |  | Y | Y | Y | 12 | 5/8/2020 | 6/4/2024 | 8/21/2024 3:32:03 PM | 2022-05-04 |  | EMS |  | Concentration | mg/L | Y | (fl. conc.) |  | Amlodipine (fl. conc.) | 88150-42-9 | Pharmaceuticals and Personal Care Products (PPCP); | Pharmaceuticals and Personal Care Products (PPCP); EMS code: AML | Chemical |  | LAB
AMP | mg/L | Ampicillin CAS: 69-53-4 | XMSP | HPLC and Tri. Quadpole MS/MS | EPA 1694 An aliquot of filtered sample containing 5% organic is injected directly. Each compound is separated by reversed phase HPLC. Analyzed by Electro Spray Ionization MS/MS detection using Triple Quadrupole MS/MS detector. Quantiitation with internal or external standard. |  | Y | Y | Y | 12 | 5/8/2020 | 6/4/2024 | 8/21/2024 3:32:03 PM | 2022-05-04 |  | EMS |  | Concentration | mg/L | Y | (fl. conc.) |  | Ampicillin (fl. conc.) | 69-53-4 | Pharmaceuticals and Personal Care Products (PPCP); | Pharmaceuticals and Personal Care Products (PPCP); EMS code: AMP | Chemical |  | LAB
AMPH | mg/L | Amphetamine CAS: 300-62-9 | XMSP | HPLC and Tri. Quadpole MS/MS | EPA 1694 An aliquot of filtered sample containing 5% organic is injected directly. Each compound is separated by reversed phase HPLC. Analyzed by Electro Spray Ionization MS/MS detection using Triple Quadrupole MS/MS detector. Quantiitation with internal or external standard. |  | Y | Y | Y | 12 | 5/8/2020 | 6/4/2024 | 8/21/2024 3:32:03 PM | 2022-05-04 |  | EMS |  | Concentration | mg/L | Y | (fl. conc.) |  | Amphetamine(fl. conc.) | 300-62-9 | Pharmaceuticals and Personal Care Products (PPCP); | Pharmaceuticals and Personal Care Products (PPCP); EMS code: AMPH | Chemical |  | LAB
AN-N | ug/g | Plant Available Nitrate reported as Nitrogen | AN-N | NO3-AVAIL-SK | Alberta Ag / APHA 4500 NO3F - Available Nitrate-N | 2 | Y | Y | Y | 14 | 10/31/2022 | 10/31/2023 | 11/23/2023 7:26:23 AM | 2020-01-13 |  | RDGIBSON |  | MassPerMass | mg/kg | Y | (mass conc.) |  | Plant Available Nitrate reported as Nitrogen (mass conc.) |  | Nutrients; | Nutrients; EMS code: AN-N | Chemical |  | LAB
ANDA | mg/L | Atrazine + n-dealkylated metabolites | E660 | EPA 8270E (mod) | EPA 8270E (mod) Miscellaneous Pesticides by GC-MS | 0.0002 | Y | Y | Y | 68 | 12/7/2021 | 9/12/2022 | 12/8/2022 12:57:05 PM | 2021-12-09 | 12/9/2021 3:22:12 PM | EMS | EMS | Concentration | mg/L | Y | (fl. conc.) |  | Atrazine + n-dealkylated metabolites (fl. conc.) | 1912-24-9, 6190-65-4, 1007-28-9 | Pesticides; | Pesticides; EMS code: ANDA | Chemical |  | LAB
ANDA | ug/g | Atrazine + n-dealkylated metabolites | ZN03 | DCM+Ac.Extr:GC/MS | DCM+Ac.Extr:GC/MS EPA 8270E (mod) | 0.071 | Y | Y | Y | 14 | 1/25/2022 | 9/8/2022 | 12/15/2022 10:50:07 AM | 2022-06-13 |  | EMS |  | MassPerMass | mg/kg | Y | (mass conc.) |  | Atrazine + n-dealkylated metabolites (mass conc.) | 1912-24-9, 6190-65-4, 1007-28-9 | Pesticides; | Pesticides; EMS code: ANDA | Chemical |  | LAB
AOX- | kg/adt | Adsorbable Organic Halides | CAL5 | Calculation (kg/adt) | Calculation (kg/adt) | 1 | Y | Y | Y | 8487 | 12/2/1996 | 7/1/2024 | 8/22/2024 1:49:14 PM | 1997-02-04 | 6/3/1998 6:51:46 PM | MPAK |  | PerAirDriedTonne | kg/adt | Y | (/adt) |  | Adsorbable Organic Halides (/adt) |  | Aggregate Organics; | Aggregate Organics; EMS code: AOX- | Chemical |  | LAB
AOX- | mg/L | Adsorbable Organic Halides | X250 | HNO3/Na2SO3,TOX Analyzer | HNO3/Na2SO3;TOX Analyzer | 2.5 | N | Y | Y | 183 |  |  | 10/24/1996 11:56:43 AM | 1996-04-12 | 6/3/1998 6:51:33 PM | EMS |  | Concentration | mg/L | Y | (fl. conc.) |  | Adsorbable Organic Halides (fl. conc.) |  | Aggregate Organics; | Aggregate Organics; EMS code: AOX- | Chemical |  | LAB
AOX- | mg/L | Adsorbable Organic Halides | X311 | Carbon adsorp, Dohrmann:TOX Analyzer | An appropriate sample aliquot is passed through two granular activated carbon (GAC) columns in series where the organo-halides are adsorbed. The columns are then rinsed with 0.8M potassium nitrate solution to remove inorganic halides. The columns are combusted individually and the generated halo-acids are collected and measured in a microcoulometric cell. | 0.01 | Y | Y | Y | 5601 | 7/17/1996 | 4/11/2023 5:10:00 PM | 7/16/2024 3:11:25 PM | 1996-04-12 | 6/3/1998 6:51:42 PM | EMS |  | Concentration | mg/L | Y | (fl. conc.) |  | Adsorbable Organic Halides (fl. conc.) |  | Aggregate Organics; | Aggregate Organics; EMS code: AOX- | Chemical |  | LAB
AOX- | mg/L | Adsorbable Organic Halides | X302 | Na2SO3,HNO3,TOX Analyzer | An aqueous sample is treated, as required, with Sodium Sulphite to remove any residual Chlorine and acidified with Nitric Acid. An aliquot of the sample, the actual dilution dependent on the estimated AOX concentration, is passed at a precisely controlled flow rate through an assembly of two granular activated carbon (GAC) columns connected in series. The GAC in the column assembly is then "washed" with a nitrate solution to desorb any inorganic halides. The GAC from each column is separately pyrolysed at high temperature in a controlled atmosphere and the resulting "HX" gases are purged through a dehydrating tube into a Chlorine Detection Cell, automatically kept at a preset potential, where the sample is trapped in an Acetic Acid electrolyte and microcoulometrically titrated with silver ions. Monitoring, integration and conversion of the cell EMF to concentration units is accomplished electronically. | 0.01 | Y | Y | Y | 9324 | 12/12/1996 | 10/31/2021 6:59:00 AM | 8/20/2024 8:39:43 AM | 1996-04-12 | 6/3/1998 6:51:41 PM | EMS |  | Concentration | mg/L | Y | (fl. conc.) |  | Adsorbable Organic Halides (fl. conc.) |  | Aggregate Organics; | Aggregate Organics; EMS code: AOX- | Chemical |  | LAB
AOX- | ug/g | Adsorbable Organic Halides | X459 | Na2SO3,HNO3,TOX Analyzer | Na2SO3;HNO3;TOX Analyzer | 0 | Y | Y | Y | 713 | 10/7/1996 | 10/9/1997 4:10:00 PM | 2/25/2009 4:42:43 PM | 1996-04-12 | 6/3/1998 6:51:41 PM | EMS |  | MassPerMass | mg/kg | Y | (mass conc.) |  | Adsorbable Organic Halides (mass conc.) |  | Aggregate Organics; | Aggregate Organics; EMS code: AOX- | Chemical |  | LAB
AOX- | t/d | Adsorbable Organic Halides | CALC | Calculation (t/d) | Calculation (t/d) | 0 | Y | Y | Y | 514 |  |  | 8/22/2024 1:49:14 PM | 1996-04-12 | 6/3/1998 6:51:42 PM | EMS |  | MassFlow | t/d | Y | (mass flow) |  | Adsorbable Organic Halides (mass flow) |  | Aggregate Organics; | Aggregate Organics; EMS code: AOX- | Chemical |  | LAB
AOX- | mg/L | Adsorbable Organic Halides | XCAL | Calculation (mg/L) | Calculation (mg/L) | 0 | Y | Y | Y | 1843 | 9/26/1996 7:00:00 AM | 2/27/2016 | 8/22/2024 1:49:14 PM | 1996-04-12 | 6/3/1998 6:51:42 PM | EMS |  | Concentration | mg/L | Y | (fl. conc.) |  | Adsorbable Organic Halides (fl. conc.) |  | Aggregate Organics; | Aggregate Organics; EMS code: AOX- | Chemical |  | LAB
AOX- | mg/L | Adsorbable Organic Halides | X249 | Pyrol:Coloumetric Titr'n | Pyrol:Coloumetric Titr'n | 0.005 | Y | Y | Y | 1909 | 9/3/1996 | 1/28/2020 8:00:00 AM | 3/6/2020 8:25:57 AM | 1996-04-12 | 6/3/1998 6:51:45 PM | EMS |  | Concentration | mg/L | Y | (fl. conc.) |  | Adsorbable Organic Halides (fl. conc.) |  | Aggregate Organics; | Aggregate Organics; EMS code: AOX- | Chemical |  | LAB
AP-I | mg/dm2/d | Part.Ash:Insoluble | EQ15 | EQUIS (mg/dm2/d) | EQUIS (mg/dm2/d) | 0.05 | N | Y | Y | 4426 |  |  | 10/23/1996 12:06:51 AM | 1996-04-12 | 6/3/1998 6:51:45 PM | EMS |  | MassFlux | mg/dm²/d | Y | (mass flux) |  | Part.Ash:Insoluble (mass flux) |  | Particulates; | Particulates; EMS code: AP-I | Chemical | Insoluble | LAB
AP-I | mg | Part.Ash:Insoluble | EQ03 | EQUIS (mg) | EQUIS (mg) | 0 | N | Y | Y | 13 |  |  | 10/22/1996 10:13:43 PM | 1996-04-12 | 6/3/1998 6:51:33 PM | EMS |  | Mass | mg | Y | (mass) |  | Part.Ash:Insoluble (mass) |  | Particulates; | Particulates; EMS code: AP-I | Chemical | Insoluble | LAB
AP-I | mg/dm2/d | Part.Ash:Insoluble | 5031 | Gravimetric-DF Insol Ash | A 10.4 cm (4.1") diameter polyethylene canister containing a collection medium is exposed to the ambient air for a period of approximately 30 days. The prepared sample (see Dustfall-Total procedure) is passed through a 0.45 um membrane filter. The residue after ignition at 550C constitutes the ashed insoluble particulate. | 0.1 | N | Y | Y | 5131 |  |  | 11/4/1996 4:56:33 PM | 1996-04-12 | 6/3/1998 6:51:33 PM | EMS |  | MassFlux | mg/dm²/d | Y | (mass flux) |  | Part.Ash:Insoluble (mass flux) |  | Particulates; | Particulates; EMS code: AP-I | Chemical | Insoluble | LAB
AP-I | mg/dm2/d | Part.Ash:Insoluble | X175 | Gravimetric | Code X175 covers a variety of gravimetric procedures. | 0.1 | Y | Y | Y | 371 |  |  | 2/4/1998 5:53:47 PM | 1996-04-12 | 6/3/1998 6:51:33 PM | EMS |  | MassFlux | mg/dm²/d | Y | (mass flux) |  | Part.Ash:Insoluble (mass flux) |  | Particulates; | Particulates; EMS code: AP-I | Chemical | Insoluble | LAB
AP-K | ug/g | Plant available potassium | AP-K | AP-K | Comm Soil Sci. Plant Anal, 25 (5&6) | 20 | Y | Y | Y | 18 | 9/1/2021 | 2/13/2023 | 1/11/2024 1:07:07 PM | 2022-06-29 |  | EMS |  | MassPerMass | mg/kg | Y | (mass conc.) |  | Plant available potassium (mass conc.) | 7440-09-7 | Nutrients; | Nutrients; EMS code: AP-K | Chemical |  | LAB
AP-P | ug/g | Plant Available Phosphate reported as Phosphorus | AP-P | PO4-AVAIL-SK | Comm. Soil Sci. Plant Anal. 25 (5&6) - Available Phosphate-P | 2 | Y | Y | Y | 18 | 9/1/2021 | 2/13/2023 | 1/11/2024 1:07:07 PM | 2020-01-13 |  | RDGIBSON |  | MassPerMass | mg/kg | Y | (mass conc.) |  | Plant Available Phosphate reported as Phosphorus (mass conc.) | 14265-44-2 | Nutrients; | Nutrients; EMS code: AP-P | Chemical |  | LAB
AP-S | mg | Part.Ash:Soluble | EQ03 | EQUIS (mg) | EQUIS (mg) | 0 | N | Y | Y | 13 |  |  | 10/22/1996 10:13:42 PM | 1996-04-12 | 6/3/1998 6:51:33 PM | EMS |  | Mass | mg | Y | (mass) |  | Part.Ash:Soluble (mass) |  | Particulates; | Particulates; EMS code: AP-S | Chemical | Soluble | LAB
AP-S | mg/dm2/d | Part.Ash:Soluble | 5041 | Gravimetric- DF Sol Ash | A 10.4 cm (4.1") diameter polyethylene canister containing a collection medium is exposed to the ambient air for a period of approximately 30 days. The sample subsequently undergoes gravimetric and/or chemical analysis in the laboratory, plus loading calculation. The prepared sample (see Dustfall-Total procedure) is passed through a 0.45 um membrane filter. A portion of the filtrate is then evaporated in a oven; the portion which ignites to constant weight at 550C constitutes the ashed soluble particulate. It should be noted that if the ashed soluble particulate procedure is completed, this will preclude phosphorus and metal analysis. | 0.05 | N | Y | Y | 5181 |  |  | 11/4/1996 4:56:33 PM | 1996-04-12 | 6/3/1998 6:51:33 PM | EMS |  | MassFlux | mg/dm²/d | Y | (mass flux) |  | Part.Ash:Soluble (mass flux) |  | Particulates; | Particulates; EMS code: AP-S | Chemical | Soluble | LAB
AP-S | mg/dm2/d | Part.Ash:Soluble | EQ15 | EQUIS (mg/dm2/d) | EQUIS (mg/dm2/d) | 0 | N | Y | Y | 3216 |  |  | 10/23/1996 12:06:51 AM | 1996-04-12 | 6/3/1998 6:51:45 PM | EMS |  | MassFlux | mg/dm²/d | Y | (mass flux) |  | Part.Ash:Soluble (mass flux) |  | Particulates; | Particulates; EMS code: AP-S | Chemical | Soluble | LAB
AP-S | mg/dm2/d | Part.Ash:Soluble | X175 | Gravimetric | Code X175 covers a variety of gravimetric procedures. | 0.1 | Y | Y | Y | 376 |  |  | 2/4/1998 6:41:38 PM | 1996-04-12 | 6/3/1998 6:51:33 PM | EMS |  | MassFlux | mg/dm²/d | Y | (mass flux) |  | Part.Ash:Soluble (mass flux) |  | Particulates; | Particulates; EMS code: AP-S | Chemical | Soluble | LAB
AP-T | mg/dm2/d | Part.Ash:Total | X175 | Gravimetric | Code X175 covers a variety of gravimetric procedures. | 0.05 | Y | Y | Y | 317 |  |  | 10/24/1996 6:31:42 AM | 1996-04-12 | 6/3/1998 6:51:33 PM | EMS |  | MassFlux | mg/dm²/d | Y | (mass flux) |  | Part.Ash:Total (mass flux) |  | Particulates; | Particulates; EMS code: AP-T | Chemical | Total | LAB
AP-T | mg/dm2/d | Part.Ash:Total | EQ15 | EQUIS (mg/dm2/d) | EQUIS (mg/dm2/d) | 0 | N | Y | Y | 55 |  |  | 10/24/1996 6:31:46 AM | 1996-04-12 | 6/3/1998 6:51:33 PM | EMS |  | MassFlux | mg/dm²/d | Y | (mass flux) |  | Part.Ash:Total (mass flux) |  | Particulates; | Particulates; EMS code: AP-T | Chemical | Total | LAB
AP-T | mg/L | Part.Ash:Total | X484 | Gravimetric Intermediate Results | A 10.4 cm (4.1") diameter polyethylene canister containing a collection medium is exposed to the ambient air for a period of approximately 30 days. The sample subsequently undergoes gravimetric and/or chemical analysis in the laboratory with intermediate results reported in units of mg and loading results reported with units of mg/d m2/d. | 0.1 | Y | Y | Y | 426 |  |  | 7/30/2024 12:34:56 PM | 1998-01-22 | 6/3/1998 6:51:47 PM | NPEPPIN |  | Concentration | mg/L | Y | (fl. conc.) |  | Part.Ash:Total (fl. conc.) |  | Particulates; | Particulates; EMS code: AP-T | Chemical | Total | LAB
AP-T | mg/dm2/d | Part.Ash:Total | CS15 | System Calculation | System Calculation (mg/dm2/d) |  | N | Y | N | 1763 |  |  | 7/11/2005 5:41:47 PM | 1996-04-12 | 6/3/1998 6:51:43 PM | EMS |  | MassFlux | mg/dm²/d | Y | (mass flux) |  | Part.Ash:Total (mass flux) |  | Particulates; | Particulates; EMS code: AP-T | Chemical | Total | LAB
AS-D | mg/L | Arsenic Dissolved | X063 | Arsine gen,535nm,Col'met | Inorganic arsenic is reduced to arsine, AsH3, by zinc in an acid medium. The arsine then passes from the generator flask through a lead acetate scrubber and into an absorber tube containing silver diethyldithiocarbamate dissolved in pyridine. In the absorber, the arsine reacts with the silver salt to form a soluble red colored complex which is measured colorimetrically at 535nm or 540 nm. | 0.001 | Y | Y | Y | 1603 |  |  | 1/17/2007 3:47:12 PM | 1996-04-12 | 6/3/1998 6:51:34 PM | EMS |  | Concentration | mg/L | Y | (fl. conc.) |  | Arsenic Dissolved (fl. conc.) | 7440-38-2 | Metals - Dissolved; | Metals - Dissolved; EMS code: AS-D | Chemical | Dissolved | LAB
AS-D | mg/L | Arsenic Dissolved | METC | MET-D-NP-CCMS-VA | Ultra trace metals in water are analyzed by CRC ICPMS, based on US EPA Method 6020B (July 2014). This procedure is intended for pristine field-filtered acid-preserved water samples. The detection limits (LOR) for this test are based on lab instrumental analysis only, not including filtration. ALS-supplied field filtration equipment does not support these LOR. Therefore, because of the high probability of false positives due to filtration, it is strongly recommended that a filtration blank be analysed to aid in data interpretation. | 2e-05 | Y | Y | Y | 1776 | 4/3/2012 | 8/15/2024 | 8/23/2024 9:50:35 AM | 2019-08-20 |  | KATFRASE |  | Concentration | mg/L | Y | (fl. conc.) |  | Arsenic Dissolved (fl. conc.) | 7440-38-2 | Metals - Dissolved; | Metals - Dissolved; EMS code: AS-D | Chemical | Dissolved | LAB
AS-D | mg/L | Arsenic Dissolved | X356 | LF,LA:HNO3:ICP | Samples are lab-filtered (0.45 um); lab-acidified with HNO3; analysis by ICP/AES. | 0.04 | Y | Y | Y | 426 | 10/17/2000 | 5/1/2002 | 3/23/2010 5:08:08 PM | 1996-04-12 | 6/3/1998 6:51:34 PM | EMS |  | Concentration | mg/L | Y | (fl. conc.) |  | Arsenic Dissolved (fl. conc.) | 7440-38-2 | Metals - Dissolved; | Metals - Dissolved; EMS code: AS-D | Chemical | Dissolved | LAB
AS-D | mg/L | Arsenic Dissolved | X170 | ICP | This code is a catch-all for a large variety of ICP methods including water and also solids digested into solutions as an intermediate result. Different procedures will have greatly different detection limits. (a) For air particulates, preliminary results are reported in mg/L and loading results as mg/dm2/d. (b) For Hi Vol samples preliminiary results are same as reported in mg/L and loading results as ug/m3. (c) Note that for Rain samples there is no filtration nor digestion steps. | 0.3 | Y | Y | Y | 9290 | 4/11/1996 6:01:00 PM | 10/29/2019 | 4/12/2024 1:44:29 PM | 1996-04-12 | 6/3/1998 6:51:41 PM | EMS |  | Concentration | mg/L | Y | (fl. conc.) |  | Arsenic Dissolved (fl. conc.) | 7440-38-2 | Metals - Dissolved; | Metals - Dissolved; EMS code: AS-D | Chemical | Dissolved | LAB
AS-D | mg/L | Arsenic Dissolved | X544 | FF,FA,HNO3,ICPMS | Analytes nebulized followed by ion entrained in plasma gas and extracted by means of a differentially pumped vacuum interface into a mass spectrometer | 0.001 | Y | Y | Y | 30 |  |  | 10/30/2014 8:31:17 AM | 2001-09-18 | 9/18/2001 1:36:23 PM | NPEPPIN | NPEPPIN | Concentration | mg/L | Y | (fl. conc.) |  | Arsenic Dissolved (fl. conc.) | 7440-38-2 | Metals - Dissolved; | Metals - Dissolved; EMS code: AS-D | Chemical | Dissolved | LAB
AS-D | mg/L | Arsenic Dissolved | 0170 | Hydride/ICP Anal. | Hydride/ICP Anal. | 0.001 | Y | Y | Y | 441 |  |  | 10/24/1996 12:12:40 PM | 1996-04-12 | 6/3/1998 6:51:34 PM | EMS |  | Concentration | mg/L | Y | (fl. conc.) |  | Arsenic Dissolved (fl. conc.) | 7440-38-2 | Metals - Dissolved; | Metals - Dissolved; EMS code: AS-D | Chemical | Dissolved | LAB
AS-D | mg/L | Arsenic Dissolved | EQ01 | EQUIS (mg/L) | EQUIS (mg/L) | 0.05 | N | Y | Y | 3530 |  |  | 10/24/1996 3:15:52 PM | 1996-04-12 | 6/3/1998 6:51:34 PM | EMS |  | Concentration | mg/L | Y | (fl. conc.) |  | Arsenic Dissolved (fl. conc.) | 7440-38-2 | Metals - Dissolved; | Metals - Dissolved; EMS code: AS-D | Chemical | Dissolved | LAB
AS-D | mg/L | Arsenic Dissolved | X089 | Nitric/Ni Dig,Graph. Fur | Nitric/Ni Dig;Graph. Fur | 0.001 | Y | Y | Y | 6 |  |  | 1/30/1997 2:12:21 PM | 1996-04-12 | 6/3/1998 6:51:34 PM | EMS |  | Concentration | mg/L | Y | (fl. conc.) |  | Arsenic Dissolved (fl. conc.) | 7440-38-2 | Metals - Dissolved; | Metals - Dissolved; EMS code: AS-D | Chemical | Dissolved | LAB
AS-D | mg/L | Arsenic Dissolved | X209 | Hom,Nitr/Perchl,Hydr,AA | Hom;Nitr/Perchl;Hydr;AA | 0.001 | Y | Y | Y | 59 |  |  | 10/24/1996 1:00:47 PM | 1996-04-12 | 6/3/1998 6:51:34 PM | EMS |  | Concentration | mg/L | Y | (fl. conc.) |  | Arsenic Dissolved (fl. conc.) | 7440-38-2 | Metals - Dissolved; | Metals - Dissolved; EMS code: AS-D | Chemical | Dissolved | LAB
AS-D | mg/L | Arsenic Dissolved | X556 | Metals in water, dissolved, "in-bottle" digestion, ICP-SFMS | The whole water sample is filtered through a 0.45 micron cellulose acetate filter paper and the filtrate is stored in a HDPE/Teflon bottle. The container is de-capped in a Class 100 workstation, acidified to 2% (v/v) HNO3 and 1% (v/v) HCl and recapped. The sample is "in-bottle" digested in a convection oven, set at 60 degrees C for 16 hours. The sample is cooled and introduced to an inductively coupled argon plasma-sector field mass spectrometer (ICP-SFMS) by pneumatic nebulization. Each respective element is measured at a specific mass to charge ratio, m/z value expressed in amu (atomic mass units), and at a specified intrument resolution. LR: Low Resolution ~300, MR: Mid Resolution ~4000 and HR: High Resolution ~10000. Ref: Method number 02-2003, National Laboratory for Environmental Testing, CCIW, Environment Canada, Burlington, Requested April 2003. | 1e-05 | Y | Y | Y | 4234 | 6/11/2003 | 12/6/2023 2:04:00 AM | 2/14/2024 9:05:32 AM | 2004-01-19 |  | UMULLER |  | Concentration | mg/L | Y | (fl. conc.) |  | Arsenic Dissolved (fl. conc.) | 7440-38-2 | Metals - Dissolved; | Metals - Dissolved; EMS code: AS-D | Chemical | Dissolved | LAB
AS-D | mg/L | Arsenic Dissolved | MSEA | Metals in Seawater (undigested) by (CRC) ICP-MS | This method describes the multi-elemental determination of trace elements by Collision/Reaction Cell (CRC) ICP-MS. The method measures ions produced by a radio-frequency inductively coupled plasma. Analyte species originating in a liquid are pneumatically nebulized and the resulting aerosol transported by argon gas into the plasma torch. The ions produced are entrained in the plasma gas and extracted, by means of a differentially pumped vacuum interface, into a mass spectrometer. The ions flow through a Collision/Reaction Cell (CRC) where depending on the mode of analysis either the ions passes through the cell unhindered (no gas - classical ICP-MS mode), or collisions/reactions chemistry takes place between the ions and a neutral gas such as hydrogen or helium. Any polyatomic ions which give interference on the analytes are removed. The ions are then sorted according to their mass-to-charge ratios by a quadrupole mass spectrometer. The ions are then sorted according to their mass-to-charge ratios by a quadrupole mass spectrometer. The ions transmitted through the quadrupole are quantified by a channel electron multiplier and the ion information processed by a data handling system. Although most polyatomic interference is removed, any additional interference must be assessed and valid corrections applied or the data flagged to indicate problems. Interference correction must include compensation for background ions contributed by the plasma gas, reagents, and constituents of the sample matrix. Instrument drift as well as suppressions or enhancements of instrument response caused by the sample matrix must be corrected for by the use of internal standards. | 0.0001 | Y | Y | Y | 4 | 4/12/2023 | 4/29/2024 | 5/16/2024 2:00:14 PM | 2009-09-08 |  | RDGIBSON |  | Concentration | mg/L | Y | (fl. conc.) |  | Arsenic Dissolved (fl. conc.) | 7440-38-2 | Metals - Dissolved; | Metals - Dissolved; EMS code: AS-D | Chemical | Dissolved | LAB
AS-D | mg/L | Arsenic Dissolved | MWID | Metals in Water by ICPMS (digested) | Metals in Water by ICPMS (digested) Based on EPA 200.8 | 0.0004 | Y | Y | Y | 31872 | 10/26/2018 5:20:00 PM | 10/26/2018 5:20:00 PM | 7/31/2024 1:46:41 PM | 2009-12-24 |  | RDGIBSON |  | Concentration | mg/L | Y | (fl. conc.) |  | Arsenic Dissolved (fl. conc.) | 7440-38-2 | Metals - Dissolved; | Metals - Dissolved; EMS code: AS-D | Chemical | Dissolved | LAB
AS-D | mg/L | Arsenic Dissolved | X562 | Dissolved metals: FF, FA, HNO3, HR-ICPMS | Ultra trace metals in water are analyzed by high resolution inductively coupled plasma mass spectrometry (HR-ICPMS), based on US EPA Method 200.8, (Oct. 1999). This procedure is intended for pristine field-filtered acid-preserved water samples. The detection limits (LOR) for this test are based on lab instrumental analysis only, not including filtration. ALS-supplied field filtration equipment does not support these LOR. Therefore, because of the high probability of false positives due to filtration, it is strongly recommended that a filtration blank be analyzed to aid in data interpretation. | 2e-05 | Y | Y | Y | 1264 | 6/8/2015 | 10/19/2023 | 11/9/2023 7:41:27 AM | 2015-10-22 |  | RDGIBSON |  | Concentration | mg/L | Y | (fl. conc.) |  | Arsenic Dissolved (fl. conc.) | 7440-38-2 | Metals - Dissolved; | Metals - Dissolved; EMS code: AS-D | Chemical | Dissolved | LAB
AS-D | mg/L | Arsenic Dissolved | 0185 | Hydr.At.Absorp.Spec(H2O) | Hydr.At.Absorp.Spec(H2O) | 0.5 | Y | Y | Y | 10200 |  |  | 5/19/2015 3:10:36 PM | 1996-04-12 | 6/3/1998 6:51:34 PM | EMS |  | Concentration | mg/L | Y | (fl. conc.) |  | Arsenic Dissolved (fl. conc.) | 7440-38-2 | Metals - Dissolved; | Metals - Dissolved; EMS code: AS-D | Chemical | Dissolved | LAB
AS-D | mg/L | Arsenic Dissolved | 0181 | HCl/K2S2O8, Hydride, ICP | HCl/K2S2O8, Hydride; ICP | 0.001 | Y | Y | Y | 1146 |  |  | 10/24/1996 2:21:46 PM | 1996-04-12 | 6/3/1998 6:51:34 PM | EMS |  | Concentration | mg/L | Y | (fl. conc.) |  | Arsenic Dissolved (fl. conc.) | 7440-38-2 | Metals - Dissolved; | Metals - Dissolved; EMS code: AS-D | Chemical | Dissolved | LAB
AS-D | mg/L | Arsenic Dissolved | F037 | K2S2O8/HCl dig, Hydride,ICP | The water sample is digested with potassium persulphate and concentrated hydrochloric acid. Samples, blanks and standards are placed in an automatic sampler and analyzed by emission spectroscopy utilizing induction coupled argon plasma as an excitation source. All forms of arsenic, selenium and antimony are broken down with acidic persulphate to arsenate, selenate and antimonate. Sodium borohydride is required to reduce the arsenate to arsenite. Hydrochloric acid reduces the selenate to selenite. Potassium iodide is required to reduce the antimonate to antimonite. Hydrides of arsenic, selenium and antimony are formed by the action of sodium borohydride in an acidic medium, and analysed by inductively coupled argon plasma optical emission spectrometry. | 0.0001 | Y | Y | Y | 230 | 8/10/1998 2:45:00 PM | 12/1/2010 | 4/18/2018 9:26:32 AM | 1997-04-02 | 6/3/1998 6:51:46 PM | MPAK |  | Concentration | mg/L | Y | (fl. conc.) |  | Arsenic Dissolved (fl. conc.) | 7440-38-2 | Metals - Dissolved; | Metals - Dissolved; EMS code: AS-D | Chemical | Dissolved | LAB
AS-D | mg/L | Arsenic Dissolved | X071 | FA,HNO3,Hydr.Gen,AA | FA;HNO3;Hydr.Gen;AA | 0.0001 | Y | Y | Y | 1794 |  |  | 3/8/2004 6:14:30 PM | 1996-04-12 | 6/3/1998 6:51:34 PM | EMS |  | Concentration | mg/L | Y | (fl. conc.) |  | Arsenic Dissolved (fl. conc.) | 7440-38-2 | Metals - Dissolved; | Metals - Dissolved; EMS code: AS-D | Chemical | Dissolved | LAB
AS-D | mg/L | Arsenic Dissolved | X095 | FF,FA,HNO3,ICP | FF;FA;HNO3;ICP | 0.001 | Y | Y | Y | 342 |  |  | 10/24/1996 12:20:01 PM | 1996-04-12 | 6/3/1998 6:51:34 PM | EMS |  | Concentration | mg/L | Y | (fl. conc.) |  | Arsenic Dissolved (fl. conc.) | 7440-38-2 | Metals - Dissolved; | Metals - Dissolved; EMS code: AS-D | Chemical | Dissolved | LAB
AS-D | mg/L | Arsenic Dissolved | X359 | LF,LA:HNO3:HVAAS | Samples are lab-filtered (0.45 um) and lab-acidified (HNO3); analysis by HVAAS. | 0.0002 | Y | Y | Y | 504 | 1/23/1998 | 6/5/2003 | 6/9/2003 10:20:17 AM | 1996-04-12 | 6/3/1998 6:51:34 PM | EMS |  | Concentration | mg/L | Y | (fl. conc.) |  | Arsenic Dissolved (fl. conc.) | 7440-38-2 | Metals - Dissolved; | Metals - Dissolved; EMS code: AS-D | Chemical | Dissolved | LAB
AS-D | mg/L | Arsenic Dissolved | X202 | FF:FA,HNO3,Hydrid Gen,AA | Samples are field-filtered (0.45 um) and field-preserved (HNO3); analyssis by gaseous hydride AA. | 0.0001 | Y | Y | Y | 1946 |  |  | 1/29/2010 9:45:12 AM | 1996-04-12 | 6/3/1998 6:51:34 PM | EMS |  | Concentration | mg/L | Y | (fl. conc.) |  | Arsenic Dissolved (fl. conc.) | 7440-38-2 | Metals - Dissolved; | Metals - Dissolved; EMS code: AS-D | Chemical | Dissolved | LAB
AS-D | mg/L | Arsenic Dissolved | F032 | Filt.,Acid,ICPMS | There is no acid digestion required for this analysis. A water sample is collected and preserved in 0.2 percent nitric acid. To analyse the extractable metals the sample is shaken, allowed to sit for at least twenty four hours, and a decanted portion is introduced to the spectrometer. When the dissolved metals are required, the samples must first be filtered through a 0.45 micron cellulose acetate filter paper and then followed by preservation in 0.2 percent nitric acid. It is important to note that different laboratories may have very different detection limits for some variables analyzed by ICPMS. | 0.0005 | Y | Y | Y | 86631 | 3/10/1998 2:03:00 PM | 8/26/2024 | 8/26/2024 4:53:05 PM | 1998-05-14 | 8/15/2019 10:51:51 AM | NPEPPIN | KATFRASE | Concentration | mg/L | Y | (fl. conc.) |  | Arsenic Dissolved (fl. conc.) | 7440-38-2 | Metals - Dissolved; | Metals - Dissolved; EMS code: AS-D | Chemical | Dissolved | LAB
AS-D | mg/L | Arsenic Dissolved | 0180 | HNO3 Dig:Hydride ICP | HNO3 Dig:Hydride ICP | 0.001 | Y | Y | Y | 5040 |  |  | 8/26/2024 2:41:32 PM | 1996-04-12 | 6/3/1998 6:51:34 PM | EMS |  | Concentration | mg/L | Y | (fl. conc.) |  | Arsenic Dissolved (fl. conc.) | 7440-38-2 | Metals - Dissolved; | Metals - Dissolved; EMS code: AS-D | Chemical | Dissolved | LAB
AS-D | mg/L | Arsenic Dissolved | F039 | Field filt + acid, ICP-AES | Dissolved and extractable water samples are ready for analysis once collected. Total samples are autoclaved with HNO3 and hydrochloric acid (HCI) in capped Teflon vials. Analysis is performed via Inductively Coupled Argon Plasma Atomic Emission Spectrometry (ICP-AES)(simultaneous multielement analysis), Graphite Furnace Atomic Absorption Spectrometry (GF-AAS)(low level analysis) and/or Flame Atomic Absorption Spectrometry (F-AAS). | 0.05 | Y | Y | Y | 5897 | 7/29/1996 6:03:00 PM | 7/5/2019 1:21:00 PM | 9/11/2019 8:47:50 AM | 1997-05-08 | 6/3/1998 6:51:46 PM | MPAK |  | Concentration | mg/L | Y | (fl. conc.) |  | Arsenic Dissolved (fl. conc.) | 7440-38-2 | Metals - Dissolved; | Metals - Dissolved; EMS code: AS-D | Chemical | Dissolved | LAB
AS-D | t/d | Arsenic Dissolved | CALC | Calculation (t/d) | Calculation (t/d) | 0 | Y | Y | Y | 2695 |  |  | 7/3/2024 12:59:25 PM | 1996-04-12 | 6/3/1998 6:51:34 PM | EMS |  | MassFlow | t/d | Y | (mass flow) |  | Arsenic Dissolved (mass flow) | 7440-38-2 | Metals - Dissolved; | Metals - Dissolved; EMS code: AS-D | Chemical | Dissolved | LAB
AS-D | mg/L | Arsenic Dissolved | X085 | LF,LA,HNO3,AA | Samples are lab-filtered (0.45 um), and lab-acidified (HNO3); analysis by AA. | 0.2 | Y | Y | Y | 101 | 1/26/1998 | 7/17/2002 | 8/6/2002 5:40:27 PM | 1996-04-12 | 6/3/1998 6:51:41 PM | EMS |  | Concentration | mg/L | Y | (fl. conc.) |  | Arsenic Dissolved (fl. conc.) | 7440-38-2 | Metals - Dissolved; | Metals - Dissolved; EMS code: AS-D | Chemical | Dissolved | LAB
AS-D | mg/L | Arsenic Dissolved | X285 | HNO3 Dig:ICP Analysis | HNO3 Dig:ICP Analysis | 0.003 | Y | Y | Y | 672 | 1/30/2002 | 11/13/2002 | 3/18/2004 10:28:33 AM | 1998-05-22 | 6/3/1998 6:51:49 PM | NPEPPIN |  | Concentration | mg/L | Y | (fl. conc.) |  | Arsenic Dissolved (fl. conc.) | 7440-38-2 | Metals - Dissolved; | Metals - Dissolved; EMS code: AS-D | Chemical | Dissolved | LAB
AS-D | mg/L | Arsenic Dissolved | F068 | HNO3dig,TeflonTube/Autoclave,ICP-MS | This method is applicable to determiniation of dissolved, extractable, and total recoverable metals in ground waters, surface waters, drinking waters, and wastewaters. Dissolved, extractable, and total recoverable sample fractions are acidified to pH <1.5 with analytical grade nitric acid (HNO3) to prevent the metals from adsorbing to the container material or precipitating. Extractable and total recoverable water samples are collected in 250 ml acid washed polyethylene bottles and are preserved with Nitric acid (0.5 ml acid for 100 mLs of sample). All preserved samples are stored at room temperature and can be held for up to 6 months. Unfiltered, un-preserved samples are stored at 4C for no more than 30 days. For dissolved and extractable water fractions, no sample preparation is necessary. Samples are simply poured in autosampler tubes (i.e. 15 ml centrifuge tubes), and placed on the autosampler and analyzed. Samples for total recoverable metals are digested with Nitric acid in special Teflon tubes in an autoclave prior to analyses. | 0.0001 | Y | Y | Y | 1398 | 6/1/1999 1:19:00 PM | 4/17/2015 1:06:00 PM | 8/11/2021 12:04:11 PM | 1999-06-28 | 6/30/1999 1:05:24 PM | NPEPPIN | NPEPPIN | Concentration | mg/L | Y | (fl. conc.) |  | Arsenic Dissolved (fl. conc.) | 7440-38-2 | Metals - Dissolved; | Metals - Dissolved; EMS code: AS-D | Chemical | Dissolved | LAB
AS-D | mg/L | Arsenic Dissolved | F023 | HNO3,GFAA | Dissolved and extractable water samples are ready for analysis once collected. Total samples are autoclaved with HNO3 and hydrochloric acid (HCl) in capped Teflon vials. Analysis is performed via Inductively Coupled Argon Plasma Atomic Emission Spectrometry (ICP-AES) (simultaneous multielement analysis), Graphite Furnace Atomic Absorption Spectrometry (GF-AAS) (low level analysis) and/or Flame Atomic Absorption Spectrometry (F-AAS). Note that extractable samples are acidified with 0.5 mL HNO3 per 100 mL of sample at time of sample collection. | 0.0005 | Y | Y | Y | 803 | 8/2/1996 4:07:00 PM | 9/16/2003 3:18:00 PM | 11/28/2003 3:22:08 PM | 1996-09-17 | 6/3/1998 6:51:45 PM | MPAK |  | Concentration | mg/L | Y | (fl. conc.) |  | Arsenic Dissolved (fl. conc.) | 7440-38-2 | Metals - Dissolved; | Metals - Dissolved; EMS code: AS-D | Chemical | Dissolved | LAB
AS-E | mg/L | Arsenic Extractable | F040 | Field Acid, lab filt, ICP-AES | Dissolved and extractable water samples are ready for analysis once collected. Total samples are autoclaved with HNO3 and hydrochloric acid (HCI) in capped Teflon vials. Analysis is performed via Inductively Coupled Argon Plasma Atomic Emission Spectrometry (ICP-AES)(simultaneous multielement analysis), Graphite Furnace Atomic Absorption Spectrometry (GF-AAS)(low level analysis) and/or Flame Atomic Absorption Spectrometry (F-AAS). | 0.05 | Y | Y | Y | 2361 | 8/5/1997 6:06:00 PM | 3/14/2008 | 9/11/2019 8:47:41 AM | 1997-05-08 | 6/3/1998 6:51:46 PM | MPAK |  | Concentration | mg/L | Y | (fl. conc.) |  | Arsenic Extractable (fl. conc.) | 7440-38-2 | Metals - Extractable; | Metals - Extractable; EMS code: AS-E | Chemical | Extractable | LAB
AS-E | mg/L | Arsenic Extractable | F069 | SEM extr,6N HCl,ICP/OES | 6.0 N Hydrochloric acid (HCI) is added to an aliquot of the sediment sample. The extract produced from the addition of the acid is then analyzed for simultaneously extracted metals (SEM) using inductively coupled plasma - optical emission spectrophotometry. The detection limits depend on the metal parameters in question and the instrumentation used. | 0.2 | Y | Y | Y | 10 | 3/26/2012 3:01:00 AM | 3/26/2012 3:01:00 AM | 11/25/2021 1:50:15 PM | 1999-07-09 |  | NPEPPIN |  | Concentration | mg/L | Y | (fl. conc.) |  | Arsenic Extractable (fl. conc.) | 7440-38-2 | Metals - Extractable; | Metals - Extractable; EMS code: AS-E | Chemical | Extractable | LAB
AS-E | mg/L | Arsenic Extractable | X553 | Metals in water, extractable, ICP-SFMS | The whole water sample is collected in a HDPE/Teflon bottle. In a Class 100 workstation it is preserved with concentrated Nitric Acid at concentration of 0.2 percent (v/v). The acid preservative extracts the metals from material in the sample and a decanted portion is introduced directly to an inductively coupled argon plasma-sector field mass spectrometer (ICP-SFMS) by pneumatic nebulization. Each respective element is measured at a specific mass to charge ratio, m/z value expressed in amu (atomic mass units), and at a specified instrument resolution. LR: Low Resolution ~300, MR: Mid Resolution ~4000 and HR: High Resolution ~10000. Ref: Method number 02-2061, National Laboratory for Environmental Testing, CCIW, Environment Canada, Burlington. Requested April 2003. | 1e-05 | Y | Y | Y | 2587 | 5/22/2003 | 9/21/2007 | 11/25/2021 1:50:15 PM | 2004-01-19 |  | UMULLER |  | Concentration | mg/L | Y | (fl. conc.) |  | Arsenic Extractable (fl. conc.) | 7440-38-2 | Metals - Extractable; | Metals - Extractable; EMS code: AS-E | Chemical | Extractable | LAB
AS-E | ug/g | Arsenic Extractable | X230 | Hom,HNO3/HCl,ICP | The sample is dried, ground to pass 100 mesh, and digested with a nitric-hydrochloric acid mixture to solubilize the solid matter and to remove the organic material by oxidation and volatilization. The completed digests are submitted to metals instrumentation personnel for metals analysis by Inductively Coupled Atomic Emission Spectrophotometry (ICP-AES), or hydride analysis by atomic absorption. | 15 | Y | Y | Y | 74 | 12/13/2011 12:02:00 PM | 12/13/2011 12:02:00 PM | 2/9/2012 12:15:56 PM | 1996-04-12 | 6/3/1998 6:51:34 PM | EMS |  | MassPerMass | mg/kg | Y | (mass conc.) |  | Arsenic Extractable (mass conc.) | 7440-38-2 | Metals - Extractable; | Metals - Extractable; EMS code: AS-E | Chemical | Extractable | LAB
AS-E | mg/L | Arsenic Extractable | X129 | Extr,Flameless AA | Extr;Flameless AA | 0.0001 | Y | Y | Y | 1018 |  |  | 8/23/1999 3:43:46 PM | 1996-04-12 | 6/3/1998 6:51:41 PM | EMS |  | Concentration | mg/L | Y | (fl. conc.) |  | Arsenic Extractable (fl. conc.) | 7440-38-2 | Metals - Extractable; | Metals - Extractable; EMS code: AS-E | Chemical | Extractable | LAB
AS-I | mg/dm2/d | Arsenic-Insoluble | 5039 | N/P Dig, ICP | Aliquot of prepared sample is digested first with nitric acid then with perchloric acid. Analysis by ICP. (a) For air particulates, preliminary results are reported in mg/L and loading results as mg/dm2/d. (b) For Hi Vol samples preliminary results are same as reported in mg/L and loading results as ug/m3. | 0.001 | Y | Y | Y | 134 | 12/5/1997 | 5/15/2000 | 11/18/2000 1:30:11 AM | 1996-04-12 | 6/3/1998 6:51:34 PM | EMS |  | MassFlux | mg/dm²/d | Y | (mass flux) |  | Arsenic-Insoluble (mass flux) | 7440-38-2 | Metals - Insoluble; | Metals - Insoluble; EMS code: AS-I | Chemical | Insoluble | LAB
AS-I | mg/dm2/d | Arsenic-Insoluble | 5035 | N/P Dig,Hydride ICP | N/P Dig;Hydride ICP | 1e-05 | Y | Y | Y | 43 |  |  | 10/23/1996 3:43:34 PM | 1996-04-12 | 6/3/1998 6:51:42 PM | EMS |  | MassFlux | mg/dm²/d | Y | (mass flux) |  | Arsenic-Insoluble (mass flux) | 7440-38-2 | Metals - Insoluble; | Metals - Insoluble; EMS code: AS-I | Chemical | Insoluble | LAB
AS-I | mg/dm2/d | Arsenic-Insoluble | EQ15 | EQUIS (mg/dm2/d) | EQUIS (mg/dm2/d) | 0.002 | N | Y | Y | 81 |  |  | 10/22/1996 10:18:48 PM | 1996-04-12 | 6/3/1998 6:51:41 PM | EMS |  | MassFlux | mg/dm²/d | Y | (mass flux) |  | Arsenic-Insoluble (mass flux) | 7440-38-2 | Metals - Insoluble; | Metals - Insoluble; EMS code: AS-I | Chemical | Insoluble | LAB
AS-L | mg/L | Leachable Arsenic after EPA TCLP extraction of solid sample | 1312 | EPA 1311/6010B (ICPMS) | This analysis is carried out in accordance with the extraction procedure outlined in "Test Methods for Evaluating Solid Waste - Physical/Chemical Methods Volume 1C" SW-846 EPA Method 1311, published by the United States Environmental Protection Agency (EPA). In summary, the sample is extracted at a 20:1 liquid to solids ratio for 16 to 20 hours using either extraction fluid #1 (glacial acetic acid, water and sodium hydroxide) or extraction fluid #2 (glacial acetic acid), depending on the pH of the original sample. The extract is then filtered through a 0.6 to 0.8 micron glass fibre filter and analysed using inductively coupled plasma - optical emission spectrophotometry (EPA Method 6010B) | 1 | Y | Y | Y | 16 |  |  | 8/24/2016 2:59:13 PM | 2011-04-27 |  | RDGIBSON |  | Concentration | mg/L | Y | (fl. conc.) |  | Leachable Arsenic after EPA TCLP extraction of solid sample (fl. conc.) | 7440-38-2 | Metals - Extractable; | Metals - Extractable; EMS code: AS-L | Chemical | Leachable | LAB
AS-S | mg/dm2/d | Arsenic-Soluble | 5045 | ICP Analysis | ICP a) For Air particulates analyzed without an acid digestion step, preliminary results are reported in mg/L and loading results as mg/dm2/d. b) For HiVol samples, preliminary results are reported as mg/L and loading results as ug/m3. | 0.005 | Y | Y | Y | 39 | 11/28/1997 | 5/15/2000 | 11/18/2000 1:30:11 AM | 1998-10-28 |  | NPEPPIN |  | MassFlux | mg/dm²/d | Y | (mass flux) |  | Arsenic-Soluble (mass flux) | 7440-38-2 | Metals - Soluble; | Metals - Soluble; EMS code: AS-S | Chemical | Soluble | LAB
AS-S | mg/dm2/d | Arsenic-Soluble | EQ15 | EQUIS (mg/dm2/d) | EQUIS (mg/dm2/d) | 0.0001 | N | Y | Y | 85 |  |  | 10/22/1996 10:18:48 PM | 1996-04-12 | 6/3/1998 6:51:34 PM | EMS |  | MassFlux | mg/dm²/d | Y | (mass flux) |  | Arsenic-Soluble (mass flux) | 7440-38-2 | Metals - Soluble; | Metals - Soluble; EMS code: AS-S | Chemical | Soluble | LAB
AS-S | mg/dm2/d | Arsenic-Soluble | X170 | ICP | This code is a catch-all for a large variety of ICP methods including water and also solids digested into solutions as an intermediate result. Different procedures will have greatly different detection limits. (a) For air particulates, preliminary results are reported in mg/L and loading results as mg/dm2/d. (b) For Hi Vol samples preliminiary results are same as reported in mg/L and loading results as ug/m3. (c) Note that for Rain samples there is no filtration nor digestion steps. | 0.005 | Y | Y | Y | 105 |  |  | 2/13/1998 2:28:22 AM | 1996-04-12 | 6/3/1998 6:51:34 PM | EMS |  | MassFlux | mg/dm²/d | Y | (mass flux) |  | Arsenic-Soluble (mass flux) | 7440-38-2 | Metals - Soluble; | Metals - Soluble; EMS code: AS-S | Chemical | Soluble | LAB
AS-S | mg/dm2/d | Arsenic-Soluble | 5043 | Hydride ICP | Hydride ICP | 5e-05 | Y | Y | Y | 43 |  |  | 10/24/1996 9:37:34 AM | 1996-04-12 | 6/3/1998 6:51:34 PM | EMS |  | MassFlux | mg/dm²/d | Y | (mass flux) |  | Arsenic-Soluble (mass flux) | 7440-38-2 | Metals - Soluble; | Metals - Soluble; EMS code: AS-S | Chemical | Soluble | LAB
AS-T | mg/L | Arsenic Total | X078 | Aqua Reg.Dig,Hydr.gen,AA | Aqua Reg.Dig;Hydr.gen;AA | 0.001 | Y | Y | Y | 16113 |  |  | 5/19/2015 3:10:36 PM | 1996-04-12 | 6/3/1998 6:51:34 PM | EMS |  | Concentration | mg/L | Y | (fl. conc.) |  | Arsenic Total (fl. conc.) | 7440-38-2 | Metals - Total; | Metals - Total; EMS code: AS-T | Chemical | Total | LAB
AS-T | ug/g | Arsenic Total | FS02 | D&G,Nitric/HCl:GF | Samples are dried at 60 C, then either sieved to -100 mesh (150 um) and rolled to homogenize, or finely ground with a mortar and pestle or a shatter box laboratory mill and rolled (depending on the submitter's requirements). The samples are then weighed into Teflon digestion liners and digested with HNO3 and HCl in a microwave oven. Samples are cooled, volumized, and allowed to settle overnight before being decanted. Analysis is then performed via Graphite Furnace Atomic Absorption Spectrometry (GF-AAS) (low level analysis). GFAA Spectrometry involves heating the sample in three or more stages. First, a low current heats the graphite tube to dry the sample. The second, or charring stage, destroys the organic matter and volatilizes other components at an intermediate temperature. Finally, a high current heats the tube to incandescence and, in an inert atmosphere, atomizes the element being determined. Additional stages may be added to drying and charring, and to cool and clean the tubes in between samples. The resultant ground-state atomic vapour absorbs monochromatic radiation from the source. A photoelectric detector measures the intensity of the transmitted radiation characteristic of the element being determined, which is inversely proportional to the concentration of ground-state atoms in the optical path over a limited range. Low level Cadmium and Lead are determined using Deuterium lamp background correction system whereas Arsenic and Selenium are determined by using Zeeman background correction system. | 0.5 | Y | Y | Y | 29 | 2/25/1998 4:36:00 PM | 10/22/2001 8:48:00 AM | 1/9/2002 1:28:48 AM | 1999-04-01 |  | NPEPPIN |  | MassPerMass | mg/kg | Y | (mass conc.) |  | Arsenic Total (mass conc.) | 7440-38-2 | Metals - Total; | Metals - Total; EMS code: AS-T | Chemical | Total | LAB
AS-T | mg/L | Arsenic Total | METN | MET-T-NP-CCMS-VA | Ultra trace metals in water are analyzed by CRC ICPMS, based on US EPA Method 6020B (July 2014). The detection limits provided can only be met for undigested samples. This procedure is intended for colorless, non-turbid, acid-preserved water samples (i.e. pristine water samples), having turbidity < 1 NTU and no odor. Where turbidity exceeds 1 NTU, and/or the sample is colored and has an odor, results may be biased low compared to true Total Metals concentrations. ALS recommends that turbidity analysis be requested on samples submitted for this test to aid with interpretation of results. Where turbidity is <1NTU, undigested metals are equivalent to total metals concentrations. | 2e-05 | Y | Y | Y | 3648 | 4/3/2012 | 8/19/2024 | 8/26/2024 1:03:06 PM | 2019-08-20 |  | KATFRASE |  | Concentration | mg/L | Y | (fl. conc.) |  | Arsenic Total (fl. conc.) | 7440-38-2 | Metals - Total; | Metals - Total; EMS code: AS-T | Chemical | Total | LAB
AS-T | ug/m3 | Arsenic Total | F088 | Aliquot,HNO3/HCl dig, ICPMS,loading | Aliquot microwave digested with 3:1 HNO3/HCL. Analysis by ICPMS. For HiVol samples preliminary result are reported in mg/L and loading results are reported in ug/m3 |  | Y | Y | Y | 565 | 8/17/2000 10:46:00 AM | 12/14/2006 | 12/30/2006 1:01:35 AM | 2000-08-16 |  | NPEPPIN |  | AirConcentration | µg/m³ | Y | (air conc.) |  | Arsenic Total (air conc.) | 7440-38-2 | Metals - Total; | Metals - Total; EMS code: AS-T | Chemical | Total | LAB
AS-T | mg/L | Arsenic Total | X561 | Total metals: FA, HNO3, HR-ICPMS (No Prep)Prep) | Ultra trace metals in water are analyzed by high resolution inductively coupled plasma mass spectrometry (HR-ICPMS), modified from US EPA 200.8, (Revision 5.5). The detection limits provided can only be met for undigested samples. This procedure is intended for colorless, non-turbid, acid-preserved water samples (i.e. pristine water samples), having turbidity < 1 NTU and no odor. Where turbidity exceeds 1 NTU, and/or the sample is colored and has an odor, results may be biased low compared to true Total Metals concentrations. ALS recommends that turbidity analysis be requested on samples submitted for this test to aid with interpretation of results. | 2e-05 | Y | Y | Y | 2756 | 6/8/2015 | 10/5/2023 | 1/31/2024 2:37:41 PM | 2015-10-22 |  | RDGIBSON |  | Concentration | mg/L | Y | (fl. conc.) |  | Arsenic Total (fl. conc.) | 7440-38-2 | Metals - Total; | Metals - Total; EMS code: AS-T | Chemical | Total | LAB
AS-T | mg/L | Arsenic Total | X351 | LA:HNO3:AA | Samples are lab-acidified (HNO3); digested HNO3; analysis by AA. | 0.2 | Y | Y | Y | 86 |  |  | 8/26/2004 11:46:08 AM | 1996-04-12 | 6/3/1998 6:51:34 PM | EMS |  | Concentration | mg/L | Y | (fl. conc.) |  | Arsenic Total (fl. conc.) | 7440-38-2 | Metals - Total; | Metals - Total; EMS code: AS-T | Chemical | Total | LAB
AS-T | mg/L | Arsenic Total | X285 | HNO3 Dig:ICP Analysis | HNO3 Dig:ICP Analysis | 0.003 | Y | Y | Y | 230 |  |  | 4/30/2004 1:59:45 PM | 1996-04-12 | 6/3/1998 6:51:34 PM | EMS |  | Concentration | mg/L | Y | (fl. conc.) |  | Arsenic Total (fl. conc.) | 7440-38-2 | Metals - Total; | Metals - Total; EMS code: AS-T | Chemical | Total | LAB
AS-T | mg/L | Arsenic Total | 0040 | HNO3 Dig: ICP Analysis | HNO3 Dig: ICP Analysis | 0.3 | Y | Y | Y | 4352 | 10/17/2000 | 6/25/2010 3:15:00 PM | 7/22/2010 2:07:51 PM | 1996-04-12 | 6/3/1998 6:51:41 PM | EMS |  | Concentration | mg/L | Y | (fl. conc.) |  | Arsenic Total (fl. conc.) | 7440-38-2 | Metals - Total; | Metals - Total; EMS code: AS-T | Chemical | Total | LAB
AS-T | ug/m3 | Arsenic Total | EQ40 | EQUIS (mg/m3) | EQUIS (mg/m3) | 0.1 | N | Y | Y | 2365 |  |  | 10/22/1996 11:59:06 PM | 1996-04-12 | 6/3/1998 6:51:34 PM | EMS |  | AirConcentration | µg/m³ | Y | (air conc.) |  | Arsenic Total (air conc.) | 7440-38-2 | Metals - Total; | Metals - Total; EMS code: AS-T | Chemical | Total | LAB
AS-T | mg/L | Arsenic Total | TCLP | Toxicity Characteristic Leaching Procedure | The TCLP procedure is designed to simulate the climatic leaching action | 0.1 | Y | Y | Y | 22 | 12/8/2009 7:49:00 PM | 9/28/2011 5:53:00 PM | 11/3/2011 11:10:18 PM | 2009-03-05 |  | JCHUDAK |  | Concentration | mg/L | Y | (fl. conc.) |  | Arsenic Total (fl. conc.) | 7440-38-2 | Metals - Total; | Metals - Total; EMS code: AS-T | Chemical | Total | LAB
AS-T | mg/L | Arsenic Total | 1349 | Tot Metals in Seawater by CRC ICPMS (BC) | Tot Metals in Seawater by CRC ICPMS (BC) | 0.0004 | Y | Y | Y | 98 | 8/16/2019 | 12/5/2021 | 12/13/2021 12:34:16 PM | 2019-09-30 |  | KATFRASE |  | Concentration | mg/L | Y | (fl. conc.) |  | Arsenic Total (fl. conc.) | 7440-38-2 | Metals - Total; | Metals - Total; EMS code: AS-T | Chemical | Total | LAB
AS-T | mg/L | Arsenic Total | F024 | HNO3/HCl,GFAA | Dissolved and extractable water samples are ready for analysis once collected. Total samples are autoclaved with HNO3 and hydrochloric acid (HCl) in capped Teflon vials. Analysis is performed via Inductively Coupled Argon Plasma Atomic Emission Spectrometry (ICP-AES) (simultaneous multielement analysis), Graphite Furnace Atomic Absorption Spectrometry (GF-AAS) (low level analysis) and/or Flame Atomic Absorption Spectrometry (F-AAS). | 0.0005 | Y | Y | Y | 942 | 8/2/1996 4:11:00 PM | 9/12/2003 3:16:00 PM | 12/1/2003 5:30:16 PM | 1996-09-17 | 6/3/1998 6:51:45 PM | MPAK |  | Concentration | mg/L | Y | (fl. conc.) |  | Arsenic Total (fl. conc.) | 7440-38-2 | Metals - Total; | Metals - Total; EMS code: AS-T | Chemical | Total | LAB
AS-T | mg/L | Arsenic Total | F038 | HNO3/HCl dig,ICP-AES | Dissolved and extractable water samples are ready for analysis once collected. Total samples are autoclaved with HNO3 and hydrochloric acid (HCI) in capped Teflon vials. Analysis is performed via Inductively Coupled Argon Plasma Atomic Emission Spectrometry (ICP-AES)(simultaneous multielement analysis), Graphite Furnace Atomic Absorption Spectrometry (GF-AAS) (low level analysis) and/or Flame Atomic Absorption Spectrometry (F-AAS). | 0.06 | Y | Y | Y | 5425 | 8/1/1996 6:04:00 PM | 5/6/2008 | 9/11/2019 8:47:42 AM | 1997-05-08 | 6/3/1998 6:51:46 PM | MPAK |  | Concentration | mg/L | Y | (fl. conc.) |  | Arsenic Total (fl. conc.) | 7440-38-2 | Metals - Total; | Metals - Total; EMS code: AS-T | Chemical | Total | LAB
AS-T | ug/g | Arsenic Total | 0190 | D&G:N/P Dig:ICP Anal. | Sample is dried and grind-mixed before nitric perchloric digestion followed by ICP analysis. | 25 | Y | Y | Y | 854 |  |  | 10/24/1996 3:05:58 PM | 1996-04-12 | 6/3/1998 6:51:34 PM | EMS |  | MassPerMass | mg/kg | Y | (mass conc.) |  | Arsenic Total (mass conc.) | 7440-38-2 | Metals - Total; | Metals - Total; EMS code: AS-T | Chemical | Total | LAB
AS-T | ug/m3 | Arsenic Total | X137 | Arsine gen,535nm,Col-Air | Arsine gen;535nm;Col-Air | 0 | Y | Y | Y | 555 |  |  | 3/9/1998 8:54:23 AM | 1996-04-12 | 6/3/1998 6:51:41 PM | EMS |  | AirConcentration | µg/m³ | Y | (air conc.) |  | Arsenic Total (air conc.) | 7440-38-2 | Metals - Total; | Metals - Total; EMS code: AS-T | Chemical | Total | LAB
AS-T | ug/g | Arsenic Total | X228 | Hom,HNO3/HCl,HVAAS | Sample is not pre-dried, but is homogenized and digested with HNO3/HCl prior to analysis by hydride vapour AA Spectrophotometry (HVAAS). Moisture determined on separate aliquot is used to calculate dry weight result. | 0.01 | Y | Y | Y | 203 |  |  | 9/23/2003 9:46:53 PM | 1996-04-12 | 6/3/1998 6:51:41 PM | EMS |  | MassPerMass | mg/kg | Y | (mass conc.) |  | Arsenic Total (mass conc.) | 7440-38-2 | Metals - Total; | Metals - Total; EMS code: AS-T | Chemical | Total | LAB
AS-T | mg/L | Arsenic Total | X452 | Hom,Nitr/Perchl,Hydr,AA | Hom;Nitr/Perchl;Hydr;AA | 0.025 | Y | Y | Y | 29 |  |  | 10/24/1996 1:00:47 PM | 1996-04-12 | 6/3/1998 6:51:41 PM | EMS |  | Concentration | mg/L | Y | (fl. conc.) |  | Arsenic Total (fl. conc.) | 7440-38-2 | Metals - Total; | Metals - Total; EMS code: AS-T | Chemical | Total | LAB
AS-T | ug/g wet | Arsenic Total | X458 | HNO3/HCL Dig:HVAAS(totl) | HNO3/HCL Dig:HVAAS(totl) | 0.05 | Y | Y | Y | 2 |  |  | 10/23/1996 11:03:49 PM | 1996-04-12 | 6/3/1998 6:51:34 PM | EMS |  | MassPerMass | µg/g | Y | (mass conc.) | (wet) | Arsenic Total (wet) (mass conc.) | 7440-38-2 | Metals - Total; | Metals - Total; EMS code: AS-T | Chemical | Total | LAB
AS-T | mg/L | Arsenic Total | F068 | HNO3dig,TeflonTube/Autoclave,ICP-MS | This method is applicable to determiniation of dissolved, extractable, and total recoverable metals in ground waters, surface waters, drinking waters, and wastewaters. Dissolved, extractable, and total recoverable sample fractions are acidified to pH <1.5 with analytical grade nitric acid (HNO3) to prevent the metals from adsorbing to the container material or precipitating. Extractable and total recoverable water samples are collected in 250 ml acid washed polyethylene bottles and are preserved with Nitric acid (0.5 ml acid for 100 mLs of sample). All preserved samples are stored at room temperature and can be held for up to 6 months. Unfiltered, un-preserved samples are stored at 4C for no more than 30 days. For dissolved and extractable water fractions, no sample preparation is necessary. Samples are simply poured in autosampler tubes (i.e. 15 ml centrifuge tubes), and placed on the autosampler and analyzed. Samples for total recoverable metals are digested with Nitric acid in special Teflon tubes in an autoclave prior to analyses. | 0.0001 | Y | Y | Y | 3948 | 6/1/1999 1:23:00 PM | 7/4/2008 9:48:00 AM | 7/30/2024 6:32:12 AM | 1999-06-28 | 6/30/1999 1:05:24 PM | NPEPPIN | NPEPPIN | Concentration | mg/L | Y | (fl. conc.) |  | Arsenic Total (fl. conc.) | 7440-38-2 | Metals - Total; | Metals - Total; EMS code: AS-T | Chemical | Total | LAB
AS-T | ug/g | Arsenic Total | 0196 | D&G: LTA/HCl: ICP | D&G: LTA/HCl: ICP | 25 | Y | Y | Y | 26 |  |  | 10/24/1996 12:50:32 PM | 1996-04-12 | 6/3/1998 6:51:34 PM | EMS |  | MassPerMass | mg/kg | Y | (mass conc.) |  | Arsenic Total (mass conc.) | 7440-38-2 | Metals - Total; | Metals - Total; EMS code: AS-T | Chemical | Total | LAB
AS-T | ug/g wet | Arsenic Total | X552 | Metals in Biota (wet), total, acid digestion, ICP-SFMS | One gram of wet, homogenized biota is weighed in a microwave TFM vessel, and 8mL of Nitric Acid and 1mL of Hydrogen Peroxide are added. The sample is digested at 200 degrees C for 15 minutes in a closed vessel high pressure microwave oven. The sample is cooled, the volume adjusted to 100mL and introduced to an inductively coupled argon plasma-sector field mass spectrometer (ICP-SFMS) by pneumatic nebulization. Each respective element is measured at a specific mass to charge ratio, m/z value expressed in amu (atomic mass units), and at a specified instrument resolution. LR: Low Resolution ~300, MR: Mid Resolution ~4000. Ref: 02-2704, National Laboratory for Environmental Testing, CCIW, Environment Canada, Burlington. Requested April 2003. | 0.002 | Y | Y | Y | 37 | 8/9/2016 | 10/16/2023 | 11/28/2023 2:21:32 PM | 2004-01-19 | 4/19/2021 1:22:00 PM | UMULLER | EMS | MassPerMass | µg/g | Y | (mass conc.) | (wet) | Arsenic Total (wet) (mass conc.) | 7440-38-2 | Metals - Total; | Metals - Total; EMS code: AS-T | Chemical | Total | LAB
AS-T | mg/L | Arsenic Total | 0181 | HCl/K2S2O8, Hydride, ICP | HCl/K2S2O8, Hydride; ICP | 0.001 | Y | Y | Y | 2390 |  |  | 2/4/2000 9:38:25 AM | 1996-04-12 | 6/3/1998 6:51:34 PM | EMS |  | Concentration | mg/L | Y | (fl. conc.) |  | Arsenic Total (fl. conc.) | 7440-38-2 | Metals - Total; | Metals - Total; EMS code: AS-T | Chemical | Total | LAB
AS-T | mg/L | Arsenic Total | MSEA | Metals in Seawater (undigested) by (CRC) ICP-MS | This method describes the multi-elemental determination of trace elements by Collision/Reaction Cell (CRC) ICP-MS. The method measures ions produced by a radio-frequency inductively coupled plasma. Analyte species originating in a liquid are pneumatically nebulized and the resulting aerosol transported by argon gas into the plasma torch. The ions produced are entrained in the plasma gas and extracted, by means of a differentially pumped vacuum interface, into a mass spectrometer. The ions flow through a Collision/Reaction Cell (CRC) where depending on the mode of analysis either the ions passes through the cell unhindered (no gas - classical ICP-MS mode), or collisions/reactions chemistry takes place between the ions and a neutral gas such as hydrogen or helium. Any polyatomic ions which give interference on the analytes are removed. The ions are then sorted according to their mass-to-charge ratios by a quadrupole mass spectrometer. The ions are then sorted according to their mass-to-charge ratios by a quadrupole mass spectrometer. The ions transmitted through the quadrupole are quantified by a channel electron multiplier and the ion information processed by a data handling system. Although most polyatomic interference is removed, any additional interference must be assessed and valid corrections applied or the data flagged to indicate problems. Interference correction must include compensation for background ions contributed by the plasma gas, reagents, and constituents of the sample matrix. Instrument drift as well as suppressions or enhancements of instrument response caused by the sample matrix must be corrected for by the use of internal standards. | 0.0001 | Y | Y | Y | 441 | 8/12/2009 12:08:00 AM | 4/29/2024 | 6/7/2024 10:03:16 AM | 2009-09-08 |  | RDGIBSON |  | Concentration | mg/L | Y | (fl. conc.) |  | Arsenic Total (fl. conc.) | 7440-38-2 | Metals - Total; | Metals - Total; EMS code: AS-T | Chemical | Total | LAB
AS-T | mg/L | Arsenic Total | X089 | Nitric/Ni Dig,Graph. Fur | Nitric/Ni Dig;Graph. Fur | 0.001 | Y | Y | Y | 6 |  |  | 1/30/1997 2:12:24 PM | 1996-04-12 | 6/3/1998 6:51:34 PM | EMS |  | Concentration | mg/L | Y | (fl. conc.) |  | Arsenic Total (fl. conc.) | 7440-38-2 | Metals - Total; | Metals - Total; EMS code: AS-T | Chemical | Total | LAB
AS-T | mg/L | Arsenic Total | F048 | Acid dig,conc'n,nebul.,ICP-OES | This method provides an automated procedure for the Inductively Coupled Argon Plasma Optical Emission Spectrophotometric (ICP-OES) determination and quantitation of trace amounts of metals in surface, ground and precipitation waters after a manual acid digestion and concentration procedure and ultrasonic nebulisation. Organically-bound metals and metal complexes are converted to metal ions by digestion with nitric acid [method details says HCl]. To increase the signal detected by the ICP, the samples are concentrated 10 times. The liquid sample is pumped onto the piezoelectric transducer, where the ultrasonic action efficiently converts it to a fine, dense aerosol. The nebulizer carrier gas flow then transports the wet aerosol through a heated U-tube, where the solvent is vaporized. Subsequently, most of the solvent vapour is removed from the aerosol stream by passing the stream through a chilled condenser. The result is a dry, analyte-laden aerosol which is introduced to the plasma. There, the high temperature desolvates the aerosol and dissociates the sample material into individual atoms and ions, which are excited to emit light at wave lengths characteristic of the elements present. Dissolved samples are field filtered through 0.45 micron cellulose acetate filter prior to acid preservation. Samples for extractable analysis are acidified with 2 mL concentrated HNO3 per litre and allowed to stand a minimum of 24 hours before analysis. Clients providing these samples are advised to preserve samples at the collection site. | 0.2 | Y | Y | Y | 29 |  |  | 7/6/2001 11:10:09 AM | 1998-11-13 |  | NPEPPIN |  | Concentration | mg/L | Y | (fl. conc.) |  | Arsenic Total (fl. conc.) | 7440-38-2 | Metals - Total; | Metals - Total; EMS code: AS-T | Chemical | Total | LAB
AS-T | mg/L | Arsenic Total | 0042 | GI HNO3 Dig: ICP | Sample aliquot taken from unfiltered unpreserved General Ions bottle; HNO3 digestion; ICP analysis. | 0.04 | Y | Y | Y | 6677 | 8/12/1993 | 7/5/2019 3:21:00 PM | 7/12/2019 6:49:13 PM | 1996-04-12 | 6/3/1998 6:51:34 PM | EMS |  | Concentration | mg/L | Y | (fl. conc.) |  | Arsenic Total (fl. conc.) | 7440-38-2 | Metals - Total; | Metals - Total; EMS code: AS-T | Chemical | Total | LAB
AS-T | mg/dm2/d | Arsenic Total | X236 | ArsineGen,535nm,Col-Air2 | ArsineGen;535nm;Col-Air2 | 0.0001 | Y | Y | Y | 289 |  |  | 10/29/2003 1:05:56 PM | 1996-04-12 | 6/3/1998 6:51:34 PM | EMS |  | MassFlux | mg/dm²/d | Y | (mass flux) |  | Arsenic Total (mass flux) | 7440-38-2 | Metals - Total; | Metals - Total; EMS code: AS-T | Chemical | Total | LAB
AS-T | ug/g | Arsenic Total | FB04 | Freeze Dry,Hom,Acid Dig,ICP | Sample moisture is determined by weighing samples before and after freeze-drying. Freeze-dried samples are ground in a Braun blender before being weighed into Teflon PFA liners and digested with concentrated HNO3, concentrated HCl, and 50% H2O2 in a microwave oven; which results in a complete sample dissolution. Samples are collected and volumized with Type I water. Analysis is performed via Inductively Coupled Argon Plasma Atomic Emission Spectrometry (ICP-AES) (simultaneous multielement analysis), Graphite Furnace Atomic Absorption Spectrometry (GF-AAS) (low level Cd and Pb analysis), Cold Vapor Atomic Fluorescence Spectrometry (CV-AFS) (Mercury analysis) and/or Flame Atomic Absorption Spectrometry (F-AAS) (K analysis). | 4 | Y | Y | Y | 241 | 8/22/1996 6:11:00 PM | 2/1/2002 2:30:00 PM | 6/23/2003 1:15:27 PM | 1997-02-27 | 6/3/1998 6:51:46 PM | MPAK |  | MassPerMass | mg/kg | Y | (mass conc.) |  | Arsenic Total (mass conc.) | 7440-38-2 | Metals - Total; | Metals - Total; EMS code: AS-T | Chemical | Total | LAB
AS-T | ug/m3 | Arsenic Total | 5012 | N/P Dig:HGA Analysis | N/P Dig:HGA Analysis | 0.2 | N | Y | Y | 628 |  |  | 10/24/1996 6:54:32 AM | 1996-04-12 | 6/3/1998 6:51:34 PM | EMS |  | AirConcentration | µg/m³ | Y | (air conc.) |  | Arsenic Total (air conc.) | 7440-38-2 | Metals - Total; | Metals - Total; EMS code: AS-T | Chemical | Total | LAB
AS-T | ug/g | Arsenic Total | X492 | Homo,HCl-HNO3 dig,ICP-OES,dry calc | Homogenize samples and take sub-samples. For Total Recoverable Metals Digest with Aqua Regia. Analyses by ICPOES (Inductively Coupled Plasma Optical Emission Spectroscopy). Perform moisture analysis on separate sub-sample and calculate dry value for trace elements. | 100 | Y | Y | Y | 8 | 8/16/2022 6:31:00 PM | 8/16/2022 6:42:00 PM | 2/25/2023 1:41:51 PM | 1998-11-13 |  | NPEPPIN |  | MassPerMass | mg/kg | Y | (mass conc.) |  | Arsenic Total (mass conc.) | 7440-38-2 | Metals - Total; | Metals - Total; EMS code: AS-T | Chemical | Total | LAB
AS-T | ppm (S) | Arsenic Total | MS61 | ALS Minerals ME-MS61 | ALS Minerals ME-MS61. A prepared sample is digested with HNO3; HF; and HClO4 in a graphite heating block to near dryness. After cooling; the sample is boiled in diluted HCl. Bulk to volume with deionized water; mixed and analyzed by ICP-AES and ICP-MS for all elements. | 0.2 | Y | Y | Y | 0 |  |  |  | 2024-01-29 |  | EMS |  | DimensionlessRatio | ppm | N | (ratio) |  | Arsenic Total (ratio) | 7440-38-2 | Metals - Total; | Metals - Total; EMS code: AS-T | Chemical | Total | LAB
AS-T | ug/g | Arsenic Total | 0194 | D&G:N/P Dig:Hydr/ICP LL | D&G:N/P Dig:Hydr/ICP LL | 0.1 | Y | Y | Y | 93 |  |  | 10/24/1996 2:08:12 PM | 1996-04-12 | 6/3/1998 6:51:34 PM | EMS |  | MassPerMass | mg/kg | Y | (mass conc.) |  | Arsenic Total (mass conc.) | 7440-38-2 | Metals - Total; | Metals - Total; EMS code: AS-T | Chemical | Total | LAB
AS-T | ug/g wet | Arsenic Total | X216 | Hom,Nitr/Perchl,Hydr,AA | Hom;Nitr/Perchl;Hydr;AA | 0.005 | Y | Y | Y | 73 |  |  | 10/24/1996 10:01:24 AM | 1996-04-12 | 6/3/1998 6:51:41 PM | EMS |  | MassPerMass | µg/g | Y | (mass conc.) | (wet) | Arsenic Total (wet) (mass conc.) | 7440-38-2 | Metals - Total; | Metals - Total; EMS code: AS-T | Chemical | Total | LAB
AS-T | ug/g | Arsenic Total | EQ02 | EQUIS (ug/g) | EQUIS (ug/g) | 0.1 | N | Y | Y | 915 |  |  | 10/23/1996 12:08:28 AM | 1996-04-12 | 6/3/1998 6:51:34 PM | EMS |  | MassPerMass | mg/kg | Y | (mass conc.) |  | Arsenic Total (mass conc.) | 7440-38-2 | Metals - Total; | Metals - Total; EMS code: AS-T | Chemical | Total | LAB
AS-T | mg/L | Arsenic Total | EQ01 | EQUIS (mg/L) | EQUIS (mg/L) | 0.001 | N | Y | Y | 4139 |  |  | 10/24/1996 3:14:37 PM | 1996-04-12 | 6/3/1998 6:51:41 PM | EMS |  | Concentration | mg/L | Y | (fl. conc.) |  | Arsenic Total (fl. conc.) | 7440-38-2 | Metals - Total; | Metals - Total; EMS code: AS-T | Chemical | Total | LAB
AS-T | t/d | Arsenic Total | CALC | Calculation (t/d) | Calculation (t/d) | 0 | Y | Y | Y | 3707 |  |  | 7/3/2024 12:59:25 PM | 1996-04-12 | 6/3/1998 6:51:34 PM | EMS |  | MassFlow | t/d | Y | (mass flow) |  | Arsenic Total (mass flow) | 7440-38-2 | Metals - Total; | Metals - Total; EMS code: AS-T | Chemical | Total | LAB
AS-T | mg/L | Arsenic Total | 0185 | Hydr.At.Absorp.Spec(H2O) | Hydr.At.Absorp.Spec(H2O) | 0.5 | Y | Y | Y | 649 | 9/23/1996 | 3/17/1999 11:20:00 AM | 3/8/2001 9:49:35 AM | 1996-04-12 | 6/3/1998 6:51:41 PM | EMS |  | Concentration | mg/L | Y | (fl. conc.) |  | Arsenic Total (fl. conc.) | 7440-38-2 | Metals - Total; | Metals - Total; EMS code: AS-T | Chemical | Total | LAB
AS-T | mg/L | Arsenic Total | 5077 | N/P Dig,Hydride ICP | N/P Dig;Hydride ICP | 0.001 | Y | Y | Y | 29 |  |  | 12/5/1996 3:04:46 PM | 1996-04-12 | 6/3/1998 6:51:42 PM | EMS |  | Concentration | mg/L | Y | (fl. conc.) |  | Arsenic Total (fl. conc.) | 7440-38-2 | Metals - Total; | Metals - Total; EMS code: AS-T | Chemical | Total | LAB
AS-T | mg/L | Arsenic Total | F037 | K2S2O8/HCl dig, Hydride,ICP | The water sample is digested with potassium persulphate and concentrated hydrochloric acid. Samples, blanks and standards are placed in an automatic sampler and analyzed by emission spectroscopy utilizing induction coupled argon plasma as an excitation source. All forms of arsenic, selenium and antimony are broken down with acidic persulphate to arsenate, selenate and antimonate. Sodium borohydride is required to reduce the arsenate to arsenite. Hydrochloric acid reduces the selenate to selenite. Potassium iodide is required to reduce the antimonate to antimonite. Hydrides of arsenic, selenium and antimony are formed by the action of sodium borohydride in an acidic medium, and analysed by inductively coupled argon plasma optical emission spectrometry. | 0.0001 | Y | Y | Y | 10059 | 7/4/1997 2:35:00 PM | 12/1/2010 | 9/11/2019 8:46:38 AM | 1997-04-02 | 6/3/1998 6:51:46 PM | MPAK |  | Concentration | mg/L | Y | (fl. conc.) |  | Arsenic Total (fl. conc.) | 7440-38-2 | Metals - Total; | Metals - Total; EMS code: AS-T | Chemical | Total | LAB
AS-T | mg/L | Arsenic Total | F121 | Acid,ICPMS | Acidified, no digestion, ICPMS | 2e-05 | Y | Y | Y | 8440 | 3/12/2010 7:57:00 PM | 8/15/2024 9:10:00 PM | 8/26/2024 8:06:03 AM | 2002-07-17 | 7/16/2012 3:28:03 PM | UMULLER | RDGIBSON | Concentration | mg/L | Y | (fl. conc.) |  | Arsenic Total (fl. conc.) | 7440-38-2 | Metals - Total; | Metals - Total; EMS code: AS-T | Chemical | Total | LAB
AS-T | ug/m3 | Arsenic Total | 5312 | Tef:N/P Dig:ICP Analysis | A nitric/perchloric acid digestion if used to bring the metals into solution prior to ICP analysis. For Hi-Vol samples preliminary results are reported as mg/L and 24 hour loading results as ug/m3. | 0.1 | Y | Y | Y | 1689 | 11/7/1997 | 5/15/2000 | 11/18/2000 1:28:58 AM | 1996-04-12 | 6/3/1998 6:51:34 PM | EMS |  | AirConcentration | µg/m³ | Y | (air conc.) |  | Arsenic Total (air conc.) | 7440-38-2 | Metals - Total; | Metals - Total; EMS code: AS-T | Chemical | Total | LAB
AS-T | ppm (S) | Arsenic Total | 6041 | Aqua Regia by ICPMS | These were sediment/soils digested with Aqua-Regia (Nitric Acid & Hydrochloric Acids) and then analysis by ICPMS. Results would be reported in mg/kg and/or %. | 0.01 | Y | Y | Y | 21 | 10/13/2019 | 8/5/2020 | 8/10/2020 3:52:06 PM | 2019-09-06 | 10/21/2019 11:54:01 AM | KATFRASE | KATFRASE | DimensionlessRatio | ppm | N | (ratio) |  | Arsenic Total (ratio) | 7440-38-2 | Metals - Total; | Metals - Total; EMS code: AS-T | Chemical | Total | LAB
AS-T | ug/m3 | Arsenic Total | 5010 | N/P Dig,Hydride ICP | N/P Dig;Hydride ICP | 0.001 | Y | Y | Y | 110 |  |  | 10/24/1996 6:37:05 AM | 1996-04-12 | 6/3/1998 6:51:34 PM | EMS |  | AirConcentration | µg/m³ | Y | (air conc.) |  | Arsenic Total (air conc.) | 7440-38-2 | Metals - Total; | Metals - Total; EMS code: AS-T | Chemical | Total | LAB
AS-T | mg/L | Arsenic Total | X545 | FA,HNO3,ICPMS | Analytes nebulized followed by ion entrained in plasma gas and extracted by means of a differentially pumped vacuum interface into a mass spectrometer | 0.001 | Y | Y | Y | 30 |  |  | 10/30/2014 8:31:17 AM | 2001-09-18 | 9/18/2001 1:36:23 PM | NPEPPIN | NPEPPIN | Concentration | mg/L | Y | (fl. conc.) |  | Arsenic Total (fl. conc.) | 7440-38-2 | Metals - Total; | Metals - Total; EMS code: AS-T | Chemical | Total | LAB
AS-T | ug/g | Arsenic Total | 0193 | Hom,N/P Dig,Hydr/ICP LL | Hom;N/P Dig;Hydr/ICP LL | 0.1 | Y | Y | Y | 207 |  |  | 10/24/1996 9:26:17 AM | 1996-04-12 | 6/3/1998 6:51:34 PM | EMS |  | MassPerMass | mg/kg | Y | (mass conc.) |  | Arsenic Total (mass conc.) | 7440-38-2 | Metals - Total; | Metals - Total; EMS code: AS-T | Chemical | Total | LAB
AS-T | ug/g | Arsenic Total | 0192 | Hom, Nitric/Perchlor,ICP | Sample already in a dry state so does not require drying before homogenization and nitric perchlorid digestion followed by ICP analysis. | 25 | Y | Y | Y | 548 | 2/13/2002 | 5/29/2002 | 1/14/2004 1:00:43 AM | 1996-04-12 | 6/3/1998 6:51:34 PM | EMS |  | MassPerMass | mg/kg | Y | (mass conc.) |  | Arsenic Total (mass conc.) | 7440-38-2 | Metals - Total; | Metals - Total; EMS code: AS-T | Chemical | Total | LAB
AS-T | ppm (S) | Arsenic Total | 6042 | Aqua Regia by ICPMS | These were sediment/soils digested with Aqua-Regia (Nitric Acid & Hydrochloric Acids) and then analysis by ICPMS. Results would be reported in ppm s. | 0.01 | Y | Y | Y | 71 | 5/14/2022 | 7/27/2024 | 7/29/2024 9:13:08 AM | 2022-06-28 |  | EMS |  | DimensionlessRatio | ppm | N | (ratio) |  | Arsenic Total (ratio) | 7440-38-2 | Metals - Total; | Metals - Total; EMS code: AS-T | Chemical | Total | LAB
AS-T | ug/g | Arsenic Total | X551 | Metals in Biota (dry), total, acid digestion, ICP-SFMS | 0.1 gram of dried, homogenized biota is weighed in a microwave TFM vessel, and 8mL of Nitric Acid and 1mL of Hydrogen Peroxide are added. The sample is digested at 200 degrees C for 15 minutes in a closed vessel high pressure microwave oven. The sample is cooled, the volume adjusted to 100mL and introduced to an iductively coupled argon plasma-sector field mass spectrometer (ICP-SFMS) by pneumatic nebulization. Each respective element is measured at a specific mass to charge ratio, m/z value expressed in amu (atomic mass units), and at a specified instrument resolution. LR: Low Resolution ~300, MR: Mid Resolution ~4000 and HR: Hich Resolution ~10000. Ref: Method number 02-2705, National Laboratory for Environmental Testing, CCIW, Environment Canada, Burlington. Requested April 2003. | 0.02 | Y | Y | Y | 57 | 8/11/2015 | 10/30/2023 | 11/28/2023 2:21:33 PM | 2004-01-19 |  | UMULLER |  | MassPerMass | mg/kg | Y | (mass conc.) |  | Arsenic Total (mass conc.) | 7440-38-2 | Metals - Total; | Metals - Total; EMS code: AS-T | Chemical | Total | LAB
AS-T | ug/g | Arsenic Total | X209 | Hom,Nitr/Perchl,Hydr,AA | Hom;Nitr/Perchl;Hydr;AA | 0.025 | Y | Y | Y | 35 |  |  | 10/24/1996 1:00:18 PM | 1996-04-12 | 6/3/1998 6:51:34 PM | EMS |  | MassPerMass | mg/kg | Y | (mass conc.) |  | Arsenic Total (mass conc.) | 7440-38-2 | Metals - Total; | Metals - Total; EMS code: AS-T | Chemical | Total | LAB
AS-T | ug/g | Arsenic Total | 0031 | ICP | ICP | 100 | Y | Y | Y | 1 |  |  | 10/23/1996 11:03:42 PM | 1996-04-12 | 6/3/1998 6:51:41 PM | EMS |  | MassPerMass | mg/kg | Y | (mass conc.) |  | Arsenic Total (mass conc.) | 7440-38-2 | Metals - Total; | Metals - Total; EMS code: AS-T | Chemical | Total | LAB
AS-T | ug/m3 | Arsenic Total | X188 | Impinger,Sel Ion Elect | Impinger;Sel Ion Elect | 0 | Y | Y | Y | 1929 |  |  | 3/9/1998 10:35:51 AM | 1996-04-12 | 6/3/1998 6:51:34 PM | EMS |  | AirConcentration | µg/m³ | Y | (air conc.) |  | Arsenic Total (air conc.) | 7440-38-2 | Metals - Total; | Metals - Total; EMS code: AS-T | Chemical | Total | LAB
AS-T | mg/L | Arsenic Total | F033 | Acid Dig,Centrifuge,ICPMS | For the determination of total recoverable elements in water or waste water, take 100 ml aliquot from a well mixed, acid preserved, sample containing not more than 0.25% (w/v) total solids and transfer to a 250 ml Griffin beaker. (If total solids are greater than 0.25% reduce the size of the aliquot by a proportionate amount). Add 1 ml of conc. nitric acid and 0.5 ml conc. hydrochloric acid. Heat on a hot plate until the volume has been reduced to 15-20ml, ensuring that the sample does not boil. Allow to cool and transfer to a 50 ml volumetric flask. Dilute to volume with ASTM type I water. Centrifuge the sample or allow to stand overnight to separate insoluble material. Prior analysis, pipette 20 ml into a 50 ml volumetric flask and dilute to volume with ASTM type I water. Analysis should be performed within seventy two hours of the completed preparation. Note that different laboratories may have very different field preservation, digestion procedures and detection limits for this method, so these should be documented. | 0.001 | Y | Y | Y | 41687 | 11/20/1997 9:30:00 AM | 8/22/2024 3:24:00 AM | 8/25/2024 9:47:36 PM | 1998-05-14 | 6/3/1998 6:51:49 PM | NPEPPIN |  | Concentration | mg/L | Y | (fl. conc.) |  | Arsenic Total (fl. conc.) | 7440-38-2 | Metals - Total; | Metals - Total; EMS code: AS-T | Chemical | Total | LAB
AS-T | ug/g wet | Arsenic Total | BPSY | MET-WET-BPSY-CCMS-VA | Metals in Animal Tissue and Vegetation (Biota) - Prescriptive. EPA 200.3/6020B This method is conducted following British Columbia Lab Manual method "Metals in Animal Tissue and Vegetation (Biota) - Prescriptive". Tissue samples are digested with nitric and hydrochloric acids; in combination with addition of hydrogen peroxide. Instrumental analysis is by collision cell inductively coupled plasma - mass spectrometry (modified from EPA Method 6020B). | 0.01 | Y | Y | Y | 2 | 11/7/2019 | 11/7/2019 | 6/10/2021 1:57:10 PM | 2021-06-04 |  | EMS |  | MassPerMass | µg/g | Y | (mass conc.) | (wet) | Arsenic Total (wet) (mass conc.) | 7440-38-2 | Metals - Total; | Metals - Total; EMS code: AS-T | Chemical | Total | LAB
AS-T | ug/g | Arsenic Total | FS82 | PBM,dry60C,10mesh,SALM-ICP-MS | PBM;dry60C;10mesh;SALM-ICP-MS | 0.2 | Y | Y | Y | 3505 | 5/21/2002 | 7/24/2024 | 7/30/2024 10:00:14 AM | 2002-11-12 |  | UMULLER |  | MassPerMass | mg/kg | Y | (mass conc.) |  | Arsenic Total (mass conc.) | 7440-38-2 | Metals - Total; | Metals - Total; EMS code: AS-T | Chemical | Total | LAB
AS-T | mg/L | Arsenic Total | X354 | LA:HNO3:HVAAS | Samples are lab-acidified with HNO3; HNO3 digestion; analysis by HVAAS. | 0.0002 | Y | Y | Y | 224 | 2/11/1998 | 8/22/2003 | 8/29/2003 3:50:14 PM | 1996-04-12 | 6/3/1998 6:51:34 PM | EMS |  | Concentration | mg/L | Y | (fl. conc.) |  | Arsenic Total (fl. conc.) | 7440-38-2 | Metals - Total; | Metals - Total; EMS code: AS-T | Chemical | Total | LAB
AS-T | ug/g | Arsenic Total | MTID | MET-DRY-CCMS-VA | Metals in Tissue by CRC ICPMS (DRY) Methodology: This method is conducted following British Columbia Lab Manual method "Metals in Animal Tissue and Vegetation (Biota) - Prescriptive". Tissue samples are homogenized and sub-sampled prior to hotblock digestion with nitric and hydrochloric acids, in combination with addition of hydrogen peroxide. Instrumental analysis is by collision cell inductively coupled plasma - mass spectrometry (modified from EPA Method 6020A). Method Limitation: This method employs a strong acid/peroxide digestion, and is intended to provide a conservative estimate of bio-available metals. Near complete recoveries are achieved for most toxicologically important metals, but elements associated with recalcitrant minerals may be only partially recovered. | 0.02 | Y | Y | Y | 122 | 8/13/2014 | 9/8/2022 | 10/26/2022 7:14:14 AM | 2015-03-09 | 3/11/2015 4:12:05 PM | RDGIBSON | EMS | MassPerMass | mg/kg | Y | (mass conc.) |  | Arsenic Total (mass conc.) | 7440-38-2 | Metals - Total; | Metals - Total; EMS code: AS-T | Chemical | Total | LAB
AS-T | mg/L | Arsenic Total | X063 | Arsine gen,535nm,Col'met | Inorganic arsenic is reduced to arsine, AsH3, by zinc in an acid medium. The arsine then passes from the generator flask through a lead acetate scrubber and into an absorber tube containing silver diethyldithiocarbamate dissolved in pyridine. In the absorber, the arsine reacts with the silver salt to form a soluble red colored complex which is measured colorimetrically at 535nm or 540 nm. | 0.001 | Y | Y | Y | 1516 |  |  | 1/25/2007 4:06:11 PM | 1996-04-12 | 6/3/1998 6:51:34 PM | EMS |  | Concentration | mg/L | Y | (fl. conc.) |  | Arsenic Total (fl. conc.) | 7440-38-2 | Metals - Total; | Metals - Total; EMS code: AS-T | Chemical | Total | LAB
AS-T | ppm (S) | Arsenic Total | MS41 | ALS Minerals ME-MS41 | ALS Minerals ME-MS41. A prepared sample is digested with aqua regia in a graphite heating block. After cooling; the resulting solution is diluted to with deionized water; mixed and analyzed by ICP-AES. Samples are then analyzed by ICP-MS for the remaining suite of elements. | 0.1 | Y | Y | Y | 0 |  |  |  | 2024-01-29 |  | EMS |  | DimensionlessRatio | ppm | N | (ratio) |  | Arsenic Total (ratio) | 7440-38-2 | Metals - Total; | Metals - Total; EMS code: AS-T | Chemical | Total | LAB
AS-T | mg/L | Arsenic Total | X555 | Metals in water, total, "in-bottle" digestion, ICP-SFMS | The whole water sample collected in a HDPE/Teflon bottle, is de-capped in a Class 100 workstation, acidified to 2% (v/v) HNO3 and 1% (v/v) HCl and recapped. The sample is "in-bottle" digested in a convection oven, set at 60 degrees C for 16 hours. The sample is cooled and introduced to an inductively coupled argon plasma-sector field mass spectrometer (ICP-SFMS) by pneumatic nebulization. Each respective element is measured at a specific mass to charge ratio, m/z value expressed in amu (atomic mass units), and at resolution. LR: Low Resolution ~300, MR: Mid Resolution ~4000 and HR: High a specified instrument Resolution ~10000. Ref: Method number 02-2003, National Laboratory for Environmental Testing, CCIW, Environment Canada, Burlington. Requested April 2003. | 1e-05 | Y | Y | Y | 17337 | 5/20/2003 | 12/6/2023 2:08:00 AM | 2/14/2024 9:05:32 AM | 2004-01-14 |  | UMULLER |  | Concentration | mg/L | Y | (fl. conc.) |  | Arsenic Total (fl. conc.) | 7440-38-2 | Metals - Total; | Metals - Total; EMS code: AS-T | Chemical | Total | LAB
AS-T | ug/g | Arsenic Total | ICPM | Metals in Soil by CRC ICPMS | This analysis is carried out using procedures from CSR Analytical Method: "Strong Acid Leachable Metals (SALM) in Soil", BC Ministry of Environment, 26 June 2009, and procedures adapted from EPA Method 200.2. The sample is manually homogenized, dried at 60 degrees Celsius, sieved through a 63 um (230 mesh) sieve, and a representative subsample of the dry material is weighed. The sample is then digested at 95 degrees Celsius for 2 hours by block digester using concentrated nitric and hydrochloric acids. Instrumental analysis is by collision cell inductively coupled plasma - mass spectrometry (modifed from EPA Method 6020A). Method Limitation: This method is not a total digestion technique. It is a very strong acid digestion that is intended to dissolve those metals that may be environmentally available. By design, elements bound in silicate structures are not normally dissolved by this procedure as they are not usually mobile in the environment. | 0.05 | Y | Y | Y | 103 | 9/19/2017 | 9/24/2023 | 1/12/2024 11:16:52 AM | 2018-01-09 |  | JKROGH |  | MassPerMass | mg/kg | Y | (mass conc.) |  | Arsenic Total (mass conc.) | 7440-38-2 | Metals - Total; | Metals - Total; EMS code: AS-T | Chemical | Total | LAB
AS-T | ug/g | Arsenic Total | FB02 | Freeze Dry,Hom,Acid Dig,GF | Sample moisture is determined by weighing samples before and after freeze-drying. Freeze-dried samples are ground in a Braun blender before being weighed into Teflon PFA liners and digested with concentrated HNO3, concentrated HCl, and 50% H2O2 in a microwave oven; which results in a complete sample dissolution. Samples are cooled and volumized with Type I water. Analysis is performed via Inductively Coupled Argon Plasma Atomic Emission Spectrometry (ICP-AES) (simultaneous multielement analysis), Graphite Furnace Atomic Absorption Spectrometry (GF-AAS) (low level Cd and Pb analysis), Cold Vapor Atomic Fluorescence Spectrometry (CV-AFS) (Mercury analysis) and/or Flame Atomic Absorption Spectrometry (F-AAS) (K analysis). | 0.5 | Y | Y | Y | 13 | 7/19/2000 2:23:00 PM | 7/4/2001 12:31:00 PM | 11/21/2001 1:09:40 AM | 1999-11-16 |  | NPEPPIN |  | MassPerMass | mg/kg | Y | (mass conc.) |  | Arsenic Total (mass conc.) | 7440-38-2 | Metals - Total; | Metals - Total; EMS code: AS-T | Chemical | Total | LAB
AS-T | mg/L | Arsenic Total | X145 | Dig, ICP | Dig; ICP | 0.05 | Y | Y | Y | 43 | 3/7/2016 | 9/8/2016 | 1/26/2017 12:01:18 PM | 2016-03-15 |  | RDGIBSON |  | Concentration | mg/L | Y | (fl. conc.) |  | Arsenic Total (fl. conc.) | 7440-38-2 | Metals - Total; | Metals - Total; EMS code: AS-T | Chemical | Total | LAB
AS-T | mg/L | Arsenic Total | X170 | ICP | This code is a catch-all for a large variety of ICP methods including water and also solids digested into solutions as an intermediate result. Different procedures will have greatly different detection limits. (a) For air particulates, preliminary results are reported in mg/L and loading results as mg/dm2/d. (b) For Hi Vol samples preliminiary results are same as reported in mg/L and loading results as ug/m3. (c) Note that for Rain samples there is no filtration nor digestion steps. | 0.04 | Y | Y | Y | 8708 | 3/28/1996 6:03:00 PM | 11/7/2018 | 7/3/2024 12:59:24 PM | 1996-04-12 | 6/3/1998 6:51:34 PM | EMS |  | Concentration | mg/L | Y | (fl. conc.) |  | Arsenic Total (fl. conc.) | 7440-38-2 | Metals - Total; | Metals - Total; EMS code: AS-T | Chemical | Total | LAB
AS-T | None | Arsenic Total | AISP | Special Analysis | Special Analysis | 0 | N | Y | N | 8 |  |  | 10/24/1996 11:23:43 AM | 1996-04-12 | 6/3/1998 6:51:34 PM | EMS |  | None | None | N | (unitless) |  | Arsenic Total (unitless) | 7440-38-2 | Metals - Total; | Metals - Total; EMS code: AS-T | Chemical | Total | LAB
AS-T | ug/g | Arsenic Total | FS01 | D&G,Nitric/HCl:ICP | Samples are dried at 60 C, then either sieved to -100mesh (150 um) and rolled to homogenize, or finely ground with a mortar and pestle or a shatter box laboratory mill and rolled (depending on the submitter's requirements). The samples are then weighed into Teflon digestion liners and digested with HNO3 and HCl in a microwave oven. Samples are cooled, volumized, and allowed to settle overnight before being decanted. Analysis is then performed via Inductively Coupled Argon Plasma Atomic Emission Spectrometry (ICP-AES) (Simultaneous multi element analysis). In ICP Spectrometry, the liquid sample is pumped to the nebulizer, where it is convertd into an aerosol. the nebulizer carrier gas flow than transports the wet aerosol past the impact bead, where the large droplets condense out and flow to the waste container. The remaining fine aerosol is carried into the plasma. There, the high temperature desolvates the aerosol and dissociates the sample material into individual atoms and ions, which are excited to emit light at wavelengths characteristic of the elements present. A diffraction grating in the spectrometer diffracts the various wavelengths of light through exit slits to the photo-multiplier tubes, which convert the light energy to electronic signals. | 8 | Y | Y | Y | 2641 | 8/30/1996 1:52:00 PM | 10/10/2007 | 5/24/2008 1:15:02 AM | 1996-11-20 | 6/3/1998 6:51:46 PM | MPAK |  | MassPerMass | mg/kg | Y | (mass conc.) |  | Arsenic Total (mass conc.) | 7440-38-2 | Metals - Total; | Metals - Total; EMS code: AS-T | Chemical | Total | LAB
AS-T | ug/m3 | Arsenic Total | 6040 | Aqua Regia Dig:ICP | Aqua regia digestion is used to bring the metals into solution prior to ICP analysis. For Hi-Vol samples, preliminary results are reported as mg/L and24 hr loadings as ug/m3 | 0.1 | Y | Y | Y | 437 | 2/13/2002 | 2/9/2005 | 2/23/2005 1:13:46 AM | 1998-05-05 | 6/3/1998 6:51:48 PM | NPEPPIN |  | AirConcentration | µg/m³ | Y | (air conc.) |  | Arsenic Total (air conc.) | 7440-38-2 | Metals - Total; | Metals - Total; EMS code: AS-T | Chemical | Total | LAB
AS-T | mg/L | Arsenic Total | ULTR | Total Metals in Water by ICPMS (Ultra) | Total Metals in Water by ICPMS (Ultra), EPA SW-846 3005A/6020A, MET-TOT-ULTRA-MS-VA This analysis is carried out using procedures adapted from "Standard Methods for the Examination of Water and Wastewater" published by the American Public Health Association, and with procedures adapted from "Test Methods for Evaluating Solid Waste" SW-846 published by the United States Environmental Protection Agency (EPA). The procedures may involve preliminary sample treatment by acid digestion, using either hotblock or microwave oven, or filtration (EPA Method 3005A). Instrumental analysis is by inductively coupled plasma - mass spectrometry (EPA Method 6020A). | 3e-05 | Y | Y | Y | 172 | 2/1/2017 | 10/26/2019 | 11/13/2019 10:17:12 AM | 2012-07-05 |  | RDGIBSON |  | Concentration | mg/L | Y | (fl. conc.) |  | Arsenic Total (fl. conc.) | 7440-38-2 | Metals - Total; | Metals - Total; EMS code: AS-T | Chemical | Total | LAB
AS-T | ug/g wet | Arsenic Total | X560 | Metals in Biota (wet), total, acid digestion, ICP-MS | Five grams of wet, homogenized biota is weighed in a centrifuge tube, and 2.5 mL of Nitric Acid and 2.5 mL of Hydrochloric acid are added. The sample is digested at room temperature for one hour and then digested at 90-95 degrees C in a water bath or heat block for one more hour. The sample is cooled and the volume is made to 50 mL with DI. The digestate is then diluted at ten fold and introduced to a collision/reaction cell inductively coupled argon plasma mass spectrometer (CRC ICPMS) by pneumatic nebulization. Each respective element is measured at a specific mass to charge ratio, m/z value expressed in amu (atomic mass units). - Maxxam Analytics (BC). Added to EMS Dec 10, 2008. | 0.01 | Y | Y | Y | 541 | 11/15/2008 4:00:00 AM | 10/30/2023 | 11/28/2023 2:21:33 PM | 2008-12-11 | 3/11/2015 4:17:10 PM | RDGIBSON | EMS | MassPerMass | µg/g | Y | (mass conc.) | (wet) | Arsenic Total (wet) (mass conc.) | 7440-38-2 | Metals - Total; | Metals - Total; EMS code: AS-T | Chemical | Total | LAB
AS-T | mg/L | Arsenic Total | F082 | Aliquot,HNO3/HCL dig,ICPMS | Aliquot microwave digested with 3:1 HNO3/HCL. Analysis by ICPMS. For HiVol samples preliminary result are reported in mg/L and loading results are reported in ug/m3 | 0.0005 | Y | Y | Y | 27442 | 7/26/2005 7:04:00 AM | 8/23/2024 | 8/26/2024 4:53:05 PM | 2000-06-13 | 8/15/2019 11:04:52 AM | NPEPPIN | KATFRASE | Concentration | mg/L | Y | (fl. conc.) |  | Arsenic Total (fl. conc.) | 7440-38-2 | Metals - Total; | Metals - Total; EMS code: AS-T | Chemical | Total | LAB
AS-T | mg/dm2/d | Arsenic Total | EQ15 | EQUIS (mg/dm2/d) | EQUIS (mg/dm2/d) | 0.0002 | N | Y | Y | 80 |  |  | 10/22/1996 10:18:48 PM | 1996-04-12 | 6/3/1998 6:51:34 PM | EMS |  | MassFlux | mg/dm²/d | Y | (mass flux) |  | Arsenic Total (mass flux) | 7440-38-2 | Metals - Total; | Metals - Total; EMS code: AS-T | Chemical | Total | LAB
AS-T | mg/dm2/d | Arsenic Total | DUST | HNO3 dig.(microwave or hot block),ICPMS | HNO3 dig.(microwave or hot block);ICPMS |  | Y | Y | Y | 705 | 7/19/2016 1:11:00 PM | 7/31/2024 | 7/31/2024 11:23:12 AM | 2008-02-07 |  | RDGIBSON |  | MassFlux | mg/dm²/d | Y | (mass flux) |  | Arsenic Total (mass flux) | 7440-38-2 | Metals - Total; | Metals - Total; EMS code: AS-T | Chemical | Total | LAB
AS-T | mg/L | Arsenic Total | 0180 | HNO3 Dig:Hydride ICP | HNO3 Dig:Hydride ICP | 0.001 | Y | Y | Y | 632 |  |  | 10/24/1996 12:12:38 PM | 1996-04-12 | 6/3/1998 6:51:34 PM | EMS |  | Concentration | mg/L | Y | (fl. conc.) |  | Arsenic Total (fl. conc.) | 7440-38-2 | Metals - Total; | Metals - Total; EMS code: AS-T | Chemical | Total | LAB
AS-T | mg/dm2/d | Arsenic Total | CS15 | System Calculation | System Calculation (mg/dm2/d) |  | N | Y | N | 44 |  |  | 11/18/2000 1:30:24 AM | 1996-04-12 | 6/3/1998 6:51:43 PM | EMS |  | MassFlux | mg/dm²/d | Y | (mass flux) |  | Arsenic Total (mass flux) | 7440-38-2 | Metals - Total; | Metals - Total; EMS code: AS-T | Chemical | Total | LAB
AS-T | ug/m3 | Arsenic Total | 5097 | N/P Dig(Filt+Wash),ICP | N/P Dig(Filt+Wash);ICP | 0.001 | Y | Y | Y | 10781 |  |  | 7/3/2024 12:59:27 PM | 1996-04-12 | 6/3/1998 6:51:34 PM | EMS |  | AirConcentration | µg/m³ | Y | (air conc.) |  | Arsenic Total (air conc.) | 7440-38-2 | Metals - Total; | Metals - Total; EMS code: AS-T | Chemical | Total | LAB
AS-T | ug/g | Arsenic Total | PB02 | PBM,HNO3/HCl dig,dry60C,100mesh, | Performance Based Method. Samples are dried at 60 C and sieved using a .149 mm (100 mesh) sieve. Solid samples are digested with the combination of nitric and hydrochloric acids with a 1:1 ratio of acids. For environmental purposes, strong acid leaches are used routinely since they have the advantage of speed, low cost, and superio reproducibility. All metals trapped within the refractory matrix are unlikely to have environmental consequences. Instrumental analysis of extracts produced using this digestion method can be performed by a variety of techniques. |  | Y | Y | Y | 52 | 5/24/2001 3:42:00 PM | 11/26/2001 10:23:00 AM | 5/26/2003 2:27:14 PM | 2001-07-18 |  | NPEPPIN |  | MassPerMass | mg/kg | Y | (mass conc.) |  | Arsenic Total (mass conc.) | 7440-38-2 | Metals - Total; | Metals - Total; EMS code: AS-T | Chemical | Total | LAB
AS-T | mg/L | Arsenic Total | X270 | HNO3:Hydride:AA | HNO3:Hydride:AA | 0.01 | Y | Y | Y | 4 |  |  | 10/24/1996 1:51:33 PM | 1996-04-12 | 6/3/1998 6:51:41 PM | EMS |  | Concentration | mg/L | Y | (fl. conc.) |  | Arsenic Total (fl. conc.) | 7440-38-2 | Metals - Total; | Metals - Total; EMS code: AS-T | Chemical | Total | LAB
AS-T | mg/L | Arsenic Total | X071 | FA,HNO3,Hydr.Gen,AA | FA;HNO3;Hydr.Gen;AA | 0.0002 | Y | Y | Y | 7780 |  |  | 8/26/2024 2:41:32 PM | 1996-04-12 | 6/3/1998 6:51:41 PM | EMS |  | Concentration | mg/L | Y | (fl. conc.) |  | Arsenic Total (fl. conc.) | 7440-38-2 | Metals - Total; | Metals - Total; EMS code: AS-T | Chemical | Total | LAB
AS-T | ug/g wet | Arsenic Total | MTIW | MET-WET-CCMS-MID-VA | Metals in Tissue by CRC ICPMS (WET) Methodology: This method is conducted following British Columbia Lab Manual method "Metals in Animal Tissue and Vegetation (Biota) - Prescriptive". Tissue samples are homogenized and sub-sampled prior to hotblock digestion with nitric and hydrochloric acids, in combination with addition of hydrogen peroxide. Instrumental analysis is by collision cell inductively coupled plasma - mass spectrometry (modified from EPA Method 6020A). Method Limitation: This method employs a strong acid/peroxide digestion, and is intended to provide a conservative estimate of bio-available metals. Near complete recoveries are achieved for most toxicologically important metals, but elements associated with recalcitrant minerals may be only partially recovered. | 0.006 | Y | Y | Y | 24 | 11/25/2018 | 12/6/2018 | 7/21/2021 2:55:07 PM | 2018-12-24 |  | EMS |  | MassPerMass | µg/g | Y | (mass conc.) | (wet) | Arsenic Total (wet) (mass conc.) | 7440-38-2 | Metals - Total; | Metals - Total; EMS code: AS-T | Chemical | Total | LAB
AS-T | mg/dm2/d | Arsenic Total | CAL1 | Calculation (mg/dm2/d) | Calculation (mg/dm2/d) | 6e-05 | Y | Y | Y | 3038 |  |  | 7/3/2024 12:59:27 PM | 1996-04-12 | 6/3/1998 6:51:42 PM | EMS |  | MassFlux | mg/dm²/d | Y | (mass flux) |  | Arsenic Total (mass flux) | 7440-38-2 | Metals - Total; | Metals - Total; EMS code: AS-T | Chemical | Total | LAB
AS-T | ug/g | Arsenic Total | X230 | Hom,HNO3/HCl,ICP | The sample is dried, ground to pass 100 mesh, and digested with a nitric-hydrochloric acid mixture to solubilize the solid matter and to remove the organic material by oxidation and volatilization. The completed digests are submitted to metals instrumentation personnel for metals analysis by Inductively Coupled Atomic Emission Spectrophotometry (ICP-AES), or hydride analysis by atomic absorption. | 25 | Y | Y | Y | 24 | 12/30/2005 | 12/30/2005 | 2/3/2006 1:13:12 AM | 1996-04-12 | 6/3/1998 6:51:34 PM | EMS |  | MassPerMass | mg/kg | Y | (mass conc.) |  | Arsenic Total (mass conc.) | 7440-38-2 | Metals - Total; | Metals - Total; EMS code: AS-T | Chemical | Total | LAB
AS-T | mg/L | Arsenic Total | MWID | Metals in Water by ICPMS (digested) | Metals in Water by ICPMS (digested) Based on EPA 200.8 | 0.0004 | Y | Y | Y | 36976 |  |  | 7/31/2024 1:46:41 PM | 2009-12-24 |  | RDGIBSON |  | Concentration | mg/L | Y | (fl. conc.) |  | Arsenic Total (fl. conc.) | 7440-38-2 | Metals - Total; | Metals - Total; EMS code: AS-T | Chemical | Total | LAB
ASHC | % | Ash content at 550?C | E205B | CSSS (2008) 28.3 (mod) | CSSS (2008) 28.3 (mod) | 1 | Y | Y | Y | 0 |  |  |  | 2023-12-12 |  | EMS |  | DimensionlessRatio | % | N | (ratio) |  | Ash content at 550?C (ratio) |  | Physical Properties; | Physical Properties; EMS code: ASHC | Physical |  | LAB
ASHT | mg/dm2/d | Particul. Total Ash | CAL1 | Calculation (mg/dm2/d) | Calculation (mg/dm2/d) | 0.1 | Y | Y | Y | 14 |  |  | 3/18/1997 3:01:02 PM | 1996-04-12 | 6/3/1998 6:51:41 PM | EMS |  | MassFlux | mg/dm²/d | Y | (mass flux) |  | Particul. Total Ash (mass flux) |  | Particulates; | Particulates; EMS code: ASHT | Chemical |  | LAB
ASUL | mg/L | Asulam | P015 | Phenoxyacid Herbicides by LC-MS/MS | Phenoxyacid Herbicides by LC-MS/MS | 1e-05 | Y | Y | Y | 68 | 12/6/2021 | 9/9/2022 | 12/8/2022 12:57:05 PM | 2021-12-09 | 12/9/2021 3:22:12 PM | EMS | EMS | Concentration | mg/L | Y | (fl. conc.) |  | Asulam (fl. conc.) | 3337-71-1 | Pesticides; | Pesticides; EMS code: ASUL | Chemical |  | LAB
ASUM | meq/L | Anion Sum | XCAL | Calculation (mg/L) | Calculation (mg/L) |  | Y | Y | Y | 4412 | 6/3/2013 | 8/15/2024 | 8/26/2024 9:09:02 AM | 2016-04-05 |  | RDGIBSON |  | Equivalents | mEq/L | Y | (equiv./vol.) |  | Anion Sum (equiv./vol.) |  | Major Ions; | Major Ions; EMS code: ASUM | Chemical |  | LAB
AT2H | mg/L | Atrazine; 2-hydroxy- | E751 | E751D Pesticides in Water by LC-MS/MS (POS1 List) | E751D Pesticides in Water by LC-MS/MS (POS1 List) (MOECC E3501) | 5e-05 | Y | Y | Y | 39 | 12/9/2021 | 12/31/2021 | 2/15/2022 9:50:04 AM | 2021-12-14 |  | EMS |  | Concentration | mg/L | Y | (fl. conc.) |  | Atrazine; 2-hydroxy- (fl. conc.) | 2163-68-0 | Pesticides; | Pesticides; EMS code: AT2H | Chemical |  | LAB
ATDP | mg/L | Atrazine-desisopropyl | E751 | E751D Pesticides in Water by LC-MS/MS (POS1 List) | E751D Pesticides in Water by LC-MS/MS (POS1 List) (MOECC E3501) | 5e-05 | Y | Y | Y | 39 | 12/9/2021 | 12/31/2021 | 2/15/2022 9:50:04 AM | 2021-12-14 |  | EMS |  | Concentration | mg/L | Y | (fl. conc.) |  | Atrazine-desisopropyl (fl. conc.) | 1007-28-9 | Pesticides; | Pesticides; EMS code: ATDP | Chemical |  | LAB
ATMT | mg/L | Atrazine + metabolites; Total | E751 | E751D Pesticides in Water by LC-MS/MS (POS1 List) | E751D Pesticides in Water by LC-MS/MS (POS1 List) (MOECC E3501) | 0.00051 | Y | Y | Y | 39 | 12/9/2021 | 12/31/2021 | 2/15/2022 9:50:04 AM | 2021-12-14 |  | EMS |  | Concentration | mg/L | Y | (fl. conc.) |  | Atrazine + metabolites; Total (fl. conc.) | 1912-24-9, 6190-65-4,1007-28-9 | Pesticides; | Pesticides; EMS code: ATMT | Chemical | Total | LAB
ATMT | ug/g | Atrazine + metabolites; Total | ZN03 | DCM+Ac.Extr:GC/MS | DCM+Ac.Extr:GC/MS EPA 8270E (mod) | 0.005 | Y | Y | Y | 2 | 11/2/2022 12:53:00 AM | 11/2/2022 4:41:00 AM | 1/7/2023 3:12:10 PM | 2023-01-06 |  | EMS |  | MassPerMass | mg/kg | Y | (mass conc.) |  | Atrazine + metabolites; Total (mass conc.) | 1912-24-9, 6190-65-4 , 1007-28-9 | Pesticides; | Pesticides; EMS code: ATMT | Chemical | Total | LAB
ATNOL | mg/L | Atenolol CAS: 29122-68-7 | XMSP | HPLC and Tri. Quadpole MS/MS | EPA 1694 An aliquot of filtered sample containing 5% organic is injected directly. Each compound is separated by reversed phase HPLC. Analyzed by Electro Spray Ionization MS/MS detection using Triple Quadrupole MS/MS detector. Quantiitation with internal or external standard. |  | Y | Y | Y | 12 | 5/8/2020 | 6/4/2024 | 8/21/2024 3:32:03 PM | 2022-05-04 |  | EMS |  | Concentration | mg/L | Y | (fl. conc.) |  | Atenolol (fl. conc.) | 29122-68-7 | Pharmaceuticals and Personal Care Products (PPCP); | Pharmaceuticals and Personal Care Products (PPCP); EMS code: ATNOL | Chemical |  | LAB
ATOR | mg/L | Atorvastatin CAS: 134523-00-5 | XMSP | HPLC and Tri. Quadpole MS/MS | EPA 1694 An aliquot of filtered sample containing 5% organic is injected directly. Each compound is separated by reversed phase HPLC. Analyzed by Electro Spray Ionization MS/MS detection using Triple Quadrupole MS/MS detector. Quantiitation with internal or external standard. |  | Y | Y | Y | 12 | 5/8/2020 | 6/4/2024 | 8/21/2024 3:32:03 PM | 2022-05-04 |  | EMS |  | Concentration | mg/L | Y | (fl. conc.) |  | Atorvastatin (fl. conc.) | 134523-00-5 | Pharmaceuticals and Personal Care Products (PPCP); | Pharmaceuticals and Personal Care Products (PPCP); EMS code: ATOR | Chemical |  | LAB
ATZD | ug/g | Atrazine-Desethyl CAS: 6190-65-4 | ZN03 | DCM+Ac.Extr:GC/MS | DCM+Ac.Extr:GC/MS EPA 8270E (mod) | 0.05 | Y | Y | Y | 12 | 1/25/2022 | 2/1/2022 | 7/6/2022 5:20:43 PM | 2022-06-13 |  | EMS |  | MassPerMass | mg/kg | Y | (mass conc.) |  | Atrazine-Desethyl (mass conc.) | 6190-65-4 | Organics; | Organics; EMS code: ATZD | Chemical |  | LAB
AU-D | mg/L | Gold Dissolved | F032 | Filt.,Acid,ICPMS | There is no acid digestion required for this analysis. A water sample is collected and preserved in 0.2 percent nitric acid. To analyse the extractable metals the sample is shaken, allowed to sit for at least twenty four hours, and a decanted portion is introduced to the spectrometer. When the dissolved metals are required, the samples must first be filtered through a 0.45 micron cellulose acetate filter paper and then followed by preservation in 0.2 percent nitric acid. It is important to note that different laboratories may have very different detection limits for some variables analyzed by ICPMS. | 0.01 | Y | Y | Y | 255 | 5/9/2003 9:40:00 AM | 11/26/2015 | 2/18/2016 11:11:17 AM | 1998-06-22 |  | NPEPPIN |  | Concentration | mg/L | Y | (fl. conc.) |  | Gold Dissolved (fl. conc.) | 7440-57-5 | Metals - Dissolved; | Metals - Dissolved; EMS code: AU-D | Chemical | Dissolved | LAB
AU-T | ppm (S) | Gold Total | 6041 | Aqua Regia by ICPMS | These were sediment/soils digested with Aqua-Regia (Nitric Acid & Hydrochloric Acids) and then analysis by ICPMS. Results would be reported in mg/kg and/or %. | 0.0002 | Y | Y | Y | 21 | 10/13/2019 | 8/5/2020 | 8/10/2020 3:52:06 PM | 2019-09-06 | 10/21/2019 11:54:07 AM | KATFRASE | KATFRASE | DimensionlessRatio | ppm | N | (ratio) |  | Gold Total (ratio) | 7440-57-5 | Metals - Total; | Metals - Total; EMS code: AU-T | Chemical | Total | LAB
AU-T | ppm (S) | Gold Total | MS41 | ALS Minerals ME-MS41 | ALS Minerals ME-MS41. A prepared sample is digested with aqua regia in a graphite heating block. After cooling; the resulting solution is diluted to with deionized water; mixed and analyzed by ICP-AES. Samples are then analyzed by ICP-MS for the remaining suite of elements. | 0.02 | Y | Y | Y | 0 |  |  |  | 2024-01-29 |  | EMS |  | DimensionlessRatio | ppm | N | (ratio) |  | Gold Total (ratio) | 7440-57-5 | Metals - Total; | Metals - Total; EMS code: AU-T | Chemical | Total | LAB
AU-T | mg/L | Gold Total | F033 | Acid Dig,Centrifuge,ICPMS | For the determination of total recoverable elements in water or waste water, take 100 ml aliquot from a well mixed, acid preserved, sample containing not more than 0.25% (w/v) total solids and transfer to a 250 ml Griffin beaker. (If total solids are greater than 0.25% reduce the size of the aliquot by a proportionate amount). Add 1 ml of conc. nitric acid and 0.5 ml conc. hydrochloric acid. Heat on a hot plate until the volume has been reduced to 15-20ml, ensuring that the sample does not boil. Allow to cool and transfer to a 50 ml volumetric flask. Dilute to volume with ASTM type I water. Centrifuge the sample or allow to stand overnight to separate insoluble material. Prior analysis, pipette 20 ml into a 50 ml volumetric flask and dilute to volume with ASTM type I water. Analysis should be performed within seventy two hours of the completed preparation. Note that different laboratories may have very different field preservation, digestion procedures and detection limits for this method, so these should be documented. | 0.01 | Y | Y | Y | 68 | 5/8/2003 9:40:00 AM | 9/12/2003 3:16:00 PM | 12/1/2003 5:30:16 PM | 1998-06-22 |  | NPEPPIN |  | Concentration | mg/L | Y | (fl. conc.) |  | Gold Total (fl. conc.) | 7440-57-5 | Metals - Total; | Metals - Total; EMS code: AU-T | Chemical | Total | LAB
AU-T | mg/L | Gold Total | F121 | Acid,ICPMS | Acidified, no digestion, ICPMS |  | Y | Y | Y | 188 | 4/6/2015 | 12/4/2015 | 2/18/2016 11:11:17 AM | 2002-07-17 |  | UMULLER |  | Concentration | mg/L | Y | (fl. conc.) |  | Gold Total (fl. conc.) | 7440-57-5 | Metals - Total; | Metals - Total; EMS code: AU-T | Chemical | Total | LAB
AU-T | ppm (S) | Gold Total | 6042 | Aqua Regia by ICPMS | These were sediment/soils digested with Aqua-Regia (Nitric Acid & Hydrochloric Acids) and then analysis by ICPMS. Results would be reported in ppm s. | 0.0002 | Y | Y | Y | 71 | 5/14/2022 | 7/27/2024 | 7/29/2024 9:13:08 AM | 2022-06-28 |  | EMS |  | DimensionlessRatio | ppm | N | (ratio) |  | Gold Total (ratio) | 7440-57-5 | Metals - Total; | Metals - Total; EMS code: AU-T | Chemical | Total | LAB
AZOX | mg/L | Azoxystrobin | E751 | E751D Pesticides in Water by LC-MS/MS (POS1 List) | E751D Pesticides in Water by LC-MS/MS (POS1 List) (MOECC E3501) | 1e-05 | Y | Y | Y | 39 | 12/9/2021 | 12/31/2021 | 2/15/2022 9:50:04 AM | 2021-12-14 |  | EMS |  | Concentration | mg/L | Y | (fl. conc.) |  | Azoxystrobin (fl. conc.) | 131860-33-8 | Pesticides; | Pesticides; EMS code: AZOX | Chemical |  | LAB
B--D | mg/L | Boron Dissolved | X170 | ICP | This code is a catch-all for a large variety of ICP methods including water and also solids digested into solutions as an intermediate result. Different procedures will have greatly different detection limits. (a) For air particulates, preliminary results are reported in mg/L and loading results as mg/dm2/d. (b) For Hi Vol samples preliminiary results are same as reported in mg/L and loading results as ug/m3. (c) Note that for Rain samples there is no filtration nor digestion steps. | 0.1 | Y | Y | Y | 13421 | 4/11/1996 6:01:00 PM | 10/29/2019 | 6/12/2020 1:21:11 PM | 1996-04-12 | 8/15/2019 11:51:15 AM | EMS | KATFRASE | Concentration | mg/L | Y | (fl. conc.) |  | Boron Dissolved (fl. conc.) | 7440-42-8 | Metals - Dissolved; | Metals - Dissolved; EMS code: B--D | Chemical | Dissolved | LAB
B--D | mg/L | Boron Dissolved | X085 | LF,LA,HNO3,AA | Samples are lab-filtered (0.45 um), and lab-acidified (HNO3); analysis by AA. | 0.01 | Y | Y | Y | 83 | 9/27/2007 | 9/18/2015 | 10/28/2015 9:54:16 AM | 1996-04-12 | 6/3/1998 6:51:34 PM | EMS |  | Concentration | mg/L | Y | (fl. conc.) |  | Boron Dissolved (fl. conc.) | 7440-42-8 | Metals - Dissolved; | Metals - Dissolved; EMS code: B--D | Chemical | Dissolved | LAB
B--D | mg/L | Boron Dissolved | MWID | Metals in Water by ICPMS (digested) | Metals in Water by ICPMS (digested) Based on EPA 200.8 | 0.1 | Y | Y | Y | 31695 | 10/26/2018 5:20:00 PM | 10/26/2018 5:20:00 PM | 7/31/2024 1:46:41 PM | 2009-12-24 |  | RDGIBSON |  | Concentration | mg/L | Y | (fl. conc.) |  | Boron Dissolved (fl. conc.) | 7440-42-8 | Metals - Dissolved; | Metals - Dissolved; EMS code: B--D | Chemical | Dissolved | LAB
B--D | mg/L | Boron Dissolved | X556 | Metals in water, dissolved, "in-bottle" digestion, ICP-SFMS | The whole water sample is filtered through a 0.45 micron cellulose acetate filter paper and the filtrate is stored in a HDPE/Teflon bottle. The container is de-capped in a Class 100 workstation, acidified to 2% (v/v) HNO3 and 1% (v/v) HCl and recapped. The sample is "in-bottle" digested in a convection oven, set at 60 degrees C for 16 hours. The sample is cooled and introduced to an inductively coupled argon plasma-sector field mass spectrometer (ICP-SFMS) by pneumatic nebulization. Each respective element is measured at a specific mass to charge ratio, m/z value expressed in amu (atomic mass units), and at a specified intrument resolution. LR: Low Resolution ~300, MR: Mid Resolution ~4000 and HR: High Resolution ~10000. Ref: Method number 02-2003, National Laboratory for Environmental Testing, CCIW, Environment Canada, Burlington, Requested April 2003. | 1e-05 | Y | Y | Y | 4234 | 6/11/2003 | 12/6/2023 2:04:00 AM | 2/14/2024 9:05:32 AM | 2004-01-19 |  | UMULLER |  | Concentration | mg/L | Y | (fl. conc.) |  | Boron Dissolved (fl. conc.) | 7440-42-8 | Metals - Dissolved; | Metals - Dissolved; EMS code: B--D | Chemical | Dissolved | LAB
B--D | mg/L | Boron Dissolved | X285 | HNO3 Dig:ICP Analysis | HNO3 Dig:ICP Analysis | 0.1 | Y | Y | Y | 651 | 1/30/2002 | 11/13/2002 | 2/19/2004 4:22:42 PM | 1996-04-12 | 6/3/1998 6:51:34 PM | EMS |  | Concentration | mg/L | Y | (fl. conc.) |  | Boron Dissolved (fl. conc.) | 7440-42-8 | Metals - Dissolved; | Metals - Dissolved; EMS code: B--D | Chemical | Dissolved | LAB
B--D | mg/L | Boron Dissolved | F032 | Filt.,Acid,ICPMS | There is no acid digestion required for this analysis. A water sample is collected and preserved in 0.2 percent nitric acid. To analyse the extractable metals the sample is shaken, allowed to sit for at least twenty four hours, and a decanted portion is introduced to the spectrometer. When the dissolved metals are required, the samples must first be filtered through a 0.45 micron cellulose acetate filter paper and then followed by preservation in 0.2 percent nitric acid. It is important to note that different laboratories may have very different detection limits for some variables analyzed by ICPMS. | 7e-05 | Y | Y | Y | 76844 | 12/19/1996 8:15:00 AM | 8/26/2024 | 8/26/2024 4:53:05 PM | 1997-01-16 | 4/29/2011 8:48:30 AM | MPAK | RDGIBSON | Concentration | mg/L | Y | (fl. conc.) |  | Boron Dissolved (fl. conc.) | 7440-42-8 | Metals - Dissolved; | Metals - Dissolved; EMS code: B--D | Chemical | Dissolved | LAB
B--D | mg/L | Boron Dissolved | X562 | Dissolved metals: FF, FA, HNO3, HR-ICPMS | Ultra trace metals in water are analyzed by high resolution inductively coupled plasma mass spectrometry (HR-ICPMS), based on US EPA Method 200.8, (Oct. 1999). This procedure is intended for pristine field-filtered acid-preserved water samples. The detection limits (LOR) for this test are based on lab instrumental analysis only, not including filtration. ALS-supplied field filtration equipment does not support these LOR. Therefore, because of the high probability of false positives due to filtration, it is strongly recommended that a filtration blank be analyzed to aid in data interpretation. | 0.005 | Y | Y | Y | 1264 | 6/8/2015 | 10/19/2023 | 11/9/2023 7:41:27 AM | 2015-10-22 |  | RDGIBSON |  | Concentration | mg/L | Y | (fl. conc.) |  | Boron Dissolved (fl. conc.) | 7440-42-8 | Metals - Dissolved; | Metals - Dissolved; EMS code: B--D | Chemical | Dissolved | LAB
B--D | mg/L | Boron Dissolved | EQ01 | EQUIS (mg/L) | EQUIS (mg/L) | 0.01 | N | Y | Y | 2069 |  |  | 10/24/1996 3:16:02 PM | 1996-04-12 | 6/3/1998 6:51:34 PM | EMS |  | Concentration | mg/L | Y | (fl. conc.) |  | Boron Dissolved (fl. conc.) | 7440-42-8 | Metals - Dissolved; | Metals - Dissolved; EMS code: B--D | Chemical | Dissolved | LAB
B--D | mg/L | Boron Dissolved | F068 | HNO3dig,TeflonTube/Autoclave,ICP-MS | This method is applicable to determiniation of dissolved, extractable, and total recoverable metals in ground waters, surface waters, drinking waters, and wastewaters. Dissolved, extractable, and total recoverable sample fractions are acidified to pH <1.5 with analytical grade nitric acid (HNO3) to prevent the metals from adsorbing to the container material or precipitating. Extractable and total recoverable water samples are collected in 250 ml acid washed polyethylene bottles and are preserved with Nitric acid (0.5 ml acid for 100 mLs of sample). All preserved samples are stored at room temperature and can be held for up to 6 months. Unfiltered, un-preserved samples are stored at 4C for no more than 30 days. For dissolved and extractable water fractions, no sample preparation is necessary. Samples are simply poured in autosampler tubes (i.e. 15 ml centrifuge tubes), and placed on the autosampler and analyzed. Samples for total recoverable metals are digested with Nitric acid in special Teflon tubes in an autoclave prior to analyses. | 0.002 | Y | Y | Y | 793 | 6/1/1999 1:19:00 PM | 2/12/2002 2:08:00 PM | 8/11/2021 12:04:11 PM | 1999-06-28 |  | NPEPPIN |  | Concentration | mg/L | Y | (fl. conc.) |  | Boron Dissolved (fl. conc.) | 7440-42-8 | Metals - Dissolved; | Metals - Dissolved; EMS code: B--D | Chemical | Dissolved | LAB
B--D | mg/L | Boron Dissolved | F039 | Field filt + acid, ICP-AES | Dissolved and extractable water samples are ready for analysis once collected. Total samples are autoclaved with HNO3 and hydrochloric acid (HCI) in capped Teflon vials. Analysis is performed via Inductively Coupled Argon Plasma Atomic Emission Spectrometry (ICP-AES)(simultaneous multielement analysis), Graphite Furnace Atomic Absorption Spectrometry (GF-AAS)(low level analysis) and/or Flame Atomic Absorption Spectrometry (F-AAS). | 0.01 | Y | Y | Y | 7281 | 7/29/1996 6:03:00 PM | 7/5/2019 1:21:00 PM | 5/16/2022 4:21:32 PM | 1997-05-08 | 6/3/1998 6:51:46 PM | MPAK |  | Concentration | mg/L | Y | (fl. conc.) |  | Boron Dissolved (fl. conc.) | 7440-42-8 | Metals - Dissolved; | Metals - Dissolved; EMS code: B--D | Chemical | Dissolved | LAB
B--D | mg/L | Boron Dissolved | MSEA | Metals in Seawater (undigested) by (CRC) ICP-MS | This method describes the multi-elemental determination of trace elements by Collision/Reaction Cell (CRC) ICP-MS. The method measures ions produced by a radio-frequency inductively coupled plasma. Analyte species originating in a liquid are pneumatically nebulized and the resulting aerosol transported by argon gas into the plasma torch. The ions produced are entrained in the plasma gas and extracted, by means of a differentially pumped vacuum interface, into a mass spectrometer. The ions flow through a Collision/Reaction Cell (CRC) where depending on the mode of analysis either the ions passes through the cell unhindered (no gas - classical ICP-MS mode), or collisions/reactions chemistry takes place between the ions and a neutral gas such as hydrogen or helium. Any polyatomic ions which give interference on the analytes are removed. The ions are then sorted according to their mass-to-charge ratios by a quadrupole mass spectrometer. The ions are then sorted according to their mass-to-charge ratios by a quadrupole mass spectrometer. The ions transmitted through the quadrupole are quantified by a channel electron multiplier and the ion information processed by a data handling system. Although most polyatomic interference is removed, any additional interference must be assessed and valid corrections applied or the data flagged to indicate problems. Interference correction must include compensation for background ions contributed by the plasma gas, reagents, and constituents of the sample matrix. Instrument drift as well as suppressions or enhancements of instrument response caused by the sample matrix must be corrected for by the use of internal standards. | 0.0004 | Y | Y | Y | 4 | 4/12/2023 | 4/29/2024 | 5/16/2024 2:00:14 PM | 2009-09-08 |  | RDGIBSON |  | Concentration | mg/L | Y | (fl. conc.) |  | Boron Dissolved (fl. conc.) | 7440-42-8 | Metals - Dissolved; | Metals - Dissolved; EMS code: B--D | Chemical | Dissolved | LAB
B--D | mg/L | Boron Dissolved | METC | MET-D-NP-CCMS-VA | Ultra trace metals in water are analyzed by CRC ICPMS, based on US EPA Method 6020B (July 2014). This procedure is intended for pristine field-filtered acid-preserved water samples. The detection limits (LOR) for this test are based on lab instrumental analysis only, not including filtration. ALS-supplied field filtration equipment does not support these LOR. Therefore, because of the high probability of false positives due to filtration, it is strongly recommended that a filtration blank be analysed to aid in data interpretation. | 0.005 | Y | Y | Y | 1776 | 4/3/2012 | 8/15/2024 | 8/23/2024 9:50:35 AM | 2019-08-20 |  | KATFRASE |  | Concentration | mg/L | Y | (fl. conc.) |  | Boron Dissolved (fl. conc.) | 7440-42-8 | Metals - Dissolved; | Metals - Dissolved; EMS code: B--D | Chemical | Dissolved | LAB
B--E | mg/L | Boron Extractable | F040 | Field Acid, lab filt, ICP-AES | Dissolved and extractable water samples are ready for analysis once collected. Total samples are autoclaved with HNO3 and hydrochloric acid (HCI) in capped Teflon vials. Analysis is performed via Inductively Coupled Argon Plasma Atomic Emission Spectrometry (ICP-AES)(simultaneous multielement analysis), Graphite Furnace Atomic Absorption Spectrometry (GF-AAS)(low level analysis) and/or Flame Atomic Absorption Spectrometry (F-AAS). | 0.01 | Y | Y | Y | 2361 | 8/5/1997 6:06:00 PM | 3/14/2008 | 9/11/2019 8:47:41 AM | 1997-05-08 | 6/3/1998 6:51:46 PM | MPAK |  | Concentration | mg/L | Y | (fl. conc.) |  | Boron Extractable (fl. conc.) | 7440-42-8 | Metals - Extractable; | Metals - Extractable; EMS code: B--E | Chemical | Extractable | LAB
B--E | mg/L | Boron Extractable | F069 | SEM extr,6N HCl,ICP/OES | 6.0 N Hydrochloric acid (HCI) is added to an aliquot of the sediment sample. The extract produced from the addition of the acid is then analyzed for simultaneously extracted metals (SEM) using inductively coupled plasma - optical emission spectrophotometry. The detection limits depend on the metal parameters in question and the instrumentation used. | 0.1 | Y | Y | Y | 10 | 3/26/2012 3:01:00 AM | 3/26/2012 3:01:00 AM | 11/25/2021 1:50:15 PM | 1999-07-09 |  | NPEPPIN |  | Concentration | mg/L | Y | (fl. conc.) |  | Boron Extractable (fl. conc.) | 7440-42-8 | Metals - Extractable; | Metals - Extractable; EMS code: B--E | Chemical | Extractable | LAB
B--E | mg/L | Boron Extractable | X553 | Metals in water, extractable, ICP-SFMS | The whole water sample is collected in a HDPE/Teflon bottle. In a Class 100 workstation it is preserved with concentrated Nitric Acid at concentration of 0.2 percent (v/v). The acid preservative extracts the metals from material in the sample and a decanted portion is introduced directly to an inductively coupled argon plasma-sector field mass spectrometer (ICP-SFMS) by pneumatic nebulization. Each respective element is measured at a specific mass to charge ratio, m/z value expressed in amu (atomic mass units), and at a specified instrument resolution. LR: Low Resolution ~300, MR: Mid Resolution ~4000 and HR: High Resolution ~10000. Ref: Method number 02-2061, National Laboratory for Environmental Testing, CCIW, Environment Canada, Burlington. Requested April 2003. | 0.0005 | Y | Y | Y | 2596 | 5/22/2003 | 9/21/2007 | 11/28/2023 2:11:53 PM | 2004-01-19 |  | UMULLER |  | Concentration | mg/L | Y | (fl. conc.) |  | Boron Extractable (fl. conc.) | 7440-42-8 | Metals - Extractable; | Metals - Extractable; EMS code: B--E | Chemical | Extractable | LAB
B--E | mg/L | Boron Extractable | F047 | 24hr HNO3 dil'n,ICP-MS | There is no acid digestion required for this analysis. A water sample is collected and preserved in 0.2 percent nitric acid. To analyse the extractable metals the sample is shaken, allowed to sit for at least twenty four hours, and a decanted portion is introduced to the spectrometer. When the dissolved metals are required, the samples must first be filtered through a 0.45 micron cellulose acetate filter paper and then followed by preservation in 0.2 percent nitric acid. It is important to note that different laboratories may have very different detection limits for some variables analyzed by ICP/MS. | 7e-05 | Y | Y | Y | 2680 | 4/16/1998 10:21:00 AM | 12/13/2011 12:02:00 PM | 9/11/2019 8:46:17 AM | 1997-10-20 | 6/3/1998 6:51:47 PM | NPEPPIN |  | Concentration | mg/L | Y | (fl. conc.) |  | Boron Extractable (fl. conc.) | 7440-42-8 | Metals - Extractable; | Metals - Extractable; EMS code: B--E | Chemical | Extractable | LAB
B--S | mg/L | Boron; soluble ion content | E6010B | EPA 6010B/CSSS 15 | See CSSS CH15/EPA 6010B/AER D50. INDUCTIVELY COUPLED PLASMA-ATOMIC EMISSION SPECTROMETRY See EPA 6010B and CSSS Ch. 15. https://www.epa.gov/sites/default/files/documents/6010b.pdf | 0.25 | Y | Y | Y | 0 |  |  |  | 2024-01-30 |  | EMS |  | Concentration | mg/L | Y | (fl. conc.) |  | Boron; soluble ion content (fl. conc.) |  | Metals - Soluble; | Metals - Soluble; EMS code: B--S | Chemical |  | LAB
B--T | mg/L | Boron Total | X285 | HNO3 Dig:ICP Analysis | HNO3 Dig:ICP Analysis | 0.008 | Y | Y | Y | 1974 | 8/10/1993 | 9/12/2003 3:16:00 PM | 4/30/2004 1:54:38 PM | 1996-04-12 | 6/3/1998 6:51:42 PM | EMS |  | Concentration | mg/L | Y | (fl. conc.) |  | Boron Total (fl. conc.) | 7440-42-8 | Metals - Total; | Metals - Total; EMS code: B--T | Chemical | Total | LAB
B--T | mg/L | Boron Total | 0042 | GI HNO3 Dig: ICP | Sample aliquot taken from unfiltered unpreserved General Ions bottle; HNO3 digestion; ICP analysis. | 0.04 | Y | Y | Y | 9848 | 9/20/1996 | 7/5/2019 3:21:00 PM | 7/12/2019 6:49:13 PM | 1996-04-12 | 6/3/1998 6:51:34 PM | EMS |  | Concentration | mg/L | Y | (fl. conc.) |  | Boron Total (fl. conc.) | 7440-42-8 | Metals - Total; | Metals - Total; EMS code: B--T | Chemical | Total | LAB
B--T | mg/L | Boron Total | METN | MET-T-NP-CCMS-VA | Ultra trace metals in water are analyzed by CRC ICPMS, based on US EPA Method 6020B (July 2014). The detection limits provided can only be met for undigested samples. This procedure is intended for colorless, non-turbid, acid-preserved water samples (i.e. pristine water samples), having turbidity < 1 NTU and no odor. Where turbidity exceeds 1 NTU, and/or the sample is colored and has an odor, results may be biased low compared to true Total Metals concentrations. ALS recommends that turbidity analysis be requested on samples submitted for this test to aid with interpretation of results. Where turbidity is <1NTU, undigested metals are equivalent to total metals concentrations. | 0.005 | Y | Y | Y | 3648 | 4/3/2012 | 8/19/2024 | 8/26/2024 1:03:06 PM | 2019-08-20 |  | KATFRASE |  | Concentration | mg/L | Y | (fl. conc.) |  | Boron Total (fl. conc.) | 7440-42-8 | Metals - Total; | Metals - Total; EMS code: B--T | Chemical | Total | LAB
B--T | ug/g | Boron Total | FS01 | D&G,Nitric/HCl:ICP | Samples are dried at 60 C, then either sieved to -100mesh (150 um) and rolled to homogenize, or finely ground with a mortar and pestle or a shatter box laboratory mill and rolled (depending on the submitter's requirements). The samples are then weighed into Teflon digestion liners and digested with HNO3 and HCl in a microwave oven. Samples are cooled, volumized, and allowed to settle overnight before being decanted. Analysis is then performed via Inductively Coupled Argon Plasma Atomic Emission Spectrometry (ICP-AES) (Simultaneous multi element analysis). In ICP Spectrometry, the liquid sample is pumped to the nebulizer, where it is convertd into an aerosol. the nebulizer carrier gas flow than transports the wet aerosol past the impact bead, where the large droplets condense out and flow to the waste container. The remaining fine aerosol is carried into the plasma. There, the high temperature desolvates the aerosol and dissociates the sample material into individual atoms and ions, which are excited to emit light at wavelengths characteristic of the elements present. A diffraction grating in the spectrometer diffracts the various wavelengths of light through exit slits to the photo-multiplier tubes, which convert the light energy to electronic signals. | 2 | Y | Y | Y | 1393 | 12/5/1997 6:01:00 PM | 2/5/2002 | 5/28/2003 2:34:19 PM | 1997-10-17 | 6/3/1998 6:51:47 PM | NPEPPIN |  | MassPerMass | mg/kg | Y | (mass conc.) |  | Boron Total (mass conc.) | 7440-42-8 | Metals - Total; | Metals - Total; EMS code: B--T | Chemical | Total | LAB
B--T | ug/g | Boron Total | X551 | Metals in Biota (dry), total, acid digestion, ICP-SFMS | 0.1 gram of dried, homogenized biota is weighed in a microwave TFM vessel, and 8mL of Nitric Acid and 1mL of Hydrogen Peroxide are added. The sample is digested at 200 degrees C for 15 minutes in a closed vessel high pressure microwave oven. The sample is cooled, the volume adjusted to 100mL and introduced to an iductively coupled argon plasma-sector field mass spectrometer (ICP-SFMS) by pneumatic nebulization. Each respective element is measured at a specific mass to charge ratio, m/z value expressed in amu (atomic mass units), and at a specified instrument resolution. LR: Low Resolution ~300, MR: Mid Resolution ~4000 and HR: Hich Resolution ~10000. Ref: Method number 02-2705, National Laboratory for Environmental Testing, CCIW, Environment Canada, Burlington. Requested April 2003. | 1 | Y | Y | Y | 49 | 8/11/2015 | 10/30/2023 | 11/28/2023 2:21:33 PM | 2015-03-11 |  | RDGIBSON |  | MassPerMass | mg/kg | Y | (mass conc.) |  | Boron Total (mass conc.) | 7440-42-8 | Metals - Total; | Metals - Total; EMS code: B--T | Chemical | Total | LAB
B--T | mg/L | Boron Total | F068 | HNO3dig,TeflonTube/Autoclave,ICP-MS | This method is applicable to determiniation of dissolved, extractable, and total recoverable metals in ground waters, surface waters, drinking waters, and wastewaters. Dissolved, extractable, and total recoverable sample fractions are acidified to pH <1.5 with analytical grade nitric acid (HNO3) to prevent the metals from adsorbing to the container material or precipitating. Extractable and total recoverable water samples are collected in 250 ml acid washed polyethylene bottles and are preserved with Nitric acid (0.5 ml acid for 100 mLs of sample). All preserved samples are stored at room temperature and can be held for up to 6 months. Unfiltered, un-preserved samples are stored at 4C for no more than 30 days. For dissolved and extractable water fractions, no sample preparation is necessary. Samples are simply poured in autosampler tubes (i.e. 15 ml centrifuge tubes), and placed on the autosampler and analyzed. Samples for total recoverable metals are digested with Nitric acid in special Teflon tubes in an autoclave prior to analyses. | 0.002 | Y | Y | Y | 1680 | 6/1/1999 1:23:00 PM | 4/21/2004 7:34:00 PM | 7/30/2024 6:32:12 AM | 1999-06-28 |  | NPEPPIN |  | Concentration | mg/L | Y | (fl. conc.) |  | Boron Total (fl. conc.) | 7440-42-8 | Metals - Total; | Metals - Total; EMS code: B--T | Chemical | Total | LAB
B--T | ug/g | Boron Total | 0190 | D&G:N/P Dig:ICP Anal. | Sample is dried and grind-mixed before nitric perchloric digestion followed by ICP analysis. | 1 | Y | Y | Y | 13 |  |  | 10/23/1996 10:17:01 PM | 1996-04-12 | 6/3/1998 6:51:34 PM | EMS |  | MassPerMass | mg/kg | Y | (mass conc.) |  | Boron Total (mass conc.) | 7440-42-8 | Metals - Total; | Metals - Total; EMS code: B--T | Chemical | Total | LAB
B--T | mg/dm2/d | Boron Total | DUST | HNO3 dig.(microwave or hot block),ICPMS | HNO3 dig.(microwave or hot block);ICPMS |  | Y | Y | Y | 420 | 7/19/2016 1:11:00 PM | 7/31/2024 | 7/31/2024 11:23:12 AM | 2008-02-07 |  | RDGIBSON |  | MassFlux | mg/dm²/d | Y | (mass flux) |  | Boron Total (mass flux) | 7440-42-8 | Metals - Total; | Metals - Total; EMS code: B--T | Chemical | Total | LAB
B--T | mg/L | Boron Total | X555 | Metals in water, total, "in-bottle" digestion, ICP-SFMS | The whole water sample collected in a HDPE/Teflon bottle, is de-capped in a Class 100 workstation, acidified to 2% (v/v) HNO3 and 1% (v/v) HCl and recapped. The sample is "in-bottle" digested in a convection oven, set at 60 degrees C for 16 hours. The sample is cooled and introduced to an inductively coupled argon plasma-sector field mass spectrometer (ICP-SFMS) by pneumatic nebulization. Each respective element is measured at a specific mass to charge ratio, m/z value expressed in amu (atomic mass units), and at resolution. LR: Low Resolution ~300, MR: Mid Resolution ~4000 and HR: High a specified instrument Resolution ~10000. Ref: Method number 02-2003, National Laboratory for Environmental Testing, CCIW, Environment Canada, Burlington. Requested April 2003. | 0.0005 | Y | Y | Y | 17325 | 5/20/2003 | 12/6/2023 2:08:00 AM | 2/14/2024 9:05:32 AM | 2004-01-14 |  | UMULLER |  | Concentration | mg/L | Y | (fl. conc.) |  | Boron Total (fl. conc.) | 7440-42-8 | Metals - Total; | Metals - Total; EMS code: B--T | Chemical | Total | LAB
B--T | mg/L | Boron Total | X561 | Total metals: FA, HNO3, HR-ICPMS (No Prep)Prep) | Ultra trace metals in water are analyzed by high resolution inductively coupled plasma mass spectrometry (HR-ICPMS), modified from US EPA 200.8, (Revision 5.5). The detection limits provided can only be met for undigested samples. This procedure is intended for colorless, non-turbid, acid-preserved water samples (i.e. pristine water samples), having turbidity < 1 NTU and no odor. Where turbidity exceeds 1 NTU, and/or the sample is colored and has an odor, results may be biased low compared to true Total Metals concentrations. ALS recommends that turbidity analysis be requested on samples submitted for this test to aid with interpretation of results. | 0.005 | Y | Y | Y | 2756 | 6/8/2015 | 10/5/2023 | 1/31/2024 2:37:41 PM | 2015-10-22 |  | RDGIBSON |  | Concentration | mg/L | Y | (fl. conc.) |  | Boron Total (fl. conc.) | 7440-42-8 | Metals - Total; | Metals - Total; EMS code: B--T | Chemical | Total | LAB
B--T | ug/g | Boron Total | FS89 | Homo,HNO3 dig,ICP-AES | Homo;HNO3 dig;ICP-AES | 2 | Y | Y | Y | 8 | 8/20/2001 12:00:00 PM | 10/31/2022 | 11/22/2022 2:40:39 PM | 2001-03-08 |  | NPEPPIN |  | MassPerMass | mg/kg | Y | (mass conc.) |  | Boron Total (mass conc.) | 7440-42-8 | Metals - Total; | Metals - Total; EMS code: B--T | Chemical | Total | LAB
B--T | mg/L | Boron Total | X145 | Dig, ICP | Dig; ICP | 0.1 | Y | Y | Y | 32 |  |  | 5/11/1998 5:23:28 AM | 1996-04-12 | 6/3/1998 6:51:45 PM | EMS |  | Concentration | mg/L | Y | (fl. conc.) |  | Boron Total (fl. conc.) | 7440-42-8 | Metals - Total; | Metals - Total; EMS code: B--T | Chemical | Total | LAB
B--T | mg/L | Boron Total | F048 | Acid dig,conc'n,nebul.,ICP-OES | This method provides an automated procedure for the Inductively Coupled Argon Plasma Optical Emission Spectrophotometric (ICP-OES) determination and quantitation of trace amounts of metals in surface, ground and precipitation waters after a manual acid digestion and concentration procedure and ultrasonic nebulisation. Organically-bound metals and metal complexes are converted to metal ions by digestion with nitric acid [method details says HCl]. To increase the signal detected by the ICP, the samples are concentrated 10 times. The liquid sample is pumped onto the piezoelectric transducer, where the ultrasonic action efficiently converts it to a fine, dense aerosol. The nebulizer carrier gas flow then transports the wet aerosol through a heated U-tube, where the solvent is vaporized. Subsequently, most of the solvent vapour is removed from the aerosol stream by passing the stream through a chilled condenser. The result is a dry, analyte-laden aerosol which is introduced to the plasma. There, the high temperature desolvates the aerosol and dissociates the sample material into individual atoms and ions, which are excited to emit light at wave lengths characteristic of the elements present. Dissolved samples are field filtered through 0.45 micron cellulose acetate filter prior to acid preservation. Samples for extractable analysis are acidified with 2 mL concentrated HNO3 per litre and allowed to stand a minimum of 24 hours before analysis. Clients providing these samples are advised to preserve samples at the collection site. | 0.1 | Y | Y | Y | 29 |  |  | 7/6/2001 11:10:09 AM | 1998-11-13 |  | NPEPPIN |  | Concentration | mg/L | Y | (fl. conc.) |  | Boron Total (fl. conc.) | 7440-42-8 | Metals - Total; | Metals - Total; EMS code: B--T | Chemical | Total | LAB
B--T | mg/L | Boron Total | EQ01 | EQUIS (mg/L) | EQUIS (mg/L) | 0.01 | N | Y | Y | 198 |  |  | 10/22/1996 10:05:42 PM | 1996-04-12 | 6/3/1998 6:51:34 PM | EMS |  | Concentration | mg/L | Y | (fl. conc.) |  | Boron Total (fl. conc.) | 7440-42-8 | Metals - Total; | Metals - Total; EMS code: B--T | Chemical | Total | LAB
B--T | ppm (S) | Boron Total | 6041 | Aqua Regia by ICPMS | These were sediment/soils digested with Aqua-Regia (Nitric Acid & Hydrochloric Acids) and then analysis by ICPMS. Results would be reported in mg/kg and/or %. | 10 | Y | Y | Y | 21 | 10/13/2019 | 8/5/2020 | 8/10/2020 3:52:06 PM | 2019-09-06 | 10/21/2019 11:54:12 AM | KATFRASE | KATFRASE | DimensionlessRatio | ppm | N | (ratio) |  | Boron Total (ratio) | 7440-42-8 | Metals - Total; | Metals - Total; EMS code: B--T | Chemical | Total | LAB
B--T | ug/g wet | Boron Total | BPSY | MET-WET-BPSY-CCMS-VA | Metals in Animal Tissue and Vegetation (Biota) - Prescriptive. EPA 200.3/6020B This method is conducted following British Columbia Lab Manual method "Metals in Animal Tissue and Vegetation (Biota) - Prescriptive". Tissue samples are digested with nitric and hydrochloric acids; in combination with addition of hydrogen peroxide. Instrumental analysis is by collision cell inductively coupled plasma - mass spectrometry (modified from EPA Method 6020B). | 0.2 | Y | Y | Y | 2 | 11/7/2019 | 11/7/2019 | 6/10/2021 1:57:10 PM | 2021-06-04 |  | EMS |  | MassPerMass | µg/g | Y | (mass conc.) | (wet) | Boron Total (wet) (mass conc.) | 7440-42-8 | Metals - Total; | Metals - Total; EMS code: B--T | Chemical | Total | LAB
B--T | mg/L | Boron Total | X349 | FA:HNO3:ICP | Samples are field-acidified with HNO3; HNO3 digestion; analysis by ICP/AES. | 0.008 | Y | Y | Y | 30 |  |  | 10/30/2014 8:31:17 AM | 1998-05-22 | 6/3/1998 6:51:49 PM | NPEPPIN |  | Concentration | mg/L | Y | (fl. conc.) |  | Boron Total (fl. conc.) | 7440-42-8 | Metals - Total; | Metals - Total; EMS code: B--T | Chemical | Total | LAB
B--T | ug/g wet | Boron Total | X560 | Metals in Biota (wet), total, acid digestion, ICP-MS | Five grams of wet, homogenized biota is weighed in a centrifuge tube, and 2.5 mL of Nitric Acid and 2.5 mL of Hydrochloric acid are added. The sample is digested at room temperature for one hour and then digested at 90-95 degrees C in a water bath or heat block for one more hour. The sample is cooled and the volume is made to 50 mL with DI. The digestate is then diluted at ten fold and introduced to a collision/reaction cell inductively coupled argon plasma mass spectrometer (CRC ICPMS) by pneumatic nebulization. Each respective element is measured at a specific mass to charge ratio, m/z value expressed in amu (atomic mass units). - Maxxam Analytics (BC). Added to EMS Dec 10, 2008. | 5 | Y | Y | Y | 541 | 11/15/2008 4:00:00 AM | 10/30/2023 | 11/28/2023 2:21:33 PM | 2008-12-11 | 3/11/2015 4:17:10 PM | RDGIBSON | EMS | MassPerMass | µg/g | Y | (mass conc.) | (wet) | Boron Total (wet) (mass conc.) | 7440-42-8 | Metals - Total; | Metals - Total; EMS code: B--T | Chemical | Total | LAB
B--T | ug/g | Boron Total | 0192 | Hom, Nitric/Perchlor,ICP | Sample already in a dry state so does not require drying before homogenization and nitric perchlorid digestion followed by ICP analysis. | 1 | Y | Y | Y | 236 |  |  | 10/24/1996 11:23:41 AM | 1996-04-12 | 6/3/1998 6:51:34 PM | EMS |  | MassPerMass | mg/kg | Y | (mass conc.) |  | Boron Total (mass conc.) | 7440-42-8 | Metals - Total; | Metals - Total; EMS code: B--T | Chemical | Total | LAB
B--T | ug/m3 | Boron Total | F088 | Aliquot,HNO3/HCl dig, ICPMS,loading | Aliquot microwave digested with 3:1 HNO3/HCL. Analysis by ICPMS. For HiVol samples preliminary result are reported in mg/L and loading results are reported in ug/m3 |  | Y | Y | Y | 565 | 8/17/2000 10:46:00 AM | 12/14/2006 | 12/30/2006 1:01:35 AM | 2000-08-16 |  | NPEPPIN |  | AirConcentration | µg/m³ | Y | (air conc.) |  | Boron Total (air conc.) | 7440-42-8 | Metals - Total; | Metals - Total; EMS code: B--T | Chemical | Total | LAB
B--T | mg/L | Boron Total | 1349 | Tot Metals in Seawater by CRC ICPMS (BC) | Tot Metals in Seawater by CRC ICPMS (BC) | 0.3 | Y | Y | Y | 98 | 8/16/2019 | 12/5/2021 | 12/13/2021 12:34:16 PM | 2019-09-30 |  | KATFRASE |  | Concentration | mg/L | Y | (fl. conc.) |  | Boron Total (fl. conc.) | 7440-42-8 | Metals - Total; | Metals - Total; EMS code: B--T | Chemical | Total | LAB
B--T | ppm (S) | Boron Total | 6042 | Aqua Regia by ICPMS | These were sediment/soils digested with Aqua-Regia (Nitric Acid & Hydrochloric Acids) and then analysis by ICPMS. Results would be reported in ppm s. | 10 | Y | Y | Y | 71 | 5/14/2022 | 7/27/2024 | 7/29/2024 9:13:08 AM | 2022-06-28 |  | EMS |  | DimensionlessRatio | ppm | N | (ratio) |  | Boron Total (ratio) | 7440-42-8 | Metals - Total; | Metals - Total; EMS code: B--T | Chemical | Total | LAB
B--T | ug/g wet | Boron Total | MTIW | MET-WET-CCMS-MID-VA | Metals in Tissue by CRC ICPMS (WET) Methodology: This method is conducted following British Columbia Lab Manual method "Metals in Animal Tissue and Vegetation (Biota) - Prescriptive". Tissue samples are homogenized and sub-sampled prior to hotblock digestion with nitric and hydrochloric acids, in combination with addition of hydrogen peroxide. Instrumental analysis is by collision cell inductively coupled plasma - mass spectrometry (modified from EPA Method 6020A). Method Limitation: This method employs a strong acid/peroxide digestion, and is intended to provide a conservative estimate of bio-available metals. Near complete recoveries are achieved for most toxicologically important metals, but elements associated with recalcitrant minerals may be only partially recovered. | 0.2 | Y | Y | Y | 24 | 11/25/2018 | 12/6/2018 | 7/21/2021 2:55:07 PM | 2018-12-24 |  | EMS |  | MassPerMass | µg/g | Y | (mass conc.) | (wet) | Boron Total (wet) (mass conc.) | 7440-42-8 | Metals - Total; | Metals - Total; EMS code: B--T | Chemical | Total | LAB
B--T | ppm (S) | Boron Total | MS41 | ALS Minerals ME-MS41 | ALS Minerals ME-MS41. A prepared sample is digested with aqua regia in a graphite heating block. After cooling; the resulting solution is diluted to with deionized water; mixed and analyzed by ICP-AES. Samples are then analyzed by ICP-MS for the remaining suite of elements. | 10 | Y | Y | Y | 0 |  |  |  | 2024-01-29 |  | EMS |  | DimensionlessRatio | ppm | N | (ratio) |  | Boron Total (ratio) | 7440-42-8 | Metals - Total; | Metals - Total; EMS code: B--T | Chemical | Total | LAB
B--T | ug/m3 | Boron Total | 5312 | Tef:N/P Dig:ICP Analysis | A nitric/perchloric acid digestion if used to bring the metals into solution prior to ICP analysis. For Hi-Vol samples preliminary results are reported as mg/L and 24 hour loading results as ug/m3. | 2 | Y | Y | Y | 1185 | 11/7/1997 | 5/15/2000 | 11/18/2000 1:28:58 AM | 1996-10-13 | 6/3/1998 6:51:46 PM | EMS |  | AirConcentration | µg/m³ | Y | (air conc.) |  | Boron Total (air conc.) | 7440-42-8 | Metals - Total; | Metals - Total; EMS code: B--T | Chemical | Total | LAB
B--T | ug/g | Boron Total | FS82 | PBM,dry60C,10mesh,SALM-ICP-MS | PBM;dry60C;10mesh;SALM-ICP-MS | 4 | Y | Y | Y | 598 | 12/7/2009 3:52:00 PM | 7/24/2024 | 7/29/2024 5:17:09 PM | 2010-03-29 |  | EMS |  | MassPerMass | mg/kg | Y | (mass conc.) |  | Boron Total (mass conc.) | 7440-42-8 | Metals - Total; | Metals - Total; EMS code: B--T | Chemical | Total | LAB
B--T | ug/m3 | Boron Total | 6040 | Aqua Regia Dig:ICP | Aqua regia digestion is used to bring the metals into solution prior to ICP analysis. For Hi-Vol samples, preliminary results are reported as mg/L and24 hr loadings as ug/m3 | 2 | Y | Y | Y | 437 | 2/13/2002 | 2/9/2005 | 2/23/2005 1:13:46 AM | 1998-05-05 | 6/3/1998 6:51:48 PM | NPEPPIN |  | AirConcentration | µg/m³ | Y | (air conc.) |  | Boron Total (air conc.) | 7440-42-8 | Metals - Total; | Metals - Total; EMS code: B--T | Chemical | Total | LAB
B--T | mg/L | Boron Total | F033 | Acid Dig,Centrifuge,ICPMS | For the determination of total recoverable elements in water or waste water, take 100 ml aliquot from a well mixed, acid preserved, sample containing not more than 0.25% (w/v) total solids and transfer to a 250 ml Griffin beaker. (If total solids are greater than 0.25% reduce the size of the aliquot by a proportionate amount). Add 1 ml of conc. nitric acid and 0.5 ml conc. hydrochloric acid. Heat on a hot plate until the volume has been reduced to 15-20ml, ensuring that the sample does not boil. Allow to cool and transfer to a 50 ml volumetric flask. Dilute to volume with ASTM type I water. Centrifuge the sample or allow to stand overnight to separate insoluble material. Prior analysis, pipette 20 ml into a 50 ml volumetric flask and dilute to volume with ASTM type I water. Analysis should be performed within seventy two hours of the completed preparation. Note that different laboratories may have very different field preservation, digestion procedures and detection limits for this method, so these should be documented. | 0.001 | Y | Y | Y | 30524 | 12/16/1996 7:35:00 AM | 8/22/2024 3:24:00 AM | 8/25/2024 9:47:36 PM | 1997-01-16 | 6/3/1998 6:51:46 PM | MPAK |  | Concentration | mg/L | Y | (fl. conc.) |  | Boron Total (fl. conc.) | 7440-42-8 | Metals - Total; | Metals - Total; EMS code: B--T | Chemical | Total | LAB
B--T | mg/L | Boron Total | X170 | ICP | This code is a catch-all for a large variety of ICP methods including water and also solids digested into solutions as an intermediate result. Different procedures will have greatly different detection limits. (a) For air particulates, preliminary results are reported in mg/L and loading results as mg/dm2/d. (b) For Hi Vol samples preliminiary results are same as reported in mg/L and loading results as ug/m3. (c) Note that for Rain samples there is no filtration nor digestion steps. | 0.1 | Y | Y | Y | 6740 | 3/28/1996 6:03:00 PM | 11/7/2018 | 10/24/2022 10:58:25 AM | 1996-04-12 | 8/15/2019 11:51:53 AM | EMS | KATFRASE | Concentration | mg/L | Y | (fl. conc.) |  | Boron Total (fl. conc.) | 7440-42-8 | Metals - Total; | Metals - Total; EMS code: B--T | Chemical | Total | LAB
B--T | ug/g | Boron Total | EQ02 | EQUIS (ug/g) | EQUIS (ug/g) | 1 | N | Y | Y | 743 |  |  | 10/23/1996 12:08:28 AM | 1996-04-12 | 6/3/1998 6:51:34 PM | EMS |  | MassPerMass | mg/kg | Y | (mass conc.) |  | Boron Total (mass conc.) | 7440-42-8 | Metals - Total; | Metals - Total; EMS code: B--T | Chemical | Total | LAB
B--T | mg/L | Boron Total | F082 | Aliquot,HNO3/HCL dig,ICPMS | Aliquot microwave digested with 3:1 HNO3/HCL. Analysis by ICPMS. For HiVol samples preliminary result are reported in mg/L and loading results are reported in ug/m3 |  | Y | Y | Y | 27601 | 1/24/2011 | 8/23/2024 | 8/26/2024 4:53:05 PM | 2000-06-13 |  | NPEPPIN |  | Concentration | mg/L | Y | (fl. conc.) |  | Boron Total (fl. conc.) | 7440-42-8 | Metals - Total; | Metals - Total; EMS code: B--T | Chemical | Total | LAB
B--T | mg/L | Boron Total | MSEA | Metals in Seawater (undigested) by (CRC) ICP-MS | This method describes the multi-elemental determination of trace elements by Collision/Reaction Cell (CRC) ICP-MS. The method measures ions produced by a radio-frequency inductively coupled plasma. Analyte species originating in a liquid are pneumatically nebulized and the resulting aerosol transported by argon gas into the plasma torch. The ions produced are entrained in the plasma gas and extracted, by means of a differentially pumped vacuum interface, into a mass spectrometer. The ions flow through a Collision/Reaction Cell (CRC) where depending on the mode of analysis either the ions passes through the cell unhindered (no gas - classical ICP-MS mode), or collisions/reactions chemistry takes place between the ions and a neutral gas such as hydrogen or helium. Any polyatomic ions which give interference on the analytes are removed. The ions are then sorted according to their mass-to-charge ratios by a quadrupole mass spectrometer. The ions are then sorted according to their mass-to-charge ratios by a quadrupole mass spectrometer. The ions transmitted through the quadrupole are quantified by a channel electron multiplier and the ion information processed by a data handling system. Although most polyatomic interference is removed, any additional interference must be assessed and valid corrections applied or the data flagged to indicate problems. Interference correction must include compensation for background ions contributed by the plasma gas, reagents, and constituents of the sample matrix. Instrument drift as well as suppressions or enhancements of instrument response caused by the sample matrix must be corrected for by the use of internal standards. | 0.0004 | Y | Y | Y | 441 | 8/12/2009 12:08:00 AM | 4/29/2024 | 6/7/2024 10:03:16 AM | 2009-09-08 |  | RDGIBSON |  | Concentration | mg/L | Y | (fl. conc.) |  | Boron Total (fl. conc.) | 7440-42-8 | Metals - Total; | Metals - Total; EMS code: B--T | Chemical | Total | LAB
B--T | mg/L | Boron Total | F121 | Acid,ICPMS | Acidified, no digestion, ICPMS | 0.05 | Y | Y | Y | 8436 | 3/12/2010 7:57:00 PM | 8/15/2024 9:10:00 PM | 8/26/2024 8:06:03 AM | 2002-07-17 | 7/16/2012 3:25:39 PM | UMULLER | RDGIBSON | Concentration | mg/L | Y | (fl. conc.) |  | Boron Total (fl. conc.) | 7440-42-8 | Metals - Total; | Metals - Total; EMS code: B--T | Chemical | Total | LAB
B--T | mg/L | Boron Total | MWID | Metals in Water by ICPMS (digested) | Metals in Water by ICPMS (digested) Based on EPA 200.8 | 0.1 | Y | Y | Y | 36355 |  |  | 7/31/2024 1:46:41 PM | 2009-12-24 |  | RDGIBSON |  | Concentration | mg/L | Y | (fl. conc.) |  | Boron Total (fl. conc.) | 7440-42-8 | Metals - Total; | Metals - Total; EMS code: B--T | Chemical | Total | LAB
B--T | ug/g | Boron Total | PB02 | PBM,HNO3/HCl dig,dry60C,100mesh, | Performance Based Method. Samples are dried at 60 C and sieved using a .149 mm (100 mesh) sieve. Solid samples are digested with the combination of nitric and hydrochloric acids with a 1:1 ratio of acids. For environmental purposes, strong acid leaches are used routinely since they have the advantage of speed, low cost, and superio reproducibility. All metals trapped within the refractory matrix are unlikely to have environmental consequences. Instrumental analysis of extracts produced using this digestion method can be performed by a variety of techniques. |  | Y | Y | Y | 10 | 9/13/2001 | 9/13/2001 | 11/6/2002 3:52:32 PM | 2001-07-18 |  | NPEPPIN |  | MassPerMass | mg/kg | Y | (mass conc.) |  | Boron Total (mass conc.) | 7440-42-8 | Metals - Total; | Metals - Total; EMS code: B--T | Chemical | Total | LAB
B--T | mg/L | Boron Total | TCLP | Toxicity Characteristic Leaching Procedure | The TCLP procedure is designed to simulate the climatic leaching action | 0.1 | Y | Y | Y | 38 | 12/8/2009 7:49:00 PM | 9/28/2011 5:53:00 PM | 8/24/2016 2:59:13 PM | 2009-03-06 |  | JCHUDAK |  | Concentration | mg/L | Y | (fl. conc.) |  | Boron Total (fl. conc.) | 7440-42-8 | Metals - Total; | Metals - Total; EMS code: B--T | Chemical | Total | LAB
B--T | ug/g | Boron Total | FB04 | Freeze Dry,Hom,Acid Dig,ICP | Sample moisture is determined by weighing samples before and after freeze-drying. Freeze-dried samples are ground in a Braun blender before being weighed into Teflon PFA liners and digested with concentrated HNO3, concentrated HCl, and 50% H2O2 in a microwave oven; which results in a complete sample dissolution. Samples are collected and volumized with Type I water. Analysis is performed via Inductively Coupled Argon Plasma Atomic Emission Spectrometry (ICP-AES) (simultaneous multielement analysis), Graphite Furnace Atomic Absorption Spectrometry (GF-AAS) (low level Cd and Pb analysis), Cold Vapor Atomic Fluorescence Spectrometry (CV-AFS) (Mercury analysis) and/or Flame Atomic Absorption Spectrometry (F-AAS) (K analysis). | 0.8 | Y | Y | Y | 195 | 8/26/1998 6:05:00 PM | 2/1/2002 2:30:00 PM | 6/23/2003 1:15:28 PM | 1997-04-02 | 6/3/1998 6:51:46 PM | MPAK |  | MassPerMass | mg/kg | Y | (mass conc.) |  | Boron Total (mass conc.) | 7440-42-8 | Metals - Total; | Metals - Total; EMS code: B--T | Chemical | Total | LAB
B--T | ug/g | Boron Total | MTID | MET-DRY-CCMS-VA | Metals in Tissue by CRC ICPMS (DRY) Methodology: This method is conducted following British Columbia Lab Manual method "Metals in Animal Tissue and Vegetation (Biota) - Prescriptive". Tissue samples are homogenized and sub-sampled prior to hotblock digestion with nitric and hydrochloric acids, in combination with addition of hydrogen peroxide. Instrumental analysis is by collision cell inductively coupled plasma - mass spectrometry (modified from EPA Method 6020A). Method Limitation: This method employs a strong acid/peroxide digestion, and is intended to provide a conservative estimate of bio-available metals. Near complete recoveries are achieved for most toxicologically important metals, but elements associated with recalcitrant minerals may be only partially recovered. | 1 | Y | Y | Y | 122 | 8/13/2014 | 9/8/2022 | 10/26/2022 7:14:14 AM | 2015-03-09 | 3/11/2015 4:12:05 PM | RDGIBSON | EMS | MassPerMass | mg/kg | Y | (mass conc.) |  | Boron Total (mass conc.) | 7440-42-8 | Metals - Total; | Metals - Total; EMS code: B--T | Chemical | Total | LAB
B--T | ug/g | Boron Total | ICPM | Metals in Soil by CRC ICPMS | This analysis is carried out using procedures from CSR Analytical Method: "Strong Acid Leachable Metals (SALM) in Soil", BC Ministry of Environment, 26 June 2009, and procedures adapted from EPA Method 200.2. The sample is manually homogenized, dried at 60 degrees Celsius, sieved through a 63 um (230 mesh) sieve, and a representative subsample of the dry material is weighed. The sample is then digested at 95 degrees Celsius for 2 hours by block digester using concentrated nitric and hydrochloric acids. Instrumental analysis is by collision cell inductively coupled plasma - mass spectrometry (modifed from EPA Method 6020A). Method Limitation: This method is not a total digestion technique. It is a very strong acid digestion that is intended to dissolve those metals that may be environmentally available. By design, elements bound in silicate structures are not normally dissolved by this procedure as they are not usually mobile in the environment. | 10 | Y | Y | Y | 103 | 9/19/2017 | 9/24/2023 | 1/12/2024 11:16:52 AM | 2018-01-09 |  | JKROGH |  | MassPerMass | mg/kg | Y | (mass conc.) |  | Boron Total (mass conc.) | 7440-42-8 | Metals - Total; | Metals - Total; EMS code: B--T | Chemical | Total | LAB
B--T | mg/L | Boron Total | F038 | HNO3/HCl dig,ICP-AES | Dissolved and extractable water samples are ready for analysis once collected. Total samples are autoclaved with HNO3 and hydrochloric acid (HCI) in capped Teflon vials. Analysis is performed via Inductively Coupled Argon Plasma Atomic Emission Spectrometry (ICP-AES)(simultaneous multielement analysis), Graphite Furnace Atomic Absorption Spectrometry (GF-AAS) (low level analysis) and/or Flame Atomic Absorption Spectrometry (F-AAS). | 0.01 | Y | Y | Y | 5398 | 8/1/1996 6:04:00 PM | 6/25/2010 3:15:00 PM | 9/11/2019 8:47:42 AM | 1997-05-08 | 6/3/1998 6:51:46 PM | MPAK |  | Concentration | mg/L | Y | (fl. conc.) |  | Boron Total (fl. conc.) | 7440-42-8 | Metals - Total; | Metals - Total; EMS code: B--T | Chemical | Total | LAB
B--T | ug/g | Boron Total | X230 | Hom,HNO3/HCl,ICP | The sample is dried, ground to pass 100 mesh, and digested with a nitric-hydrochloric acid mixture to solubilize the solid matter and to remove the organic material by oxidation and volatilization. The completed digests are submitted to metals instrumentation personnel for metals analysis by Inductively Coupled Atomic Emission Spectrophotometry (ICP-AES), or hydride analysis by atomic absorption. | 1 | Y | Y | Y | 36 | 2/12/2002 | 5/29/2002 | 1/14/2004 1:00:43 AM | 1998-11-05 |  | NPEPPIN |  | MassPerMass | mg/kg | Y | (mass conc.) |  | Boron Total (mass conc.) | 7440-42-8 | Metals - Total; | Metals - Total; EMS code: B--T | Chemical | Total | LAB
B--T | ug/g wet | Boron Total | X552 | Metals in Biota (wet), total, acid digestion, ICP-SFMS | One gram of wet, homogenized biota is weighed in a microwave TFM vessel, and 8mL of Nitric Acid and 1mL of Hydrogen Peroxide are added. The sample is digested at 200 degrees C for 15 minutes in a closed vessel high pressure microwave oven. The sample is cooled, the volume adjusted to 100mL and introduced to an inductively coupled argon plasma-sector field mass spectrometer (ICP-SFMS) by pneumatic nebulization. Each respective element is measured at a specific mass to charge ratio, m/z value expressed in amu (atomic mass units), and at a specified instrument resolution. LR: Low Resolution ~300, MR: Mid Resolution ~4000. Ref: 02-2704, National Laboratory for Environmental Testing, CCIW, Environment Canada, Burlington. Requested April 2003. | 0.2 | Y | Y | Y | 37 | 8/9/2016 | 10/16/2023 | 11/28/2023 2:21:32 PM | 2015-03-11 |  | RDGIBSON |  | MassPerMass | µg/g | Y | (mass conc.) | (wet) | Boron Total (wet) (mass conc.) | 7440-42-8 | Metals - Total; | Metals - Total; EMS code: B--T | Chemical | Total | LAB
B002 | ug/g | Bromomethane | MS02 | Purge&Trap GC/MS (VOC-S) | Purge&Trap GC/MS (VOC-S) | 0.012 | Y | Y | Y | 11 | 11/19/2004 | 11/29/2004 1:55:00 PM | 1/6/2005 1:01:54 AM | 1996-04-12 | 6/3/1998 6:51:34 PM | EMS |  | MassPerMass | mg/kg | Y | (mass conc.) |  | Bromomethane (mass conc.) | 74-83-9 | Volatile Organic Compounds (VOCs); | Volatile Organic Compounds (VOCs); EMS code: B002 | Chemical |  | LAB
B002 | mg/L | Bromomethane | MS01 | Purge&Trap GC/MS (VOC-W) | Purge&Trap GC/MS (VOC-W) | 2.4 | Y | Y | Y | 246 | 10/8/2004 | 7/26/2012 7:12:00 AM | 9/26/2012 1:07:08 PM | 1996-04-12 | 3/18/2003 2:49:40 PM | EMS | UMULLER | Concentration | mg/L | Y | (fl. conc.) |  | Bromomethane (fl. conc.) | 74-83-9 | Volatile Organic Compounds (VOCs); | Volatile Organic Compounds (VOCs); EMS code: B002 | Chemical |  | LAB
B002 | mg/L | Bromomethane | MS04 | Static Headspace GC/MS (Water) | (Water) A static headspace procedure is used as a means of driving the analytes out of the matrix into a headspace. Once equilibrium is achieved, a fixed volume of headspeace is taken for subsequent determination by Gas Chromatography - Mass Spectroscopy (GC/MS) using Selected Ion Monitoring. | 0.001 | Y | Y | Y | 51 | 8/13/2009 | 6/28/2024 12:29:00 AM | 7/9/2024 3:31:02 PM | 2009-12-18 |  | RDGIBSON |  | Concentration | mg/L | Y | (fl. conc.) |  | Bromomethane (fl. conc.) | 74-83-9 | Volatile Organic Compounds (VOCs); | Volatile Organic Compounds (VOCs); EMS code: B002 | Chemical |  | LAB
B005 | mg/L | Bromobenzene | MS04 | Static Headspace GC/MS (Water) | (Water) A static headspace procedure is used as a means of driving the analytes out of the matrix into a headspace. Once equilibrium is achieved, a fixed volume of headspeace is taken for subsequent determination by Gas Chromatography - Mass Spectroscopy (GC/MS) using Selected Ion Monitoring. | 0.002 | Y | Y | Y | 8 | 10/5/2019 2:07:00 AM | 6/28/2024 12:29:00 AM | 7/9/2024 3:31:02 PM | 2021-04-07 |  | RDGIBSON |  | Concentration | mg/L | Y | (fl. conc.) |  | Bromobenzene (fl. conc.) | 108-86-1 | Volatile Organic Compounds (VOCs); | Volatile Organic Compounds (VOCs); EMS code: B005 | Chemical |  | LAB
B006 | % | 4-Bromofluorobenzene | F007 | Static Headspace:HRGC/LRMS | Static headspace analysis is a method for the determination of volatile compounds in liquids and/or solids by analysing the vapour phase of a sample in a closed system; in this case a sealed vial. Static Headspace is sometimes referred to as Equilibrium Headspace because the vapour phase above the sample in a sealed vial is brought into thermodynamic equilibrium with the sample prior to analysis. Volatile organic compounds are thermally extracted from the sample matrix by Static Headspace sampling and introduced directly to a thick-film megabore GC column. Compounds eluting from the GC column are introduced into a Mass Spectrometer and target compounds are identified by a comparison of both retention times and mass spectra with those of reference standards stored in a computerized data system. Reference retention times and mass spectra for analytes are obtained by the measurement of calibration standards under precisely the same extraction conditions used for samples. The concentration of each identified component is measured by relating the response of the sample quantitation mass ions to the response of the quantitation mass ions of a series of external standards. A deuterated surrogate recovery standard is added to provide information on matrix effects. | 1 | Y | Y | Y | 96 | 3/20/2007 | 12/22/2014 | 1/29/2015 2:12:57 PM | 2012-06-25 |  | RDGIBSON |  | DimensionlessRatio | % | N | (ratio) |  | 4-Bromofluorobenzene (ratio) | 460-00-4 | Organics; | Organics; EMS code: B006 | Chemical |  | LAB
B006 | % | 4-Bromofluorobenzene | MS04 | Static Headspace GC/MS (Water) | (Water) A static headspace procedure is used as a means of driving the analytes out of the matrix into a headspace. Once equilibrium is achieved, a fixed volume of headspeace is taken for subsequent determination by Gas Chromatography - Mass Spectroscopy (GC/MS) using Selected Ion Monitoring. | 1 | Y | Y | Y | 46 | 1/9/2015 | 5/22/2015 | 7/3/2015 3:19:03 PM | 2015-04-15 |  | RDGIBSON |  | DimensionlessRatio | % | N | (ratio) |  | 4-Bromofluorobenzene (ratio) | 460-00-4 | Organics; | Organics; EMS code: B006 | Chemical |  | LAB
B006 | mg/L | 4-Bromofluorobenzene | E611 | E611S: EPA 8260D (mod) | E611S: Dimethyl selenide and dimethyl diselenide are analyzed by static headspace GC-MS. Samples are prepared in headspace vials and are heated and agitated on the headspace autosampler; causing VOCs to partition between the aqueous phase and the headspace in accordance with Henry`s law. | 0.001 | Y | Y | Y | 0 |  |  |  | 2024-01-15 | 1/15/2024 2:17:10 PM | EMS | EMS | Concentration | mg/L | Y | (fl. conc.) |  | 4-Bromofluorobenzene (fl. conc.) | 460-00-4 | Organics; | Organics; EMS code: B006 | Chemical |  | LAB
B006 | % | 4-Bromofluorobenzene | SURR | Surrogate Recovery | Surrogate Recovery | 1 | Y | Y | Y | 1333 | 11/1/2011 | 8/14/2024 | 8/16/2024 4:05:21 PM | 2015-07-07 |  | RDGIBSON |  | DimensionlessRatio | % | N | (ratio) |  | 4-Bromofluorobenzene (ratio) | 460-00-4 | Organics; | Organics; EMS code: B006 | Chemical |  | SURROGATE_RESULT
B006 | mg/L | 4-Bromofluorobenzene | 8260D | EPA 8260D (mod) | EPA 8260D (mod) VOLATILE ORGANIC COMPOUNDS BY GAS CHROMATOGRAPHY/MASS SPECTROMETRY This method is used to determine volatile organic compounds (VOCs) in a variety of solid waste matrices. | 0.005 | Y | Y | Y | 0 |  |  |  | 2024-01-24 |  | RDGIBSON |  | Concentration | mg/L | Y | (fl. conc.) |  | 4-Bromofluorobenzene (fl. conc.) | 460-00-4 | Organics; | Organics; EMS code: B006 | Chemical |  | LAB
B006 | mg/L | 4-Bromofluorobenzene | MS01 | Purge&Trap GC/MS (VOC-W) | Purge&Trap GC/MS (VOC-W) |  | Y | Y | Y | 785 | 1/15/2007 4:50:00 PM | 6/4/2024 10:24:00 AM | 6/13/2024 5:01:27 PM | 2003-06-17 |  | UMULLER |  | Concentration | mg/L | Y | (fl. conc.) |  | 4-Bromofluorobenzene (fl. conc.) | 460-00-4 | Organics; | Organics; EMS code: B006 | Chemical |  | LAB
B007 | ug/g | Bromophos | ZN03 | DCM+Ac.Extr:GC/MS | DCM+Ac.Extr:GC/MS EPA 8270E (mod) | 0.01 | Y | Y | Y | 19 | 4/20/2002 10:25:00 AM | 10/16/2002 3:20:00 PM | 2/25/2004 1:09:43 AM | 1996-04-12 | 11/16/1999 2:27:40 PM | EMS | NPEPPIN | MassPerMass | mg/kg | Y | (mass conc.) |  | Bromophos (mass conc.) | 2104-96-3 | Pesticides; | Pesticides; EMS code: B007 | Chemical |  | LAB
B007 | mg/L | Bromophos | ZN01 | DCM+Ac.Extr:GC/MS | DCM+Ac.Extr:GC/MS | 0.0001 | Y | Y | Y | 55 | 10/3/2002 | 3/29/2012 10:15:00 PM | 5/3/2012 9:27:49 AM | 1996-04-12 | 11/16/1999 4:06:27 PM | EMS | NPEPPIN | Concentration | mg/L | Y | (fl. conc.) |  | Bromophos (fl. conc.) | 2104-96-3 | Pesticides; | Pesticides; EMS code: B007 | Chemical |  | LAB
B008 | mg/L | Bromophos Ethyl | ZN01 | DCM+Ac.Extr:GC/MS | DCM+Ac.Extr:GC/MS | 0.0003 | Y | Y | Y | 43 | 7/29/2005 7:10:00 AM | 3/29/2012 10:15:00 PM | 5/3/2012 9:27:49 AM | 1996-04-12 | 11/16/1999 4:06:27 PM | EMS | NPEPPIN | Concentration | mg/L | Y | (fl. conc.) |  | Bromophos Ethyl (fl. conc.) | 4824-78-6 | Pesticides; | Pesticides; EMS code: B008 | Chemical |  | LAB
B009 | mg/L | Bromoxynil | P015 | Phenoxyacid Herbicides by LC-MS/MS | Phenoxyacid Herbicides by LC-MS/MS | 0.5 | Y | Y | Y | 71 | 10/9/2019 | 9/9/2022 | 12/8/2022 12:57:05 PM | 2019-10-17 |  | KATFRASE |  | Concentration | mg/L | Y | (fl. conc.) |  | Bromoxynil (fl. conc.) | 1689-84-5 | Pesticides; | Pesticides; EMS code: B009 | Chemical |  | LAB
B009 | ug/g | Bromoxynil | FS15 | DCM/MEOH acid extr'n,DCM,methn,Florisil,HRGC-EC | Sediment samples are shaken with a mixture of Dichloromethane, Methanol and Sulphuric Acid, filtered, then extracted with Dichloromethane. After concentration, the extract is methylated using ethereal Diazomethane. Extracts are cleaned up as required using Florisil. The extracts are analysed for Acid Extractable Herbicides using High Resolution Gas Chromatography with Electron Capture detection. Positive results may be verified by use of a secondary confirmation GC column. |  | Y | Y | Y | 27 | 4/4/2002 | 6/24/2003 | 2/26/2004 1:05:26 AM | 1998-12-30 |  | NPEPPIN |  | MassPerMass | mg/kg | Y | (mass conc.) |  | Bromoxynil (mass conc.) | 1689-84-5 | Pesticides; | Pesticides; EMS code: B009 | Chemical |  | LAB
B009 | mg/L | Bromoxynil | AEWA | H2SO4 acid,CHCl2 extr.florisil. GC/MS | Effluent and water samples are acidified with Sulphuric Acid and extracted with Dichloromethane. After concentration, the extract is methylated using ethereal Diazomethane. Extracts are cleaned up as required using Florisil. The extracts are analyzed for Acid Extractable Herbicides using GC/MS. | 0.0002 | Y | Y | Y | 6 | 6/16/2003 | 7/29/2005 7:10:00 AM | 6/2/2009 12:00:57 PM | 1998-11-17 |  | NPEPPIN |  | Concentration | mg/L | Y | (fl. conc.) |  | Bromoxynil (fl. conc.) | 1689-84-5 | Pesticides; | Pesticides; EMS code: B009 | Chemical |  | LAB
B009 | ug/g | Bromoxynil | ZN03 | DCM+Ac.Extr:GC/MS | DCM+Ac.Extr:GC/MS EPA 8270E (mod) | 0.0005 | Y | Y | Y | 17 | 1/14/2022 | 11/2/2022 4:41:00 AM | 1/11/2024 1:07:07 PM | 2022-07-27 |  | EMS |  | MassPerMass | mg/kg | Y | (mass conc.) |  | Bromoxynil (mass conc.) | 1689-84-5 | Pesticides; | Pesticides; EMS code: B009 | Chemical |  | LAB
B010 | % | Bromofluorobenzene | SURR | Surrogate Recovery | Surrogate Recovery | 0 | Y | Y | Y | 72 | 2/7/2002 10:55:00 AM | 11/29/2004 1:55:00 PM | 1/6/2005 1:01:54 AM | 1996-04-12 | 6/3/1998 6:51:34 PM | EMS |  | DimensionlessRatio | % | N | (ratio) |  | Bromofluorobenzene (ratio) | 460-00-4 | Organics; | Organics; EMS code: B010 | Chemical |  | SURROGATE_RESULT
B012 | mg/L | Bromodichloromethane | E611 | E611S: EPA 8260D (mod) | E611S: Dimethyl selenide and dimethyl diselenide are analyzed by static headspace GC-MS. Samples are prepared in headspace vials and are heated and agitated on the headspace autosampler; causing VOCs to partition between the aqueous phase and the headspace in accordance with Henry`s law. | 0.001 | Y | Y | Y | 1 | 12/14/2023 | 12/14/2023 | 1/16/2024 8:06:04 AM | 2024-01-15 | 1/15/2024 2:17:10 PM | EMS | EMS | Concentration | mg/L | Y | (fl. conc.) |  | Bromodichloromethane (fl. conc.) | 75-27-4 | Volatile Organic Compounds (VOCs); | Volatile Organic Compounds (VOCs); EMS code: B012 | Chemical |  | LAB
B012 | mg/L | Bromodichloromethane | MS04 | Static Headspace GC/MS (Water) | (Water) A static headspace procedure is used as a means of driving the analytes out of the matrix into a headspace. Once equilibrium is achieved, a fixed volume of headspeace is taken for subsequent determination by Gas Chromatography - Mass Spectroscopy (GC/MS) using Selected Ion Monitoring. | 0.001 | Y | Y | Y | 92 | 8/13/2009 | 6/28/2024 12:29:00 AM | 7/23/2024 11:41:49 AM | 2009-12-18 |  | RDGIBSON |  | Concentration | mg/L | Y | (fl. conc.) |  | Bromodichloromethane (fl. conc.) | 75-27-4 | Volatile Organic Compounds (VOCs); | Volatile Organic Compounds (VOCs); EMS code: B012 | Chemical |  | LAB
B012 | mg/L | Bromodichloromethane | P016 | Trihalomethanes | Trihalomethanes | 0.001 | Y | Y | Y | 57 | 3/24/2023 10:24:00 PM | 12/17/2023 11:36:00 AM | 12/29/2023 9:48:33 AM | 1996-04-12 | 6/3/1998 6:51:34 PM | EMS |  | Concentration | mg/L | Y | (fl. conc.) |  | Bromodichloromethane (fl. conc.) | 75-27-4 | Volatile Organic Compounds (VOCs); | Volatile Organic Compounds (VOCs); EMS code: B012 | Chemical |  | LAB
B012 | ug/g | Bromodichloromethane | MS02 | Purge&Trap GC/MS (VOC-S) | Purge&Trap GC/MS (VOC-S) | 0.002 | Y | Y | Y | 27 | 11/19/2004 | 10/28/2022 5:04:00 PM | 1/11/2024 1:07:07 PM | 1996-04-12 | 6/3/1998 6:51:45 PM | EMS |  | MassPerMass | mg/kg | Y | (mass conc.) |  | Bromodichloromethane (mass conc.) | 75-27-4 | Volatile Organic Compounds (VOCs); | Volatile Organic Compounds (VOCs); EMS code: B012 | Chemical |  | LAB
B012 | mg/L | Bromodichloromethane | MS01 | Purge&Trap GC/MS (VOC-W) | Purge&Trap GC/MS (VOC-W) | 0.4 | Y | Y | Y | 548 | 4/15/2002 | 6/4/2024 10:24:00 AM | 6/13/2024 5:01:27 PM | 1996-04-12 | 3/18/2003 2:49:40 PM | EMS | UMULLER | Concentration | mg/L | Y | (fl. conc.) |  | Bromodichloromethane (fl. conc.) | 75-27-4 | Volatile Organic Compounds (VOCs); | Volatile Organic Compounds (VOCs); EMS code: B012 | Chemical |  | LAB
B012 | ug/g | Bromodichloromethane | FS04 | Static Headspace:HRGC/LRMS | Static Headspace anaylsis is a method for the determination of volatile compounds in liquids and/or solids by analysing the vapour phase of a sample in a closed system; in this case a sealed vial. Static Headspace is sometimes referred to as Equilibrium Headspace because the vapour phase above the sample in a sealed vial is brought into thermodynamic equilibrium with the sample prior to analysis. Volatile organic compounds are thermally extracted from the sample matrix by Static Headspace sampling and introduced directly to a thick-film megabore GC column. Compounds eluting from the GC column are introduced into a Mass Spectrometer and target compounds are identified by a comparison of both retention times and mass spectra with those of reference standards stored in a computerized data system. Reference retention times and mass spectra for analytes are obtained by the measurement of calibration standards under precisely the same extraction conditions used for samples. The concentration of each identified component is measured by relating the response of the sample quantitation mass ions to the response of the quantitation mass ions of a series of external standards. A deuterated surrogate recovery standard is added to provide information on matrix effects. | 0.05 | Y | Y | Y | 44 | 6/30/1999 4:41:00 PM | 5/22/2015 | 7/3/2015 3:19:03 PM | 1996-09-16 | 6/3/1998 6:51:44 PM | MPAK |  | MassPerMass | mg/kg | Y | (mass conc.) |  | Bromodichloromethane (mass conc.) | 75-27-4 | Volatile Organic Compounds (VOCs); | Volatile Organic Compounds (VOCs); EMS code: B012 | Chemical |  | LAB
B012 | mg/L | Bromodichloromethane | X384 | Purge and trap GC/MS | The sample is analyzed by purge and trap gas chromatography with detection by mass spectrometry. The procedure involves purging the volatile compounds from the sample with an inert gas, and trapping them on a solid sorbent. When purging of the sample is complete, the trap is heated and the compounds are transferred to the gas chromatographic column. Analysis is then accomplished by separation of the components by gas chromatography with detection by mass spectrometry. | 0.001 | Y | Y | Y | 112 | 8/5/2004 10:31:00 AM | 7/26/2012 7:12:00 AM | 9/26/2012 1:07:08 PM | 1996-04-12 | 6/3/1998 6:51:45 PM | EMS |  | Concentration | mg/L | Y | (fl. conc.) |  | Bromodichloromethane (fl. conc.) | 75-27-4 | Volatile Organic Compounds (VOCs); | Volatile Organic Compounds (VOCs); EMS code: B012 | Chemical |  | LAB
B012 | mg/L | Bromodichloromethane | F007 | Static Headspace:HRGC/LRMS | Static headspace analysis is a method for the determination of volatile compounds in liquids and/or solids by analysing the vapour phase of a sample in a closed system; in this case a sealed vial. Static Headspace is sometimes referred to as Equilibrium Headspace because the vapour phase above the sample in a sealed vial is brought into thermodynamic equilibrium with the sample prior to analysis. Volatile organic compounds are thermally extracted from the sample matrix by Static Headspace sampling and introduced directly to a thick-film megabore GC column. Compounds eluting from the GC column are introduced into a Mass Spectrometer and target compounds are identified by a comparison of both retention times and mass spectra with those of reference standards stored in a computerized data system. Reference retention times and mass spectra for analytes are obtained by the measurement of calibration standards under precisely the same extraction conditions used for samples. The concentration of each identified component is measured by relating the response of the sample quantitation mass ions to the response of the quantitation mass ions of a series of external standards. A deuterated surrogate recovery standard is added to provide information on matrix effects. | 0.005 | Y | Y | Y | 381 | 9/28/1998 3:08:00 PM | 8/14/2024 | 8/16/2024 4:05:21 PM | 1996-09-16 | 6/3/1998 6:51:44 PM | MPAK |  | Concentration | mg/L | Y | (fl. conc.) |  | Bromodichloromethane (fl. conc.) | 75-27-4 | Volatile Organic Compounds (VOCs); | Volatile Organic Compounds (VOCs); EMS code: B012 | Chemical |  | LAB
B013 | mg/L | Bromoform | P016 | Trihalomethanes | Trihalomethanes | 0.001 | Y | Y | Y | 57 | 3/24/2023 10:24:00 PM | 12/17/2023 11:36:00 AM | 12/29/2023 9:48:33 AM | 1996-04-12 | 6/3/1998 6:51:34 PM | EMS |  | Concentration | mg/L | Y | (fl. conc.) |  | Bromoform (fl. conc.) | 75-25-2 | Volatile Organic Compounds (VOCs); | Volatile Organic Compounds (VOCs); EMS code: B013 | Chemical |  | LAB
B013 | mg/L | Bromoform | F007 | Static Headspace:HRGC/LRMS | Static headspace analysis is a method for the determination of volatile compounds in liquids and/or solids by analysing the vapour phase of a sample in a closed system; in this case a sealed vial. Static Headspace is sometimes referred to as Equilibrium Headspace because the vapour phase above the sample in a sealed vial is brought into thermodynamic equilibrium with the sample prior to analysis. Volatile organic compounds are thermally extracted from the sample matrix by Static Headspace sampling and introduced directly to a thick-film megabore GC column. Compounds eluting from the GC column are introduced into a Mass Spectrometer and target compounds are identified by a comparison of both retention times and mass spectra with those of reference standards stored in a computerized data system. Reference retention times and mass spectra for analytes are obtained by the measurement of calibration standards under precisely the same extraction conditions used for samples. The concentration of each identified component is measured by relating the response of the sample quantitation mass ions to the response of the quantitation mass ions of a series of external standards. A deuterated surrogate recovery standard is added to provide information on matrix effects. | 0.01 | Y | Y | Y | 381 | 9/28/1998 3:08:00 PM | 8/14/2024 | 8/16/2024 4:05:21 PM | 1996-09-16 | 6/3/1998 6:51:44 PM | MPAK |  | Concentration | mg/L | Y | (fl. conc.) |  | Bromoform (fl. conc.) | 75-25-2 | Volatile Organic Compounds (VOCs); | Volatile Organic Compounds (VOCs); EMS code: B013 | Chemical |  | LAB
B013 | mg/L | Bromoform | MS04 | Static Headspace GC/MS (Water) | (Water) A static headspace procedure is used as a means of driving the analytes out of the matrix into a headspace. Once equilibrium is achieved, a fixed volume of headspeace is taken for subsequent determination by Gas Chromatography - Mass Spectroscopy (GC/MS) using Selected Ion Monitoring. | 0.001 | Y | Y | Y | 75 | 8/13/2009 | 6/28/2024 12:29:00 AM | 7/9/2024 3:31:02 PM | 2009-12-18 |  | RDGIBSON |  | Concentration | mg/L | Y | (fl. conc.) |  | Bromoform (fl. conc.) | 75-25-2 | Volatile Organic Compounds (VOCs); | Volatile Organic Compounds (VOCs); EMS code: B013 | Chemical |  | LAB
B013 | ug/g | Bromoform | MS02 | Purge&Trap GC/MS (VOC-S) | Purge&Trap GC/MS (VOC-S) | 0.002 | Y | Y | Y | 27 | 11/19/2004 | 10/28/2022 5:04:00 PM | 1/11/2024 1:07:07 PM | 1996-04-12 | 6/3/1998 6:51:34 PM | EMS |  | MassPerMass | mg/kg | Y | (mass conc.) |  | Bromoform (mass conc.) | 75-25-2 | Volatile Organic Compounds (VOCs); | Volatile Organic Compounds (VOCs); EMS code: B013 | Chemical |  | LAB
B013 | mg/L | Bromoform | MS01 | Purge&Trap GC/MS (VOC-W) | Purge&Trap GC/MS (VOC-W) | 0.3 | Y | Y | Y | 546 | 4/15/2002 | 6/4/2024 10:24:00 AM | 6/13/2024 5:01:27 PM | 1996-04-12 | 3/18/2003 2:49:40 PM | EMS | UMULLER | Concentration | mg/L | Y | (fl. conc.) |  | Bromoform (fl. conc.) | 75-25-2 | Volatile Organic Compounds (VOCs); | Volatile Organic Compounds (VOCs); EMS code: B013 | Chemical |  | LAB
B013 | mg/L | Bromoform | E611 | E611S: EPA 8260D (mod) | E611S: Dimethyl selenide and dimethyl diselenide are analyzed by static headspace GC-MS. Samples are prepared in headspace vials and are heated and agitated on the headspace autosampler; causing VOCs to partition between the aqueous phase and the headspace in accordance with Henry`s law. | 0.001 | Y | Y | Y | 1 | 12/14/2023 | 12/14/2023 | 1/16/2024 8:06:04 AM | 2024-01-15 | 1/15/2024 2:17:10 PM | EMS | EMS | Concentration | mg/L | Y | (fl. conc.) |  | Bromoform (fl. conc.) | 75-25-2 | Volatile Organic Compounds (VOCs); | Volatile Organic Compounds (VOCs); EMS code: B013 | Chemical |  | LAB
B013 | mg/L | Bromoform | X384 | Purge and trap GC/MS | The sample is analyzed by purge and trap gas chromatography with detection by mass spectrometry. The procedure involves purging the volatile compounds from the sample with an inert gas, and trapping them on a solid sorbent. When purging of the sample is complete, the trap is heated and the compounds are transferred to the gas chromatographic column. Analysis is then accomplished by separation of the components by gas chromatography with detection by mass spectrometry. | 0.001 | Y | Y | Y | 112 | 8/5/2004 10:31:00 AM | 7/26/2012 7:12:00 AM | 9/26/2012 1:07:08 PM | 1996-04-12 | 6/3/1998 6:51:34 PM | EMS |  | Concentration | mg/L | Y | (fl. conc.) |  | Bromoform (fl. conc.) | 75-25-2 | Volatile Organic Compounds (VOCs); | Volatile Organic Compounds (VOCs); EMS code: B013 | Chemical |  | LAB
B013 | ug/g | Bromoform | FS04 | Static Headspace:HRGC/LRMS | Static Headspace anaylsis is a method for the determination of volatile compounds in liquids and/or solids by analysing the vapour phase of a sample in a closed system; in this case a sealed vial. Static Headspace is sometimes referred to as Equilibrium Headspace because the vapour phase above the sample in a sealed vial is brought into thermodynamic equilibrium with the sample prior to analysis. Volatile organic compounds are thermally extracted from the sample matrix by Static Headspace sampling and introduced directly to a thick-film megabore GC column. Compounds eluting from the GC column are introduced into a Mass Spectrometer and target compounds are identified by a comparison of both retention times and mass spectra with those of reference standards stored in a computerized data system. Reference retention times and mass spectra for analytes are obtained by the measurement of calibration standards under precisely the same extraction conditions used for samples. The concentration of each identified component is measured by relating the response of the sample quantitation mass ions to the response of the quantitation mass ions of a series of external standards. A deuterated surrogate recovery standard is added to provide information on matrix effects. | 0.1 | Y | Y | Y | 44 | 6/30/1999 5:05:00 PM | 5/22/2015 | 7/3/2015 3:19:03 PM | 1996-09-16 | 6/3/1998 6:51:44 PM | MPAK |  | MassPerMass | mg/kg | Y | (mass conc.) |  | Bromoform (mass conc.) | 75-25-2 | Volatile Organic Compounds (VOCs); | Volatile Organic Compounds (VOCs); EMS code: B013 | Chemical |  | LAB
B014 | ug/g | Bromacil | ZN03 | DCM+Ac.Extr:GC/MS | DCM+Ac.Extr:GC/MS EPA 8270E (mod) | 0.01 | Y | Y | Y | 17 | 1/14/2022 | 11/2/2022 4:41:00 AM | 1/11/2024 1:07:07 PM | 1999-11-16 |  | NPEPPIN |  | MassPerMass | mg/kg | Y | (mass conc.) |  | Bromacil (mass conc.) | 314-40-9 | Pesticides; | Pesticides; EMS code: B014 | Chemical |  | LAB
B014 | mg/L | Bromacil | ZN01 | DCM+Ac.Extr:GC/MS | DCM+Ac.Extr:GC/MS | 0.0003 | Y | Y | Y | 27 | 11/1/2008 | 3/29/2012 10:15:00 PM | 5/3/2012 9:27:49 AM | 1996-04-12 | 11/16/1999 4:06:27 PM | EMS | NPEPPIN | Concentration | mg/L | Y | (fl. conc.) |  | Bromacil (fl. conc.) | 314-40-9 | Pesticides; | Pesticides; EMS code: B014 | Chemical |  | LAB
B014 | mg/L | Bromacil | P015 | Phenoxyacid Herbicides by LC-MS/MS | Phenoxyacid Herbicides by LC-MS/MS | 0.0001 | Y | Y | Y | 68 | 12/6/2021 | 9/9/2022 | 12/8/2022 12:57:05 PM | 2021-12-14 |  | EMS |  | Concentration | mg/L | Y | (fl. conc.) |  | Bromacil (fl. conc.) | 314-40-9 | Pesticides; | Pesticides; EMS code: B014 | Chemical |  | LAB
B015 | mg/L | Benfluralin | ZN01 | DCM+Ac.Extr:GC/MS | DCM+Ac.Extr:GC/MS | 0.0001 | Y | Y | Y | 28 | 7/28/2005 7:03:00 AM | 3/29/2012 10:15:00 PM | 5/3/2012 9:27:49 AM | 1996-04-12 | 11/16/1999 4:06:27 PM | EMS | NPEPPIN | Concentration | mg/L | Y | (fl. conc.) |  | Benfluralin (fl. conc.) | 1861-40-1 | Pesticides; | Pesticides; EMS code: B015 | Chemical |  | LAB
B017 | mg/L | BHC; alpha | ZN01 | DCM+Ac.Extr:GC/MS | DCM+Ac.Extr:GC/MS | 0.0003 | Y | Y | Y | 52 | 10/3/2002 | 3/29/2012 10:15:00 PM | 5/3/2012 9:27:49 AM | 1996-04-12 | 11/16/1999 4:06:27 PM | EMS | NPEPPIN | Concentration | mg/L | Y | (fl. conc.) |  | BHC; alpha (fl. conc.) | 319-84-6 | Pesticides; | Pesticides; EMS code: B017 | Chemical |  | LAB
B017 | ug/g | BHC; alpha | ZN03 | DCM+Ac.Extr:GC/MS | DCM+Ac.Extr:GC/MS EPA 8270E (mod) | 0.03 | Y | Y | Y | 28 | 4/20/2002 10:10:00 AM | 11/2/2022 4:41:00 AM | 1/7/2023 3:12:10 PM | 1996-04-12 | 11/16/1999 2:27:40 PM | EMS | NPEPPIN | MassPerMass | mg/kg | Y | (mass conc.) |  | BHC; alpha (mass conc.) | 319-84-6 | Pesticides; | Pesticides; EMS code: B017 | Chemical |  | LAB
B017 | mg/L | BHC; alpha | F027 | Hexane Extr,GC/ECD | Effluent samples are solvent extracted with Hexane and the extract dried over Sodium Sulfate, concentrated and cleaned up, as required, using Florisil, activated Copper metal, concentrated Sulfuric Acid and Mercury, to remove interferences. The extracts are then analysed using megabore column Gas Chromatography with Electron Capture detection. | 5e-05 | Y | Y | Y | 20 | 3/3/2000 7:36:00 AM | 1/22/2002 10:58:00 AM | 6/18/2003 2:53:29 PM | 2002-06-05 |  | UMULLER |  | Concentration | mg/L | Y | (fl. conc.) |  | BHC; alpha (fl. conc.) | 319-84-6 | Pesticides; | Pesticides; EMS code: B017 | Chemical |  | LAB
B018 | mg/L | BHC; beta | ZN01 | DCM+Ac.Extr:GC/MS | DCM+Ac.Extr:GC/MS | 0.0003 | Y | Y | Y | 52 | 10/3/2002 | 3/29/2012 10:15:00 PM | 5/3/2012 9:27:49 AM | 1996-04-12 | 11/16/1999 4:06:27 PM | EMS | NPEPPIN | Concentration | mg/L | Y | (fl. conc.) |  | BHC; beta (fl. conc.) | 319-85-7 | Pesticides; | Pesticides; EMS code: B018 | Chemical |  | LAB
B018 | mg/L | BHC; beta | F027 | Hexane Extr,GC/ECD | Effluent samples are solvent extracted with Hexane and the extract dried over Sodium Sulfate, concentrated and cleaned up, as required, using Florisil, activated Copper metal, concentrated Sulfuric Acid and Mercury, to remove interferences. The extracts are then analysed using megabore column Gas Chromatography with Electron Capture detection. | 5e-05 | Y | Y | Y | 20 | 3/3/2000 7:36:00 AM | 1/22/2002 10:58:00 AM | 6/18/2003 2:53:29 PM | 2002-06-05 |  | UMULLER |  | Concentration | mg/L | Y | (fl. conc.) |  | BHC; beta (fl. conc.) | 319-85-7 | Pesticides; | Pesticides; EMS code: B018 | Chemical |  | LAB
B018 | ug/g | BHC; beta | ZN03 | DCM+Ac.Extr:GC/MS | DCM+Ac.Extr:GC/MS EPA 8270E (mod) | 0.03 | Y | Y | Y | 28 | 4/20/2002 10:10:00 AM | 11/2/2022 4:41:00 AM | 1/7/2023 3:12:10 PM | 1996-04-12 | 11/16/1999 2:27:40 PM | EMS | NPEPPIN | MassPerMass | mg/kg | Y | (mass conc.) |  | BHC; beta (mass conc.) | 319-85-7 | Pesticides; | Pesticides; EMS code: B018 | Chemical |  | LAB
B019 | mg/L | BHC; delta | ZN01 | DCM+Ac.Extr:GC/MS | DCM+Ac.Extr:GC/MS | 0.0005 | Y | Y | Y | 1 | 7/28/2005 7:03:00 AM | 7/28/2005 7:03:00 AM | 6/2/2009 12:00:57 PM | 2003-06-16 |  | UMULLER |  | Concentration | mg/L | Y | (fl. conc.) |  | BHC; delta (fl. conc.) | 319-86-8 | Pesticides; | Pesticides; EMS code: B019 | Chemical |  | LAB
B019 | ug/g | BHC; delta | ZN03 | DCM+Ac.Extr:GC/MS | DCM+Ac.Extr:GC/MS EPA 8270E (mod) | 0.05 | Y | Y | Y | 2 | 11/2/2022 12:53:00 AM | 11/2/2022 4:41:00 AM | 1/7/2023 3:12:10 PM | 2003-06-16 |  | UMULLER |  | MassPerMass | mg/kg | Y | (mass conc.) |  | BHC; delta (mass conc.) | 319-86-8 | Pesticides; | Pesticides; EMS code: B019 | Chemical |  | LAB
B020 | mg/L | Benzene | MS04 | Static Headspace GC/MS (Water) | (Water) A static headspace procedure is used as a means of driving the analytes out of the matrix into a headspace. Once equilibrium is achieved, a fixed volume of headspeace is taken for subsequent determination by Gas Chromatography - Mass Spectroscopy (GC/MS) using Selected Ion Monitoring. | 0.0005 | Y | Y | Y | 129 | 8/13/2009 | 6/28/2024 12:29:00 AM | 7/30/2024 6:32:12 AM | 2009-12-23 |  | RDGIBSON |  | Concentration | mg/L | Y | (fl. conc.) |  | Benzene (fl. conc.) | 71-43-2 | Petroleum Hydrocarbons; | Petroleum Hydrocarbons; EMS code: B020 | Chemical |  | LAB
B020 | ug/g | Benzene | MS03 | Static Headspace GC/MS (Soil) | (Soil) A static headspace procedure is used as a means of driving the analytes out of the matrix into a headspace. Once equilibrium is achieved, a fixed volume of headspeace is taken for subsequent determination by Gas Chromatography - Mass Spectroscopy (GC/MS) using Selected Ion Monitoring. | 0.005 | Y | Y | Y | 228 |  |  | 4/30/2021 11:32:40 AM | 2009-12-18 |  | RDGIBSON |  | MassPerMass | mg/kg | Y | (mass conc.) |  | Benzene (mass conc.) | 71-43-2 | Petroleum Hydrocarbons; | Petroleum Hydrocarbons; EMS code: B020 | Chemical |  | LAB
B020 | mg/L | Benzene | 8260D | EPA 8260D (mod) | EPA 8260D (mod) VOLATILE ORGANIC COMPOUNDS BY GAS CHROMATOGRAPHY/MASS SPECTROMETRY This method is used to determine volatile organic compounds (VOCs) in a variety of solid waste matrices. | 0.005 | Y | Y | Y | 0 |  |  |  | 2024-01-24 |  | RDGIBSON |  | Concentration | mg/L | Y | (fl. conc.) |  | Benzene (fl. conc.) | 71-43-2 | Petroleum Hydrocarbons; | Petroleum Hydrocarbons; EMS code: B020 | Chemical |  | LAB
B020 | ug/g | Benzene | X263 | Purge and trap GC/MS | Purge and trap GC/MS | 0.001 | Y | Y | Y | 20 | 3/29/2006 | 4/4/2006 | 4/27/2006 1:02:35 AM | 1996-04-12 | 6/3/1998 6:51:41 PM | EMS |  | MassPerMass | mg/kg | Y | (mass conc.) |  | Benzene (mass conc.) | 71-43-2 | Petroleum Hydrocarbons; | Petroleum Hydrocarbons; EMS code: B020 | Chemical |  | LAB
B020 | ug/g | Benzene | FS04 | Static Headspace:HRGC/LRMS | Static Headspace anaylsis is a method for the determination of volatile compounds in liquids and/or solids by analysing the vapour phase of a sample in a closed system; in this case a sealed vial. Static Headspace is sometimes referred to as Equilibrium Headspace because the vapour phase above the sample in a sealed vial is brought into thermodynamic equilibrium with the sample prior to analysis. Volatile organic compounds are thermally extracted from the sample matrix by Static Headspace sampling and introduced directly to a thick-film megabore GC column. Compounds eluting from the GC column are introduced into a Mass Spectrometer and target compounds are identified by a comparison of both retention times and mass spectra with those of reference standards stored in a computerized data system. Reference retention times and mass spectra for analytes are obtained by the measurement of calibration standards under precisely the same extraction conditions used for samples. The concentration of each identified component is measured by relating the response of the sample quantitation mass ions to the response of the quantitation mass ions of a series of external standards. A deuterated surrogate recovery standard is added to provide information on matrix effects. | 0.01 | Y | Y | Y | 89 | 7/2/1999 2:39:00 PM | 3/7/2013 3:50:00 AM | 3/19/2013 1:15:54 PM | 1996-09-17 | 6/3/1998 6:51:45 PM | MPAK |  | MassPerMass | mg/kg | Y | (mass conc.) |  | Benzene (mass conc.) | 71-43-2 | Petroleum Hydrocarbons; | Petroleum Hydrocarbons; EMS code: B020 | Chemical |  | LAB
B020 | mg/L | Benzene | X384 | Purge and trap GC/MS | The sample is analyzed by purge and trap gas chromatography with detection by mass spectrometry. The procedure involves purging the volatile compounds from the sample with an inert gas, and trapping them on a solid sorbent. When purging of the sample is complete, the trap is heated and the compounds are transferred to the gas chromatographic column. Analysis is then accomplished by separation of the components by gas chromatography with detection by mass spectrometry. | 0.0005 | Y | Y | Y | 134 | 4/27/2000 | 2/17/2006 | 3/31/2006 1:00:58 AM | 1996-04-12 | 6/3/1998 6:51:34 PM | EMS |  | Concentration | mg/L | Y | (fl. conc.) |  | Benzene (fl. conc.) | 71-43-2 | Petroleum Hydrocarbons; | Petroleum Hydrocarbons; EMS code: B020 | Chemical |  | LAB
B020 | ug/g | Benzene | MS02 | Purge&Trap GC/MS (VOC-S) | Purge&Trap GC/MS (VOC-S) | 0.001 | Y | Y | Y | 117 | 11/19/2004 | 10/28/2022 5:04:00 PM | 7/19/2024 8:20:35 AM | 1996-04-12 | 6/3/1998 6:51:34 PM | EMS |  | MassPerMass | mg/kg | Y | (mass conc.) |  | Benzene (mass conc.) | 71-43-2 | Petroleum Hydrocarbons; | Petroleum Hydrocarbons; EMS code: B020 | Chemical |  | LAB
B020 | mg/L | Benzene | F007 | Static Headspace:HRGC/LRMS | Static headspace analysis is a method for the determination of volatile compounds in liquids and/or solids by analysing the vapour phase of a sample in a closed system; in this case a sealed vial. Static Headspace is sometimes referred to as Equilibrium Headspace because the vapour phase above the sample in a sealed vial is brought into thermodynamic equilibrium with the sample prior to analysis. Volatile organic compounds are thermally extracted from the sample matrix by Static Headspace sampling and introduced directly to a thick-film megabore GC column. Compounds eluting from the GC column are introduced into a Mass Spectrometer and target compounds are identified by a comparison of both retention times and mass spectra with those of reference standards stored in a computerized data system. Reference retention times and mass spectra for analytes are obtained by the measurement of calibration standards under precisely the same extraction conditions used for samples. The concentration of each identified component is measured by relating the response of the sample quantitation mass ions to the response of the quantitation mass ions of a series of external standards. A deuterated surrogate recovery standard is added to provide information on matrix effects. | 0.0005 | Y | Y | Y | 1408 | 4/21/1998 2:34:00 PM | 8/14/2024 | 8/16/2024 4:05:21 PM | 1996-09-17 | 6/3/1998 6:51:45 PM | MPAK |  | Concentration | mg/L | Y | (fl. conc.) |  | Benzene (fl. conc.) | 71-43-2 | Petroleum Hydrocarbons; | Petroleum Hydrocarbons; EMS code: B020 | Chemical |  | LAB
B020 | mg/L | Benzene | MS01 | Purge&Trap GC/MS (VOC-W) | Purge&Trap GC/MS (VOC-W) | 0.2 | Y | Y | Y | 1107 | 4/15/2002 | 6/4/2024 10:24:00 AM | 6/13/2024 5:01:27 PM | 1996-04-12 | 3/18/2003 2:49:40 PM | EMS | UMULLER | Concentration | mg/L | Y | (fl. conc.) |  | Benzene (fl. conc.) | 71-43-2 | Petroleum Hydrocarbons; | Petroleum Hydrocarbons; EMS code: B020 | Chemical |  | LAB
B020 | ug/g | Benzene | MEOH | MeOH extrac,GC/FID,nC6-nC10 | MeOH extrac;GC/FID;nC6-nC10. EPA8260B; 5021; 5035; BC MOE | 0.005 | Y | Y | Y | 49 | 5/21/2015 | 8/11/2019 | 8/13/2019 10:56:07 AM | 2015-05-26 |  | RDGIBSON |  | MassPerMass | mg/kg | Y | (mass conc.) |  | Benzene (mass conc.) | 71-43-2 | Petroleum Hydrocarbons; | Petroleum Hydrocarbons; EMS code: B020 | Chemical |  | LAB
B021 | mg/L | Ethylbenzene | 8260D | EPA 8260D (mod) | EPA 8260D (mod) VOLATILE ORGANIC COMPOUNDS BY GAS CHROMATOGRAPHY/MASS SPECTROMETRY This method is used to determine volatile organic compounds (VOCs) in a variety of solid waste matrices. | 0.005 | Y | Y | Y | 0 |  |  |  | 2024-01-25 |  | RDGIBSON |  | Concentration | mg/L | Y | (fl. conc.) |  | Ethylbenzene (fl. conc.) | 100-41-4 | Petroleum Hydrocarbons; | Petroleum Hydrocarbons; EMS code: B021 | Chemical |  | LAB
B021 | ug/g | Ethylbenzene | MEOH | MeOH extrac,GC/FID,nC6-nC10 | MeOH extrac;GC/FID;nC6-nC10. EPA8260B; 5021; 5035; BC MOE | 0.015 | Y | Y | Y | 49 | 5/21/2015 | 8/11/2019 | 8/13/2019 10:56:07 AM | 2015-05-26 |  | RDGIBSON |  | MassPerMass | mg/kg | Y | (mass conc.) |  | Ethylbenzene (mass conc.) | 100-41-4 | Petroleum Hydrocarbons; | Petroleum Hydrocarbons; EMS code: B021 | Chemical |  | LAB
B021 | mg/L | Ethylbenzene | F007 | Static Headspace:HRGC/LRMS | Static headspace analysis is a method for the determination of volatile compounds in liquids and/or solids by analysing the vapour phase of a sample in a closed system; in this case a sealed vial. Static Headspace is sometimes referred to as Equilibrium Headspace because the vapour phase above the sample in a sealed vial is brought into thermodynamic equilibrium with the sample prior to analysis. Volatile organic compounds are thermally extracted from the sample matrix by Static Headspace sampling and introduced directly to a thick-film megabore GC column. Compounds eluting from the GC column are introduced into a Mass Spectrometer and target compounds are identified by a comparison of both retention times and mass spectra with those of reference standards stored in a computerized data system. Reference retention times and mass spectra for analytes are obtained by the measurement of calibration standards under precisely the same extraction conditions used for samples. The concentration of each identified component is measured by relating the response of the sample quantitation mass ions to the response of the quantitation mass ions of a series of external standards. A deuterated surrogate recovery standard is added to provide information on matrix effects. | 0.0005 | Y | Y | Y | 1409 | 4/21/1998 2:34:00 PM | 8/14/2024 | 8/16/2024 4:05:21 PM | 1996-09-17 | 6/3/1998 6:51:45 PM | MPAK |  | Concentration | mg/L | Y | (fl. conc.) |  | Ethylbenzene (fl. conc.) | 100-41-4 | Petroleum Hydrocarbons; | Petroleum Hydrocarbons; EMS code: B021 | Chemical |  | LAB
B021 | mg/L | Ethylbenzene | X384 | Purge and trap GC/MS | The sample is analyzed by purge and trap gas chromatography with detection by mass spectrometry. The procedure involves purging the volatile compounds from the sample with an inert gas, and trapping them on a solid sorbent. When purging of the sample is complete, the trap is heated and the compounds are transferred to the gas chromatographic column. Analysis is then accomplished by separation of the components by gas chromatography with detection by mass spectrometry. | 0.0005 | Y | Y | Y | 125 | 1/24/2001 | 2/17/2006 | 3/31/2006 1:00:58 AM | 1996-04-12 | 6/3/1998 6:51:34 PM | EMS |  | Concentration | mg/L | Y | (fl. conc.) |  | Ethylbenzene (fl. conc.) | 100-41-4 | Petroleum Hydrocarbons; | Petroleum Hydrocarbons; EMS code: B021 | Chemical |  | LAB
B021 | mg/L | Ethylbenzene | MS04 | Static Headspace GC/MS (Water) | (Water) A static headspace procedure is used as a means of driving the analytes out of the matrix into a headspace. Once equilibrium is achieved, a fixed volume of headspeace is taken for subsequent determination by Gas Chromatography - Mass Spectroscopy (GC/MS) using Selected Ion Monitoring. | 0.0005 | Y | Y | Y | 108 | 8/13/2009 | 6/28/2024 12:29:00 AM | 7/30/2024 6:32:12 AM | 2009-12-23 |  | RDGIBSON |  | Concentration | mg/L | Y | (fl. conc.) |  | Ethylbenzene (fl. conc.) | 100-41-4 | Petroleum Hydrocarbons; | Petroleum Hydrocarbons; EMS code: B021 | Chemical |  | LAB
B021 | mg/L | Ethylbenzene | MS01 | Purge&Trap GC/MS (VOC-W) | Purge&Trap GC/MS (VOC-W) | 0.4 | Y | Y | Y | 1107 | 4/15/2002 | 6/4/2024 10:24:00 AM | 6/13/2024 5:01:27 PM | 1996-04-12 | 5/20/2003 9:29:14 AM | EMS | UMULLER | Concentration | mg/L | Y | (fl. conc.) |  | Ethylbenzene (fl. conc.) | 100-41-4 | Petroleum Hydrocarbons; | Petroleum Hydrocarbons; EMS code: B021 | Chemical |  | LAB
B021 | ug/g | Ethylbenzene | X263 | Purge and trap GC/MS | Purge and trap GC/MS | 0.001 | Y | Y | Y | 20 | 3/29/2006 | 4/4/2006 | 4/27/2006 1:02:35 AM | 1996-04-12 | 6/3/1998 6:51:41 PM | EMS |  | MassPerMass | mg/kg | Y | (mass conc.) |  | Ethylbenzene (mass conc.) | 100-41-4 | Petroleum Hydrocarbons; | Petroleum Hydrocarbons; EMS code: B021 | Chemical |  | LAB
B021 | ug/g | Ethylbenzene | FS04 | Static Headspace:HRGC/LRMS | Static Headspace anaylsis is a method for the determination of volatile compounds in liquids and/or solids by analysing the vapour phase of a sample in a closed system; in this case a sealed vial. Static Headspace is sometimes referred to as Equilibrium Headspace because the vapour phase above the sample in a sealed vial is brought into thermodynamic equilibrium with the sample prior to analysis. Volatile organic compounds are thermally extracted from the sample matrix by Static Headspace sampling and introduced directly to a thick-film megabore GC column. Compounds eluting from the GC column are introduced into a Mass Spectrometer and target compounds are identified by a comparison of both retention times and mass spectra with those of reference standards stored in a computerized data system. Reference retention times and mass spectra for analytes are obtained by the measurement of calibration standards under precisely the same extraction conditions used for samples. The concentration of each identified component is measured by relating the response of the sample quantitation mass ions to the response of the quantitation mass ions of a series of external standards. A deuterated surrogate recovery standard is added to provide information on matrix effects. | 0.01 | Y | Y | Y | 89 | 7/2/1999 2:40:00 PM | 3/7/2013 3:50:00 AM | 3/19/2013 1:15:54 PM | 1996-09-17 | 6/3/1998 6:51:45 PM | MPAK |  | MassPerMass | mg/kg | Y | (mass conc.) |  | Ethylbenzene (mass conc.) | 100-41-4 | Petroleum Hydrocarbons; | Petroleum Hydrocarbons; EMS code: B021 | Chemical |  | LAB
B021 | ug/g | Ethylbenzene | MS02 | Purge&Trap GC/MS (VOC-S) | Purge&Trap GC/MS (VOC-S) | 0.002 | Y | Y | Y | 117 | 11/19/2004 | 10/28/2022 5:04:00 PM | 7/19/2024 8:20:35 AM | 1996-04-12 | 6/3/1998 6:51:34 PM | EMS |  | MassPerMass | mg/kg | Y | (mass conc.) |  | Ethylbenzene (mass conc.) | 100-41-4 | Petroleum Hydrocarbons; | Petroleum Hydrocarbons; EMS code: B021 | Chemical |  | LAB
B021 | ug/g | Ethylbenzene | MS03 | Static Headspace GC/MS (Soil) | (Soil) A static headspace procedure is used as a means of driving the analytes out of the matrix into a headspace. Once equilibrium is achieved, a fixed volume of headspeace is taken for subsequent determination by Gas Chromatography - Mass Spectroscopy (GC/MS) using Selected Ion Monitoring. | 0.01 | Y | Y | Y | 228 |  |  | 4/30/2021 11:32:40 AM | 2009-12-18 |  | RDGIBSON |  | MassPerMass | mg/kg | Y | (mass conc.) |  | Ethylbenzene (mass conc.) | 100-41-4 | Petroleum Hydrocarbons; | Petroleum Hydrocarbons; EMS code: B021 | Chemical |  | LAB
B022 | ug/g | 1;2 Dichlorobenzene | MS02 | Purge&Trap GC/MS (VOC-S) | Purge&Trap GC/MS (VOC-S) | 0.003 | Y | Y | Y | 26 | 11/19/2004 | 2/4/2022 | 1/11/2024 1:07:07 PM | 1996-04-12 | 6/3/1998 6:51:45 PM | EMS |  | MassPerMass | mg/kg | Y | (mass conc.) |  | 1;2 Dichlorobenzene (mass conc.) | 95-50-1 | Volatile Organic Compounds (VOCs); | Volatile Organic Compounds (VOCs); EMS code: B022 | Chemical |  | LAB
B022 | mg/L | 1;2 Dichlorobenzene | F007 | Static Headspace:HRGC/LRMS | Static headspace analysis is a method for the determination of volatile compounds in liquids and/or solids by analysing the vapour phase of a sample in a closed system; in this case a sealed vial. Static Headspace is sometimes referred to as Equilibrium Headspace because the vapour phase above the sample in a sealed vial is brought into thermodynamic equilibrium with the sample prior to analysis. Volatile organic compounds are thermally extracted from the sample matrix by Static Headspace sampling and introduced directly to a thick-film megabore GC column. Compounds eluting from the GC column are introduced into a Mass Spectrometer and target compounds are identified by a comparison of both retention times and mass spectra with those of reference standards stored in a computerized data system. Reference retention times and mass spectra for analytes are obtained by the measurement of calibration standards under precisely the same extraction conditions used for samples. The concentration of each identified component is measured by relating the response of the sample quantitation mass ions to the response of the quantitation mass ions of a series of external standards. A deuterated surrogate recovery standard is added to provide information on matrix effects. | 0.001 | Y | Y | Y | 329 | 4/15/2011 | 8/14/2024 | 8/16/2024 4:05:21 PM | 2012-07-09 |  | RDGIBSON |  | Concentration | mg/L | Y | (fl. conc.) |  | 1;2 Dichlorobenzene (fl. conc.) | 95-50-1 | Volatile Organic Compounds (VOCs); | Volatile Organic Compounds (VOCs); EMS code: B022 | Chemical |  | LAB
B022 | mg/L | 1;2 Dichlorobenzene | MS04 | Static Headspace GC/MS (Water) | (Water) A static headspace procedure is used as a means of driving the analytes out of the matrix into a headspace. Once equilibrium is achieved, a fixed volume of headspeace is taken for subsequent determination by Gas Chromatography - Mass Spectroscopy (GC/MS) using Selected Ion Monitoring. | 0.0005 | Y | Y | Y | 66 | 8/13/2009 | 6/28/2024 12:29:00 AM | 7/9/2024 3:31:02 PM | 2009-12-18 |  | RDGIBSON |  | Concentration | mg/L | Y | (fl. conc.) |  | 1;2 Dichlorobenzene (fl. conc.) | 95-50-1 | Volatile Organic Compounds (VOCs); | Volatile Organic Compounds (VOCs); EMS code: B022 | Chemical |  | LAB
B022 | ug/g | 1;2 Dichlorobenzene | FS04 | Static Headspace:HRGC/LRMS | Static Headspace anaylsis is a method for the determination of volatile compounds in liquids and/or solids by analysing the vapour phase of a sample in a closed system; in this case a sealed vial. Static Headspace is sometimes referred to as Equilibrium Headspace because the vapour phase above the sample in a sealed vial is brought into thermodynamic equilibrium with the sample prior to analysis. Volatile organic compounds are thermally extracted from the sample matrix by Static Headspace sampling and introduced directly to a thick-film megabore GC column. Compounds eluting from the GC column are introduced into a Mass Spectrometer and target compounds are identified by a comparison of both retention times and mass spectra with those of reference standards stored in a computerized data system. Reference retention times and mass spectra for analytes are obtained by the measurement of calibration standards under precisely the same extraction conditions used for samples. The concentration of each identified component is measured by relating the response of the sample quantitation mass ions to the response of the quantitation mass ions of a series of external standards. A deuterated surrogate recovery standard is added to provide information on matrix effects. | 0.05 | Y | Y | Y | 43 | 5/21/2015 | 5/22/2015 | 7/3/2015 3:19:03 PM | 2015-07-03 |  | RDGIBSON |  | MassPerMass | mg/kg | Y | (mass conc.) |  | 1;2 Dichlorobenzene (mass conc.) | 95-50-1 | Volatile Organic Compounds (VOCs); | Volatile Organic Compounds (VOCs); EMS code: B022 | Chemical |  | LAB
B022 | mg/L | 1;2 Dichlorobenzene | X384 | Purge and trap GC/MS | The sample is analyzed by purge and trap gas chromatography with detection by mass spectrometry. The procedure involves purging the volatile compounds from the sample with an inert gas, and trapping them on a solid sorbent. When purging of the sample is complete, the trap is heated and the compounds are transferred to the gas chromatographic column. Analysis is then accomplished by separation of the components by gas chromatography with detection by mass spectrometry. | 0.001 | Y | Y | Y | 1 |  |  | 10/23/1996 11:03:45 PM | 1996-04-12 | 6/3/1998 6:51:34 PM | EMS |  | Concentration | mg/L | Y | (fl. conc.) |  | 1;2 Dichlorobenzene (fl. conc.) | 95-50-1 | Volatile Organic Compounds (VOCs); | Volatile Organic Compounds (VOCs); EMS code: B022 | Chemical |  | LAB
B022 | mg/L | 1;2 Dichlorobenzene | MS01 | Purge&Trap GC/MS (VOC-W) | Purge&Trap GC/MS (VOC-W) | 0.6 | Y | Y | Y | 193 | 4/15/2002 | 4/1/2011 | 7/9/2012 2:09:46 PM | 1996-04-12 | 3/18/2003 2:49:40 PM | EMS | UMULLER | Concentration | mg/L | Y | (fl. conc.) |  | 1;2 Dichlorobenzene (fl. conc.) | 95-50-1 | Volatile Organic Compounds (VOCs); | Volatile Organic Compounds (VOCs); EMS code: B022 | Chemical |  | LAB
B025 | ug/g | 1;4-Dichlorobenzene | MS02 | Purge&Trap GC/MS (VOC-S) | Purge&Trap GC/MS (VOC-S) | 0.003 | Y | Y | Y | 27 | 11/19/2004 | 10/28/2022 5:04:00 PM | 1/11/2024 1:07:07 PM | 1996-04-12 | 6/3/1998 6:51:34 PM | EMS |  | MassPerMass | mg/kg | Y | (mass conc.) |  | 1;4-Dichlorobenzene (mass conc.) | 106-46-7 | Volatile Organic Compounds (VOCs); | Volatile Organic Compounds (VOCs); EMS code: B025 | Chemical |  | LAB
B025 | mg/L | 1;4-Dichlorobenzene | MS01 | Purge&Trap GC/MS (VOC-W) | Purge&Trap GC/MS (VOC-W) | 0.6 | Y | Y | Y | 551 | 4/15/2002 | 6/4/2024 10:24:00 AM | 6/13/2024 5:01:27 PM | 1996-04-12 | 3/18/2003 2:49:40 PM | EMS | UMULLER | Concentration | mg/L | Y | (fl. conc.) |  | 1;4-Dichlorobenzene (fl. conc.) | 106-46-7 | Volatile Organic Compounds (VOCs); | Volatile Organic Compounds (VOCs); EMS code: B025 | Chemical |  | LAB
B025 | mg/L | 1;4-Dichlorobenzene | X384 | Purge and trap GC/MS | The sample is analyzed by purge and trap gas chromatography with detection by mass spectrometry. The procedure involves purging the volatile compounds from the sample with an inert gas, and trapping them on a solid sorbent. When purging of the sample is complete, the trap is heated and the compounds are transferred to the gas chromatographic column. Analysis is then accomplished by separation of the components by gas chromatography with detection by mass spectrometry. | 0.001 | Y | Y | Y | 120 | 4/27/2000 | 7/26/2012 7:12:00 AM | 9/26/2012 1:07:08 PM | 1996-04-12 | 6/3/1998 6:51:45 PM | EMS |  | Concentration | mg/L | Y | (fl. conc.) |  | 1;4-Dichlorobenzene (fl. conc.) | 106-46-7 | Volatile Organic Compounds (VOCs); | Volatile Organic Compounds (VOCs); EMS code: B025 | Chemical |  | LAB
B025 | ug/g | 1;4-Dichlorobenzene | FS04 | Static Headspace:HRGC/LRMS | Static Headspace anaylsis is a method for the determination of volatile compounds in liquids and/or solids by analysing the vapour phase of a sample in a closed system; in this case a sealed vial. Static Headspace is sometimes referred to as Equilibrium Headspace because the vapour phase above the sample in a sealed vial is brought into thermodynamic equilibrium with the sample prior to analysis. Volatile organic compounds are thermally extracted from the sample matrix by Static Headspace sampling and introduced directly to a thick-film megabore GC column. Compounds eluting from the GC column are introduced into a Mass Spectrometer and target compounds are identified by a comparison of both retention times and mass spectra with those of reference standards stored in a computerized data system. Reference retention times and mass spectra for analytes are obtained by the measurement of calibration standards under precisely the same extraction conditions used for samples. The concentration of each identified component is measured by relating the response of the sample quantitation mass ions to the response of the quantitation mass ions of a series of external standards. A deuterated surrogate recovery standard is added to provide information on matrix effects. | 0.05 | Y | Y | Y | 43 | 5/21/2015 | 5/22/2015 | 7/3/2015 3:19:03 PM | 2015-07-03 |  | RDGIBSON |  | MassPerMass | mg/kg | Y | (mass conc.) |  | 1;4-Dichlorobenzene (mass conc.) | 106-46-7 | Volatile Organic Compounds (VOCs); | Volatile Organic Compounds (VOCs); EMS code: B025 | Chemical |  | LAB
B025 | mg/L | 1;4-Dichlorobenzene | F007 | Static Headspace:HRGC/LRMS | Static headspace analysis is a method for the determination of volatile compounds in liquids and/or solids by analysing the vapour phase of a sample in a closed system; in this case a sealed vial. Static Headspace is sometimes referred to as Equilibrium Headspace because the vapour phase above the sample in a sealed vial is brought into thermodynamic equilibrium with the sample prior to analysis. Volatile organic compounds are thermally extracted from the sample matrix by Static Headspace sampling and introduced directly to a thick-film megabore GC column. Compounds eluting from the GC column are introduced into a Mass Spectrometer and target compounds are identified by a comparison of both retention times and mass spectra with those of reference standards stored in a computerized data system. Reference retention times and mass spectra for analytes are obtained by the measurement of calibration standards under precisely the same extraction conditions used for samples. The concentration of each identified component is measured by relating the response of the sample quantitation mass ions to the response of the quantitation mass ions of a series of external standards. A deuterated surrogate recovery standard is added to provide information on matrix effects. | 0.001 | Y | Y | Y | 330 | 4/15/2011 | 8/14/2024 | 8/16/2024 4:05:21 PM | 2012-03-28 |  | RDGIBSON |  | Concentration | mg/L | Y | (fl. conc.) |  | 1;4-Dichlorobenzene (fl. conc.) | 106-46-7 | Volatile Organic Compounds (VOCs); | Volatile Organic Compounds (VOCs); EMS code: B025 | Chemical |  | LAB
B025 | mg/L | 1;4-Dichlorobenzene | MS04 | Static Headspace GC/MS (Water) | (Water) A static headspace procedure is used as a means of driving the analytes out of the matrix into a headspace. Once equilibrium is achieved, a fixed volume of headspeace is taken for subsequent determination by Gas Chromatography - Mass Spectroscopy (GC/MS) using Selected Ion Monitoring. | 0.0005 | Y | Y | Y | 69 | 8/13/2009 | 6/28/2024 12:29:00 AM | 7/9/2024 3:31:02 PM | 2009-12-18 |  | RDGIBSON |  | Concentration | mg/L | Y | (fl. conc.) |  | 1;4-Dichlorobenzene (fl. conc.) | 106-46-7 | Volatile Organic Compounds (VOCs); | Volatile Organic Compounds (VOCs); EMS code: B025 | Chemical |  | LAB
B026 | mg/L | 1;3-Dichlorobenzene | MS01 | Purge&Trap GC/MS (VOC-W) | Purge&Trap GC/MS (VOC-W) | 0.6 | Y | Y | Y | 552 | 4/15/2002 | 6/4/2024 10:24:00 AM | 6/13/2024 5:01:27 PM | 1996-04-12 | 3/18/2003 2:49:40 PM | EMS | UMULLER | Concentration | mg/L | Y | (fl. conc.) |  | 1;3-Dichlorobenzene (fl. conc.) | 541-73-1 | Volatile Organic Compounds (VOCs); | Volatile Organic Compounds (VOCs); EMS code: B026 | Chemical |  | LAB
B026 | ug/g | 1;3-Dichlorobenzene | FS04 | Static Headspace:HRGC/LRMS | Static Headspace anaylsis is a method for the determination of volatile compounds in liquids and/or solids by analysing the vapour phase of a sample in a closed system; in this case a sealed vial. Static Headspace is sometimes referred to as Equilibrium Headspace because the vapour phase above the sample in a sealed vial is brought into thermodynamic equilibrium with the sample prior to analysis. Volatile organic compounds are thermally extracted from the sample matrix by Static Headspace sampling and introduced directly to a thick-film megabore GC column. Compounds eluting from the GC column are introduced into a Mass Spectrometer and target compounds are identified by a comparison of both retention times and mass spectra with those of reference standards stored in a computerized data system. Reference retention times and mass spectra for analytes are obtained by the measurement of calibration standards under precisely the same extraction conditions used for samples. The concentration of each identified component is measured by relating the response of the sample quantitation mass ions to the response of the quantitation mass ions of a series of external standards. A deuterated surrogate recovery standard is added to provide information on matrix effects. | 0.05 | Y | Y | Y | 43 | 5/21/2015 | 5/22/2015 | 7/3/2015 3:19:03 PM | 2015-07-03 |  | RDGIBSON |  | MassPerMass | mg/kg | Y | (mass conc.) |  | 1;3-Dichlorobenzene (mass conc.) | 541-73-1 | Volatile Organic Compounds (VOCs); | Volatile Organic Compounds (VOCs); EMS code: B026 | Chemical |  | LAB
B026 | mg/L | 1;3-Dichlorobenzene | F007 | Static Headspace:HRGC/LRMS | Static headspace analysis is a method for the determination of volatile compounds in liquids and/or solids by analysing the vapour phase of a sample in a closed system; in this case a sealed vial. Static Headspace is sometimes referred to as Equilibrium Headspace because the vapour phase above the sample in a sealed vial is brought into thermodynamic equilibrium with the sample prior to analysis. Volatile organic compounds are thermally extracted from the sample matrix by Static Headspace sampling and introduced directly to a thick-film megabore GC column. Compounds eluting from the GC column are introduced into a Mass Spectrometer and target compounds are identified by a comparison of both retention times and mass spectra with those of reference standards stored in a computerized data system. Reference retention times and mass spectra for analytes are obtained by the measurement of calibration standards under precisely the same extraction conditions used for samples. The concentration of each identified component is measured by relating the response of the sample quantitation mass ions to the response of the quantitation mass ions of a series of external standards. A deuterated surrogate recovery standard is added to provide information on matrix effects. | 0.001 | Y | Y | Y | 329 | 4/15/2011 | 8/14/2024 | 8/16/2024 4:05:21 PM | 2012-03-28 |  | RDGIBSON |  | Concentration | mg/L | Y | (fl. conc.) |  | 1;3-Dichlorobenzene (fl. conc.) | 541-73-1 | Volatile Organic Compounds (VOCs); | Volatile Organic Compounds (VOCs); EMS code: B026 | Chemical |  | LAB
B026 | ug/g | 1;3-Dichlorobenzene | MS02 | Purge&Trap GC/MS (VOC-S) | Purge&Trap GC/MS (VOC-S) | 0.003 | Y | Y | Y | 27 | 11/19/2004 | 10/28/2022 5:04:00 PM | 1/11/2024 1:07:07 PM | 1996-04-12 | 6/3/1998 6:51:34 PM | EMS |  | MassPerMass | mg/kg | Y | (mass conc.) |  | 1;3-Dichlorobenzene (mass conc.) | 541-73-1 | Volatile Organic Compounds (VOCs); | Volatile Organic Compounds (VOCs); EMS code: B026 | Chemical |  | LAB
B026 | mg/L | 1;3-Dichlorobenzene | X384 | Purge and trap GC/MS | The sample is analyzed by purge and trap gas chromatography with detection by mass spectrometry. The procedure involves purging the volatile compounds from the sample with an inert gas, and trapping them on a solid sorbent. When purging of the sample is complete, the trap is heated and the compounds are transferred to the gas chromatographic column. Analysis is then accomplished by separation of the components by gas chromatography with detection by mass spectrometry. | 0.001 | Y | Y | Y | 110 | 7/29/2005 7:02:00 AM | 7/26/2012 7:12:00 AM | 9/26/2012 1:07:08 PM | 1996-04-12 | 6/3/1998 6:51:34 PM | EMS |  | Concentration | mg/L | Y | (fl. conc.) |  | 1;3-Dichlorobenzene (fl. conc.) | 541-73-1 | Volatile Organic Compounds (VOCs); | Volatile Organic Compounds (VOCs); EMS code: B026 | Chemical |  | LAB
B026 | mg/L | 1;3-Dichlorobenzene | MS04 | Static Headspace GC/MS (Water) | (Water) A static headspace procedure is used as a means of driving the analytes out of the matrix into a headspace. Once equilibrium is achieved, a fixed volume of headspeace is taken for subsequent determination by Gas Chromatography - Mass Spectroscopy (GC/MS) using Selected Ion Monitoring. | 0.0005 | Y | Y | Y | 70 | 8/13/2009 | 6/28/2024 12:29:00 AM | 7/9/2024 3:31:02 PM | 2009-12-18 |  | RDGIBSON |  | Concentration | mg/L | Y | (fl. conc.) |  | 1;3-Dichlorobenzene (fl. conc.) | 541-73-1 | Volatile Organic Compounds (VOCs); | Volatile Organic Compounds (VOCs); EMS code: B026 | Chemical |  | LAB
B027 | mg/L | Trichlorobenzene-1;4 | X384 | Purge and trap GC/MS | The sample is analyzed by purge and trap gas chromatography with detection by mass spectrometry. The procedure involves purging the volatile compounds from the sample with an inert gas, and trapping them on a solid sorbent. When purging of the sample is complete, the trap is heated and the compounds are transferred to the gas chromatographic column. Analysis is then accomplished by separation of the components by gas chromatography with detection by mass spectrometry. | 0.0002 | N | Y | Y | 1 |  |  | 10/23/1996 11:03:45 PM | 1996-04-12 | 5/29/2000 4:10:34 PM | EMS | NPEPPIN | Concentration | mg/L | Y | (fl. conc.) |  | Trichlorobenzene-1;4 (fl. conc.) | 120-82-1 | Organics; | Organics; EMS code: B027 | Chemical |  | LAB
B028 | mg/L | Dibromochloromethane | MS04 | Static Headspace GC/MS (Water) | (Water) A static headspace procedure is used as a means of driving the analytes out of the matrix into a headspace. Once equilibrium is achieved, a fixed volume of headspeace is taken for subsequent determination by Gas Chromatography - Mass Spectroscopy (GC/MS) using Selected Ion Monitoring. | 0.001 | Y | Y | Y | 7 |  |  | 3/10/2023 3:28:43 PM | 2009-12-18 |  | RDGIBSON |  | Concentration | mg/L | Y | (fl. conc.) |  | Dibromochloromethane (fl. conc.) | 124-48-1 | Organics; | Organics; EMS code: B028 | Chemical |  | LAB
B028 | mg/L | Dibromochloromethane | P016 | Trihalomethanes | Trihalomethanes | 0.005 | Y | Y | Y | 3 | 3/24/2023 10:24:00 PM | 12/17/2023 11:36:00 AM | 12/29/2023 9:48:33 AM | 2023-03-31 |  | EMS |  | Concentration | mg/L | Y | (fl. conc.) |  | Dibromochloromethane (fl. conc.) | 124-48-1 | Organics; | Organics; EMS code: B028 | Chemical |  | LAB
B028 | ug/g | Dibromochloromethane | MS02 | Purge&Trap GC/MS (VOC-S) | Purge&Trap GC/MS (VOC-S) | 0.1 | Y | Y | Y | 1 | 10/28/2022 5:04:00 PM | 10/28/2022 5:04:00 PM | 1/7/2023 3:12:09 PM | 2022-12-29 |  | EMS |  | MassPerMass | mg/kg | Y | (mass conc.) |  | Dibromochloromethane (mass conc.) | 124-48-1 | Organics; | Organics; EMS code: B028 | Chemical |  | LAB
B028 | mg/L | Dibromochloromethane | MS01 | Purge&Trap GC/MS (VOC-W) | Purge&Trap GC/MS (VOC-W) | 0.001 | Y | Y | Y | 458 | 7/29/2005 7:02:00 AM | 6/4/2024 10:24:00 AM | 6/13/2024 5:01:27 PM | 2003-06-17 |  | UMULLER |  | Concentration | mg/L | Y | (fl. conc.) |  | Dibromochloromethane (fl. conc.) | 124-48-1 | Organics; | Organics; EMS code: B028 | Chemical |  | LAB
B030 | mg/L | Butylate | ZN01 | DCM+Ac.Extr:GC/MS | DCM+Ac.Extr:GC/MS | 0.0005 | Y | Y | Y | 28 | 7/28/2005 7:03:00 AM | 3/29/2012 10:15:00 PM | 5/3/2012 9:27:49 AM | 1996-04-12 | 11/16/1999 4:06:27 PM | EMS | NPEPPIN | Concentration | mg/L | Y | (fl. conc.) |  | Butylate (fl. conc.) | 2008-41-5 | Pesticides; | Pesticides; EMS code: B030 | Chemical |  | LAB
B031 | % | Dibromophenol | SURR | Surrogate Recovery | Surrogate Recovery | 30 | Y | Y | Y | 106 |  |  | 3/18/1997 2:04:17 PM | 1996-04-12 | 6/3/1998 6:51:34 PM | EMS |  | DimensionlessRatio | % | N | (ratio) |  | Dibromophenol (ratio) | 615-58-7 | Organics; | Organics; EMS code: B031 | Chemical |  | SURROGATE_RESULT
B032 | % | Tribromophenol | SURR | Surrogate Recovery | Surrogate Recovery | 30 | Y | Y | Y | 156 | 4/4/2002 | 6/24/2003 | 2/26/2004 1:05:26 AM | 1996-04-12 | 6/3/1998 6:51:34 PM | EMS |  | DimensionlessRatio | % | N | (ratio) |  | Tribromophenol (ratio) | 25376-38-9 | Organics; | Organics;EMS code: B032 | Chemical |  | SURROGATE_RESULT
B037 | mg/L | 1;2-Dibromoethane | MS01 | Purge&Trap GC/MS (VOC-W) | Purge&Trap GC/MS (VOC-W) | 0.001 | Y | Y | Y | 365 | 1/15/2007 4:50:00 PM | 6/4/2024 10:24:00 AM | 6/13/2024 5:01:27 PM | 2003-06-17 |  | UMULLER |  | Concentration | mg/L | Y | (fl. conc.) |  | 1;2-Dibromoethane (fl. conc.) | 106-93-4 | Volatile Organic Compounds (VOCs); | Volatile Organic Compounds (VOCs); EMS code: B037 | Chemical |  | LAB
B037 | mg/L | 1;2-Dibromoethane | MS04 | Static Headspace GC/MS (Water) | (Water) A static headspace procedure is used as a means of driving the analytes out of the matrix into a headspace. Once equilibrium is achieved, a fixed volume of headspeace is taken for subsequent determination by Gas Chromatography - Mass Spectroscopy (GC/MS) using Selected Ion Monitoring. | 0.0005 | Y | Y | Y | 33 | 8/13/2009 | 6/28/2024 12:29:00 AM | 7/9/2024 3:31:02 PM | 2010-01-13 | 1/13/2010 11:39:41 AM | RDGIBSON | RDGIBSON | Concentration | mg/L | Y | (fl. conc.) |  | 1;2-Dibromoethane (fl. conc.) | 106-93-4 | Volatile Organic Compounds (VOCs); | Volatile Organic Compounds (VOCs); EMS code: B037 | Chemical |  | LAB
B037 | ug/g | 1;2-Dibromoethane | MS02 | Purge&Trap GC/MS (VOC-S) | Purge&Trap GC/MS (VOC-S) | 0.1 | Y | Y | Y | 1 | 10/28/2022 5:04:00 PM | 10/28/2022 5:04:00 PM | 1/7/2023 3:12:09 PM | 2022-12-29 |  | EMS |  | MassPerMass | mg/kg | Y | (mass conc.) |  | 1;2-Dibromoethane (mass conc.) | 106-93-4 | Volatile Organic Compounds (VOCs); | Volatile Organic Compounds (VOCs); EMS code: B037 | Chemical |  | LAB
B039 | mg/L | 1;2-Dichlorobenzene | MS01 | Purge&Trap GC/MS (VOC-W) | Purge&Trap GC/MS (VOC-W) | 0.001 | Y | Y | Y | 433 | 7/29/2005 7:02:00 AM | 6/4/2024 10:24:00 AM | 6/13/2024 5:01:27 PM | 2003-06-17 |  | UMULLER |  | Concentration | mg/L | Y | (fl. conc.) |  | 1;2-Dichlorobenzene (fl. conc.) | 95-50-1 | Volatile Organic Compounds (VOCs); | Volatile Organic Compounds (VOCs); EMS code: B039 | Chemical |  | LAB
B039 | ug/g | 1;2-Dichlorobenzene | MS02 | Purge&Trap GC/MS (VOC-S) | Purge&Trap GC/MS (VOC-S) | 0.05 | Y | Y | Y | 1 | 10/28/2022 5:04:00 PM | 10/28/2022 5:04:00 PM | 1/7/2023 3:12:09 PM | 2022-12-29 |  | EMS |  | MassPerMass | mg/kg | Y | (mass conc.) |  | 1;2-Dichlorobenzene (mass conc.) | 95-50-1 | Volatile Organic Compounds (VOCs); | Volatile Organic Compounds (VOCs); EMS code: B039 | Chemical |  | LAB
B040 | mg/L | 1;4-Dichlorobenzene-d4 (IS) | MS01 | Purge&Trap GC/MS (VOC-W) | Purge&Trap GC/MS (VOC-W) |  | Y | Y | Y | 58 | 1/15/2007 4:50:00 PM | 6/2/2022 6:05:00 AM | 6/3/2022 5:00:16 PM | 2003-06-17 |  | UMULLER |  | Concentration | mg/L | Y | (fl. conc.) |  | 1;4-Dichlorobenzene-d4 (IS) (fl. conc.) | 3855-82-1 | Volatile Organic Compounds (VOCs); | Volatile Organic Compounds (VOCs); EMS code: B040 | Chemical |  | SURROGATE_RESULT
B043 | % | 1;4-Difluorobenzene (IS) | F007 | Static Headspace:HRGC/LRMS | Static headspace analysis is a method for the determination of volatile compounds in liquids and/or solids by analysing the vapour phase of a sample in a closed system; in this case a sealed vial. Static Headspace is sometimes referred to as Equilibrium Headspace because the vapour phase above the sample in a sealed vial is brought into thermodynamic equilibrium with the sample prior to analysis. Volatile organic compounds are thermally extracted from the sample matrix by Static Headspace sampling and introduced directly to a thick-film megabore GC column. Compounds eluting from the GC column are introduced into a Mass Spectrometer and target compounds are identified by a comparison of both retention times and mass spectra with those of reference standards stored in a computerized data system. Reference retention times and mass spectra for analytes are obtained by the measurement of calibration standards under precisely the same extraction conditions used for samples. The concentration of each identified component is measured by relating the response of the sample quantitation mass ions to the response of the quantitation mass ions of a series of external standards. A deuterated surrogate recovery standard is added to provide information on matrix effects. | 1 | Y | Y | Y | 61 | 10/12/2010 | 12/22/2014 | 1/29/2015 2:12:57 PM | 2012-06-25 | 6/25/2012 10:39:25 AM | RDGIBSON | RDGIBSON | DimensionlessRatio | % | N | (ratio) |  | 1;4-Difluorobenzene (IS) (ratio) | 540-36-3 | Organics; | Organics;EMS code: B043 | Chemical |  | LAB
B043 | mg/L | 1;4-Difluorobenzene (IS) | E611 | E611S: EPA 8260D (mod) | E611S: Dimethyl selenide and dimethyl diselenide are analyzed by static headspace GC-MS. Samples are prepared in headspace vials and are heated and agitated on the headspace autosampler; causing VOCs to partition between the aqueous phase and the headspace in accordance with Henry`s law. | 0.001 | Y | Y | Y | 0 |  |  |  | 2024-01-15 | 1/15/2024 2:17:10 PM | EMS | EMS | Concentration | mg/L | Y | (fl. conc.) |  | 1;4-Difluorobenzene (IS) (fl. conc.) | 540-36-3 | Organics; | Organics;EMS code: B043 | Chemical |  | LAB
B043 | % | 1;4-Difluorobenzene (IS) | MS04 | Static Headspace GC/MS (Water) | (Water) A static headspace procedure is used as a means of driving the analytes out of the matrix into a headspace. Once equilibrium is achieved, a fixed volume of headspeace is taken for subsequent determination by Gas Chromatography - Mass Spectroscopy (GC/MS) using Selected Ion Monitoring. | 1 | Y | Y | Y | 46 | 1/9/2015 | 5/22/2015 | 7/3/2015 3:19:03 PM | 2015-04-15 |  | RDGIBSON |  | DimensionlessRatio | % | N | (ratio) |  | 1;4-Difluorobenzene (IS) (ratio) | 540-36-3 | Organics; | Organics;EMS code: B043 | Chemical |  | LAB
B043 | % | 1;4-Difluorobenzene (IS) | SURR | Surrogate Recovery | Surrogate Recovery | 1 | Y | Y | Y | 1333 | 11/1/2011 | 8/14/2024 | 8/16/2024 4:05:21 PM | 2015-07-08 |  | RDGIBSON |  | DimensionlessRatio | % | N | (ratio) |  | 1;4-Difluorobenzene (IS) (ratio) | 540-36-3 | Organics; | Organics;EMS code: B043 | Chemical |  | SURROGATE_RESULT
B043 | mg/L | 1;4-Difluorobenzene (IS) | 8260D | EPA 8260D (mod) | EPA 8260D (mod) VOLATILE ORGANIC COMPOUNDS BY GAS CHROMATOGRAPHY/MASS SPECTROMETRY This method is used to determine volatile organic compounds (VOCs) in a variety of solid waste matrices. | 0.001 | Y | Y | Y | 0 |  |  |  | 2024-01-25 |  | RDGIBSON |  | Concentration | mg/L | Y | (fl. conc.) |  | 1;4-Difluorobenzene (IS) (fl. conc.) | 540-36-3 | Organics; | Organics;EMS code: B043 | Chemical |  | LAB
B044 | mg/L | Bifenox | ZN01 | DCM+Ac.Extr:GC/MS | DCM+Ac.Extr:GC/MS | 0.0005 | Y | Y | Y | 1 | 7/29/2005 7:10:00 AM | 7/29/2005 7:10:00 AM | 6/2/2009 12:00:58 PM | 2003-06-16 |  | UMULLER |  | Concentration | mg/L | Y | (fl. conc.) |  | Bifenox (fl. conc.) | 42576-02-3 | Pesticides; | Pesticides; EMS code: B044 | Chemical |  | LAB
B046 | mg/L | Butalin | ZN01 | DCM+Ac.Extr:GC/MS | DCM+Ac.Extr:GC/MS | 0.0005 | Y | Y | Y | 1 | 7/29/2005 7:10:00 AM | 7/29/2005 7:10:00 AM | 6/2/2009 12:00:58 PM | 2003-06-16 |  | UMULLER |  | Concentration | mg/L | Y | (fl. conc.) |  | Butalin (fl. conc.) | 33629-47-9 | Pesticides; | Pesticides; EMS code: B046 | Chemical |  | LAB
B047 | mg/L | Bromopropylate | ZN01 | DCM+Ac.Extr:GC/MS | DCM+Ac.Extr:GC/MS | 0.0005 | Y | Y | Y | 1 | 7/29/2005 7:10:00 AM | 7/29/2005 7:10:00 AM | 6/2/2009 12:00:58 PM | 2003-06-16 |  | UMULLER |  | Concentration | mg/L | Y | (fl. conc.) |  | Bromopropylate (fl. conc.) | 18181-80-1 | Pesticides; | Pesticides; EMS code: B047 | Chemical |  | LAB
B048 | % | 2,4,6-Tribromophenol | SURR | Surrogate Recovery | Surrogate Recovery | 1 | Y | Y | Y | 378 | 6/8/2019 | 8/14/2024 | 8/16/2024 4:05:21 PM | 2019-05-15 |  | KATFRASE |  | DimensionlessRatio | % | N | (ratio) |  | 2,4,6-Tribromophenol (ratio) | 118-79-6 | Organics; | Organics;EMS code: B048 | Chemical |  | SURROGATE_RESULT
B049 | % | 2,4-Dichlorophenol-d3 | SURR | Surrogate Recovery | Surrogate Recovery | 1 | Y | Y | Y | 346 | 6/8/2019 | 8/14/2024 | 8/16/2024 4:05:21 PM | 2019-05-16 |  | KATFRASE |  | DimensionlessRatio | % | N | (ratio) |  | 2,4-Dichlorophenol-d3 (ratio) | 93951-74-7 | Organics; | Organics;EMS code: B049 | Chemical |  | SURROGATE_RESULT
B050 | % | 2-Chlorophenol-d4 | SURR | Surrogate Recovery | Surrogate Recovery | 1 | Y | Y | Y | 346 | 6/8/2019 | 8/14/2024 | 8/16/2024 4:05:21 PM | 2019-05-16 |  | KATFRASE |  | DimensionlessRatio | % | N | (ratio) |  | 2-Chlorophenol-d4 (ratio) | 93951-73-6 | Organics; | Organics;EMS code: B050 | Chemical |  | SURROGATE_RESULT
B051 | ug/g | Bendiocarb | ZN03 | DCM+Ac.Extr:GC/MS | DCM+Ac.Extr:GC/MS EPA 8270E (mod) | 0.05 | Y | Y | Y | 14 | 1/25/2022 | 9/8/2022 | 12/15/2022 10:50:07 AM | 2022-06-13 |  | EMS |  | MassPerMass | mg/kg | Y | (mass conc.) |  | Bendiocarb (mass conc.) | 22781-23-3 | Pesticides; | Pesticides; EMS code: B051 | Chemical |  | LAB
B051 | mg/L | Bendiocarb | E660 | EPA 8270E (mod) | EPA 8270E (mod) Miscellaneous Pesticides by GC-MS | 0.0005 | Y | Y | Y | 68 | 12/7/2021 | 9/12/2022 | 12/8/2022 12:57:05 PM | 2021-12-09 | 12/9/2021 3:22:12 PM | EMS | EMS | Concentration | mg/L | Y | (fl. conc.) |  | Bendiocarb (fl. conc.) | 22781-23-3 | Pesticides; | Pesticides; EMS code: B051 | Chemical |  | LAB
B051 | mg/L | Bendiocarb | ZN01 | DCM+Ac.Extr:GC/MS | DCM+Ac.Extr:GC/MS | 0.02 | Y | Y | Y | 1 | 7/29/2005 7:10:00 AM | 7/29/2005 7:10:00 AM | 6/2/2009 12:00:58 PM | 2009-02-20 |  | JCHUDAK |  | Concentration | mg/L | Y | (fl. conc.) |  | Bendiocarb (fl. conc.) | 22781-23-3 | Pesticides; | Pesticides; EMS code: B051 | Chemical |  | LAB
B052 | mg/L | BPMC | ZN01 | DCM+Ac.Extr:GC/MS | DCM+Ac.Extr:GC/MS | 0.02 | Y | Y | Y | 1 | 7/29/2005 7:10:00 AM | 7/29/2005 7:10:00 AM | 6/2/2009 12:00:58 PM | 2009-02-20 |  | JCHUDAK |  | Concentration | mg/L | Y | (fl. conc.) |  | BPMC (fl. conc.) | 3766-81-2 | Pesticides; | Pesticides; EMS code: B052 | Chemical |  | LAB
B053 | pg/L | 1,2,3,4,6,7,8-HpCDD | E1613B | EPA 1613B | EPA 1613B | 1.4 | Y | Y | Y | 1 | 2/3/2022 | 2/3/2022 | 7/6/2022 4:59:13 PM | 2022-06-15 |  | EMS |  | Concentration | pg/L | Y | (fl. conc.) |  | 1,2,3,4,6,7,8-HpCDD (fl. conc.) | 35822-46-9 | Dioxins & Furans; | Dioxins & Furans; EMS code: B053 | Chemical |  | LAB
B053 | pg/g | 1,2,3,4,6,7,8-HpCDD | 1613B | USEPA 1613B | USEPA 1613B | 0.17 | Y | Y | Y | 4 | 3/20/2020 | 3/21/2020 | 2/21/2023 2:10:21 PM | 2023-02-06 |  | EMS |  | MassPerMass | pg/g | Y | (mass conc.) |  | 1,2,3,4,6,7,8-HpCDD (mass conc.) | 35822-46-9 | Dioxins & Furans; | Dioxins & Furans; EMS code: B053 | Chemical |  | LAB
B053 | pg/g wet | 1,2,3,4,6,7,8-HpCDD | 1101 | DX-1613B | DCM Extraction, Column clean-up, isotope dilution GC/HRMS. Method Reference: USEPA 1613B. Methodology: Samples are filtered if required. Solids are extracted by Soxhlet using toluene. The liquid portion is extracted by liquid/liquid extraction using dichloromethane. The extracts are prepared using column chromatography, reduced in volume and analyzed by isotope-dilution GC/HRMS. | 0.038 | Y | Y | Y | 1 | 2/2/2022 | 2/2/2022 | 7/6/2022 5:20:42 PM | 2019-06-03 |  | KATFRASE |  | MassPerMass | pg/g | Y | (mass conc.) | (wet) | 1,2,3,4,6,7,8-HpCDD (wet) (mass conc.) | 35822-46-9 | Dioxins & Furans; | Dioxins & Furans; EMS code: B053 | Chemical |  | LAB
B056 | pg/L | 1,2,3,4,6,7,8-HpCDF | E1613B | EPA 1613B | EPA 1613B | 0.76 | Y | Y | Y | 1 | 2/3/2022 | 2/3/2022 | 7/6/2022 4:59:13 PM | 2022-06-15 |  | EMS |  | Concentration | pg/L | Y | (fl. conc.) |  | 1,2,3,4,6,7,8-HpCDF (fl. conc.) | 67562-39-4 | Dioxins & Furans; | Dioxins & Furans; EMS code: B056 | Chemical |  | LAB
B056 | pg/g wet | 1,2,3,4,6,7,8-HpCDF | 1101 | DX-1613B | DCM Extraction, Column clean-up, isotope dilution GC/HRMS. Method Reference: USEPA 1613B. Methodology: Samples are filtered if required. Solids are extracted by Soxhlet using toluene. The liquid portion is extracted by liquid/liquid extraction using dichloromethane. The extracts are prepared using column chromatography, reduced in volume and analyzed by isotope-dilution GC/HRMS. | 0.025 | Y | Y | Y | 1 | 2/2/2022 | 2/2/2022 | 7/6/2022 5:20:42 PM | 2019-06-03 |  | KATFRASE |  | MassPerMass | pg/g | Y | (mass conc.) | (wet) | 1,2,3,4,6,7,8-HpCDF (wet) (mass conc.) | 67562-39-4 | Dioxins & Furans; | Dioxins & Furans; EMS code: B056 | Chemical |  | LAB
B056 | pg/g | 1,2,3,4,6,7,8-HpCDF | 1613B | USEPA 1613B | USEPA 1613B | 0.068 | Y | Y | Y | 4 | 3/20/2020 | 3/21/2020 | 2/21/2023 2:10:21 PM | 2023-02-06 |  | EMS |  | MassPerMass | pg/g | Y | (mass conc.) |  | 1,2,3,4,6,7,8-HpCDF (mass conc.) | 67562-39-4 | Dioxins & Furans; | Dioxins & Furans; EMS code: B056 | Chemical |  | LAB
B059 | pg/L | 1,2,3,4,7,8,9-HpCDF | E1613B | EPA 1613B | EPA 1613B | 1.2 | Y | Y | Y | 1 | 2/3/2022 | 2/3/2022 | 7/6/2022 4:59:13 PM | 2022-06-15 |  | EMS |  | Concentration | pg/L | Y | (fl. conc.) |  | 1,2,3,4,7,8,9-HpCDF (fl. conc.) | 55673-89-7 | Dioxins & Furans; | Dioxins & Furans; EMS code: B059 | Chemical |  | LAB
B059 | pg/g | 1,2,3,4,7,8,9-HpCDF | 1613B | USEPA 1613B | USEPA 1613B | 0.069 | Y | Y | Y | 4 | 3/20/2020 | 3/21/2020 | 2/21/2023 2:10:21 PM | 2023-02-06 |  | EMS |  | MassPerMass | pg/g | Y | (mass conc.) |  | 1,2,3,4,7,8,9-HpCDF (mass conc.) | 55673-89-7 | Dioxins & Furans; | Dioxins & Furans; EMS code: B059 | Chemical |  | LAB
B059 | pg/g wet | 1,2,3,4,7,8,9-HpCDF | 1101 | DX-1613B | DCM Extraction, Column clean-up, isotope dilution GC/HRMS. Method Reference: USEPA 1613B. Methodology: Samples are filtered if required. Solids are extracted by Soxhlet using toluene. The liquid portion is extracted by liquid/liquid extraction using dichloromethane. The extracts are prepared using column chromatography, reduced in volume and analyzed by isotope-dilution GC/HRMS. | 0.033 | Y | Y | Y | 1 | 2/2/2022 | 2/2/2022 | 7/6/2022 5:20:42 PM | 2019-06-03 |  | KATFRASE |  | MassPerMass | pg/g | Y | (mass conc.) | (wet) | 1,2,3,4,7,8,9-HpCDF (wet) (mass conc.) | 55673-89-7 | Dioxins & Furans; | Dioxins & Furans; EMS code: B059 | Chemical |  | LAB
B062 | pg/L | 1,2,3,4,7,8-HxCDD | E1613B | EPA 1613B | EPA 1613B | 0.6 | Y | Y | Y | 1 | 2/3/2022 | 2/3/2022 | 7/6/2022 4:59:13 PM | 2022-06-15 |  | EMS |  | Concentration | pg/L | Y | (fl. conc.) |  | 1,2,3,4,7,8-HxCDD (fl. conc.) | 39227-28-6 | Dioxins & Furans; | Dioxins & Furans; EMS code: B062 | Chemical |  | LAB
B062 | pg/g | 1,2,3,4,7,8-HxCDD | 1613B | USEPA 1613B | USEPA 1613B | 0.096 | Y | Y | Y | 4 | 3/20/2020 | 3/21/2020 | 2/21/2023 2:10:21 PM | 2023-02-06 |  | EMS |  | MassPerMass | pg/g | Y | (mass conc.) |  | 1,2,3,4,7,8-HxCDD (mass conc.) | 39227-28-6 | Dioxins & Furans; | Dioxins & Furans; EMS code: B062 | Chemical |  | LAB
B062 | pg/g wet | 1,2,3,4,7,8-HxCDD | 1101 | DX-1613B | DCM Extraction, Column clean-up, isotope dilution GC/HRMS. Method Reference: USEPA 1613B. Methodology: Samples are filtered if required. Solids are extracted by Soxhlet using toluene. The liquid portion is extracted by liquid/liquid extraction using dichloromethane. The extracts are prepared using column chromatography, reduced in volume and analyzed by isotope-dilution GC/HRMS. | 0.044 | Y | Y | Y | 1 | 2/2/2022 | 2/2/2022 | 7/6/2022 5:20:42 PM | 2019-06-04 |  | KATFRASE |  | MassPerMass | pg/g | Y | (mass conc.) | (wet) | 1,2,3,4,7,8-HxCDD (wet) (mass conc.) | 39227-28-6 | Dioxins & Furans; | Dioxins & Furans; EMS code: B062 | Chemical |  | LAB
B065 | pg/g wet | 1,2,3,4,7,8-HxCDF | 1101 | DX-1613B | DCM Extraction, Column clean-up, isotope dilution GC/HRMS. Method Reference: USEPA 1613B. Methodology: Samples are filtered if required. Solids are extracted by Soxhlet using toluene. The liquid portion is extracted by liquid/liquid extraction using dichloromethane. The extracts are prepared using column chromatography, reduced in volume and analyzed by isotope-dilution GC/HRMS. | 0.028 | Y | Y | Y | 1 | 2/2/2022 | 2/2/2022 | 7/6/2022 5:20:42 PM | 2019-06-04 |  | KATFRASE |  | MassPerMass | pg/g | Y | (mass conc.) | (wet) | 1,2,3,4,7,8-HxCDF (wet) (mass conc.) | 70648-26-9 | Dioxins & Furans; | Dioxins & Furans; EMS code: B065 | Chemical |  | LAB
B065 | pg/L | 1,2,3,4,7,8-HxCDF | E1613B | EPA 1613B | EPA 1613B | 0.6 | Y | Y | Y | 1 | 2/3/2022 | 2/3/2022 | 7/6/2022 4:59:13 PM | 2022-06-15 |  | EMS |  | Concentration | pg/L | Y | (fl. conc.) |  | 1,2,3,4,7,8-HxCDF (fl. conc.) | 70648-26-9 | Dioxins & Furans; | Dioxins & Furans; EMS code: B065 | Chemical |  | LAB
B065 | pg/g | 1,2,3,4,7,8-HxCDF | 1613B | USEPA 1613B | USEPA 1613B | 0.05 | Y | Y | Y | 4 | 3/20/2020 | 3/21/2020 | 2/21/2023 2:10:21 PM | 2023-02-06 |  | EMS |  | MassPerMass | pg/g | Y | (mass conc.) |  | 1,2,3,4,7,8-HxCDF (mass conc.) | 70648-26-9 | Dioxins & Furans; | Dioxins & Furans; EMS code: B065 | Chemical |  | LAB
B068 | pg/L | 1,2,3,6,7,8-HxCDD | E1613B | EPA 1613B | EPA 1613B | 0.57 | Y | Y | Y | 1 | 2/3/2022 | 2/3/2022 | 7/6/2022 4:59:13 PM | 2022-06-15 |  | EMS |  | Concentration | pg/L | Y | (fl. conc.) |  | 1,2,3,6,7,8-HxCDD (fl. conc.) | 57653-85-7 | Dioxins & Furans; | Dioxins & Furans; EMS code: B068 | Chemical |  | LAB
B068 | pg/g | 1,2,3,6,7,8-HxCDD | 1613B | USEPA 1613B | USEPA 1613B | 0.099 | Y | Y | Y | 4 | 3/20/2020 | 3/21/2020 | 2/21/2023 2:10:21 PM | 2023-02-06 |  | EMS |  | MassPerMass | pg/g | Y | (mass conc.) |  | 1,2,3,6,7,8-HxCDD (mass conc.) | 57653-85-7 | Dioxins & Furans; | Dioxins & Furans; EMS code: B068 | Chemical |  | LAB
B068 | pg/g wet | 1,2,3,6,7,8-HxCDD | 1101 | DX-1613B | DCM Extraction, Column clean-up, isotope dilution GC/HRMS. Method Reference: USEPA 1613B. Methodology: Samples are filtered if required. Solids are extracted by Soxhlet using toluene. The liquid portion is extracted by liquid/liquid extraction using dichloromethane. The extracts are prepared using column chromatography, reduced in volume and analyzed by isotope-dilution GC/HRMS. | 0.043 | Y | Y | Y | 1 | 2/2/2022 | 2/2/2022 | 7/6/2022 5:20:42 PM | 2019-06-04 |  | KATFRASE |  | MassPerMass | pg/g | Y | (mass conc.) | (wet) | 1,2,3,6,7,8-HxCDD (wet) (mass conc.) | 57653-85-7 | Dioxins & Furans; | Dioxins & Furans; EMS code: B068 | Chemical |  | LAB
B071 | pg/g wet | 1,2,3,6,7,8-HxCDF | 1101 | DX-1613B | DCM Extraction, Column clean-up, isotope dilution GC/HRMS. Method Reference: USEPA 1613B. Methodology: Samples are filtered if required. Solids are extracted by Soxhlet using toluene. The liquid portion is extracted by liquid/liquid extraction using dichloromethane. The extracts are prepared using column chromatography, reduced in volume and analyzed by isotope-dilution GC/HRMS. | 0.027 | Y | Y | Y | 1 | 2/2/2022 | 2/2/2022 | 7/6/2022 5:20:42 PM | 2019-06-04 |  | KATFRASE |  | MassPerMass | pg/g | Y | (mass conc.) | (wet) | 1,2,3,6,7,8-HxCDF (wet) (mass conc.) | 57117-44-9 | Dioxins & Furans; | Dioxins & Furans; EMS code: B071 | Chemical |  | LAB
B071 | pg/g | 1,2,3,6,7,8-HxCDF | 1613B | USEPA 1613B | USEPA 1613B | 0.051 | Y | Y | Y | 4 | 3/20/2020 | 3/21/2020 | 2/21/2023 2:10:21 PM | 2023-02-06 |  | EMS |  | MassPerMass | pg/g | Y | (mass conc.) |  | 1,2,3,6,7,8-HxCDF (mass conc.) | 57117-44-9 | Dioxins & Furans; | Dioxins & Furans; EMS code: B071 | Chemical |  | LAB
B071 | pg/L | 1,2,3,6,7,8-HxCDF | E1613B | EPA 1613B | EPA 1613B | 0.54 | Y | Y | Y | 1 | 2/3/2022 | 2/3/2022 | 7/6/2022 4:59:13 PM | 2022-06-15 |  | EMS |  | Concentration | pg/L | Y | (fl. conc.) |  | 1,2,3,6,7,8-HxCDF (fl. conc.) | 57117-44-9 | Dioxins & Furans; | Dioxins & Furans; EMS code: B071 | Chemical |  | LAB
B074 | pg/g | 1,2,3,7,8,9-HxCDD | 1613B | USEPA 1613B | USEPA 1613B | 0.097 | Y | Y | Y | 4 | 3/20/2020 | 3/21/2020 | 2/21/2023 2:10:21 PM | 2023-02-06 |  | EMS |  | MassPerMass | pg/g | Y | (mass conc.) |  | 1,2,3,7,8,9-HxCDD (mass conc.) | 19408-74-3 | Dioxins & Furans; | Dioxins & Furans; EMS code: B074 | Chemical |  | LAB
B074 | pg/g wet | 1,2,3,7,8,9-HxCDD | 1101 | DX-1613B | DCM Extraction, Column clean-up, isotope dilution GC/HRMS. Method Reference: USEPA 1613B. Methodology: Samples are filtered if required. Solids are extracted by Soxhlet using toluene. The liquid portion is extracted by liquid/liquid extraction using dichloromethane. The extracts are prepared using column chromatography, reduced in volume and analyzed by isotope-dilution GC/HRMS. | 0.044 | Y | Y | Y | 1 | 2/2/2022 | 2/2/2022 | 7/6/2022 5:20:42 PM | 2019-06-04 |  | KATFRASE |  | MassPerMass | pg/g | Y | (mass conc.) | (wet) | 1,2,3,7,8,9-HxCDD (wet) (mass conc.) | 19408-74-3 | Dioxins & Furans; | Dioxins & Furans; EMS code: B074 | Chemical |  | LAB
B074 | pg/L | 1,2,3,7,8,9-HxCDD | E1613B | EPA 1613B | EPA 1613B | 0.54 | Y | Y | Y | 1 | 2/3/2022 | 2/3/2022 | 7/6/2022 4:59:13 PM | 2022-06-15 |  | EMS |  | Concentration | pg/L | Y | (fl. conc.) |  | 1,2,3,7,8,9-HxCDD (fl. conc.) | 19408-74-3 | Dioxins & Furans; | Dioxins & Furans; EMS code: B074 | Chemical |  | LAB
B077 | pg/g | 1,2,3,7,8,9-HxCDF | 1613B | USEPA 1613B | USEPA 1613B | 0.058 | Y | Y | Y | 4 | 3/20/2020 | 3/21/2020 | 2/21/2023 2:10:21 PM | 2023-02-06 |  | EMS |  | MassPerMass | pg/g | Y | (mass conc.) |  | 1,2,3,7,8,9-HxCDF (mass conc.) | 72918-21-9 | Dioxins & Furans; | Dioxins & Furans; EMS code: B077 | Chemical |  | LAB
B077 | pg/L | 1,2,3,7,8,9-HxCDF | E1613B | EPA 1613B | EPA 1613B | 0.93 | Y | Y | Y | 1 | 2/3/2022 | 2/3/2022 | 7/6/2022 4:59:13 PM | 2022-06-15 |  | EMS |  | Concentration | pg/L | Y | (fl. conc.) |  | 1,2,3,7,8,9-HxCDF (fl. conc.) | 72918-21-9 | Dioxins & Furans; | Dioxins & Furans; EMS code: B077 | Chemical |  | LAB
B077 | pg/g wet | 1,2,3,7,8,9-HxCDF | 1101 | DX-1613B | DCM Extraction, Column clean-up, isotope dilution GC/HRMS. Method Reference: USEPA 1613B. Methodology: Samples are filtered if required. Solids are extracted by Soxhlet using toluene. The liquid portion is extracted by liquid/liquid extraction using dichloromethane. The extracts are prepared using column chromatography, reduced in volume and analyzed by isotope-dilution GC/HRMS. | 0.034 | Y | Y | Y | 1 | 2/2/2022 | 2/2/2022 | 7/6/2022 5:20:43 PM | 2019-06-04 |  | KATFRASE |  | MassPerMass | pg/g | Y | (mass conc.) | (wet) | 1,2,3,7,8,9-HxCDF (wet) (mass conc.) | 72918-21-9 | Dioxins & Furans; | Dioxins & Furans; EMS code: B077 | Chemical |  | LAB
B080 | pg/g wet | 1,2,3,7,8-PeCDD | 1101 | DX-1613B | DCM Extraction, Column clean-up, isotope dilution GC/HRMS. Method Reference: USEPA 1613B. Methodology: Samples are filtered if required. Solids are extracted by Soxhlet using toluene. The liquid portion is extracted by liquid/liquid extraction using dichloromethane. The extracts are prepared using column chromatography, reduced in volume and analyzed by isotope-dilution GC/HRMS. | 0.04 | Y | Y | Y | 1 | 2/2/2022 | 2/2/2022 | 7/6/2022 5:20:43 PM | 2019-06-04 |  | KATFRASE |  | MassPerMass | pg/g | Y | (mass conc.) | (wet) | 1,2,3,7,8-PeCDD (wet) (mass conc.) | 40321-76-4 | Dioxins & Furans; | Dioxins & Furans; EMS code: B080 | Chemical |  | LAB
B080 | pg/L | 1,2,3,7,8-PeCDD | E1613B | EPA 1613B | EPA 1613B | 0.42 | Y | Y | Y | 1 | 2/3/2022 | 2/3/2022 | 7/6/2022 4:59:13 PM | 2022-06-15 |  | EMS |  | Concentration | pg/L | Y | (fl. conc.) |  | 1,2,3,7,8-PeCDD (fl. conc.) | 40321-76-4 | Dioxins & Furans; | Dioxins & Furans; EMS code: B080 | Chemical |  | LAB
B080 | pg/g | 1,2,3,7,8-PeCDD | 1613B | USEPA 1613B | USEPA 1613B | 0.039 | Y | Y | Y | 4 | 3/20/2020 | 3/21/2020 | 2/21/2023 2:10:21 PM | 2023-02-06 |  | EMS |  | MassPerMass | pg/g | Y | (mass conc.) |  | 1,2,3,7,8-PeCDD (mass conc.) | 40321-76-4 | Dioxins & Furans; | Dioxins & Furans; EMS code: B080 | Chemical |  | LAB
B083 | pg/g wet | 1,2,3,7,8-PeCDF | 1101 | DX-1613B | DCM Extraction, Column clean-up, isotope dilution GC/HRMS. Method Reference: USEPA 1613B. Methodology: Samples are filtered if required. Solids are extracted by Soxhlet using toluene. The liquid portion is extracted by liquid/liquid extraction using dichloromethane. The extracts are prepared using column chromatography, reduced in volume and analyzed by isotope-dilution GC/HRMS. | 0.036 | Y | Y | Y | 1 | 2/2/2022 | 2/2/2022 | 7/6/2022 5:20:43 PM | 2019-06-04 |  | KATFRASE |  | MassPerMass | pg/g | Y | (mass conc.) | (wet) | 1,2,3,7,8-PeCDF (wet) (mass conc.) | 57117-41-6 | Dioxins & Furans; | Dioxins & Furans; EMS code: B083 | Chemical |  | LAB
B083 | pg/g | 1,2,3,7,8-PeCDF | 1613B | USEPA 1613B | USEPA 1613B | 0.13 | Y | Y | Y | 4 | 3/20/2020 | 3/21/2020 | 2/21/2023 2:10:21 PM | 2023-02-06 |  | EMS |  | MassPerMass | pg/g | Y | (mass conc.) |  | 1,2,3,7,8-PeCDF (mass conc.) | 57117-41-6 | Dioxins & Furans; | Dioxins & Furans; EMS code: B083 | Chemical |  | LAB
B083 | pg/L | 1,2,3,7,8-PeCDF | E1613B | EPA 1613B | EPA 1613B | 0.53 | Y | Y | Y | 1 | 2/3/2022 | 2/3/2022 | 7/6/2022 4:59:13 PM | 2022-06-15 |  | EMS |  | Concentration | pg/L | Y | (fl. conc.) |  | 1,2,3,7,8-PeCDF (fl. conc.) | 57117-41-6 | Dioxins & Furans; | Dioxins & Furans; EMS code: B083 | Chemical |  | LAB
B086 | % | 13C12-1,2,3,4,6,7,8-HpCDD | SURR | Surrogate Recovery | Surrogate Recovery | 1 | Y | Y | Y | 6 | 3/20/2020 | 2/3/2022 | 2/21/2023 2:10:21 PM | 2022-06-15 |  | EMS |  | DimensionlessRatio | % | N | (ratio) |  | 13C12-1,2,3,4,6,7,8-HpCDD (ratio) | 109719-83-7 | Dioxins & Furans; | Dioxins & Furans; EMS code: B086 | Chemical |  | SURROGATE_RESULT
B087 | % | 13C12-1,2,3,4,6,7,8-HpCDF | SURR | Surrogate Recovery | Surrogate Recovery | 1 | Y | Y | Y | 6 | 3/20/2020 | 2/3/2022 | 2/21/2023 2:10:21 PM | 2022-06-15 |  | EMS |  | DimensionlessRatio | % | N | (ratio) |  | 13C12-1,2,3,4,6,7,8-HpCDF (ratio) | 109719-84-8 | Dioxins & Furans; | Dioxins & Furans; EMS code: B087 | Chemical |  | SURROGATE_RESULT
B088 | % | 13C12-1,2,3,4,7,8,9-HpCDF | SURR | Surrogate Recovery | Surrogate Recovery | 1 | Y | Y | Y | 6 | 3/20/2020 | 2/3/2022 | 2/21/2023 2:10:21 PM | 2022-06-15 |  | EMS |  | DimensionlessRatio | % | N | (ratio) |  | 13C12-1,2,3,4,7,8,9-HpCDF (ratio) | 109719-94-0 | Dioxins & Furans; | Dioxins & Furans; EMS code: B088 | Chemical |  | SURROGATE_RESULT
B089 | % | 13C12-1,2,3,4,7,8-HxCDD | SURR | Surrogate Recovery | Surrogate Recovery | 1 | Y | Y | Y | 6 | 3/20/2020 | 2/3/2022 | 2/21/2023 2:10:21 PM | 2022-06-15 |  | EMS |  | DimensionlessRatio | % | N | (ratio) |  | 13C12-1,2,3,4,7,8-HxCDD (ratio) | 109719-80-4 | Dioxins & Furans; | Dioxins & Furans; EMS code: B089 | Chemical |  | SURROGATE_RESULT
B090 | % | 13C12-1,2,3,4,7,8-HxCDF | SURR | Surrogate Recovery | Surrogate Recovery | 1 | Y | Y | Y | 6 | 3/20/2020 | 2/3/2022 | 2/21/2023 2:10:21 PM | 2022-06-15 |  | EMS |  | DimensionlessRatio | % | N | (ratio) |  | 13C12-1,2,3,4,7,8-HxCDF (ratio) | 114423-98-2 | Dioxins & Furans; | Dioxins & Furans; EMS code: B090 | Chemical |  | SURROGATE_RESULT
B091 | % | 13C12-1,2,3,6,7,8-HxCDD | SURR | Surrogate Recovery | Surrogate Recovery | 1 | Y | Y | Y | 6 | 3/20/2020 | 2/3/2022 | 2/21/2023 2:10:21 PM | 2022-06-15 |  | EMS |  | DimensionlessRatio | % | N | (ratio) |  | 13C12-1,2,3,6,7,8-HxCDD (ratio) | 109719-81-5 | Dioxins & Furans; | Dioxins & Furans; EMS code: B091 | Chemical |  | SURROGATE_RESULT
B092 | % | 13C12-1,2,3,6,7,8-HxCDF | SURR | Surrogate Recovery | Surrogate Recovery | 1 | Y | Y | Y | 6 | 3/20/2020 | 2/3/2022 | 2/21/2023 2:10:21 PM | 2022-06-15 |  | EMS |  | DimensionlessRatio | % | N | (ratio) |  | 13C12-1,2,3,6,7,8-HxCDF (ratio) | 116843-03-9 | Dioxins & Furans; | Dioxins & Furans; EMS code: B092 | Chemical |  | SURROGATE_RESULT
B093 | % | 13C12-1,2,3,7,8,9-HxCDF | SURR | Surrogate Recovery | Surrogate Recovery | 1 | Y | Y | Y | 6 | 3/20/2020 | 2/3/2022 | 2/21/2023 2:10:21 PM | 2022-06-15 |  | EMS |  | DimensionlessRatio | % | N | (ratio) |  | 13C12-1,2,3,7,8,9-HxCDF (ratio) | 116843-04-0 | Dioxins & Furans; | Dioxins & Furans; EMS code: B093 | Chemical |  | SURROGATE_RESULT
B094 | % | 13C12-1,2,3,7,8-PeCDD | SURR | Surrogate Recovery | Surrogate Recovery | 1 | Y | Y | Y | 6 | 3/20/2020 | 2/3/2022 | 2/21/2023 2:10:21 PM | 2022-06-15 |  | EMS |  | DimensionlessRatio | % | N | (ratio) |  | 13C12-1,2,3,7,8-PeCDD (ratio) | 109719-79-1 | Dioxins & Furans; | Dioxins & Furans; EMS code: B094 | Chemical |  | SURROGATE_RESULT
B095 | % | 13C12-1,2,3,7,8-PeCDF | SURR | Surrogate Recovery | Surrogate Recovery | 1 | Y | Y | Y | 6 | 3/20/2020 | 2/3/2022 | 2/21/2023 2:10:21 PM | 2022-06-15 |  | EMS |  | DimensionlessRatio | % | N | (ratio) |  | 13C12-1,2,3,7,8-PeCDF (ratio) | 109719-77-9 | Dioxins & Furans; | Dioxins & Furans; EMS code: B095 | Chemical |  | SURROGATE_RESULT
B096 | % | 13C12-2,3,4,6,7,8-HxCDF | SURR | Surrogate Recovery | Surrogate Recovery | 1 | Y | Y | Y | 6 | 3/20/2020 | 2/3/2022 | 2/21/2023 2:10:21 PM | 2022-06-15 |  | EMS |  | DimensionlessRatio | % | N | (ratio) |  | 13C12-2,3,4,6,7,8-HxCDF (ratio) | 116843-05-1 | Dioxins & Furans; | Dioxins & Furans; EMS code: B096 | Chemical |  | SURROGATE_RESULT
B097 | % | 13C12-2,3,4,7,8-PeCDF | SURR | Surrogate Recovery | Surrogate Recovery | 1 | Y | Y | Y | 6 | 3/20/2020 | 2/3/2022 | 2/21/2023 2:10:21 PM | 2022-06-15 |  | EMS |  | DimensionlessRatio | % | N | (ratio) |  | 13C12-2,3,4,7,8-PeCDF (ratio) | 116843-02-8 | Dioxins & Furans; | Dioxins & Furans; EMS code: B097 | Chemical |  | SURROGATE_RESULT
B098 | % | 13C12-2,3,7,8-TCDD | SURR | Surrogate Recovery | Surrogate Recovery | 1 | Y | Y | Y | 6 | 3/20/2020 | 2/3/2022 | 2/21/2023 2:10:21 PM | 2022-06-15 |  | EMS |  | DimensionlessRatio | % | N | (ratio) |  | 13C12-2,3,7,8-TCDD (ratio) | 76523-40-5 | Dioxins & Furans; | Dioxins & Furans; EMS code: B098 | Chemical |  | SURROGATE_RESULT
B099 | % | 13C12-2,3,7,8-TCDF | SURR | Surrogate Recovery | Surrogate Recovery | 1 | Y | Y | Y | 6 | 3/20/2020 | 2/3/2022 | 2/21/2023 2:10:21 PM | 2022-06-15 |  | EMS |  | DimensionlessRatio | % | N | (ratio) |  | 13C12-2,3,7,8-TCDF (ratio) | 89059-46-1 | Dioxins & Furans; | Dioxins & Furans; EMS code: B099 | Chemical |  | SURROGATE_RESULT
B101 | % | 13C12-OCDD | SURR | Surrogate Recovery | Surrogate Recovery | 1 | Y | Y | Y | 6 | 3/20/2020 | 2/3/2022 | 2/21/2023 2:10:21 PM | 2022-06-15 |  | EMS |  | DimensionlessRatio | % | N | (ratio) |  | 13C12-OCDD (ratio) | 114423-97-1 | Dioxins & Furans; | Dioxins & Furans; EMS code: B101 | Chemical |  | SURROGATE_RESULT
B102 | pg/g | 2,3,4,6,7,8-HxCDF | 1613B | USEPA 1613B | USEPA 1613B | 0.05 | Y | Y | Y | 4 | 3/20/2020 | 3/21/2020 | 2/21/2023 2:10:21 PM | 2023-02-06 |  | EMS |  | MassPerMass | pg/g | Y | (mass conc.) |  | 2,3,4,6,7,8-HxCDF (mass conc.) | 60851-34-5 | Dioxins & Furans; | Dioxins & Furans; EMS code: B102 | Chemical |  | LAB
B102 | pg/g wet | 2,3,4,6,7,8-HxCDF | 1101 | DX-1613B | DCM Extraction, Column clean-up, isotope dilution GC/HRMS. Method Reference: USEPA 1613B. Methodology: Samples are filtered if required. Solids are extracted by Soxhlet using toluene. The liquid portion is extracted by liquid/liquid extraction using dichloromethane. The extracts are prepared using column chromatography, reduced in volume and analyzed by isotope-dilution GC/HRMS. | 0.028 | Y | Y | Y | 1 | 2/2/2022 | 2/2/2022 | 7/6/2022 5:20:43 PM | 2019-06-04 |  | KATFRASE |  | MassPerMass | pg/g | Y | (mass conc.) | (wet) | 2,3,4,6,7,8-HxCDF (wet) (mass conc.) | 60851-34-5 | Dioxins & Furans; | Dioxins & Furans; EMS code: B102 | Chemical |  | LAB
B102 | pg/L | 2,3,4,6,7,8-HxCDF | E1613B | EPA 1613B | EPA 1613B | 0.58 | Y | Y | Y | 1 | 2/3/2022 | 2/3/2022 | 7/6/2022 4:59:13 PM | 2022-06-15 |  | EMS |  | Concentration | pg/L | Y | (fl. conc.) |  | 2,3,4,6,7,8-HxCDF (fl. conc.) | 60851-34-5 | Dioxins & Furans; | Dioxins & Furans; EMS code: B102 | Chemical |  | LAB
B105 | pg/g | 2,3,4,7,8-PeCDF | 1613B | USEPA 1613B | USEPA 1613B | 0.11 | Y | Y | Y | 4 | 3/20/2020 | 3/21/2020 | 2/21/2023 2:10:21 PM | 2023-02-06 |  | EMS |  | MassPerMass | pg/g | Y | (mass conc.) |  | 2,3,4,7,8-PeCDF (mass conc.) | 57117-31-4 | Dioxins & Furans; | Dioxins & Furans; EMS code: B105 | Chemical |  | LAB
B105 | pg/g wet | 2,3,4,7,8-PeCDF | 1101 | DX-1613B | DCM Extraction, Column clean-up, isotope dilution GC/HRMS. Method Reference: USEPA 1613B. Methodology: Samples are filtered if required. Solids are extracted by Soxhlet using toluene. The liquid portion is extracted by liquid/liquid extraction using dichloromethane. The extracts are prepared using column chromatography, reduced in volume and analyzed by isotope-dilution GC/HRMS. | 0.03 | Y | Y | Y | 1 | 2/2/2022 | 2/2/2022 | 7/6/2022 5:20:43 PM | 2019-06-04 |  | KATFRASE |  | MassPerMass | pg/g | Y | (mass conc.) | (wet) | 2,3,4,7,8-PeCDF (wet) (mass conc.) | 57117-31-4 | Dioxins & Furans; | Dioxins & Furans; EMS code: B105 | Chemical |  | LAB
B105 | pg/L | 2,3,4,7,8-PeCDF | E1613B | EPA 1613B | EPA 1613B | 0.39 | Y | Y | Y | 1 | 2/3/2022 | 2/3/2022 | 7/6/2022 4:59:13 PM | 2022-06-15 |  | EMS |  | Concentration | pg/L | Y | (fl. conc.) |  | 2,3,4,7,8-PeCDF (fl. conc.) | 57117-31-4 | Dioxins & Furans; | Dioxins & Furans; EMS code: B105 | Chemical |  | LAB
B108 | pg/g wet | 2,3,7,8-TCDD | 1101 | DX-1613B | DCM Extraction, Column clean-up, isotope dilution GC/HRMS. Method Reference: USEPA 1613B. Methodology: Samples are filtered if required. Solids are extracted by Soxhlet using toluene. The liquid portion is extracted by liquid/liquid extraction using dichloromethane. The extracts are prepared using column chromatography, reduced in volume and analyzed by isotope-dilution GC/HRMS. | 0.075 | Y | Y | Y | 1 | 2/2/2022 | 2/2/2022 | 7/6/2022 5:20:43 PM | 2019-06-04 |  | KATFRASE |  | MassPerMass | pg/g | Y | (mass conc.) | (wet) | 2,3,7,8-TCDD (wet) (mass conc.) | 1746-01-6 | Dioxins & Furans; | Dioxins & Furans; EMS code: B108 | Chemical |  | LAB
B108 | pg/L | 2,3,7,8-TCDD | E1613B | EPA 1613B | EPA 1613B | 0.51 | Y | Y | Y | 1 | 2/3/2022 | 2/3/2022 | 7/6/2022 4:59:13 PM | 2022-06-15 |  | EMS |  | Concentration | pg/L | Y | (fl. conc.) |  | 2,3,7,8-TCDD (fl. conc.) | 1746-01-6 | Dioxins & Furans; | Dioxins & Furans; EMS code: B108 | Chemical |  | LAB
B108 | pg/g | 2,3,7,8-TCDD | 1613B | USEPA 1613B | USEPA 1613B | 0.07 | Y | Y | Y | 4 | 3/20/2020 | 3/21/2020 | 2/21/2023 2:10:21 PM | 2023-02-06 |  | EMS |  | MassPerMass | pg/g | Y | (mass conc.) |  | 2,3,7,8-TCDD (mass conc.) | 1746-01-6 | Dioxins & Furans; | Dioxins & Furans; EMS code: B108 | Chemical |  | LAB
B111 | pg/g wet | 2,3,7,8-TCDF | 1101 | DX-1613B | DCM Extraction, Column clean-up, isotope dilution GC/HRMS. Method Reference: USEPA 1613B. Methodology: Samples are filtered if required. Solids are extracted by Soxhlet using toluene. The liquid portion is extracted by liquid/liquid extraction using dichloromethane. The extracts are prepared using column chromatography, reduced in volume and analyzed by isotope-dilution GC/HRMS. | 0.08 | Y | Y | Y | 1 | 2/2/2022 | 2/2/2022 | 7/6/2022 5:20:43 PM | 2019-06-04 |  | KATFRASE |  | MassPerMass | pg/g | Y | (mass conc.) | (wet) | 2,3,7,8-TCDF (wet) (mass conc.) | 51207-31-9 | Dioxins & Furans; | Dioxins & Furans; EMS code: B111 | Chemical |  | LAB
B111 | pg/L | 2,3,7,8-TCDF | E1613B | EPA 1613B | EPA 1613B | 0.31 | Y | Y | Y | 1 | 2/3/2022 | 2/3/2022 | 7/6/2022 4:59:13 PM | 2022-06-15 |  | EMS |  | Concentration | pg/L | Y | (fl. conc.) |  | 2,3,7,8-TCDF (fl. conc.) | 51207-31-9 | Dioxins & Furans; | Dioxins & Furans; EMS code: B111 | Chemical |  | LAB
B111 | pg/g | 2,3,7,8-TCDF | 1613B | USEPA 1613B | USEPA 1613B | 0.061 | Y | Y | Y | 4 | 3/20/2020 | 3/21/2020 | 2/21/2023 2:10:21 PM | 2023-02-06 |  | EMS |  | MassPerMass | pg/g | Y | (mass conc.) |  | 2,3,7,8-TCDF (mass conc.) | 51207-31-9 | Dioxins & Furans; | Dioxins & Furans; EMS code: B111 | Chemical |  | LAB
B114 | % | 37Cl4-2,3,7,8-TCDD (Cleanup) | SURR | Surrogate Recovery | Surrogate Recovery |  | Y | Y | Y | 6 | 3/20/2020 | 2/3/2022 | 2/21/2023 2:10:21 PM | 2022-07-06 |  | EMS |  | DimensionlessRatio | % | N | (ratio) |  | 37Cl4-2,3,7,8-TCDD (Cleanup) (ratio) | 85508-50-5 | Dioxins & Furans; | Dioxins & Furans; EMS code: B114 | Chemical |  | SURROGATE_RESULT
B115 | pg/L | Lower Bound PCDD/F TEQ (WHO 2005) | E1613B | EPA 1613B | EPA 1613B |  | Y | Y | Y | 1 | 2/3/2022 | 2/3/2022 | 7/6/2022 4:59:13 PM | 2022-06-15 |  | EMS |  | Concentration | pg/L | Y | (fl. conc.) |  | Lower Bound PCDD/F TEQ (WHO 2005) (fl. conc.) |  |  | EMS code: B115 | Chemical |  | LAB
B115 | pg/g wet | Lower Bound PCDD/F TEQ (WHO 2005) | 1101 | DX-1613B | DCM Extraction, Column clean-up, isotope dilution GC/HRMS. Method Reference: USEPA 1613B. Methodology: Samples are filtered if required. Solids are extracted by Soxhlet using toluene. The liquid portion is extracted by liquid/liquid extraction using dichloromethane. The extracts are prepared using column chromatography, reduced in volume and analyzed by isotope-dilution GC/HRMS. |  | Y | Y | Y | 1 | 2/2/2022 | 2/2/2022 | 7/6/2022 5:20:43 PM | 2019-06-04 |  | KATFRASE |  | MassPerMass | pg/g | Y | (mass conc.) | (wet) | Lower Bound PCDD/F TEQ (WHO 2005) (wet) (mass conc.) |  |  | EMS code: B115 | Chemical |  | LAB
B115 | pg/g | Lower Bound PCDD/F TEQ (WHO 2005) | 1613B | USEPA 1613B | USEPA 1613B |  | Y | Y | Y | 4 | 3/20/2020 | 3/21/2020 | 2/21/2023 2:10:21 PM | 2023-02-06 |  | EMS |  | MassPerMass | pg/g | Y | (mass conc.) |  | Lower Bound PCDD/F TEQ (WHO 2005) (mass conc.) |  |  | EMS code: B115 | Chemical |  | LAB
B116 | pg/g | Mid Point PCDD/F TEQ (WHO 2005) | 1613B | USEPA 1613B | USEPA 1613B |  | Y | Y | Y | 4 | 3/20/2020 | 3/21/2020 | 2/21/2023 2:10:22 PM | 2023-02-06 |  | EMS |  | MassPerMass | pg/g | Y | (mass conc.) |  | Mid Point PCDD/F TEQ (WHO 2005) (mass conc.) |  |  | EMS code: B116 | Chemical |  | LAB
B116 | pg/L | Mid Point PCDD/F TEQ (WHO 2005) | E1613B | EPA 1613B | EPA 1613B |  | Y | Y | Y | 1 | 2/3/2022 | 2/3/2022 | 7/6/2022 4:59:13 PM | 2022-06-15 |  | EMS |  | Concentration | pg/L | Y | (fl. conc.) |  | Mid Point PCDD/F TEQ (WHO 2005) (fl. conc.) |  |  | EMS code: B116 | Chemical |  | LAB
B116 | pg/g wet | Mid Point PCDD/F TEQ (WHO 2005) | 1101 | DX-1613B | DCM Extraction, Column clean-up, isotope dilution GC/HRMS. Method Reference: USEPA 1613B. Methodology: Samples are filtered if required. Solids are extracted by Soxhlet using toluene. The liquid portion is extracted by liquid/liquid extraction using dichloromethane. The extracts are prepared using column chromatography, reduced in volume and analyzed by isotope-dilution GC/HRMS. |  | Y | Y | Y | 1 | 2/2/2022 | 2/2/2022 | 7/6/2022 5:20:43 PM | 2019-06-04 |  | KATFRASE |  | MassPerMass | pg/g | Y | (mass conc.) | (wet) | Mid Point PCDD/F TEQ (WHO 2005) (wet) (mass conc.) |  |  | EMS code: B116 | Chemical |  | LAB
B117 | pg/g | OCDD | 1613B | USEPA 1613B | USEPA 1613B | 0.13 | Y | Y | Y | 4 | 3/20/2020 | 3/21/2020 | 2/21/2023 2:10:22 PM | 2023-02-06 |  | EMS |  | MassPerMass | pg/g | Y | (mass conc.) |  | OCDD (mass conc.) | 3268-87-9 | Dioxins & Furans; | Dioxins & Furans; EMS code: B117 | Chemical |  | LAB
B117 | pg/g wet | OCDD | 1101 | DX-1613B | DCM Extraction, Column clean-up, isotope dilution GC/HRMS. Method Reference: USEPA 1613B. Methodology: Samples are filtered if required. Solids are extracted by Soxhlet using toluene. The liquid portion is extracted by liquid/liquid extraction using dichloromethane. The extracts are prepared using column chromatography, reduced in volume and analyzed by isotope-dilution GC/HRMS. | 0.034 | Y | Y | Y | 1 | 2/2/2022 | 2/2/2022 | 7/6/2022 5:20:43 PM | 2019-06-04 |  | KATFRASE |  | MassPerMass | pg/g | Y | (mass conc.) | (wet) | OCDD (wet) (mass conc.) | 3268-87-9 | Dioxins & Furans; | Dioxins & Furans; EMS code: B117 | Chemical |  | LAB
B117 | pg/L | OCDD | E1613B | EPA 1613B | EPA 1613B | 2.6 | Y | Y | Y | 1 | 2/3/2022 | 2/3/2022 | 7/6/2022 4:59:14 PM | 2022-06-15 |  | EMS |  | Concentration | pg/L | Y | (fl. conc.) |  | OCDD (fl. conc.) | 3268-87-9 | Dioxins & Furans; | Dioxins & Furans; EMS code: B117 | Chemical |  | LAB
B120 | pg/g | OCDF | 1613B | USEPA 1613B | USEPA 1613B | 0.054 | Y | Y | Y | 4 | 3/20/2020 | 3/21/2020 | 2/21/2023 2:10:22 PM | 2023-02-06 |  | EMS |  | MassPerMass | pg/g | Y | (mass conc.) |  | OCDF (mass conc.) | 39001-02-0 | Dioxins & Furans; | Dioxins & Furans; EMS code: B120 | Chemical |  | LAB
B120 | pg/g wet | OCDF | 1101 | DX-1613B | DCM Extraction, Column clean-up, isotope dilution GC/HRMS. Method Reference: USEPA 1613B. Methodology: Samples are filtered if required. Solids are extracted by Soxhlet using toluene. The liquid portion is extracted by liquid/liquid extraction using dichloromethane. The extracts are prepared using column chromatography, reduced in volume and analyzed by isotope-dilution GC/HRMS. | 0.038 | Y | Y | Y | 1 | 2/2/2022 | 2/2/2022 | 7/6/2022 5:20:43 PM | 2019-06-04 |  | KATFRASE |  | MassPerMass | pg/g | Y | (mass conc.) | (wet) | OCDF (wet) (mass conc.) | 39001-02-0 | Dioxins & Furans; | Dioxins & Furans; EMS code: B120 | Chemical |  | LAB
B120 | pg/L | OCDF | E1613B | EPA 1613B | EPA 1613B | 1.4 | Y | Y | Y | 1 | 2/3/2022 | 2/3/2022 | 7/6/2022 4:59:14 PM | 2022-06-15 |  | EMS |  | Concentration | pg/L | Y | (fl. conc.) |  | OCDF (fl. conc.) | 39001-02-0 | Dioxins & Furans; | Dioxins & Furans; EMS code: B120 | Chemical |  | LAB
B123 | None | Total HpCDD # Homologues | 1101 | DX-1613B | DCM Extraction, Column clean-up, isotope dilution GC/HRMS. Method Reference: USEPA 1613B. Methodology: Samples are filtered if required. Solids are extracted by Soxhlet using toluene. The liquid portion is extracted by liquid/liquid extraction using dichloromethane. The extracts are prepared using column chromatography, reduced in volume and analyzed by isotope-dilution GC/HRMS. |  | Y | Y | Y | 6 | 3/20/2020 | 2/3/2022 | 2/21/2023 2:10:22 PM | 2019-06-04 |  | KATFRASE |  | None | None | N | (unitless) |  | Total HpCDD # Homologues (unitless) | 37871-00-4 | Dioxins & Furans; | Dioxins & Furans; EMS code: B123 | Chemical | Total | LAB
B124 | None | Total HpCDF # Homologues | 1101 | DX-1613B | DCM Extraction, Column clean-up, isotope dilution GC/HRMS. Method Reference: USEPA 1613B. Methodology: Samples are filtered if required. Solids are extracted by Soxhlet using toluene. The liquid portion is extracted by liquid/liquid extraction using dichloromethane. The extracts are prepared using column chromatography, reduced in volume and analyzed by isotope-dilution GC/HRMS. |  | Y | Y | Y | 6 | 3/20/2020 | 2/3/2022 | 2/21/2023 2:10:22 PM | 2019-06-04 |  | KATFRASE |  | None | None | N | (unitless) |  | Total HpCDF # Homologues (unitless) | 38998-75-3 | Dioxins & Furans; | Dioxins & Furans; EMS code: B124 | Chemical | Total | LAB
B125 | None | Total HxCDD # Homologues | 1101 | DX-1613B | DCM Extraction, Column clean-up, isotope dilution GC/HRMS. Method Reference: USEPA 1613B. Methodology: Samples are filtered if required. Solids are extracted by Soxhlet using toluene. The liquid portion is extracted by liquid/liquid extraction using dichloromethane. The extracts are prepared using column chromatography, reduced in volume and analyzed by isotope-dilution GC/HRMS. |  | Y | Y | Y | 6 | 3/20/2020 | 2/3/2022 | 2/21/2023 2:10:22 PM | 2019-06-04 |  | KATFRASE |  | None | None | N | (unitless) |  | Total HxCDD # Homologues (unitless) | 34465-46-8 | Dioxins & Furans; | Dioxins & Furans; EMS code: B125 | Chemical | Total | LAB
B126 | None | Total HxCDF # Homologues | 1101 | DX-1613B | DCM Extraction, Column clean-up, isotope dilution GC/HRMS. Method Reference: USEPA 1613B. Methodology: Samples are filtered if required. Solids are extracted by Soxhlet using toluene. The liquid portion is extracted by liquid/liquid extraction using dichloromethane. The extracts are prepared using column chromatography, reduced in volume and analyzed by isotope-dilution GC/HRMS. |  | Y | Y | Y | 6 | 3/20/2020 | 2/3/2022 | 2/21/2023 2:10:22 PM | 2019-06-04 |  | KATFRASE |  | None | None | N | (unitless) |  | Total HxCDF # Homologues (unitless) | 55684-94-1 | Dioxins & Furans; | Dioxins & Furans; EMS code: B126 | Chemical | Total | LAB
B127 | None | Total PeCDD # Homologues | 1101 | DX-1613B | DCM Extraction, Column clean-up, isotope dilution GC/HRMS. Method Reference: USEPA 1613B. Methodology: Samples are filtered if required. Solids are extracted by Soxhlet using toluene. The liquid portion is extracted by liquid/liquid extraction using dichloromethane. The extracts are prepared using column chromatography, reduced in volume and analyzed by isotope-dilution GC/HRMS. |  | Y | Y | Y | 6 | 3/20/2020 | 2/3/2022 | 2/21/2023 2:10:22 PM | 2019-06-04 |  | KATFRASE |  | None | None | N | (unitless) |  | Total PeCDD # Homologues (unitless) | 36088-22-9 | Dioxins & Furans; | Dioxins & Furans; EMS code: B127 | Chemical | Total | LAB
B128 | None | Total PeCDF # Homologues | 1101 | DX-1613B | DCM Extraction, Column clean-up, isotope dilution GC/HRMS. Method Reference: USEPA 1613B. Methodology: Samples are filtered if required. Solids are extracted by Soxhlet using toluene. The liquid portion is extracted by liquid/liquid extraction using dichloromethane. The extracts are prepared using column chromatography, reduced in volume and analyzed by isotope-dilution GC/HRMS. |  | Y | Y | Y | 6 | 3/20/2020 | 2/3/2022 | 2/21/2023 2:10:22 PM | 2019-06-04 |  | KATFRASE |  | None | None | N | (unitless) |  | Total PeCDF # Homologues (unitless) | 30402-15-4 | Dioxins & Furans; | Dioxins & Furans; EMS code: B128 | Chemical | Total | LAB
B129 | None | Total TCDD # Homologues | 1101 | DX-1613B | DCM Extraction, Column clean-up, isotope dilution GC/HRMS. Method Reference: USEPA 1613B. Methodology: Samples are filtered if required. Solids are extracted by Soxhlet using toluene. The liquid portion is extracted by liquid/liquid extraction using dichloromethane. The extracts are prepared using column chromatography, reduced in volume and analyzed by isotope-dilution GC/HRMS. |  | Y | Y | Y | 6 | 3/20/2020 | 2/3/2022 | 2/21/2023 2:10:22 PM | 2019-06-04 |  | KATFRASE |  | None | None | N | (unitless) |  | Total TCDD # Homologues (unitless) | 41903-57-5 | Dioxins & Furans; | Dioxins & Furans; EMS code: B129 | Chemical | Total | LAB
B130 | None | Total TCDF # Homologues | 1101 | DX-1613B | DCM Extraction, Column clean-up, isotope dilution GC/HRMS. Method Reference: USEPA 1613B. Methodology: Samples are filtered if required. Solids are extracted by Soxhlet using toluene. The liquid portion is extracted by liquid/liquid extraction using dichloromethane. The extracts are prepared using column chromatography, reduced in volume and analyzed by isotope-dilution GC/HRMS. |  | Y | Y | Y | 6 | 3/20/2020 | 2/3/2022 | 2/21/2023 2:10:22 PM | 2019-06-04 |  | KATFRASE |  | None | None | N | (unitless) |  | Total TCDF # Homologues (unitless) | 55722-27-5 | Dioxins & Furans; | Dioxins & Furans; EMS code: B130 | Chemical | Total | LAB
B131 | pg/g wet | Total-HpCDD | 1101 | DX-1613B | DCM Extraction, Column clean-up, isotope dilution GC/HRMS. Method Reference: USEPA 1613B. Methodology: Samples are filtered if required. Solids are extracted by Soxhlet using toluene. The liquid portion is extracted by liquid/liquid extraction using dichloromethane. The extracts are prepared using column chromatography, reduced in volume and analyzed by isotope-dilution GC/HRMS. | 0.038 | Y | Y | Y | 1 | 2/2/2022 | 2/2/2022 | 7/6/2022 5:20:43 PM | 2019-06-04 |  | KATFRASE |  | MassPerMass | pg/g | Y | (mass conc.) | (wet) | Total-HpCDD (wet) (mass conc.) | 37871-00-4 | Dioxins & Furans; | Dioxins & Furans; EMS code: B131 | Chemical |  | LAB
B131 | pg/g | Total-HpCDD | 1613B | USEPA 1613B | USEPA 1613B | 0.17 | Y | Y | Y | 4 | 3/20/2020 | 3/21/2020 | 2/21/2023 2:10:22 PM | 2023-02-06 |  | EMS |  | MassPerMass | pg/g | Y | (mass conc.) |  | Total-HpCDD (mass conc.) | 37871-00-4 | Dioxins & Furans; | Dioxins & Furans; EMS code: B131 | Chemical |  | LAB
B131 | pg/L | Total-HpCDD | E1613B | EPA 1613B | EPA 1613B | 1.4 | Y | Y | Y | 1 | 2/3/2022 | 2/3/2022 | 7/6/2022 4:59:14 PM | 2022-06-15 |  | EMS |  | Concentration | pg/L | Y | (fl. conc.) |  | Total-HpCDD (fl. conc.) | 37871-00-4 | Dioxins & Furans; | Dioxins & Furans; EMS code: B131 | Chemical |  | LAB
B134 | pg/g | Total-HpCDF | 1613B | USEPA 1613B | USEPA 1613B | 0.069 | Y | Y | Y | 4 | 3/20/2020 | 3/21/2020 | 2/21/2023 2:10:22 PM | 2023-02-06 |  | EMS |  | MassPerMass | pg/g | Y | (mass conc.) |  | Total-HpCDF (mass conc.) | 38998-75-3 | Dioxins & Furans; | Dioxins & Furans; EMS code: B134 | Chemical | Total | LAB
B134 | pg/L | Total-HpCDF | E1613B | EPA 1613B | EPA 1613B | 1.2 | Y | Y | Y | 1 | 2/3/2022 | 2/3/2022 | 7/6/2022 4:59:14 PM | 2022-06-15 |  | EMS |  | Concentration | pg/L | Y | (fl. conc.) |  | Total-HpCDF (fl. conc.) | 38998-75-3 | Dioxins & Furans; | Dioxins & Furans; EMS code: B134 | Chemical | Total | LAB
B134 | pg/g wet | Total-HpCDF | 1101 | DX-1613B | DCM Extraction, Column clean-up, isotope dilution GC/HRMS. Method Reference: USEPA 1613B. Methodology: Samples are filtered if required. Solids are extracted by Soxhlet using toluene. The liquid portion is extracted by liquid/liquid extraction using dichloromethane. The extracts are prepared using column chromatography, reduced in volume and analyzed by isotope-dilution GC/HRMS. | 0.033 | Y | Y | Y | 1 | 2/2/2022 | 2/2/2022 | 7/6/2022 5:20:43 PM | 2019-06-04 |  | KATFRASE |  | MassPerMass | pg/g | Y | (mass conc.) | (wet) | Total-HpCDF (wet) (mass conc.) | 38998-75-3 | Dioxins & Furans; | Dioxins & Furans; EMS code: B134 | Chemical | Total | LAB
B137 | pg/L | Total-HxCDD | E1613B | EPA 1613B | EPA 1613B | 0.6 | Y | Y | Y | 1 | 2/3/2022 | 2/3/2022 | 7/6/2022 4:59:14 PM | 2022-06-15 |  | EMS |  | Concentration | pg/L | Y | (fl. conc.) |  | Total-HxCDD (fl. conc.) | 34465-46-8 | Dioxins & Furans; | Dioxins & Furans; EMS code: B137 | Chemical |  | LAB
B137 | pg/g wet | Total-HxCDD | 1101 | DX-1613B | DCM Extraction, Column clean-up, isotope dilution GC/HRMS. Method Reference: USEPA 1613B. Methodology: Samples are filtered if required. Solids are extracted by Soxhlet using toluene. The liquid portion is extracted by liquid/liquid extraction using dichloromethane. The extracts are prepared using column chromatography, reduced in volume and analyzed by isotope-dilution GC/HRMS. | 0.044 | Y | Y | Y | 1 | 2/2/2022 | 2/2/2022 | 7/6/2022 5:20:43 PM | 2019-06-04 |  | KATFRASE |  | MassPerMass | pg/g | Y | (mass conc.) | (wet) | Total-HxCDD (wet) (mass conc.) | 34465-46-8 | Dioxins & Furans; | Dioxins & Furans; EMS code: B137 | Chemical |  | LAB
B137 | pg/g | Total-HxCDD | 1613B | USEPA 1613B | USEPA 1613B | 0.099 | Y | Y | Y | 4 | 3/20/2020 | 3/21/2020 | 2/21/2023 2:10:22 PM | 2023-02-06 |  | EMS |  | MassPerMass | pg/g | Y | (mass conc.) |  | Total-HxCDD (mass conc.) | 34465-46-8 | Dioxins & Furans; | Dioxins & Furans; EMS code: B137 | Chemical |  | LAB
B140 | pg/g wet | Total-HxCDF | 1101 | DX-1613B | DCM Extraction, Column clean-up, isotope dilution GC/HRMS. Method Reference: USEPA 1613B. Methodology: Samples are filtered if required. Solids are extracted by Soxhlet using toluene. The liquid portion is extracted by liquid/liquid extraction using dichloromethane. The extracts are prepared using column chromatography, reduced in volume and analyzed by isotope-dilution GC/HRMS. | 0.034 | Y | Y | Y | 1 | 2/2/2022 | 2/2/2022 | 7/6/2022 5:20:43 PM | 2019-06-04 |  | KATFRASE |  | MassPerMass | pg/g | Y | (mass conc.) | (wet) | Total-HxCDF (wet) (mass conc.) | 55684-94-1 | Dioxins & Furans; | Dioxins & Furans; EMS code: B140 | Chemical |  | LAB
B140 | pg/L | Total-HxCDF | E1613B | EPA 1613B | EPA 1613B | 0.93 | Y | Y | Y | 1 | 2/3/2022 | 2/3/2022 | 7/6/2022 4:59:14 PM | 2022-06-15 |  | EMS |  | Concentration | pg/L | Y | (fl. conc.) |  | Total-HxCDF (fl. conc.) | 55684-94-1 | Dioxins & Furans; | Dioxins & Furans; EMS code: B140 | Chemical |  | LAB
B140 | pg/g | Total-HxCDF | 1613B | USEPA 1613B | USEPA 1613B | 0.058 | Y | Y | Y | 4 | 3/20/2020 | 3/21/2020 | 2/21/2023 2:10:22 PM | 2023-02-06 |  | EMS |  | MassPerMass | pg/g | Y | (mass conc.) |  | Total-HxCDF (mass conc.) | 55684-94-1 | Dioxins & Furans; | Dioxins & Furans; EMS code: B140 | Chemical |  | LAB
B143 | pg/L | Total-PeCDD | E1613B | EPA 1613B | EPA 1613B | 0.42 | Y | Y | Y | 1 | 2/3/2022 | 2/3/2022 | 7/6/2022 4:59:14 PM | 2022-06-15 |  | EMS |  | Concentration | pg/L | Y | (fl. conc.) |  | Total-PeCDD (fl. conc.) | 36088-22-9 | Dioxins & Furans; | Dioxins & Furans; EMS code: B143 | Chemical | Total | LAB
B143 | pg/g wet | Total-PeCDD | 1101 | DX-1613B | DCM Extraction, Column clean-up, isotope dilution GC/HRMS. Method Reference: USEPA 1613B. Methodology: Samples are filtered if required. Solids are extracted by Soxhlet using toluene. The liquid portion is extracted by liquid/liquid extraction using dichloromethane. The extracts are prepared using column chromatography, reduced in volume and analyzed by isotope-dilution GC/HRMS. | 0.04 | Y | Y | Y | 1 | 2/2/2022 | 2/2/2022 | 7/6/2022 5:20:43 PM | 2019-06-04 |  | KATFRASE |  | MassPerMass | pg/g | Y | (mass conc.) | (wet) | Total-PeCDD (wet) (mass conc.) | 36088-22-9 | Dioxins & Furans; | Dioxins & Furans; EMS code: B143 | Chemical | Total | LAB
B143 | pg/g | Total-PeCDD | 1613B | USEPA 1613B | USEPA 1613B | 0.039 | Y | Y | Y | 4 | 3/20/2020 | 3/21/2020 | 2/21/2023 2:10:22 PM | 2023-02-06 |  | EMS |  | MassPerMass | pg/g | Y | (mass conc.) |  | Total-PeCDD (mass conc.) | 36088-22-9 | Dioxins & Furans; | Dioxins & Furans; EMS code: B143 | Chemical | Total | LAB
B146 | pg/g wet | Total-PeCDF | 1101 | DX-1613B | DCM Extraction, Column clean-up, isotope dilution GC/HRMS. Method Reference: USEPA 1613B. Methodology: Samples are filtered if required. Solids are extracted by Soxhlet using toluene. The liquid portion is extracted by liquid/liquid extraction using dichloromethane. The extracts are prepared using column chromatography, reduced in volume and analyzed by isotope-dilution GC/HRMS. | 0.036 | Y | Y | Y | 1 | 2/2/2022 | 2/2/2022 | 7/6/2022 5:20:43 PM | 2019-06-04 |  | KATFRASE |  | MassPerMass | pg/g | Y | (mass conc.) | (wet) | Total-PeCDF (wet) (mass conc.) | 30402-15-4 | Dioxins & Furans; | Dioxins & Furans; EMS code: B146 | Chemical |  | LAB
B146 | pg/g | Total-PeCDF | 1613B | USEPA 1613B | USEPA 1613B | 0.13 | Y | Y | Y | 4 | 3/20/2020 | 3/21/2020 | 2/21/2023 2:10:22 PM | 2023-02-06 |  | EMS |  | MassPerMass | pg/g | Y | (mass conc.) |  | Total-PeCDF (mass conc.) | 30402-15-4 | Dioxins & Furans; | Dioxins & Furans; EMS code: B146 | Chemical |  | LAB
B146 | pg/L | Total-PeCDF | E1613B | EPA 1613B | EPA 1613B | 0.53 | Y | Y | Y | 1 | 2/3/2022 | 2/3/2022 | 7/6/2022 4:59:14 PM | 2022-06-15 |  | EMS |  | Concentration | pg/L | Y | (fl. conc.) |  | Total-PeCDF (fl. conc.) | 30402-15-4 | Dioxins & Furans; | Dioxins & Furans; EMS code: B146 | Chemical |  | LAB
B149 | pg/L | Total-TCDD | E1613B | EPA 1613B | EPA 1613B | 0.51 | Y | Y | Y | 1 | 2/3/2022 | 2/3/2022 | 7/6/2022 4:59:14 PM | 2022-06-15 |  | EMS |  | Concentration | pg/L | Y | (fl. conc.) |  | Total-TCDD (fl. conc.) | 41903-57-5 | Dioxins & Furans; | Dioxins & Furans; EMS code: B149 | Chemical | Total | LAB
B149 | pg/g | Total-TCDD | 1613B | USEPA 1613B | USEPA 1613B | 0.07 | Y | Y | Y | 4 | 3/20/2020 | 3/21/2020 | 2/21/2023 2:10:22 PM | 2023-02-06 |  | EMS |  | MassPerMass | pg/g | Y | (mass conc.) |  | Total-TCDD (mass conc.) | 41903-57-5 | Dioxins & Furans; | Dioxins & Furans; EMS code: B149 | Chemical | Total | LAB
B149 | pg/g wet | Total-TCDD | 1101 | DX-1613B | DCM Extraction, Column clean-up, isotope dilution GC/HRMS. Method Reference: USEPA 1613B. Methodology: Samples are filtered if required. Solids are extracted by Soxhlet using toluene. The liquid portion is extracted by liquid/liquid extraction using dichloromethane. The extracts are prepared using column chromatography, reduced in volume and analyzed by isotope-dilution GC/HRMS. | 0.075 | Y | Y | Y | 1 | 2/2/2022 | 2/2/2022 | 7/6/2022 5:20:43 PM | 2019-06-04 |  | KATFRASE |  | MassPerMass | pg/g | Y | (mass conc.) | (wet) | Total-TCDD (wet) (mass conc.) | 41903-57-5 | Dioxins & Furans; | Dioxins & Furans; EMS code: B149 | Chemical | Total | LAB
B152 | pg/L | Total-TCDF | E1613B | EPA 1613B | EPA 1613B | 0.31 | Y | Y | Y | 1 | 2/3/2022 | 2/3/2022 | 7/6/2022 4:59:14 PM | 2022-06-15 |  | EMS |  | Concentration | pg/L | Y | (fl. conc.) |  | Total-TCDF (fl. conc.) | 55722-27-5 | Dioxins & Furans; | Dioxins & Furans; EMS code: B152 | Chemical | Total | LAB
B152 | pg/g wet | Total-TCDF | 1101 | DX-1613B | DCM Extraction, Column clean-up, isotope dilution GC/HRMS. Method Reference: USEPA 1613B. Methodology: Samples are filtered if required. Solids are extracted by Soxhlet using toluene. The liquid portion is extracted by liquid/liquid extraction using dichloromethane. The extracts are prepared using column chromatography, reduced in volume and analyzed by isotope-dilution GC/HRMS. | 0.08 | Y | Y | Y | 1 | 2/2/2022 | 2/2/2022 | 7/6/2022 5:20:43 PM | 2019-06-04 |  | KATFRASE |  | MassPerMass | pg/g | Y | (mass conc.) | (wet) | Total-TCDF (wet) (mass conc.) | 55722-27-5 | Dioxins & Furans; | Dioxins & Furans; EMS code: B152 | Chemical | Total | LAB
B152 | pg/g | Total-TCDF | 1613B | USEPA 1613B | USEPA 1613B | 0.061 | Y | Y | Y | 4 | 3/20/2020 | 3/21/2020 | 2/21/2023 2:10:22 PM | 2023-02-06 |  | EMS |  | MassPerMass | pg/g | Y | (mass conc.) |  | Total-TCDF (mass conc.) | 55722-27-5 | Dioxins & Furans; | Dioxins & Furans; EMS code: B152 | Chemical | Total | LAB
B155 | pg/L | Upper Bound PCDD/F TEQ (WHO 2005) | E1613B | EPA 1613B | EPA 1613B |  | Y | Y | Y | 1 | 2/3/2022 | 2/3/2022 | 7/6/2022 4:59:14 PM | 2022-06-15 |  | EMS |  | Concentration | pg/L | Y | (fl. conc.) |  | Upper Bound PCDD/F TEQ (WHO 2005) (fl. conc.) |  |  | EMS code: B155 | Chemical |  | LAB
B155 | pg/g | Upper Bound PCDD/F TEQ (WHO 2005) | 1613B | USEPA 1613B | USEPA 1613B |  | Y | Y | Y | 4 | 3/20/2020 | 3/21/2020 | 2/21/2023 2:10:22 PM | 2023-02-06 |  | EMS |  | MassPerMass | pg/g | Y | (mass conc.) |  | Upper Bound PCDD/F TEQ (WHO 2005) (mass conc.) |  |  | EMS code: B155 | Chemical |  | LAB
B155 | pg/g wet | Upper Bound PCDD/F TEQ (WHO 2005) | 1101 | DX-1613B | DCM Extraction, Column clean-up, isotope dilution GC/HRMS. Method Reference: USEPA 1613B. Methodology: Samples are filtered if required. Solids are extracted by Soxhlet using toluene. The liquid portion is extracted by liquid/liquid extraction using dichloromethane. The extracts are prepared using column chromatography, reduced in volume and analyzed by isotope-dilution GC/HRMS. |  | Y | Y | Y | 1 | 2/2/2022 | 2/2/2022 | 7/6/2022 5:20:43 PM | 2019-06-04 |  | KATFRASE |  | MassPerMass | pg/g | Y | (mass conc.) | (wet) | Upper Bound PCDD/F TEQ (WHO 2005) (wet) (mass conc.) |  |  | EMS code: B155 | Chemical |  | LAB
B156 | % | 13C12 BDE 100 | SURR | Surrogate Recovery | Surrogate Recovery | 1 | Y | Y | Y | 36 | 3/27/2020 | 11/26/2023 | 3/15/2024 4:40:24 PM | 2022-08-30 |  | EMS |  | DimensionlessRatio | % | N | (ratio) |  | 13C12 BDE 100 (ratio) | 189084-64-8 | Polybrominated Diphenyl Ethers (PBDEs); | Polybrominated Diphenyl Ethers (PBDEs); EMS code: B156 | Chemical |  | SURROGATE_RESULT
B157 | % | 13C12 BDE 126 | SURR | Surrogate Recovery | Surrogate Recovery | 1 | Y | Y | Y | 36 | 3/27/2020 | 11/26/2023 | 3/15/2024 4:40:24 PM | 2022-08-30 |  | EMS |  | DimensionlessRatio | % | N | (ratio) |  | 13C12 BDE 126 (ratio) | 366791-32-4 | Polybrominated Diphenyl Ethers (PBDEs); | Polybrominated Diphenyl Ethers (PBDEs); EMS code: B157 | Chemical |  | SURROGATE_RESULT
B158 | % | 13C12 BDE 138 Cleanup | SURR | Surrogate Recovery | Surrogate Recovery | 1 | Y | Y | Y | 36 | 3/27/2020 | 11/26/2023 | 3/15/2024 4:40:24 PM | 2022-08-30 |  | EMS |  | DimensionlessRatio | % | N | (ratio) |  | 13C12 BDE 138 Cleanup (ratio) | 182677-30-1 | Polybrominated Diphenyl Ethers (PBDEs); | Polybrominated Diphenyl Ethers (PBDEs); EMS code: B158 | Chemical |  | SURROGATE_RESULT
B159 | % | 13C12 BDE 15 | SURR | Surrogate Recovery | Surrogate Recovery | 1 | Y | Y | Y | 36 | 3/27/2020 | 11/26/2023 | 3/15/2024 4:40:24 PM | 2022-08-30 |  | EMS |  | DimensionlessRatio | % | N | (ratio) |  | 13C12 BDE 15 (ratio) | 2050-47-7 | Polybrominated Diphenyl Ethers (PBDEs); | Polybrominated Diphenyl Ethers (PBDEs); EMS code: B159 | Chemical |  | SURROGATE_RESULT
B160 | % | 13C12 BDE 153 | SURR | Surrogate Recovery | Surrogate Recovery | 1 | Y | Y | Y | 36 | 3/27/2020 | 11/26/2023 | 3/15/2024 4:40:24 PM | 2022-08-30 |  | EMS |  | DimensionlessRatio | % | N | (ratio) |  | 13C12 BDE 153 (ratio) | 68631-49-2 | Polybrominated Diphenyl Ethers (PBDEs); | Polybrominated Diphenyl Ethers (PBDEs); EMS code: B160 | Chemical |  | SURROGATE_RESULT
B161 | % | 13C12 BDE 154 | SURR | Surrogate Recovery | Surrogate Recovery | 1 | Y | Y | Y | 36 | 3/27/2020 | 11/26/2023 | 3/15/2024 4:40:24 PM | 2022-08-30 |  | EMS |  | DimensionlessRatio | % | N | (ratio) |  | 13C12 BDE 154 (ratio) | 207122-15-4 | Polybrominated Diphenyl Ethers (PBDEs); | Polybrominated Diphenyl Ethers (PBDEs); EMS code: B161 | Chemical |  | SURROGATE_RESULT
B162 | % | 13C12 BDE 183 | SURR | Surrogate Recovery | Surrogate Recovery | 1 | Y | Y | Y | 36 | 3/27/2020 | 11/26/2023 | 3/15/2024 4:40:24 PM | 2022-08-30 |  | EMS |  | DimensionlessRatio | % | N | (ratio) |  | 13C12 BDE 183 (ratio) | 207122-16-5 | Polybrominated Diphenyl Ethers (PBDEs); | Polybrominated Diphenyl Ethers (PBDEs); EMS code: B162 | Chemical |  | SURROGATE_RESULT
B163 | % | 13C12 BDE 197 | SURR | Surrogate Recovery | Surrogate Recovery | 1 | Y | Y | Y | 36 | 3/27/2020 | 11/26/2023 | 3/15/2024 4:40:24 PM | 2022-08-30 |  | EMS |  | DimensionlessRatio | % | N | (ratio) |  | 13C12 BDE 197 (ratio) | 117964-21-3 | Polybrominated Diphenyl Ethers (PBDEs); | Polybrominated Diphenyl Ethers (PBDEs); EMS code: B163 | Chemical |  | SURROGATE_RESULT
B164 | % | 13C12 BDE 207 | SURR | Surrogate Recovery | Surrogate Recovery | 1 | Y | Y | Y | 36 | 3/27/2020 | 11/26/2023 | 3/15/2024 4:40:24 PM | 2022-08-30 |  | EMS |  | DimensionlessRatio | % | N | (ratio) |  | 13C12 BDE 207 (ratio) | 437701-79-6 | Polybrominated Diphenyl Ethers (PBDEs); | Polybrominated Diphenyl Ethers (PBDEs); EMS code: B164 | Chemical |  | SURROGATE_RESULT
B165 | % | 13C12 BDE 209 | SURR | Surrogate Recovery | Surrogate Recovery | 1 | Y | Y | Y | 36 | 3/27/2020 | 11/26/2023 | 3/15/2024 4:40:24 PM | 2022-08-30 |  | EMS |  | DimensionlessRatio | % | N | (ratio) |  | 13C12 BDE 209 (ratio) | 1163-19-5 | Polybrominated Diphenyl Ethers (PBDEs); | Polybrominated Diphenyl Ethers (PBDEs); EMS code: B165 | Chemical |  | SURROGATE_RESULT
B166 | % | 13C12 BDE 28 | SURR | Surrogate Recovery | Surrogate Recovery | 1 | Y | Y | Y | 36 | 3/27/2020 | 11/26/2023 | 3/15/2024 4:40:24 PM | 2022-08-30 |  | EMS |  | DimensionlessRatio | % | N | (ratio) |  | 13C12 BDE 28 (ratio) | 41318-75-6 | Polybrominated Diphenyl Ethers (PBDEs); | Polybrominated Diphenyl Ethers (PBDEs); EMS code: B166 | Chemical |  | SURROGATE_RESULT
B167 | % | 13C12 BDE 47 | SURR | Surrogate Recovery | Surrogate Recovery | 1 | Y | Y | Y | 36 | 3/27/2020 | 11/26/2023 | 3/15/2024 4:40:24 PM | 2022-08-30 |  | EMS |  | DimensionlessRatio | % | N | (ratio) |  | 13C12 BDE 47 (ratio) | 5436-43-1 | Polybrominated Diphenyl Ethers (PBDEs); | Polybrominated Diphenyl Ethers (PBDEs); EMS code: B167 | Chemical |  | SURROGATE_RESULT
B168 | % | 13C12 BDE 77 | SURR | Surrogate Recovery | Surrogate Recovery | 1 | Y | Y | Y | 36 | 3/27/2020 | 11/26/2023 | 3/15/2024 4:40:24 PM | 2022-08-30 |  | EMS |  | DimensionlessRatio | % | N | (ratio) |  | 13C12 BDE 77 (ratio) | 93703-48-1 | Polybrominated Diphenyl Ethers (PBDEs); | Polybrominated Diphenyl Ethers (PBDEs); EMS code: B168 | Chemical |  | SURROGATE_RESULT
B169 | % | 13C12 BDE 99 | SURR | Surrogate Recovery | Surrogate Recovery | 1 | Y | Y | Y | 36 | 3/27/2020 | 11/26/2023 | 3/15/2024 4:40:24 PM | 2022-08-30 |  | EMS |  | DimensionlessRatio | % | N | (ratio) |  | 13C12 BDE 99 (ratio) | 60348-60-9 | Polybrominated Diphenyl Ethers (PBDEs); | Polybrominated Diphenyl Ethers (PBDEs); EMS code: B169 | Chemical |  | SURROGATE_RESULT
B170 | % | 13C6 HBB | SURR | Surrogate Recovery | Surrogate Recovery | 1 | Y | Y | Y | 36 | 3/27/2020 | 11/26/2023 | 3/15/2024 4:40:24 PM | 2022-08-30 |  | EMS |  | DimensionlessRatio | % | N | (ratio) |  | 13C6 HBB (ratio) | 2483735-50-6 | Volatile Organic Compounds (VOCs); | Volatile Organic Compounds (VOCs); EMS code: B170 | Chemical |  | SURROGATE_RESULT
B171 | pg/L | BDE 10 | PBDE | PBDE-1614-HRMS-BU | PBDE-1614-HRMS-BU (USEPA 1614) Samples are liquid/liquid extracted with dichloromethane. Extracts are prepared by column chromatography, reduced in volume and analyzed by isotope-dilution GC/HRMS.(ALS - Polybrominated Diphenyl Ethers (PBDEs) are flame retardant chemicals added to products so they won't catch fire or burn so easily if they are exposed to flame or high heat. | 0.1 | Y | Y | Y | 25 | 10/18/2023 | 11/26/2023 | 3/15/2024 4:40:24 PM | 2024-02-05 |  | RDGIBSON |  | Concentration | pg/L | Y | (fl. conc.) |  | BDE 10 (fl. conc.) | 51930-04-2 | Polybrominated Diphenyl Ethers (PBDEs); | Polybrominated Diphenyl Ethers (PBDEs); EMS code: B171 | Chemical |  | LAB
B171 | pg/g | BDE 10 | E1614 | USEPA 1614 | USEPA 1614 | 0.3 | Y | Y | Y | 11 | 3/27/2020 | 3/30/2020 | 2/21/2023 2:10:31 PM | 2023-02-07 |  | EMS |  | MassPerMass | pg/g | Y | (mass conc.) |  | BDE 10 (mass conc.) | 51930-04-2 | Polybrominated Diphenyl Ethers (PBDEs); | Polybrominated Diphenyl Ethers (PBDEs); EMS code: B171 | Chemical |  | LAB
B174 | pg/g | BDE 100 | E1614 | USEPA 1614 | USEPA 1614 | 0.32 | Y | Y | Y | 11 | 3/27/2020 | 3/30/2020 | 2/21/2023 2:10:31 PM | 2023-02-07 |  | EMS |  | MassPerMass | pg/g | Y | (mass conc.) |  | BDE 100 (mass conc.) | 189084-64-8 | Polybrominated Diphenyl Ethers (PBDEs); | Polybrominated Diphenyl Ethers (PBDEs); EMS code: B174 | Chemical |  | LAB
B174 | pg/L | BDE 100 | PBDE | PBDE-1614-HRMS-BU | PBDE-1614-HRMS-BU (USEPA 1614) Samples are liquid/liquid extracted with dichloromethane. Extracts are prepared by column chromatography, reduced in volume and analyzed by isotope-dilution GC/HRMS.(ALS - Polybrominated Diphenyl Ethers (PBDEs) are flame retardant chemicals added to products so they won't catch fire or burn so easily if they are exposed to flame or high heat. | 0.18 | Y | Y | Y | 25 | 10/18/2023 | 11/26/2023 | 3/15/2024 4:40:24 PM | 2024-02-05 |  | EMS |  | Concentration | pg/L | Y | (fl. conc.) |  | BDE 100 (fl. conc.) | 189084-64-8 | Polybrominated Diphenyl Ethers (PBDEs); | Polybrominated Diphenyl Ethers (PBDEs); EMS code: B174 | Chemical |  | LAB
B177 | pg/L | BDE 105 | PBDE | PBDE-1614-HRMS-BU | PBDE-1614-HRMS-BU (USEPA 1614) Samples are liquid/liquid extracted with dichloromethane. Extracts are prepared by column chromatography, reduced in volume and analyzed by isotope-dilution GC/HRMS.(ALS - Polybrominated Diphenyl Ethers (PBDEs) are flame retardant chemicals added to products so they won't catch fire or burn so easily if they are exposed to flame or high heat. | 0.5 | Y | Y | Y | 25 | 10/18/2023 | 11/26/2023 | 3/15/2024 4:40:24 PM | 2024-02-05 |  | EMS |  | Concentration | pg/L | Y | (fl. conc.) |  | BDE 105 (fl. conc.) | 373594-78-6 | Polybrominated Diphenyl Ethers (PBDEs); | Polybrominated Diphenyl Ethers (PBDEs); EMS code: B177 | Chemical |  | LAB
B177 | pg/g | BDE 105 | E1614 | USEPA 1614 | USEPA 1614 | 0.63 | Y | Y | Y | 11 | 3/27/2020 | 3/30/2020 | 2/21/2023 2:10:31 PM | 2023-02-07 |  | EMS |  | MassPerMass | pg/g | Y | (mass conc.) |  | BDE 105 (mass conc.) | 373594-78-6 | Polybrominated Diphenyl Ethers (PBDEs); | Polybrominated Diphenyl Ethers (PBDEs); EMS code: B177 | Chemical |  | LAB
B180 | pg/L | BDE 116 | PBDE | PBDE-1614-HRMS-BU | PBDE-1614-HRMS-BU (USEPA 1614) Samples are liquid/liquid extracted with dichloromethane. Extracts are prepared by column chromatography, reduced in volume and analyzed by isotope-dilution GC/HRMS.(ALS - Polybrominated Diphenyl Ethers (PBDEs) are flame retardant chemicals added to products so they won't catch fire or burn so easily if they are exposed to flame or high heat. | 0.63 | Y | Y | Y | 25 | 10/18/2023 | 11/26/2023 | 3/15/2024 4:40:24 PM | 2024-02-05 |  | EMS |  | Concentration | pg/L | Y | (fl. conc.) |  | BDE 116 (fl. conc.) | 189084-65-9 | Polybrominated Diphenyl Ethers (PBDEs); | Polybrominated Diphenyl Ethers (PBDEs); EMS code: B180 | Chemical |  | LAB
B180 | pg/g | BDE 116 | E1614 | USEPA 1614 | USEPA 1614 | 1 | Y | Y | Y | 11 | 3/27/2020 | 3/30/2020 | 2/21/2023 2:10:31 PM | 2023-02-07 |  | EMS |  | MassPerMass | pg/g | Y | (mass conc.) |  | BDE 116 (mass conc.) | 189084-65-9 | Polybrominated Diphenyl Ethers (PBDEs); | Polybrominated Diphenyl Ethers (PBDEs); EMS code: B180 | Chemical |  | LAB
B183 | pg/L | BDE 118 | PBDE | PBDE-1614-HRMS-BU | PBDE-1614-HRMS-BU (USEPA 1614) Samples are liquid/liquid extracted with dichloromethane. Extracts are prepared by column chromatography, reduced in volume and analyzed by isotope-dilution GC/HRMS.(ALS - Polybrominated Diphenyl Ethers (PBDEs) are flame retardant chemicals added to products so they won't catch fire or burn so easily if they are exposed to flame or high heat. | 0.45 | Y | Y | Y | 25 | 10/18/2023 | 11/26/2023 | 3/15/2024 4:40:24 PM | 2024-02-05 |  | EMS |  | Concentration | pg/L | Y | (fl. conc.) |  | BDE 118 (fl. conc.) | 446254-80-4 | Polybrominated Diphenyl Ethers (PBDEs); | Polybrominated Diphenyl Ethers (PBDEs); EMS code: B183 | Chemical |  | LAB
B183 | pg/g | BDE 118 | E1614 | USEPA 1614 | USEPA 1614 | 0.69 | Y | Y | Y | 11 | 3/27/2020 | 3/30/2020 | 2/21/2023 2:10:31 PM | 2023-02-07 |  | EMS |  | MassPerMass | pg/g | Y | (mass conc.) |  | BDE 118 (mass conc.) | 446254-80-4 | Polybrominated Diphenyl Ethers (PBDEs); | Polybrominated Diphenyl Ethers (PBDEs); EMS code: B183 | Chemical |  | LAB
B185 | pg/g | BDE 119/120 | E1614 | USEPA 1614 | USEPA 1614 | 0.66 | Y | Y | Y | 11 | 3/27/2020 | 3/30/2020 | 2/21/2023 2:10:31 PM | 2023-02-07 |  | EMS |  | MassPerMass | pg/g | Y | (mass conc.) |  | BDE 119/120 (mass conc.) | 189084-66-0, 417727-71-0 | Polybrominated Diphenyl Ethers (PBDEs); | Polybrominated Diphenyl Ethers (PBDEs); EMS code: B185 | Chemical |  | LAB
B185 | pg/L | BDE 119/120 | PBDE | PBDE-1614-HRMS-BU | PBDE-1614-HRMS-BU (USEPA 1614) Samples are liquid/liquid extracted with dichloromethane. Extracts are prepared by column chromatography, reduced in volume and analyzed by isotope-dilution GC/HRMS.(ALS - Polybrominated Diphenyl Ethers (PBDEs) are flame retardant chemicals added to products so they won't catch fire or burn so easily if they are exposed to flame or high heat. | 0.36 | Y | Y | Y | 25 | 10/18/2023 | 11/26/2023 | 3/15/2024 4:40:24 PM | 2024-02-05 |  | EMS |  | Concentration | pg/L | Y | (fl. conc.) |  | BDE 119/120 (fl. conc.) | 189084-66-0, 417727-71-0 | Polybrominated Diphenyl Ethers (PBDEs); | Polybrominated Diphenyl Ethers (PBDEs); EMS code: B185 | Chemical |  | LAB
B187 | pg/L | BDE 12/13 | PBDE | PBDE-1614-HRMS-BU | PBDE-1614-HRMS-BU (USEPA 1614) Samples are liquid/liquid extracted with dichloromethane. Extracts are prepared by column chromatography, reduced in volume and analyzed by isotope-dilution GC/HRMS.(ALS - Polybrominated Diphenyl Ethers (PBDEs) are flame retardant chemicals added to products so they won't catch fire or burn so easily if they are exposed to flame or high heat. | 0.066 | Y | Y | Y | 25 | 10/18/2023 | 11/26/2023 | 3/15/2024 4:40:24 PM | 2024-02-05 |  | EMS |  | Concentration | pg/L | Y | (fl. conc.) |  | BDE 12/13 (fl. conc.) | 189084-59-1, 83694-71-7 | Polybrominated Diphenyl Ethers (PBDEs); | Polybrominated Diphenyl Ethers (PBDEs); EMS code: B187 | Chemical |  | LAB
B187 | pg/g | BDE 12/13 | E1614 | USEPA 1614 | USEPA 1614 | 0.18 | Y | Y | Y | 11 | 3/27/2020 | 3/30/2020 | 2/21/2023 2:10:31 PM | 2023-02-07 |  | EMS |  | MassPerMass | pg/g | Y | (mass conc.) |  | BDE 12/13 (mass conc.) | 189084-59-1, 83694-71-7 | Polybrominated Diphenyl Ethers (PBDEs); | Polybrominated Diphenyl Ethers (PBDEs); EMS code: B187 | Chemical |  | LAB
B190 | pg/L | BDE 126 | PBDE | PBDE-1614-HRMS-BU | PBDE-1614-HRMS-BU (USEPA 1614) Samples are liquid/liquid extracted with dichloromethane. Extracts are prepared by column chromatography, reduced in volume and analyzed by isotope-dilution GC/HRMS.(ALS - Polybrominated Diphenyl Ethers (PBDEs) are flame retardant chemicals added to products so they won't catch fire or burn so easily if they are exposed to flame or high heat. | 0.24 | Y | Y | Y | 25 | 10/18/2023 | 11/26/2023 | 3/15/2024 4:40:24 PM | 2024-02-05 |  | EMS |  | Concentration | pg/L | Y | (fl. conc.) |  | BDE 126 (fl. conc.) | 366791-32-4 | Polybrominated Diphenyl Ethers (PBDEs); | Polybrominated Diphenyl Ethers (PBDEs); EMS code: B190 | Chemical |  | LAB
B190 | pg/g | BDE 126 | E1614 | USEPA 1614 | USEPA 1614 | 0.39 | Y | Y | Y | 11 | 3/27/2020 | 3/30/2020 | 2/21/2023 2:10:31 PM | 2023-02-07 |  | EMS |  | MassPerMass | pg/g | Y | (mass conc.) |  | BDE 126 (mass conc.) | 366791-32-4 | Polybrominated Diphenyl Ethers (PBDEs); | Polybrominated Diphenyl Ethers (PBDEs); EMS code: B190 | Chemical |  | LAB
B193 | pg/g | BDE 128 | E1614 | USEPA 1614 | USEPA 1614 | 3.5 | Y | Y | Y | 11 | 3/27/2020 | 3/30/2020 | 2/21/2023 2:10:31 PM | 2023-02-07 |  | EMS |  | MassPerMass | pg/g | Y | (mass conc.) |  | BDE 128 (mass conc.) | 182677-28-7 | Polybrominated Diphenyl Ethers (PBDEs); | Polybrominated Diphenyl Ethers (PBDEs); EMS code: B193 | Chemical |  | LAB
B193 | pg/L | BDE 128 | PBDE | PBDE-1614-HRMS-BU | PBDE-1614-HRMS-BU (USEPA 1614) Samples are liquid/liquid extracted with dichloromethane. Extracts are prepared by column chromatography, reduced in volume and analyzed by isotope-dilution GC/HRMS.(ALS - Polybrominated Diphenyl Ethers (PBDEs) are flame retardant chemicals added to products so they won't catch fire or burn so easily if they are exposed to flame or high heat. | 2.7 | Y | Y | Y | 25 | 10/18/2023 | 11/26/2023 | 3/15/2024 4:40:24 PM | 2024-02-05 |  | EMS |  | Concentration | pg/L | Y | (fl. conc.) |  | BDE 128 (fl. conc.) | 182677-28-7 | Polybrominated Diphenyl Ethers (PBDEs); | Polybrominated Diphenyl Ethers (PBDEs); EMS code: B193 | Chemical |  | LAB
B196 | pg/g | BDE 138/166 | E1614 | USEPA 1614 | USEPA 1614 | 3 | Y | Y | Y | 11 | 3/27/2020 | 3/30/2020 | 2/21/2023 2:10:31 PM | 2023-02-07 |  | EMS |  | MassPerMass | pg/g | Y | (mass conc.) |  | BDE 138/166 (mass conc.) | 182677-30-1, 189084-58-0 | Polybrominated Diphenyl Ethers (PBDEs); | Polybrominated Diphenyl Ethers (PBDEs); EMS code: B196 | Chemical |  | LAB
B196 | pg/L | BDE 138/166 | PBDE | PBDE-1614-HRMS-BU | PBDE-1614-HRMS-BU (USEPA 1614) Samples are liquid/liquid extracted with dichloromethane. Extracts are prepared by column chromatography, reduced in volume and analyzed by isotope-dilution GC/HRMS.(ALS - Polybrominated Diphenyl Ethers (PBDEs) are flame retardant chemicals added to products so they won't catch fire or burn so easily if they are exposed to flame or high heat. | 2.2 | Y | Y | Y | 25 | 10/18/2023 | 11/26/2023 | 3/15/2024 4:40:24 PM | 2024-02-05 |  | EMS |  | Concentration | pg/L | Y | (fl. conc.) |  | BDE 138/166 (fl. conc.) | 182677-30-1, 189084-58-0 | Polybrominated Diphenyl Ethers (PBDEs); | Polybrominated Diphenyl Ethers (PBDEs); EMS code: B196 | Chemical |  | LAB
B199 | pg/L | BDE 140 | PBDE | PBDE-1614-HRMS-BU | PBDE-1614-HRMS-BU (USEPA 1614) Samples are liquid/liquid extracted with dichloromethane. Extracts are prepared by column chromatography, reduced in volume and analyzed by isotope-dilution GC/HRMS.(ALS - Polybrominated Diphenyl Ethers (PBDEs) are flame retardant chemicals added to products so they won't catch fire or burn so easily if they are exposed to flame or high heat. | 1.1 | Y | Y | Y | 25 | 10/18/2023 | 11/26/2023 | 3/15/2024 4:40:24 PM | 2024-02-05 |  | EMS |  | Concentration | pg/L | Y | (fl. conc.) |  | BDE 140 (fl. conc.) | 243982-83-4 | Polybrominated Diphenyl Ethers (PBDEs); | Polybrominated Diphenyl Ethers (PBDEs); EMS code: B199 | Chemical |  | LAB
B199 | pg/g | BDE 140 | E1614 | USEPA 1614 | USEPA 1614 | 1.8 | Y | Y | Y | 11 | 3/27/2020 | 3/30/2020 | 2/21/2023 2:10:31 PM | 2023-02-07 |  | EMS |  | MassPerMass | pg/g | Y | (mass conc.) |  | BDE 140 (mass conc.) | 243982-83-4 | Polybrominated Diphenyl Ethers (PBDEs); | Polybrominated Diphenyl Ethers (PBDEs); EMS code: B199 | Chemical |  | LAB
B202 | pg/g | BDE 15 | E1614 | USEPA 1614 | USEPA 1614 | 0.15 | Y | Y | Y | 11 | 3/27/2020 | 3/30/2020 | 2/21/2023 2:10:31 PM | 2023-02-07 |  | EMS |  | MassPerMass | pg/g | Y | (mass conc.) |  | BDE 15 (mass conc.) | 2050-47-7 | Polybrominated Diphenyl Ethers (PBDEs); | Polybrominated Diphenyl Ethers (PBDEs); EMS code: B202 | Chemical |  | LAB
B202 | pg/L | BDE 15 | PBDE | PBDE-1614-HRMS-BU | PBDE-1614-HRMS-BU (USEPA 1614) Samples are liquid/liquid extracted with dichloromethane. Extracts are prepared by column chromatography, reduced in volume and analyzed by isotope-dilution GC/HRMS.(ALS - Polybrominated Diphenyl Ethers (PBDEs) are flame retardant chemicals added to products so they won't catch fire or burn so easily if they are exposed to flame or high heat. | 0.055 | Y | Y | Y | 25 | 10/18/2023 | 11/26/2023 | 3/15/2024 4:40:24 PM | 2024-02-05 |  | EMS |  | Concentration | pg/L | Y | (fl. conc.) |  | BDE 15 (fl. conc.) | 2050-47-7 | Polybrominated Diphenyl Ethers (PBDEs); | Polybrominated Diphenyl Ethers (PBDEs); EMS code: B202 | Chemical |  | LAB
B205 | pg/L | BDE 153 | PBDE | PBDE-1614-HRMS-BU | PBDE-1614-HRMS-BU (USEPA 1614) Samples are liquid/liquid extracted with dichloromethane. Extracts are prepared by column chromatography, reduced in volume and analyzed by isotope-dilution GC/HRMS.(ALS - Polybrominated Diphenyl Ethers (PBDEs) are flame retardant chemicals added to products so they won't catch fire or burn so easily if they are exposed to flame or high heat. | 1.4 | Y | Y | Y | 25 | 10/18/2023 | 11/26/2023 | 3/15/2024 4:40:24 PM | 2024-02-05 |  | EMS |  | Concentration | pg/L | Y | (fl. conc.) |  | BDE 153 (fl. conc.) | 68631-49-2 | Polybrominated Diphenyl Ethers (PBDEs); | Polybrominated Diphenyl Ethers (PBDEs); EMS code: B205 | Chemical |  | LAB
B205 | pg/g | BDE 153 | E1614 | USEPA 1614 | USEPA 1614 | 1.5 | Y | Y | Y | 11 | 3/27/2020 | 3/30/2020 | 2/21/2023 2:10:31 PM | 2023-02-07 |  | EMS |  | MassPerMass | pg/g | Y | (mass conc.) |  | BDE 153 (mass conc.) | 68631-49-2 | Polybrominated Diphenyl Ethers (PBDEs); | Polybrominated Diphenyl Ethers (PBDEs); EMS code: B205 | Chemical |  | LAB
B208 | pg/L | BDE 154 | PBDE | PBDE-1614-HRMS-BU | PBDE-1614-HRMS-BU (USEPA 1614) Samples are liquid/liquid extracted with dichloromethane. Extracts are prepared by column chromatography, reduced in volume and analyzed by isotope-dilution GC/HRMS.(ALS - Polybrominated Diphenyl Ethers (PBDEs) are flame retardant chemicals added to products so they won't catch fire or burn so easily if they are exposed to flame or high heat. | 0.44 | Y | Y | Y | 25 | 10/18/2023 | 11/26/2023 | 3/15/2024 4:40:24 PM | 2024-02-05 |  | EMS |  | Concentration | pg/L | Y | (fl. conc.) |  | BDE 154 (fl. conc.) | 207122-15-4 | Polybrominated Diphenyl Ethers (PBDEs); | Polybrominated Diphenyl Ethers (PBDEs); EMS code: B208 | Chemical |  | LAB
B208 | pg/g | BDE 154 | E1614 | USEPA 1614 | USEPA 1614 | 1.1 | Y | Y | Y | 11 | 3/27/2020 | 3/30/2020 | 2/21/2023 2:10:31 PM | 2023-02-07 |  | EMS |  | MassPerMass | pg/g | Y | (mass conc.) |  | BDE 154 (mass conc.) | 207122-15-4 | Polybrominated Diphenyl Ethers (PBDEs); | Polybrominated Diphenyl Ethers (PBDEs); EMS code: B208 | Chemical |  | LAB
B211 | pg/L | BDE 155 | PBDE | PBDE-1614-HRMS-BU | PBDE-1614-HRMS-BU (USEPA 1614) Samples are liquid/liquid extracted with dichloromethane. Extracts are prepared by column chromatography, reduced in volume and analyzed by isotope-dilution GC/HRMS.(ALS - Polybrominated Diphenyl Ethers (PBDEs) are flame retardant chemicals added to products so they won't catch fire or burn so easily if they are exposed to flame or high heat. | 0.35 | Y | Y | Y | 25 | 10/18/2023 | 11/26/2023 | 3/15/2024 4:40:24 PM | 2024-02-05 |  | EMS |  | Concentration | pg/L | Y | (fl. conc.) |  | BDE 155 (fl. conc.) | 35854-94-5 | Polybrominated Diphenyl Ethers (PBDEs); | Polybrominated Diphenyl Ethers (PBDEs); EMS code: B211 | Chemical |  | LAB
B211 | pg/g | BDE 155 | E1614 | USEPA 1614 | USEPA 1614 | 0.7 | Y | Y | Y | 11 | 3/27/2020 | 3/30/2020 | 2/21/2023 2:10:31 PM | 2023-02-07 |  | EMS |  | MassPerMass | pg/g | Y | (mass conc.) |  | BDE 155 (mass conc.) | 35854-94-5 | Polybrominated Diphenyl Ethers (PBDEs); | Polybrominated Diphenyl Ethers (PBDEs); EMS code: B211 | Chemical |  | LAB
B214 | pg/g | BDE 156 | E1614 | USEPA 1614 | USEPA 1614 | 3.8 | Y | Y | Y | 11 | 3/27/2020 | 3/30/2020 | 2/21/2023 2:10:31 PM | 2023-02-07 |  | EMS |  | MassPerMass | pg/g | Y | (mass conc.) |  | BDE 156 (mass conc.) | 405237-85-6 | Polybrominated Diphenyl Ethers (PBDEs); | Polybrominated Diphenyl Ethers (PBDEs); EMS code: B214 | Chemical |  | LAB
B214 | pg/L | BDE 156 | PBDE | PBDE-1614-HRMS-BU | PBDE-1614-HRMS-BU (USEPA 1614) Samples are liquid/liquid extracted with dichloromethane. Extracts are prepared by column chromatography, reduced in volume and analyzed by isotope-dilution GC/HRMS.(ALS - Polybrominated Diphenyl Ethers (PBDEs) are flame retardant chemicals added to products so they won't catch fire or burn so easily if they are exposed to flame or high heat. | 3.3 | Y | Y | Y | 25 | 10/18/2023 | 11/26/2023 | 3/15/2024 4:40:24 PM | 2024-02-05 |  | EMS |  | Concentration | pg/L | Y | (fl. conc.) |  | BDE 156 (fl. conc.) | 405237-85-6 | Polybrominated Diphenyl Ethers (PBDEs); | Polybrominated Diphenyl Ethers (PBDEs); EMS code: B214 | Chemical |  | LAB
B217 | pg/L | BDE 17/25 | PBDE | PBDE-1614-HRMS-BU | PBDE-1614-HRMS-BU (USEPA 1614) Samples are liquid/liquid extracted with dichloromethane. Extracts are prepared by column chromatography, reduced in volume and analyzed by isotope-dilution GC/HRMS.(ALS - Polybrominated Diphenyl Ethers (PBDEs) are flame retardant chemicals added to products so they won't catch fire or burn so easily if they are exposed to flame or high heat. | 0.25 | Y | Y | Y | 25 | 10/18/2023 | 11/26/2023 | 3/15/2024 4:40:24 PM | 2024-02-05 |  | EMS |  | Concentration | pg/L | Y | (fl. conc.) |  | BDE 17/25 (fl. conc.) | 147217-75-2, 147217-77-4 | Polybrominated Diphenyl Ethers (PBDEs); | Polybrominated Diphenyl Ethers (PBDEs); EMS code: B217 | Chemical |  | LAB
B217 | pg/g | BDE 17/25 | E1614 | USEPA 1614 | USEPA 1614 | 0.53 | Y | Y | Y | 11 | 3/27/2020 | 3/30/2020 | 2/21/2023 2:10:31 PM | 2023-02-07 |  | EMS |  | MassPerMass | pg/g | Y | (mass conc.) |  | BDE 17/25 (mass conc.) | 147217-75-2, 147217-77-4 | Polybrominated Diphenyl Ethers (PBDEs); | Polybrominated Diphenyl Ethers (PBDEs); EMS code: B217 | Chemical |  | LAB
B220 | pg/g | BDE 181 | E1614 | USEPA 1614 | USEPA 1614 | 3.3 | Y | Y | Y | 11 | 3/27/2020 | 3/30/2020 | 2/21/2023 2:10:31 PM | 2023-02-07 |  | EMS |  | MassPerMass | pg/g | Y | (mass conc.) |  | BDE 181 (mass conc.) | 189084-67-1 | Polybrominated Diphenyl Ethers (PBDEs); | Polybrominated Diphenyl Ethers (PBDEs); EMS code: B220 | Chemical |  | LAB
B220 | pg/L | BDE 181 | PBDE | PBDE-1614-HRMS-BU | PBDE-1614-HRMS-BU (USEPA 1614) Samples are liquid/liquid extracted with dichloromethane. Extracts are prepared by column chromatography, reduced in volume and analyzed by isotope-dilution GC/HRMS.(ALS - Polybrominated Diphenyl Ethers (PBDEs) are flame retardant chemicals added to products so they won't catch fire or burn so easily if they are exposed to flame or high heat. | 1.3 | Y | Y | Y | 25 | 10/18/2023 | 11/26/2023 | 3/15/2024 4:40:24 PM | 2024-02-05 |  | EMS |  | Concentration | pg/L | Y | (fl. conc.) |  | BDE 181 (fl. conc.) | 189084-67-1 | Polybrominated Diphenyl Ethers (PBDEs); | Polybrominated Diphenyl Ethers (PBDEs); EMS code: B220 | Chemical |  | LAB
B223 | pg/L | BDE 183 | PBDE | PBDE-1614-HRMS-BU | PBDE-1614-HRMS-BU (USEPA 1614) Samples are liquid/liquid extracted with dichloromethane. Extracts are prepared by column chromatography, reduced in volume and analyzed by isotope-dilution GC/HRMS.(ALS - Polybrominated Diphenyl Ethers (PBDEs) are flame retardant chemicals added to products so they won't catch fire or burn so easily if they are exposed to flame or high heat. | 0.79 | Y | Y | Y | 25 | 10/18/2023 | 11/26/2023 | 3/15/2024 4:40:24 PM | 2024-02-05 |  | EMS |  | Concentration | pg/L | Y | (fl. conc.) |  | BDE 183 (fl. conc.) | 207122-16-5 | Polybrominated Diphenyl Ethers (PBDEs); | Polybrominated Diphenyl Ethers (PBDEs); EMS code: B223 | Chemical |  | LAB
B223 | pg/g | BDE 183 | E1614 | USEPA 1614 | USEPA 1614 | 2.2 | Y | Y | Y | 11 | 3/27/2020 | 3/30/2020 | 2/21/2023 2:10:31 PM | 2023-02-07 |  | EMS |  | MassPerMass | pg/g | Y | (mass conc.) |  | BDE 183 (mass conc.) | 207122-16-5 | Polybrominated Diphenyl Ethers (PBDEs); | Polybrominated Diphenyl Ethers (PBDEs); EMS code: B223 | Chemical |  | LAB
B226 | pg/L | BDE 184 | PBDE | PBDE-1614-HRMS-BU | PBDE-1614-HRMS-BU (USEPA 1614) Samples are liquid/liquid extracted with dichloromethane. Extracts are prepared by column chromatography, reduced in volume and analyzed by isotope-dilution GC/HRMS.(ALS - Polybrominated Diphenyl Ethers (PBDEs) are flame retardant chemicals added to products so they won't catch fire or burn so easily if they are exposed to flame or high heat. | 0.52 | Y | Y | Y | 25 | 10/18/2023 | 11/26/2023 | 3/15/2024 4:40:25 PM | 2024-02-05 |  | EMS |  | Concentration | pg/L | Y | (fl. conc.) |  | BDE 184 (fl. conc.) | 117948-63-7 | Polybrominated Diphenyl Ethers (PBDEs); | Polybrominated Diphenyl Ethers (PBDEs); EMS code: B226 | Chemical |  | LAB
B226 | pg/g | BDE 184 | E1614 | USEPA 1614 | USEPA 1614 | 1.5 | Y | Y | Y | 11 | 3/27/2020 | 3/30/2020 | 2/21/2023 2:10:31 PM | 2023-02-07 |  | EMS |  | MassPerMass | pg/g | Y | (mass conc.) |  | BDE 184 (mass conc.) | 117948-63-7 | Polybrominated Diphenyl Ethers (PBDEs); | Polybrominated Diphenyl Ethers (PBDEs); EMS code: B226 | Chemical |  | LAB
B229 | pg/L | BDE 190 | PBDE | PBDE-1614-HRMS-BU | PBDE-1614-HRMS-BU (USEPA 1614) Samples are liquid/liquid extracted with dichloromethane. Extracts are prepared by column chromatography, reduced in volume and analyzed by isotope-dilution GC/HRMS.(ALS - Polybrominated Diphenyl Ethers (PBDEs) are flame retardant chemicals added to products so they won't catch fire or burn so easily if they are exposed to flame or high heat. | 2 | Y | Y | Y | 25 | 10/18/2023 | 11/26/2023 | 3/15/2024 4:40:25 PM | 2024-02-05 |  | EMS |  | Concentration | pg/L | Y | (fl. conc.) |  | BDE 190 (fl. conc.) | 189084-68-2 | Polybrominated Diphenyl Ethers (PBDEs); | Polybrominated Diphenyl Ethers (PBDEs); EMS code: B229 | Chemical |  | LAB
B229 | pg/g | BDE 190 | E1614 | USEPA 1614 | USEPA 1614 | 4 | Y | Y | Y | 11 | 3/27/2020 | 3/30/2020 | 2/21/2023 2:10:31 PM | 2023-02-07 |  | EMS |  | MassPerMass | pg/g | Y | (mass conc.) |  | BDE 190 (mass conc.) | 189084-68-2 | Polybrominated Diphenyl Ethers (PBDEs); | Polybrominated Diphenyl Ethers (PBDEs); EMS code: B229 | Chemical |  | LAB
B234 | pg/g | BDE 191 | E1614 | USEPA 1614 | USEPA 1614 | 3.2 | Y | Y | Y | 11 | 3/27/2020 | 3/30/2020 | 2/21/2023 2:10:31 PM | 2023-02-07 |  | EMS |  | MassPerMass | pg/g | Y | (mass conc.) |  | BDE 191 (mass conc.) | 446255-30-7 | Polybrominated Diphenyl Ethers (PBDEs); | Polybrominated Diphenyl Ethers (PBDEs); EMS code: B234 | Chemical |  | LAB
B234 | pg/L | BDE 191 | PBDE | PBDE-1614-HRMS-BU | PBDE-1614-HRMS-BU (USEPA 1614) Samples are liquid/liquid extracted with dichloromethane. Extracts are prepared by column chromatography, reduced in volume and analyzed by isotope-dilution GC/HRMS.(ALS - Polybrominated Diphenyl Ethers (PBDEs) are flame retardant chemicals added to products so they won't catch fire or burn so easily if they are exposed to flame or high heat. | 1.3 | Y | Y | Y | 25 | 10/18/2023 | 11/26/2023 | 3/15/2024 4:40:25 PM | 2024-02-05 |  | EMS |  | Concentration | pg/L | Y | (fl. conc.) |  | BDE 191 (fl. conc.) | 446255-30-7 | Polybrominated Diphenyl Ethers (PBDEs); | Polybrominated Diphenyl Ethers (PBDEs); EMS code: B234 | Chemical |  | LAB
B235 | pg/g | BDE 196 | E1614 | USEPA 1614 | USEPA 1614 | 3.7 | Y | Y | Y | 11 | 3/27/2020 | 3/30/2020 | 2/21/2023 2:10:31 PM | 2023-02-07 |  | EMS |  | MassPerMass | pg/g | Y | (mass conc.) |  | BDE 196 (mass conc.) | 446255-39-6 | Polybrominated Diphenyl Ethers (PBDEs); | Polybrominated Diphenyl Ethers (PBDEs); EMS code: B235 | Chemical |  | LAB
B235 | pg/L | BDE 196 | PBDE | PBDE-1614-HRMS-BU | PBDE-1614-HRMS-BU (USEPA 1614) Samples are liquid/liquid extracted with dichloromethane. Extracts are prepared by column chromatography, reduced in volume and analyzed by isotope-dilution GC/HRMS.(ALS - Polybrominated Diphenyl Ethers (PBDEs) are flame retardant chemicals added to products so they won't catch fire or burn so easily if they are exposed to flame or high heat. | 1.3 | Y | Y | Y | 25 | 10/18/2023 | 11/26/2023 | 3/15/2024 4:40:25 PM | 2024-02-05 |  | EMS |  | Concentration | pg/L | Y | (fl. conc.) |  | BDE 196 (fl. conc.) | 446255-39-6 | Polybrominated Diphenyl Ethers (PBDEs); | Polybrominated Diphenyl Ethers (PBDEs); EMS code: B235 | Chemical |  | LAB
B238 | pg/L | BDE 197 | PBDE | PBDE-1614-HRMS-BU | PBDE-1614-HRMS-BU (USEPA 1614) Samples are liquid/liquid extracted with dichloromethane. Extracts are prepared by column chromatography, reduced in volume and analyzed by isotope-dilution GC/HRMS.(ALS - Polybrominated Diphenyl Ethers (PBDEs) are flame retardant chemicals added to products so they won't catch fire or burn so easily if they are exposed to flame or high heat. | 1.1 | Y | Y | Y | 25 | 10/18/2023 | 11/26/2023 | 3/15/2024 4:40:25 PM | 2024-02-05 |  | EMS |  | Concentration | pg/L | Y | (fl. conc.) |  | BDE 197 (fl. conc.) | 117964-21-3 | Polybrominated Diphenyl Ethers (PBDEs); | Polybrominated Diphenyl Ethers (PBDEs); EMS code: B238 | Chemical |  | LAB
B238 | pg/g | BDE 197 | E1614 | USEPA 1614 | USEPA 1614 | 3.2 | Y | Y | Y | 11 | 3/27/2020 | 3/30/2020 | 2/21/2023 2:10:31 PM | 2023-02-07 |  | EMS |  | MassPerMass | pg/g | Y | (mass conc.) |  | BDE 197 (mass conc.) | 117964-21-3 | Polybrominated Diphenyl Ethers (PBDEs); | Polybrominated Diphenyl Ethers (PBDEs); EMS code: B238 | Chemical |  | LAB
B241 | pg/L | BDE 203 | PBDE | PBDE-1614-HRMS-BU | PBDE-1614-HRMS-BU (USEPA 1614) Samples are liquid/liquid extracted with dichloromethane. Extracts are prepared by column chromatography, reduced in volume and analyzed by isotope-dilution GC/HRMS.(ALS - Polybrominated Diphenyl Ethers (PBDEs) are flame retardant chemicals added to products so they won't catch fire or burn so easily if they are exposed to flame or high heat. | 1.6 | Y | Y | Y | 25 | 10/18/2023 | 11/26/2023 | 3/15/2024 4:40:25 PM | 2024-02-05 |  | EMS |  | Concentration | pg/L | Y | (fl. conc.) |  | BDE 203 (fl. conc.) | 337513-72-1 | Polybrominated Diphenyl Ethers (PBDEs); | Polybrominated Diphenyl Ethers (PBDEs); EMS code: B241 | Chemical |  | LAB
B241 | pg/g | BDE 203 | E1614 | USEPA 1614 | USEPA 1614 | 4.9 | Y | Y | Y | 11 | 3/27/2020 | 3/30/2020 | 2/21/2023 2:10:31 PM | 2023-02-07 |  | EMS |  | MassPerMass | pg/g | Y | (mass conc.) |  | BDE 203 (mass conc.) | 337513-72-1 | Polybrominated Diphenyl Ethers (PBDEs); | Polybrominated Diphenyl Ethers (PBDEs); EMS code: B241 | Chemical |  | LAB
B244 | pg/L | BDE 206 | PBDE | PBDE-1614-HRMS-BU | PBDE-1614-HRMS-BU (USEPA 1614) Samples are liquid/liquid extracted with dichloromethane. Extracts are prepared by column chromatography, reduced in volume and analyzed by isotope-dilution GC/HRMS.(ALS - Polybrominated Diphenyl Ethers (PBDEs) are flame retardant chemicals added to products so they won't catch fire or burn so easily if they are exposed to flame or high heat. | 2.4 | Y | Y | Y | 25 | 10/18/2023 | 11/26/2023 | 3/15/2024 4:40:25 PM | 2024-02-05 |  | EMS |  | Concentration | pg/L | Y | (fl. conc.) |  | BDE 206 (fl. conc.) | 63387-28-0 | Polybrominated Diphenyl Ethers (PBDEs); | Polybrominated Diphenyl Ethers (PBDEs); EMS code: B244 | Chemical |  | LAB
B244 | pg/g | BDE 206 | E1614 | USEPA 1614 | USEPA 1614 | 11 | Y | Y | Y | 11 | 3/27/2020 | 3/30/2020 | 2/21/2023 2:10:31 PM | 2023-02-07 |  | EMS |  | MassPerMass | pg/g | Y | (mass conc.) |  | BDE 206 (mass conc.) | 63387-28-0 | Polybrominated Diphenyl Ethers (PBDEs); | Polybrominated Diphenyl Ethers (PBDEs); EMS code: B244 | Chemical |  | LAB
B247 | pg/g | BDE 207 | E1614 | USEPA 1614 | USEPA 1614 | 9 | Y | Y | Y | 11 | 3/27/2020 | 3/30/2020 | 2/21/2023 2:10:31 PM | 2023-02-09 |  | EMS |  | MassPerMass | pg/g | Y | (mass conc.) |  | BDE 207 (mass conc.) | 437701-79-6 | Polybrominated Diphenyl Ethers (PBDEs); | Polybrominated Diphenyl Ethers (PBDEs); EMS code: B247 | Chemical |  | LAB
B247 | pg/L | BDE 207 | PBDE | PBDE-1614-HRMS-BU | PBDE-1614-HRMS-BU (USEPA 1614) Samples are liquid/liquid extracted with dichloromethane. Extracts are prepared by column chromatography, reduced in volume and analyzed by isotope-dilution GC/HRMS.(ALS - Polybrominated Diphenyl Ethers (PBDEs) are flame retardant chemicals added to products so they won't catch fire or burn so easily if they are exposed to flame or high heat. | 2.2 | Y | Y | Y | 25 | 10/18/2023 | 11/26/2023 | 3/15/2024 4:40:25 PM | 2024-02-05 |  | EMS |  | Concentration | pg/L | Y | (fl. conc.) |  | BDE 207 (fl. conc.) | 437701-79-6 | Polybrominated Diphenyl Ethers (PBDEs); | Polybrominated Diphenyl Ethers (PBDEs); EMS code: B247 | Chemical |  | LAB
B250 | pg/L | BDE 208 | PBDE | PBDE-1614-HRMS-BU | PBDE-1614-HRMS-BU (USEPA 1614) Samples are liquid/liquid extracted with dichloromethane. Extracts are prepared by column chromatography, reduced in volume and analyzed by isotope-dilution GC/HRMS.(ALS - Polybrominated Diphenyl Ethers (PBDEs) are flame retardant chemicals added to products so they won't catch fire or burn so easily if they are exposed to flame or high heat. | 2.2 | Y | Y | Y | 25 | 10/18/2023 | 11/26/2023 | 3/15/2024 4:40:25 PM | 2024-02-05 |  | EMS |  | Concentration | pg/L | Y | (fl. conc.) |  | BDE 208 (fl. conc.) | 437701-78-5 | Polybrominated Diphenyl Ethers (PBDEs); | Polybrominated Diphenyl Ethers (PBDEs); EMS code: B250 | Chemical |  | LAB
B250 | pg/g | BDE 208 | E1614 | USEPA 1614 | USEPA 1614 | 10 | Y | Y | Y | 11 | 3/27/2020 | 3/30/2020 | 2/21/2023 2:10:31 PM | 2023-02-09 |  | EMS |  | MassPerMass | pg/g | Y | (mass conc.) |  | BDE 208 (mass conc.) | 437701-78-5 | Polybrominated Diphenyl Ethers (PBDEs); | Polybrominated Diphenyl Ethers (PBDEs); EMS code: B250 | Chemical |  | LAB
B253 | pg/L | BDE 209 | PBDE | PBDE-1614-HRMS-BU | PBDE-1614-HRMS-BU (USEPA 1614) Samples are liquid/liquid extracted with dichloromethane. Extracts are prepared by column chromatography, reduced in volume and analyzed by isotope-dilution GC/HRMS.(ALS - Polybrominated Diphenyl Ethers (PBDEs) are flame retardant chemicals added to products so they won't catch fire or burn so easily if they are exposed to flame or high heat. | 14 | Y | Y | Y | 25 | 10/18/2023 | 11/26/2023 | 3/15/2024 4:40:25 PM | 2024-02-05 |  | EMS |  | Concentration | pg/L | Y | (fl. conc.) |  | BDE 209 (fl. conc.) | 1163-19-5 | Polybrominated Diphenyl Ethers (PBDEs); | Polybrominated Diphenyl Ethers (PBDEs); EMS code: B253 | Chemical |  | LAB
B253 | pg/g | BDE 209 | E1614 | USEPA 1614 | USEPA 1614 | 16 | Y | Y | Y | 11 | 3/27/2020 | 3/30/2020 | 2/21/2023 2:10:31 PM | 2023-02-09 |  | EMS |  | MassPerMass | pg/g | Y | (mass conc.) |  | BDE 209 (mass conc.) | 1163-19-5 | Polybrominated Diphenyl Ethers (PBDEs); | Polybrominated Diphenyl Ethers (PBDEs); EMS code: B253 | Chemical |  | LAB
B256 | pg/L | BDE 28/33 | PBDE | PBDE-1614-HRMS-BU | PBDE-1614-HRMS-BU (USEPA 1614) Samples are liquid/liquid extracted with dichloromethane. Extracts are prepared by column chromatography, reduced in volume and analyzed by isotope-dilution GC/HRMS.(ALS - Polybrominated Diphenyl Ethers (PBDEs) are flame retardant chemicals added to products so they won't catch fire or burn so easily if they are exposed to flame or high heat. | 0.25 | Y | Y | Y | 25 | 10/18/2023 | 11/26/2023 | 3/15/2024 4:40:25 PM | 2024-02-05 |  | EMS |  | Concentration | pg/L | Y | (fl. conc.) |  | BDE 28/33 (fl. conc.) | 41318-75-6, 49690-94-0 | Polybrominated Diphenyl Ethers (PBDEs); | Polybrominated Diphenyl Ethers (PBDEs); EMS code: B256 | Chemical |  | LAB
B256 | pg/g | BDE 28/33 | E1614 | USEPA 1614 | USEPA 1614 | 0.52 | Y | Y | Y | 11 | 3/27/2020 | 3/30/2020 | 2/21/2023 2:10:31 PM | 2023-02-09 |  | EMS |  | MassPerMass | pg/g | Y | (mass conc.) |  | BDE 28/33 (mass conc.) | 41318-75-6, 49690-94-0 | Polybrominated Diphenyl Ethers (PBDEs); | Polybrominated Diphenyl Ethers (PBDEs); EMS code: B256 | Chemical |  | LAB
B259 | pg/g | BDE 30 | E1614 | USEPA 1614 | USEPA 1614 | 0.59 | Y | Y | Y | 11 | 3/27/2020 | 3/30/2020 | 2/21/2023 2:10:31 PM | 2023-02-09 |  | EMS |  | MassPerMass | pg/g | Y | (mass conc.) |  | BDE 30 (mass conc.) | 155999-95-4 | Polybrominated Diphenyl Ethers (PBDEs); | Polybrominated Diphenyl Ethers (PBDEs); EMS code: B259 | Chemical |  | LAB
B259 | pg/L | BDE 30 | PBDE | PBDE-1614-HRMS-BU | PBDE-1614-HRMS-BU (USEPA 1614) Samples are liquid/liquid extracted with dichloromethane. Extracts are prepared by column chromatography, reduced in volume and analyzed by isotope-dilution GC/HRMS.(ALS - Polybrominated Diphenyl Ethers (PBDEs) are flame retardant chemicals added to products so they won't catch fire or burn so easily if they are exposed to flame or high heat. | 0.25 | Y | Y | Y | 25 | 10/18/2023 | 11/26/2023 | 3/15/2024 4:40:25 PM | 2024-02-05 |  | EMS |  | Concentration | pg/L | Y | (fl. conc.) |  | BDE 30 (fl. conc.) | 155999-95-4 | Polybrominated Diphenyl Ethers (PBDEs); | Polybrominated Diphenyl Ethers (PBDEs); EMS code: B259 | Chemical |  | LAB
B262 | pg/L | BDE 32 | PBDE | PBDE-1614-HRMS-BU | PBDE-1614-HRMS-BU (USEPA 1614) Samples are liquid/liquid extracted with dichloromethane. Extracts are prepared by column chromatography, reduced in volume and analyzed by isotope-dilution GC/HRMS.(ALS - Polybrominated Diphenyl Ethers (PBDEs) are flame retardant chemicals added to products so they won't catch fire or burn so easily if they are exposed to flame or high heat. | 0.19 | Y | Y | Y | 25 | 10/18/2023 | 11/26/2023 | 3/15/2024 4:40:25 PM | 2024-02-05 |  | EMS |  | Concentration | pg/L | Y | (fl. conc.) |  | BDE 32 (fl. conc.) | 189084-60-4 | Polybrominated Diphenyl Ethers (PBDEs); | Polybrominated Diphenyl Ethers (PBDEs); EMS code: B262 | Chemical |  | LAB
B262 | pg/g | BDE 32 | E1614 | USEPA 1614 | USEPA 1614 | 0.42 | Y | Y | Y | 11 | 3/27/2020 | 3/30/2020 | 2/21/2023 2:10:31 PM | 2023-02-09 |  | EMS |  | MassPerMass | pg/g | Y | (mass conc.) |  | BDE 32 (mass conc.) | 189084-60-4 | Polybrominated Diphenyl Ethers (PBDEs); | Polybrominated Diphenyl Ethers (PBDEs); EMS code: B262 | Chemical |  | LAB
B265 | pg/L | BDE 35 | PBDE | PBDE-1614-HRMS-BU | PBDE-1614-HRMS-BU (USEPA 1614) Samples are liquid/liquid extracted with dichloromethane. Extracts are prepared by column chromatography, reduced in volume and analyzed by isotope-dilution GC/HRMS.(ALS - Polybrominated Diphenyl Ethers (PBDEs) are flame retardant chemicals added to products so they won't catch fire or burn so easily if they are exposed to flame or high heat. | 0.18 | Y | Y | Y | 25 | 10/18/2023 | 11/26/2023 | 3/15/2024 4:40:25 PM | 2024-02-05 |  | EMS |  | Concentration | pg/L | Y | (fl. conc.) |  | BDE 35 (fl. conc.) | 147217-80-9 | Polybrominated Diphenyl Ethers (PBDEs); | Polybrominated Diphenyl Ethers (PBDEs); EMS code: B265 | Chemical |  | LAB
B265 | pg/g | BDE 35 | E1614 | USEPA 1614 | USEPA 1614 | 0.36 | Y | Y | Y | 11 | 3/27/2020 | 3/30/2020 | 2/21/2023 2:10:31 PM | 2023-02-09 |  | EMS |  | MassPerMass | pg/g | Y | (mass conc.) |  | BDE 35 (mass conc.) | 147217-80-9 | Polybrominated Diphenyl Ethers (PBDEs); | Polybrominated Diphenyl Ethers (PBDEs); EMS code: B265 | Chemical |  | LAB
B268 | pg/g | BDE 37 | E1614 | USEPA 1614 | USEPA 1614 | 0.36 | Y | Y | Y | 11 | 3/27/2020 | 3/30/2020 | 2/21/2023 2:10:31 PM | 2023-02-09 |  | EMS |  | MassPerMass | pg/g | Y | (mass conc.) |  | BDE 37 (mass conc.) | 147217-81-0 | Polybrominated Diphenyl Ethers (PBDEs); | Polybrominated Diphenyl Ethers (PBDEs); EMS code: B268 | Chemical |  | LAB
B268 | pg/L | BDE 37 | PBDE | PBDE-1614-HRMS-BU | PBDE-1614-HRMS-BU (USEPA 1614) Samples are liquid/liquid extracted with dichloromethane. Extracts are prepared by column chromatography, reduced in volume and analyzed by isotope-dilution GC/HRMS.(ALS - Polybrominated Diphenyl Ethers (PBDEs) are flame retardant chemicals added to products so they won't catch fire or burn so easily if they are exposed to flame or high heat. | 0.18 | Y | Y | Y | 25 | 10/18/2023 | 11/26/2023 | 3/15/2024 4:40:25 PM | 2024-02-05 |  | EMS |  | Concentration | pg/L | Y | (fl. conc.) |  | BDE 37 (fl. conc.) | 147217-81-0 | Polybrominated Diphenyl Ethers (PBDEs); | Polybrominated Diphenyl Ethers (PBDEs); EMS code: B268 | Chemical |  | LAB
B271 | pg/g | BDE 47 | E1614 | USEPA 1614 | USEPA 1614 | 0.33 | Y | Y | Y | 11 | 3/27/2020 | 3/30/2020 | 2/21/2023 2:10:31 PM | 2023-02-09 |  | EMS |  | MassPerMass | pg/g | Y | (mass conc.) |  | BDE 47 (mass conc.) | 5436-43-1 | Polybrominated Diphenyl Ethers (PBDEs); | Polybrominated Diphenyl Ethers (PBDEs); EMS code: B271 | Chemical |  | LAB
B271 | pg/L | BDE 47 | PBDE | PBDE-1614-HRMS-BU | PBDE-1614-HRMS-BU (USEPA 1614) Samples are liquid/liquid extracted with dichloromethane. Extracts are prepared by column chromatography, reduced in volume and analyzed by isotope-dilution GC/HRMS.(ALS - Polybrominated Diphenyl Ethers (PBDEs) are flame retardant chemicals added to products so they won't catch fire or burn so easily if they are exposed to flame or high heat. | 0.19 | Y | Y | Y | 25 | 10/18/2023 | 11/26/2023 | 3/15/2024 4:40:25 PM | 2024-02-05 |  | EMS |  | Concentration | pg/L | Y | (fl. conc.) |  | BDE 47 (fl. conc.) | 5436-43-1 | Polybrominated Diphenyl Ethers (PBDEs); | Polybrominated Diphenyl Ethers (PBDEs); EMS code: B271 | Chemical |  | LAB
B274 | pg/g | BDE 49 | E1614 | USEPA 1614 | USEPA 1614 | 0.54 | Y | Y | Y | 11 | 3/27/2020 | 3/30/2020 | 2/21/2023 2:10:31 PM | 2023-02-09 |  | EMS |  | MassPerMass | pg/g | Y | (mass conc.) |  | BDE 49 (mass conc.) | 243982-82-3 | Polybrominated Diphenyl Ethers (PBDEs); | Polybrominated Diphenyl Ethers (PBDEs); EMS code: B274 | Chemical |  | LAB
B274 | pg/L | BDE 49 | PBDE | PBDE-1614-HRMS-BU | PBDE-1614-HRMS-BU (USEPA 1614) Samples are liquid/liquid extracted with dichloromethane. Extracts are prepared by column chromatography, reduced in volume and analyzed by isotope-dilution GC/HRMS.(ALS - Polybrominated Diphenyl Ethers (PBDEs) are flame retardant chemicals added to products so they won't catch fire or burn so easily if they are exposed to flame or high heat. | 0.23 | Y | Y | Y | 25 | 10/18/2023 | 11/26/2023 | 3/15/2024 4:40:25 PM | 2024-02-05 |  | EMS |  | Concentration | pg/L | Y | (fl. conc.) |  | BDE 49 (fl. conc.) | 243982-82-3 | Polybrominated Diphenyl Ethers (PBDEs); | Polybrominated Diphenyl Ethers (PBDEs); EMS code: B274 | Chemical |  | LAB
B277 | pg/L | BDE 51 | PBDE | PBDE-1614-HRMS-BU | PBDE-1614-HRMS-BU (USEPA 1614) Samples are liquid/liquid extracted with dichloromethane. Extracts are prepared by column chromatography, reduced in volume and analyzed by isotope-dilution GC/HRMS.(ALS - Polybrominated Diphenyl Ethers (PBDEs) are flame retardant chemicals added to products so they won't catch fire or burn so easily if they are exposed to flame or high heat. | 0.15 | Y | Y | Y | 25 | 10/18/2023 | 11/26/2023 | 3/15/2024 4:40:25 PM | 2024-02-05 |  | EMS |  | Concentration | pg/L | Y | (fl. conc.) |  | BDE 51 (fl. conc.) | 189084-57-9 | Polybrominated Diphenyl Ethers (PBDEs); | Polybrominated Diphenyl Ethers (PBDEs); EMS code: B277 | Chemical |  | LAB
B277 | pg/g | BDE 51 | E1614 | USEPA 1614 | USEPA 1614 | 0.35 | Y | Y | Y | 11 | 3/27/2020 | 3/30/2020 | 2/21/2023 2:10:31 PM | 2023-02-09 |  | EMS |  | MassPerMass | pg/g | Y | (mass conc.) |  | BDE 51 (mass conc.) | 189084-57-9 | Polybrominated Diphenyl Ethers (PBDEs); | Polybrominated Diphenyl Ethers (PBDEs); EMS code: B277 | Chemical |  | LAB
B280 | pg/L | BDE 66 | PBDE | PBDE-1614-HRMS-BU | PBDE-1614-HRMS-BU (USEPA 1614) Samples are liquid/liquid extracted with dichloromethane. Extracts are prepared by column chromatography, reduced in volume and analyzed by isotope-dilution GC/HRMS.(ALS - Polybrominated Diphenyl Ethers (PBDEs) are flame retardant chemicals added to products so they won't catch fire or burn so easily if they are exposed to flame or high heat. | 0.29 | Y | Y | Y | 25 | 10/18/2023 | 11/26/2023 | 3/15/2024 4:40:25 PM | 2024-02-05 |  | EMS |  | Concentration | pg/L | Y | (fl. conc.) |  | BDE 66 (fl. conc.) | 189084-61-5 | Polybrominated Diphenyl Ethers (PBDEs); | Polybrominated Diphenyl Ethers (PBDEs); EMS code: B280 | Chemical |  | LAB
B280 | pg/g | BDE 66 | E1614 | USEPA 1614 | USEPA 1614 | 0.54 | Y | Y | Y | 11 | 3/27/2020 | 3/30/2020 | 2/21/2023 2:10:31 PM | 2023-02-09 |  | EMS |  | MassPerMass | pg/g | Y | (mass conc.) |  | BDE 66 (mass conc.) | 189084-61-5 | Polybrominated Diphenyl Ethers (PBDEs); | Polybrominated Diphenyl Ethers (PBDEs); EMS code: B280 | Chemical |  | LAB
B283 | pg/g | BDE 7 | E1614 | USEPA 1614 | USEPA 1614 | 0.28 | Y | Y | Y | 11 | 3/27/2020 | 3/30/2020 | 2/21/2023 2:10:31 PM | 2023-02-09 |  | EMS |  | MassPerMass | pg/g | Y | (mass conc.) |  | BDE 7 (mass conc.) | 171977-44-9 | Polybrominated Diphenyl Ethers (PBDEs); | Polybrominated Diphenyl Ethers (PBDEs); EMS code: B283 | Chemical |  | LAB
B283 | pg/L | BDE 7 | PBDE | PBDE-1614-HRMS-BU | PBDE-1614-HRMS-BU (USEPA 1614) Samples are liquid/liquid extracted with dichloromethane. Extracts are prepared by column chromatography, reduced in volume and analyzed by isotope-dilution GC/HRMS.(ALS - Polybrominated Diphenyl Ethers (PBDEs) are flame retardant chemicals added to products so they won't catch fire or burn so easily if they are exposed to flame or high heat. | 0.11 | Y | Y | Y | 25 | 10/18/2023 | 11/26/2023 | 3/15/2024 4:40:25 PM | 2024-02-05 |  | EMS |  | Concentration | pg/L | Y | (fl. conc.) |  | BDE 7 (fl. conc.) | 171977-44-9 | Polybrominated Diphenyl Ethers (PBDEs); | Polybrominated Diphenyl Ethers (PBDEs); EMS code: B283 | Chemical |  | LAB
B286 | pg/L | BDE 71 | PBDE | PBDE-1614-HRMS-BU | PBDE-1614-HRMS-BU (USEPA 1614) Samples are liquid/liquid extracted with dichloromethane. Extracts are prepared by column chromatography, reduced in volume and analyzed by isotope-dilution GC/HRMS.(ALS - Polybrominated Diphenyl Ethers (PBDEs) are flame retardant chemicals added to products so they won't catch fire or burn so easily if they are exposed to flame or high heat. | 0.23 | Y | Y | Y | 25 | 10/18/2023 | 11/26/2023 | 3/15/2024 4:40:25 PM | 2024-02-05 |  | EMS |  | Concentration | pg/L | Y | (fl. conc.) |  | BDE 71 (fl. conc.) | 189084-62-6 | Polybrominated Diphenyl Ethers (PBDEs); | Polybrominated Diphenyl Ethers (PBDEs); EMS code: B286 | Chemical |  | LAB
B286 | pg/g | BDE 71 | E1614 | USEPA 1614 | USEPA 1614 | 0.51 | Y | Y | Y | 11 | 3/27/2020 | 3/30/2020 | 2/21/2023 2:10:31 PM | 2023-02-09 |  | EMS |  | MassPerMass | pg/g | Y | (mass conc.) |  | BDE 71 (mass conc.) | 189084-62-6 | Polybrominated Diphenyl Ethers (PBDEs); | Polybrominated Diphenyl Ethers (PBDEs); EMS code: B286 | Chemical |  | LAB
B289 | pg/g | BDE 75 | E1614 | USEPA 1614 | USEPA 1614 | 0.44 | Y | Y | Y | 11 | 3/27/2020 | 3/30/2020 | 2/21/2023 2:10:31 PM | 2023-02-09 |  | EMS |  | MassPerMass | pg/g | Y | (mass conc.) |  | BDE 75 (mass conc.) | 189084-63-7 | Polybrominated Diphenyl Ethers (PBDEs); | Polybrominated Diphenyl Ethers (PBDEs); EMS code: B289 | Chemical |  | LAB
B289 | pg/L | BDE 75 | PBDE | PBDE-1614-HRMS-BU | PBDE-1614-HRMS-BU (USEPA 1614) Samples are liquid/liquid extracted with dichloromethane. Extracts are prepared by column chromatography, reduced in volume and analyzed by isotope-dilution GC/HRMS.(ALS - Polybrominated Diphenyl Ethers (PBDEs) are flame retardant chemicals added to products so they won't catch fire or burn so easily if they are exposed to flame or high heat. | 0.18 | Y | Y | Y | 25 | 10/18/2023 | 11/26/2023 | 3/15/2024 4:40:25 PM | 2024-02-05 |  | EMS |  | Concentration | pg/L | Y | (fl. conc.) |  | BDE 75 (fl. conc.) | 189084-63-7 | Polybrominated Diphenyl Ethers (PBDEs); | Polybrominated Diphenyl Ethers (PBDEs); EMS code: B289 | Chemical |  | LAB
B292 | pg/g | BDE 77 | E1614 | USEPA 1614 | USEPA 1614 | 0.42 | Y | Y | Y | 11 | 3/27/2020 | 3/30/2020 | 2/21/2023 2:10:31 PM | 2023-02-09 |  | EMS |  | MassPerMass | pg/g | Y | (mass conc.) |  | BDE 77 (mass conc.) | 189084-63-7 | Polybrominated Diphenyl Ethers (PBDEs); | Polybrominated Diphenyl Ethers (PBDEs); EMS code: B292 | Chemical |  | LAB
B292 | pg/L | BDE 77 | PBDE | PBDE-1614-HRMS-BU | PBDE-1614-HRMS-BU (USEPA 1614) Samples are liquid/liquid extracted with dichloromethane. Extracts are prepared by column chromatography, reduced in volume and analyzed by isotope-dilution GC/HRMS.(ALS - Polybrominated Diphenyl Ethers (PBDEs) are flame retardant chemicals added to products so they won't catch fire or burn so easily if they are exposed to flame or high heat. | 0.18 | Y | Y | Y | 25 | 10/18/2023 | 11/26/2023 | 3/15/2024 4:40:25 PM | 2024-02-05 |  | EMS |  | Concentration | pg/L | Y | (fl. conc.) |  | BDE 77 (fl. conc.) | 189084-63-7 | Polybrominated Diphenyl Ethers (PBDEs); | Polybrominated Diphenyl Ethers (PBDEs); EMS code: B292 | Chemical |  | LAB
B297 | pg/L | BDE 79 | PBDE | PBDE-1614-HRMS-BU | PBDE-1614-HRMS-BU (USEPA 1614) Samples are liquid/liquid extracted with dichloromethane. Extracts are prepared by column chromatography, reduced in volume and analyzed by isotope-dilution GC/HRMS.(ALS - Polybrominated Diphenyl Ethers (PBDEs) are flame retardant chemicals added to products so they won't catch fire or burn so easily if they are exposed to flame or high heat. | 0.17 | Y | Y | Y | 25 | 10/18/2023 | 11/26/2023 | 3/15/2024 4:40:25 PM | 2024-02-05 |  | EMS |  | Concentration | pg/L | Y | (fl. conc.) |  | BDE 79 (fl. conc.) | 446254-48-4 | Polybrominated Diphenyl Ethers (PBDEs); | Polybrominated Diphenyl Ethers (PBDEs); EMS code: B297 | Chemical |  | LAB
B297 | pg/g | BDE 79 | E1614 | USEPA 1614 | USEPA 1614 | 0.37 | Y | Y | Y | 11 | 3/27/2020 | 3/30/2020 | 2/21/2023 2:10:31 PM | 2023-02-09 |  | EMS |  | MassPerMass | pg/g | Y | (mass conc.) |  | BDE 79 (mass conc.) | 446254-48-4 | Polybrominated Diphenyl Ethers (PBDEs); | Polybrominated Diphenyl Ethers (PBDEs); EMS code: B297 | Chemical |  | LAB
B298 | pg/L | BDE 8/11 | PBDE | PBDE-1614-HRMS-BU | PBDE-1614-HRMS-BU (USEPA 1614) Samples are liquid/liquid extracted with dichloromethane. Extracts are prepared by column chromatography, reduced in volume and analyzed by isotope-dilution GC/HRMS.(ALS - Polybrominated Diphenyl Ethers (PBDEs) are flame retardant chemicals added to products so they won't catch fire or burn so easily if they are exposed to flame or high heat. | 0.074 | Y | Y | Y | 25 | 10/18/2023 | 11/26/2023 | 3/15/2024 4:40:25 PM | 2024-02-05 |  | EMS |  | Concentration | pg/L | Y | (fl. conc.) |  | BDE 8/11 (fl. conc.) | 147217-71-8, 6903-63-5 | Polybrominated Diphenyl Ethers (PBDEs); | Polybrominated Diphenyl Ethers (PBDEs); EMS code: B298 | Chemical |  | LAB
B298 | pg/g | BDE 8/11 | E1614 | USEPA 1614 | USEPA 1614 | 0.2 | Y | Y | Y | 11 | 3/27/2020 | 3/30/2020 | 2/21/2023 2:10:31 PM | 2023-02-09 |  | EMS |  | MassPerMass | pg/g | Y | (mass conc.) |  | BDE 8/11 (mass conc.) | 147217-71-8, 6903-63-5 | Polybrominated Diphenyl Ethers (PBDEs); | Polybrominated Diphenyl Ethers (PBDEs); EMS code: B298 | Chemical |  | LAB
B301 | pg/g | BDE 85 | E1614 | USEPA 1614 | USEPA 1614 | 0.58 | Y | Y | Y | 11 | 3/27/2020 | 3/30/2020 | 2/21/2023 2:10:31 PM | 2023-02-09 |  | EMS |  | MassPerMass | pg/g | Y | (mass conc.) |  | BDE 85 (mass conc.) | 182346-21-0 | Polybrominated Diphenyl Ethers (PBDEs); | Polybrominated Diphenyl Ethers (PBDEs); EMS code: B301 | Chemical |  | LAB
B301 | pg/L | BDE 85 | PBDE | PBDE-1614-HRMS-BU | PBDE-1614-HRMS-BU (USEPA 1614) Samples are liquid/liquid extracted with dichloromethane. Extracts are prepared by column chromatography, reduced in volume and analyzed by isotope-dilution GC/HRMS.(ALS - Polybrominated Diphenyl Ethers (PBDEs) are flame retardant chemicals added to products so they won't catch fire or burn so easily if they are exposed to flame or high heat. | 0.4 | Y | Y | Y | 25 | 10/18/2023 | 11/26/2023 | 3/15/2024 4:40:25 PM | 2024-02-05 |  | EMS |  | Concentration | pg/L | Y | (fl. conc.) |  | BDE 85 (fl. conc.) | 182346-21-0 | Polybrominated Diphenyl Ethers (PBDEs); | Polybrominated Diphenyl Ethers (PBDEs); EMS code: B301 | Chemical |  | LAB
B304 | pg/L | BDE 99 | PBDE | PBDE-1614-HRMS-BU | PBDE-1614-HRMS-BU (USEPA 1614) Samples are liquid/liquid extracted with dichloromethane. Extracts are prepared by column chromatography, reduced in volume and analyzed by isotope-dilution GC/HRMS.(ALS - Polybrominated Diphenyl Ethers (PBDEs) are flame retardant chemicals added to products so they won't catch fire or burn so easily if they are exposed to flame or high heat. | 0.24 | Y | Y | Y | 25 | 10/18/2023 | 11/26/2023 | 3/15/2024 4:40:25 PM | 2024-02-05 |  | EMS |  | Concentration | pg/L | Y | (fl. conc.) |  | BDE 99 (fl. conc.) | 60348-60-9 | Polybrominated Diphenyl Ethers (PBDEs); | Polybrominated Diphenyl Ethers (PBDEs); EMS code: B304 | Chemical |  | LAB
B304 | pg/g | BDE 99 | E1614 | USEPA 1614 | USEPA 1614 | 0.48 | Y | Y | Y | 11 | 3/27/2020 | 3/30/2020 | 2/21/2023 2:10:31 PM | 2023-02-09 |  | EMS |  | MassPerMass | pg/g | Y | (mass conc.) |  | BDE 99 (mass conc.) | 60348-60-9 | Polybrominated Diphenyl Ethers (PBDEs); | Polybrominated Diphenyl Ethers (PBDEs); EMS code: B304 | Chemical |  | LAB
B307 | pg/L | HBB | PBDE | PBDE-1614-HRMS-BU | PBDE-1614-HRMS-BU (USEPA 1614) Samples are liquid/liquid extracted with dichloromethane. Extracts are prepared by column chromatography, reduced in volume and analyzed by isotope-dilution GC/HRMS.(ALS - Polybrominated Diphenyl Ethers (PBDEs) are flame retardant chemicals added to products so they won't catch fire or burn so easily if they are exposed to flame or high heat. | 0.27 | Y | Y | Y | 25 | 10/18/2023 | 11/26/2023 | 3/15/2024 4:40:25 PM | 2024-02-05 |  | EMS |  | Concentration | pg/L | Y | (fl. conc.) |  | HBB (fl. conc.) | 87-82-1 |  | EMS code: B307 | Chemical |  | LAB
B307 | pg/g | HBB | E1614 | USEPA 1614 | USEPA 1614 | 0.67 | Y | Y | Y | 11 | 3/27/2020 | 3/30/2020 | 2/21/2023 2:10:31 PM | 2023-02-09 |  | EMS |  | MassPerMass | pg/g | Y | (mass conc.) |  | HBB (mass conc.) | 87-82-1 |  | EMS code: B307 | Chemical |  | LAB
B310 | pg/L | PBEB | PBDE | PBDE-1614-HRMS-BU | PBDE-1614-HRMS-BU (USEPA 1614) Samples are liquid/liquid extracted with dichloromethane. Extracts are prepared by column chromatography, reduced in volume and analyzed by isotope-dilution GC/HRMS.(ALS - Polybrominated Diphenyl Ethers (PBDEs) are flame retardant chemicals added to products so they won't catch fire or burn so easily if they are exposed to flame or high heat. | 0.098 | Y | Y | Y | 25 | 10/18/2023 | 11/26/2023 | 3/15/2024 4:40:25 PM | 2024-02-05 |  | EMS |  | Concentration | pg/L | Y | (fl. conc.) |  | PBEB (fl. conc.) | 85-22-3 |  | EMS code: B310 | Chemical |  | LAB
B310 | pg/g | PBEB | E1614 | USEPA 1614 | USEPA 1614 | 0.56 | Y | Y | Y | 11 | 3/27/2020 | 3/30/2020 | 2/21/2023 2:10:31 PM | 2023-02-09 |  | EMS |  | MassPerMass | pg/g | Y | (mass conc.) |  | PBEB (mass conc.) | 85-22-3 |  | EMS code: B310 | Chemical |  | LAB
BA-D | mg/L | Barium Dissolved | F039 | Field filt + acid, ICP-AES | Dissolved and extractable water samples are ready for analysis once collected. Total samples are autoclaved with HNO3 and hydrochloric acid (HCI) in capped Teflon vials. Analysis is performed via Inductively Coupled Argon Plasma Atomic Emission Spectrometry (ICP-AES)(simultaneous multielement analysis), Graphite Furnace Atomic Absorption Spectrometry (GF-AAS)(low level analysis) and/or Flame Atomic Absorption Spectrometry (F-AAS). | 0.001 | Y | Y | Y | 6513 | 7/29/1996 6:03:00 PM | 7/5/2019 1:21:00 PM | 9/11/2019 8:47:50 AM | 1997-05-08 | 6/3/1998 6:51:46 PM | MPAK |  | Concentration | mg/L | Y | (fl. conc.) |  | Barium Dissolved (fl. conc.) | 7440-39-3 | Metals - Dissolved; | Metals - Dissolved; EMS code: BA-D | Chemical | Dissolved | LAB
BA-D | mg/L | Barium Dissolved | X285 | HNO3 Dig:ICP Analysis | HNO3 Dig:ICP Analysis | 0.01 | Y | Y | Y | 742 | 1/30/2002 | 9/16/2003 3:18:00 PM | 3/18/2004 10:28:33 AM | 1996-04-12 | 6/3/1998 6:51:34 PM | EMS |  | Concentration | mg/L | Y | (fl. conc.) |  | Barium Dissolved (fl. conc.) | 7440-39-3 | Metals - Dissolved; | Metals - Dissolved; EMS code: BA-D | Chemical | Dissolved | LAB
BA-D | mg/L | Barium Dissolved | 0020 | HNO3 Dig:AA Analysis | HNO3 Dig:AA Analysis | 0.1 | Y | Y | Y | 8 | 9/29/2000 | 12/1/2010 | 12/9/2010 3:55:26 PM | 1996-04-12 | 6/3/1998 6:51:34 PM | EMS |  | Concentration | mg/L | Y | (fl. conc.) |  | Barium Dissolved (fl. conc.) | 7440-39-3 | Metals - Dissolved; | Metals - Dissolved; EMS code: BA-D | Chemical | Dissolved | LAB
BA-D | mg/L | Barium Dissolved | X085 | LF,LA,HNO3,AA | Samples are lab-filtered (0.45 um), and lab-acidified (HNO3); analysis by AA. | 0.005 | Y | Y | Y | 207 | 12/12/1996 | 9/18/2015 | 10/28/2015 9:54:16 AM | 1996-04-12 | 6/3/1998 6:51:45 PM | EMS |  | Concentration | mg/L | Y | (fl. conc.) |  | Barium Dissolved (fl. conc.) | 7440-39-3 | Metals - Dissolved; | Metals - Dissolved; EMS code: BA-D | Chemical | Dissolved | LAB
BA-D | mg/L | Barium Dissolved | X556 | Metals in water, dissolved, "in-bottle" digestion, ICP-SFMS | The whole water sample is filtered through a 0.45 micron cellulose acetate filter paper and the filtrate is stored in a HDPE/Teflon bottle. The container is de-capped in a Class 100 workstation, acidified to 2% (v/v) HNO3 and 1% (v/v) HCl and recapped. The sample is "in-bottle" digested in a convection oven, set at 60 degrees C for 16 hours. The sample is cooled and introduced to an inductively coupled argon plasma-sector field mass spectrometer (ICP-SFMS) by pneumatic nebulization. Each respective element is measured at a specific mass to charge ratio, m/z value expressed in amu (atomic mass units), and at a specified intrument resolution. LR: Low Resolution ~300, MR: Mid Resolution ~4000 and HR: High Resolution ~10000. Ref: Method number 02-2003, National Laboratory for Environmental Testing, CCIW, Environment Canada, Burlington, Requested April 2003. | 5e-05 | Y | Y | Y | 4234 | 6/11/2003 | 12/6/2023 2:04:00 AM | 2/14/2024 9:05:32 AM | 2004-01-19 |  | UMULLER |  | Concentration | mg/L | Y | (fl. conc.) |  | Barium Dissolved (fl. conc.) | 7440-39-3 | Metals - Dissolved; | Metals - Dissolved; EMS code: BA-D | Chemical | Dissolved | LAB
BA-D | mg/L | Barium Dissolved | MSEA | Metals in Seawater (undigested) by (CRC) ICP-MS | This method describes the multi-elemental determination of trace elements by Collision/Reaction Cell (CRC) ICP-MS. The method measures ions produced by a radio-frequency inductively coupled plasma. Analyte species originating in a liquid are pneumatically nebulized and the resulting aerosol transported by argon gas into the plasma torch. The ions produced are entrained in the plasma gas and extracted, by means of a differentially pumped vacuum interface, into a mass spectrometer. The ions flow through a Collision/Reaction Cell (CRC) where depending on the mode of analysis either the ions passes through the cell unhindered (no gas - classical ICP-MS mode), or collisions/reactions chemistry takes place between the ions and a neutral gas such as hydrogen or helium. Any polyatomic ions which give interference on the analytes are removed. The ions are then sorted according to their mass-to-charge ratios by a quadrupole mass spectrometer. The ions are then sorted according to their mass-to-charge ratios by a quadrupole mass spectrometer. The ions transmitted through the quadrupole are quantified by a channel electron multiplier and the ion information processed by a data handling system. Although most polyatomic interference is removed, any additional interference must be assessed and valid corrections applied or the data flagged to indicate problems. Interference correction must include compensation for background ions contributed by the plasma gas, reagents, and constituents of the sample matrix. Instrument drift as well as suppressions or enhancements of instrument response caused by the sample matrix must be corrected for by the use of internal standards. | 2e-05 | Y | Y | Y | 4 | 4/12/2023 | 4/29/2024 | 5/16/2024 2:00:14 PM | 2009-09-08 |  | RDGIBSON |  | Concentration | mg/L | Y | (fl. conc.) |  | Barium Dissolved (fl. conc.) | 7440-39-3 | Metals - Dissolved; | Metals - Dissolved; EMS code: BA-D | Chemical | Dissolved | LAB
BA-D | mg/L | Barium Dissolved | MWID | Metals in Water by ICPMS (digested) | Metals in Water by ICPMS (digested) Based on EPA 200.8 | 0.001 | Y | Y | Y | 31785 | 10/26/2018 5:20:00 PM | 10/26/2018 5:20:00 PM | 7/31/2024 1:46:41 PM | 2009-12-24 |  | RDGIBSON |  | Concentration | mg/L | Y | (fl. conc.) |  | Barium Dissolved (fl. conc.) | 7440-39-3 | Metals - Dissolved; | Metals - Dissolved; EMS code: BA-D | Chemical | Dissolved | LAB
BA-D | mg/L | Barium Dissolved | F032 | Filt.,Acid,ICPMS | There is no acid digestion required for this analysis. A water sample is collected and preserved in 0.2 percent nitric acid. To analyse the extractable metals the sample is shaken, allowed to sit for at least twenty four hours, and a decanted portion is introduced to the spectrometer. When the dissolved metals are required, the samples must first be filtered through a 0.45 micron cellulose acetate filter paper and then followed by preservation in 0.2 percent nitric acid. It is important to note that different laboratories may have very different detection limits for some variables analyzed by ICPMS. | 1e-05 | Y | Y | Y | 81174 | 12/19/1996 8:15:00 AM | 8/26/2024 | 8/26/2024 4:53:05 PM | 1997-01-16 | 4/29/2011 8:47:13 AM | MPAK | RDGIBSON | Concentration | mg/L | Y | (fl. conc.) |  | Barium Dissolved (fl. conc.) | 7440-39-3 | Metals - Dissolved; | Metals - Dissolved; EMS code: BA-D | Chemical | Dissolved | LAB
BA-D | mg/L | Barium Dissolved | X562 | Dissolved metals: FF, FA, HNO3, HR-ICPMS | Ultra trace metals in water are analyzed by high resolution inductively coupled plasma mass spectrometry (HR-ICPMS), based on US EPA Method 200.8, (Oct. 1999). This procedure is intended for pristine field-filtered acid-preserved water samples. The detection limits (LOR) for this test are based on lab instrumental analysis only, not including filtration. ALS-supplied field filtration equipment does not support these LOR. Therefore, because of the high probability of false positives due to filtration, it is strongly recommended that a filtration blank be analyzed to aid in data interpretation. | 2e-05 | Y | Y | Y | 1264 | 6/8/2015 | 10/19/2023 | 11/9/2023 7:41:27 AM | 2015-10-22 |  | RDGIBSON |  | Concentration | mg/L | Y | (fl. conc.) |  | Barium Dissolved (fl. conc.) | 7440-39-3 | Metals - Dissolved; | Metals - Dissolved; EMS code: BA-D | Chemical | Dissolved | LAB
BA-D | mg/L | Barium Dissolved | F068 | HNO3dig,TeflonTube/Autoclave,ICP-MS | This method is applicable to determiniation of dissolved, extractable, and total recoverable metals in ground waters, surface waters, drinking waters, and wastewaters. Dissolved, extractable, and total recoverable sample fractions are acidified to pH <1.5 with analytical grade nitric acid (HNO3) to prevent the metals from adsorbing to the container material or precipitating. Extractable and total recoverable water samples are collected in 250 ml acid washed polyethylene bottles and are preserved with Nitric acid (0.5 ml acid for 100 mLs of sample). All preserved samples are stored at room temperature and can be held for up to 6 months. Unfiltered, un-preserved samples are stored at 4C for no more than 30 days. For dissolved and extractable water fractions, no sample preparation is necessary. Samples are simply poured in autosampler tubes (i.e. 15 ml centrifuge tubes), and placed on the autosampler and analyzed. Samples for total recoverable metals are digested with Nitric acid in special Teflon tubes in an autoclave prior to analyses. | 2e-05 | Y | Y | Y | 1397 | 6/1/1999 1:19:00 PM | 4/17/2015 1:06:00 PM | 8/11/2021 12:04:11 PM | 1999-06-28 |  | NPEPPIN |  | Concentration | mg/L | Y | (fl. conc.) |  | Barium Dissolved (fl. conc.) | 7440-39-3 | Metals - Dissolved; | Metals - Dissolved; EMS code: BA-D | Chemical | Dissolved | LAB
BA-D | mg/L | Barium Dissolved | EQ01 | EQUIS (mg/L) | EQUIS (mg/L) | 0.01 | N | Y | Y | 967 |  |  | 10/23/1996 12:07:03 AM | 1996-04-12 | 6/3/1998 6:51:34 PM | EMS |  | Concentration | mg/L | Y | (fl. conc.) |  | Barium Dissolved (fl. conc.) | 7440-39-3 | Metals - Dissolved; | Metals - Dissolved; EMS code: BA-D | Chemical | Dissolved | LAB
BA-D | mg/L | Barium Dissolved | X170 | ICP | This code is a catch-all for a large variety of ICP methods including water and also solids digested into solutions as an intermediate result. Different procedures will have greatly different detection limits. (a) For air particulates, preliminary results are reported in mg/L and loading results as mg/dm2/d. (b) For Hi Vol samples preliminiary results are same as reported in mg/L and loading results as ug/m3. (c) Note that for Rain samples there is no filtration nor digestion steps. | 0.02 | Y | Y | Y | 19293 | 4/11/1996 6:01:00 PM | 11/6/2018 | 8/26/2024 2:41:32 PM | 1996-04-12 | 8/15/2019 11:52:29 AM | EMS | KATFRASE | Concentration | mg/L | Y | (fl. conc.) |  | Barium Dissolved (fl. conc.) | 7440-39-3 | Metals - Dissolved; | Metals - Dissolved; EMS code: BA-D | Chemical | Dissolved | LAB
BA-D | mg/L | Barium Dissolved | METC | MET-D-NP-CCMS-VA | Ultra trace metals in water are analyzed by CRC ICPMS, based on US EPA Method 6020B (July 2014). This procedure is intended for pristine field-filtered acid-preserved water samples. The detection limits (LOR) for this test are based on lab instrumental analysis only, not including filtration. ALS-supplied field filtration equipment does not support these LOR. Therefore, because of the high probability of false positives due to filtration, it is strongly recommended that a filtration blank be analysed to aid in data interpretation. | 2e-05 | Y | Y | Y | 1776 | 4/3/2012 | 8/15/2024 | 8/23/2024 9:50:35 AM | 2019-08-20 |  | KATFRASE |  | Concentration | mg/L | Y | (fl. conc.) |  | Barium Dissolved (fl. conc.) | 7440-39-3 | Metals - Dissolved; | Metals - Dissolved; EMS code: BA-D | Chemical | Dissolved | LAB
BA-E | mg/L | Barium Extractable | F047 | 24hr HNO3 dil'n,ICP-MS | There is no acid digestion required for this analysis. A water sample is collected and preserved in 0.2 percent nitric acid. To analyse the extractable metals the sample is shaken, allowed to sit for at least twenty four hours, and a decanted portion is introduced to the spectrometer. When the dissolved metals are required, the samples must first be filtered through a 0.45 micron cellulose acetate filter paper and then followed by preservation in 0.2 percent nitric acid. It is important to note that different laboratories may have very different detection limits for some variables analyzed by ICP/MS. | 1e-05 | Y | Y | Y | 2678 | 4/16/1998 10:21:00 AM | 12/13/2011 12:02:00 PM | 9/11/2019 8:46:17 AM | 1997-10-20 | 6/3/1998 6:51:47 PM | NPEPPIN |  | Concentration | mg/L | Y | (fl. conc.) |  | Barium Extractable (fl. conc.) | 7440-39-3 | Metals - Extractable; | Metals - Extractable; EMS code: BA-E | Chemical | Extractable | LAB
BA-E | mg/L | Barium Extractable | F069 | SEM extr,6N HCl,ICP/OES | 6.0 N Hydrochloric acid (HCI) is added to an aliquot of the sediment sample. The extract produced from the addition of the acid is then analyzed for simultaneously extracted metals (SEM) using inductively coupled plasma - optical emission spectrophotometry. The detection limits depend on the metal parameters in question and the instrumentation used. | 0.01 | Y | Y | Y | 10 | 3/26/2012 3:01:00 AM | 3/26/2012 3:01:00 AM | 11/25/2021 1:50:15 PM | 1999-07-09 |  | NPEPPIN |  | Concentration | mg/L | Y | (fl. conc.) |  | Barium Extractable (fl. conc.) | 7440-39-3 | Metals - Extractable; | Metals - Extractable; EMS code: BA-E | Chemical | Extractable | LAB
BA-E | mg/L | Barium Extractable | X553 | Metals in water, extractable, ICP-SFMS | The whole water sample is collected in a HDPE/Teflon bottle. In a Class 100 workstation it is preserved with concentrated Nitric Acid at concentration of 0.2 percent (v/v). The acid preservative extracts the metals from material in the sample and a decanted portion is introduced directly to an inductively coupled argon plasma-sector field mass spectrometer (ICP-SFMS) by pneumatic nebulization. Each respective element is measured at a specific mass to charge ratio, m/z value expressed in amu (atomic mass units), and at a specified instrument resolution. LR: Low Resolution ~300, MR: Mid Resolution ~4000 and HR: High Resolution ~10000. Ref: Method number 02-2061, National Laboratory for Environmental Testing, CCIW, Environment Canada, Burlington. Requested April 2003. | 5e-05 | Y | Y | Y | 2596 | 5/22/2003 | 9/21/2007 | 11/28/2023 2:11:53 PM | 2004-01-19 |  | UMULLER |  | Concentration | mg/L | Y | (fl. conc.) |  | Barium Extractable (fl. conc.) | 7440-39-3 | Metals - Extractable; | Metals - Extractable; EMS code: BA-E | Chemical | Extractable | LAB
BA-E | mg/L | Barium Extractable | F040 | Field Acid, lab filt, ICP-AES | Dissolved and extractable water samples are ready for analysis once collected. Total samples are autoclaved with HNO3 and hydrochloric acid (HCI) in capped Teflon vials. Analysis is performed via Inductively Coupled Argon Plasma Atomic Emission Spectrometry (ICP-AES)(simultaneous multielement analysis), Graphite Furnace Atomic Absorption Spectrometry (GF-AAS)(low level analysis) and/or Flame Atomic Absorption Spectrometry (F-AAS). | 0.001 | Y | Y | Y | 2361 | 8/5/1997 6:06:00 PM | 3/14/2008 | 9/11/2019 8:47:41 AM | 1997-05-08 | 6/3/1998 6:51:46 PM | MPAK |  | Concentration | mg/L | Y | (fl. conc.) |  | Barium Extractable (fl. conc.) | 7440-39-3 | Metals - Extractable; | Metals - Extractable; EMS code: BA-E | Chemical | Extractable | LAB
BA-O | % | Barium (as BaO) | ICP06 | ALS Minerals ME-ICP06 | ALS Minerals ME-ICP06. A prepared sample is fused with a lithium tetraborate/lithium metaborate flux and poured into a graphite crucible and fused in a furnace at 1000°C. The resulting melt is then cooled and dissolved in 100 mL of 4% HNO3/2% HCl followed by ICP-AES. | 0.01 | Y | Y | Y | 0 |  |  |  | 2024-01-29 |  | EMS |  | DimensionlessRatio | % | N | (ratio) |  | Barium (as BaO) (ratio) |  |  | EMS code: BA-O | Chemical |  | LAB
BA-T | mg/L | Barium Total | 0040 | HNO3 Dig: ICP Analysis | HNO3 Dig: ICP Analysis | 0.01 | Y | Y | Y | 7400 | 5/8/2003 9:40:00 AM | 6/25/2010 3:15:00 PM | 7/22/2010 2:09:06 PM | 1996-04-12 | 6/3/1998 6:51:34 PM | EMS |  | Concentration | mg/L | Y | (fl. conc.) |  | Barium Total (fl. conc.) | 7440-39-3 | Metals - Total; | Metals - Total; EMS code: BA-T | Chemical | Total | LAB
BA-T | ppm (S) | Barium Total | 6042 | Aqua Regia by ICPMS | These were sediment/soils digested with Aqua-Regia (Nitric Acid & Hydrochloric Acids) and then analysis by ICPMS. Results would be reported in ppm s. | 10 | Y | Y | Y | 71 | 5/14/2022 | 7/27/2024 | 7/29/2024 9:13:08 AM | 2022-06-28 |  | EMS |  | DimensionlessRatio | ppm | N | (ratio) |  | Barium Total (ratio) | 7440-39-3 | Metals - Total; | Metals - Total; EMS code: BA-T | Chemical | Total | LAB
BA-T | mg/L | Barium Total | MSEA | Metals in Seawater (undigested) by (CRC) ICP-MS | This method describes the multi-elemental determination of trace elements by Collision/Reaction Cell (CRC) ICP-MS. The method measures ions produced by a radio-frequency inductively coupled plasma. Analyte species originating in a liquid are pneumatically nebulized and the resulting aerosol transported by argon gas into the plasma torch. The ions produced are entrained in the plasma gas and extracted, by means of a differentially pumped vacuum interface, into a mass spectrometer. The ions flow through a Collision/Reaction Cell (CRC) where depending on the mode of analysis either the ions passes through the cell unhindered (no gas - classical ICP-MS mode), or collisions/reactions chemistry takes place between the ions and a neutral gas such as hydrogen or helium. Any polyatomic ions which give interference on the analytes are removed. The ions are then sorted according to their mass-to-charge ratios by a quadrupole mass spectrometer. The ions are then sorted according to their mass-to-charge ratios by a quadrupole mass spectrometer. The ions transmitted through the quadrupole are quantified by a channel electron multiplier and the ion information processed by a data handling system. Although most polyatomic interference is removed, any additional interference must be assessed and valid corrections applied or the data flagged to indicate problems. Interference correction must include compensation for background ions contributed by the plasma gas, reagents, and constituents of the sample matrix. Instrument drift as well as suppressions or enhancements of instrument response caused by the sample matrix must be corrected for by the use of internal standards. | 2e-05 | Y | Y | Y | 441 | 8/12/2009 12:08:00 AM | 4/29/2024 | 6/7/2024 10:03:16 AM | 2009-09-08 |  | RDGIBSON |  | Concentration | mg/L | Y | (fl. conc.) |  | Barium Total (fl. conc.) | 7440-39-3 | Metals - Total; | Metals - Total; EMS code: BA-T | Chemical | Total | LAB
BA-T | ug/g | Barium Total | FB04 | Freeze Dry,Hom,Acid Dig,ICP | Sample moisture is determined by weighing samples before and after freeze-drying. Freeze-dried samples are ground in a Braun blender before being weighed into Teflon PFA liners and digested with concentrated HNO3, concentrated HCl, and 50% H2O2 in a microwave oven; which results in a complete sample dissolution. Samples are collected and volumized with Type I water. Analysis is performed via Inductively Coupled Argon Plasma Atomic Emission Spectrometry (ICP-AES) (simultaneous multielement analysis), Graphite Furnace Atomic Absorption Spectrometry (GF-AAS) (low level Cd and Pb analysis), Cold Vapor Atomic Fluorescence Spectrometry (CV-AFS) (Mercury analysis) and/or Flame Atomic Absorption Spectrometry (F-AAS) (K analysis). | 0.08 | Y | Y | Y | 245 | 8/22/1996 6:11:00 PM | 12/1/2010 | 12/9/2010 3:55:26 PM | 1997-02-27 | 6/3/1998 6:51:46 PM | MPAK |  | MassPerMass | mg/kg | Y | (mass conc.) |  | Barium Total (mass conc.) | 7440-39-3 | Metals - Total; | Metals - Total; EMS code: BA-T | Chemical | Total | LAB
BA-T | mg/L | Barium Total | 0200 | 20%A.R.Dig:AA Anal. | 20%A.R.Dig:AA Anal. | 0.5 | Y | Y | Y | 116 |  |  | 10/24/1996 9:48:13 AM | 1996-04-12 | 6/3/1998 6:51:41 PM | EMS |  | Concentration | mg/L | Y | (fl. conc.) |  | Barium Total (fl. conc.) | 7440-39-3 | Metals - Total; | Metals - Total; EMS code: BA-T | Chemical | Total | LAB
BA-T | ug/g | Barium Total | ICPM | Metals in Soil by CRC ICPMS | This analysis is carried out using procedures from CSR Analytical Method: "Strong Acid Leachable Metals (SALM) in Soil", BC Ministry of Environment, 26 June 2009, and procedures adapted from EPA Method 200.2. The sample is manually homogenized, dried at 60 degrees Celsius, sieved through a 63 um (230 mesh) sieve, and a representative subsample of the dry material is weighed. The sample is then digested at 95 degrees Celsius for 2 hours by block digester using concentrated nitric and hydrochloric acids. Instrumental analysis is by collision cell inductively coupled plasma - mass spectrometry (modifed from EPA Method 6020A). Method Limitation: This method is not a total digestion technique. It is a very strong acid digestion that is intended to dissolve those metals that may be environmentally available. By design, elements bound in silicate structures are not normally dissolved by this procedure as they are not usually mobile in the environment. | 0.5 | Y | Y | Y | 103 | 9/19/2017 | 9/24/2023 | 1/12/2024 11:16:52 AM | 2018-01-09 |  | JKROGH |  | MassPerMass | mg/kg | Y | (mass conc.) |  | Barium Total (mass conc.) | 7440-39-3 | Metals - Total; | Metals - Total; EMS code: BA-T | Chemical | Total | LAB
BA-T | mg/L | Barium Total | X555 | Metals in water, total, "in-bottle" digestion, ICP-SFMS | The whole water sample collected in a HDPE/Teflon bottle, is de-capped in a Class 100 workstation, acidified to 2% (v/v) HNO3 and 1% (v/v) HCl and recapped. The sample is "in-bottle" digested in a convection oven, set at 60 degrees C for 16 hours. The sample is cooled and introduced to an inductively coupled argon plasma-sector field mass spectrometer (ICP-SFMS) by pneumatic nebulization. Each respective element is measured at a specific mass to charge ratio, m/z value expressed in amu (atomic mass units), and at resolution. LR: Low Resolution ~300, MR: Mid Resolution ~4000 and HR: High a specified instrument Resolution ~10000. Ref: Method number 02-2003, National Laboratory for Environmental Testing, CCIW, Environment Canada, Burlington. Requested April 2003. | 5e-05 | Y | Y | Y | 17343 | 5/20/2003 | 12/6/2023 2:08:00 AM | 2/14/2024 9:05:32 AM | 2004-01-14 |  | UMULLER |  | Concentration | mg/L | Y | (fl. conc.) |  | Barium Total (fl. conc.) | 7440-39-3 | Metals - Total; | Metals - Total; EMS code: BA-T | Chemical | Total | LAB
BA-T | ug/g | Barium Total | FS01 | D&G,Nitric/HCl:ICP | Samples are dried at 60 C, then either sieved to -100mesh (150 um) and rolled to homogenize, or finely ground with a mortar and pestle or a shatter box laboratory mill and rolled (depending on the submitter's requirements). The samples are then weighed into Teflon digestion liners and digested with HNO3 and HCl in a microwave oven. Samples are cooled, volumized, and allowed to settle overnight before being decanted. Analysis is then performed via Inductively Coupled Argon Plasma Atomic Emission Spectrometry (ICP-AES) (Simultaneous multi element analysis). In ICP Spectrometry, the liquid sample is pumped to the nebulizer, where it is convertd into an aerosol. the nebulizer carrier gas flow than transports the wet aerosol past the impact bead, where the large droplets condense out and flow to the waste container. The remaining fine aerosol is carried into the plasma. There, the high temperature desolvates the aerosol and dissociates the sample material into individual atoms and ions, which are excited to emit light at wavelengths characteristic of the elements present. A diffraction grating in the spectrometer diffracts the various wavelengths of light through exit slits to the photo-multiplier tubes, which convert the light energy to electronic signals. | 0.2 | Y | Y | Y | 2642 | 8/30/1996 1:52:00 PM | 10/10/2007 | 5/24/2008 1:15:02 AM | 1996-11-20 | 6/3/1998 6:51:46 PM | MPAK |  | MassPerMass | mg/kg | Y | (mass conc.) |  | Barium Total (mass conc.) | 7440-39-3 | Metals - Total; | Metals - Total; EMS code: BA-T | Chemical | Total | LAB
BA-T | ug/g | Barium Total | X492 | Homo,HCl-HNO3 dig,ICP-OES,dry calc | Homogenize samples and take sub-samples. For Total Recoverable Metals Digest with Aqua Regia. Analyses by ICPOES (Inductively Coupled Plasma Optical Emission Spectroscopy). Perform moisture analysis on separate sub-sample and calculate dry value for trace elements. | 1 | Y | Y | Y | 1 |  |  | 8/24/2016 2:59:13 PM | 1998-11-13 |  | NPEPPIN |  | MassPerMass | mg/kg | Y | (mass conc.) |  | Barium Total (mass conc.) | 7440-39-3 | Metals - Total; | Metals - Total; EMS code: BA-T | Chemical | Total | LAB
BA-T | mg/dm2/d | Barium Total | DUST | HNO3 dig.(microwave or hot block),ICPMS | HNO3 dig.(microwave or hot block);ICPMS |  | Y | Y | Y | 420 | 7/19/2016 1:11:00 PM | 7/31/2024 | 7/31/2024 11:23:12 AM | 2008-02-07 |  | RDGIBSON |  | MassFlux | mg/dm²/d | Y | (mass flux) |  | Barium Total (mass flux) | 7440-39-3 | Metals - Total; | Metals - Total; EMS code: BA-T | Chemical | Total | LAB
BA-T | ug/g wet | Barium Total | X552 | Metals in Biota (wet), total, acid digestion, ICP-SFMS | One gram of wet, homogenized biota is weighed in a microwave TFM vessel, and 8mL of Nitric Acid and 1mL of Hydrogen Peroxide are added. The sample is digested at 200 degrees C for 15 minutes in a closed vessel high pressure microwave oven. The sample is cooled, the volume adjusted to 100mL and introduced to an inductively coupled argon plasma-sector field mass spectrometer (ICP-SFMS) by pneumatic nebulization. Each respective element is measured at a specific mass to charge ratio, m/z value expressed in amu (atomic mass units), and at a specified instrument resolution. LR: Low Resolution ~300, MR: Mid Resolution ~4000. Ref: 02-2704, National Laboratory for Environmental Testing, CCIW, Environment Canada, Burlington. Requested April 2003. | 0.005 | Y | Y | Y | 37 | 8/9/2016 | 10/16/2023 | 11/28/2023 2:21:32 PM | 2004-01-19 | 4/19/2021 1:22:00 PM | UMULLER | EMS | MassPerMass | µg/g | Y | (mass conc.) | (wet) | Barium Total (wet) (mass conc.) | 7440-39-3 | Metals - Total; | Metals - Total; EMS code: BA-T | Chemical | Total | LAB
BA-T | mg/L | Barium Total | F048 | Acid dig,conc'n,nebul.,ICP-OES | This method provides an automated procedure for the Inductively Coupled Argon Plasma Optical Emission Spectrophotometric (ICP-OES) determination and quantitation of trace amounts of metals in surface, ground and precipitation waters after a manual acid digestion and concentration procedure and ultrasonic nebulisation. Organically-bound metals and metal complexes are converted to metal ions by digestion with nitric acid [method details says HCl]. To increase the signal detected by the ICP, the samples are concentrated 10 times. The liquid sample is pumped onto the piezoelectric transducer, where the ultrasonic action efficiently converts it to a fine, dense aerosol. The nebulizer carrier gas flow then transports the wet aerosol through a heated U-tube, where the solvent is vaporized. Subsequently, most of the solvent vapour is removed from the aerosol stream by passing the stream through a chilled condenser. The result is a dry, analyte-laden aerosol which is introduced to the plasma. There, the high temperature desolvates the aerosol and dissociates the sample material into individual atoms and ions, which are excited to emit light at wave lengths characteristic of the elements present. Dissolved samples are field filtered through 0.45 micron cellulose acetate filter prior to acid preservation. Samples for extractable analysis are acidified with 2 mL concentrated HNO3 per litre and allowed to stand a minimum of 24 hours before analysis. Clients providing these samples are advised to preserve samples at the collection site. | 0.0001 | Y | Y | Y | 6159 | 7/4/1997 2:35:00 PM | 4/23/2003 | 9/11/2019 8:46:38 AM | 1997-10-20 | 6/3/1998 6:51:47 PM | NPEPPIN |  | Concentration | mg/L | Y | (fl. conc.) |  | Barium Total (fl. conc.) | 7440-39-3 | Metals - Total; | Metals - Total; EMS code: BA-T | Chemical | Total | LAB
BA-T | mg/L | Barium Total | X145 | Dig, ICP | Dig; ICP | 0.01 | Y | Y | Y | 654 |  |  | 10/1/2003 9:40:52 AM | 1996-04-12 | 6/3/1998 6:51:34 PM | EMS |  | Concentration | mg/L | Y | (fl. conc.) |  | Barium Total (fl. conc.) | 7440-39-3 | Metals - Total; | Metals - Total; EMS code: BA-T | Chemical | Total | LAB
BA-T | mg/L | Barium Total | F068 | HNO3dig,TeflonTube/Autoclave,ICP-MS | This method is applicable to determiniation of dissolved, extractable, and total recoverable metals in ground waters, surface waters, drinking waters, and wastewaters. Dissolved, extractable, and total recoverable sample fractions are acidified to pH <1.5 with analytical grade nitric acid (HNO3) to prevent the metals from adsorbing to the container material or precipitating. Extractable and total recoverable water samples are collected in 250 ml acid washed polyethylene bottles and are preserved with Nitric acid (0.5 ml acid for 100 mLs of sample). All preserved samples are stored at room temperature and can be held for up to 6 months. Unfiltered, un-preserved samples are stored at 4C for no more than 30 days. For dissolved and extractable water fractions, no sample preparation is necessary. Samples are simply poured in autosampler tubes (i.e. 15 ml centrifuge tubes), and placed on the autosampler and analyzed. Samples for total recoverable metals are digested with Nitric acid in special Teflon tubes in an autoclave prior to analyses. | 2e-05 | Y | Y | Y | 1777 | 6/1/1999 1:23:00 PM | 7/4/2008 9:48:00 AM | 7/30/2024 6:32:12 AM | 1999-06-28 |  | NPEPPIN |  | Concentration | mg/L | Y | (fl. conc.) |  | Barium Total (fl. conc.) | 7440-39-3 | Metals - Total; | Metals - Total; EMS code: BA-T | Chemical | Total | LAB
BA-T | ug/g | Barium Total | FS82 | PBM,dry60C,10mesh,SALM-ICP-MS | PBM;dry60C;10mesh;SALM-ICP-MS | 0.1 | Y | Y | Y | 3648 | 5/21/2002 | 7/24/2024 | 7/29/2024 5:17:09 PM | 2002-11-12 |  | UMULLER |  | MassPerMass | mg/kg | Y | (mass conc.) |  | Barium Total (mass conc.) | 7440-39-3 | Metals - Total; | Metals - Total; EMS code: BA-T | Chemical | Total | LAB
BA-T | mg/L | Barium Total | METN | MET-T-NP-CCMS-VA | Ultra trace metals in water are analyzed by CRC ICPMS, based on US EPA Method 6020B (July 2014). The detection limits provided can only be met for undigested samples. This procedure is intended for colorless, non-turbid, acid-preserved water samples (i.e. pristine water samples), having turbidity < 1 NTU and no odor. Where turbidity exceeds 1 NTU, and/or the sample is colored and has an odor, results may be biased low compared to true Total Metals concentrations. ALS recommends that turbidity analysis be requested on samples submitted for this test to aid with interpretation of results. Where turbidity is <1NTU, undigested metals are equivalent to total metals concentrations. | 2e-05 | Y | Y | Y | 3648 | 4/3/2012 | 8/19/2024 | 8/26/2024 1:03:06 PM | 2019-08-20 |  | KATFRASE |  | Concentration | mg/L | Y | (fl. conc.) |  | Barium Total (fl. conc.) | 7440-39-3 | Metals - Total; | Metals - Total; EMS code: BA-T | Chemical | Total | LAB
BA-T | mg/L | Barium Total | X170 | ICP | This code is a catch-all for a large variety of ICP methods including water and also solids digested into solutions as an intermediate result. Different procedures will have greatly different detection limits. (a) For air particulates, preliminary results are reported in mg/L and loading results as mg/dm2/d. (b) For Hi Vol samples preliminiary results are same as reported in mg/L and loading results as ug/m3. (c) Note that for Rain samples there is no filtration nor digestion steps. | 0.02 | Y | Y | Y | 10995 | 3/28/1996 6:03:00 PM | 11/7/2018 | 8/26/2024 2:41:32 PM | 1996-04-12 | 8/15/2019 11:52:56 AM | EMS | KATFRASE | Concentration | mg/L | Y | (fl. conc.) |  | Barium Total (fl. conc.) | 7440-39-3 | Metals - Total; | Metals - Total; EMS code: BA-T | Chemical | Total | LAB
BA-T | mg/L | Barium Total | F033 | Acid Dig,Centrifuge,ICPMS | For the determination of total recoverable elements in water or waste water, take 100 ml aliquot from a well mixed, acid preserved, sample containing not more than 0.25% (w/v) total solids and transfer to a 250 ml Griffin beaker. (If total solids are greater than 0.25% reduce the size of the aliquot by a proportionate amount). Add 1 ml of conc. nitric acid and 0.5 ml conc. hydrochloric acid. Heat on a hot plate until the volume has been reduced to 15-20ml, ensuring that the sample does not boil. Allow to cool and transfer to a 50 ml volumetric flask. Dilute to volume with ASTM type I water. Centrifuge the sample or allow to stand overnight to separate insoluble material. Prior analysis, pipette 20 ml into a 50 ml volumetric flask and dilute to volume with ASTM type I water. Analysis should be performed within seventy two hours of the completed preparation. Note that different laboratories may have very different field preservation, digestion procedures and detection limits for this method, so these should be documented. | 0.0001 | Y | Y | Y | 40410 | 12/16/1996 7:42:00 AM | 8/22/2024 3:24:00 AM | 8/25/2024 9:47:36 PM | 1997-01-16 | 6/3/1998 6:51:46 PM | MPAK |  | Concentration | mg/L | Y | (fl. conc.) |  | Barium Total (fl. conc.) | 7440-39-3 | Metals - Total; | Metals - Total; EMS code: BA-T | Chemical | Total | LAB
BA-T | ug/g wet | Barium Total | BPSY | MET-WET-BPSY-CCMS-VA | Metals in Animal Tissue and Vegetation (Biota) - Prescriptive. EPA 200.3/6020B This method is conducted following British Columbia Lab Manual method "Metals in Animal Tissue and Vegetation (Biota) - Prescriptive". Tissue samples are digested with nitric and hydrochloric acids; in combination with addition of hydrogen peroxide. Instrumental analysis is by collision cell inductively coupled plasma - mass spectrometry (modified from EPA Method 6020B). | 0.01 | Y | Y | Y | 2 | 11/7/2019 | 11/7/2019 | 6/10/2021 1:57:10 PM | 2021-06-04 |  | EMS |  | MassPerMass | µg/g | Y | (mass conc.) | (wet) | Barium Total (wet) (mass conc.) | 7440-39-3 | Metals - Total; | Metals - Total; EMS code: BA-T | Chemical | Total | LAB
BA-T | ug/g | Barium Total | X551 | Metals in Biota (dry), total, acid digestion, ICP-SFMS | 0.1 gram of dried, homogenized biota is weighed in a microwave TFM vessel, and 8mL of Nitric Acid and 1mL of Hydrogen Peroxide are added. The sample is digested at 200 degrees C for 15 minutes in a closed vessel high pressure microwave oven. The sample is cooled, the volume adjusted to 100mL and introduced to an iductively coupled argon plasma-sector field mass spectrometer (ICP-SFMS) by pneumatic nebulization. Each respective element is measured at a specific mass to charge ratio, m/z value expressed in amu (atomic mass units), and at a specified instrument resolution. LR: Low Resolution ~300, MR: Mid Resolution ~4000 and HR: Hich Resolution ~10000. Ref: Method number 02-2705, National Laboratory for Environmental Testing, CCIW, Environment Canada, Burlington. Requested April 2003. | 0.05 | Y | Y | Y | 49 | 8/11/2015 | 10/30/2023 | 11/28/2023 2:21:33 PM | 2004-01-19 |  | UMULLER |  | MassPerMass | mg/kg | Y | (mass conc.) |  | Barium Total (mass conc.) | 7440-39-3 | Metals - Total; | Metals - Total; EMS code: BA-T | Chemical | Total | LAB
BA-T | ug/g | Barium Total | 0196 | D&G: LTA/HCl: ICP | D&G: LTA/HCl: ICP | 1 | Y | Y | Y | 27 |  |  | 10/24/1996 2:08:12 PM | 1996-04-12 | 6/3/1998 6:51:34 PM | EMS |  | MassPerMass | mg/kg | Y | (mass conc.) |  | Barium Total (mass conc.) | 7440-39-3 | Metals - Total; | Metals - Total; EMS code: BA-T | Chemical | Total | LAB
BA-T | ug/g | Barium Total | PB02 | PBM,HNO3/HCl dig,dry60C,100mesh, | Performance Based Method. Samples are dried at 60 C and sieved using a .149 mm (100 mesh) sieve. Solid samples are digested with the combination of nitric and hydrochloric acids with a 1:1 ratio of acids. For environmental purposes, strong acid leaches are used routinely since they have the advantage of speed, low cost, and superio reproducibility. All metals trapped within the refractory matrix are unlikely to have environmental consequences. Instrumental analysis of extracts produced using this digestion method can be performed by a variety of techniques. |  | Y | Y | Y | 52 | 5/24/2001 3:42:00 PM | 11/26/2001 10:23:00 AM | 5/26/2003 2:27:23 PM | 2001-07-18 |  | NPEPPIN |  | MassPerMass | mg/kg | Y | (mass conc.) |  | Barium Total (mass conc.) | 7440-39-3 | Metals - Total; | Metals - Total; EMS code: BA-T | Chemical | Total | LAB
BA-T | mg/L | Barium Total | TCLP | Toxicity Characteristic Leaching Procedure | The TCLP procedure is designed to simulate the climatic leaching action | 0.1 | Y | Y | Y | 38 | 12/8/2009 7:49:00 PM | 9/28/2011 5:53:00 PM | 8/24/2016 2:59:13 PM | 2009-03-05 |  | JCHUDAK |  | Concentration | mg/L | Y | (fl. conc.) |  | Barium Total (fl. conc.) | 7440-39-3 | Metals - Total; | Metals - Total; EMS code: BA-T | Chemical | Total | LAB
BA-T | ug/g | Barium Total | 0192 | Hom, Nitric/Perchlor,ICP | Sample already in a dry state so does not require drying before homogenization and nitric perchlorid digestion followed by ICP analysis. | 1 | Y | Y | Y | 708 |  |  | 10/24/1996 12:50:25 PM | 1996-04-12 | 6/3/1998 6:51:34 PM | EMS |  | MassPerMass | mg/kg | Y | (mass conc.) |  | Barium Total (mass conc.) | 7440-39-3 | Metals - Total; | Metals - Total; EMS code: BA-T | Chemical | Total | LAB
BA-T | ug/g | Barium Total | FS89 | Homo,HNO3 dig,ICP-AES | Homo;HNO3 dig;ICP-AES | 0.2 | Y | Y | Y | 4 | 1/5/2001 | 8/20/2001 12:00:00 PM | 10/23/2002 1:42:40 PM | 2001-03-08 |  | NPEPPIN |  | MassPerMass | mg/kg | Y | (mass conc.) |  | Barium Total (mass conc.) | 7440-39-3 | Metals - Total; | Metals - Total; EMS code: BA-T | Chemical | Total | LAB
BA-T | mg/L | Barium Total | EQ01 | EQUIS (mg/L) | EQUIS (mg/L) | 0.01 | N | Y | Y | 69 |  |  | 10/22/1996 11:26:55 PM | 1996-04-12 | 6/3/1998 6:51:34 PM | EMS |  | Concentration | mg/L | Y | (fl. conc.) |  | Barium Total (fl. conc.) | 7440-39-3 | Metals - Total; | Metals - Total; EMS code: BA-T | Chemical | Total | LAB
BA-T | ug/g | Barium Total | X230 | Hom,HNO3/HCl,ICP | The sample is dried, ground to pass 100 mesh, and digested with a nitric-hydrochloric acid mixture to solubilize the solid matter and to remove the organic material by oxidation and volatilization. The completed digests are submitted to metals instrumentation personnel for metals analysis by Inductively Coupled Atomic Emission Spectrophotometry (ICP-AES), or hydride analysis by atomic absorption. | 50 | Y | Y | Y | 121 | 2/12/2002 | 12/30/2005 | 2/3/2006 1:13:12 AM | 1996-04-12 | 6/3/1998 6:51:34 PM | EMS |  | MassPerMass | mg/kg | Y | (mass conc.) |  | Barium Total (mass conc.) | 7440-39-3 | Metals - Total; | Metals - Total; EMS code: BA-T | Chemical | Total | LAB
BA-T | ppm (S) | Barium Total | 6041 | Aqua Regia by ICPMS | These were sediment/soils digested with Aqua-Regia (Nitric Acid & Hydrochloric Acids) and then analysis by ICPMS. Results would be reported in mg/kg and/or %. | 0.5 | Y | Y | Y | 21 | 10/13/2019 | 8/5/2020 | 8/10/2020 3:52:06 PM | 2019-09-06 | 10/21/2019 11:54:17 AM | KATFRASE | KATFRASE | DimensionlessRatio | ppm | N | (ratio) |  | Barium Total (ratio) | 7440-39-3 | Metals - Total; | Metals - Total; EMS code: BA-T | Chemical | Total | LAB
BA-T | mg/L | Barium Total | 0020 | HNO3 Dig:AA Analysis | HNO3 Dig:AA Analysis | 0.1 | Y | Y | Y | 85 |  |  | 8/26/2004 11:46:09 AM | 1996-04-12 | 6/3/1998 6:51:34 PM | EMS |  | Concentration | mg/L | Y | (fl. conc.) |  | Barium Total (fl. conc.) | 7440-39-3 | Metals - Total; | Metals - Total; EMS code: BA-T | Chemical | Total | LAB
BA-T | ppm (S) | Barium Total | MS61 | ALS Minerals ME-MS61 | ALS Minerals ME-MS61. A prepared sample is digested with HNO3; HF; and HClO4 in a graphite heating block to near dryness. After cooling; the sample is boiled in diluted HCl. Bulk to volume with deionized water; mixed and analyzed by ICP-AES and ICP-MS for all elements. | 10 | Y | Y | Y | 0 |  |  |  | 2024-01-29 |  | EMS |  | DimensionlessRatio | ppm | N | (ratio) |  | Barium Total (ratio) | 7440-39-3 | Metals - Total; | Metals - Total; EMS code: BA-T | Chemical | Total | LAB
BA-T | ppm (S) | Barium Total | MS41 | ALS Minerals ME-MS41 | ALS Minerals ME-MS41. A prepared sample is digested with aqua regia in a graphite heating block. After cooling; the resulting solution is diluted to with deionized water; mixed and analyzed by ICP-AES. Samples are then analyzed by ICP-MS for the remaining suite of elements. | 10 | Y | Y | Y | 0 |  |  |  | 2024-01-29 |  | EMS |  | DimensionlessRatio | ppm | N | (ratio) |  | Barium Total (ratio) | 7440-39-3 | Metals - Total; | Metals - Total; EMS code: BA-T | Chemical | Total | LAB
BA-T | mg/L | Barium Total | MWID | Metals in Water by ICPMS (digested) | Metals in Water by ICPMS (digested) Based on EPA 200.8 | 0.001 | Y | Y | Y | 36810 |  |  | 7/31/2024 1:46:41 PM | 2009-12-24 |  | RDGIBSON |  | Concentration | mg/L | Y | (fl. conc.) |  | Barium Total (fl. conc.) | 7440-39-3 | Metals - Total; | Metals - Total; EMS code: BA-T | Chemical | Total | LAB
BA-T | mg/L | Barium Total | X285 | HNO3 Dig:ICP Analysis | HNO3 Dig:ICP Analysis | 0.01 | Y | Y | Y | 690 | 8/10/1993 | 10/5/1993 | 4/30/2004 1:59:45 PM | 1998-05-22 | 6/3/1998 6:51:49 PM | NPEPPIN |  | Concentration | mg/L | Y | (fl. conc.) |  | Barium Total (fl. conc.) | 7440-39-3 | Metals - Total; | Metals - Total; EMS code: BA-T | Chemical | Total | LAB
BA-T | mg/L | Barium Total | 0042 | GI HNO3 Dig: ICP | Sample aliquot taken from unfiltered unpreserved General Ions bottle; HNO3 digestion; ICP analysis. | 0.001 | Y | Y | Y | 6204 | 9/20/1996 | 7/5/2019 3:21:00 PM | 7/12/2019 6:49:13 PM | 1996-04-12 | 6/3/1998 6:51:34 PM | EMS |  | Concentration | mg/L | Y | (fl. conc.) |  | Barium Total (fl. conc.) | 7440-39-3 | Metals - Total; | Metals - Total; EMS code: BA-T | Chemical | Total | LAB
BA-T | mg/L | Barium Total | F121 | Acid,ICPMS | Acidified, no digestion, ICPMS | 2e-05 | Y | Y | Y | 8436 | 3/12/2010 7:57:00 PM | 8/15/2024 9:10:00 PM | 8/26/2024 8:06:03 AM | 2002-07-17 | 7/16/2012 3:27:54 PM | UMULLER | RDGIBSON | Concentration | mg/L | Y | (fl. conc.) |  | Barium Total (fl. conc.) | 7440-39-3 | Metals - Total; | Metals - Total; EMS code: BA-T | Chemical | Total | LAB
BA-T | mg/L | Barium Total | X561 | Total metals: FA, HNO3, HR-ICPMS (No Prep)Prep) | Ultra trace metals in water are analyzed by high resolution inductively coupled plasma mass spectrometry (HR-ICPMS), modified from US EPA 200.8, (Revision 5.5). The detection limits provided can only be met for undigested samples. This procedure is intended for colorless, non-turbid, acid-preserved water samples (i.e. pristine water samples), having turbidity < 1 NTU and no odor. Where turbidity exceeds 1 NTU, and/or the sample is colored and has an odor, results may be biased low compared to true Total Metals concentrations. ALS recommends that turbidity analysis be requested on samples submitted for this test to aid with interpretation of results. | 2e-05 | Y | Y | Y | 2756 | 6/8/2015 | 10/5/2023 | 1/31/2024 2:37:41 PM | 2015-10-22 |  | RDGIBSON |  | Concentration | mg/L | Y | (fl. conc.) |  | Barium Total (fl. conc.) | 7440-39-3 | Metals - Total; | Metals - Total; EMS code: BA-T | Chemical | Total | LAB
BA-T | ug/m3 | Barium Total | F088 | Aliquot,HNO3/HCl dig, ICPMS,loading | Aliquot microwave digested with 3:1 HNO3/HCL. Analysis by ICPMS. For HiVol samples preliminary result are reported in mg/L and loading results are reported in ug/m3 |  | Y | Y | Y | 557 | 8/17/2000 10:46:00 AM | 12/14/2006 | 12/30/2006 1:01:35 AM | 2000-08-16 |  | NPEPPIN |  | AirConcentration | µg/m³ | Y | (air conc.) |  | Barium Total (air conc.) | 7440-39-3 | Metals - Total; | Metals - Total; EMS code: BA-T | Chemical | Total | LAB
BA-T | ug/m3 | Barium Total | 6040 | Aqua Regia Dig:ICP | Aqua regia digestion is used to bring the metals into solution prior to ICP analysis. For Hi-Vol samples, preliminary results are reported as mg/L and24 hr loadings as ug/m3 | 0.0005 | Y | Y | Y | 437 | 2/13/2002 | 2/9/2005 | 2/23/2005 1:13:47 AM | 2004-02-10 |  | UMULLER |  | AirConcentration | µg/m³ | Y | (air conc.) |  | Barium Total (air conc.) | 7440-39-3 | Metals - Total; | Metals - Total; EMS code: BA-T | Chemical | Total | LAB
BA-T | mg/L | Barium Total | X349 | FA:HNO3:ICP | Samples are field-acidified with HNO3; HNO3 digestion; analysis by ICP/AES. | 0.01 | Y | Y | Y | 30 |  |  | 10/30/2014 8:31:17 AM | 1998-05-22 | 6/3/1998 6:51:49 PM | NPEPPIN |  | Concentration | mg/L | Y | (fl. conc.) |  | Barium Total (fl. conc.) | 7440-39-3 | Metals - Total; | Metals - Total; EMS code: BA-T | Chemical | Total | LAB
BA-T | mg/L | Barium Total | F082 | Aliquot,HNO3/HCL dig,ICPMS | Aliquot microwave digested with 3:1 HNO3/HCL. Analysis by ICPMS. For HiVol samples preliminary result are reported in mg/L and loading results are reported in ug/m3 |  | Y | Y | Y | 27657 | 7/26/2005 7:04:00 AM | 8/23/2024 | 8/26/2024 4:53:05 PM | 2000-06-13 |  | NPEPPIN |  | Concentration | mg/L | Y | (fl. conc.) |  | Barium Total (fl. conc.) | 7440-39-3 | Metals - Total; | Metals - Total; EMS code: BA-T | Chemical | Total | LAB
BA-T | ug/g wet | Barium Total | MTIW | MET-WET-CCMS-MID-VA | Metals in Tissue by CRC ICPMS (WET) Methodology: This method is conducted following British Columbia Lab Manual method "Metals in Animal Tissue and Vegetation (Biota) - Prescriptive". Tissue samples are homogenized and sub-sampled prior to hotblock digestion with nitric and hydrochloric acids, in combination with addition of hydrogen peroxide. Instrumental analysis is by collision cell inductively coupled plasma - mass spectrometry (modified from EPA Method 6020A). Method Limitation: This method employs a strong acid/peroxide digestion, and is intended to provide a conservative estimate of bio-available metals. Near complete recoveries are achieved for most toxicologically important metals, but elements associated with recalcitrant minerals may be only partially recovered. | 0.01 | Y | Y | Y | 24 | 11/25/2018 | 12/6/2018 | 7/21/2021 2:55:07 PM | 2018-12-24 |  | EMS |  | MassPerMass | µg/g | Y | (mass conc.) | (wet) | Barium Total (wet) (mass conc.) | 7440-39-3 | Metals - Total; | Metals - Total; EMS code: BA-T | Chemical | Total | LAB
BA-T | ug/g | Barium Total | EQ02 | EQUIS (ug/g) | EQUIS (ug/g) | 1 | N | Y | Y | 760 |  |  | 10/23/1996 12:08:28 AM | 1996-04-12 | 6/3/1998 6:51:34 PM | EMS |  | MassPerMass | mg/kg | Y | (mass conc.) |  | Barium Total (mass conc.) | 7440-39-3 | Metals - Total; | Metals - Total; EMS code: BA-T | Chemical | Total | LAB
BA-T | mg/L | Barium Total | 1349 | Tot Metals in Seawater by CRC ICPMS (BC) | Tot Metals in Seawater by CRC ICPMS (BC) | 0.001 | Y | Y | Y | 98 | 8/16/2019 | 12/5/2021 | 12/13/2021 12:34:16 PM | 2019-09-30 |  | KATFRASE |  | Concentration | mg/L | Y | (fl. conc.) |  | Barium Total (fl. conc.) | 7440-39-3 | Metals - Total; | Metals - Total; EMS code: BA-T | Chemical | Total | LAB
BA-T | ug/g | Barium Total | MTID | MET-DRY-CCMS-VA | Metals in Tissue by CRC ICPMS (DRY) Methodology: This method is conducted following British Columbia Lab Manual method "Metals in Animal Tissue and Vegetation (Biota) - Prescriptive". Tissue samples are homogenized and sub-sampled prior to hotblock digestion with nitric and hydrochloric acids, in combination with addition of hydrogen peroxide. Instrumental analysis is by collision cell inductively coupled plasma - mass spectrometry (modified from EPA Method 6020A). Method Limitation: This method employs a strong acid/peroxide digestion, and is intended to provide a conservative estimate of bio-available metals. Near complete recoveries are achieved for most toxicologically important metals, but elements associated with recalcitrant minerals may be only partially recovered. | 0.05 | Y | Y | Y | 122 | 8/13/2014 | 9/13/2022 | 10/26/2022 7:14:14 AM | 2015-03-09 | 3/11/2015 4:12:05 PM | RDGIBSON | EMS | MassPerMass | mg/kg | Y | (mass conc.) |  | Barium Total (mass conc.) | 7440-39-3 | Metals - Total; | Metals - Total; EMS code: BA-T | Chemical | Total | LAB
BA-T | ug/g wet | Barium Total | X560 | Metals in Biota (wet), total, acid digestion, ICP-MS | Five grams of wet, homogenized biota is weighed in a centrifuge tube, and 2.5 mL of Nitric Acid and 2.5 mL of Hydrochloric acid are added. The sample is digested at room temperature for one hour and then digested at 90-95 degrees C in a water bath or heat block for one more hour. The sample is cooled and the volume is made to 50 mL with DI. The digestate is then diluted at ten fold and introduced to a collision/reaction cell inductively coupled argon plasma mass spectrometer (CRC ICPMS) by pneumatic nebulization. Each respective element is measured at a specific mass to charge ratio, m/z value expressed in amu (atomic mass units). - Maxxam Analytics (BC). Added to EMS Dec 10, 2008. | 0.1 | Y | Y | Y | 541 | 11/15/2008 4:00:00 AM | 10/30/2023 | 11/28/2023 2:21:33 PM | 2008-12-11 | 3/11/2015 4:17:10 PM | RDGIBSON | EMS | MassPerMass | µg/g | Y | (mass conc.) | (wet) | Barium Total (wet) (mass conc.) | 7440-39-3 | Metals - Total; | Metals - Total; EMS code: BA-T | Chemical | Total | LAB
BA-T | ug/g | Barium Total | 0190 | D&G:N/P Dig:ICP Anal. | Sample is dried and grind-mixed before nitric perchloric digestion followed by ICP analysis. | 1 | Y | Y | Y | 949 |  |  | 10/24/1996 3:06:16 PM | 1996-04-12 | 6/3/1998 6:51:41 PM | EMS |  | MassPerMass | mg/kg | Y | (mass conc.) |  | Barium Total (mass conc.) | 7440-39-3 | Metals - Total; | Metals - Total; EMS code: BA-T | Chemical | Total | LAB
BA-T | ppm (S) | Barium Total | MS81 | ALS Minerals: ME-MS81 | ALS Minerals: ME-MS81. A prepared sample (0.100 g) is added to lithium metaborate/lithium tetraborate flux; mixed well and fused in a furnace at 1025°C. The resulting melt is then cooled and dissolved in an acid mixture containing nitric; hydrochloric and hydrofluoric acids. This solution is then analyzed by inductively coupled plasma - mass spectrometry. | 0.5 | Y | Y | Y | 0 |  |  |  | 2024-01-29 |  | EMS |  | DimensionlessRatio | ppm | N | (ratio) |  | Barium Total (ratio) | 7440-39-3 | Metals - Total; | Metals - Total; EMS code: BA-T | Chemical | Total | LAB
BA-T | mg/L | Barium Total | F038 | HNO3/HCl dig,ICP-AES | Dissolved and extractable water samples are ready for analysis once collected. Total samples are autoclaved with HNO3 and hydrochloric acid (HCI) in capped Teflon vials. Analysis is performed via Inductively Coupled Argon Plasma Atomic Emission Spectrometry (ICP-AES)(simultaneous multielement analysis), Graphite Furnace Atomic Absorption Spectrometry (GF-AAS) (low level analysis) and/or Flame Atomic Absorption Spectrometry (F-AAS). | 0.001 | Y | Y | Y | 5400 | 8/1/1996 6:04:00 PM | 5/6/2008 | 9/11/2019 8:47:42 AM | 1997-05-08 | 6/3/1998 6:51:46 PM | MPAK |  | Concentration | mg/L | Y | (fl. conc.) |  | Barium Total (fl. conc.) | 7440-39-3 | Metals - Total; | Metals - Total; EMS code: BA-T | Chemical | Total | LAB
BBTF | % | 2-Bromobenzotrifluoride | TEHW | Total Ext-Hydrocarbons in Water by GCFID | Total Ext-Hydrocarbons in Water by GCFID, BC Lab Manual | 1 | Y | Y | Y | 14 | 5/21/2015 | 5/21/2015 | 3/10/2023 3:28:43 PM | 2015-05-26 |  | RDGIBSON |  | DimensionlessRatio | % | N | (ratio) |  | 2-Bromobenzotrifluoride (ratio) | 392-83-6 |  | EMS code: BBTF | Chemical |  | LAB
BBTF | % | 2-Bromobenzotrifluoride | SURR | Surrogate Recovery | Surrogate Recovery | 1 | Y | Y | Y | 1499 | 5/5/2015 | 4/16/2022 | 4/18/2022 1:17:01 PM | 2015-08-10 | 4/21/2022 1:43:21 PM | RDGIBSON | EMS | DimensionlessRatio | % | N | (ratio) |  | 2-Bromobenzotrifluoride (ratio) | 392-83-6 |  | EMS code: BBTF | Chemical |  | SURROGATE_RESULT
BBTFEP | % | Bromobenzotrifluoride, 2 (EPH surr) | SURR | Surrogate Recovery | Surrogate Recovery |  | Y | Y | Y | 793 | 12/18/2019 | 8/21/2024 | 8/22/2024 5:27:05 PM | 2022-04-21 | 4/21/2022 1:43:21 PM | EMS | EMS | DimensionlessRatio | % | N | (ratio) |  | Bromobenzotrifluoride, 2 (EPH surr) (ratio) | 392-83-6 |  | EMS code: BBTFEP | Chemical |  | SURROGATE_RESULT
BBTFF2 | % | Bromobenzotrifluoride, 2-F2-F4 | SURR | Surrogate Recovery | Surrogate Recovery |  | Y | Y | Y | 29 | 12/30/2021 | 2/7/2022 | 1/11/2024 1:07:07 PM | 2022-04-21 | 4/21/2022 1:43:21 PM | EMS | EMS | DimensionlessRatio | % | N | (ratio) |  | Bromobenzotrifluoride, 2-F2-F4 (ratio) | 392-83-6 |  | EMS code: BBTFF2 | Chemical |  | SURROGATE_RESULT
BCOL | CFU/100mL | Background Colonies | X022 | Membrane Filter Count | Non-turbid water samples are passed through a 0.45 um membrane filter which is placed on a selective medium for 18 - 24 hours incubation at 35 degrees C for growth of total coliform colonies, of fecal coliform colonies, fecal streptococci colonies, enterococci colonies, or of E. coli colonies. | 200 | Y | Y | Y | 229 | 7/3/2013 1:05:00 PM | 12/7/2023 | 12/29/2023 9:48:33 AM | 2013-06-26 |  | RDGIBSON |  | MicrobialFormation | CFU/100mL | N | (microb.) |  | Background Colonies (microb.) |  | Bacteriological; | Bacteriological; EMS code: BCOL | Chemical |  | LAB
BE-B | mg/L | beta-BHC | X268 | CH2Cl2 ext:GC/MS | CH2Cl2 ext:GC/MS |  | Y | Y | Y | 16 | 12/29/2021 | 9/2/2022 | 11/10/2022 9:54:06 AM | 2022-04-08 |  | EMS |  | Concentration | mg/L | Y | (fl. conc.) |  | beta-BHC (fl. conc.) | 319-85-7 | Pesticides; | Pesticides; EMS code: BE-B | Chemical |  | LAB
BE-B | mg/L | beta-BHC | MW03 | OCP1-LL-SF-ECD-VA | This analysis is carried out using procedures adapted from "Test Methods for Evaluating Solid Waste" SW-846, Methods 3510, 3630, 3660 and 8081, published by the United States Environmental Protection Agency (EPA). The procedure involves extraction of the entire water sample with dichloromethane. The extract then undergoes one or more of the following clean-up procedures (if required): silica gel clean-up and/or sulphur clean-up. The final extract is analyzed by dual capillary column gas chromatography with electron capture detection (GC/ECD). | 0.0001 | Y | Y | Y | 10 | 9/20/2016 | 9/21/2016 | 8/15/2019 2:42:34 PM | 2019-07-08 |  | KATFRASE |  | Concentration | mg/L | Y | (fl. conc.) |  | beta-BHC (fl. conc.) | 319-85-7 | Pesticides; | Pesticides; EMS code: BE-B | Chemical |  | LAB
BE-B | mg/L | beta-BHC | E660 | EPA 8270E (mod) | EPA 8270E (mod) Miscellaneous Pesticides by GC-MS | 5e-06 | Y | Y | Y | 2 | 9/14/2022 | 9/14/2022 | 12/8/2022 9:48:13 AM | 2022-12-06 |  | EMS |  | Concentration | mg/L | Y | (fl. conc.) |  | beta-BHC (fl. conc.) | 319-85-7 | Pesticides; | Pesticides; EMS code: BE-B | Chemical |  | LAB
BE-B | ug/g | beta-BHC | ZN03 | DCM+Ac.Extr:GC/MS | DCM+Ac.Extr:GC/MS EPA 8270E (mod) | 0.02 | Y | Y | Y | 17 | 1/17/2022 | 11/14/2022 | 1/11/2024 1:07:07 PM | 2022-06-23 |  | EMS |  | MassPerMass | mg/kg | Y | (mass conc.) |  | beta-BHC (mass conc.) | 319-85-7 | Pesticides; | Pesticides; EMS code: BE-B | Chemical |  | LAB
BE-B | ug/g | beta-BHC | E8081 | EPA 8081 | See methods EPA 3540, 3630, 3660, 8081 | 0.002 | Y | Y | Y | 9 | 3/18/2020 | 3/23/2020 | 2/22/2023 12:51:03 PM | 2022-07-29 |  | EMS |  | MassPerMass | mg/kg | Y | (mass conc.) |  | beta-BHC (mass conc.) | 319-85-7 | Pesticides; | Pesticides; EMS code: BE-B | Chemical |  | LAB
BE-D | mg/L | Beryllium Dissolved | MWID | Metals in Water by ICPMS (digested) | Metals in Water by ICPMS (digested) Based on EPA 200.8 | 0.0002 | Y | Y | Y | 31771 | 10/26/2018 5:20:00 PM | 10/26/2018 5:20:00 PM | 7/31/2024 1:46:41 PM | 2009-12-24 |  | RDGIBSON |  | Concentration | mg/L | Y | (fl. conc.) |  | Beryllium Dissolved (fl. conc.) | 7440-41-7 | Metals - Dissolved; | Metals - Dissolved; EMS code: BE-D | Chemical | Dissolved | LAB
BE-D | mg/L | Beryllium Dissolved | METC | MET-D-NP-CCMS-VA | Ultra trace metals in water are analyzed by CRC ICPMS, based on US EPA Method 6020B (July 2014). This procedure is intended for pristine field-filtered acid-preserved water samples. The detection limits (LOR) for this test are based on lab instrumental analysis only, not including filtration. ALS-supplied field filtration equipment does not support these LOR. Therefore, because of the high probability of false positives due to filtration, it is strongly recommended that a filtration blank be analysed to aid in data interpretation. | 1e-05 | Y | Y | Y | 1776 | 4/3/2012 | 8/15/2024 | 8/23/2024 9:50:35 AM | 2019-08-20 |  | KATFRASE |  | Concentration | mg/L | Y | (fl. conc.) |  | Beryllium Dissolved (fl. conc.) | 7440-41-7 | Metals - Dissolved; | Metals - Dissolved; EMS code: BE-D | Chemical | Dissolved | LAB
BE-D | mg/L | Beryllium Dissolved | F068 | HNO3dig,TeflonTube/Autoclave,ICP-MS | This method is applicable to determiniation of dissolved, extractable, and total recoverable metals in ground waters, surface waters, drinking waters, and wastewaters. Dissolved, extractable, and total recoverable sample fractions are acidified to pH <1.5 with analytical grade nitric acid (HNO3) to prevent the metals from adsorbing to the container material or precipitating. Extractable and total recoverable water samples are collected in 250 ml acid washed polyethylene bottles and are preserved with Nitric acid (0.5 ml acid for 100 mLs of sample). All preserved samples are stored at room temperature and can be held for up to 6 months. Unfiltered, un-preserved samples are stored at 4C for no more than 30 days. For dissolved and extractable water fractions, no sample preparation is necessary. Samples are simply poured in autosampler tubes (i.e. 15 ml centrifuge tubes), and placed on the autosampler and analyzed. Samples for total recoverable metals are digested with Nitric acid in special Teflon tubes in an autoclave prior to analyses. | 2e-06 | Y | Y | Y | 1398 | 6/1/1999 1:19:00 PM | 4/17/2015 1:06:00 PM | 8/11/2021 12:04:11 PM | 1999-06-28 |  | NPEPPIN |  | Concentration | mg/L | Y | (fl. conc.) |  | Beryllium Dissolved (fl. conc.) | 7440-41-7 | Metals - Dissolved; | Metals - Dissolved; EMS code: BE-D | Chemical | Dissolved | LAB
BE-D | mg/L | Beryllium Dissolved | MSEA | Metals in Seawater (undigested) by (CRC) ICP-MS | This method describes the multi-elemental determination of trace elements by Collision/Reaction Cell (CRC) ICP-MS. The method measures ions produced by a radio-frequency inductively coupled plasma. Analyte species originating in a liquid are pneumatically nebulized and the resulting aerosol transported by argon gas into the plasma torch. The ions produced are entrained in the plasma gas and extracted, by means of a differentially pumped vacuum interface, into a mass spectrometer. The ions flow through a Collision/Reaction Cell (CRC) where depending on the mode of analysis either the ions passes through the cell unhindered (no gas - classical ICP-MS mode), or collisions/reactions chemistry takes place between the ions and a neutral gas such as hydrogen or helium. Any polyatomic ions which give interference on the analytes are removed. The ions are then sorted according to their mass-to-charge ratios by a quadrupole mass spectrometer. The ions are then sorted according to their mass-to-charge ratios by a quadrupole mass spectrometer. The ions transmitted through the quadrupole are quantified by a channel electron multiplier and the ion information processed by a data handling system. Although most polyatomic interference is removed, any additional interference must be assessed and valid corrections applied or the data flagged to indicate problems. Interference correction must include compensation for background ions contributed by the plasma gas, reagents, and constituents of the sample matrix. Instrument drift as well as suppressions or enhancements of instrument response caused by the sample matrix must be corrected for by the use of internal standards. | 3e-05 | Y | Y | Y | 4 | 4/12/2023 | 4/29/2024 | 5/16/2024 2:00:14 PM | 2009-09-08 |  | RDGIBSON |  | Concentration | mg/L | Y | (fl. conc.) |  | Beryllium Dissolved (fl. conc.) | 7440-41-7 | Metals - Dissolved; | Metals - Dissolved; EMS code: BE-D | Chemical | Dissolved | LAB
BE-D | mg/L | Beryllium Dissolved | X170 | ICP | This code is a catch-all for a large variety of ICP methods including water and also solids digested into solutions as an intermediate result. Different procedures will have greatly different detection limits. (a) For air particulates, preliminary results are reported in mg/L and loading results as mg/dm2/d. (b) For Hi Vol samples preliminiary results are same as reported in mg/L and loading results as ug/m3. (c) Note that for Rain samples there is no filtration nor digestion steps. | 0.001 | Y | Y | Y | 7471 | 4/11/1996 6:01:00 PM | 11/6/2018 | 6/12/2020 1:21:11 PM | 1996-04-12 | 6/3/1998 6:51:34 PM | EMS |  | Concentration | mg/L | Y | (fl. conc.) |  | Beryllium Dissolved (fl. conc.) | 7440-41-7 | Metals - Dissolved; | Metals - Dissolved; EMS code: BE-D | Chemical | Dissolved | LAB
BE-D | mg/L | Beryllium Dissolved | F039 | Field filt + acid, ICP-AES | Dissolved and extractable water samples are ready for analysis once collected. Total samples are autoclaved with HNO3 and hydrochloric acid (HCI) in capped Teflon vials. Analysis is performed via Inductively Coupled Argon Plasma Atomic Emission Spectrometry (ICP-AES)(simultaneous multielement analysis), Graphite Furnace Atomic Absorption Spectrometry (GF-AAS)(low level analysis) and/or Flame Atomic Absorption Spectrometry (F-AAS). | 0.001 | Y | Y | Y | 5976 | 7/29/1996 6:03:00 PM | 7/5/2019 1:21:00 PM | 9/11/2019 8:47:50 AM | 1997-05-08 | 6/3/1998 6:51:46 PM | MPAK |  | Concentration | mg/L | Y | (fl. conc.) |  | Beryllium Dissolved (fl. conc.) | 7440-41-7 | Metals - Dissolved; | Metals - Dissolved; EMS code: BE-D | Chemical | Dissolved | LAB
BE-D | mg/L | Beryllium Dissolved | X556 | Metals in water, dissolved, "in-bottle" digestion, ICP-SFMS | The whole water sample is filtered through a 0.45 micron cellulose acetate filter paper and the filtrate is stored in a HDPE/Teflon bottle. The container is de-capped in a Class 100 workstation, acidified to 2% (v/v) HNO3 and 1% (v/v) HCl and recapped. The sample is "in-bottle" digested in a convection oven, set at 60 degrees C for 16 hours. The sample is cooled and introduced to an inductively coupled argon plasma-sector field mass spectrometer (ICP-SFMS) by pneumatic nebulization. Each respective element is measured at a specific mass to charge ratio, m/z value expressed in amu (atomic mass units), and at a specified intrument resolution. LR: Low Resolution ~300, MR: Mid Resolution ~4000 and HR: High Resolution ~10000. Ref: Method number 02-2003, National Laboratory for Environmental Testing, CCIW, Environment Canada, Burlington, Requested April 2003. | 1e-06 | Y | Y | Y | 4233 | 6/11/2003 | 12/6/2023 2:04:00 AM | 2/14/2024 9:05:32 AM | 2004-01-19 |  | UMULLER |  | Concentration | mg/L | Y | (fl. conc.) |  | Beryllium Dissolved (fl. conc.) | 7440-41-7 | Metals - Dissolved; | Metals - Dissolved; EMS code: BE-D | Chemical | Dissolved | LAB
BE-D | mg/L | Beryllium Dissolved | F032 | Filt.,Acid,ICPMS | There is no acid digestion required for this analysis. A water sample is collected and preserved in 0.2 percent nitric acid. To analyse the extractable metals the sample is shaken, allowed to sit for at least twenty four hours, and a decanted portion is introduced to the spectrometer. When the dissolved metals are required, the samples must first be filtered through a 0.45 micron cellulose acetate filter paper and then followed by preservation in 0.2 percent nitric acid. It is important to note that different laboratories may have very different detection limits for some variables analyzed by ICPMS. | 0.001 | Y | Y | Y | 74899 | 3/10/1998 2:03:00 PM | 8/26/2024 | 8/26/2024 4:53:05 PM | 1997-10-20 | 8/15/2019 10:53:04 AM | NPEPPIN | KATFRASE | Concentration | mg/L | Y | (fl. conc.) |  | Beryllium Dissolved (fl. conc.) | 7440-41-7 | Metals - Dissolved; | Metals - Dissolved; EMS code: BE-D | Chemical | Dissolved | LAB
BE-D | mg/L | Beryllium Dissolved | X285 | HNO3 Dig:ICP Analysis | HNO3 Dig:ICP Analysis | 0.001 | Y | Y | Y | 669 | 1/30/2002 | 11/13/2002 | 3/18/2004 10:28:34 AM | 1998-05-22 | 6/3/1998 6:51:49 PM | NPEPPIN |  | Concentration | mg/L | Y | (fl. conc.) |  | Beryllium Dissolved (fl. conc.) | 7440-41-7 | Metals - Dissolved; | Metals - Dissolved; EMS code: BE-D | Chemical | Dissolved | LAB
BE-D | mg/L | Beryllium Dissolved | X562 | Dissolved metals: FF, FA, HNO3, HR-ICPMS | Ultra trace metals in water are analyzed by high resolution inductively coupled plasma mass spectrometry (HR-ICPMS), based on US EPA Method 200.8, (Oct. 1999). This procedure is intended for pristine field-filtered acid-preserved water samples. The detection limits (LOR) for this test are based on lab instrumental analysis only, not including filtration. ALS-supplied field filtration equipment does not support these LOR. Therefore, because of the high probability of false positives due to filtration, it is strongly recommended that a filtration blank be analyzed to aid in data interpretation. | 1e-05 | Y | Y | Y | 1264 | 6/8/2015 | 10/19/2023 | 11/9/2023 7:41:27 AM | 2015-10-22 |  | RDGIBSON |  | Concentration | mg/L | Y | (fl. conc.) |  | Beryllium Dissolved (fl. conc.) | 7440-41-7 | Metals - Dissolved; | Metals - Dissolved; EMS code: BE-D | Chemical | Dissolved | LAB
BE-E | mg/L | Beryllium Extractable | X553 | Metals in water, extractable, ICP-SFMS | The whole water sample is collected in a HDPE/Teflon bottle. In a Class 100 workstation it is preserved with concentrated Nitric Acid at concentration of 0.2 percent (v/v). The acid preservative extracts the metals from material in the sample and a decanted portion is introduced directly to an inductively coupled argon plasma-sector field mass spectrometer (ICP-SFMS) by pneumatic nebulization. Each respective element is measured at a specific mass to charge ratio, m/z value expressed in amu (atomic mass units), and at a specified instrument resolution. LR: Low Resolution ~300, MR: Mid Resolution ~4000 and HR: High Resolution ~10000. Ref: Method number 02-2061, National Laboratory for Environmental Testing, CCIW, Environment Canada, Burlington. Requested April 2003. | 1e-06 | Y | Y | Y | 2597 | 5/22/2003 | 9/21/2007 | 11/28/2023 2:11:53 PM | 2004-01-19 |  | UMULLER |  | Concentration | mg/L | Y | (fl. conc.) |  | Beryllium Extractable (fl. conc.) | 7440-41-7 | Metals - Extractable; | Metals - Extractable; EMS code: BE-E | Chemical | Extractable | LAB
BE-E | mg/L | Beryllium Extractable | F047 | 24hr HNO3 dil'n,ICP-MS | There is no acid digestion required for this analysis. A water sample is collected and preserved in 0.2 percent nitric acid. To analyse the extractable metals the sample is shaken, allowed to sit for at least twenty four hours, and a decanted portion is introduced to the spectrometer. When the dissolved metals are required, the samples must first be filtered through a 0.45 micron cellulose acetate filter paper and then followed by preservation in 0.2 percent nitric acid. It is important to note that different laboratories may have very different detection limits for some variables analyzed by ICP/MS. | 2e-06 | Y | Y | Y | 2676 | 4/16/1998 10:21:00 AM | 6/26/2003 | 9/11/2019 8:46:17 AM | 1997-10-20 | 6/3/1998 6:51:47 PM | NPEPPIN |  | Concentration | mg/L | Y | (fl. conc.) |  | Beryllium Extractable (fl. conc.) | 7440-41-7 | Metals - Extractable; | Metals - Extractable; EMS code: BE-E | Chemical | Extractable | LAB
BE-E | mg/L | Beryllium Extractable | F040 | Field Acid, lab filt, ICP-AES | Dissolved and extractable water samples are ready for analysis once collected. Total samples are autoclaved with HNO3 and hydrochloric acid (HCI) in capped Teflon vials. Analysis is performed via Inductively Coupled Argon Plasma Atomic Emission Spectrometry (ICP-AES)(simultaneous multielement analysis), Graphite Furnace Atomic Absorption Spectrometry (GF-AAS)(low level analysis) and/or Flame Atomic Absorption Spectrometry (F-AAS). | 0.001 | Y | Y | Y | 2350 | 8/5/1997 6:06:00 PM | 3/14/2008 | 9/11/2019 8:47:41 AM | 1997-05-08 | 6/3/1998 6:51:46 PM | MPAK |  | Concentration | mg/L | Y | (fl. conc.) |  | Beryllium Extractable (fl. conc.) | 7440-41-7 | Metals - Extractable; | Metals - Extractable; EMS code: BE-E | Chemical | Extractable | LAB
BE-E | mg/L | Beryllium Extractable | F069 | SEM extr,6N HCl,ICP/OES | 6.0 N Hydrochloric acid (HCI) is added to an aliquot of the sediment sample. The extract produced from the addition of the acid is then analyzed for simultaneously extracted metals (SEM) using inductively coupled plasma - optical emission spectrophotometry. The detection limits depend on the metal parameters in question and the instrumentation used. | 0.005 | Y | Y | Y | 1 |  |  | 11/25/2021 1:50:15 PM | 1999-07-09 |  | NPEPPIN |  | Concentration | mg/L | Y | (fl. conc.) |  | Beryllium Extractable (fl. conc.) | 7440-41-7 | Metals - Extractable; | Metals - Extractable; EMS code: BE-E | Chemical | Extractable | LAB
BE-T | ug/g wet | Beryllium Total | BPSY | MET-WET-BPSY-CCMS-VA | Metals in Animal Tissue and Vegetation (Biota) - Prescriptive. EPA 200.3/6020B This method is conducted following British Columbia Lab Manual method "Metals in Animal Tissue and Vegetation (Biota) - Prescriptive". Tissue samples are digested with nitric and hydrochloric acids; in combination with addition of hydrogen peroxide. Instrumental analysis is by collision cell inductively coupled plasma - mass spectrometry (modified from EPA Method 6020B). | 0.002 | Y | Y | Y | 2 | 11/7/2019 | 11/7/2019 | 6/10/2021 1:57:10 PM | 2021-06-04 |  | EMS |  | MassPerMass | µg/g | Y | (mass conc.) | (wet) | Beryllium Total (wet) (mass conc.) | 7440-41-7 | Metals - Total; | Metals - Total; EMS code: BE-T | Chemical | Total | LAB
BE-T | mg/L | Beryllium Total | F082 | Aliquot,HNO3/HCL dig,ICPMS | Aliquot microwave digested with 3:1 HNO3/HCL. Analysis by ICPMS. For HiVol samples preliminary result are reported in mg/L and loading results are reported in ug/m3 | 0.001 | Y | Y | Y | 26977 | 7/26/2005 7:04:00 AM | 8/23/2024 | 8/26/2024 4:53:05 PM | 2000-06-13 | 8/15/2019 11:05:28 AM | NPEPPIN | KATFRASE | Concentration | mg/L | Y | (fl. conc.) |  | Beryllium Total (fl. conc.) | 7440-41-7 | Metals - Total; | Metals - Total; EMS code: BE-T | Chemical | Total | LAB
BE-T | ug/g | Beryllium Total | FS01 | D&G,Nitric/HCl:ICP | Samples are dried at 60 C, then either sieved to -100mesh (150 um) and rolled to homogenize, or finely ground with a mortar and pestle or a shatter box laboratory mill and rolled (depending on the submitter's requirements). The samples are then weighed into Teflon digestion liners and digested with HNO3 and HCl in a microwave oven. Samples are cooled, volumized, and allowed to settle overnight before being decanted. Analysis is then performed via Inductively Coupled Argon Plasma Atomic Emission Spectrometry (ICP-AES) (Simultaneous multi element analysis). In ICP Spectrometry, the liquid sample is pumped to the nebulizer, where it is convertd into an aerosol. the nebulizer carrier gas flow than transports the wet aerosol past the impact bead, where the large droplets condense out and flow to the waste container. The remaining fine aerosol is carried into the plasma. There, the high temperature desolvates the aerosol and dissociates the sample material into individual atoms and ions, which are excited to emit light at wavelengths characteristic of the elements present. A diffraction grating in the spectrometer diffracts the various wavelengths of light through exit slits to the photo-multiplier tubes, which convert the light energy to electronic signals. | 0.2 | Y | Y | Y | 2642 | 8/30/1996 1:52:00 PM | 10/10/2007 | 5/24/2008 1:15:02 AM | 1996-06-19 | 6/3/1998 6:51:44 PM | MPAK |  | MassPerMass | mg/kg | Y | (mass conc.) |  | Beryllium Total (mass conc.) | 7440-41-7 | Metals - Total; | Metals - Total; EMS code: BE-T | Chemical | Total | LAB
BE-T | mg/L | Beryllium Total | F048 | Acid dig,conc'n,nebul.,ICP-OES | This method provides an automated procedure for the Inductively Coupled Argon Plasma Optical Emission Spectrophotometric (ICP-OES) determination and quantitation of trace amounts of metals in surface, ground and precipitation waters after a manual acid digestion and concentration procedure and ultrasonic nebulisation. Organically-bound metals and metal complexes are converted to metal ions by digestion with nitric acid [method details says HCl]. To increase the signal detected by the ICP, the samples are concentrated 10 times. The liquid sample is pumped onto the piezoelectric transducer, where the ultrasonic action efficiently converts it to a fine, dense aerosol. The nebulizer carrier gas flow then transports the wet aerosol through a heated U-tube, where the solvent is vaporized. Subsequently, most of the solvent vapour is removed from the aerosol stream by passing the stream through a chilled condenser. The result is a dry, analyte-laden aerosol which is introduced to the plasma. There, the high temperature desolvates the aerosol and dissociates the sample material into individual atoms and ions, which are excited to emit light at wave lengths characteristic of the elements present. Dissolved samples are field filtered through 0.45 micron cellulose acetate filter prior to acid preservation. Samples for extractable analysis are acidified with 2 mL concentrated HNO3 per litre and allowed to stand a minimum of 24 hours before analysis. Clients providing these samples are advised to preserve samples at the collection site. | 5e-05 | Y | Y | Y | 6159 | 7/4/1997 2:35:00 PM | 4/23/2003 | 9/11/2019 8:46:38 AM | 1997-10-20 | 6/3/1998 6:51:47 PM | NPEPPIN |  | Concentration | mg/L | Y | (fl. conc.) |  | Beryllium Total (fl. conc.) | 7440-41-7 | Metals - Total; | Metals - Total; EMS code: BE-T | Chemical | Total | LAB
BE-T | ug/g | Beryllium Total | FS82 | PBM,dry60C,10mesh,SALM-ICP-MS | PBM;dry60C;10mesh;SALM-ICP-MS | 0.1 | Y | Y | Y | 3648 | 5/21/2002 | 7/24/2024 | 7/29/2024 5:17:09 PM | 2002-11-12 |  | UMULLER |  | MassPerMass | mg/kg | Y | (mass conc.) |  | Beryllium Total (mass conc.) | 7440-41-7 | Metals - Total; | Metals - Total; EMS code: BE-T | Chemical | Total | LAB
BE-T | ug/g | Beryllium Total | 0192 | Hom, Nitric/Perchlor,ICP | Sample already in a dry state so does not require drying before homogenization and nitric perchlorid digestion followed by ICP analysis. | 1 | Y | Y | Y | 236 |  |  | 10/24/1996 11:23:41 AM | 1996-04-12 | 6/3/1998 6:51:41 PM | EMS |  | MassPerMass | mg/kg | Y | (mass conc.) |  | Beryllium Total (mass conc.) | 7440-41-7 | Metals - Total; | Metals - Total; EMS code: BE-T | Chemical | Total | LAB
BE-T | mg/L | Beryllium Total | 1349 | Tot Metals in Seawater by CRC ICPMS (BC) | Tot Metals in Seawater by CRC ICPMS (BC) | 0.0005 | Y | Y | Y | 98 | 8/16/2019 | 12/5/2021 | 12/13/2021 12:34:16 PM | 2019-09-30 |  | KATFRASE |  | Concentration | mg/L | Y | (fl. conc.) |  | Beryllium Total (fl. conc.) | 7440-41-7 | Metals - Total; | Metals - Total; EMS code: BE-T | Chemical | Total | LAB
BE-T | ug/g | Beryllium Total | MTID | MET-DRY-CCMS-VA | Metals in Tissue by CRC ICPMS (DRY) Methodology: This method is conducted following British Columbia Lab Manual method "Metals in Animal Tissue and Vegetation (Biota) - Prescriptive". Tissue samples are homogenized and sub-sampled prior to hotblock digestion with nitric and hydrochloric acids, in combination with addition of hydrogen peroxide. Instrumental analysis is by collision cell inductively coupled plasma - mass spectrometry (modified from EPA Method 6020A). Method Limitation: This method employs a strong acid/peroxide digestion, and is intended to provide a conservative estimate of bio-available metals. Near complete recoveries are achieved for most toxicologically important metals, but elements associated with recalcitrant minerals may be only partially recovered. | 0.01 | Y | Y | Y | 122 | 8/13/2014 | 9/8/2022 | 10/26/2022 7:14:14 AM | 2015-03-09 | 3/11/2015 4:12:05 PM | RDGIBSON | EMS | MassPerMass | mg/kg | Y | (mass conc.) |  | Beryllium Total (mass conc.) | 7440-41-7 | Metals - Total; | Metals - Total; EMS code: BE-T | Chemical | Total | LAB
BE-T | ug/g | Beryllium Total | FB04 | Freeze Dry,Hom,Acid Dig,ICP | Sample moisture is determined by weighing samples before and after freeze-drying. Freeze-dried samples are ground in a Braun blender before being weighed into Teflon PFA liners and digested with concentrated HNO3, concentrated HCl, and 50% H2O2 in a microwave oven; which results in a complete sample dissolution. Samples are collected and volumized with Type I water. Analysis is performed via Inductively Coupled Argon Plasma Atomic Emission Spectrometry (ICP-AES) (simultaneous multielement analysis), Graphite Furnace Atomic Absorption Spectrometry (GF-AAS) (low level Cd and Pb analysis), Cold Vapor Atomic Fluorescence Spectrometry (CV-AFS) (Mercury analysis) and/or Flame Atomic Absorption Spectrometry (F-AAS) (K analysis). | 0.08 | Y | Y | Y | 241 | 8/22/1996 6:11:00 PM | 2/1/2002 2:30:00 PM | 6/23/2003 1:15:28 PM | 1996-06-19 | 6/3/1998 6:51:44 PM | MPAK |  | MassPerMass | mg/kg | Y | (mass conc.) |  | Beryllium Total (mass conc.) | 7440-41-7 | Metals - Total; | Metals - Total; EMS code: BE-T | Chemical | Total | LAB
BE-T | mg/L | Beryllium Total | TCLP | Toxicity Characteristic Leaching Procedure | The TCLP procedure is designed to simulate the climatic leaching action | 0.1 | Y | Y | Y | 22 | 12/8/2009 7:49:00 PM | 9/28/2011 5:53:00 PM | 11/3/2011 11:10:18 PM | 2009-03-05 |  | JCHUDAK |  | Concentration | mg/L | Y | (fl. conc.) |  | Beryllium Total (fl. conc.) | 7440-41-7 | Metals - Total; | Metals - Total; EMS code: BE-T | Chemical | Total | LAB
BE-T | mg/L | Beryllium Total | F038 | HNO3/HCl dig,ICP-AES | Dissolved and extractable water samples are ready for analysis once collected. Total samples are autoclaved with HNO3 and hydrochloric acid (HCI) in capped Teflon vials. Analysis is performed via Inductively Coupled Argon Plasma Atomic Emission Spectrometry (ICP-AES)(simultaneous multielement analysis), Graphite Furnace Atomic Absorption Spectrometry (GF-AAS) (low level analysis) and/or Flame Atomic Absorption Spectrometry (F-AAS). | 0.001 | Y | Y | Y | 5400 | 8/1/1996 6:04:00 PM | 12/1/2010 | 9/11/2019 8:47:42 AM | 1997-05-08 | 6/3/1998 6:51:46 PM | MPAK |  | Concentration | mg/L | Y | (fl. conc.) |  | Beryllium Total (fl. conc.) | 7440-41-7 | Metals - Total; | Metals - Total; EMS code: BE-T | Chemical | Total | LAB
BE-T | mg/L | Beryllium Total | X555 | Metals in water, total, "in-bottle" digestion, ICP-SFMS | The whole water sample collected in a HDPE/Teflon bottle, is de-capped in a Class 100 workstation, acidified to 2% (v/v) HNO3 and 1% (v/v) HCl and recapped. The sample is "in-bottle" digested in a convection oven, set at 60 degrees C for 16 hours. The sample is cooled and introduced to an inductively coupled argon plasma-sector field mass spectrometer (ICP-SFMS) by pneumatic nebulization. Each respective element is measured at a specific mass to charge ratio, m/z value expressed in amu (atomic mass units), and at resolution. LR: Low Resolution ~300, MR: Mid Resolution ~4000 and HR: High a specified instrument Resolution ~10000. Ref: Method number 02-2003, National Laboratory for Environmental Testing, CCIW, Environment Canada, Burlington. Requested April 2003. | 1e-06 | Y | Y | Y | 17334 | 5/20/2003 | 12/6/2023 2:08:00 AM | 2/14/2024 9:05:32 AM | 2004-01-14 |  | UMULLER |  | Concentration | mg/L | Y | (fl. conc.) |  | Beryllium Total (fl. conc.) | 7440-41-7 | Metals - Total; | Metals - Total; EMS code: BE-T | Chemical | Total | LAB
BE-T | mg/L | Beryllium Total | X285 | HNO3 Dig:ICP Analysis | HNO3 Dig:ICP Analysis | 0.001 | Y | Y | Y | 595 | 8/10/1993 | 10/5/1993 | 4/30/2004 1:54:38 PM | 1998-05-22 | 6/3/1998 6:51:49 PM | NPEPPIN |  | Concentration | mg/L | Y | (fl. conc.) |  | Beryllium Total (fl. conc.) | 7440-41-7 | Metals - Total; | Metals - Total; EMS code: BE-T | Chemical | Total | LAB
BE-T | ug/g wet | Beryllium Total | X552 | Metals in Biota (wet), total, acid digestion, ICP-SFMS | One gram of wet, homogenized biota is weighed in a microwave TFM vessel, and 8mL of Nitric Acid and 1mL of Hydrogen Peroxide are added. The sample is digested at 200 degrees C for 15 minutes in a closed vessel high pressure microwave oven. The sample is cooled, the volume adjusted to 100mL and introduced to an inductively coupled argon plasma-sector field mass spectrometer (ICP-SFMS) by pneumatic nebulization. Each respective element is measured at a specific mass to charge ratio, m/z value expressed in amu (atomic mass units), and at a specified instrument resolution. LR: Low Resolution ~300, MR: Mid Resolution ~4000. Ref: 02-2704, National Laboratory for Environmental Testing, CCIW, Environment Canada, Burlington. Requested April 2003. | 0.0001 | Y | Y | Y | 37 | 8/9/2016 | 10/16/2023 | 11/28/2023 2:21:32 PM | 2004-01-19 | 4/19/2021 1:22:00 PM | UMULLER | EMS | MassPerMass | µg/g | Y | (mass conc.) | (wet) | Beryllium Total (wet) (mass conc.) | 7440-41-7 | Metals - Total; | Metals - Total; EMS code: BE-T | Chemical | Total | LAB
BE-T | ug/m3 | Beryllium Total | 5312 | Tef:N/P Dig:ICP Analysis | A nitric/perchloric acid digestion if used to bring the metals into solution prior to ICP analysis. For Hi-Vol samples preliminary results are reported as mg/L and 24 hour loading results as ug/m3. | 0.0001 | Y | Y | Y | 862 | 11/7/1997 | 3/29/2000 | 9/28/2000 1:34:40 PM | 1996-10-13 | 6/3/1998 6:51:46 PM | EMS |  | AirConcentration | µg/m³ | Y | (air conc.) |  | Beryllium Total (air conc.) | 7440-41-7 | Metals - Total; | Metals - Total; EMS code: BE-T | Chemical | Total | LAB
BE-T | mg/L | Beryllium Total | MSEA | Metals in Seawater (undigested) by (CRC) ICP-MS | This method describes the multi-elemental determination of trace elements by Collision/Reaction Cell (CRC) ICP-MS. The method measures ions produced by a radio-frequency inductively coupled plasma. Analyte species originating in a liquid are pneumatically nebulized and the resulting aerosol transported by argon gas into the plasma torch. The ions produced are entrained in the plasma gas and extracted, by means of a differentially pumped vacuum interface, into a mass spectrometer. The ions flow through a Collision/Reaction Cell (CRC) where depending on the mode of analysis either the ions passes through the cell unhindered (no gas - classical ICP-MS mode), or collisions/reactions chemistry takes place between the ions and a neutral gas such as hydrogen or helium. Any polyatomic ions which give interference on the analytes are removed. The ions are then sorted according to their mass-to-charge ratios by a quadrupole mass spectrometer. The ions are then sorted according to their mass-to-charge ratios by a quadrupole mass spectrometer. The ions transmitted through the quadrupole are quantified by a channel electron multiplier and the ion information processed by a data handling system. Although most polyatomic interference is removed, any additional interference must be assessed and valid corrections applied or the data flagged to indicate problems. Interference correction must include compensation for background ions contributed by the plasma gas, reagents, and constituents of the sample matrix. Instrument drift as well as suppressions or enhancements of instrument response caused by the sample matrix must be corrected for by the use of internal standards. | 3e-05 | Y | Y | Y | 441 | 8/12/2009 12:08:00 AM | 4/29/2024 | 6/7/2024 10:03:16 AM | 2009-09-08 |  | RDGIBSON |  | Concentration | mg/L | Y | (fl. conc.) |  | Beryllium Total (fl. conc.) | 7440-41-7 | Metals - Total; | Metals - Total; EMS code: BE-T | Chemical | Total | LAB
BE-T | ppm (S) | Beryllium Total | MS61 | ALS Minerals ME-MS61 | ALS Minerals ME-MS61. A prepared sample is digested with HNO3; HF; and HClO4 in a graphite heating block to near dryness. After cooling; the sample is boiled in diluted HCl. Bulk to volume with deionized water; mixed and analyzed by ICP-AES and ICP-MS for all elements. | 0.05 | Y | Y | Y | 0 |  |  |  | 2024-01-29 |  | EMS |  | DimensionlessRatio | ppm | N | (ratio) |  | Beryllium Total (ratio) | 7440-41-7 | Metals - Total; | Metals - Total; EMS code: BE-T | Chemical | Total | LAB
BE-T | mg/L | Beryllium Total | F121 | Acid,ICPMS | Acidified, no digestion, ICPMS | 1e-05 | Y | Y | Y | 8412 | 3/12/2010 7:57:00 PM | 8/15/2024 9:10:00 PM | 8/26/2024 8:06:03 AM | 2002-07-17 | 7/16/2012 3:27:46 PM | UMULLER | RDGIBSON | Concentration | mg/L | Y | (fl. conc.) |  | Beryllium Total (fl. conc.) | 7440-41-7 | Metals - Total; | Metals - Total; EMS code: BE-T | Chemical | Total | LAB
BE-T | ug/g | Beryllium Total | EQ02 | EQUIS (ug/g) | EQUIS (ug/g) | 1 | N | Y | Y | 743 |  |  | 10/23/1996 12:08:28 AM | 1996-04-12 | 6/3/1998 6:51:34 PM | EMS |  | MassPerMass | mg/kg | Y | (mass conc.) |  | Beryllium Total (mass conc.) | 7440-41-7 | Metals - Total; | Metals - Total; EMS code: BE-T | Chemical | Total | LAB
BE-T | ug/g | Beryllium Total | PB02 | PBM,HNO3/HCl dig,dry60C,100mesh, | Performance Based Method. Samples are dried at 60 C and sieved using a .149 mm (100 mesh) sieve. Solid samples are digested with the combination of nitric and hydrochloric acids with a 1:1 ratio of acids. For environmental purposes, strong acid leaches are used routinely since they have the advantage of speed, low cost, and superio reproducibility. All metals trapped within the refractory matrix are unlikely to have environmental consequences. Instrumental analysis of extracts produced using this digestion method can be performed by a variety of techniques. |  | Y | Y | Y | 52 | 5/24/2001 3:42:00 PM | 11/26/2001 10:23:00 AM | 5/26/2003 2:27:26 PM | 2001-07-18 |  | NPEPPIN |  | MassPerMass | mg/kg | Y | (mass conc.) |  | Beryllium Total (mass conc.) | 7440-41-7 | Metals - Total; | Metals - Total; EMS code: BE-T | Chemical | Total | LAB
BE-T | mg/L | Beryllium Total | 0040 | HNO3 Dig: ICP Analysis | HNO3 Dig: ICP Analysis | 0.001 | Y | Y | Y | 4841 | 5/8/2003 9:40:00 AM | 6/25/2010 3:15:00 PM | 7/22/2010 2:11:39 PM | 1996-04-12 | 6/3/1998 6:51:41 PM | EMS |  | Concentration | mg/L | Y | (fl. conc.) |  | Beryllium Total (fl. conc.) | 7440-41-7 | Metals - Total; | Metals - Total; EMS code: BE-T | Chemical | Total | LAB
BE-T | mg/L | Beryllium Total | 0042 | GI HNO3 Dig: ICP | Sample aliquot taken from unfiltered unpreserved General Ions bottle; HNO3 digestion; ICP analysis. | 0.001 | Y | Y | Y | 6259 | 9/20/1996 | 7/5/2019 3:21:00 PM | 7/12/2019 6:49:13 PM | 1996-04-12 | 6/3/1998 6:51:34 PM | EMS |  | Concentration | mg/L | Y | (fl. conc.) |  | Beryllium Total (fl. conc.) | 7440-41-7 | Metals - Total; | Metals - Total; EMS code: BE-T | Chemical | Total | LAB
BE-T | mg/L | Beryllium Total | METN | MET-T-NP-CCMS-VA | Ultra trace metals in water are analyzed by CRC ICPMS, based on US EPA Method 6020B (July 2014). The detection limits provided can only be met for undigested samples. This procedure is intended for colorless, non-turbid, acid-preserved water samples (i.e. pristine water samples), having turbidity < 1 NTU and no odor. Where turbidity exceeds 1 NTU, and/or the sample is colored and has an odor, results may be biased low compared to true Total Metals concentrations. ALS recommends that turbidity analysis be requested on samples submitted for this test to aid with interpretation of results. Where turbidity is <1NTU, undigested metals are equivalent to total metals concentrations. | 1e-05 | Y | Y | Y | 3648 | 4/3/2012 | 8/19/2024 | 8/26/2024 1:03:06 PM | 2019-08-20 |  | KATFRASE |  | Concentration | mg/L | Y | (fl. conc.) |  | Beryllium Total (fl. conc.) | 7440-41-7 | Metals - Total; | Metals - Total; EMS code: BE-T | Chemical | Total | LAB
BE-T | ppm (S) | Beryllium Total | 6041 | Aqua Regia by ICPMS | These were sediment/soils digested with Aqua-Regia (Nitric Acid & Hydrochloric Acids) and then analysis by ICPMS. Results would be reported in mg/kg and/or %. | 0.01 | Y | Y | Y | 21 | 10/13/2019 | 8/5/2020 | 8/10/2020 3:52:06 PM | 2019-09-06 | 10/21/2019 11:54:21 AM | KATFRASE | KATFRASE | DimensionlessRatio | ppm | N | (ratio) |  | Beryllium Total (ratio) | 7440-41-7 | Metals - Total; | Metals - Total; EMS code: BE-T | Chemical | Total | LAB
BE-T | ug/g | Beryllium Total | 0190 | D&G:N/P Dig:ICP Anal. | Sample is dried and grind-mixed before nitric perchloric digestion followed by ICP analysis. | 1 | Y | Y | Y | 13 |  |  | 10/23/1996 10:17:01 PM | 1996-04-12 | 6/3/1998 6:51:34 PM | EMS |  | MassPerMass | mg/kg | Y | (mass conc.) |  | Beryllium Total (mass conc.) | 7440-41-7 | Metals - Total; | Metals - Total; EMS code: BE-T | Chemical | Total | LAB
BE-T | ug/g | Beryllium Total | ICPM | Metals in Soil by CRC ICPMS | This analysis is carried out using procedures from CSR Analytical Method: "Strong Acid Leachable Metals (SALM) in Soil", BC Ministry of Environment, 26 June 2009, and procedures adapted from EPA Method 200.2. The sample is manually homogenized, dried at 60 degrees Celsius, sieved through a 63 um (230 mesh) sieve, and a representative subsample of the dry material is weighed. The sample is then digested at 95 degrees Celsius for 2 hours by block digester using concentrated nitric and hydrochloric acids. Instrumental analysis is by collision cell inductively coupled plasma - mass spectrometry (modifed from EPA Method 6020A). Method Limitation: This method is not a total digestion technique. It is a very strong acid digestion that is intended to dissolve those metals that may be environmentally available. By design, elements bound in silicate structures are not normally dissolved by this procedure as they are not usually mobile in the environment. | 0.1 | Y | Y | Y | 103 | 9/19/2017 | 9/24/2023 | 1/12/2024 11:16:52 AM | 2018-01-09 |  | JKROGH |  | MassPerMass | mg/kg | Y | (mass conc.) |  | Beryllium Total (mass conc.) | 7440-41-7 | Metals - Total; | Metals - Total; EMS code: BE-T | Chemical | Total | LAB
BE-T | ug/m3 | Beryllium Total | F088 | Aliquot,HNO3/HCl dig, ICPMS,loading | Aliquot microwave digested with 3:1 HNO3/HCL. Analysis by ICPMS. For HiVol samples preliminary result are reported in mg/L and loading results are reported in ug/m3 |  | Y | Y | Y | 565 | 8/17/2000 10:46:00 AM | 12/14/2006 | 12/30/2006 1:01:36 AM | 2000-08-16 |  | NPEPPIN |  | AirConcentration | µg/m³ | Y | (air conc.) |  | Beryllium Total (air conc.) | 7440-41-7 | Metals - Total; | Metals - Total; EMS code: BE-T | Chemical | Total | LAB
BE-T | ppm (S) | Beryllium Total | 6042 | Aqua Regia by ICPMS | These were sediment/soils digested with Aqua-Regia (Nitric Acid & Hydrochloric Acids) and then analysis by ICPMS. Results would be reported in ppm s. | 0.01 | Y | Y | Y | 71 | 5/14/2022 | 7/27/2024 | 7/29/2024 9:13:08 AM | 2022-06-28 |  | EMS |  | DimensionlessRatio | ppm | N | (ratio) |  | Beryllium Total (ratio) | 7440-41-7 | Metals - Total; | Metals - Total; EMS code: BE-T | Chemical | Total | LAB
BE-T | ppm (S) | Beryllium Total | MS41 | ALS Minerals ME-MS41 | ALS Minerals ME-MS41. A prepared sample is digested with aqua regia in a graphite heating block. After cooling; the resulting solution is diluted to with deionized water; mixed and analyzed by ICP-AES. Samples are then analyzed by ICP-MS for the remaining suite of elements. | 0.05 | Y | Y | Y | 0 |  |  |  | 2024-01-29 |  | EMS |  | DimensionlessRatio | ppm | N | (ratio) |  | Beryllium Total (ratio) | 7440-41-7 | Metals - Total; | Metals - Total; EMS code: BE-T | Chemical | Total | LAB
BE-T | mg/L | Beryllium Total | MWID | Metals in Water by ICPMS (digested) | Metals in Water by ICPMS (digested) Based on EPA 200.8 | 0.0002 | Y | Y | Y | 36804 |  |  | 7/31/2024 1:46:41 PM | 2009-12-24 |  | RDGIBSON |  | Concentration | mg/L | Y | (fl. conc.) |  | Beryllium Total (fl. conc.) | 7440-41-7 | Metals - Total; | Metals - Total; EMS code: BE-T | Chemical | Total | LAB
BE-T | ug/g wet | Beryllium Total | MTIW | MET-WET-CCMS-MID-VA | Metals in Tissue by CRC ICPMS (WET) Methodology: This method is conducted following British Columbia Lab Manual method "Metals in Animal Tissue and Vegetation (Biota) - Prescriptive". Tissue samples are homogenized and sub-sampled prior to hotblock digestion with nitric and hydrochloric acids, in combination with addition of hydrogen peroxide. Instrumental analysis is by collision cell inductively coupled plasma - mass spectrometry (modified from EPA Method 6020A). Method Limitation: This method employs a strong acid/peroxide digestion, and is intended to provide a conservative estimate of bio-available metals. Near complete recoveries are achieved for most toxicologically important metals, but elements associated with recalcitrant minerals may be only partially recovered. | 0.002 | Y | Y | Y | 24 | 11/25/2018 | 12/6/2018 | 7/21/2021 2:55:07 PM | 2018-12-24 |  | EMS |  | MassPerMass | µg/g | Y | (mass conc.) | (wet) | Beryllium Total (wet) (mass conc.) | 7440-41-7 | Metals - Total; | Metals - Total; EMS code: BE-T | Chemical | Total | LAB
BE-T | ug/m3 | Beryllium Total | 6040 | Aqua Regia Dig:ICP | Aqua regia digestion is used to bring the metals into solution prior to ICP analysis. For Hi-Vol samples, preliminary results are reported as mg/L and24 hr loadings as ug/m3 | 0.0001 | Y | Y | Y | 437 | 2/13/2002 | 2/9/2005 | 2/23/2005 1:13:47 AM | 1998-05-05 | 6/3/1998 6:51:48 PM | NPEPPIN |  | AirConcentration | µg/m³ | Y | (air conc.) |  | Beryllium Total (air conc.) | 7440-41-7 | Metals - Total; | Metals - Total; EMS code: BE-T | Chemical | Total | LAB
BE-T | mg/L | Beryllium Total | X349 | FA:HNO3:ICP | Samples are field-acidified with HNO3; HNO3 digestion; analysis by ICP/AES. | 0.001 | Y | Y | Y | 30 |  |  | 10/30/2014 8:31:17 AM | 1998-05-22 | 6/3/1998 6:51:49 PM | NPEPPIN |  | Concentration | mg/L | Y | (fl. conc.) |  | Beryllium Total (fl. conc.) | 7440-41-7 | Metals - Total; | Metals - Total; EMS code: BE-T | Chemical | Total | LAB
BE-T | mg/L | Beryllium Total | F033 | Acid Dig,Centrifuge,ICPMS | For the determination of total recoverable elements in water or waste water, take 100 ml aliquot from a well mixed, acid preserved, sample containing not more than 0.25% (w/v) total solids and transfer to a 250 ml Griffin beaker. (If total solids are greater than 0.25% reduce the size of the aliquot by a proportionate amount). Add 1 ml of conc. nitric acid and 0.5 ml conc. hydrochloric acid. Heat on a hot plate until the volume has been reduced to 15-20ml, ensuring that the sample does not boil. Allow to cool and transfer to a 50 ml volumetric flask. Dilute to volume with ASTM type I water. Centrifuge the sample or allow to stand overnight to separate insoluble material. Prior analysis, pipette 20 ml into a 50 ml volumetric flask and dilute to volume with ASTM type I water. Analysis should be performed within seventy two hours of the completed preparation. Note that different laboratories may have very different field preservation, digestion procedures and detection limits for this method, so these should be documented. | 0.001 | Y | Y | Y | 37153 | 11/20/1997 9:31:00 AM | 8/22/2024 3:24:00 AM | 8/25/2024 9:47:36 PM | 1998-05-14 | 6/3/1998 6:51:49 PM | NPEPPIN |  | Concentration | mg/L | Y | (fl. conc.) |  | Beryllium Total (fl. conc.) | 7440-41-7 | Metals - Total; | Metals - Total; EMS code: BE-T | Chemical | Total | LAB
BE-T | ug/g | Beryllium Total | X230 | Hom,HNO3/HCl,ICP | The sample is dried, ground to pass 100 mesh, and digested with a nitric-hydrochloric acid mixture to solubilize the solid matter and to remove the organic material by oxidation and volatilization. The completed digests are submitted to metals instrumentation personnel for metals analysis by Inductively Coupled Atomic Emission Spectrophotometry (ICP-AES), or hydride analysis by atomic absorption. | 1 | Y | Y | Y | 41 | 2/12/2002 | 12/30/2005 | 2/3/2006 1:13:12 AM | 1998-11-05 |  | NPEPPIN |  | MassPerMass | mg/kg | Y | (mass conc.) |  | Beryllium Total (mass conc.) | 7440-41-7 | Metals - Total; | Metals - Total; EMS code: BE-T | Chemical | Total | LAB
BE-T | mg/L | Beryllium Total | F068 | HNO3dig,TeflonTube/Autoclave,ICP-MS | This method is applicable to determiniation of dissolved, extractable, and total recoverable metals in ground waters, surface waters, drinking waters, and wastewaters. Dissolved, extractable, and total recoverable sample fractions are acidified to pH <1.5 with analytical grade nitric acid (HNO3) to prevent the metals from adsorbing to the container material or precipitating. Extractable and total recoverable water samples are collected in 250 ml acid washed polyethylene bottles and are preserved with Nitric acid (0.5 ml acid for 100 mLs of sample). All preserved samples are stored at room temperature and can be held for up to 6 months. Unfiltered, un-preserved samples are stored at 4C for no more than 30 days. For dissolved and extractable water fractions, no sample preparation is necessary. Samples are simply poured in autosampler tubes (i.e. 15 ml centrifuge tubes), and placed on the autosampler and analyzed. Samples for total recoverable metals are digested with Nitric acid in special Teflon tubes in an autoclave prior to analyses. | 2e-06 | Y | Y | Y | 1779 | 6/1/1999 1:24:00 PM | 7/4/2008 9:48:00 AM | 7/30/2024 6:32:12 AM | 1999-06-28 |  | NPEPPIN |  | Concentration | mg/L | Y | (fl. conc.) |  | Beryllium Total (fl. conc.) | 7440-41-7 | Metals - Total; | Metals - Total; EMS code: BE-T | Chemical | Total | LAB
BE-T | ug/g wet | Beryllium Total | X560 | Metals in Biota (wet), total, acid digestion, ICP-MS | Five grams of wet, homogenized biota is weighed in a centrifuge tube, and 2.5 mL of Nitric Acid and 2.5 mL of Hydrochloric acid are added. The sample is digested at room temperature for one hour and then digested at 90-95 degrees C in a water bath or heat block for one more hour. The sample is cooled and the volume is made to 50 mL with DI. The digestate is then diluted at ten fold and introduced to a collision/reaction cell inductively coupled argon plasma mass spectrometer (CRC ICPMS) by pneumatic nebulization. Each respective element is measured at a specific mass to charge ratio, m/z value expressed in amu (atomic mass units). - Maxxam Analytics (BC). Added to EMS Dec 10, 2008. | 0.1 | Y | Y | Y | 541 | 11/15/2008 4:00:00 AM | 10/30/2023 | 11/28/2023 2:21:33 PM | 2008-12-11 | 3/11/2015 4:17:10 PM | RDGIBSON | EMS | MassPerMass | µg/g | Y | (mass conc.) | (wet) | Beryllium Total (wet) (mass conc.) | 7440-41-7 | Metals - Total; | Metals - Total; EMS code: BE-T | Chemical | Total | LAB
BE-T | mg/L | Beryllium Total | X170 | ICP | This code is a catch-all for a large variety of ICP methods including water and also solids digested into solutions as an intermediate result. Different procedures will have greatly different detection limits. (a) For air particulates, preliminary results are reported in mg/L and loading results as mg/dm2/d. (b) For Hi Vol samples preliminiary results are same as reported in mg/L and loading results as ug/m3. (c) Note that for Rain samples there is no filtration nor digestion steps. | 0.001 | Y | Y | Y | 3951 | 3/28/1996 6:03:00 PM | 11/7/2018 | 10/24/2022 10:58:25 AM | 1996-04-12 | 6/3/1998 6:51:34 PM | EMS |  | Concentration | mg/L | Y | (fl. conc.) |  | Beryllium Total (fl. conc.) | 7440-41-7 | Metals - Total; | Metals - Total; EMS code: BE-T | Chemical | Total | LAB
BE-T | mg/dm2/d | Beryllium Total | DUST | HNO3 dig.(microwave or hot block),ICPMS | HNO3 dig.(microwave or hot block);ICPMS |  | Y | Y | Y | 420 | 7/19/2016 1:11:00 PM | 7/31/2024 | 7/31/2024 11:23:12 AM | 2008-02-07 |  | RDGIBSON |  | MassFlux | mg/dm²/d | Y | (mass flux) |  | Beryllium Total (mass flux) | 7440-41-7 | Metals - Total; | Metals - Total; EMS code: BE-T | Chemical | Total | LAB
BE-T | ug/g | Beryllium Total | X551 | Metals in Biota (dry), total, acid digestion, ICP-SFMS | 0.1 gram of dried, homogenized biota is weighed in a microwave TFM vessel, and 8mL of Nitric Acid and 1mL of Hydrogen Peroxide are added. The sample is digested at 200 degrees C for 15 minutes in a closed vessel high pressure microwave oven. The sample is cooled, the volume adjusted to 100mL and introduced to an iductively coupled argon plasma-sector field mass spectrometer (ICP-SFMS) by pneumatic nebulization. Each respective element is measured at a specific mass to charge ratio, m/z value expressed in amu (atomic mass units), and at a specified instrument resolution. LR: Low Resolution ~300, MR: Mid Resolution ~4000 and HR: Hich Resolution ~10000. Ref: Method number 02-2705, National Laboratory for Environmental Testing, CCIW, Environment Canada, Burlington. Requested April 2003. | 0.001 | Y | Y | Y | 49 | 8/11/2015 | 10/30/2023 | 11/28/2023 2:21:33 PM | 2004-01-19 |  | UMULLER |  | MassPerMass | mg/kg | Y | (mass conc.) |  | Beryllium Total (mass conc.) | 7440-41-7 | Metals - Total; | Metals - Total; EMS code: BE-T | Chemical | Total | LAB
BE-T | mg/L | Beryllium Total | X561 | Total metals: FA, HNO3, HR-ICPMS (No Prep)Prep) | Ultra trace metals in water are analyzed by high resolution inductively coupled plasma mass spectrometry (HR-ICPMS), modified from US EPA 200.8, (Revision 5.5). The detection limits provided can only be met for undigested samples. This procedure is intended for colorless, non-turbid, acid-preserved water samples (i.e. pristine water samples), having turbidity < 1 NTU and no odor. Where turbidity exceeds 1 NTU, and/or the sample is colored and has an odor, results may be biased low compared to true Total Metals concentrations. ALS recommends that turbidity analysis be requested on samples submitted for this test to aid with interpretation of results. | 1e-05 | Y | Y | Y | 2756 | 6/8/2015 | 10/5/2023 | 1/31/2024 2:37:41 PM | 2015-10-22 |  | RDGIBSON |  | Concentration | mg/L | Y | (fl. conc.) |  | Beryllium Total (fl. conc.) | 7440-41-7 | Metals - Total; | Metals - Total; EMS code: BE-T | Chemical | Total | LAB
BI-D | mg/L | Bismuth Dissolved | METC | MET-D-NP-CCMS-VA | Ultra trace metals in water are analyzed by CRC ICPMS, based on US EPA Method 6020B (July 2014). This procedure is intended for pristine field-filtered acid-preserved water samples. The detection limits (LOR) for this test are based on lab instrumental analysis only, not including filtration. ALS-supplied field filtration equipment does not support these LOR. Therefore, because of the high probability of false positives due to filtration, it is strongly recommended that a filtration blank be analysed to aid in data interpretation. | 5e-06 | Y | Y | Y | 1776 | 4/3/2012 | 8/15/2024 | 8/23/2024 9:50:35 AM | 2019-08-20 |  | KATFRASE |  | Concentration | mg/L | Y | (fl. conc.) |  | Bismuth Dissolved (fl. conc.) | 7440-69-9 | Metals - Dissolved; | Metals - Dissolved; EMS code: BI-D | Chemical | Dissolved | LAB
BI-D | mg/L | Bismuth Dissolved | X556 | Metals in water, dissolved, "in-bottle" digestion, ICP-SFMS | The whole water sample is filtered through a 0.45 micron cellulose acetate filter paper and the filtrate is stored in a HDPE/Teflon bottle. The container is de-capped in a Class 100 workstation, acidified to 2% (v/v) HNO3 and 1% (v/v) HCl and recapped. The sample is "in-bottle" digested in a convection oven, set at 60 degrees C for 16 hours. The sample is cooled and introduced to an inductively coupled argon plasma-sector field mass spectrometer (ICP-SFMS) by pneumatic nebulization. Each respective element is measured at a specific mass to charge ratio, m/z value expressed in amu (atomic mass units), and at a specified intrument resolution. LR: Low Resolution ~300, MR: Mid Resolution ~4000 and HR: High Resolution ~10000. Ref: Method number 02-2003, National Laboratory for Environmental Testing, CCIW, Environment Canada, Burlington, Requested April 2003. | 1e-05 | Y | Y | Y | 4233 | 6/11/2003 | 12/6/2023 2:04:00 AM | 2/14/2024 9:05:32 AM | 2004-01-19 |  | UMULLER |  | Concentration | mg/L | Y | (fl. conc.) |  | Bismuth Dissolved (fl. conc.) | 7440-69-9 | Metals - Dissolved; | Metals - Dissolved; EMS code: BI-D | Chemical | Dissolved | LAB
BI-D | mg/L | Bismuth Dissolved | X562 | Dissolved metals: FF, FA, HNO3, HR-ICPMS | Ultra trace metals in water are analyzed by high resolution inductively coupled plasma mass spectrometry (HR-ICPMS), based on US EPA Method 200.8, (Oct. 1999). This procedure is intended for pristine field-filtered acid-preserved water samples. The detection limits (LOR) for this test are based on lab instrumental analysis only, not including filtration. ALS-supplied field filtration equipment does not support these LOR. Therefore, because of the high probability of false positives due to filtration, it is strongly recommended that a filtration blank be analyzed to aid in data interpretation. | 5e-06 | Y | Y | Y | 1264 | 6/8/2015 | 10/19/2023 | 11/9/2023 7:41:27 AM | 2015-10-22 |  | RDGIBSON |  | Concentration | mg/L | Y | (fl. conc.) |  | Bismuth Dissolved (fl. conc.) | 7440-69-9 | Metals - Dissolved; | Metals - Dissolved; EMS code: BI-D | Chemical | Dissolved | LAB
BI-D | mg/L | Bismuth Dissolved | X285 | HNO3 Dig:ICP Analysis | HNO3 Dig:ICP Analysis | 0.01 | Y | Y | Y | 564 | 8/21/2002 | 11/13/2002 | 2/19/2004 9:57:19 AM | 1998-05-22 | 6/3/1998 6:51:49 PM | NPEPPIN |  | Concentration | mg/L | Y | (fl. conc.) |  | Bismuth Dissolved (fl. conc.) | 7440-69-9 | Metals - Dissolved; | Metals - Dissolved; EMS code: BI-D | Chemical | Dissolved | LAB
BI-D | mg/L | Bismuth Dissolved | MWID | Metals in Water by ICPMS (digested) | Metals in Water by ICPMS (digested) Based on EPA 200.8 | 0.001 | Y | Y | Y | 31673 | 10/26/2018 5:20:00 PM | 10/26/2018 5:20:00 PM | 7/31/2024 1:46:41 PM | 2009-12-24 |  | RDGIBSON |  | Concentration | mg/L | Y | (fl. conc.) |  | Bismuth Dissolved (fl. conc.) | 7440-69-9 | Metals - Dissolved; | Metals - Dissolved; EMS code: BI-D | Chemical | Dissolved | LAB
BI-D | mg/L | Bismuth Dissolved | MSEA | Metals in Seawater (undigested) by (CRC) ICP-MS | This method describes the multi-elemental determination of trace elements by Collision/Reaction Cell (CRC) ICP-MS. The method measures ions produced by a radio-frequency inductively coupled plasma. Analyte species originating in a liquid are pneumatically nebulized and the resulting aerosol transported by argon gas into the plasma torch. The ions produced are entrained in the plasma gas and extracted, by means of a differentially pumped vacuum interface, into a mass spectrometer. The ions flow through a Collision/Reaction Cell (CRC) where depending on the mode of analysis either the ions passes through the cell unhindered (no gas - classical ICP-MS mode), or collisions/reactions chemistry takes place between the ions and a neutral gas such as hydrogen or helium. Any polyatomic ions which give interference on the analytes are removed. The ions are then sorted according to their mass-to-charge ratios by a quadrupole mass spectrometer. The ions are then sorted according to their mass-to-charge ratios by a quadrupole mass spectrometer. The ions transmitted through the quadrupole are quantified by a channel electron multiplier and the ion information processed by a data handling system. Although most polyatomic interference is removed, any additional interference must be assessed and valid corrections applied or the data flagged to indicate problems. Interference correction must include compensation for background ions contributed by the plasma gas, reagents, and constituents of the sample matrix. Instrument drift as well as suppressions or enhancements of instrument response caused by the sample matrix must be corrected for by the use of internal standards. | 1e-05 | Y | Y | Y | 4 | 4/12/2023 | 4/29/2024 | 5/16/2024 2:00:14 PM | 2009-09-08 |  | RDGIBSON |  | Concentration | mg/L | Y | (fl. conc.) |  | Bismuth Dissolved (fl. conc.) | 7440-69-9 | Metals - Dissolved; | Metals - Dissolved; EMS code: BI-D | Chemical | Dissolved | LAB
BI-D | mg/L | Bismuth Dissolved | F039 | Field filt + acid, ICP-AES | Dissolved and extractable water samples are ready for analysis once collected. Total samples are autoclaved with HNO3 and hydrochloric acid (HCI) in capped Teflon vials. Analysis is performed via Inductively Coupled Argon Plasma Atomic Emission Spectrometry (ICP-AES)(simultaneous multielement analysis), Graphite Furnace Atomic Absorption Spectrometry (GF-AAS)(low level analysis) and/or Flame Atomic Absorption Spectrometry (F-AAS). | 0.01 | Y | Y | Y | 1480 | 1/18/2002 | 7/5/2019 1:21:00 PM | 7/12/2019 6:49:16 PM | 1998-05-21 | 6/3/1998 6:51:49 PM | NPEPPIN |  | Concentration | mg/L | Y | (fl. conc.) |  | Bismuth Dissolved (fl. conc.) | 7440-69-9 | Metals - Dissolved; | Metals - Dissolved; EMS code: BI-D | Chemical | Dissolved | LAB
BI-D | mg/L | Bismuth Dissolved | F068 | HNO3dig,TeflonTube/Autoclave,ICP-MS | This method is applicable to determiniation of dissolved, extractable, and total recoverable metals in ground waters, surface waters, drinking waters, and wastewaters. Dissolved, extractable, and total recoverable sample fractions are acidified to pH <1.5 with analytical grade nitric acid (HNO3) to prevent the metals from adsorbing to the container material or precipitating. Extractable and total recoverable water samples are collected in 250 ml acid washed polyethylene bottles and are preserved with Nitric acid (0.5 ml acid for 100 mLs of sample). All preserved samples are stored at room temperature and can be held for up to 6 months. Unfiltered, un-preserved samples are stored at 4C for no more than 30 days. For dissolved and extractable water fractions, no sample preparation is necessary. Samples are simply poured in autosampler tubes (i.e. 15 ml centrifuge tubes), and placed on the autosampler and analyzed. Samples for total recoverable metals are digested with Nitric acid in special Teflon tubes in an autoclave prior to analyses. | 2e-05 | Y | Y | Y | 1398 | 6/1/1999 1:20:00 PM | 4/17/2015 1:06:00 PM | 8/11/2021 12:04:11 PM | 1999-06-28 |  | NPEPPIN |  | Concentration | mg/L | Y | (fl. conc.) |  | Bismuth Dissolved (fl. conc.) | 7440-69-9 | Metals - Dissolved; | Metals - Dissolved; EMS code: BI-D | Chemical | Dissolved | LAB
BI-D | mg/L | Bismuth Dissolved | X170 | ICP | This code is a catch-all for a large variety of ICP methods including water and also solids digested into solutions as an intermediate result. Different procedures will have greatly different detection limits. (a) For air particulates, preliminary results are reported in mg/L and loading results as mg/dm2/d. (b) For Hi Vol samples preliminiary results are same as reported in mg/L and loading results as ug/m3. (c) Note that for Rain samples there is no filtration nor digestion steps. | 0.02 | Y | Y | Y | 6135 | 12/3/1996 | 11/6/2018 | 6/12/2020 1:21:11 PM | 1996-04-12 | 6/3/1998 6:51:34 PM | EMS |  | Concentration | mg/L | Y | (fl. conc.) |  | Bismuth Dissolved (fl. conc.) | 7440-69-9 | Metals - Dissolved; | Metals - Dissolved; EMS code: BI-D | Chemical | Dissolved | LAB
BI-D | mg/L | Bismuth Dissolved | F032 | Filt.,Acid,ICPMS | There is no acid digestion required for this analysis. A water sample is collected and preserved in 0.2 percent nitric acid. To analyse the extractable metals the sample is shaken, allowed to sit for at least twenty four hours, and a decanted portion is introduced to the spectrometer. When the dissolved metals are required, the samples must first be filtered through a 0.45 micron cellulose acetate filter paper and then followed by preservation in 0.2 percent nitric acid. It is important to note that different laboratories may have very different detection limits for some variables analyzed by ICPMS. | 0.01 | Y | Y | Y | 73357 | 3/10/1998 2:03:00 PM | 8/26/2024 | 8/26/2024 4:53:05 PM | 1998-05-14 | 4/29/2011 8:47:55 AM | NPEPPIN | RDGIBSON | Concentration | mg/L | Y | (fl. conc.) |  | Bismuth Dissolved (fl. conc.) | 7440-69-9 | Metals - Dissolved; | Metals - Dissolved; EMS code: BI-D | Chemical | Dissolved | LAB
BI-E | mg/L | Bismuth Extractable | X553 | Metals in water, extractable, ICP-SFMS | The whole water sample is collected in a HDPE/Teflon bottle. In a Class 100 workstation it is preserved with concentrated Nitric Acid at concentration of 0.2 percent (v/v). The acid preservative extracts the metals from material in the sample and a decanted portion is introduced directly to an inductively coupled argon plasma-sector field mass spectrometer (ICP-SFMS) by pneumatic nebulization. Each respective element is measured at a specific mass to charge ratio, m/z value expressed in amu (atomic mass units), and at a specified instrument resolution. LR: Low Resolution ~300, MR: Mid Resolution ~4000 and HR: High Resolution ~10000. Ref: Method number 02-2061, National Laboratory for Environmental Testing, CCIW, Environment Canada, Burlington. Requested April 2003. | 1e-06 | Y | Y | Y | 2583 | 5/22/2003 | 9/21/2007 | 4/18/2018 9:28:02 AM | 2004-01-19 |  | UMULLER |  | Concentration | mg/L | Y | (fl. conc.) |  | Bismuth Extractable (fl. conc.) | 7440-69-9 | Metals - Extractable; | Metals - Extractable; EMS code: BI-E | Chemical | Extractable | LAB
BI-L | mg/L | Bismuth Leachable | MEND | MEND (mod) EPA 300.1 (mod) | MEND (mod) EPA 300.1 (mod). Inorganic anions are analyzed by obtaining an extract from a soil sample that has been added in a 3:1 ratio of deionized water to soil; then leached and the extract filtered prior to analysis; which is then analyzed by Ion Chromatography with conductivity and/or UV detection. | 0.0005 | Y | Y | Y | 0 |  |  |  | 2024-01-29 |  | RDGIBSON |  | Concentration | mg/L | Y | (fl. conc.) |  | Bismuth Leachable (fl. conc.) |  |  | EMS code: BI-L | Chemical |  | LAB
BI-T | mg/L | Bismuth Total | F048 | Acid dig,conc'n,nebul.,ICP-OES | This method provides an automated procedure for the Inductively Coupled Argon Plasma Optical Emission Spectrophotometric (ICP-OES) determination and quantitation of trace amounts of metals in surface, ground and precipitation waters after a manual acid digestion and concentration procedure and ultrasonic nebulisation. Organically-bound metals and metal complexes are converted to metal ions by digestion with nitric acid [method details says HCl]. To increase the signal detected by the ICP, the samples are concentrated 10 times. The liquid sample is pumped onto the piezoelectric transducer, where the ultrasonic action efficiently converts it to a fine, dense aerosol. The nebulizer carrier gas flow then transports the wet aerosol through a heated U-tube, where the solvent is vaporized. Subsequently, most of the solvent vapour is removed from the aerosol stream by passing the stream through a chilled condenser. The result is a dry, analyte-laden aerosol which is introduced to the plasma. There, the high temperature desolvates the aerosol and dissociates the sample material into individual atoms and ions, which are excited to emit light at wave lengths characteristic of the elements present. Dissolved samples are field filtered through 0.45 micron cellulose acetate filter prior to acid preservation. Samples for extractable analysis are acidified with 2 mL concentrated HNO3 per litre and allowed to stand a minimum of 24 hours before analysis. Clients providing these samples are advised to preserve samples at the collection site. | 0.1 | Y | Y | Y | 29 |  |  | 7/6/2001 11:10:10 AM | 1998-11-13 |  | NPEPPIN |  | Concentration | mg/L | Y | (fl. conc.) |  | Bismuth Total (fl. conc.) | 7440-69-9 | Metals - Total; | Metals - Total; EMS code: BI-T | Chemical | Total | LAB
BI-T | ug/g | Bismuth Total | FS01 | D&G,Nitric/HCl:ICP | Samples are dried at 60 C, then either sieved to -100mesh (150 um) and rolled to homogenize, or finely ground with a mortar and pestle or a shatter box laboratory mill and rolled (depending on the submitter's requirements). The samples are then weighed into Teflon digestion liners and digested with HNO3 and HCl in a microwave oven. Samples are cooled, volumized, and allowed to settle overnight before being decanted. Analysis is then performed via Inductively Coupled Argon Plasma Atomic Emission Spectrometry (ICP-AES) (Simultaneous multi element analysis). In ICP Spectrometry, the liquid sample is pumped to the nebulizer, where it is convertd into an aerosol. the nebulizer carrier gas flow than transports the wet aerosol past the impact bead, where the large droplets condense out and flow to the waste container. The remaining fine aerosol is carried into the plasma. There, the high temperature desolvates the aerosol and dissociates the sample material into individual atoms and ions, which are excited to emit light at wavelengths characteristic of the elements present. A diffraction grating in the spectrometer diffracts the various wavelengths of light through exit slits to the photo-multiplier tubes, which convert the light energy to electronic signals. | 10 | Y | Y | Y | 271 | 1/31/2002 | 3/22/2002 | 4/1/2004 1:19:34 AM | 2004-02-10 |  | UMULLER |  | MassPerMass | mg/kg | Y | (mass conc.) |  | Bismuth Total (mass conc.) | 7440-69-9 | Metals - Total; | Metals - Total; EMS code: BI-T | Chemical | Total | LAB
BI-T | mg/L | Bismuth Total | MSEA | Metals in Seawater (undigested) by (CRC) ICP-MS | This method describes the multi-elemental determination of trace elements by Collision/Reaction Cell (CRC) ICP-MS. The method measures ions produced by a radio-frequency inductively coupled plasma. Analyte species originating in a liquid are pneumatically nebulized and the resulting aerosol transported by argon gas into the plasma torch. The ions produced are entrained in the plasma gas and extracted, by means of a differentially pumped vacuum interface, into a mass spectrometer. The ions flow through a Collision/Reaction Cell (CRC) where depending on the mode of analysis either the ions passes through the cell unhindered (no gas - classical ICP-MS mode), or collisions/reactions chemistry takes place between the ions and a neutral gas such as hydrogen or helium. Any polyatomic ions which give interference on the analytes are removed. The ions are then sorted according to their mass-to-charge ratios by a quadrupole mass spectrometer. The ions are then sorted according to their mass-to-charge ratios by a quadrupole mass spectrometer. The ions transmitted through the quadrupole are quantified by a channel electron multiplier and the ion information processed by a data handling system. Although most polyatomic interference is removed, any additional interference must be assessed and valid corrections applied or the data flagged to indicate problems. Interference correction must include compensation for background ions contributed by the plasma gas, reagents, and constituents of the sample matrix. Instrument drift as well as suppressions or enhancements of instrument response caused by the sample matrix must be corrected for by the use of internal standards. | 1e-05 | Y | Y | Y | 441 | 8/12/2009 12:08:00 AM | 4/29/2024 | 6/7/2024 10:03:16 AM | 2009-09-08 |  | RDGIBSON |  | Concentration | mg/L | Y | (fl. conc.) |  | Bismuth Total (fl. conc.) | 7440-69-9 | Metals - Total; | Metals - Total; EMS code: BI-T | Chemical | Total | LAB
BI-T | ug/m3 | Bismuth Total | 6040 | Aqua Regia Dig:ICP | Aqua regia digestion is used to bring the metals into solution prior to ICP analysis. For Hi-Vol samples, preliminary results are reported as mg/L and24 hr loadings as ug/m3 | 0.01 | Y | Y | Y | 437 | 2/13/2002 | 2/9/2005 | 2/23/2005 1:13:47 AM | 1998-05-05 | 6/3/1998 6:51:48 PM | NPEPPIN |  | AirConcentration | µg/m³ | Y | (air conc.) |  | Bismuth Total (air conc.) | 7440-69-9 | Metals - Total; | Metals - Total; EMS code: BI-T | Chemical | Total | LAB
BI-T | mg/L | Bismuth Total | X170 | ICP | This code is a catch-all for a large variety of ICP methods including water and also solids digested into solutions as an intermediate result. Different procedures will have greatly different detection limits. (a) For air particulates, preliminary results are reported in mg/L and loading results as mg/dm2/d. (b) For Hi Vol samples preliminiary results are same as reported in mg/L and loading results as ug/m3. (c) Note that for Rain samples there is no filtration nor digestion steps. | 0.02 | Y | Y | Y | 2645 | 3/29/2001 | 11/7/2018 | 10/24/2022 10:58:25 AM | 1996-04-12 | 6/3/1998 6:51:34 PM | EMS |  | Concentration | mg/L | Y | (fl. conc.) |  | Bismuth Total (fl. conc.) | 7440-69-9 | Metals - Total; | Metals - Total; EMS code: BI-T | Chemical | Total | LAB
BI-T | ppm (S) | Bismuth Total | MS61 | ALS Minerals ME-MS61 | ALS Minerals ME-MS61. A prepared sample is digested with HNO3; HF; and HClO4 in a graphite heating block to near dryness. After cooling; the sample is boiled in diluted HCl. Bulk to volume with deionized water; mixed and analyzed by ICP-AES and ICP-MS for all elements. | 0.01 | Y | Y | Y | 0 |  |  |  | 2024-01-29 |  | EMS |  | DimensionlessRatio | ppm | N | (ratio) |  | Bismuth Total (ratio) | 7440-69-9 | Metals - Total; | Metals - Total; EMS code: BI-T | Chemical | Total | LAB
BI-T | ug/g | Bismuth Total | PB02 | PBM,HNO3/HCl dig,dry60C,100mesh, | Performance Based Method. Samples are dried at 60 C and sieved using a .149 mm (100 mesh) sieve. Solid samples are digested with the combination of nitric and hydrochloric acids with a 1:1 ratio of acids. For environmental purposes, strong acid leaches are used routinely since they have the advantage of speed, low cost, and superio reproducibility. All metals trapped within the refractory matrix are unlikely to have environmental consequences. Instrumental analysis of extracts produced using this digestion method can be performed by a variety of techniques. |  | Y | Y | Y | 52 | 5/24/2001 3:42:00 PM | 11/26/2001 10:23:00 AM | 5/26/2003 2:27:28 PM | 2001-10-11 |  | NPEPPIN |  | MassPerMass | mg/kg | Y | (mass conc.) |  | Bismuth Total (mass conc.) | 7440-69-9 | Metals - Total; | Metals - Total; EMS code: BI-T | Chemical | Total | LAB
BI-T | ug/g | Bismuth Total | ICPM | Metals in Soil by CRC ICPMS | This analysis is carried out using procedures from CSR Analytical Method: "Strong Acid Leachable Metals (SALM) in Soil", BC Ministry of Environment, 26 June 2009, and procedures adapted from EPA Method 200.2. The sample is manually homogenized, dried at 60 degrees Celsius, sieved through a 63 um (230 mesh) sieve, and a representative subsample of the dry material is weighed. The sample is then digested at 95 degrees Celsius for 2 hours by block digester using concentrated nitric and hydrochloric acids. Instrumental analysis is by collision cell inductively coupled plasma - mass spectrometry (modifed from EPA Method 6020A). Method Limitation: This method is not a total digestion technique. It is a very strong acid digestion that is intended to dissolve those metals that may be environmentally available. By design, elements bound in silicate structures are not normally dissolved by this procedure as they are not usually mobile in the environment. | 0.1 | Y | Y | Y | 103 | 9/19/2017 | 9/24/2023 | 1/12/2024 11:16:52 AM | 2018-01-09 |  | JKROGH |  | MassPerMass | mg/kg | Y | (mass conc.) |  | Bismuth Total (mass conc.) | 7440-69-9 | Metals - Total; | Metals - Total; EMS code: BI-T | Chemical | Total | LAB
BI-T | ug/g | Bismuth Total | X492 | Homo,HCl-HNO3 dig,ICP-OES,dry calc | Homogenize samples and take sub-samples. For Total Recoverable Metals Digest with Aqua Regia. Analyses by ICPOES (Inductively Coupled Plasma Optical Emission Spectroscopy). Perform moisture analysis on separate sub-sample and calculate dry value for trace elements. | 10 | Y | Y | Y | 448 | 6/20/2005 | 10/10/2007 | 5/24/2008 1:15:02 AM | 1998-11-13 |  | NPEPPIN |  | MassPerMass | mg/kg | Y | (mass conc.) |  | Bismuth Total (mass conc.) | 7440-69-9 | Metals - Total; | Metals - Total; EMS code: BI-T | Chemical | Total | LAB
BI-T | ppm (S) | Bismuth Total | MS41 | ALS Minerals ME-MS41 | ALS Minerals ME-MS41. A prepared sample is digested with aqua regia in a graphite heating block. After cooling; the resulting solution is diluted to with deionized water; mixed and analyzed by ICP-AES. Samples are then analyzed by ICP-MS for the remaining suite of elements. | 0.01 | Y | Y | Y | 0 |  |  |  | 2024-01-29 |  | EMS |  | DimensionlessRatio | ppm | N | (ratio) |  | Bismuth Total (ratio) | 7440-69-9 | Metals - Total; | Metals - Total; EMS code: BI-T | Chemical | Total | LAB
BI-T | mg/L | Bismuth Total | F033 | Acid Dig,Centrifuge,ICPMS | For the determination of total recoverable elements in water or waste water, take 100 ml aliquot from a well mixed, acid preserved, sample containing not more than 0.25% (w/v) total solids and transfer to a 250 ml Griffin beaker. (If total solids are greater than 0.25% reduce the size of the aliquot by a proportionate amount). Add 1 ml of conc. nitric acid and 0.5 ml conc. hydrochloric acid. Heat on a hot plate until the volume has been reduced to 15-20ml, ensuring that the sample does not boil. Allow to cool and transfer to a 50 ml volumetric flask. Dilute to volume with ASTM type I water. Centrifuge the sample or allow to stand overnight to separate insoluble material. Prior analysis, pipette 20 ml into a 50 ml volumetric flask and dilute to volume with ASTM type I water. Analysis should be performed within seventy two hours of the completed preparation. Note that different laboratories may have very different field preservation, digestion procedures and detection limits for this method, so these should be documented. | 0.01 | Y | Y | Y | 36503 | 11/20/1997 9:31:00 AM | 8/22/2024 3:24:00 AM | 8/25/2024 9:47:36 PM | 1998-05-14 | 6/3/1998 6:51:49 PM | NPEPPIN |  | Concentration | mg/L | Y | (fl. conc.) |  | Bismuth Total (fl. conc.) | 7440-69-9 | Metals - Total; | Metals - Total; EMS code: BI-T | Chemical | Total | LAB
BI-T | ug/m3 | Bismuth Total | 5312 | Tef:N/P Dig:ICP Analysis | A nitric/perchloric acid digestion if used to bring the metals into solution prior to ICP analysis. For Hi-Vol samples preliminary results are reported as mg/L and 24 hour loading results as ug/m3. | 0.01 | Y | Y | Y | 1185 | 11/7/1997 | 3/29/2000 | 9/28/2000 1:34:40 PM | 1996-10-13 | 6/3/1998 6:51:46 PM | EMS |  | AirConcentration | µg/m³ | Y | (air conc.) |  | Bismuth Total (air conc.) | 7440-69-9 | Metals - Total; | Metals - Total; EMS code: BI-T | Chemical | Total | LAB
BI-T | ug/g | Bismuth Total | FS82 | PBM,dry60C,10mesh,SALM-ICP-MS | PBM;dry60C;10mesh;SALM-ICP-MS | 0.1 | Y | Y | Y | 2279 | 2/4/2002 | 7/24/2024 | 7/29/2024 5:17:09 PM | 2002-11-12 |  | UMULLER |  | MassPerMass | mg/kg | Y | (mass conc.) |  | Bismuth Total (mass conc.) | 7440-69-9 | Metals - Total; | Metals - Total; EMS code: BI-T | Chemical | Total | LAB
BI-T | ug/m3 | Bismuth Total | F088 | Aliquot,HNO3/HCl dig, ICPMS,loading | Aliquot microwave digested with 3:1 HNO3/HCL. Analysis by ICPMS. For HiVol samples preliminary result are reported in mg/L and loading results are reported in ug/m3 |  | Y | Y | Y | 565 | 8/17/2000 10:46:00 AM | 12/14/2006 | 12/30/2006 1:01:36 AM | 2000-08-16 |  | NPEPPIN |  | AirConcentration | µg/m³ | Y | (air conc.) |  | Bismuth Total (air conc.) | 7440-69-9 | Metals - Total; | Metals - Total; EMS code: BI-T | Chemical | Total | LAB
BI-T | mg/L | Bismuth Total | X561 | Total metals: FA, HNO3, HR-ICPMS (No Prep)Prep) | Ultra trace metals in water are analyzed by high resolution inductively coupled plasma mass spectrometry (HR-ICPMS), modified from US EPA 200.8, (Revision 5.5). The detection limits provided can only be met for undigested samples. This procedure is intended for colorless, non-turbid, acid-preserved water samples (i.e. pristine water samples), having turbidity < 1 NTU and no odor. Where turbidity exceeds 1 NTU, and/or the sample is colored and has an odor, results may be biased low compared to true Total Metals concentrations. ALS recommends that turbidity analysis be requested on samples submitted for this test to aid with interpretation of results. | 5e-06 | Y | Y | Y | 2756 | 6/8/2015 | 10/5/2023 | 1/31/2024 2:37:41 PM | 2015-10-22 |  | RDGIBSON |  | Concentration | mg/L | Y | (fl. conc.) |  | Bismuth Total (fl. conc.) | 7440-69-9 | Metals - Total; | Metals - Total; EMS code: BI-T | Chemical | Total | LAB
BI-T | ppm (S) | Bismuth Total | 6041 | Aqua Regia by ICPMS | These were sediment/soils digested with Aqua-Regia (Nitric Acid & Hydrochloric Acids) and then analysis by ICPMS. Results would be reported in mg/kg and/or %. | 0.0005 | Y | Y | Y | 21 | 10/13/2019 | 8/5/2020 | 8/10/2020 3:52:06 PM | 2019-09-06 | 10/21/2019 11:54:26 AM | KATFRASE | KATFRASE | DimensionlessRatio | ppm | N | (ratio) |  | Bismuth Total (ratio) | 7440-69-9 | Metals - Total; | Metals - Total; EMS code: BI-T | Chemical | Total | LAB
BI-T | ug/g wet | Bismuth Total | BPSY | MET-WET-BPSY-CCMS-VA | Metals in Animal Tissue and Vegetation (Biota) - Prescriptive. EPA 200.3/6020B This method is conducted following British Columbia Lab Manual method "Metals in Animal Tissue and Vegetation (Biota) - Prescriptive". Tissue samples are digested with nitric and hydrochloric acids; in combination with addition of hydrogen peroxide. Instrumental analysis is by collision cell inductively coupled plasma - mass spectrometry (modified from EPA Method 6020B). | 0.002 | Y | Y | Y | 2 | 11/7/2019 | 11/7/2019 | 6/10/2021 1:57:10 PM | 2021-06-04 |  | EMS |  | MassPerMass | µg/g | Y | (mass conc.) | (wet) | Bismuth Total (wet) (mass conc.) | 7440-69-9 | Metals - Total; | Metals - Total; EMS code: BI-T | Chemical | Total | LAB
BI-T | mg/L | Bismuth Total | X349 | FA:HNO3:ICP | Samples are field-acidified with HNO3; HNO3 digestion; analysis by ICP/AES. | 0.01 | Y | Y | Y | 30 |  |  | 10/30/2014 8:31:17 AM | 1998-05-22 | 6/3/1998 6:51:49 PM | NPEPPIN |  | Concentration | mg/L | Y | (fl. conc.) |  | Bismuth Total (fl. conc.) | 7440-69-9 | Metals - Total; | Metals - Total; EMS code: BI-T | Chemical | Total | LAB
BI-T | mg/L | Bismuth Total | 0040 | HNO3 Dig: ICP Analysis | HNO3 Dig: ICP Analysis | 0.02 | Y | Y | Y | 1824 | 6/9/2003 12:01:00 PM | 6/25/2010 3:15:00 PM | 7/22/2010 2:12:20 PM | 1996-04-12 | 6/3/1998 6:51:34 PM | EMS |  | Concentration | mg/L | Y | (fl. conc.) |  | Bismuth Total (fl. conc.) | 7440-69-9 | Metals - Total; | Metals - Total; EMS code: BI-T | Chemical | Total | LAB
BI-T | ug/g | Bismuth Total | X551 | Metals in Biota (dry), total, acid digestion, ICP-SFMS | 0.1 gram of dried, homogenized biota is weighed in a microwave TFM vessel, and 8mL of Nitric Acid and 1mL of Hydrogen Peroxide are added. The sample is digested at 200 degrees C for 15 minutes in a closed vessel high pressure microwave oven. The sample is cooled, the volume adjusted to 100mL and introduced to an iductively coupled argon plasma-sector field mass spectrometer (ICP-SFMS) by pneumatic nebulization. Each respective element is measured at a specific mass to charge ratio, m/z value expressed in amu (atomic mass units), and at a specified instrument resolution. LR: Low Resolution ~300, MR: Mid Resolution ~4000 and HR: Hich Resolution ~10000. Ref: Method number 02-2705, National Laboratory for Environmental Testing, CCIW, Environment Canada, Burlington. Requested April 2003. | 0.001 | Y | Y | Y | 46 | 8/11/2015 | 10/30/2023 | 11/28/2023 2:21:33 PM | 2004-01-19 |  | UMULLER |  | MassPerMass | mg/kg | Y | (mass conc.) |  | Bismuth Total (mass conc.) | 7440-69-9 | Metals - Total; | Metals - Total; EMS code: BI-T | Chemical | Total | LAB
BI-T | mg/L | Bismuth Total | 1349 | Tot Metals in Seawater by CRC ICPMS (BC) | Tot Metals in Seawater by CRC ICPMS (BC) | 0.0005 | Y | Y | Y | 98 | 8/16/2019 | 12/5/2021 | 12/13/2021 12:34:16 PM | 2019-09-30 |  | KATFRASE |  | Concentration | mg/L | Y | (fl. conc.) |  | Bismuth Total (fl. conc.) | 7440-69-9 | Metals - Total; | Metals - Total; EMS code: BI-T | Chemical | Total | LAB
BI-T | mg/L | Bismuth Total | 5038 | N/P Dig, ICP | Aliquot of prepared sample is digested first with nitric acid then with perchloric acid. Analysis by ICP. (a) For air particulates, preliminary results are reported in mg/L and loading results as mg/dm2/d. (b) For Hi Vol samples preliminary results are same as reported in mg/L and loading results as ug/m3. | 0.024 | Y | Y | Y | 4 | 12/1/2010 | 12/1/2010 | 12/9/2010 3:55:26 PM | 1998-02-17 | 6/3/1998 6:51:48 PM | NPEPPIN |  | Concentration | mg/L | Y | (fl. conc.) |  | Bismuth Total (fl. conc.) | 7440-69-9 | Metals - Total; | Metals - Total; EMS code: BI-T | Chemical | Total | LAB
BI-T | ppm (S) | Bismuth Total | 6042 | Aqua Regia by ICPMS | These were sediment/soils digested with Aqua-Regia (Nitric Acid & Hydrochloric Acids) and then analysis by ICPMS. Results would be reported in ppm s. | 0.0001 | Y | Y | Y | 71 | 5/14/2022 | 7/27/2024 | 7/29/2024 9:13:08 AM | 2022-06-28 |  | EMS |  | DimensionlessRatio | ppm | N | (ratio) |  | Bismuth Total (ratio) | 7440-69-9 | Metals - Total; | Metals - Total; EMS code: BI-T | Chemical | Total | LAB
BI-T | mg/dm2/d | Bismuth Total | DUST | HNO3 dig.(microwave or hot block),ICPMS | HNO3 dig.(microwave or hot block);ICPMS |  | Y | Y | Y | 267 | 7/19/2016 1:11:00 PM | 7/31/2024 | 7/31/2024 11:23:12 AM | 2008-02-07 |  | RDGIBSON |  | MassFlux | mg/dm²/d | Y | (mass flux) |  | Bismuth Total (mass flux) | 7440-69-9 | Metals - Total; | Metals - Total; EMS code: BI-T | Chemical | Total | LAB
BI-T | mg/L | Bismuth Total | F082 | Aliquot,HNO3/HCL dig,ICPMS | Aliquot microwave digested with 3:1 HNO3/HCL. Analysis by ICPMS. For HiVol samples preliminary result are reported in mg/L and loading results are reported in ug/m3 |  | Y | Y | Y | 26928 | 7/26/2005 7:04:00 AM | 8/23/2024 | 8/26/2024 4:53:05 PM | 2000-06-13 |  | NPEPPIN |  | Concentration | mg/L | Y | (fl. conc.) |  | Bismuth Total (fl. conc.) | 7440-69-9 | Metals - Total; | Metals - Total; EMS code: BI-T | Chemical | Total | LAB
BI-T | mg/L | Bismuth Total | F038 | HNO3/HCl dig,ICP-AES | Dissolved and extractable water samples are ready for analysis once collected. Total samples are autoclaved with HNO3 and hydrochloric acid (HCI) in capped Teflon vials. Analysis is performed via Inductively Coupled Argon Plasma Atomic Emission Spectrometry (ICP-AES)(simultaneous multielement analysis), Graphite Furnace Atomic Absorption Spectrometry (GF-AAS) (low level analysis) and/or Flame Atomic Absorption Spectrometry (F-AAS). | 0.01 | Y | Y | Y | 27 | 1/31/2002 | 3/22/2002 | 2/13/2004 1:01:28 AM | 1998-05-21 | 6/3/1998 6:51:49 PM | NPEPPIN |  | Concentration | mg/L | Y | (fl. conc.) |  | Bismuth Total (fl. conc.) | 7440-69-9 | Metals - Total; | Metals - Total; EMS code: BI-T | Chemical | Total | LAB
BI-T | ug/g wet | Bismuth Total | MTIW | MET-WET-CCMS-MID-VA | Metals in Tissue by CRC ICPMS (WET) Methodology: This method is conducted following British Columbia Lab Manual method "Metals in Animal Tissue and Vegetation (Biota) - Prescriptive". Tissue samples are homogenized and sub-sampled prior to hotblock digestion with nitric and hydrochloric acids, in combination with addition of hydrogen peroxide. Instrumental analysis is by collision cell inductively coupled plasma - mass spectrometry (modified from EPA Method 6020A). Method Limitation: This method employs a strong acid/peroxide digestion, and is intended to provide a conservative estimate of bio-available metals. Near complete recoveries are achieved for most toxicologically important metals, but elements associated with recalcitrant minerals may be only partially recovered. | 0.002 | Y | Y | Y | 24 | 11/25/2018 | 12/6/2018 | 7/21/2021 2:55:07 PM | 2018-12-24 |  | EMS |  | MassPerMass | µg/g | Y | (mass conc.) | (wet) | Bismuth Total (wet) (mass conc.) | 7440-69-9 | Metals - Total; | Metals - Total; EMS code: BI-T | Chemical | Total | LAB
BI-T | ug/g | Bismuth Total | MTID | MET-DRY-CCMS-VA | Metals in Tissue by CRC ICPMS (DRY) Methodology: This method is conducted following British Columbia Lab Manual method "Metals in Animal Tissue and Vegetation (Biota) - Prescriptive". Tissue samples are homogenized and sub-sampled prior to hotblock digestion with nitric and hydrochloric acids, in combination with addition of hydrogen peroxide. Instrumental analysis is by collision cell inductively coupled plasma - mass spectrometry (modified from EPA Method 6020A). Method Limitation: This method employs a strong acid/peroxide digestion, and is intended to provide a conservative estimate of bio-available metals. Near complete recoveries are achieved for most toxicologically important metals, but elements associated with recalcitrant minerals may be only partially recovered. | 0.01 | Y | Y | Y | 122 | 8/13/2014 | 9/8/2022 | 10/26/2022 7:14:14 AM | 2015-03-09 | 3/11/2015 4:12:05 PM | RDGIBSON | EMS | MassPerMass | mg/kg | Y | (mass conc.) |  | Bismuth Total (mass conc.) | 7440-69-9 | Metals - Total; | Metals - Total; EMS code: BI-T | Chemical | Total | LAB
BI-T | mg/L | Bismuth Total | 0042 | GI HNO3 Dig: ICP | Sample aliquot taken from unfiltered unpreserved General Ions bottle; HNO3 digestion; ICP analysis. | 0.02 | Y | Y | Y | 6258 | 9/20/1996 | 7/5/2019 3:21:00 PM | 7/12/2019 6:49:13 PM | 1996-04-12 | 6/3/1998 6:51:34 PM | EMS |  | Concentration | mg/L | Y | (fl. conc.) |  | Bismuth Total (fl. conc.) | 7440-69-9 | Metals - Total; | Metals - Total; EMS code: BI-T | Chemical | Total | LAB
BI-T | mg/L | Bismuth Total | MWID | Metals in Water by ICPMS (digested) | Metals in Water by ICPMS (digested) Based on EPA 200.8 | 0.001 | Y | Y | Y | 36785 |  |  | 7/31/2024 1:46:41 PM | 2009-12-24 |  | RDGIBSON |  | Concentration | mg/L | Y | (fl. conc.) |  | Bismuth Total (fl. conc.) | 7440-69-9 | Metals - Total; | Metals - Total; EMS code: BI-T | Chemical | Total | LAB
BI-T | mg/L | Bismuth Total | X555 | Metals in water, total, "in-bottle" digestion, ICP-SFMS | The whole water sample collected in a HDPE/Teflon bottle, is de-capped in a Class 100 workstation, acidified to 2% (v/v) HNO3 and 1% (v/v) HCl and recapped. The sample is "in-bottle" digested in a convection oven, set at 60 degrees C for 16 hours. The sample is cooled and introduced to an inductively coupled argon plasma-sector field mass spectrometer (ICP-SFMS) by pneumatic nebulization. Each respective element is measured at a specific mass to charge ratio, m/z value expressed in amu (atomic mass units), and at resolution. LR: Low Resolution ~300, MR: Mid Resolution ~4000 and HR: High a specified instrument Resolution ~10000. Ref: Method number 02-2003, National Laboratory for Environmental Testing, CCIW, Environment Canada, Burlington. Requested April 2003. | 1e-06 | Y | Y | Y | 17341 | 5/20/2003 | 12/6/2023 2:08:00 AM | 2/14/2024 9:05:32 AM | 2004-01-14 |  | UMULLER |  | Concentration | mg/L | Y | (fl. conc.) |  | Bismuth Total (fl. conc.) | 7440-69-9 | Metals - Total; | Metals - Total; EMS code: BI-T | Chemical | Total | LAB
BI-T | ug/g wet | Bismuth Total | X552 | Metals in Biota (wet), total, acid digestion, ICP-SFMS | One gram of wet, homogenized biota is weighed in a microwave TFM vessel, and 8mL of Nitric Acid and 1mL of Hydrogen Peroxide are added. The sample is digested at 200 degrees C for 15 minutes in a closed vessel high pressure microwave oven. The sample is cooled, the volume adjusted to 100mL and introduced to an inductively coupled argon plasma-sector field mass spectrometer (ICP-SFMS) by pneumatic nebulization. Each respective element is measured at a specific mass to charge ratio, m/z value expressed in amu (atomic mass units), and at a specified instrument resolution. LR: Low Resolution ~300, MR: Mid Resolution ~4000. Ref: 02-2704, National Laboratory for Environmental Testing, CCIW, Environment Canada, Burlington. Requested April 2003. | 0.0001 | Y | Y | Y | 37 | 8/9/2016 | 10/16/2023 | 11/28/2023 2:21:32 PM | 2004-01-19 | 4/19/2021 1:22:00 PM | UMULLER | EMS | MassPerMass | µg/g | Y | (mass conc.) | (wet) | Bismuth Total (wet) (mass conc.) | 7440-69-9 | Metals - Total; | Metals - Total; EMS code: BI-T | Chemical | Total | LAB
BI-T | ug/g wet | Bismuth Total | X560 | Metals in Biota (wet), total, acid digestion, ICP-MS | Five grams of wet, homogenized biota is weighed in a centrifuge tube, and 2.5 mL of Nitric Acid and 2.5 mL of Hydrochloric acid are added. The sample is digested at room temperature for one hour and then digested at 90-95 degrees C in a water bath or heat block for one more hour. The sample is cooled and the volume is made to 50 mL with DI. The digestate is then diluted at ten fold and introduced to a collision/reaction cell inductively coupled argon plasma mass spectrometer (CRC ICPMS) by pneumatic nebulization. Each respective element is measured at a specific mass to charge ratio, m/z value expressed in amu (atomic mass units). - Maxxam Analytics (BC). Added to EMS Dec 10, 2008. | 0.1 | Y | Y | Y | 189 | 11/15/2008 4:00:00 AM | 10/30/2023 | 11/28/2023 2:21:33 PM | 2008-12-11 | 3/11/2015 4:17:10 PM | RDGIBSON | EMS | MassPerMass | µg/g | Y | (mass conc.) | (wet) | Bismuth Total (wet) (mass conc.) | 7440-69-9 | Metals - Total; | Metals - Total; EMS code: BI-T | Chemical | Total | LAB
BI-T | mg/L | Bismuth Total | X285 | HNO3 Dig:ICP Analysis | HNO3 Dig:ICP Analysis | 0.01 | Y | Y | Y | 565 | 8/10/1993 | 10/5/1993 | 4/30/2004 11:32:37 AM | 1998-05-22 | 6/3/1998 6:51:49 PM | NPEPPIN |  | Concentration | mg/L | Y | (fl. conc.) |  | Bismuth Total (fl. conc.) | 7440-69-9 | Metals - Total; | Metals - Total; EMS code: BI-T | Chemical | Total | LAB
BI-T | mg/L | Bismuth Total | F068 | HNO3dig,TeflonTube/Autoclave,ICP-MS | This method is applicable to determiniation of dissolved, extractable, and total recoverable metals in ground waters, surface waters, drinking waters, and wastewaters. Dissolved, extractable, and total recoverable sample fractions are acidified to pH <1.5 with analytical grade nitric acid (HNO3) to prevent the metals from adsorbing to the container material or precipitating. Extractable and total recoverable water samples are collected in 250 ml acid washed polyethylene bottles and are preserved with Nitric acid (0.5 ml acid for 100 mLs of sample). All preserved samples are stored at room temperature and can be held for up to 6 months. Unfiltered, un-preserved samples are stored at 4C for no more than 30 days. For dissolved and extractable water fractions, no sample preparation is necessary. Samples are simply poured in autosampler tubes (i.e. 15 ml centrifuge tubes), and placed on the autosampler and analyzed. Samples for total recoverable metals are digested with Nitric acid in special Teflon tubes in an autoclave prior to analyses. | 2e-05 | Y | Y | Y | 1779 | 6/1/1999 1:24:00 PM | 7/4/2008 9:48:00 AM | 7/30/2024 6:32:12 AM | 1999-06-28 |  | NPEPPIN |  | Concentration | mg/L | Y | (fl. conc.) |  | Bismuth Total (fl. conc.) | 7440-69-9 | Metals - Total; | Metals - Total; EMS code: BI-T | Chemical | Total | LAB
BI-T | mg/L | Bismuth Total | METN | MET-T-NP-CCMS-VA | Ultra trace metals in water are analyzed by CRC ICPMS, based on US EPA Method 6020B (July 2014). The detection limits provided can only be met for undigested samples. This procedure is intended for colorless, non-turbid, acid-preserved water samples (i.e. pristine water samples), having turbidity < 1 NTU and no odor. Where turbidity exceeds 1 NTU, and/or the sample is colored and has an odor, results may be biased low compared to true Total Metals concentrations. ALS recommends that turbidity analysis be requested on samples submitted for this test to aid with interpretation of results. Where turbidity is <1NTU, undigested metals are equivalent to total metals concentrations. | 5e-06 | Y | Y | Y | 3648 | 4/3/2012 | 8/19/2024 | 8/26/2024 1:03:06 PM | 2019-08-20 |  | KATFRASE |  | Concentration | mg/L | Y | (fl. conc.) |  | Bismuth Total (fl. conc.) | 7440-69-9 | Metals - Total; | Metals - Total; EMS code: BI-T | Chemical | Total | LAB
BI-T | mg/L | Bismuth Total | F121 | Acid,ICPMS | Acidified, no digestion, ICPMS | 1e-05 | Y | Y | Y | 8413 | 3/12/2010 7:57:00 PM | 8/15/2024 9:10:00 PM | 8/26/2024 8:06:03 AM | 2002-07-17 | 7/16/2012 3:27:39 PM | UMULLER | RDGIBSON | Concentration | mg/L | Y | (fl. conc.) |  | Bismuth Total (fl. conc.) | 7440-69-9 | Metals - Total; | Metals - Total; EMS code: BI-T | Chemical | Total | LAB
BIRB | CFU/mL | Iron Related Bacteria | BART | BIOLOGICAL ACTIVITY REACTION TEST | METHOD FOR THE DETERMINATION OF BACTERIAL POPULATIONS BY THE BIOLOGICAL ACTIVITY REACTION TEST? APPLICATION: This method determines the activity of bacterial populations in water, including drinking water, surface water (ponds, lakes, creeks etc.), raw water, ground water, wastewater, swimming pools and whirlpools. PRINCIPLE: Aliquots of water samples are inoculated in Biological Activity Reaction Test (BART?) vials specific to the desired organism type and incubated for eight days at room temperature. The reaction pattern is observed each day of the incubation period to determine the community activity and characteristics. Each vial type contains selective culture media specific to enriching and isolating the desired microbial community and therefore a different reaction pattern. The concentration, as colony forming unit per millilitre (cfu/mL), of the bacterial population can be extrapolated from the activity level based on the time delay for when a positive reaction is first observed in the reaction vial. METHOD SOURCES: The CARO-EDM method is performed without significant modification according to: 1. Droycon Bioconcepts Inc., Biological Activity Reaction Test BART? User Manual, 2004 Edition Supporting Documentation: CARO-EDM Microbiology Procedures Manual, CE-PM-009-current, BART-SOFT Version 6. Equipment Manuals: Autoclave ? Sanyo, Model 2400, or equivalent. DETECTION LIMITS: Reporting Detection Limits (RDL) are dictated by the method and are based on an 8 day incubation period and may be lowered if the incubation period is extended. Table 2 lists the RDLs of the BART tests commonly performed by CARO-EDM. These detection limits are derived from the BART-SOFT version 6 population calculator. Refer to the CARO LIMS database for RDLs for those parameters not listed here. LAB-BART? Type RDL (cfu/ml)* IRB=34, SRB=31, HAB=65, DN=18. | 2 | Y | Y | Y | 16 | 7/22/2013 4:24:00 PM | 2/24/2021 3:45:00 PM | 3/5/2021 5:01:06 PM | 2014-12-15 |  | RDGIBSON |  | MicrobialFormation | CFU/mL | N | (microb.) |  | Iron Related Bacteria (microb.) |  | Bacteriological; | Bacteriological; EMS code: BIRB | Chemical |  | LAB
BOD- | mg/L | Carbenaceous plus nitrogenous Biochemical Oxygen Demand | X483 | Seeded Dil No Nit Inh | The biochemical oxygen demand is determined by measuring the amount of dissolved oxygen (DO) that is consumed in the sample under aerobic conditions during the incubation period (five days at 20 C) after seeding material has been added. A polarographic membrane electrode that has been calibrated in saturated air is used to acquire all dissolved oxygen measurements. No nitrogen inhibitor added. A modified Winkler procedure may be substituted for the DO probe procedure. Performance Base Method published circa Dec. 2000. | 5 | Y | Y | Y | 5 |  |  | 1/20/2022 4:45:11 PM | 2000-11-27 |  | NPEPPIN |  | Concentration | mg/L | Y | (fl. conc.) |  | Carbenaceous plus nitrogenous Biochemical Oxygen Demand (fl. conc.) |  | Aggregate Organics; | Aggregate Organics; EMS code: BOD- | Chemical |  | LAB
BOD- | mg/L | Carbenaceous plus nitrogenous Biochemical Oxygen Demand | X482 | Unseeded Dil No Nit Inh 5 day | The biochemical oxygen demand is determined by measuring the amount of dissolved oxygen (DO) that is consumed in the sample under aerobic conditions during the incubation period (five days at 20 C). A polarographic membrane electrode that has been calibrated in saturated air is used to acquire all dissolved oxygen measurements. No nitrogen inhibitor added. No seed is added. A modified Winkler procedure may be substituted for the DO probe procedure. Performance Base Method published circa Dec. 2000. | 5 | Y | Y | Y | 5 |  |  | 4/24/2017 4:51:48 PM | 2000-11-27 |  | NPEPPIN |  | Concentration | mg/L | Y | (fl. conc.) |  | Carbenaceous plus nitrogenous Biochemical Oxygen Demand (fl. conc.) |  | Aggregate Organics; | Aggregate Organics; EMS code: BOD- | Chemical |  | LAB
BODS | mg/L | Biochemical Oxygen Demand - Soluble | XFGF | Filter Glass Fibre | Filter sample through a 1.2 um glass fiber filter paper | 10 | Y | Y | Y | 6 | 12/3/2009 2:46:00 PM | 2/8/2010 2:41:00 PM | 3/29/2010 2:40:34 PM | 2010-02-01 | 2/1/2010 12:00:49 PM | RDGIBSON | RDGIBSON | Concentration | mg/L | Y | (fl. conc.) |  | Biochemical Oxygen Demand - Soluble (fl. conc.) |  | Aggregate Organics; | Aggregate Organics; EMS code: BODS | Chemical | Soluble | LAB
BODS | mg/L | Biochemical Oxygen Demand - Soluble | X483 | Seeded Dil No Nit Inh | The biochemical oxygen demand is determined by measuring the amount of dissolved oxygen (DO) that is consumed in the sample under aerobic conditions during the incubation period (five days at 20 C) after seeding material has been added. A polarographic membrane electrode that has been calibrated in saturated air is used to acquire all dissolved oxygen measurements. No nitrogen inhibitor added. A modified Winkler procedure may be substituted for the DO probe procedure. Performance Base Method published circa Dec. 2000. | 10 | Y | Y | Y | 359 | 3/10/2010 12:42:00 PM | 3/24/2021 | 3/24/2021 5:01:11 PM | 2010-04-01 |  | RDGIBSON |  | Concentration | mg/L | Y | (fl. conc.) |  | Biochemical Oxygen Demand - Soluble (fl. conc.) |  | Aggregate Organics; | Aggregate Organics; EMS code: BODS | Chemical | Soluble | LAB
BOSC | mg/L | Boscalid | E751 | E751D Pesticides in Water by LC-MS/MS (POS1 List) | E751D Pesticides in Water by LC-MS/MS (POS1 List) (MOECC E3501) | 5e-05 | Y | Y | Y | 39 | 12/9/2021 | 12/31/2021 | 2/15/2022 9:50:05 AM | 2021-12-14 |  | EMS |  | Concentration | mg/L | Y | (fl. conc.) |  | Boscalid (fl. conc.) | 188425-85-6 | Pesticides; | Pesticides; EMS code: BOSC | Chemical |  | LAB
BPA | ug/g | Bisphenol A (CAS: 80-05-7) | A849 | J. of Chrom. A849 (1999) 467-482 | Sample is derivatized and injected directly. Each compound is separated by reversed phase HPLC and analyzed by Electro Spray Ionization MS/MS detection using Triple Quadrupole MS/MS detector. Internal standards are used for quantitation. | 0.05 | Y | Y | Y | 8 | 3/14/2020 | 3/16/2020 | 2/21/2023 2:10:31 PM | 2022-08-08 |  | EMS |  | MassPerMass | mg/kg | Y | (mass conc.) |  | Bisphenol A (mass conc.) | 80-05-7 | Pharmaceuticals and Personal Care Products (PPCP); | Pharmaceuticals and Personal Care Products (PPCP); EMS code: BPA | Chemical |  | LAB
BPTP | mg/L | B(a)P Total Potency Equivalent | HEGC | Hexane extraction: GC/MS | Hexane extraction: GC/MS | 1e-05 | Y | Y | Y | 13 | 6/5/2023 | 7/28/2024 | 8/6/2024 3:27:04 PM | 2023-12-12 |  | EMS |  | Concentration | mg/L | Y | (fl. conc.) |  | B(a)P Total Potency Equivalent (fl. conc.) |  | Toxicity; | Toxicity; EMS code: BPTP | Chemical | Total | LAB
BPTP | ug/g | B(a)P Total Potency Equivalent | FS87 | DCM,Ac/Hex extr,PAM,HRGC/LRMS | DCM;Ac/Hex extr;PAM;HRGC/LRMS | 0.02 | Y | Y | Y | 78 | 9/20/2017 | 9/9/2022 | 1/11/2024 1:07:07 PM | 2018-01-09 |  | JKROGH |  | MassPerMass | mg/kg | Y | (mass conc.) |  | B(a)P Total Potency Equivalent (mass conc.) |  | Toxicity; | Toxicity; EMS code: BPTP | Chemical | Total | LAB
BR-E | mg/L | Bromide Extractable | BRAW | CRC ICPMS - run RAW for Br and I, unfiltered, unpreserved | This method is applicable to water samples as well as digests of solid samples with the exception of Br and I which is only applicable to un-digested un-preserved water samples. Water samples are analyzed for dissolved metals following sample filtration or for total metals following sample preservation or digestion with the exception of water samples for Bromine and Iodine which does not require any sample preparation and only cold preservation. This method describes the multi-elemental determination of trace elements by Collision/Reaction Cell (CRC) ICP-MS. Ions produced by radio-frequency inductively coupled plasma are measured. Analyte species originating in a liquid are pneumatically nebulized and the resulting aerosol transported by argon gas into the plasma torch. The ions produced are entrained in the plasma gas and extracted, by means of a differentially pumped vacuum interface, into a mass spectrometer. The ions flow through a Collision/Reaction Cell (CRC) where, depending on the mode of analysis, they either pass through the cell unhindered (no gas - classical ICP-MS mode) or collision/reaction chemistry takes place between the ions and an introduced gas such as hydrogen or helium. The ions are then sorted according to their mass-to-charge ratios by a quadrupole mass spectrometer. The ions transmitted through the quadrupole are quantified by a channel electron multiplier and the ion information is processed by a data handling system. Although most polyatomic interferences are removed, any additional interferences must be assessed and valid corrections applied or the data flagged to indicate problems. Interference correction must include compensation for background ions contributed by the plasma gas, reagents, and constituents of the sample matrix. Instrument drift as well as suppression or enhancement of instrument response caused by the sample matrix must be corrected for by the use of internal standards. | 0.01 | Y | Y | Y | 4053 | 3/16/2012 3:11:00 PM | 6/26/2024 12:29:00 PM | 7/9/2024 3:31:02 PM | 2012-04-13 |  | RDGIBSON |  | Concentration | mg/L | Y | (fl. conc.) |  | Bromide Extractable (fl. conc.) | 24959-67-9 | Metals - Extractable; | Metals - Extractable; EMS code: BR-E | Chemical | Extractable | LAB
BR-L | mg/L | Bromide Leachable | MEND | MEND (mod) EPA 300.1 (mod) | MEND (mod) EPA 300.1 (mod). Inorganic anions are analyzed by obtaining an extract from a soil sample that has been added in a 3:1 ratio of deionized water to soil; then leached and the extract filtered prior to analysis; which is then analyzed by Ion Chromatography with conductivity and/or UV detection. | 0.05 | Y | Y | Y | 0 |  |  |  | 2024-01-24 |  | RDGIBSON |  | Concentration | mg/L | Y | (fl. conc.) |  | Bromide Leachable (fl. conc.) |  |  | EMS code: BR-L | Chemical |  | LAB
BROD | mg/L | Brodifacoum | P015 | Phenoxyacid Herbicides by LC-MS/MS | Phenoxyacid Herbicides by LC-MS/MS | 1e-05 | Y | Y | Y | 42 | 12/8/2021 | 9/9/2022 | 12/8/2022 12:57:05 PM | 2021-12-09 | 12/9/2021 3:22:12 PM | EMS | EMS | Concentration | mg/L | Y | (fl. conc.) |  | Brodifacoum (fl. conc.) | 56073-10-0 | Pesticides; | Pesticides; EMS code: BROD | Chemical |  | LAB
BSRB | CFU/mL | Sulfate Reducing Bactera | BART | BIOLOGICAL ACTIVITY REACTION TEST | METHOD FOR THE DETERMINATION OF BACTERIAL POPULATIONS BY THE BIOLOGICAL ACTIVITY REACTION TEST? APPLICATION: This method determines the activity of bacterial populations in water, including drinking water, surface water (ponds, lakes, creeks etc.), raw water, ground water, wastewater, swimming pools and whirlpools. PRINCIPLE: Aliquots of water samples are inoculated in Biological Activity Reaction Test (BART?) vials specific to the desired organism type and incubated for eight days at room temperature. The reaction pattern is observed each day of the incubation period to determine the community activity and characteristics. Each vial type contains selective culture media specific to enriching and isolating the desired microbial community and therefore a different reaction pattern. The concentration, as colony forming unit per millilitre (cfu/mL), of the bacterial population can be extrapolated from the activity level based on the time delay for when a positive reaction is first observed in the reaction vial. METHOD SOURCES: The CARO-EDM method is performed without significant modification according to: 1. Droycon Bioconcepts Inc., Biological Activity Reaction Test BART? User Manual, 2004 Edition Supporting Documentation: CARO-EDM Microbiology Procedures Manual, CE-PM-009-current, BART-SOFT Version 6. Equipment Manuals: Autoclave ? Sanyo, Model 2400, or equivalent. DETECTION LIMITS: Reporting Detection Limits (RDL) are dictated by the method and are based on an 8 day incubation period and may be lowered if the incubation period is extended. Table 2 lists the RDLs of the BART tests commonly performed by CARO-EDM. These detection limits are derived from the BART-SOFT version 6 population calculator. Refer to the CARO LIMS database for RDLs for those parameters not listed here. LAB-BART? Type RDL (cfu/ml)* IRB=34, SRB=31, HAB=65, DN=18. | 8 | Y | Y | Y | 16 | 7/22/2013 4:27:00 PM | 2/24/2021 3:45:00 PM | 3/5/2021 5:01:06 PM | 2014-12-15 |  | RDGIBSON |  | MicrobialFormation | CFU/mL | N | (microb.) |  | Sulfate Reducing Bactera (microb.) |  | Bacteriological; | Bacteriological; EMS code: BSRB | Chemical |  | LAB
BTEX | mg/L | BTEX Total | XCAL | Calculation (mg/L) | Calculation (mg/L) | 0.0011 | Y | Y | Y | 64 | 2/14/2013 | 10/13/2019 | 10/15/2019 5:53:11 PM | 2013-02-21 |  | RDGIBSON |  | Concentration | mg/L | Y | (fl. conc.) |  | BTEX Total (fl. conc.) |  | Volatile Organic Compounds (VOCs); | Volatile Organic Compounds (VOCs); EMS code: BTEX | Chemical | Total | LAB
BTEX | mg/L | BTEX Total | 8260D | EPA 8260D (mod) | EPA 8260D (mod) VOLATILE ORGANIC COMPOUNDS BY GAS CHROMATOGRAPHY/MASS SPECTROMETRY This method is used to determine volatile organic compounds (VOCs) in a variety of solid waste matrices. | 0.012 | Y | Y | Y | 0 |  |  |  | 2024-01-24 |  | RDGIBSON |  | Concentration | mg/L | Y | (fl. conc.) |  | BTEX Total (fl. conc.) |  | Volatile Organic Compounds (VOCs); | Volatile Organic Compounds (VOCs); EMS code: BTEX | Chemical | Total | LAB
BTEX | mg/L | BTEX Total | F007 | Static Headspace:HRGC/LRMS | Static headspace analysis is a method for the determination of volatile compounds in liquids and/or solids by analysing the vapour phase of a sample in a closed system; in this case a sealed vial. Static Headspace is sometimes referred to as Equilibrium Headspace because the vapour phase above the sample in a sealed vial is brought into thermodynamic equilibrium with the sample prior to analysis. Volatile organic compounds are thermally extracted from the sample matrix by Static Headspace sampling and introduced directly to a thick-film megabore GC column. Compounds eluting from the GC column are introduced into a Mass Spectrometer and target compounds are identified by a comparison of both retention times and mass spectra with those of reference standards stored in a computerized data system. Reference retention times and mass spectra for analytes are obtained by the measurement of calibration standards under precisely the same extraction conditions used for samples. The concentration of each identified component is measured by relating the response of the sample quantitation mass ions to the response of the quantitation mass ions of a series of external standards. A deuterated surrogate recovery standard is added to provide information on matrix effects. | 0.001 | Y | Y | Y | 69 | 12/3/2021 | 5/16/2024 | 6/6/2024 10:23:03 AM | 2021-12-14 |  | EMS |  | Concentration | mg/L | Y | (fl. conc.) |  | BTEX Total (fl. conc.) |  | Volatile Organic Compounds (VOCs); | Volatile Organic Compounds (VOCs); EMS code: BTEX | Chemical | Total | LAB
BTEXS | mg/L | BTEX+Styrene Total | F007 | Static Headspace:HRGC/LRMS | Static headspace analysis is a method for the determination of volatile compounds in liquids and/or solids by analysing the vapour phase of a sample in a closed system; in this case a sealed vial. Static Headspace is sometimes referred to as Equilibrium Headspace because the vapour phase above the sample in a sealed vial is brought into thermodynamic equilibrium with the sample prior to analysis. Volatile organic compounds are thermally extracted from the sample matrix by Static Headspace sampling and introduced directly to a thick-film megabore GC column. Compounds eluting from the GC column are introduced into a Mass Spectrometer and target compounds are identified by a comparison of both retention times and mass spectra with those of reference standards stored in a computerized data system. Reference retention times and mass spectra for analytes are obtained by the measurement of calibration standards under precisely the same extraction conditions used for samples. The concentration of each identified component is measured by relating the response of the sample quantitation mass ions to the response of the quantitation mass ions of a series of external standards. A deuterated surrogate recovery standard is added to provide information on matrix effects. |  | Y | Y | Y | 7 | 7/25/2023 | 5/16/2024 | 6/6/2024 10:23:03 AM | 2022-04-06 |  | EMS |  | Concentration | mg/L | Y | (fl. conc.) |  | BTEX+Styrene Total (fl. conc.) |  |  | EMS code: BTEXS | Chemical | Total | LAB
BUAC | ug/g | Butachlor | ZN03 | DCM+Ac.Extr:GC/MS | DCM+Ac.Extr:GC/MS EPA 8270E (mod) | 0.005 | Y | Y | Y | 2 | 11/2/2022 12:53:00 AM | 11/2/2022 4:41:00 AM | 1/7/2023 3:12:10 PM | 2022-12-29 |  | EMS |  | MassPerMass | mg/kg | Y | (mass conc.) |  | Butachlor (mass conc.) | 23184-66-9 |  | EMS code: BUAC | Chemical |  | LAB
BUOH | mg/L | 1-Butanol CAS: 71-36-3 | E-8015 | EPA 8015D | EPA Method 8015D (SW-846): Nonhalogenated Organics Using GC/FID |  | Y | Y | Y | 14 | 12/29/2021 | 1/12/2022 | 9/2/2022 6:21:52 PM | 2022-04-04 |  | EMS |  | Concentration | mg/L | Y | (fl. conc.) |  | 1-Butanol (fl. conc.) | 71-36-3 | Organics; | Organics; EMS code: BUOH | Chemical |  | LAB
BUOH | ug/g | 1-Butanol CAS: 71-36-3 | E8015S | EPA 8015 D in soil | See EPA 8015D | 10 | Y | Y | Y | 15 | 1/12/2022 | 2/1/2022 | 1/11/2024 1:07:07 PM | 2022-06-23 |  | EMS |  | MassPerMass | mg/kg | Y | (mass conc.) |  | 1-Butanol (mass conc.) | 71-36-3 | Organics; | Organics; EMS code: BUOH | Chemical |  | LAB
BUTA | mg/L | Butyric Acid | VFAD | Volatile Fatty Acids by GC/FID | Volatile Fatty Acids by GC/FID; Standard Methods for the Examination of Water and Wastewater; 21st Edition, Method 5560 D; pages 5-59 to 5-61. 5560 D. Gas Chromatographic Method Principle: This gas chromatographic procedure may be usedto determine the individual concentrations of the following fattyacids: acetic; propionic; butyric; isobutyric; valeric; andisovaleric acids in digester sludge samples. Caproic and heptanoicacids also may be determined; but with less certaintybecause of their reduced solubility in water. Volatile fatty acidsare generally soluble in digester sludge and exist in either theprotonated or unprotonated form; depending on the pH of thesludge sample. Acetic and propionic acids usually are in higherconcentrations than the other fatty acids in digester sludge. Allacids must be converted to the protonated (volatile) form beforeinjection into the gas chromatograph (GC) so vaporization mayoccur. The fatty acids must also be separated from the solidsmaterial in the sample before introduction into the GC to minimizedegradation of the GC column. Digester sludge samples areprepared by acidification; centrifugation; and filtration. The sampleis analyzed by direct injection into a gas chromatographequipped with a flame ionization detector after both centrifugationand filtration.b. Interference: The blank amount may be high relative tosample analyte. High blanks may be caused by a buildup ofcontaminants in the injector and guard column or by samplecarry-over. The analyst must be aware of the blank levels andcorrect unacceptably high blank levels that are above the reportinglimits.c. Minimum detectable concentration: The method detectionlevel has been determined by the USEPA method1 in reagentwater to be about 3 mg/L for acetic acid and about 1 mg/L for allother target compounds. | 2 | Y | Y | Y | 5 | 3/14/2017 10:30:00 AM | 11/19/2017 5:10:00 PM | 6/20/2018 9:02:11 AM | 2018-06-19 |  | RDGIBSON |  | Concentration | mg/L | Y | (fl. conc.) |  | Butyric Acid (fl. conc.) | 107-92-6 | Organics; | Organics; EMS code: BUTA | Chemical |  | LAB
BZF | mg/L | Bezafibrate CAS: 41859-67-0 | XMSP | HPLC and Tri. Quadpole MS/MS | EPA 1694 An aliquot of filtered sample containing 5% organic is injected directly. Each compound is separated by reversed phase HPLC. Analyzed by Electro Spray Ionization MS/MS detection using Triple Quadrupole MS/MS detector. Quantiitation with internal or external standard. |  | Y | Y | Y | 12 | 5/8/2020 | 6/4/2024 | 8/21/2024 3:32:03 PM | 2022-05-04 |  | EMS |  | Concentration | mg/L | Y | (fl. conc.) |  | Bezafibrate (fl. conc.) | 41859-67-0 | Organics; | Organics; EMS code: BZF | Chemical |  | LAB
C--T | mg/L | Carbon Total | 0910 | Beckman Comb/IR Det | Beckman Comb/IR Det | 3 | Y | Y | Y | 1952 |  |  | 2/7/2000 3:11:47 PM | 1996-04-12 | 6/3/1998 6:51:34 PM | EMS |  | Concentration | mg/L | Y | (fl. conc.) |  | Carbon Total (fl. conc.) |  | Carbon; | Carbon; EMS code: C--T | Chemical | Total | LAB
C--T | mg/L | Carbon Total | X171 | Carbon Anal.,Infra Red | Analysis is comprised of 2 phases, the determination of Total Carbon (TC) or Dissolved Carbon (DC) and the determination of Total Inorganic Carbon (TIC) or Dissolved Inorganic Carbon (DIC). Total Organic Carbon (TOC) content, or Dissolved Organic Carbon (DOC) is then calculated by difference. A series of External Standard solutions, Potassium Biphthalate for TC (DC) and Sodium Carbonate/Sodium Bicarbonate for TIC (DIC) is used for instrument calibration. During TC (DC) analysis, a small portion of the sample (or standard) solution is introduced into a combustion tube, containing a Platinum catalyst at an elevated temperature with an Oxygen gas carrier, via an automated autosampler/injector. The sample's TC (DC) component is combusted/oxidized to form CO2. The CO2 flows through an IC reaction vessel, is cooled and dried by an electronic dehumidifier and passes through a halogen scrubber into a Non-dispersive Infrared Detector (NDIR). The NDIR produces an analog signal over a constant time base that a data processor converts into a peak, the area of which is proportional to the TC concentration of the sample. The TC (DC) concentration of a sample is determined electronically from a calibration curve expressing the relationship between the peak area of a standard and its concentration. Similarly, for the TIC (DIC) phase of the analysis, another small portion of sample (or standard) solution is introduced directly into the IC reaction vessel, containing a Phosphoric Acid solution through which carrier gas is bubbled. However, here only the IC component of the sample reacts to form CO2. The TOC (DOC) concentration is automatically determined by the data processor from the difference between TC (DC) and TIC (DOC) concentrations. | 0.5 | Y | Y | Y | 4 | 3/13/2000 8:28:00 AM | 5/19/2000 4:35:00 PM | 8/11/2000 2:17:34 AM | 1996-04-12 | 6/3/1998 6:51:34 PM | EMS |  | Concentration | mg/L | Y | (fl. conc.) |  | Carbon Total (fl. conc.) |  | Carbon; | Carbon; EMS code: C--T | Chemical | Total | LAB
C--T | % | Carbon Total | IR07 | ALS Minerals C-IR07 | ALS Minerals C-IR07. Sample is heated to 1350ºC in an induction furnace with oxygen stream. Carbon dioxide released from the sample are measured by an IR detection system (Leco analyzer). | 0.01 | Y | Y | Y | 0 |  |  |  | 2024-01-29 |  | RDGIBSON |  | DimensionlessRatio | % | N | (ratio) |  | Carbon Total (ratio) |  | Carbon; | Carbon; EMS code: C--T | Chemical | Total | LAB
C--T | % | Carbon Total | FS93 | Oxid to CO2,NDIR,Calc'n | Total Carbon is determined by high temperature oxidation of carbon to carbon diooxide which is then measured by means of a nondispersive infrared analyzer. Inorganic Carbon is determined by reaction with phosphoric acid to convert all carbonates to carbon dioxide which is also measured by means of a nondispersive infrared analyzer. Total Organic Carbon is determined as the difference between Total and Inorganic Carbons. | 0.05 | Y | Y | Y | 31 | 1/13/2022 | 2/20/2024 | 2/26/2024 6:09:04 PM | 2001-05-22 |  | NPEPPIN |  | DimensionlessRatio | % | N | (ratio) |  | Carbon Total (ratio) |  | Carbon; | Carbon; EMS code: C--T | Chemical | Total | LAB
C--T | mg/L | Carbon Total | CS01 | System Calculation | System Calculation (mg/L) |  | N | Y | N | 11041 |  |  | 7/24/2024 7:36:07 PM | 1996-04-12 | 6/3/1998 6:51:43 PM | EMS |  | Concentration | mg/L | Y | (fl. conc.) |  | Carbon Total (fl. conc.) |  | Carbon; | Carbon; EMS code: C--T | Chemical | Total | LAB
C--T | % | Carbon Total | IR08 | ALS Minerals ME-IR08 | ALS Minerals ME-IR08. A prepared sample is heated to approximately 1350°C in an induction furnace with oxygen stream. Sulfur and carbon dioxide released from the sample are measured by IR (Leco analyzer) | 0.01 | Y | Y | Y | 0 |  |  |  | 2024-01-29 |  | RDGIBSON |  | DimensionlessRatio | % | N | (ratio) |  | Carbon Total (ratio) |  | Carbon; | Carbon; EMS code: C--T | Chemical | Total | LAB
C--T | mg/L | Carbon Total | 0920 | Dohrmann Comb/FID | Dohrmann Comb/FID | 1 | Y | Y | Y | 1 |  |  | 10/23/1996 11:16:26 PM | 1996-04-12 | 6/3/1998 6:51:34 PM | EMS |  | Concentration | mg/L | Y | (fl. conc.) |  | Carbon Total (fl. conc.) |  | Carbon; | Carbon; EMS code: C--T | Chemical | Total | LAB
C--T | ug/g | Carbon Total | CS02 | System Calculation | System Calculation (ug/g) |  | N | Y | N | 406 |  |  | 8/13/2016 5:13:53 PM | 1996-04-12 | 6/3/1998 6:51:43 PM | EMS |  | MassPerMass | mg/kg | Y | (mass conc.) |  | Carbon Total (mass conc.) |  | Carbon; | Carbon; EMS code: C--T | Chemical | Total | LAB
C--T | ug/g | Carbon Total | EQ02 | EQUIS (ug/g) | EQUIS (ug/g) | 100 | N | Y | Y | 198 |  |  | 10/23/1996 12:08:28 AM | 1996-04-12 | 6/3/1998 6:51:34 PM | EMS |  | MassPerMass | mg/kg | Y | (mass conc.) |  | Carbon Total (mass conc.) |  | Carbon; | Carbon; EMS code: C--T | Chemical | Total | LAB
C--T | ug/g | Carbon Total | 0790 | D/G:Leco Carbon | D/G:Leco Carbon | 1000 | Y | Y | Y | 856 | 4/28/2000 7:31:00 PM | 7/5/2001 7:30:00 AM | 12/14/2001 1:06:36 AM | 1996-04-12 | 6/3/1998 6:51:34 PM | EMS |  | MassPerMass | mg/kg | Y | (mass conc.) |  | Carbon Total (mass conc.) |  | Carbon; | Carbon; EMS code: C--T | Chemical | Total | LAB
C001 | mg/L | Carbaryl | P045 | DCM Extr:HPLC | DCM Extr:HPLC | 0.0005 | Y | Y | Y | 1 | 7/29/2005 7:10:00 AM | 7/29/2005 7:10:00 AM | 6/2/2009 12:00:59 PM | 1996-04-12 | 6/3/1998 6:51:42 PM | EMS |  | Concentration | mg/L | Y | (fl. conc.) |  | Carbaryl (fl. conc.) | 63-25-2 | Pesticides; | Pesticides; EMS code: C001 | Chemical |  | LAB
C001 | ug/g | Carbaryl | ZN03 | DCM+Ac.Extr:GC/MS | DCM+Ac.Extr:GC/MS EPA 8270E (mod) | 0.05 | Y | Y | Y | 14 | 1/25/2022 | 9/8/2022 | 12/15/2022 10:50:07 AM | 2003-06-16 |  | UMULLER |  | MassPerMass | mg/kg | Y | (mass conc.) |  | Carbaryl (mass conc.) | 63-25-2 | Pesticides; | Pesticides; EMS code: C001 | Chemical |  | LAB
C001 | ug/g | Carbaryl | FS61 | DCM +MeOH extr,HRGC-NPD | Sediment samples are shaken with a mixture of Dichloromethane and Methanol, filtered then re-extracted with Dichloromethane. The extract is concentrated and analysed using High Resolution Gas Chromatography with a Thermionic Specific Detector (TSD), alternately referred to as a Nitrogen Phosphorus Detector (NPD). | 0.01 | Y | Y | Y | 34 | 4/20/2002 3:40:00 PM | 7/16/2002 3:20:00 PM | 2/26/2004 1:05:26 AM | 1998-10-02 |  | NPEPPIN |  | MassPerMass | mg/kg | Y | (mass conc.) |  | Carbaryl (mass conc.) | 63-25-2 | Pesticides; | Pesticides; EMS code: C001 | Chemical |  | LAB
C001 | mg/L | Carbaryl | E751 | E751D Pesticides in Water by LC-MS/MS (POS1 List) | E751D Pesticides in Water by LC-MS/MS (POS1 List) (MOECC E3501) | 5e-05 | Y | Y | Y | 39 | 12/9/2021 | 12/31/2021 | 2/15/2022 9:50:05 AM | 2021-12-14 |  | EMS |  | Concentration | mg/L | Y | (fl. conc.) |  | Carbaryl (fl. conc.) | 63-25-2 | Pesticides; | Pesticides; EMS code: C001 | Chemical |  | LAB
C001 | mg/L | Carbaryl | E660 | EPA 8270E (mod) | EPA 8270E (mod) Miscellaneous Pesticides by GC-MS | 0.0002 | Y | Y | Y | 68 | 12/7/2021 | 9/12/2022 | 12/8/2022 12:57:05 PM | 2021-12-09 | 12/9/2021 3:22:12 PM | EMS | EMS | Concentration | mg/L | Y | (fl. conc.) |  | Carbaryl (fl. conc.) | 63-25-2 | Pesticides; | Pesticides; EMS code: C001 | Chemical |  | LAB
C001 | mg/L | Carbaryl | F056 | DCM extr,HRGC-NPD | Effluent samples are extracted with Dichloromethane. The extract is concentrated and analysed using High Resolution Gas Chromatography with a Thermionic Specific Detector (TSD), alternately referred to as a Nitrogen Phosphorus Detector (NPD). | 0.005 | Y | Y | Y | 12 | 4/3/2000 10:51:00 AM | 2/16/2001 3:36:00 PM | 9/11/2002 9:34:56 AM | 1998-10-02 | 10/2/1998 10:50:33 AM | NPEPPIN | NPEPPIN | Concentration | mg/L | Y | (fl. conc.) |  | Carbaryl (fl. conc.) | 63-25-2 | Pesticides; | Pesticides; EMS code: C001 | Chemical |  | LAB
C002 | mg/L | Chlorbenside | ZN01 | DCM+Ac.Extr:GC/MS | DCM+Ac.Extr:GC/MS | 0.0001 | Y | Y | Y | 28 | 7/28/2005 7:03:00 AM | 3/29/2012 10:15:00 PM | 5/3/2012 9:27:49 AM | 1996-04-12 | 11/16/1999 4:06:27 PM | EMS | NPEPPIN | Concentration | mg/L | Y | (fl. conc.) |  | Chlorbenside (fl. conc.) |  | Pesticides; | Pesticides; EMS code: C002 | Chemical |  | LAB
C003 | ug/g | Chloroethane | MS02 | Purge&Trap GC/MS (VOC-S) | Purge&Trap GC/MS (VOC-S) | 0.012 | Y | Y | Y | 26 | 11/19/2004 | 2/4/2022 | 1/11/2024 1:07:07 PM | 1996-04-12 | 6/3/1998 6:51:45 PM | EMS |  | MassPerMass | mg/kg | Y | (mass conc.) |  | Chloroethane (mass conc.) | 75-00-3 | Volatile Organic Compounds (VOCs); | Volatile Organic Compounds (VOCs); EMS code: C003 | Chemical |  | LAB
C003 | ug/g | Chloroethane | FS04 | Static Headspace:HRGC/LRMS | Static Headspace anaylsis is a method for the determination of volatile compounds in liquids and/or solids by analysing the vapour phase of a sample in a closed system; in this case a sealed vial. Static Headspace is sometimes referred to as Equilibrium Headspace because the vapour phase above the sample in a sealed vial is brought into thermodynamic equilibrium with the sample prior to analysis. Volatile organic compounds are thermally extracted from the sample matrix by Static Headspace sampling and introduced directly to a thick-film megabore GC column. Compounds eluting from the GC column are introduced into a Mass Spectrometer and target compounds are identified by a comparison of both retention times and mass spectra with those of reference standards stored in a computerized data system. Reference retention times and mass spectra for analytes are obtained by the measurement of calibration standards under precisely the same extraction conditions used for samples. The concentration of each identified component is measured by relating the response of the sample quantitation mass ions to the response of the quantitation mass ions of a series of external standards. A deuterated surrogate recovery standard is added to provide information on matrix effects. | 0.1 | Y | Y | Y | 43 | 5/21/2015 | 5/22/2015 | 7/3/2015 3:19:03 PM | 2015-07-03 |  | RDGIBSON |  | MassPerMass | mg/kg | Y | (mass conc.) |  | Chloroethane (mass conc.) | 75-00-3 | Volatile Organic Compounds (VOCs); | Volatile Organic Compounds (VOCs); EMS code: C003 | Chemical |  | LAB
C003 | mg/L | Chloroethane | F007 | Static Headspace:HRGC/LRMS | Static headspace analysis is a method for the determination of volatile compounds in liquids and/or solids by analysing the vapour phase of a sample in a closed system; in this case a sealed vial. Static Headspace is sometimes referred to as Equilibrium Headspace because the vapour phase above the sample in a sealed vial is brought into thermodynamic equilibrium with the sample prior to analysis. Volatile organic compounds are thermally extracted from the sample matrix by Static Headspace sampling and introduced directly to a thick-film megabore GC column. Compounds eluting from the GC column are introduced into a Mass Spectrometer and target compounds are identified by a comparison of both retention times and mass spectra with those of reference standards stored in a computerized data system. Reference retention times and mass spectra for analytes are obtained by the measurement of calibration standards under precisely the same extraction conditions used for samples. The concentration of each identified component is measured by relating the response of the sample quantitation mass ions to the response of the quantitation mass ions of a series of external standards. A deuterated surrogate recovery standard is added to provide information on matrix effects. | 0.001 | Y | Y | Y | 329 | 4/15/2011 | 8/14/2024 | 8/16/2024 4:05:21 PM | 2012-03-28 |  | RDGIBSON |  | Concentration | mg/L | Y | (fl. conc.) |  | Chloroethane (fl. conc.) | 75-00-3 | Volatile Organic Compounds (VOCs); | Volatile Organic Compounds (VOCs); EMS code: C003 | Chemical |  | LAB
C003 | mg/L | Chloroethane | MS04 | Static Headspace GC/MS (Water) | (Water) A static headspace procedure is used as a means of driving the analytes out of the matrix into a headspace. Once equilibrium is achieved, a fixed volume of headspeace is taken for subsequent determination by Gas Chromatography - Mass Spectroscopy (GC/MS) using Selected Ion Monitoring. | 0.001 | Y | Y | Y | 70 | 8/13/2009 | 6/28/2024 12:29:00 AM | 7/9/2024 3:31:02 PM | 2009-12-23 |  | RDGIBSON |  | Concentration | mg/L | Y | (fl. conc.) |  | Chloroethane (fl. conc.) | 75-00-3 | Volatile Organic Compounds (VOCs); | Volatile Organic Compounds (VOCs); EMS code: C003 | Chemical |  | LAB
C003 | mg/L | Chloroethane | X384 | Purge and trap GC/MS | The sample is analyzed by purge and trap gas chromatography with detection by mass spectrometry. The procedure involves purging the volatile compounds from the sample with an inert gas, and trapping them on a solid sorbent. When purging of the sample is complete, the trap is heated and the compounds are transferred to the gas chromatographic column. Analysis is then accomplished by separation of the components by gas chromatography with detection by mass spectrometry. | 0.001 | Y | Y | Y | 114 | 4/30/2002 | 7/26/2012 7:12:00 AM | 9/26/2012 1:07:08 PM | 1996-04-12 | 6/3/1998 6:51:34 PM | EMS |  | Concentration | mg/L | Y | (fl. conc.) |  | Chloroethane (fl. conc.) | 75-00-3 | Volatile Organic Compounds (VOCs); | Volatile Organic Compounds (VOCs); EMS code: C003 | Chemical |  | LAB
C003 | mg/L | Chloroethane | MS01 | Purge&Trap GC/MS (VOC-W) | Purge&Trap GC/MS (VOC-W) | 2.3 | Y | Y | Y | 431 | 4/15/2002 | 6/4/2024 10:24:00 AM | 6/13/2024 5:01:27 PM | 1996-04-12 | 3/18/2003 2:49:40 PM | EMS | UMULLER | Concentration | mg/L | Y | (fl. conc.) |  | Chloroethane (fl. conc.) | 75-00-3 | Volatile Organic Compounds (VOCs); | Volatile Organic Compounds (VOCs); EMS code: C003 | Chemical |  | LAB
C004 | ug/g | Vinyl Chloride | MS02 | Purge&Trap GC/MS (VOC-S) | Purge&Trap GC/MS (VOC-S) | 0.015 | Y | Y | Y | 27 | 11/19/2004 | 10/28/2022 5:04:00 PM | 1/11/2024 1:07:07 PM | 1996-04-12 | 6/3/1998 6:51:34 PM | EMS |  | MassPerMass | mg/kg | Y | (mass conc.) |  | Vinyl Chloride (mass conc.) | 75-01-4 | Volatile Organic Compounds (VOCs); | Volatile Organic Compounds (VOCs); EMS code: C004 | Chemical |  | LAB
C004 | mg/L | Vinyl Chloride | MS01 | Purge&Trap GC/MS (VOC-W) | Purge&Trap GC/MS (VOC-W) | 2.9 | Y | Y | Y | 544 | 4/15/2002 | 6/4/2024 10:24:00 AM | 6/13/2024 5:01:27 PM | 1996-04-12 | 3/18/2003 2:49:40 PM | EMS | UMULLER | Concentration | mg/L | Y | (fl. conc.) |  | Vinyl Chloride (fl. conc.) | 75-01-4 | Volatile Organic Compounds (VOCs); | Volatile Organic Compounds (VOCs); EMS code: C004 | Chemical |  | LAB
C004 | mg/L | Vinyl Chloride | MS04 | Static Headspace GC/MS (Water) | (Water) A static headspace procedure is used as a means of driving the analytes out of the matrix into a headspace. Once equilibrium is achieved, a fixed volume of headspeace is taken for subsequent determination by Gas Chromatography - Mass Spectroscopy (GC/MS) using Selected Ion Monitoring. | 0.0005 | Y | Y | Y | 70 | 8/13/2009 | 6/28/2024 12:29:00 AM | 7/9/2024 3:31:02 PM | 2009-12-23 |  | RDGIBSON |  | Concentration | mg/L | Y | (fl. conc.) |  | Vinyl Chloride (fl. conc.) | 75-01-4 | Volatile Organic Compounds (VOCs); | Volatile Organic Compounds (VOCs); EMS code: C004 | Chemical |  | LAB
C004 | ug/g | Vinyl Chloride | FS04 | Static Headspace:HRGC/LRMS | Static Headspace anaylsis is a method for the determination of volatile compounds in liquids and/or solids by analysing the vapour phase of a sample in a closed system; in this case a sealed vial. Static Headspace is sometimes referred to as Equilibrium Headspace because the vapour phase above the sample in a sealed vial is brought into thermodynamic equilibrium with the sample prior to analysis. Volatile organic compounds are thermally extracted from the sample matrix by Static Headspace sampling and introduced directly to a thick-film megabore GC column. Compounds eluting from the GC column are introduced into a Mass Spectrometer and target compounds are identified by a comparison of both retention times and mass spectra with those of reference standards stored in a computerized data system. Reference retention times and mass spectra for analytes are obtained by the measurement of calibration standards under precisely the same extraction conditions used for samples. The concentration of each identified component is measured by relating the response of the sample quantitation mass ions to the response of the quantitation mass ions of a series of external standards. A deuterated surrogate recovery standard is added to provide information on matrix effects. | 0.1 | Y | Y | Y | 43 | 5/21/2015 | 5/22/2015 | 7/3/2015 3:19:03 PM | 2015-07-03 |  | RDGIBSON |  | MassPerMass | mg/kg | Y | (mass conc.) |  | Vinyl Chloride (mass conc.) | 75-01-4 | Volatile Organic Compounds (VOCs); | Volatile Organic Compounds (VOCs); EMS code: C004 | Chemical |  | LAB
C004 | mg/L | Vinyl Chloride | F007 | Static Headspace:HRGC/LRMS | Static headspace analysis is a method for the determination of volatile compounds in liquids and/or solids by analysing the vapour phase of a sample in a closed system; in this case a sealed vial. Static Headspace is sometimes referred to as Equilibrium Headspace because the vapour phase above the sample in a sealed vial is brought into thermodynamic equilibrium with the sample prior to analysis. Volatile organic compounds are thermally extracted from the sample matrix by Static Headspace sampling and introduced directly to a thick-film megabore GC column. Compounds eluting from the GC column are introduced into a Mass Spectrometer and target compounds are identified by a comparison of both retention times and mass spectra with those of reference standards stored in a computerized data system. Reference retention times and mass spectra for analytes are obtained by the measurement of calibration standards under precisely the same extraction conditions used for samples. The concentration of each identified component is measured by relating the response of the sample quantitation mass ions to the response of the quantitation mass ions of a series of external standards. A deuterated surrogate recovery standard is added to provide information on matrix effects. | 0.001 | Y | Y | Y | 329 | 4/15/2011 | 8/14/2024 | 8/16/2024 4:05:21 PM | 2012-03-28 |  | RDGIBSON |  | Concentration | mg/L | Y | (fl. conc.) |  | Vinyl Chloride (fl. conc.) | 75-01-4 | Volatile Organic Compounds (VOCs); | Volatile Organic Compounds (VOCs); EMS code: C004 | Chemical |  | LAB
C004 | mg/L | Vinyl Chloride | X384 | Purge and trap GC/MS | The sample is analyzed by purge and trap gas chromatography with detection by mass spectrometry. The procedure involves purging the volatile compounds from the sample with an inert gas, and trapping them on a solid sorbent. When purging of the sample is complete, the trap is heated and the compounds are transferred to the gas chromatographic column. Analysis is then accomplished by separation of the components by gas chromatography with detection by mass spectrometry. | 0.001 | Y | Y | Y | 119 | 4/27/2000 | 7/26/2012 7:12:00 AM | 9/26/2012 1:07:08 PM | 1996-04-12 | 6/3/1998 6:51:34 PM | EMS |  | Concentration | mg/L | Y | (fl. conc.) |  | Vinyl Chloride (fl. conc.) | 75-01-4 | Volatile Organic Compounds (VOCs); | Volatile Organic Compounds (VOCs); EMS code: C004 | Chemical |  | LAB
C005 | mg/L | Captan | ZN01 | DCM+Ac.Extr:GC/MS | DCM+Ac.Extr:GC/MS | 0.001 | Y | Y | Y | 28 | 7/28/2005 7:03:00 AM | 3/29/2012 10:15:00 PM | 5/3/2012 9:27:49 AM | 1996-04-12 | 11/16/1999 4:06:27 PM | EMS | NPEPPIN | Concentration | mg/L | Y | (fl. conc.) |  | Captan (fl. conc.) | 133-06-2 | Pesticides; | Pesticides; EMS code: C005 | Chemical |  | LAB
C005 | ug/g | Captan | ZN03 | DCM+Ac.Extr:GC/MS | DCM+Ac.Extr:GC/MS EPA 8270E (mod) | 0.1 | Y | Y | Y | 2 | 11/2/2022 12:53:00 AM | 11/2/2022 4:41:00 AM | 1/7/2023 3:12:10 PM | 1996-04-12 | 11/16/1999 2:27:40 PM | EMS | NPEPPIN | MassPerMass | mg/kg | Y | (mass conc.) |  | Captan (mass conc.) | 133-06-2 | Pesticides; | Pesticides; EMS code: C005 | Chemical |  | LAB
C007 | mg/L | Carbophenthion | ZN01 | DCM+Ac.Extr:GC/MS | DCM+Ac.Extr:GC/MS | 0.0003 | Y | Y | Y | 52 | 10/3/2002 | 3/29/2012 10:15:00 PM | 5/3/2012 9:27:49 AM | 1996-04-12 | 11/16/1999 4:06:27 PM | EMS | NPEPPIN | Concentration | mg/L | Y | (fl. conc.) |  | Carbophenthion (fl. conc.) |  | Pesticides; | Pesticides; EMS code: C007 | Chemical |  | LAB
C007 | ug/g | Carbophenthion | X260 | Sox-Ac/CH2Cl2 ext:GC/NPD | Sox-Ac/CH2Cl2 ext:GC/NPD | 0.01 | Y | Y | Y | 97 | 4/20/2002 10:25:00 AM | 10/16/2002 3:20:00 PM | 2/25/2004 1:09:44 AM | 1996-04-12 | 6/3/1998 6:51:45 PM | EMS |  | MassPerMass | mg/kg | Y | (mass conc.) |  | Carbophenthion (mass conc.) |  | Pesticides; | Pesticides; EMS code: C007 | Chemical |  | LAB
C009 | mg/L | Chlorfenvinphos(e) | ZN01 | DCM+Ac.Extr:GC/MS | DCM+Ac.Extr:GC/MS | 0.0005 | Y | Y | Y | 4 | 10/3/2002 | 10/10/2003 | 3/24/2004 1:04:39 AM | 1999-11-17 |  | NPEPPIN |  | Concentration | mg/L | Y | (fl. conc.) |  | Chlorfenvinphos(e) (fl. conc.) | 18708-86-6 | Pesticides; | Pesticides; EMS code: C009 | Chemical |  | LAB
C009 | ug/g | Chlorfenvinphos(e) | ZN03 | DCM+Ac.Extr:GC/MS | DCM+Ac.Extr:GC/MS EPA 8270E (mod) | 0.05 | Y | Y | Y | 19 | 4/20/2002 10:25:00 AM | 10/16/2002 3:20:00 PM | 2/25/2004 1:09:44 AM | 1999-11-16 |  | NPEPPIN |  | MassPerMass | mg/kg | Y | (mass conc.) |  | Chlorfenvinphos(e) (mass conc.) | 18708-86-6 | Pesticides; | Pesticides; EMS code: C009 | Chemical |  | LAB
C010 | ug/g | Chlorobenzene | MS02 | Purge&Trap GC/MS (VOC-S) | Purge&Trap GC/MS (VOC-S) | 0.003 | Y | Y | Y | 27 | 11/19/2004 | 10/28/2022 5:04:00 PM | 1/11/2024 1:07:07 PM | 1996-04-12 | 6/3/1998 6:51:34 PM | EMS |  | MassPerMass | mg/kg | Y | (mass conc.) |  | Chlorobenzene (mass conc.) | 108-90-7 | Volatile Organic Compounds (VOCs); | Volatile Organic Compounds (VOCs); EMS code: C010 | Chemical |  | LAB
C010 | mg/L | Chlorobenzene | X384 | Purge and trap GC/MS | The sample is analyzed by purge and trap gas chromatography with detection by mass spectrometry. The procedure involves purging the volatile compounds from the sample with an inert gas, and trapping them on a solid sorbent. When purging of the sample is complete, the trap is heated and the compounds are transferred to the gas chromatographic column. Analysis is then accomplished by separation of the components by gas chromatography with detection by mass spectrometry. | 0.001 | Y | Y | Y | 120 | 4/27/2000 | 7/26/2012 7:12:00 AM | 9/26/2012 1:07:08 PM | 1996-04-12 | 6/3/1998 6:51:34 PM | EMS |  | Concentration | mg/L | Y | (fl. conc.) |  | Chlorobenzene (fl. conc.) | 108-90-7 | Volatile Organic Compounds (VOCs); | Volatile Organic Compounds (VOCs); EMS code: C010 | Chemical |  | LAB
C010 | ug/g | Chlorobenzene | FS04 | Static Headspace:HRGC/LRMS | Static Headspace anaylsis is a method for the determination of volatile compounds in liquids and/or solids by analysing the vapour phase of a sample in a closed system; in this case a sealed vial. Static Headspace is sometimes referred to as Equilibrium Headspace because the vapour phase above the sample in a sealed vial is brought into thermodynamic equilibrium with the sample prior to analysis. Volatile organic compounds are thermally extracted from the sample matrix by Static Headspace sampling and introduced directly to a thick-film megabore GC column. Compounds eluting from the GC column are introduced into a Mass Spectrometer and target compounds are identified by a comparison of both retention times and mass spectra with those of reference standards stored in a computerized data system. Reference retention times and mass spectra for analytes are obtained by the measurement of calibration standards under precisely the same extraction conditions used for samples. The concentration of each identified component is measured by relating the response of the sample quantitation mass ions to the response of the quantitation mass ions of a series of external standards. A deuterated surrogate recovery standard is added to provide information on matrix effects. | 0.05 | Y | Y | Y | 44 | 6/30/1999 4:41:00 PM | 5/22/2015 | 7/3/2015 3:19:03 PM | 1996-09-16 | 6/3/1998 6:51:45 PM | MPAK |  | MassPerMass | mg/kg | Y | (mass conc.) |  | Chlorobenzene (mass conc.) | 108-90-7 | Volatile Organic Compounds (VOCs); | Volatile Organic Compounds (VOCs); EMS code: C010 | Chemical |  | LAB
C010 | mg/L | Chlorobenzene | F007 | Static Headspace:HRGC/LRMS | Static headspace analysis is a method for the determination of volatile compounds in liquids and/or solids by analysing the vapour phase of a sample in a closed system; in this case a sealed vial. Static Headspace is sometimes referred to as Equilibrium Headspace because the vapour phase above the sample in a sealed vial is brought into thermodynamic equilibrium with the sample prior to analysis. Volatile organic compounds are thermally extracted from the sample matrix by Static Headspace sampling and introduced directly to a thick-film megabore GC column. Compounds eluting from the GC column are introduced into a Mass Spectrometer and target compounds are identified by a comparison of both retention times and mass spectra with those of reference standards stored in a computerized data system. Reference retention times and mass spectra for analytes are obtained by the measurement of calibration standards under precisely the same extraction conditions used for samples. The concentration of each identified component is measured by relating the response of the sample quantitation mass ions to the response of the quantitation mass ions of a series of external standards. A deuterated surrogate recovery standard is added to provide information on matrix effects. | 0.005 | Y | Y | Y | 380 | 9/28/1998 3:07:00 PM | 8/14/2024 | 8/16/2024 4:05:21 PM | 1996-09-16 | 6/3/1998 6:51:45 PM | MPAK |  | Concentration | mg/L | Y | (fl. conc.) |  | Chlorobenzene (fl. conc.) | 108-90-7 | Volatile Organic Compounds (VOCs); | Volatile Organic Compounds (VOCs); EMS code: C010 | Chemical |  | LAB
C010 | mg/L | Chlorobenzene | MS04 | Static Headspace GC/MS (Water) | (Water) A static headspace procedure is used as a means of driving the analytes out of the matrix into a headspace. Once equilibrium is achieved, a fixed volume of headspeace is taken for subsequent determination by Gas Chromatography - Mass Spectroscopy (GC/MS) using Selected Ion Monitoring. | 0.0005 | Y | Y | Y | 66 | 8/13/2009 | 6/28/2024 12:29:00 AM | 7/9/2024 3:31:02 PM | 2009-12-18 |  | RDGIBSON |  | Concentration | mg/L | Y | (fl. conc.) |  | Chlorobenzene (fl. conc.) | 108-90-7 | Volatile Organic Compounds (VOCs); | Volatile Organic Compounds (VOCs); EMS code: C010 | Chemical |  | LAB
C010 | mg/L | Chlorobenzene | MS01 | Purge&Trap GC/MS (VOC-W) | Purge&Trap GC/MS (VOC-W) | 0.6 | Y | Y | Y | 544 | 4/15/2002 | 6/4/2024 10:24:00 AM | 6/13/2024 5:01:27 PM | 1996-04-12 | 3/18/2003 2:49:40 PM | EMS | UMULLER | Concentration | mg/L | Y | (fl. conc.) |  | Chlorobenzene (fl. conc.) | 108-90-7 | Volatile Organic Compounds (VOCs); | Volatile Organic Compounds (VOCs); EMS code: C010 | Chemical |  | LAB
C011 | mg/L | alpha-Chlordane | X268 | CH2Cl2 ext:GC/MS | CH2Cl2 ext:GC/MS |  | Y | Y | Y | 16 | 12/29/2021 | 9/2/2022 | 11/10/2022 9:54:06 AM | 2022-04-08 |  | EMS |  | Concentration | mg/L | Y | (fl. conc.) |  | alpha-Chlordane (fl. conc.) | 5103-71-9 | Pesticides; | Pesticides; EMS code: C011 | Chemical |  | LAB
C011 | mg/L | alpha-Chlordane | E660 | EPA 8270E (mod) | EPA 8270E (mod) Miscellaneous Pesticides by GC-MS | 2e-06 | Y | Y | Y | 2 | 9/14/2022 | 9/14/2022 | 12/8/2022 9:48:13 AM | 2022-12-06 |  | EMS |  | Concentration | mg/L | Y | (fl. conc.) |  | alpha-Chlordane (fl. conc.) | 5103-71-9 | Pesticides; | Pesticides; EMS code: C011 | Chemical |  | LAB
C011 | ug/g | alpha-Chlordane | X213 | Hex/AcExt,Part,Chrom,GC | Hex/AcExt;Part;Chrom;GC | 0.001 | Y | Y | Y | 187 | 4/20/2002 10:10:00 AM | 10/13/2002 10:30:00 AM | 2/26/2004 1:05:26 AM | 1996-04-12 | 6/3/1998 6:51:34 PM | EMS |  | MassPerMass | mg/kg | Y | (mass conc.) |  | alpha-Chlordane (mass conc.) | 5103-71-9 | Pesticides; | Pesticides; EMS code: C011 | Chemical |  | LAB
C011 | ug/g | alpha-Chlordane | E8081 | EPA 8081 | See methods EPA 3540, 3630, 3660, 8081 | 0.001 | Y | Y | Y | 9 | 3/18/2020 | 3/23/2020 | 2/22/2023 12:51:03 PM | 2022-07-29 |  | EMS |  | MassPerMass | mg/kg | Y | (mass conc.) |  | alpha-Chlordane (mass conc.) | 5103-71-9 | Pesticides; | Pesticides; EMS code: C011 | Chemical |  | LAB
C011 | ug/g | alpha-Chlordane | ZN03 | DCM+Ac.Extr:GC/MS | DCM+Ac.Extr:GC/MS EPA 8270E (mod) | 0.05 | Y | Y | Y | 17 | 1/17/2022 | 11/14/2022 | 1/11/2024 1:07:07 PM | 1996-04-12 | 11/16/1999 2:27:40 PM | EMS | NPEPPIN | MassPerMass | mg/kg | Y | (mass conc.) |  | alpha-Chlordane (mass conc.) | 5103-71-9 | Pesticides; | Pesticides; EMS code: C011 | Chemical |  | LAB
C011 | ug/g wet | alpha-Chlordane | X220 | Hex/AcExt,Part,Chrom,GC | Hex/AcExt;Part;Chrom;GC | 0.0002 | Y | Y | Y | 73 |  |  | 10/24/1996 10:01:24 AM | 1996-04-12 | 6/3/1998 6:51:41 PM | EMS |  | MassPerMass | µg/g | Y | (mass conc.) | (wet) | alpha-Chlordane (wet) (mass conc.) | 5103-71-9 | Pesticides; | Pesticides; EMS code: C011 | Chemical |  | LAB
C011 | mg/L | alpha-Chlordane | ZN01 | DCM+Ac.Extr:GC/MS | DCM+Ac.Extr:GC/MS | 0.0005 | Y | Y | Y | 52 | 10/3/2002 | 3/29/2012 10:15:00 PM | 5/3/2012 9:27:49 AM | 1996-04-12 | 11/16/1999 4:06:27 PM | EMS | NPEPPIN | Concentration | mg/L | Y | (fl. conc.) |  | alpha-Chlordane (fl. conc.) | 5103-71-9 | Pesticides; | Pesticides; EMS code: C011 | Chemical |  | LAB
C011 | mg/L | alpha-Chlordane | MW03 | OCP1-LL-SF-ECD-VA | This analysis is carried out using procedures adapted from "Test Methods for Evaluating Solid Waste" SW-846, Methods 3510, 3630, 3660 and 8081, published by the United States Environmental Protection Agency (EPA). The procedure involves extraction of the entire water sample with dichloromethane. The extract then undergoes one or more of the following clean-up procedures (if required): silica gel clean-up and/or sulphur clean-up. The final extract is analyzed by dual capillary column gas chromatography with electron capture detection (GC/ECD). | 5e-05 | Y | Y | Y | 10 | 9/20/2016 | 9/21/2016 | 8/15/2019 2:42:34 PM | 2019-08-15 |  | KATFRASE |  | Concentration | mg/L | Y | (fl. conc.) |  | alpha-Chlordane (fl. conc.) | 5103-71-9 | Pesticides; | Pesticides; EMS code: C011 | Chemical |  | LAB
C012 | ug/g | gamma-Chlordane | E8081 | EPA 8081 | See methods EPA 3540, 3630, 3660, 8081 | 0.001 | Y | Y | Y | 9 | 3/18/2020 | 3/23/2020 | 2/22/2023 12:51:03 PM | 2022-07-29 |  | EMS |  | MassPerMass | mg/kg | Y | (mass conc.) |  | gamma-Chlordane (mass conc.) | 5103-74-2 | Pesticides; | Pesticides; EMS code: C012 | Chemical |  | LAB
C012 | ug/g | gamma-Chlordane | X213 | Hex/AcExt,Part,Chrom,GC | Hex/AcExt;Part;Chrom;GC | 0.001 | Y | Y | Y | 187 | 4/20/2002 10:10:00 AM | 10/13/2002 10:30:00 AM | 2/26/2004 1:05:27 AM | 1996-04-12 | 6/3/1998 6:51:34 PM | EMS |  | MassPerMass | mg/kg | Y | (mass conc.) |  | gamma-Chlordane (mass conc.) | 5103-74-2 | Pesticides; | Pesticides; EMS code: C012 | Chemical |  | LAB
C012 | mg/L | gamma-Chlordane | X268 | CH2Cl2 ext:GC/MS | CH2Cl2 ext:GC/MS |  | Y | Y | Y | 16 | 12/29/2021 | 9/2/2022 | 11/10/2022 9:54:06 AM | 2022-04-08 |  | EMS |  | Concentration | mg/L | Y | (fl. conc.) |  | gamma-Chlordane (fl. conc.) | 5103-74-2 | Pesticides; | Pesticides; EMS code: C012 | Chemical |  | LAB
C012 | mg/L | gamma-Chlordane | MW03 | OCP1-LL-SF-ECD-VA | This analysis is carried out using procedures adapted from "Test Methods for Evaluating Solid Waste" SW-846, Methods 3510, 3630, 3660 and 8081, published by the United States Environmental Protection Agency (EPA). The procedure involves extraction of the entire water sample with dichloromethane. The extract then undergoes one or more of the following clean-up procedures (if required): silica gel clean-up and/or sulphur clean-up. The final extract is analyzed by dual capillary column gas chromatography with electron capture detection (GC/ECD). | 5e-05 | Y | Y | Y | 10 | 9/20/2016 | 9/21/2016 | 8/15/2019 2:42:34 PM | 2019-08-15 |  | KATFRASE |  | Concentration | mg/L | Y | (fl. conc.) |  | gamma-Chlordane (fl. conc.) | 5103-74-2 | Pesticides; | Pesticides; EMS code: C012 | Chemical |  | LAB
C012 | ug/g | gamma-Chlordane | ZN03 | DCM+Ac.Extr:GC/MS | DCM+Ac.Extr:GC/MS EPA 8270E (mod) | 0.05 | Y | Y | Y | 17 | 1/17/2022 | 11/14/2022 | 1/11/2024 1:07:07 PM | 1996-04-12 | 11/16/1999 2:27:40 PM | EMS | NPEPPIN | MassPerMass | mg/kg | Y | (mass conc.) |  | gamma-Chlordane (mass conc.) | 5103-74-2 | Pesticides; | Pesticides; EMS code: C012 | Chemical |  | LAB
C012 | ug/g wet | gamma-Chlordane | X220 | Hex/AcExt,Part,Chrom,GC | Hex/AcExt;Part;Chrom;GC | 0.0002 | Y | Y | Y | 73 |  |  | 10/24/1996 10:01:25 AM | 1996-04-12 | 6/3/1998 6:51:41 PM | EMS |  | MassPerMass | µg/g | Y | (mass conc.) | (wet) | gamma-Chlordane (wet) (mass conc.) | 5103-74-2 | Pesticides; | Pesticides; EMS code: C012 | Chemical |  | LAB
C012 | mg/L | gamma-Chlordane | ZN01 | DCM+Ac.Extr:GC/MS | DCM+Ac.Extr:GC/MS | 0.0005 | Y | Y | Y | 52 | 10/3/2002 | 3/29/2012 10:15:00 PM | 5/3/2012 9:27:49 AM | 1996-04-12 | 11/16/1999 4:06:27 PM | EMS | NPEPPIN | Concentration | mg/L | Y | (fl. conc.) |  | gamma-Chlordane (fl. conc.) | 5103-74-2 | Pesticides; | Pesticides; EMS code: C012 | Chemical |  | LAB
C012 | mg/L | gamma-Chlordane | E660 | EPA 8270E (mod) | EPA 8270E (mod) Miscellaneous Pesticides by GC-MS | 2e-06 | Y | Y | Y | 2 | 9/14/2022 | 9/14/2022 | 12/8/2022 9:48:13 AM | 2022-12-06 |  | EMS |  | Concentration | mg/L | Y | (fl. conc.) |  | gamma-Chlordane (fl. conc.) | 5103-74-2 | Pesticides; | Pesticides; EMS code: C012 | Chemical |  | LAB
C013 | ug/g | Total Chlordane | ZN03 | DCM+Ac.Extr:GC/MS | DCM+Ac.Extr:GC/MS EPA 8270E (mod) | 0.03 | Y | Y | Y | 19 | 1/17/2022 | 11/14/2022 | 1/11/2024 1:07:07 PM | 2022-06-21 |  | EMS |  | MassPerMass | mg/kg | Y | (mass conc.) |  | Total Chlordane (mass conc.) | 5103-71-9, 5103-74-2 | Pesticides; | Pesticides; EMS code: C013 | Chemical | Total | LAB
C013 | mg/L | Total Chlordane | X268 | CH2Cl2 ext:GC/MS | CH2Cl2 ext:GC/MS |  | Y | Y | Y | 16 | 12/29/2021 | 9/2/2022 | 11/10/2022 9:54:06 AM | 2022-04-08 |  | EMS |  | Concentration | mg/L | Y | (fl. conc.) |  | Total Chlordane (fl. conc.) | 5103-71-9, 5103-74-2 | Pesticides; | Pesticides; EMS code: C013 | Chemical | Total | LAB
C013 | mg/L | Total Chlordane | E660 | EPA 8270E (mod) | EPA 8270E (mod) Miscellaneous Pesticides by GC-MS | 3e-06 | Y | Y | Y | 2 | 9/14/2022 | 9/14/2022 | 12/8/2022 9:48:13 AM | 2022-12-06 |  | EMS |  | Concentration | mg/L | Y | (fl. conc.) |  | Total Chlordane (fl. conc.) | 5103-71-9, 5103-74-2 | Pesticides; | Pesticides; EMS code: C013 | Chemical | Total | LAB
C014 | mg/L | Chlorthiophos | ZN01 | DCM+Ac.Extr:GC/MS | DCM+Ac.Extr:GC/MS | 0.0003 | Y | Y | Y | 28 | 7/29/2005 7:10:00 AM | 3/29/2012 10:15:00 PM | 5/3/2012 9:27:49 AM | 1996-04-12 | 11/16/1999 4:06:27 PM | EMS | NPEPPIN | Concentration | mg/L | Y | (fl. conc.) |  | Chlorthiophos (fl. conc.) | 60238-56-4 | Pesticides; | Pesticides; EMS code: C014 | Chemical |  | LAB
C015 | mg/L | Chlormephos | ZN01 | DCM+Ac.Extr:GC/MS | DCM+Ac.Extr:GC/MS | 0.0005 | Y | Y | Y | 28 | 7/29/2005 7:10:00 AM | 3/29/2012 10:15:00 PM | 5/3/2012 9:27:49 AM | 1996-04-12 | 11/16/1999 4:06:27 PM | EMS | NPEPPIN | Concentration | mg/L | Y | (fl. conc.) |  | Chlormephos (fl. conc.) | 24934-91-6 | Pesticides; | Pesticides; EMS code: C015 | Chemical |  | LAB
C016 | mg/L | Chlorothalonil | ZN01 | DCM+Ac.Extr:GC/MS | DCM+Ac.Extr:GC/MS | 0.001 | Y | Y | Y | 28 | 7/28/2005 7:03:00 AM | 3/29/2012 10:15:00 PM | 5/3/2012 9:27:50 AM | 1996-04-12 | 11/16/1999 4:06:27 PM | EMS | NPEPPIN | Concentration | mg/L | Y | (fl. conc.) |  | Chlorothalonil (fl. conc.) | 1897-45-6 | Pesticides; | Pesticides; EMS code: C016 | Chemical |  | LAB
C016 | ug/g | Chlorothalonil | ZN03 | DCM+Ac.Extr:GC/MS | DCM+Ac.Extr:GC/MS EPA 8270E (mod) | 0.1 | Y | Y | Y | 2 | 11/2/2022 12:53:00 AM | 11/2/2022 4:41:00 AM | 1/7/2023 3:12:10 PM | 1996-04-12 | 11/16/1999 2:27:40 PM | EMS | NPEPPIN | MassPerMass | mg/kg | Y | (mass conc.) |  | Chlorothalonil (mass conc.) | 1897-45-6 | Pesticides; | Pesticides; EMS code: C016 | Chemical |  | LAB
C017 | mg/L | Chlorpyrifos | E660 | EPA 8270E (mod) | EPA 8270E (mod) Miscellaneous Pesticides by GC-MS | 0.0001 | Y | Y | Y | 68 | 12/7/2021 | 9/12/2022 | 12/8/2022 12:57:05 PM | 2021-12-09 | 12/9/2021 3:22:12 PM | EMS | EMS | Concentration | mg/L | Y | (fl. conc.) |  | Chlorpyrifos (fl. conc.) | 2921-88-2 | Pesticides; | Pesticides; EMS code: C017 | Chemical |  | LAB
C017 | ug/g | Chlorpyrifos | ZN03 | DCM+Ac.Extr:GC/MS | DCM+Ac.Extr:GC/MS EPA 8270E (mod) | 0.02 | Y | Y | Y | 54 | 4/20/2002 10:25:00 AM | 11/2/2022 4:41:00 AM | 1/7/2023 3:12:10 PM | 1996-04-12 | 11/16/1999 2:27:40 PM | EMS | NPEPPIN | MassPerMass | mg/kg | Y | (mass conc.) |  | Chlorpyrifos (mass conc.) | 2921-88-2 | Pesticides; | Pesticides; EMS code: C017 | Chemical |  | LAB
C017 | ug/g | Chlorpyrifos | FS16 | DCM/MEOH extr'n,DCM,conc'n,HRGC-PFPD/N,P det'n | Sediment samples are shaken with a mixture of Dichloromethane and Methanol, filtered, then extracted with Dichloromethane. The extract is then concentrated and analysed using High Resolution Gas Chromatography with Pulsed Flame Photometric Detection. Confirmation, if required, is by High Resolution Gas Chromatography with Nitrogen Phosphorus Detection. | 0.05 | Y | Y | Y | 1 | 4/24/2002 4:00:00 PM | 4/24/2002 4:00:00 PM | 2/12/2004 1:02:34 AM | 1997-10-20 | 6/3/1998 6:51:47 PM | NPEPPIN |  | MassPerMass | mg/kg | Y | (mass conc.) |  | Chlorpyrifos (mass conc.) | 2921-88-2 | Pesticides; | Pesticides; EMS code: C017 | Chemical |  | LAB
C017 | mg/L | Chlorpyrifos | F053 | DCM extr'n,conc'n,HGRC-PFPD/N,P det'n | Effluent samples are extracted with Dichloromethane. The extract is then concentrated and analysed using High Resolution Gas Chromatography with Pulsed Flame Photometric Dectection. Confirmation, if required, is by High Resolution Gas Chromatography with Nitrogen Phosphorus Detection. | 5e-05 | Y | Y | Y | 31 | 3/3/2000 7:36:00 AM | 1/22/2002 10:26:00 AM | 6/18/2003 2:53:31 PM | 1997-10-20 | 6/3/1998 6:51:47 PM | NPEPPIN |  | Concentration | mg/L | Y | (fl. conc.) |  | Chlorpyrifos (fl. conc.) | 2921-88-2 | Pesticides; | Pesticides; EMS code: C017 | Chemical |  | LAB
C017 | mg/L | Chlorpyrifos | ZN01 | DCM+Ac.Extr:GC/MS | DCM+Ac.Extr:GC/MS | 0.0002 | Y | Y | Y | 52 | 10/3/2002 | 3/29/2012 10:15:00 PM | 5/3/2012 9:27:50 AM | 1996-04-12 | 11/16/1999 4:06:27 PM | EMS | NPEPPIN | Concentration | mg/L | Y | (fl. conc.) |  | Chlorpyrifos (fl. conc.) | 2921-88-2 | Pesticides; | Pesticides; EMS code: C017 | Chemical |  | LAB
C019 | mg/L | Chlorpyrifos Methyl | ZN01 | DCM+Ac.Extr:GC/MS | DCM+Ac.Extr:GC/MS | 0.0001 | Y | Y | Y | 1 | 7/29/2005 7:10:00 AM | 7/29/2005 7:10:00 AM | 6/2/2009 12:00:59 PM | 1996-04-12 | 11/16/1999 4:06:27 PM | EMS | NPEPPIN | Concentration | mg/L | Y | (fl. conc.) |  | Chlorpyrifos Methyl (fl. conc.) |  |  | EMS code: C019 | Chemical |  | LAB
C020 | mg/L | Chlorophenols | X268 | CH2Cl2 ext:GC/MS | CH2Cl2 ext:GC/MS |  | Y | Y | Y | 12 | 6/7/2024 | 6/7/2024 | 6/26/2024 4:28:10 PM | 2022-04-20 |  | EMS |  | Concentration | mg/L | Y | (fl. conc.) |  | Chlorophenols (fl. conc.) |  |  | EMS code: C020 | Chemical |  | LAB
C021 | mg/L | 1;1-Dichloroethane | F007 | Static Headspace:HRGC/LRMS | Static headspace analysis is a method for the determination of volatile compounds in liquids and/or solids by analysing the vapour phase of a sample in a closed system; in this case a sealed vial. Static Headspace is sometimes referred to as Equilibrium Headspace because the vapour phase above the sample in a sealed vial is brought into thermodynamic equilibrium with the sample prior to analysis. Volatile organic compounds are thermally extracted from the sample matrix by Static Headspace sampling and introduced directly to a thick-film megabore GC column. Compounds eluting from the GC column are introduced into a Mass Spectrometer and target compounds are identified by a comparison of both retention times and mass spectra with those of reference standards stored in a computerized data system. Reference retention times and mass spectra for analytes are obtained by the measurement of calibration standards under precisely the same extraction conditions used for samples. The concentration of each identified component is measured by relating the response of the sample quantitation mass ions to the response of the quantitation mass ions of a series of external standards. A deuterated surrogate recovery standard is added to provide information on matrix effects. | 0.005 | Y | Y | Y | 380 | 9/28/1998 3:08:00 PM | 8/14/2024 | 8/16/2024 4:05:21 PM | 1996-09-16 | 6/3/1998 6:51:45 PM | MPAK |  | Concentration | mg/L | Y | (fl. conc.) |  | 1;1-Dichloroethane (fl. conc.) | 75-34-3 | Volatile Organic Compounds (VOCs); | Volatile Organic Compounds (VOCs); EMS code: C021 | Chemical |  | LAB
C021 | ug/g | 1;1-Dichloroethane | MS02 | Purge&Trap GC/MS (VOC-S) | Purge&Trap GC/MS (VOC-S) | 0.003 | Y | Y | Y | 27 | 11/19/2004 | 10/28/2022 5:04:00 PM | 1/11/2024 1:07:07 PM | 1996-04-12 | 6/3/1998 6:51:34 PM | EMS |  | MassPerMass | mg/kg | Y | (mass conc.) |  | 1;1-Dichloroethane (mass conc.) | 75-34-3 | Volatile Organic Compounds (VOCs); | Volatile Organic Compounds (VOCs); EMS code: C021 | Chemical |  | LAB
C021 | mg/L | 1;1-Dichloroethane | MS04 | Static Headspace GC/MS (Water) | (Water) A static headspace procedure is used as a means of driving the analytes out of the matrix into a headspace. Once equilibrium is achieved, a fixed volume of headspeace is taken for subsequent determination by Gas Chromatography - Mass Spectroscopy (GC/MS) using Selected Ion Monitoring. | 0.0007 | Y | Y | Y | 70 | 8/13/2009 | 6/28/2024 12:29:00 AM | 7/9/2024 3:31:02 PM | 2009-12-18 |  | RDGIBSON |  | Concentration | mg/L | Y | (fl. conc.) |  | 1;1-Dichloroethane (fl. conc.) | 75-34-3 | Volatile Organic Compounds (VOCs); | Volatile Organic Compounds (VOCs); EMS code: C021 | Chemical |  | LAB
C021 | mg/L | 1;1-Dichloroethane | X384 | Purge and trap GC/MS | The sample is analyzed by purge and trap gas chromatography with detection by mass spectrometry. The procedure involves purging the volatile compounds from the sample with an inert gas, and trapping them on a solid sorbent. When purging of the sample is complete, the trap is heated and the compounds are transferred to the gas chromatographic column. Analysis is then accomplished by separation of the components by gas chromatography with detection by mass spectrometry. | 0.001 | Y | Y | Y | 121 | 4/27/2000 | 7/26/2012 7:12:00 AM | 9/26/2012 1:07:08 PM | 1996-04-12 | 6/3/1998 6:51:34 PM | EMS |  | Concentration | mg/L | Y | (fl. conc.) |  | 1;1-Dichloroethane (fl. conc.) | 75-34-3 | Volatile Organic Compounds (VOCs); | Volatile Organic Compounds (VOCs); EMS code: C021 | Chemical |  | LAB
C021 | ug/g | 1;1-Dichloroethane | FS04 | Static Headspace:HRGC/LRMS | Static Headspace anaylsis is a method for the determination of volatile compounds in liquids and/or solids by analysing the vapour phase of a sample in a closed system; in this case a sealed vial. Static Headspace is sometimes referred to as Equilibrium Headspace because the vapour phase above the sample in a sealed vial is brought into thermodynamic equilibrium with the sample prior to analysis. Volatile organic compounds are thermally extracted from the sample matrix by Static Headspace sampling and introduced directly to a thick-film megabore GC column. Compounds eluting from the GC column are introduced into a Mass Spectrometer and target compounds are identified by a comparison of both retention times and mass spectra with those of reference standards stored in a computerized data system. Reference retention times and mass spectra for analytes are obtained by the measurement of calibration standards under precisely the same extraction conditions used for samples. The concentration of each identified component is measured by relating the response of the sample quantitation mass ions to the response of the quantitation mass ions of a series of external standards. A deuterated surrogate recovery standard is added to provide information on matrix effects. | 0.05 | Y | Y | Y | 44 | 6/30/1999 4:41:00 PM | 5/22/2015 | 7/3/2015 3:19:03 PM | 1996-09-16 | 6/3/1998 6:51:45 PM | MPAK |  | MassPerMass | mg/kg | Y | (mass conc.) |  | 1;1-Dichloroethane (mass conc.) | 75-34-3 | Volatile Organic Compounds (VOCs); | Volatile Organic Compounds (VOCs); EMS code: C021 | Chemical |  | LAB
C021 | mg/L | 1;1-Dichloroethane | MS01 | Purge&Trap GC/MS (VOC-W) | Purge&Trap GC/MS (VOC-W) | 0.5 | Y | Y | Y | 544 | 4/15/2002 | 6/4/2024 10:24:00 AM | 6/13/2024 5:01:27 PM | 1996-04-12 | 3/18/2003 2:49:40 PM | EMS | UMULLER | Concentration | mg/L | Y | (fl. conc.) |  | 1;1-Dichloroethane (fl. conc.) | 75-34-3 | Volatile Organic Compounds (VOCs); | Volatile Organic Compounds (VOCs); EMS code: C021 | Chemical |  | LAB
C022 | mg/L | 1;2-Dichloroethane | MS01 | Purge&Trap GC/MS (VOC-W) | Purge&Trap GC/MS (VOC-W) | 0.4 | Y | Y | Y | 544 | 4/15/2002 | 6/4/2024 10:24:00 AM | 6/13/2024 5:01:27 PM | 1996-04-12 | 3/18/2003 2:49:40 PM | EMS | UMULLER | Concentration | mg/L | Y | (fl. conc.) |  | 1;2-Dichloroethane (fl. conc.) | 107-06-2 | Volatile Organic Compounds (VOCs); | Volatile Organic Compounds (VOCs); EMS code: C022 | Chemical |  | LAB
C022 | mg/L | 1;2-Dichloroethane | MS04 | Static Headspace GC/MS (Water) | (Water) A static headspace procedure is used as a means of driving the analytes out of the matrix into a headspace. Once equilibrium is achieved, a fixed volume of headspeace is taken for subsequent determination by Gas Chromatography - Mass Spectroscopy (GC/MS) using Selected Ion Monitoring. | 0.0005 | Y | Y | Y | 69 | 8/13/2009 | 6/28/2024 12:29:00 AM | 7/9/2024 3:31:02 PM | 2009-12-18 |  | RDGIBSON |  | Concentration | mg/L | Y | (fl. conc.) |  | 1;2-Dichloroethane (fl. conc.) | 107-06-2 | Volatile Organic Compounds (VOCs); | Volatile Organic Compounds (VOCs); EMS code: C022 | Chemical |  | LAB
C022 | mg/L | 1;2-Dichloroethane | F007 | Static Headspace:HRGC/LRMS | Static headspace analysis is a method for the determination of volatile compounds in liquids and/or solids by analysing the vapour phase of a sample in a closed system; in this case a sealed vial. Static Headspace is sometimes referred to as Equilibrium Headspace because the vapour phase above the sample in a sealed vial is brought into thermodynamic equilibrium with the sample prior to analysis. Volatile organic compounds are thermally extracted from the sample matrix by Static Headspace sampling and introduced directly to a thick-film megabore GC column. Compounds eluting from the GC column are introduced into a Mass Spectrometer and target compounds are identified by a comparison of both retention times and mass spectra with those of reference standards stored in a computerized data system. Reference retention times and mass spectra for analytes are obtained by the measurement of calibration standards under precisely the same extraction conditions used for samples. The concentration of each identified component is measured by relating the response of the sample quantitation mass ions to the response of the quantitation mass ions of a series of external standards. A deuterated surrogate recovery standard is added to provide information on matrix effects. | 0.005 | Y | Y | Y | 381 | 9/28/1998 3:08:00 PM | 8/14/2024 | 8/16/2024 4:05:21 PM | 1996-09-16 | 6/3/1998 6:51:45 PM | MPAK |  | Concentration | mg/L | Y | (fl. conc.) |  | 1;2-Dichloroethane (fl. conc.) | 107-06-2 | Volatile Organic Compounds (VOCs); | Volatile Organic Compounds (VOCs); EMS code: C022 | Chemical |  | LAB
C022 | mg/L | 1;2-Dichloroethane | X384 | Purge and trap GC/MS | The sample is analyzed by purge and trap gas chromatography with detection by mass spectrometry. The procedure involves purging the volatile compounds from the sample with an inert gas, and trapping them on a solid sorbent. When purging of the sample is complete, the trap is heated and the compounds are transferred to the gas chromatographic column. Analysis is then accomplished by separation of the components by gas chromatography with detection by mass spectrometry. | 0.001 | Y | Y | Y | 119 | 4/27/2000 | 7/26/2012 7:12:00 AM | 9/26/2012 1:07:08 PM | 1996-04-12 | 6/3/1998 6:51:34 PM | EMS |  | Concentration | mg/L | Y | (fl. conc.) |  | 1;2-Dichloroethane (fl. conc.) | 107-06-2 | Volatile Organic Compounds (VOCs); | Volatile Organic Compounds (VOCs); EMS code: C022 | Chemical |  | LAB
C022 | ug/g | 1;2-Dichloroethane | FS04 | Static Headspace:HRGC/LRMS | Static Headspace anaylsis is a method for the determination of volatile compounds in liquids and/or solids by analysing the vapour phase of a sample in a closed system; in this case a sealed vial. Static Headspace is sometimes referred to as Equilibrium Headspace because the vapour phase above the sample in a sealed vial is brought into thermodynamic equilibrium with the sample prior to analysis. Volatile organic compounds are thermally extracted from the sample matrix by Static Headspace sampling and introduced directly to a thick-film megabore GC column. Compounds eluting from the GC column are introduced into a Mass Spectrometer and target compounds are identified by a comparison of both retention times and mass spectra with those of reference standards stored in a computerized data system. Reference retention times and mass spectra for analytes are obtained by the measurement of calibration standards under precisely the same extraction conditions used for samples. The concentration of each identified component is measured by relating the response of the sample quantitation mass ions to the response of the quantitation mass ions of a series of external standards. A deuterated surrogate recovery standard is added to provide information on matrix effects. | 0.05 | Y | Y | Y | 44 | 6/30/1999 4:41:00 PM | 5/22/2015 | 7/3/2015 3:19:03 PM | 1996-09-16 | 6/3/1998 6:51:45 PM | MPAK |  | MassPerMass | mg/kg | Y | (mass conc.) |  | 1;2-Dichloroethane (mass conc.) | 107-06-2 | Volatile Organic Compounds (VOCs); | Volatile Organic Compounds (VOCs); EMS code: C022 | Chemical |  | LAB
C022 | ug/g | 1;2-Dichloroethane | MS02 | Purge&Trap GC/MS (VOC-S) | Purge&Trap GC/MS (VOC-S) | 0.002 | Y | Y | Y | 27 | 11/19/2004 | 10/28/2022 5:04:00 PM | 1/11/2024 1:07:07 PM | 1996-04-12 | 6/3/1998 6:51:45 PM | EMS |  | MassPerMass | mg/kg | Y | (mass conc.) |  | 1;2-Dichloroethane (mass conc.) | 107-06-2 | Volatile Organic Compounds (VOCs); | Volatile Organic Compounds (VOCs); EMS code: C022 | Chemical |  | LAB
C023 | mg/L | trans1;2-Dichloroethylene | MS04 | Static Headspace GC/MS (Water) | (Water) A static headspace procedure is used as a means of driving the analytes out of the matrix into a headspace. Once equilibrium is achieved, a fixed volume of headspeace is taken for subsequent determination by Gas Chromatography - Mass Spectroscopy (GC/MS) using Selected Ion Monitoring. | 0.0005 | Y | Y | Y | 56 | 8/13/2009 | 7/9/2013 2:45:00 AM | 3/10/2023 3:28:43 PM | 2009-12-18 | 12/24/2009 10:04:12 AM | RDGIBSON | EMS | Concentration | mg/L | Y | (fl. conc.) |  | trans1;2-Dichloroethylene (fl. conc.) | 156-60-5 | Volatile Organic Compounds (VOCs); | Volatile Organic Compounds (VOCs); EMS code: C023 | Chemical |  | LAB
C023 | mg/L | trans1;2-Dichloroethylene | MS01 | Purge&Trap GC/MS (VOC-W) | Purge&Trap GC/MS (VOC-W) | 0.4 | Y | Y | Y | 185 | 4/15/2002 | 4/1/2011 | 7/9/2012 2:09:47 PM | 1996-04-12 | 3/18/2003 2:49:40 PM | EMS | UMULLER | Concentration | mg/L | Y | (fl. conc.) |  | trans1;2-Dichloroethylene (fl. conc.) | 156-60-5 | Volatile Organic Compounds (VOCs); | Volatile Organic Compounds (VOCs); EMS code: C023 | Chemical |  | LAB
C023 | ug/g | trans1;2-Dichloroethylene | FS04 | Static Headspace:HRGC/LRMS | Static Headspace anaylsis is a method for the determination of volatile compounds in liquids and/or solids by analysing the vapour phase of a sample in a closed system; in this case a sealed vial. Static Headspace is sometimes referred to as Equilibrium Headspace because the vapour phase above the sample in a sealed vial is brought into thermodynamic equilibrium with the sample prior to analysis. Volatile organic compounds are thermally extracted from the sample matrix by Static Headspace sampling and introduced directly to a thick-film megabore GC column. Compounds eluting from the GC column are introduced into a Mass Spectrometer and target compounds are identified by a comparison of both retention times and mass spectra with those of reference standards stored in a computerized data system. Reference retention times and mass spectra for analytes are obtained by the measurement of calibration standards under precisely the same extraction conditions used for samples. The concentration of each identified component is measured by relating the response of the sample quantitation mass ions to the response of the quantitation mass ions of a series of external standards. A deuterated surrogate recovery standard is added to provide information on matrix effects. | 0.05 | Y | Y | Y | 44 | 6/30/1999 4:41:00 PM | 5/22/2015 | 7/3/2015 3:19:03 PM | 1996-09-16 | 6/3/1998 6:51:45 PM | MPAK |  | MassPerMass | mg/kg | Y | (mass conc.) |  | trans1;2-Dichloroethylene (mass conc.) | 156-60-5 | Volatile Organic Compounds (VOCs); | Volatile Organic Compounds (VOCs); EMS code: C023 | Chemical |  | LAB
C023 | mg/L | trans1;2-Dichloroethylene | F007 | Static Headspace:HRGC/LRMS | Static headspace analysis is a method for the determination of volatile compounds in liquids and/or solids by analysing the vapour phase of a sample in a closed system; in this case a sealed vial. Static Headspace is sometimes referred to as Equilibrium Headspace because the vapour phase above the sample in a sealed vial is brought into thermodynamic equilibrium with the sample prior to analysis. Volatile organic compounds are thermally extracted from the sample matrix by Static Headspace sampling and introduced directly to a thick-film megabore GC column. Compounds eluting from the GC column are introduced into a Mass Spectrometer and target compounds are identified by a comparison of both retention times and mass spectra with those of reference standards stored in a computerized data system. Reference retention times and mass spectra for analytes are obtained by the measurement of calibration standards under precisely the same extraction conditions used for samples. The concentration of each identified component is measured by relating the response of the sample quantitation mass ions to the response of the quantitation mass ions of a series of external standards. A deuterated surrogate recovery standard is added to provide information on matrix effects. | 0.005 | Y | Y | Y | 380 | 9/28/1998 3:08:00 PM | 8/14/2024 | 8/16/2024 4:05:21 PM | 1996-09-16 | 6/3/1998 6:51:44 PM | MPAK |  | Concentration | mg/L | Y | (fl. conc.) |  | trans1;2-Dichloroethylene (fl. conc.) | 156-60-5 | Volatile Organic Compounds (VOCs); | Volatile Organic Compounds (VOCs); EMS code: C023 | Chemical |  | LAB
C023 | ug/g | trans1;2-Dichloroethylene | MS02 | Purge&Trap GC/MS (VOC-S) | Purge&Trap GC/MS (VOC-S) | 0.002 | Y | Y | Y | 26 | 11/19/2004 | 2/4/2022 | 1/11/2024 1:07:07 PM | 1996-04-12 | 6/3/1998 6:51:45 PM | EMS |  | MassPerMass | mg/kg | Y | (mass conc.) |  | trans1;2-Dichloroethylene (mass conc.) | 156-60-5 | Volatile Organic Compounds (VOCs); | Volatile Organic Compounds (VOCs); EMS code: C023 | Chemical |  | LAB
C023 | mg/L | trans1;2-Dichloroethylene | X384 | Purge and trap GC/MS | The sample is analyzed by purge and trap gas chromatography with detection by mass spectrometry. The procedure involves purging the volatile compounds from the sample with an inert gas, and trapping them on a solid sorbent. When purging of the sample is complete, the trap is heated and the compounds are transferred to the gas chromatographic column. Analysis is then accomplished by separation of the components by gas chromatography with detection by mass spectrometry. | 0.001 | Y | Y | Y | 10 |  |  | 3/19/2004 8:18:43 AM | 1996-04-12 | 6/3/1998 6:51:34 PM | EMS |  | Concentration | mg/L | Y | (fl. conc.) |  | trans1;2-Dichloroethylene (fl. conc.) | 156-60-5 | Volatile Organic Compounds (VOCs); | Volatile Organic Compounds (VOCs); EMS code: C023 | Chemical |  | LAB
C024 | mg/L | 1;1-Dichloroethylene | X384 | Purge and trap GC/MS | The sample is analyzed by purge and trap gas chromatography with detection by mass spectrometry. The procedure involves purging the volatile compounds from the sample with an inert gas, and trapping them on a solid sorbent. When purging of the sample is complete, the trap is heated and the compounds are transferred to the gas chromatographic column. Analysis is then accomplished by separation of the components by gas chromatography with detection by mass spectrometry. | 0.001 | Y | Y | Y | 1 |  |  | 10/23/1996 11:03:45 PM | 1996-04-12 | 6/3/1998 6:51:34 PM | EMS |  | Concentration | mg/L | Y | (fl. conc.) |  | 1;1-Dichloroethylene (fl. conc.) | 75-35-4 | Volatile Organic Compounds (VOCs); | Volatile Organic Compounds (VOCs); EMS code: C024 | Chemical |  | LAB
C024 | mg/L | 1;1-Dichloroethylene | MS04 | Static Headspace GC/MS (Water) | (Water) A static headspace procedure is used as a means of driving the analytes out of the matrix into a headspace. Once equilibrium is achieved, a fixed volume of headspeace is taken for subsequent determination by Gas Chromatography - Mass Spectroscopy (GC/MS) using Selected Ion Monitoring. | 0.0005 | Y | Y | Y | 70 | 8/13/2009 | 6/28/2024 12:29:00 AM | 7/9/2024 3:31:02 PM | 2009-12-18 |  | RDGIBSON |  | Concentration | mg/L | Y | (fl. conc.) |  | 1;1-Dichloroethylene (fl. conc.) | 75-35-4 | Volatile Organic Compounds (VOCs); | Volatile Organic Compounds (VOCs); EMS code: C024 | Chemical |  | LAB
C024 | ug/g | 1;1-Dichloroethylene | MS02 | Purge&Trap GC/MS (VOC-S) | Purge&Trap GC/MS (VOC-S) | 0.002 | Y | Y | Y | 27 | 11/19/2004 | 10/28/2022 5:04:00 PM | 1/11/2024 1:07:07 PM | 1996-04-12 | 6/3/1998 6:51:34 PM | EMS |  | MassPerMass | mg/kg | Y | (mass conc.) |  | 1;1-Dichloroethylene (mass conc.) | 75-35-4 | Volatile Organic Compounds (VOCs); | Volatile Organic Compounds (VOCs); EMS code: C024 | Chemical |  | LAB
C024 | ug/g | 1;1-Dichloroethylene | FS04 | Static Headspace:HRGC/LRMS | Static Headspace anaylsis is a method for the determination of volatile compounds in liquids and/or solids by analysing the vapour phase of a sample in a closed system; in this case a sealed vial. Static Headspace is sometimes referred to as Equilibrium Headspace because the vapour phase above the sample in a sealed vial is brought into thermodynamic equilibrium with the sample prior to analysis. Volatile organic compounds are thermally extracted from the sample matrix by Static Headspace sampling and introduced directly to a thick-film megabore GC column. Compounds eluting from the GC column are introduced into a Mass Spectrometer and target compounds are identified by a comparison of both retention times and mass spectra with those of reference standards stored in a computerized data system. Reference retention times and mass spectra for analytes are obtained by the measurement of calibration standards under precisely the same extraction conditions used for samples. The concentration of each identified component is measured by relating the response of the sample quantitation mass ions to the response of the quantitation mass ions of a series of external standards. A deuterated surrogate recovery standard is added to provide information on matrix effects. | 0.05 | Y | Y | Y | 44 | 6/30/1999 4:41:00 PM | 5/22/2015 | 7/3/2015 3:19:03 PM | 1996-09-16 | 6/3/1998 6:51:44 PM | MPAK |  | MassPerMass | mg/kg | Y | (mass conc.) |  | 1;1-Dichloroethylene (mass conc.) | 75-35-4 | Volatile Organic Compounds (VOCs); | Volatile Organic Compounds (VOCs); EMS code: C024 | Chemical |  | LAB
C024 | mg/L | 1;1-Dichloroethylene | MS01 | Purge&Trap GC/MS (VOC-W) | Purge&Trap GC/MS (VOC-W) | 0.4 | Y | Y | Y | 311 | 4/15/2002 | 6/4/2024 10:24:00 AM | 6/13/2024 5:01:27 PM | 1996-04-12 | 3/18/2003 2:49:40 PM | EMS | UMULLER | Concentration | mg/L | Y | (fl. conc.) |  | 1;1-Dichloroethylene (fl. conc.) | 75-35-4 | Volatile Organic Compounds (VOCs); | Volatile Organic Compounds (VOCs); EMS code: C024 | Chemical |  | LAB
C024 | mg/L | 1;1-Dichloroethylene | F007 | Static Headspace:HRGC/LRMS | Static headspace analysis is a method for the determination of volatile compounds in liquids and/or solids by analysing the vapour phase of a sample in a closed system; in this case a sealed vial. Static Headspace is sometimes referred to as Equilibrium Headspace because the vapour phase above the sample in a sealed vial is brought into thermodynamic equilibrium with the sample prior to analysis. Volatile organic compounds are thermally extracted from the sample matrix by Static Headspace sampling and introduced directly to a thick-film megabore GC column. Compounds eluting from the GC column are introduced into a Mass Spectrometer and target compounds are identified by a comparison of both retention times and mass spectra with those of reference standards stored in a computerized data system. Reference retention times and mass spectra for analytes are obtained by the measurement of calibration standards under precisely the same extraction conditions used for samples. The concentration of each identified component is measured by relating the response of the sample quantitation mass ions to the response of the quantitation mass ions of a series of external standards. A deuterated surrogate recovery standard is added to provide information on matrix effects. | 0.005 | Y | Y | Y | 380 | 9/28/1998 3:08:00 PM | 8/14/2024 | 8/16/2024 4:05:21 PM | 1996-09-16 | 6/3/1998 6:51:44 PM | MPAK |  | Concentration | mg/L | Y | (fl. conc.) |  | 1;1-Dichloroethylene (fl. conc.) | 75-35-4 | Volatile Organic Compounds (VOCs); | Volatile Organic Compounds (VOCs); EMS code: C024 | Chemical |  | LAB
C025 | ug/g | 1;2-Dichloropropane | MS02 | Purge&Trap GC/MS (VOC-S) | Purge&Trap GC/MS (VOC-S) | 0.05 | Y | Y | Y | 16 | 1/14/2022 | 10/28/2022 5:04:00 PM | 1/11/2024 1:07:07 PM | 2022-07-27 |  | EMS |  | MassPerMass | mg/kg | Y | (mass conc.) |  | 1;2-Dichloropropane (mass conc.) | 78-87-5 | Volatile Organic Compounds (VOCs); | Volatile Organic Compounds (VOCs); EMS code: C025 | Chemical |  | LAB
C025 | mg/L | 1;2-Dichloropropane | MS01 | Purge&Trap GC/MS (VOC-W) | Purge&Trap GC/MS (VOC-W) | 0.0001 | Y | Y | Y | 450 | 4/15/2002 | 6/4/2024 10:24:00 AM | 6/13/2024 5:01:27 PM | 2003-03-18 | 10/7/2010 1:39:36 PM | UMULLER | RDGIBSON | Concentration | mg/L | Y | (fl. conc.) |  | 1;2-Dichloropropane (fl. conc.) | 78-87-5 | Volatile Organic Compounds (VOCs); | Volatile Organic Compounds (VOCs); EMS code: C025 | Chemical |  | LAB
C025 | ug/g | 1;2-Dichloropropane | FS04 | Static Headspace:HRGC/LRMS | Static Headspace anaylsis is a method for the determination of volatile compounds in liquids and/or solids by analysing the vapour phase of a sample in a closed system; in this case a sealed vial. Static Headspace is sometimes referred to as Equilibrium Headspace because the vapour phase above the sample in a sealed vial is brought into thermodynamic equilibrium with the sample prior to analysis. Volatile organic compounds are thermally extracted from the sample matrix by Static Headspace sampling and introduced directly to a thick-film megabore GC column. Compounds eluting from the GC column are introduced into a Mass Spectrometer and target compounds are identified by a comparison of both retention times and mass spectra with those of reference standards stored in a computerized data system. Reference retention times and mass spectra for analytes are obtained by the measurement of calibration standards under precisely the same extraction conditions used for samples. The concentration of each identified component is measured by relating the response of the sample quantitation mass ions to the response of the quantitation mass ions of a series of external standards. A deuterated surrogate recovery standard is added to provide information on matrix effects. | 0.1 | Y | Y | Y | 55 | 6/30/1999 5:05:00 PM | 5/22/2015 | 7/3/2015 3:19:03 PM | 1996-09-16 | 6/3/1998 6:51:45 PM | MPAK |  | MassPerMass | mg/kg | Y | (mass conc.) |  | 1;2-Dichloropropane (mass conc.) | 78-87-5 | Volatile Organic Compounds (VOCs); | Volatile Organic Compounds (VOCs); EMS code: C025 | Chemical |  | LAB
C025 | mg/L | 1;2-Dichloropropane | X384 | Purge and trap GC/MS | The sample is analyzed by purge and trap gas chromatography with detection by mass spectrometry. The procedure involves purging the volatile compounds from the sample with an inert gas, and trapping them on a solid sorbent. When purging of the sample is complete, the trap is heated and the compounds are transferred to the gas chromatographic column. Analysis is then accomplished by separation of the components by gas chromatography with detection by mass spectrometry. | 0.001 | Y | Y | Y | 181 | 8/5/2004 10:31:00 AM | 7/6/2010 7:02:00 AM | 7/19/2010 1:33:13 PM | 1996-04-12 | 6/3/1998 6:51:34 PM | EMS |  | Concentration | mg/L | Y | (fl. conc.) |  | 1;2-Dichloropropane (fl. conc.) | 78-87-5 | Volatile Organic Compounds (VOCs); | Volatile Organic Compounds (VOCs); EMS code: C025 | Chemical |  | LAB
C025 | mg/L | 1;2-Dichloropropane | MS04 | Static Headspace GC/MS (Water) | (Water) A static headspace procedure is used as a means of driving the analytes out of the matrix into a headspace. Once equilibrium is achieved, a fixed volume of headspeace is taken for subsequent determination by Gas Chromatography - Mass Spectroscopy (GC/MS) using Selected Ion Monitoring. | 0.0005 | Y | Y | Y | 70 | 8/13/2009 | 6/28/2024 12:29:00 AM | 7/9/2024 3:31:02 PM | 2009-12-18 |  | RDGIBSON |  | Concentration | mg/L | Y | (fl. conc.) |  | 1;2-Dichloropropane (fl. conc.) | 78-87-5 | Volatile Organic Compounds (VOCs); | Volatile Organic Compounds (VOCs); EMS code: C025 | Chemical |  | LAB
C025 | mg/L | 1;2-Dichloropropane | F007 | Static Headspace:HRGC/LRMS | Static headspace analysis is a method for the determination of volatile compounds in liquids and/or solids by analysing the vapour phase of a sample in a closed system; in this case a sealed vial. Static Headspace is sometimes referred to as Equilibrium Headspace because the vapour phase above the sample in a sealed vial is brought into thermodynamic equilibrium with the sample prior to analysis. Volatile organic compounds are thermally extracted from the sample matrix by Static Headspace sampling and introduced directly to a thick-film megabore GC column. Compounds eluting from the GC column are introduced into a Mass Spectrometer and target compounds are identified by a comparison of both retention times and mass spectra with those of reference standards stored in a computerized data system. Reference retention times and mass spectra for analytes are obtained by the measurement of calibration standards under precisely the same extraction conditions used for samples. The concentration of each identified component is measured by relating the response of the sample quantitation mass ions to the response of the quantitation mass ions of a series of external standards. A deuterated surrogate recovery standard is added to provide information on matrix effects. | 0.01 | Y | Y | Y | 380 | 9/28/1998 3:08:00 PM | 8/14/2024 | 8/16/2024 4:05:21 PM | 1996-09-16 | 6/3/1998 6:51:45 PM | MPAK |  | Concentration | mg/L | Y | (fl. conc.) |  | 1;2-Dichloropropane (fl. conc.) | 78-87-5 | Volatile Organic Compounds (VOCs); | Volatile Organic Compounds (VOCs); EMS code: C025 | Chemical |  | LAB
C026 | mg/L | 1;3-Dichloropropane | MS04 | Static Headspace GC/MS (Water) | (Water) A static headspace procedure is used as a means of driving the analytes out of the matrix into a headspace. Once equilibrium is achieved, a fixed volume of headspeace is taken for subsequent determination by Gas Chromatography - Mass Spectroscopy (GC/MS) using Selected Ion Monitoring. | 0.001 | Y | Y | Y | 22 | 10/5/2019 2:07:00 AM | 4/12/2024 1:31:00 AM | 5/1/2024 2:57:07 PM | 2021-04-07 |  | RDGIBSON |  | Concentration | mg/L | Y | (fl. conc.) |  | 1;3-Dichloropropane (fl. conc.) |  | Volatile Organic Compounds (VOCs); | Volatile Organic Compounds (VOCs); EMS code: C026 | Chemical |  | LAB
C027 | mg/L | cis-1;3-Dichloropropene | MS04 | Static Headspace GC/MS (Water) | (Water) A static headspace procedure is used as a means of driving the analytes out of the matrix into a headspace. Once equilibrium is achieved, a fixed volume of headspeace is taken for subsequent determination by Gas Chromatography - Mass Spectroscopy (GC/MS) using Selected Ion Monitoring. | 0.001 | Y | Y | Y | 69 | 8/13/2009 | 6/28/2024 12:29:00 AM | 7/9/2024 3:31:02 PM | 2009-12-23 |  | RDGIBSON |  | Concentration | mg/L | Y | (fl. conc.) |  | cis-1;3-Dichloropropene (fl. conc.) | 10061-01-5 | Volatile Organic Compounds (VOCs); | Volatile Organic Compounds (VOCs); EMS code: C027 | Chemical |  | LAB
C027 | mg/L | cis-1;3-Dichloropropene | X384 | Purge and trap GC/MS | The sample is analyzed by purge and trap gas chromatography with detection by mass spectrometry. The procedure involves purging the volatile compounds from the sample with an inert gas, and trapping them on a solid sorbent. When purging of the sample is complete, the trap is heated and the compounds are transferred to the gas chromatographic column. Analysis is then accomplished by separation of the components by gas chromatography with detection by mass spectrometry. | 0.001 | Y | Y | Y | 112 | 8/5/2004 10:31:00 AM | 7/26/2012 7:12:00 AM | 9/26/2012 1:07:08 PM | 1996-04-12 | 6/3/1998 6:51:45 PM | EMS |  | Concentration | mg/L | Y | (fl. conc.) |  | cis-1;3-Dichloropropene (fl. conc.) | 10061-01-5 | Volatile Organic Compounds (VOCs); | Volatile Organic Compounds (VOCs); EMS code: C027 | Chemical |  | LAB
C027 | mg/L | cis-1;3-Dichloropropene | MS01 | Purge&Trap GC/MS (VOC-W) | Purge&Trap GC/MS (VOC-W) | 0.7 | Y | Y | Y | 378 | 4/15/2002 | 6/1/2015 12:01:00 AM | 1/24/2016 3:03:56 PM | 1996-04-12 | 3/18/2003 2:49:40 PM | EMS | UMULLER | Concentration | mg/L | Y | (fl. conc.) |  | cis-1;3-Dichloropropene (fl. conc.) | 10061-01-5 | Volatile Organic Compounds (VOCs); | Volatile Organic Compounds (VOCs); EMS code: C027 | Chemical |  | LAB
C027 | ug/g | cis-1;3-Dichloropropene | FS04 | Static Headspace:HRGC/LRMS | Static Headspace anaylsis is a method for the determination of volatile compounds in liquids and/or solids by analysing the vapour phase of a sample in a closed system; in this case a sealed vial. Static Headspace is sometimes referred to as Equilibrium Headspace because the vapour phase above the sample in a sealed vial is brought into thermodynamic equilibrium with the sample prior to analysis. Volatile organic compounds are thermally extracted from the sample matrix by Static Headspace sampling and introduced directly to a thick-film megabore GC column. Compounds eluting from the GC column are introduced into a Mass Spectrometer and target compounds are identified by a comparison of both retention times and mass spectra with those of reference standards stored in a computerized data system. Reference retention times and mass spectra for analytes are obtained by the measurement of calibration standards under precisely the same extraction conditions used for samples. The concentration of each identified component is measured by relating the response of the sample quantitation mass ions to the response of the quantitation mass ions of a series of external standards. A deuterated surrogate recovery standard is added to provide information on matrix effects. | 0.05 | Y | Y | Y | 44 | 6/30/1999 4:41:00 PM | 5/22/2015 | 7/3/2015 3:19:04 PM | 1996-09-16 | 6/3/1998 6:51:45 PM | MPAK |  | MassPerMass | mg/kg | Y | (mass conc.) |  | cis-1;3-Dichloropropene (mass conc.) | 10061-01-5 | Volatile Organic Compounds (VOCs); | Volatile Organic Compounds (VOCs); EMS code: C027 | Chemical |  | LAB
C027 | ug/g | cis-1;3-Dichloropropene | MS02 | Purge&Trap GC/MS (VOC-S) | Purge&Trap GC/MS (VOC-S) | 0.003 | Y | Y | Y | 26 | 11/19/2004 | 2/4/2022 | 1/11/2024 1:07:07 PM | 1996-04-12 | 6/3/1998 6:51:34 PM | EMS |  | MassPerMass | mg/kg | Y | (mass conc.) |  | cis-1;3-Dichloropropene (mass conc.) | 10061-01-5 | Volatile Organic Compounds (VOCs); | Volatile Organic Compounds (VOCs); EMS code: C027 | Chemical |  | LAB
C027 | mg/L | cis-1;3-Dichloropropene | F007 | Static Headspace:HRGC/LRMS | Static headspace analysis is a method for the determination of volatile compounds in liquids and/or solids by analysing the vapour phase of a sample in a closed system; in this case a sealed vial. Static Headspace is sometimes referred to as Equilibrium Headspace because the vapour phase above the sample in a sealed vial is brought into thermodynamic equilibrium with the sample prior to analysis. Volatile organic compounds are thermally extracted from the sample matrix by Static Headspace sampling and introduced directly to a thick-film megabore GC column. Compounds eluting from the GC column are introduced into a Mass Spectrometer and target compounds are identified by a comparison of both retention times and mass spectra with those of reference standards stored in a computerized data system. Reference retention times and mass spectra for analytes are obtained by the measurement of calibration standards under precisely the same extraction conditions used for samples. The concentration of each identified component is measured by relating the response of the sample quantitation mass ions to the response of the quantitation mass ions of a series of external standards. A deuterated surrogate recovery standard is added to provide information on matrix effects. | 0.005 | Y | Y | Y | 381 | 9/28/1998 3:08:00 PM | 8/14/2024 | 8/16/2024 4:05:21 PM | 1996-09-16 | 6/3/1998 6:51:45 PM | MPAK |  | Concentration | mg/L | Y | (fl. conc.) |  | cis-1;3-Dichloropropene (fl. conc.) | 10061-01-5 | Volatile Organic Compounds (VOCs); | Volatile Organic Compounds (VOCs); EMS code: C027 | Chemical |  | LAB
C028 | ug/g | trans1;3-Dichloropropene | MS02 | Purge&Trap GC/MS (VOC-S) | Purge&Trap GC/MS (VOC-S) | 0.006 | Y | Y | Y | 26 | 11/19/2004 | 2/4/2022 | 1/11/2024 1:07:07 PM | 1996-04-12 | 6/3/1998 6:51:34 PM | EMS |  | MassPerMass | mg/kg | Y | (mass conc.) |  | trans1;3-Dichloropropene (mass conc.) | 10061-02-6 | Volatile Organic Compounds (VOCs); | Volatile Organic Compounds (VOCs); EMS code: C028 | Chemical |  | LAB
C028 | mg/L | trans1;3-Dichloropropene | F007 | Static Headspace:HRGC/LRMS | Static headspace analysis is a method for the determination of volatile compounds in liquids and/or solids by analysing the vapour phase of a sample in a closed system; in this case a sealed vial. Static Headspace is sometimes referred to as Equilibrium Headspace because the vapour phase above the sample in a sealed vial is brought into thermodynamic equilibrium with the sample prior to analysis. Volatile organic compounds are thermally extracted from the sample matrix by Static Headspace sampling and introduced directly to a thick-film megabore GC column. Compounds eluting from the GC column are introduced into a Mass Spectrometer and target compounds are identified by a comparison of both retention times and mass spectra with those of reference standards stored in a computerized data system. Reference retention times and mass spectra for analytes are obtained by the measurement of calibration standards under precisely the same extraction conditions used for samples. The concentration of each identified component is measured by relating the response of the sample quantitation mass ions to the response of the quantitation mass ions of a series of external standards. A deuterated surrogate recovery standard is added to provide information on matrix effects. | 0.005 | Y | Y | Y | 381 | 9/28/1998 3:08:00 PM | 8/14/2024 | 8/16/2024 4:05:21 PM | 1996-09-16 | 6/3/1998 6:51:45 PM | MPAK |  | Concentration | mg/L | Y | (fl. conc.) |  | trans1;3-Dichloropropene (fl. conc.) | 10061-02-6 | Volatile Organic Compounds (VOCs); | Volatile Organic Compounds (VOCs); EMS code: C028 | Chemical |  | LAB
C028 | mg/L | trans1;3-Dichloropropene | X384 | Purge and trap GC/MS | The sample is analyzed by purge and trap gas chromatography with detection by mass spectrometry. The procedure involves purging the volatile compounds from the sample with an inert gas, and trapping them on a solid sorbent. When purging of the sample is complete, the trap is heated and the compounds are transferred to the gas chromatographic column. Analysis is then accomplished by separation of the components by gas chromatography with detection by mass spectrometry. | 0.001 | Y | Y | Y | 112 | 8/5/2004 10:31:00 AM | 7/26/2012 7:12:00 AM | 9/26/2012 1:07:08 PM | 1996-04-12 | 6/3/1998 6:51:34 PM | EMS |  | Concentration | mg/L | Y | (fl. conc.) |  | trans1;3-Dichloropropene (fl. conc.) | 10061-02-6 | Volatile Organic Compounds (VOCs); | Volatile Organic Compounds (VOCs); EMS code: C028 | Chemical |  | LAB
C028 | mg/L | trans1;3-Dichloropropene | MS01 | Purge&Trap GC/MS (VOC-W) | Purge&Trap GC/MS (VOC-W) | 1.1 | Y | Y | Y | 378 | 4/15/2002 | 6/1/2015 12:01:00 AM | 1/24/2016 3:03:56 PM | 1996-04-12 | 3/18/2003 2:49:40 PM | EMS | UMULLER | Concentration | mg/L | Y | (fl. conc.) |  | trans1;3-Dichloropropene (fl. conc.) | 10061-02-6 | Volatile Organic Compounds (VOCs); | Volatile Organic Compounds (VOCs); EMS code: C028 | Chemical |  | LAB
C028 | ug/g | trans1;3-Dichloropropene | FS04 | Static Headspace:HRGC/LRMS | Static Headspace anaylsis is a method for the determination of volatile compounds in liquids and/or solids by analysing the vapour phase of a sample in a closed system; in this case a sealed vial. Static Headspace is sometimes referred to as Equilibrium Headspace because the vapour phase above the sample in a sealed vial is brought into thermodynamic equilibrium with the sample prior to analysis. Volatile organic compounds are thermally extracted from the sample matrix by Static Headspace sampling and introduced directly to a thick-film megabore GC column. Compounds eluting from the GC column are introduced into a Mass Spectrometer and target compounds are identified by a comparison of both retention times and mass spectra with those of reference standards stored in a computerized data system. Reference retention times and mass spectra for analytes are obtained by the measurement of calibration standards under precisely the same extraction conditions used for samples. The concentration of each identified component is measured by relating the response of the sample quantitation mass ions to the response of the quantitation mass ions of a series of external standards. A deuterated surrogate recovery standard is added to provide information on matrix effects. | 0.05 | Y | Y | Y | 44 | 6/30/1999 4:41:00 PM | 5/22/2015 | 7/3/2015 3:19:04 PM | 1996-09-16 | 6/3/1998 6:51:45 PM | MPAK |  | MassPerMass | mg/kg | Y | (mass conc.) |  | trans1;3-Dichloropropene (mass conc.) | 10061-02-6 | Volatile Organic Compounds (VOCs); | Volatile Organic Compounds (VOCs); EMS code: C028 | Chemical |  | LAB
C028 | mg/L | trans1;3-Dichloropropene | MS04 | Static Headspace GC/MS (Water) | (Water) A static headspace procedure is used as a means of driving the analytes out of the matrix into a headspace. Once equilibrium is achieved, a fixed volume of headspeace is taken for subsequent determination by Gas Chromatography - Mass Spectroscopy (GC/MS) using Selected Ion Monitoring. | 0.001 | Y | Y | Y | 69 | 8/13/2009 | 6/28/2024 12:29:00 AM | 7/9/2024 3:31:02 PM | 2009-12-23 |  | RDGIBSON |  | Concentration | mg/L | Y | (fl. conc.) |  | trans1;3-Dichloropropene (fl. conc.) | 10061-02-6 | Volatile Organic Compounds (VOCs); | Volatile Organic Compounds (VOCs); EMS code: C028 | Chemical |  | LAB
C029 | ug/g | cis-1;2-DiClethylene | FS04 | Static Headspace:HRGC/LRMS | Static Headspace anaylsis is a method for the determination of volatile compounds in liquids and/or solids by analysing the vapour phase of a sample in a closed system; in this case a sealed vial. Static Headspace is sometimes referred to as Equilibrium Headspace because the vapour phase above the sample in a sealed vial is brought into thermodynamic equilibrium with the sample prior to analysis. Volatile organic compounds are thermally extracted from the sample matrix by Static Headspace sampling and introduced directly to a thick-film megabore GC column. Compounds eluting from the GC column are introduced into a Mass Spectrometer and target compounds are identified by a comparison of both retention times and mass spectra with those of reference standards stored in a computerized data system. Reference retention times and mass spectra for analytes are obtained by the measurement of calibration standards under precisely the same extraction conditions used for samples. The concentration of each identified component is measured by relating the response of the sample quantitation mass ions to the response of the quantitation mass ions of a series of external standards. A deuterated surrogate recovery standard is added to provide information on matrix effects. | 0.05 | Y | Y | Y | 43 | 5/21/2015 | 5/22/2015 | 7/3/2015 3:19:04 PM | 2015-07-03 |  | RDGIBSON |  | MassPerMass | mg/kg | Y | (mass conc.) |  | cis-1;2-DiClethylene (mass conc.) | 156-59-2 | Volatile Organic Compounds (VOCs); | Volatile Organic Compounds (VOCs); EMS code: C029 | Chemical |  | LAB
C029 | ug/g | cis-1;2-DiClethylene | MS02 | Purge&Trap GC/MS (VOC-S) | Purge&Trap GC/MS (VOC-S) | 0.05 | Y | Y | Y | 15 | 1/14/2022 | 2/4/2022 | 1/11/2024 1:07:07 PM | 2022-07-27 |  | EMS |  | MassPerMass | mg/kg | Y | (mass conc.) |  | cis-1;2-DiClethylene (mass conc.) | 156-59-2 | Volatile Organic Compounds (VOCs); | Volatile Organic Compounds (VOCs); EMS code: C029 | Chemical |  | LAB
C029 | mg/L | cis-1;2-DiClethylene | F007 | Static Headspace:HRGC/LRMS | Static headspace analysis is a method for the determination of volatile compounds in liquids and/or solids by analysing the vapour phase of a sample in a closed system; in this case a sealed vial. Static Headspace is sometimes referred to as Equilibrium Headspace because the vapour phase above the sample in a sealed vial is brought into thermodynamic equilibrium with the sample prior to analysis. Volatile organic compounds are thermally extracted from the sample matrix by Static Headspace sampling and introduced directly to a thick-film megabore GC column. Compounds eluting from the GC column are introduced into a Mass Spectrometer and target compounds are identified by a comparison of both retention times and mass spectra with those of reference standards stored in a computerized data system. Reference retention times and mass spectra for analytes are obtained by the measurement of calibration standards under precisely the same extraction conditions used for samples. The concentration of each identified component is measured by relating the response of the sample quantitation mass ions to the response of the quantitation mass ions of a series of external standards. A deuterated surrogate recovery standard is added to provide information on matrix effects. | 0.001 | Y | Y | Y | 329 | 4/15/2011 | 8/14/2024 | 8/16/2024 4:05:21 PM | 2012-03-28 |  | RDGIBSON |  | Concentration | mg/L | Y | (fl. conc.) |  | cis-1;2-DiClethylene (fl. conc.) | 156-59-2 | Volatile Organic Compounds (VOCs); | Volatile Organic Compounds (VOCs); EMS code: C029 | Chemical |  | LAB
C029 | mg/L | cis-1;2-DiClethylene | X384 | Purge and trap GC/MS | The sample is analyzed by purge and trap gas chromatography with detection by mass spectrometry. The procedure involves purging the volatile compounds from the sample with an inert gas, and trapping them on a solid sorbent. When purging of the sample is complete, the trap is heated and the compounds are transferred to the gas chromatographic column. Analysis is then accomplished by separation of the components by gas chromatography with detection by mass spectrometry. | 0.001 | Y | Y | Y | 69 |  |  | 1/16/2007 2:58:21 PM | 1996-04-12 | 6/3/1998 6:51:45 PM | EMS |  | Concentration | mg/L | Y | (fl. conc.) |  | cis-1;2-DiClethylene (fl. conc.) | 156-59-2 | Volatile Organic Compounds (VOCs); | Volatile Organic Compounds (VOCs); EMS code: C029 | Chemical |  | LAB
C029 | mg/L | cis-1;2-DiClethylene | MS01 | Purge&Trap GC/MS (VOC-W) | Purge&Trap GC/MS (VOC-W) |  | Y | Y | Y | 89 | 4/15/2002 | 4/1/2011 | 7/9/2012 2:09:47 PM | 2003-03-18 | 3/18/2003 2:49:40 PM | UMULLER | UMULLER | Concentration | mg/L | Y | (fl. conc.) |  | cis-1;2-DiClethylene (fl. conc.) | 156-59-2 | Volatile Organic Compounds (VOCs); | Volatile Organic Compounds (VOCs); EMS code: C029 | Chemical |  | LAB
C029 | mg/L | cis-1;2-DiClethylene | MS04 | Static Headspace GC/MS (Water) | (Water) A static headspace procedure is used as a means of driving the analytes out of the matrix into a headspace. Once equilibrium is achieved, a fixed volume of headspeace is taken for subsequent determination by Gas Chromatography - Mass Spectroscopy (GC/MS) using Selected Ion Monitoring. | 0.0005 | Y | Y | Y | 51 | 8/13/2009 | 6/28/2024 12:29:00 AM | 7/9/2024 3:31:02 PM | 2010-01-13 | 1/13/2010 10:44:05 AM | RDGIBSON | RDGIBSON | Concentration | mg/L | Y | (fl. conc.) |  | cis-1;2-DiClethylene (fl. conc.) | 156-59-2 | Volatile Organic Compounds (VOCs); | Volatile Organic Compounds (VOCs); EMS code: C029 | Chemical |  | LAB
C030 | ug/g | cis-Nonachlor | E8081 | EPA 8081 | See methods EPA 3540, 3630, 3660, 8081 | 0.001 | Y | Y | Y | 9 | 3/18/2020 | 3/23/2020 | 2/22/2023 12:51:03 PM | 2022-07-29 |  | EMS |  | MassPerMass | mg/kg | Y | (mass conc.) |  | cis-Nonachlor (mass conc.) |  | Pesticides; | Pesticides; EMS code: C030 | Chemical |  | LAB
C030 | mg/L | cis-Nonachlor | MW03 | OCP1-LL-SF-ECD-VA | This analysis is carried out using procedures adapted from "Test Methods for Evaluating Solid Waste" SW-846, Methods 3510, 3630, 3660 and 8081, published by the United States Environmental Protection Agency (EPA). The procedure involves extraction of the entire water sample with dichloromethane. The extract then undergoes one or more of the following clean-up procedures (if required): silica gel clean-up and/or sulphur clean-up. The final extract is analyzed by dual capillary column gas chromatography with electron capture detection (GC/ECD). | 5e-05 | Y | Y | Y | 10 | 9/20/2016 | 9/21/2016 | 8/15/2019 2:42:34 PM | 2019-07-08 |  | KATFRASE |  | Concentration | mg/L | Y | (fl. conc.) |  | cis-Nonachlor (fl. conc.) |  | Pesticides; | Pesticides; EMS code: C030 | Chemical |  | LAB
C032 | mg/L | Chloroform | F007 | Static Headspace:HRGC/LRMS | Static headspace analysis is a method for the determination of volatile compounds in liquids and/or solids by analysing the vapour phase of a sample in a closed system; in this case a sealed vial. Static Headspace is sometimes referred to as Equilibrium Headspace because the vapour phase above the sample in a sealed vial is brought into thermodynamic equilibrium with the sample prior to analysis. Volatile organic compounds are thermally extracted from the sample matrix by Static Headspace sampling and introduced directly to a thick-film megabore GC column. Compounds eluting from the GC column are introduced into a Mass Spectrometer and target compounds are identified by a comparison of both retention times and mass spectra with those of reference standards stored in a computerized data system. Reference retention times and mass spectra for analytes are obtained by the measurement of calibration standards under precisely the same extraction conditions used for samples. The concentration of each identified component is measured by relating the response of the sample quantitation mass ions to the response of the quantitation mass ions of a series of external standards. A deuterated surrogate recovery standard is added to provide information on matrix effects. | 0.005 | Y | Y | Y | 381 | 9/28/1998 3:08:00 PM | 8/14/2024 | 8/16/2024 4:05:21 PM | 1996-09-16 | 6/3/1998 6:51:44 PM | MPAK |  | Concentration | mg/L | Y | (fl. conc.) |  | Chloroform (fl. conc.) | 67-66-3 | Volatile Organic Compounds (VOCs); | Volatile Organic Compounds (VOCs); EMS code: C032 | Chemical |  | LAB
C032 | mg/L | Chloroform | X384 | Purge and trap GC/MS | The sample is analyzed by purge and trap gas chromatography with detection by mass spectrometry. The procedure involves purging the volatile compounds from the sample with an inert gas, and trapping them on a solid sorbent. When purging of the sample is complete, the trap is heated and the compounds are transferred to the gas chromatographic column. Analysis is then accomplished by separation of the components by gas chromatography with detection by mass spectrometry. | 0.001 | Y | Y | Y | 112 | 8/5/2004 10:31:00 AM | 7/26/2012 7:12:00 AM | 9/26/2012 1:07:08 PM | 1996-04-12 | 6/3/1998 6:51:45 PM | EMS |  | Concentration | mg/L | Y | (fl. conc.) |  | Chloroform (fl. conc.) | 67-66-3 | Volatile Organic Compounds (VOCs); | Volatile Organic Compounds (VOCs); EMS code: C032 | Chemical |  | LAB
C032 | mg/L | Chloroform | MS04 | Static Headspace GC/MS (Water) | (Water) A static headspace procedure is used as a means of driving the analytes out of the matrix into a headspace. Once equilibrium is achieved, a fixed volume of headspeace is taken for subsequent determination by Gas Chromatography - Mass Spectroscopy (GC/MS) using Selected Ion Monitoring. | 0.001 | Y | Y | Y | 99 | 8/13/2009 | 6/28/2024 12:29:00 AM | 7/23/2024 11:41:49 AM | 2009-12-18 |  | RDGIBSON |  | Concentration | mg/L | Y | (fl. conc.) |  | Chloroform (fl. conc.) | 67-66-3 | Volatile Organic Compounds (VOCs); | Volatile Organic Compounds (VOCs); EMS code: C032 | Chemical |  | LAB
C032 | ug/g | Chloroform | FS04 | Static Headspace:HRGC/LRMS | Static Headspace anaylsis is a method for the determination of volatile compounds in liquids and/or solids by analysing the vapour phase of a sample in a closed system; in this case a sealed vial. Static Headspace is sometimes referred to as Equilibrium Headspace because the vapour phase above the sample in a sealed vial is brought into thermodynamic equilibrium with the sample prior to analysis. Volatile organic compounds are thermally extracted from the sample matrix by Static Headspace sampling and introduced directly to a thick-film megabore GC column. Compounds eluting from the GC column are introduced into a Mass Spectrometer and target compounds are identified by a comparison of both retention times and mass spectra with those of reference standards stored in a computerized data system. Reference retention times and mass spectra for analytes are obtained by the measurement of calibration standards under precisely the same extraction conditions used for samples. The concentration of each identified component is measured by relating the response of the sample quantitation mass ions to the response of the quantitation mass ions of a series of external standards. A deuterated surrogate recovery standard is added to provide information on matrix effects. | 0.05 | Y | Y | Y | 44 | 6/30/1999 4:41:00 PM | 5/22/2015 | 7/3/2015 3:19:04 PM | 1996-09-16 | 6/3/1998 6:51:44 PM | MPAK |  | MassPerMass | mg/kg | Y | (mass conc.) |  | Chloroform (mass conc.) | 67-66-3 | Volatile Organic Compounds (VOCs); | Volatile Organic Compounds (VOCs); EMS code: C032 | Chemical |  | LAB
C032 | ug/g | Chloroform | MS02 | Purge&Trap GC/MS (VOC-S) | Purge&Trap GC/MS (VOC-S) | 0.002 | Y | Y | Y | 27 | 11/19/2004 | 10/28/2022 5:04:00 PM | 1/11/2024 1:07:07 PM | 1996-04-12 | 6/3/1998 6:51:34 PM | EMS |  | MassPerMass | mg/kg | Y | (mass conc.) |  | Chloroform (mass conc.) | 67-66-3 | Volatile Organic Compounds (VOCs); | Volatile Organic Compounds (VOCs); EMS code: C032 | Chemical |  | LAB
C032 | mg/L | Chloroform | X252 | Iso-octaneExt:GC/ECD | Iso-octaneExt:GC/ECD | 0.0005 | Y | Y | Y | 11 |  |  | 10/23/1996 11:56:37 PM | 1996-04-12 | 6/3/1998 6:51:34 PM | EMS |  | Concentration | mg/L | Y | (fl. conc.) |  | Chloroform (fl. conc.) | 67-66-3 | Volatile Organic Compounds (VOCs); | Volatile Organic Compounds (VOCs); EMS code: C032 | Chemical |  | LAB
C032 | mg/L | Chloroform | E611 | E611S: EPA 8260D (mod) | E611S: Dimethyl selenide and dimethyl diselenide are analyzed by static headspace GC-MS. Samples are prepared in headspace vials and are heated and agitated on the headspace autosampler; causing VOCs to partition between the aqueous phase and the headspace in accordance with Henry`s law. | 0.001 | Y | Y | Y | 1 | 12/14/2023 | 12/14/2023 | 1/16/2024 8:06:04 AM | 2024-01-15 | 1/15/2024 2:17:10 PM | EMS | EMS | Concentration | mg/L | Y | (fl. conc.) |  | Chloroform (fl. conc.) | 67-66-3 | Volatile Organic Compounds (VOCs); | Volatile Organic Compounds (VOCs); EMS code: C032 | Chemical |  | LAB
C032 | mg/L | Chloroform | P016 | Trihalomethanes | Trihalomethanes | 0.001 | Y | Y | Y | 68 | 3/24/2023 10:24:00 PM | 12/17/2023 11:36:00 AM | 12/29/2023 9:48:33 AM | 1996-04-12 | 6/3/1998 6:51:45 PM | EMS |  | Concentration | mg/L | Y | (fl. conc.) |  | Chloroform (fl. conc.) | 67-66-3 | Volatile Organic Compounds (VOCs); | Volatile Organic Compounds (VOCs); EMS code: C032 | Chemical |  | LAB
C032 | mg/L | Chloroform | MS01 | Purge&Trap GC/MS (VOC-W) | Purge&Trap GC/MS (VOC-W) | 0.4 | Y | Y | Y | 548 | 4/15/2002 | 6/4/2024 10:24:00 AM | 6/13/2024 5:01:27 PM | 1996-04-12 | 3/18/2003 2:49:40 PM | EMS | UMULLER | Concentration | mg/L | Y | (fl. conc.) |  | Chloroform (fl. conc.) | 67-66-3 | Volatile Organic Compounds (VOCs); | Volatile Organic Compounds (VOCs); EMS code: C032 | Chemical |  | LAB
C033 | mg/L | Chlorodibromomethane | MS04 | Static Headspace GC/MS (Water) | (Water) A static headspace procedure is used as a means of driving the analytes out of the matrix into a headspace. Once equilibrium is achieved, a fixed volume of headspeace is taken for subsequent determination by Gas Chromatography - Mass Spectroscopy (GC/MS) using Selected Ion Monitoring. | 0.001 | Y | Y | Y | 75 | 8/13/2009 | 6/28/2024 12:29:00 AM | 7/9/2024 3:31:02 PM | 2010-01-13 | 1/13/2010 1:00:13 PM | RDGIBSON | RDGIBSON | Concentration | mg/L | Y | (fl. conc.) |  | Chlorodibromomethane (fl. conc.) | 124-48-1 | Volatile Organic Compounds (VOCs); | Volatile Organic Compounds (VOCs); EMS code: C033 | Chemical |  | LAB
C033 | ug/g | Chlorodibromomethane | FS04 | Static Headspace:HRGC/LRMS | Static Headspace anaylsis is a method for the determination of volatile compounds in liquids and/or solids by analysing the vapour phase of a sample in a closed system; in this case a sealed vial. Static Headspace is sometimes referred to as Equilibrium Headspace because the vapour phase above the sample in a sealed vial is brought into thermodynamic equilibrium with the sample prior to analysis. Volatile organic compounds are thermally extracted from the sample matrix by Static Headspace sampling and introduced directly to a thick-film megabore GC column. Compounds eluting from the GC column are introduced into a Mass Spectrometer and target compounds are identified by a comparison of both retention times and mass spectra with those of reference standards stored in a computerized data system. Reference retention times and mass spectra for analytes are obtained by the measurement of calibration standards under precisely the same extraction conditions used for samples. The concentration of each identified component is measured by relating the response of the sample quantitation mass ions to the response of the quantitation mass ions of a series of external standards. A deuterated surrogate recovery standard is added to provide information on matrix effects. | 0.05 | Y | Y | Y | 44 | 6/30/1999 4:41:00 PM | 5/22/2015 | 7/3/2015 3:19:04 PM | 1996-09-16 | 6/3/1998 6:51:44 PM | MPAK |  | MassPerMass | mg/kg | Y | (mass conc.) |  | Chlorodibromomethane (mass conc.) | 124-48-1 | Volatile Organic Compounds (VOCs); | Volatile Organic Compounds (VOCs); EMS code: C033 | Chemical |  | LAB
C033 | mg/L | Chlorodibromomethane | F007 | Static Headspace:HRGC/LRMS | Static headspace analysis is a method for the determination of volatile compounds in liquids and/or solids by analysing the vapour phase of a sample in a closed system; in this case a sealed vial. Static Headspace is sometimes referred to as Equilibrium Headspace because the vapour phase above the sample in a sealed vial is brought into thermodynamic equilibrium with the sample prior to analysis. Volatile organic compounds are thermally extracted from the sample matrix by Static Headspace sampling and introduced directly to a thick-film megabore GC column. Compounds eluting from the GC column are introduced into a Mass Spectrometer and target compounds are identified by a comparison of both retention times and mass spectra with those of reference standards stored in a computerized data system. Reference retention times and mass spectra for analytes are obtained by the measurement of calibration standards under precisely the same extraction conditions used for samples. The concentration of each identified component is measured by relating the response of the sample quantitation mass ions to the response of the quantitation mass ions of a series of external standards. A deuterated surrogate recovery standard is added to provide information on matrix effects. | 0.005 | Y | Y | Y | 381 | 9/28/1998 3:08:00 PM | 8/14/2024 | 8/16/2024 4:05:21 PM | 1996-09-16 | 6/3/1998 6:51:44 PM | MPAK |  | Concentration | mg/L | Y | (fl. conc.) |  | Chlorodibromomethane (fl. conc.) | 124-48-1 | Volatile Organic Compounds (VOCs); | Volatile Organic Compounds (VOCs); EMS code: C033 | Chemical |  | LAB
C033 | mg/L | Chlorodibromomethane | P016 | Trihalomethanes | Trihalomethanes | 0.001 | Y | Y | Y | 54 |  |  | 10/24/1996 1:43:46 PM | 1996-04-12 | 6/3/1998 6:51:45 PM | EMS |  | Concentration | mg/L | Y | (fl. conc.) |  | Chlorodibromomethane (fl. conc.) | 124-48-1 | Volatile Organic Compounds (VOCs); | Volatile Organic Compounds (VOCs); EMS code: C033 | Chemical |  | LAB
C033 | ug/g | Chlorodibromomethane | MS02 | Purge&Trap GC/MS (VOC-S) | Purge&Trap GC/MS (VOC-S) | 0.002 | Y | Y | Y | 26 | 11/19/2004 | 2/4/2022 | 1/11/2024 1:07:07 PM | 1996-04-12 | 6/3/1998 6:51:34 PM | EMS |  | MassPerMass | mg/kg | Y | (mass conc.) |  | Chlorodibromomethane (mass conc.) | 124-48-1 | Volatile Organic Compounds (VOCs); | Volatile Organic Compounds (VOCs); EMS code: C033 | Chemical |  | LAB
C033 | mg/L | Chlorodibromomethane | E611 | E611S: EPA 8260D (mod) | E611S: Dimethyl selenide and dimethyl diselenide are analyzed by static headspace GC-MS. Samples are prepared in headspace vials and are heated and agitated on the headspace autosampler; causing VOCs to partition between the aqueous phase and the headspace in accordance with Henry`s law. | 0.001 | Y | Y | Y | 1 | 12/14/2023 | 12/14/2023 | 1/16/2024 8:06:04 AM | 2024-01-15 | 1/15/2024 2:17:10 PM | EMS | EMS | Concentration | mg/L | Y | (fl. conc.) |  | Chlorodibromomethane (fl. conc.) | 124-48-1 | Volatile Organic Compounds (VOCs); | Volatile Organic Compounds (VOCs); EMS code: C033 | Chemical |  | LAB
C033 | mg/L | Chlorodibromomethane | X384 | Purge and trap GC/MS | The sample is analyzed by purge and trap gas chromatography with detection by mass spectrometry. The procedure involves purging the volatile compounds from the sample with an inert gas, and trapping them on a solid sorbent. When purging of the sample is complete, the trap is heated and the compounds are transferred to the gas chromatographic column. Analysis is then accomplished by separation of the components by gas chromatography with detection by mass spectrometry. | 0.001 | Y | Y | Y | 112 | 8/5/2004 10:31:00 AM | 7/26/2012 7:12:00 AM | 9/26/2012 1:07:08 PM | 1996-04-12 | 6/3/1998 6:51:34 PM | EMS |  | Concentration | mg/L | Y | (fl. conc.) |  | Chlorodibromomethane (fl. conc.) | 124-48-1 | Volatile Organic Compounds (VOCs); | Volatile Organic Compounds (VOCs); EMS code: C033 | Chemical |  | LAB
C033 | mg/L | Chlorodibromomethane | MS01 | Purge&Trap GC/MS (VOC-W) | Purge&Trap GC/MS (VOC-W) | 0.4 | Y | Y | Y | 163 | 4/15/2002 | 4/1/2011 | 7/9/2012 2:09:47 PM | 1996-04-12 | 3/18/2003 2:49:40 PM | EMS | UMULLER | Concentration | mg/L | Y | (fl. conc.) |  | Chlorodibromomethane (fl. conc.) | 124-48-1 | Volatile Organic Compounds (VOCs); | Volatile Organic Compounds (VOCs); EMS code: C033 | Chemical |  | LAB
C034 | ug/g | Carbon Tetrachloride | FS04 | Static Headspace:HRGC/LRMS | Static Headspace anaylsis is a method for the determination of volatile compounds in liquids and/or solids by analysing the vapour phase of a sample in a closed system; in this case a sealed vial. Static Headspace is sometimes referred to as Equilibrium Headspace because the vapour phase above the sample in a sealed vial is brought into thermodynamic equilibrium with the sample prior to analysis. Volatile organic compounds are thermally extracted from the sample matrix by Static Headspace sampling and introduced directly to a thick-film megabore GC column. Compounds eluting from the GC column are introduced into a Mass Spectrometer and target compounds are identified by a comparison of both retention times and mass spectra with those of reference standards stored in a computerized data system. Reference retention times and mass spectra for analytes are obtained by the measurement of calibration standards under precisely the same extraction conditions used for samples. The concentration of each identified component is measured by relating the response of the sample quantitation mass ions to the response of the quantitation mass ions of a series of external standards. A deuterated surrogate recovery standard is added to provide information on matrix effects. | 0.05 | Y | Y | Y | 44 | 6/30/1999 4:41:00 PM | 5/22/2015 | 7/3/2015 3:19:04 PM | 1996-09-16 | 6/3/1998 6:51:45 PM | MPAK |  | MassPerMass | mg/kg | Y | (mass conc.) |  | Carbon Tetrachloride (mass conc.) | 56-23-5 | Volatile Organic Compounds (VOCs); | Volatile Organic Compounds (VOCs); EMS code: C034 | Chemical |  | LAB
C034 | ug/g | Carbon Tetrachloride | MS02 | Purge&Trap GC/MS (VOC-S) | Purge&Trap GC/MS (VOC-S) | 0.002 | Y | Y | Y | 27 | 11/19/2004 | 10/28/2022 5:04:00 PM | 1/11/2024 1:07:07 PM | 1996-04-12 | 6/3/1998 6:51:45 PM | EMS |  | MassPerMass | mg/kg | Y | (mass conc.) |  | Carbon Tetrachloride (mass conc.) | 56-23-5 | Volatile Organic Compounds (VOCs); | Volatile Organic Compounds (VOCs); EMS code: C034 | Chemical |  | LAB
C034 | mg/L | Carbon Tetrachloride | MS04 | Static Headspace GC/MS (Water) | (Water) A static headspace procedure is used as a means of driving the analytes out of the matrix into a headspace. Once equilibrium is achieved, a fixed volume of headspeace is taken for subsequent determination by Gas Chromatography - Mass Spectroscopy (GC/MS) using Selected Ion Monitoring. | 0.003 | Y | Y | Y | 70 | 8/13/2009 | 6/28/2024 12:29:00 AM | 7/9/2024 3:31:02 PM | 2009-12-18 |  | RDGIBSON |  | Concentration | mg/L | Y | (fl. conc.) |  | Carbon Tetrachloride (fl. conc.) | 56-23-5 | Volatile Organic Compounds (VOCs); | Volatile Organic Compounds (VOCs); EMS code: C034 | Chemical |  | LAB
C034 | mg/L | Carbon Tetrachloride | F007 | Static Headspace:HRGC/LRMS | Static headspace analysis is a method for the determination of volatile compounds in liquids and/or solids by analysing the vapour phase of a sample in a closed system; in this case a sealed vial. Static Headspace is sometimes referred to as Equilibrium Headspace because the vapour phase above the sample in a sealed vial is brought into thermodynamic equilibrium with the sample prior to analysis. Volatile organic compounds are thermally extracted from the sample matrix by Static Headspace sampling and introduced directly to a thick-film megabore GC column. Compounds eluting from the GC column are introduced into a Mass Spectrometer and target compounds are identified by a comparison of both retention times and mass spectra with those of reference standards stored in a computerized data system. Reference retention times and mass spectra for analytes are obtained by the measurement of calibration standards under precisely the same extraction conditions used for samples. The concentration of each identified component is measured by relating the response of the sample quantitation mass ions to the response of the quantitation mass ions of a series of external standards. A deuterated surrogate recovery standard is added to provide information on matrix effects. | 0.005 | Y | Y | Y | 380 | 9/28/1998 3:08:00 PM | 8/14/2024 | 8/16/2024 4:05:21 PM | 1996-09-16 | 6/3/1998 6:51:45 PM | MPAK |  | Concentration | mg/L | Y | (fl. conc.) |  | Carbon Tetrachloride (fl. conc.) | 56-23-5 | Volatile Organic Compounds (VOCs); | Volatile Organic Compounds (VOCs); EMS code: C034 | Chemical |  | LAB
C034 | mg/L | Carbon Tetrachloride | MS01 | Purge&Trap GC/MS (VOC-W) | Purge&Trap GC/MS (VOC-W) | 0.3 | Y | Y | Y | 631 | 4/15/2002 | 6/4/2024 10:24:00 AM | 6/13/2024 5:01:27 PM | 1996-04-12 | 3/18/2003 2:49:40 PM | EMS | UMULLER | Concentration | mg/L | Y | (fl. conc.) |  | Carbon Tetrachloride (fl. conc.) | 56-23-5 | Volatile Organic Compounds (VOCs); | Volatile Organic Compounds (VOCs); EMS code: C034 | Chemical |  | LAB
C035 | mg/L | 2-Chlorophenol | F028 | Acid,Extr.MeCl,HRGC/LRMS:TIM,SIM | A measured volume of sample is acidified and extracted with Methylene Chloride. For samples which have a high organic content that potentially produce chromatographic interferences an initial pre-extraction under basic conditions is performed. The sample extract is then evaporated to small volume and analysed by High Resolution Gas Chromatography / Low Resolution Mass Spectrometry (HRGC/LRMS). LRMS data is acquired in Total Ion Mode (TIM) and quantitation is performed using a single mass ion. Positive Identification of the various target compounds is made by comparison of absolute retention times to those of standards combined with computer based mass spectral mass/intensity library searching and matching routines. | 0.005 | Y | Y | Y | 104 | 4/10/1997 3:18:00 PM | 11/1/2022 4:39:00 AM | 4/21/2023 4:26:45 PM | 1997-01-14 | 6/3/1998 6:51:46 PM | MPAK |  | Concentration | mg/L | Y | (fl. conc.) |  | 2-Chlorophenol (fl. conc.) |  | Organics; | Organics; EMS code: C035 | Chemical |  | LAB
C035 | ug/g | 2-Chlorophenol | 1332 | Phenolics by Tumbler/GC-MS | Phenolics by Tumbler/GC-MS | 0.02 | Y | Y | Y | 30 | 9/9/2019 | 9/22/2022 8:16:00 PM | 1/11/2024 1:07:07 PM | 2019-09-30 |  | KATFRASE |  | MassPerMass | mg/kg | Y | (mass conc.) |  | 2-Chlorophenol (mass conc.) |  | Organics; | Organics; EMS code: C035 | Chemical |  | LAB
C035 | mg/L | 2-Chlorophenol | X268 | CH2Cl2 ext:GC/MS | CH2Cl2 ext:GC/MS | 0.0001 | Y | Y | Y | 26 | 12/28/2021 | 6/7/2024 | 6/26/2024 4:28:10 PM | 2012-06-25 |  | RDGIBSON |  | Concentration | mg/L | Y | (fl. conc.) |  | 2-Chlorophenol (fl. conc.) |  | Organics; | Organics; EMS code: C035 | Chemical |  | LAB
C036 | mg/L | 4-Chloro-3-Methylphenol | F028 | Acid,Extr.MeCl,HRGC/LRMS:TIM,SIM | A measured volume of sample is acidified and extracted with Methylene Chloride. For samples which have a high organic content that potentially produce chromatographic interferences an initial pre-extraction under basic conditions is performed. The sample extract is then evaporated to small volume and analysed by High Resolution Gas Chromatography / Low Resolution Mass Spectrometry (HRGC/LRMS). LRMS data is acquired in Total Ion Mode (TIM) and quantitation is performed using a single mass ion. Positive Identification of the various target compounds is made by comparison of absolute retention times to those of standards combined with computer based mass spectral mass/intensity library searching and matching routines. | 0.005 | Y | Y | Y | 82 | 4/10/1997 3:17:00 PM | 11/1/2022 4:39:00 AM | 4/21/2023 4:26:45 PM | 1997-01-14 | 6/3/1998 6:51:46 PM | MPAK |  | Concentration | mg/L | Y | (fl. conc.) |  | 4-Chloro-3-Methylphenol (fl. conc.) |  | Organics; | Organics; EMS code: C036 | Chemical |  | LAB
C036 | mg/L | 4-Chloro-3-Methylphenol | X268 | CH2Cl2 ext:GC/MS | CH2Cl2 ext:GC/MS | 0.0005 | Y | Y | Y | 26 | 12/28/2021 | 6/7/2024 | 6/26/2024 4:28:10 PM | 2012-06-25 |  | RDGIBSON |  | Concentration | mg/L | Y | (fl. conc.) |  | 4-Chloro-3-Methylphenol (fl. conc.) |  | Organics; | Organics; EMS code: C036 | Chemical |  | LAB
C036 | ug/g | 4-Chloro-3-Methylphenol | 1332 | Phenolics by Tumbler/GC-MS | Phenolics by Tumbler/GC-MS | 0.02 | Y | Y | Y | 30 | 9/9/2019 | 9/22/2022 8:16:00 PM | 1/11/2024 1:07:07 PM | 2019-09-30 |  | KATFRASE |  | MassPerMass | mg/kg | Y | (mass conc.) |  | 4-Chloro-3-Methylphenol (mass conc.) |  | Organics; | Organics; EMS code: C036 | Chemical |  | LAB
C039 | mg/L | 3-hydroxy Carbofuran | P045 | DCM Extr:HPLC | DCM Extr:HPLC | 0.0005 | Y | Y | Y | 1 | 7/29/2005 7:10:00 AM | 7/29/2005 7:10:00 AM | 6/2/2009 12:01:00 PM | 1996-04-12 | 6/3/1998 6:51:41 PM | EMS |  | Concentration | mg/L | Y | (fl. conc.) |  | 3-hydroxy Carbofuran (fl. conc.) |  | Pesticides; | Pesticides; EMS code: C039 | Chemical |  | LAB
C040 | mg/L | Carbofuran | P045 | DCM Extr:HPLC | DCM Extr:HPLC | 0.0005 | Y | Y | Y | 1 | 7/29/2005 7:10:00 AM | 7/29/2005 7:10:00 AM | 6/2/2009 12:01:00 PM | 1996-04-12 | 6/3/1998 6:51:41 PM | EMS |  | Concentration | mg/L | Y | (fl. conc.) |  | Carbofuran (fl. conc.) | 1563-66-2 | Pesticides; | Pesticides; EMS code: C040 | Chemical |  | LAB
C040 | ug/g | Carbofuran | ZN03 | DCM+Ac.Extr:GC/MS | DCM+Ac.Extr:GC/MS EPA 8270E (mod) | 0.05 | Y | Y | Y | 14 | 1/25/2022 | 9/8/2022 | 12/15/2022 10:50:07 AM | 2003-06-16 |  | UMULLER |  | MassPerMass | mg/kg | Y | (mass conc.) |  | Carbofuran (mass conc.) | 1563-66-2 | Pesticides; | Pesticides; EMS code: C040 | Chemical |  | LAB
C040 | mg/L | Carbofuran | E660 | EPA 8270E (mod) | EPA 8270E (mod) Miscellaneous Pesticides by GC-MS | 0.0002 | Y | Y | Y | 68 | 12/7/2021 | 9/12/2022 | 12/8/2022 12:57:05 PM | 2021-12-09 | 12/9/2021 3:22:12 PM | EMS | EMS | Concentration | mg/L | Y | (fl. conc.) |  | Carbofuran (fl. conc.) | 1563-66-2 | Pesticides; | Pesticides; EMS code: C040 | Chemical |  | LAB
C042 | mg/L | 2-Chloroethylvinyl E | MS01 | Purge&Trap GC/MS (VOC-W) | Purge&Trap GC/MS (VOC-W) | 2.9 | Y | Y | Y | 10 |  |  | 3/19/2004 8:18:43 AM | 1996-04-12 | 3/18/2003 2:49:40 PM | EMS | UMULLER | Concentration | mg/L | Y | (fl. conc.) |  | 2-Chloroethylvinyl E (fl. conc.) | 110-75-8 | Organics; | Organics; EMS code: C042 | Chemical |  | LAB
C042 | mg/L | 2-Chloroethylvinyl E | F007 | Static Headspace:HRGC/LRMS | Static headspace analysis is a method for the determination of volatile compounds in liquids and/or solids by analysing the vapour phase of a sample in a closed system; in this case a sealed vial. Static Headspace is sometimes referred to as Equilibrium Headspace because the vapour phase above the sample in a sealed vial is brought into thermodynamic equilibrium with the sample prior to analysis. Volatile organic compounds are thermally extracted from the sample matrix by Static Headspace sampling and introduced directly to a thick-film megabore GC column. Compounds eluting from the GC column are introduced into a Mass Spectrometer and target compounds are identified by a comparison of both retention times and mass spectra with those of reference standards stored in a computerized data system. Reference retention times and mass spectra for analytes are obtained by the measurement of calibration standards under precisely the same extraction conditions used for samples. The concentration of each identified component is measured by relating the response of the sample quantitation mass ions to the response of the quantitation mass ions of a series of external standards. A deuterated surrogate recovery standard is added to provide information on matrix effects. | 0.01 | Y | Y | Y | 13 | 9/28/1998 3:08:00 PM | 12/19/2001 4:41:00 PM | 6/18/2003 2:53:34 PM | 1998-02-17 | 6/3/1998 6:51:47 PM | NPEPPIN |  | Concentration | mg/L | Y | (fl. conc.) |  | 2-Chloroethylvinyl E (fl. conc.) | 110-75-8 | Organics; | Organics; EMS code: C042 | Chemical |  | LAB
C046 | mg/L | Clopyralid | AEWA | H2SO4 acid,CHCl2 extr.florisil. GC/MS | Effluent and water samples are acidified with Sulphuric Acid and extracted with Dichloromethane. After concentration, the extract is methylated using ethereal Diazomethane. Extracts are cleaned up as required using Florisil. The extracts are analyzed for Acid Extractable Herbicides using GC/MS. | 0.0005 | Y | Y | Y | 1 | 7/29/2005 7:10:00 AM | 7/29/2005 7:10:00 AM | 6/2/2009 12:01:00 PM | 2003-06-17 |  | UMULLER |  | Concentration | mg/L | Y | (fl. conc.) |  | Clopyralid (fl. conc.) | 1702-17-6 | Pesticides; | Pesticides; EMS code: C046 | Chemical |  | LAB
C046 | mg/L | Clopyralid | P015 | Phenoxyacid Herbicides by LC-MS/MS | Phenoxyacid Herbicides by LC-MS/MS | 0.0001 | Y | Y | Y | 68 | 12/6/2021 | 9/9/2022 | 12/8/2022 12:57:05 PM | 2021-12-14 |  | EMS |  | Concentration | mg/L | Y | (fl. conc.) |  | Clopyralid (fl. conc.) | 1702-17-6 | Pesticides; | Pesticides; EMS code: C046 | Chemical |  | LAB
C046 | ug/g | Clopyralid | ZN03 | DCM+Ac.Extr:GC/MS | DCM+Ac.Extr:GC/MS EPA 8270E (mod) | 0.001 | Y | Y | Y | 15 | 1/14/2022 | 2/2/2022 | 1/11/2024 1:07:07 PM | 2022-07-27 |  | EMS |  | MassPerMass | mg/kg | Y | (mass conc.) |  | Clopyralid (mass conc.) | 1702-17-6 | Pesticides; | Pesticides; EMS code: C046 | Chemical |  | LAB
C048 | mg/L | 12-Chlorodehydroabietic Acid | RFAS | Separation of resin and fatty acids (RFAs) in water by GC/MS | Scientific Principle The separation of resin and fatty acids (RFAs) in water samples is performed using gas chromatography/mass spectrometry (GC/MS) in the Selected Ion Monitoring (SIM) mode. The separation is based on compound size; boiling point and polarity. Analyte concentration is calculated using the internal standard method. RFA compounds are extracted from water samples using a liquid-liquid solvent extraction with DCM. Extraction occurs after the sample is brought to a pH of ~ 2 and placed in a rotary tumbler (at approx. 40 rpm) for 1 hour. The resulting extract is transferred to a vial; dried with acidified baked sodium sulfate; methylated with diazomethane and concentrated to 1mL under nitrogen prior to analysis. The resulting extracts are vialed and fortified with labeled internal standard compounds prior to analysis by gas chromatography with a (50%-Phenyl)-methylpolysiloxane (DB-17) column and a mass spectrometer detector in SIM Mode. | 0.006 | Y | Y | Y | 1 | 9/17/2021 7:49:00 PM | 9/17/2021 7:49:00 PM | 2/1/2022 2:38:14 PM | 2022-01-19 |  | EMS |  | Concentration | mg/L | Y | (fl. conc.) |  | 12-Chlorodehydroabietic Acid (fl. conc.) | 65310-45-4 | Organics; | Organics; EMS code: C048 | Chemical |  | LAB
C048 | mg/L | 12-Chlorodehydroabietic Acid | E677 | Resins & Fatty Acids are methylated then analyzed by GCMS | EPA 8270E; AB Env. 129; NCASI RA/FA-85.02 (mod) Resins & Fatty Acids are methylated then analyzed by GCMS. Total fatty acids is the sum of arachidic acid; behenic acid; lauric acid; lignoceric acid; linoleic acid; linolenic acid; myristic acid; oleic acid; palmitic acid; and stearic acid. Total resin acids is the sum of abietic acid; 12-chlorodehydroabietic acid; 14-chlorodehydroabietic acid; dehydroabietic acid; 12;14-dichlorodehydroabietic acid; isopimaric acid; levopimaric acid; neoabietic acid; palustric acid; pimaric acid and sandaracopimaric acid. Total resin acids (BC WWQG) is the sum of abietic acid; neoabietic acid; pimaric acid; isopimaric acid; and sandaracopimaric acid. | 0.001 | Y | Y | Y | 12 | 11/28/2022 | 6/15/2024 | 7/8/2024 11:10:02 AM | 2021-12-30 |  | EMS |  | Concentration | mg/L | Y | (fl. conc.) |  | 12-Chlorodehydroabietic Acid (fl. conc.) | 65310-45-4 | Organics; | Organics; EMS code: C048 | Chemical |  | LAB
C048 | mg/L | 12-Chlorodehydroabietic Acid | P030 | Resin Acids | The water sample is acidified and extracted with dichloro-methane. The raw extract is concentrated and derivatized with diazomethane to produce the corresponding methyl ester derivatives. (Suitable substitute derivatives may be produced by other available techniques.) If required, the extracts are cleaned up by Florisil column chromatography. The derivatives are analyzed by gas chromatography with flame ionization detection. | 0.003 | Y | Y | Y | 12 | 8/26/2014 | 8/15/2017 | 8/15/2017 4:37:09 PM | 2015-03-11 |  | RDGIBSON |  | Concentration | mg/L | Y | (fl. conc.) |  | 12-Chlorodehydroabietic Acid (fl. conc.) | 65310-45-4 | Organics; | Organics; EMS code: C048 | Chemical |  | LAB
C048 | mg/L | 12-Chlorodehydroabietic Acid | F009 | pH 9,Diethyl Ether:HRGC/LRMS | Effluent samples are first adjusted to pH 9 and then extracted with Diethyl Ether. The extracts are rotary evaporated and the resulting residues, containing the Resin Acids and other solvent soluble organics are then redissolved in Methanol, transferred to small glass vials, methylated with Diazomethane then evaporated to dryness under Nitrogen. [Methylation Reaction R-COOH(Resin/Fatty Acids) + CH2N2(Diazomethane) = R-COOCH3(Methyl ester) + N2) The residue is re-dissolved in Methyl t-Butyl Ether (MTBE) and injected on a High Resolution Gas Chromatograph/Low Resolution Mass Spectrometer (HRGC/LRMS) system for analysis. LRMS data is acquired in Total Ion Mode (TIM) and quantitation is performed using selected mass ions. Positive identification of the various Resin Acids is made on the basis of a comparison of absolute retention times to those of the external standards and computer based mass spectral mass/intensity library matching routines. | 0.005 | Y | Y | Y | 135 | 12/27/1996 4:16:00 PM | 1/8/2002 12:49:00 PM | 6/7/2024 5:30:18 PM | 1997-01-13 | 6/3/1998 6:51:46 PM | MPAK |  | Concentration | mg/L | Y | (fl. conc.) |  | 12-Chlorodehydroabietic Acid (fl. conc.) | 65310-45-4 | Organics; | Organics; EMS code: C048 | Chemical |  | LAB
C048 | ug/g | 12-Chlorodehydroabietic Acid | FS05 | MeOH,DCM extr:HRGC/LRMS | Sediment samples are extracted by shaking with Methanol in a screw cap centrifuge bottle. The extracts are rotary evaporated and the resulting residues, containing the Resin Acids and other solvent soluble organics, are then redissolved in Methanol, transferred to small glass vials, methylated with Diazomethane then evaporated to dryness under Nitrogen. [Methylation Reaction R-COOH(Resin/Fatty Acids) + CH2N2 (Diazomethane) = R-COOCH3 (Methyl ester) + N2] The residue is re-dissolved in Methyl t-Butyl Ether (MTBE) and injected on a High Resolution Gas Chromatograph / Low Resolution Mass Spectrometer (HRGC/LRMS) system for analysis. LRMS data is acquired in Total Ion Mode (TIM) and quantitation is performed using selected mass ions. Positive identification of the various Resin Acids is made on the basis of a comparison of absolute retention times to those of the external standards and computer based mass spectral mass/intensity library matching routines. | 0.2 | Y | Y | Y | 1 | 9/14/2001 3:29:00 PM | 9/14/2001 3:29:00 PM | 2/17/2003 3:39:55 PM | 1997-01-13 | 6/3/1998 6:51:46 PM | MPAK |  | MassPerMass | mg/kg | Y | (mass conc.) |  | 12-Chlorodehydroabietic Acid (mass conc.) | 65310-45-4 | Organics; | Organics; EMS code: C048 | Chemical |  | LAB
C049 | mg/L | 14-Chlorodehydroabietic Acid | P030 | Resin Acids | The water sample is acidified and extracted with dichloro-methane. The raw extract is concentrated and derivatized with diazomethane to produce the corresponding methyl ester derivatives. (Suitable substitute derivatives may be produced by other available techniques.) If required, the extracts are cleaned up by Florisil column chromatography. The derivatives are analyzed by gas chromatography with flame ionization detection. | 0.003 | Y | Y | Y | 12 | 8/26/2014 | 8/15/2017 | 8/15/2017 4:37:09 PM | 2015-03-11 |  | RDGIBSON |  | Concentration | mg/L | Y | (fl. conc.) |  | 14-Chlorodehydroabietic Acid (fl. conc.) | 65281-76-7 | Organics; | Organics; EMS code: C049 | Chemical |  | LAB
C049 | mg/L | 14-Chlorodehydroabietic Acid | RFAS | Separation of resin and fatty acids (RFAs) in water by GC/MS | Scientific Principle The separation of resin and fatty acids (RFAs) in water samples is performed using gas chromatography/mass spectrometry (GC/MS) in the Selected Ion Monitoring (SIM) mode. The separation is based on compound size; boiling point and polarity. Analyte concentration is calculated using the internal standard method. RFA compounds are extracted from water samples using a liquid-liquid solvent extraction with DCM. Extraction occurs after the sample is brought to a pH of ~ 2 and placed in a rotary tumbler (at approx. 40 rpm) for 1 hour. The resulting extract is transferred to a vial; dried with acidified baked sodium sulfate; methylated with diazomethane and concentrated to 1mL under nitrogen prior to analysis. The resulting extracts are vialed and fortified with labeled internal standard compounds prior to analysis by gas chromatography with a (50%-Phenyl)-methylpolysiloxane (DB-17) column and a mass spectrometer detector in SIM Mode. | 0.006 | Y | Y | Y | 1 | 9/17/2021 7:49:00 PM | 9/17/2021 7:49:00 PM | 2/1/2022 2:38:14 PM | 2022-01-19 |  | EMS |  | Concentration | mg/L | Y | (fl. conc.) |  | 14-Chlorodehydroabietic Acid (fl. conc.) | 65281-76-7 | Organics; | Organics; EMS code: C049 | Chemical |  | LAB
C049 | mg/L | 14-Chlorodehydroabietic Acid | F009 | pH 9,Diethyl Ether:HRGC/LRMS | Effluent samples are first adjusted to pH 9 and then extracted with Diethyl Ether. The extracts are rotary evaporated and the resulting residues, containing the Resin Acids and other solvent soluble organics are then redissolved in Methanol, transferred to small glass vials, methylated with Diazomethane then evaporated to dryness under Nitrogen. [Methylation Reaction R-COOH(Resin/Fatty Acids) + CH2N2(Diazomethane) = R-COOCH3(Methyl ester) + N2) The residue is re-dissolved in Methyl t-Butyl Ether (MTBE) and injected on a High Resolution Gas Chromatograph/Low Resolution Mass Spectrometer (HRGC/LRMS) system for analysis. LRMS data is acquired in Total Ion Mode (TIM) and quantitation is performed using selected mass ions. Positive identification of the various Resin Acids is made on the basis of a comparison of absolute retention times to those of the external standards and computer based mass spectral mass/intensity library matching routines. | 0.005 | Y | Y | Y | 135 | 12/27/1996 4:16:00 PM | 1/8/2002 12:49:00 PM | 6/7/2024 5:30:18 PM | 1997-01-13 | 6/3/1998 6:51:46 PM | MPAK |  | Concentration | mg/L | Y | (fl. conc.) |  | 14-Chlorodehydroabietic Acid (fl. conc.) | 65281-76-7 | Organics; | Organics; EMS code: C049 | Chemical |  | LAB
C049 | ug/g | 14-Chlorodehydroabietic Acid | FS05 | MeOH,DCM extr:HRGC/LRMS | Sediment samples are extracted by shaking with Methanol in a screw cap centrifuge bottle. The extracts are rotary evaporated and the resulting residues, containing the Resin Acids and other solvent soluble organics, are then redissolved in Methanol, transferred to small glass vials, methylated with Diazomethane then evaporated to dryness under Nitrogen. [Methylation Reaction R-COOH(Resin/Fatty Acids) + CH2N2 (Diazomethane) = R-COOCH3 (Methyl ester) + N2] The residue is re-dissolved in Methyl t-Butyl Ether (MTBE) and injected on a High Resolution Gas Chromatograph / Low Resolution Mass Spectrometer (HRGC/LRMS) system for analysis. LRMS data is acquired in Total Ion Mode (TIM) and quantitation is performed using selected mass ions. Positive identification of the various Resin Acids is made on the basis of a comparison of absolute retention times to those of the external standards and computer based mass spectral mass/intensity library matching routines. | 0.2 | Y | Y | Y | 1 | 9/14/2001 3:29:00 PM | 9/14/2001 3:29:00 PM | 2/17/2003 3:39:55 PM | 1997-01-13 | 6/3/1998 6:51:46 PM | MPAK |  | MassPerMass | mg/kg | Y | (mass conc.) |  | 14-Chlorodehydroabietic Acid (mass conc.) | 65281-76-7 | Organics; | Organics; EMS code: C049 | Chemical |  | LAB
C049 | mg/L | 14-Chlorodehydroabietic Acid | E677 | Resins & Fatty Acids are methylated then analyzed by GCMS | EPA 8270E; AB Env. 129; NCASI RA/FA-85.02 (mod) Resins & Fatty Acids are methylated then analyzed by GCMS. Total fatty acids is the sum of arachidic acid; behenic acid; lauric acid; lignoceric acid; linoleic acid; linolenic acid; myristic acid; oleic acid; palmitic acid; and stearic acid. Total resin acids is the sum of abietic acid; 12-chlorodehydroabietic acid; 14-chlorodehydroabietic acid; dehydroabietic acid; 12;14-dichlorodehydroabietic acid; isopimaric acid; levopimaric acid; neoabietic acid; palustric acid; pimaric acid and sandaracopimaric acid. Total resin acids (BC WWQG) is the sum of abietic acid; neoabietic acid; pimaric acid; isopimaric acid; and sandaracopimaric acid. | 0.001 | Y | Y | Y | 12 | 11/28/2022 | 6/15/2024 | 7/8/2024 11:10:02 AM | 2021-12-30 |  | EMS |  | Concentration | mg/L | Y | (fl. conc.) |  | 14-Chlorodehydroabietic Acid (fl. conc.) | 65281-76-7 | Organics; | Organics; EMS code: C049 | Chemical |  | LAB
C050 | ug/g | Chlorodehydroabietic | FS05 | MeOH,DCM extr:HRGC/LRMS | Sediment samples are extracted by shaking with Methanol in a screw cap centrifuge bottle. The extracts are rotary evaporated and the resulting residues, containing the Resin Acids and other solvent soluble organics, are then redissolved in Methanol, transferred to small glass vials, methylated with Diazomethane then evaporated to dryness under Nitrogen. [Methylation Reaction R-COOH(Resin/Fatty Acids) + CH2N2 (Diazomethane) = R-COOCH3 (Methyl ester) + N2] The residue is re-dissolved in Methyl t-Butyl Ether (MTBE) and injected on a High Resolution Gas Chromatograph / Low Resolution Mass Spectrometer (HRGC/LRMS) system for analysis. LRMS data is acquired in Total Ion Mode (TIM) and quantitation is performed using selected mass ions. Positive identification of the various Resin Acids is made on the basis of a comparison of absolute retention times to those of the external standards and computer based mass spectral mass/intensity library matching routines. | 0.2 | Y | Y | Y | 4 | 6/21/2002 12:00:00 PM | 12/3/2002 | 3/27/2004 1:09:19 AM | 1996-09-16 | 6/3/1998 6:51:44 PM | MPAK |  | MassPerMass | mg/kg | Y | (mass conc.) |  | Chlorodehydroabietic (mass conc.) | 65310-45-4, 65281-76-7 | Organics; | Organics; EMS code: C050 | Chemical |  | LAB
C050 | mg/L | Chlorodehydroabietic | P030 | Resin Acids | The water sample is acidified and extracted with dichloro-methane. The raw extract is concentrated and derivatized with diazomethane to produce the corresponding methyl ester derivatives. (Suitable substitute derivatives may be produced by other available techniques.) If required, the extracts are cleaned up by Florisil column chromatography. The derivatives are analyzed by gas chromatography with flame ionization detection. | 0.001 | Y | Y | Y | 66 |  |  | 3/18/1997 2:04:18 PM | 1996-04-12 | 6/3/1998 6:51:34 PM | EMS |  | Concentration | mg/L | Y | (fl. conc.) |  | Chlorodehydroabietic (fl. conc.) | 65310-45-4, 65281-76-7 | Organics; | Organics; EMS code: C050 | Chemical |  | LAB
C050 | mg/L | Chlorodehydroabietic | F009 | pH 9,Diethyl Ether:HRGC/LRMS | Effluent samples are first adjusted to pH 9 and then extracted with Diethyl Ether. The extracts are rotary evaporated and the resulting residues, containing the Resin Acids and other solvent soluble organics are then redissolved in Methanol, transferred to small glass vials, methylated with Diazomethane then evaporated to dryness under Nitrogen. [Methylation Reaction R-COOH(Resin/Fatty Acids) + CH2N2(Diazomethane) = R-COOCH3(Methyl ester) + N2) The residue is re-dissolved in Methyl t-Butyl Ether (MTBE) and injected on a High Resolution Gas Chromatograph/Low Resolution Mass Spectrometer (HRGC/LRMS) system for analysis. LRMS data is acquired in Total Ion Mode (TIM) and quantitation is performed using selected mass ions. Positive identification of the various Resin Acids is made on the basis of a comparison of absolute retention times to those of the external standards and computer based mass spectral mass/intensity library matching routines. | 0.005 | Y | Y | Y | 67 | 1/31/2002 | 5/13/2003 | 4/1/2004 1:08:45 AM | 1996-09-16 | 6/3/1998 6:51:44 PM | MPAK |  | Concentration | mg/L | Y | (fl. conc.) |  | Chlorodehydroabietic (fl. conc.) | 65310-45-4, 65281-76-7 | Organics; | Organics; EMS code: C050 | Chemical |  | LAB
C050 | mg/L | Chlorodehydroabietic | P035 | DCM-DIAZO-FLOR | DCM-DIAZO-FLOR | 0.0005 | Y | Y | Y | 22 |  |  | 10/24/1996 12:44:23 PM | 1996-04-12 | 6/3/1998 6:51:34 PM | EMS |  | Concentration | mg/L | Y | (fl. conc.) |  | Chlorodehydroabietic (fl. conc.) | 65310-45-4, 65281-76-7 | Organics; | Organics; EMS code: C050 | Chemical |  | LAB
C050 | mg/L | Chlorodehydroabietic | X239 | Gas Chromatography | Gas Chromatography | 0.001 | Y | Y | Y | 12 |  |  | 10/23/1996 11:56:36 PM | 1996-04-12 | 6/3/1998 6:51:34 PM | EMS |  | Concentration | mg/L | Y | (fl. conc.) |  | Chlorodehydroabietic (fl. conc.) | 65310-45-4, 65281-76-7 | Organics; | Organics; EMS code: C050 | Chemical |  | LAB
C051 | mg/L | Chlorfenson | ZN01 | DCM+Ac.Extr:GC/MS | DCM+Ac.Extr:GC/MS | 0.0002 | Y | Y | Y | 28 | 7/28/2005 7:03:00 AM | 3/29/2012 10:15:00 PM | 5/3/2012 9:27:50 AM | 1996-04-12 | 11/16/1999 4:06:27 PM | EMS | NPEPPIN | Concentration | mg/L | Y | (fl. conc.) |  | Chlorfenson (fl. conc.) |  | Pesticides; | Pesticides; EMS code: C051 | Chemical |  | LAB
C052 | mg/L | Chlorfenvinphos | ZN01 | DCM+Ac.Extr:GC/MS | DCM+Ac.Extr:GC/MS | 0.0005 | Y | Y | Y | 28 | 7/29/2005 7:10:00 AM | 3/29/2012 10:15:00 PM | 5/3/2012 9:27:50 AM | 1996-04-12 | 11/16/1999 4:06:27 PM | EMS | NPEPPIN | Concentration | mg/L | Y | (fl. conc.) |  | Chlorfenvinphos (fl. conc.) | 470-90-6 | Pesticides; | Pesticides; EMS code: C052 | Chemical |  | LAB
C054 | mg/L | 3-Chlorophenol | X268 | CH2Cl2 ext:GC/MS | CH2Cl2 ext:GC/MS | 0.0001 | Y | Y | Y | 26 | 12/28/2021 | 6/7/2024 | 6/26/2024 4:28:10 PM | 2012-06-25 |  | RDGIBSON |  | Concentration | mg/L | Y | (fl. conc.) |  | 3-Chlorophenol (fl. conc.) |  | Organics; | Organics; EMS code: C054 | Chemical |  | LAB
C054 | ug/g | 3-Chlorophenol | 1332 | Phenolics by Tumbler/GC-MS | Phenolics by Tumbler/GC-MS | 0.02 | Y | Y | Y | 23 | 9/9/2019 | 2/2/2022 | 1/11/2024 1:07:07 PM | 2019-09-30 |  | KATFRASE |  | MassPerMass | mg/kg | Y | (mass conc.) |  | 3-Chlorophenol (mass conc.) |  | Organics; | Organics; EMS code: C054 | Chemical |  | LAB
C055 | mg/L | 4-Chlorophenol | X268 | CH2Cl2 ext:GC/MS | CH2Cl2 ext:GC/MS | 0.0001 | Y | Y | Y | 26 | 12/28/2021 | 6/7/2024 | 6/26/2024 4:28:10 PM | 2012-06-25 |  | RDGIBSON |  | Concentration | mg/L | Y | (fl. conc.) |  | 4-Chlorophenol (fl. conc.) |  | Organics; | Organics; EMS code: C055 | Chemical |  | LAB
C055 | ug/g | 4-Chlorophenol | 1332 | Phenolics by Tumbler/GC-MS | Phenolics by Tumbler/GC-MS | 0.02 | Y | Y | Y | 23 | 9/9/2019 | 2/2/2022 | 1/11/2024 1:07:07 PM | 2019-09-30 |  | KATFRASE |  | MassPerMass | mg/kg | Y | (mass conc.) |  | 4-Chlorophenol (mass conc.) |  | Organics; | Organics; EMS code: C055 | Chemical |  | LAB
C058 | mg/L | 2-Chloroethylvinylether | MS01 | Purge&Trap GC/MS (VOC-W) | Purge&Trap GC/MS (VOC-W) | 0.001 | Y | Y | Y | 99 | 7/29/2005 7:02:00 AM | 7/26/2012 7:12:00 AM | 9/26/2012 1:07:08 PM | 2003-06-17 |  | UMULLER |  | Concentration | mg/L | Y | (fl. conc.) |  | 2-Chloroethylvinylether (fl. conc.) | 110-75-8 | Organics; | Organics; EMS code: C058 | Chemical |  | LAB
C059 | mg/L | Dichlorodifluoromethane | MS04 | Static Headspace GC/MS (Water) | (Water) A static headspace procedure is used as a means of driving the analytes out of the matrix into a headspace. Once equilibrium is achieved, a fixed volume of headspeace is taken for subsequent determination by Gas Chromatography - Mass Spectroscopy (GC/MS) using Selected Ion Monitoring. | 0.002 | Y | Y | Y | 12 | 7/13/2011 | 6/28/2024 12:29:00 AM | 7/9/2024 3:31:02 PM | 2012-02-08 |  | RDGIBSON |  | Concentration | mg/L | Y | (fl. conc.) |  | Dichlorodifluoromethane (fl. conc.) | 75-71-8 | Chlorofluorocarbon (CFCs); | Chlorofluorocarbon (CFCs); EMS code: C059 | Chemical |  | LAB
C060 | mg/L | Chlorpropham | ZN01 | DCM+Ac.Extr:GC/MS | DCM+Ac.Extr:GC/MS | 0.0002 | Y | Y | Y | 28 | 7/28/2005 7:03:00 AM | 3/29/2012 10:15:00 PM | 5/3/2012 9:27:50 AM | 1996-04-12 | 11/16/1999 4:06:27 PM | EMS | NPEPPIN | Concentration | mg/L | Y | (fl. conc.) |  | Chlorpropham (fl. conc.) | 101-21-3 | Pesticides; | Pesticides; EMS code: C060 | Chemical |  | LAB
C061 | mg/L | Chlorthal-dimethyl | ZN01 | DCM+Ac.Extr:GC/MS | DCM+Ac.Extr:GC/MS | 0.0001 | Y | Y | Y | 1 | 7/28/2005 7:03:00 AM | 7/28/2005 7:03:00 AM | 6/2/2009 12:01:00 PM | 1996-04-12 | 6/3/1998 6:51:42 PM | EMS |  | Concentration | mg/L | Y | (fl. conc.) |  | Chlorthal-dimethyl (fl. conc.) | 1861-32-1 | Pesticides; | Pesticides; EMS code: C061 | Chemical |  | LAB
C062 | mg/L | 1;1-Dichloroethene | MS01 | Purge&Trap GC/MS (VOC-W) | Purge&Trap GC/MS (VOC-W) | 0.001 | Y | Y | Y | 332 | 7/29/2005 7:02:00 AM | 5/27/2016 6:28:00 PM | 12/30/2023 10:26:22 AM | 2003-06-17 |  | UMULLER |  | Concentration | mg/L | Y | (fl. conc.) |  | 1;1-Dichloroethene (fl. conc.) | 75-35-4 | Volatile Organic Compounds (VOCs); | Volatile Organic Compounds (VOCs); EMS code: C062 | Chemical |  | LAB
C063 | mg/L | cis-1;2-Dichloroethene | MS01 | Purge&Trap GC/MS (VOC-W) | Purge&Trap GC/MS (VOC-W) | 0.001 | Y | Y | Y | 433 | 7/29/2005 7:02:00 AM | 6/4/2024 10:24:00 AM | 6/13/2024 5:01:27 PM | 2003-06-17 |  | UMULLER |  | Concentration | mg/L | Y | (fl. conc.) |  | cis-1;2-Dichloroethene (fl. conc.) | 156-59-2 | Volatile Organic Compounds (VOCs); | Volatile Organic Compounds (VOCs); EMS code: C063 | Chemical |  | LAB
C063 | ug/g | cis-1;2-Dichloroethene | MS02 | Purge&Trap GC/MS (VOC-S) | Purge&Trap GC/MS (VOC-S) | 0.05 | Y | Y | Y | 1 | 10/28/2022 5:04:00 PM | 10/28/2022 5:04:00 PM | 1/7/2023 3:12:09 PM | 2022-12-29 |  | EMS |  | MassPerMass | mg/kg | Y | (mass conc.) |  | cis-1;2-Dichloroethene (mass conc.) | 156-59-2 | Volatile Organic Compounds (VOCs); | Volatile Organic Compounds (VOCs); EMS code: C063 | Chemical |  | LAB
C063 | mg/L | cis-1;2-Dichloroethene | MS04 | Static Headspace GC/MS (Water) | (Water) A static headspace procedure is used as a means of driving the analytes out of the matrix into a headspace. Once equilibrium is achieved, a fixed volume of headspeace is taken for subsequent determination by Gas Chromatography - Mass Spectroscopy (GC/MS) using Selected Ion Monitoring. | 0.0005 | Y | Y | Y | 19 |  |  | 3/10/2023 3:28:43 PM | 2009-12-18 |  | RDGIBSON |  | Concentration | mg/L | Y | (fl. conc.) |  | cis-1;2-Dichloroethene (fl. conc.) | 156-59-2 | Volatile Organic Compounds (VOCs); | Volatile Organic Compounds (VOCs); EMS code: C063 | Chemical |  | LAB
C064 | mg/L | trans-1;2-Dichloroethene | MS04 | Static Headspace GC/MS (Water) | (Water) A static headspace procedure is used as a means of driving the analytes out of the matrix into a headspace. Once equilibrium is achieved, a fixed volume of headspeace is taken for subsequent determination by Gas Chromatography - Mass Spectroscopy (GC/MS) using Selected Ion Monitoring. | 0.001 | Y | Y | Y | 8 | 10/5/2019 2:07:00 AM | 6/28/2024 12:29:00 AM | 7/9/2024 3:31:02 PM | 2021-04-07 |  | RDGIBSON |  | Concentration | mg/L | Y | (fl. conc.) |  | trans-1;2-Dichloroethene (fl. conc.) | 156-60-5 | Volatile Organic Compounds (VOCs); | Volatile Organic Compounds (VOCs); EMS code: C064 | Chemical |  | LAB
C064 | ug/g | trans-1;2-Dichloroethene | MS02 | Purge&Trap GC/MS (VOC-S) | Purge&Trap GC/MS (VOC-S) | 0.05 | Y | Y | Y | 1 | 10/28/2022 5:04:00 PM | 10/28/2022 5:04:00 PM | 1/7/2023 3:12:09 PM | 2022-12-29 |  | EMS |  | MassPerMass | mg/kg | Y | (mass conc.) |  | trans-1;2-Dichloroethene (mass conc.) | 156-60-5 | Volatile Organic Compounds (VOCs); | Volatile Organic Compounds (VOCs); EMS code: C064 | Chemical |  | LAB
C064 | mg/L | trans-1;2-Dichloroethene | MS01 | Purge&Trap GC/MS (VOC-W) | Purge&Trap GC/MS (VOC-W) | 0.001 | Y | Y | Y | 458 | 7/29/2005 7:02:00 AM | 6/4/2024 10:24:00 AM | 6/13/2024 5:01:27 PM | 2003-06-17 |  | UMULLER |  | Concentration | mg/L | Y | (fl. conc.) |  | trans-1;2-Dichloroethene (fl. conc.) | 156-60-5 | Volatile Organic Compounds (VOCs); | Volatile Organic Compounds (VOCs); EMS code: C064 | Chemical |  | LAB
C067 | ug/g | alpha-BHC | ZN03 | DCM+Ac.Extr:GC/MS | DCM+Ac.Extr:GC/MS EPA 8270E (mod) | 0.02 | Y | Y | Y | 17 | 1/17/2022 | 11/14/2022 | 1/11/2024 1:07:07 PM | 2022-06-23 |  | EMS |  | MassPerMass | mg/kg | Y | (mass conc.) |  | alpha-BHC (mass conc.) | 319-84-6 | Pesticides; | Pesticides; EMS code: C067 | Chemical |  | LAB
C067 | mg/L | alpha-BHC | E660 | EPA 8270E (mod) | EPA 8270E (mod) Miscellaneous Pesticides by GC-MS | 5e-06 | Y | Y | Y | 2 | 9/14/2022 | 9/14/2022 | 12/8/2022 9:48:13 AM | 2022-12-06 |  | EMS |  | Concentration | mg/L | Y | (fl. conc.) |  | alpha-BHC (fl. conc.) | 319-84-6 | Pesticides; | Pesticides; EMS code: C067 | Chemical |  | LAB
C067 | mg/L | alpha-BHC | MW03 | OCP1-LL-SF-ECD-VA | This analysis is carried out using procedures adapted from "Test Methods for Evaluating Solid Waste" SW-846, Methods 3510, 3630, 3660 and 8081, published by the United States Environmental Protection Agency (EPA). The procedure involves extraction of the entire water sample with dichloromethane. The extract then undergoes one or more of the following clean-up procedures (if required): silica gel clean-up and/or sulphur clean-up. The final extract is analyzed by dual capillary column gas chromatography with electron capture detection (GC/ECD). | 5e-05 | Y | Y | Y | 10 | 9/20/2016 | 9/21/2016 | 8/15/2019 2:42:34 PM | 2019-07-08 |  | KATFRASE |  | Concentration | mg/L | Y | (fl. conc.) |  | alpha-BHC (fl. conc.) | 319-84-6 | Pesticides; | Pesticides; EMS code: C067 | Chemical |  | LAB
C067 | mg/L | alpha-BHC | X268 | CH2Cl2 ext:GC/MS | CH2Cl2 ext:GC/MS |  | Y | Y | Y | 16 | 12/29/2021 | 9/2/2022 | 11/10/2022 9:54:06 AM | 2022-04-08 |  | EMS |  | Concentration | mg/L | Y | (fl. conc.) |  | alpha-BHC (fl. conc.) | 319-84-6 | Pesticides; | Pesticides; EMS code: C067 | Chemical |  | LAB
C067 | ug/g | alpha-BHC | E8081 | EPA 8081 | See methods EPA 3540, 3630, 3660, 8081 | 0.001 | Y | Y | Y | 9 | 3/18/2020 | 3/23/2020 | 2/22/2023 12:51:03 PM | 2022-07-29 |  | EMS |  | MassPerMass | mg/kg | Y | (mass conc.) |  | alpha-BHC (mass conc.) | 319-84-6 | Pesticides; | Pesticides; EMS code: C067 | Chemical |  | LAB
C070 | ug/g | Chloromethane | FS04 | Static Headspace:HRGC/LRMS | Static Headspace anaylsis is a method for the determination of volatile compounds in liquids and/or solids by analysing the vapour phase of a sample in a closed system; in this case a sealed vial. Static Headspace is sometimes referred to as Equilibrium Headspace because the vapour phase above the sample in a sealed vial is brought into thermodynamic equilibrium with the sample prior to analysis. Volatile organic compounds are thermally extracted from the sample matrix by Static Headspace sampling and introduced directly to a thick-film megabore GC column. Compounds eluting from the GC column are introduced into a Mass Spectrometer and target compounds are identified by a comparison of both retention times and mass spectra with those of reference standards stored in a computerized data system. Reference retention times and mass spectra for analytes are obtained by the measurement of calibration standards under precisely the same extraction conditions used for samples. The concentration of each identified component is measured by relating the response of the sample quantitation mass ions to the response of the quantitation mass ions of a series of external standards. A deuterated surrogate recovery standard is added to provide information on matrix effects. | 0.1 | Y | Y | Y | 43 | 5/21/2015 | 5/22/2015 | 7/3/2015 3:19:04 PM | 2015-07-03 |  | RDGIBSON |  | MassPerMass | mg/kg | Y | (mass conc.) |  | Chloromethane (mass conc.) | 74-87-3 | Volatile Organic Compounds (VOCs); | Volatile Organic Compounds (VOCs); EMS code: C070 | Chemical |  | LAB
C070 | mg/L | Chloromethane | X384 | Purge and trap GC/MS | The sample is analyzed by purge and trap gas chromatography with detection by mass spectrometry. The procedure involves purging the volatile compounds from the sample with an inert gas, and trapping them on a solid sorbent. When purging of the sample is complete, the trap is heated and the compounds are transferred to the gas chromatographic column. Analysis is then accomplished by separation of the components by gas chromatography with detection by mass spectrometry. | 0.001 | Y | Y | Y | 112 | 8/5/2004 10:31:00 AM | 7/26/2012 7:12:00 AM | 9/26/2012 1:07:08 PM | 1996-04-12 | 6/3/1998 6:51:34 PM | EMS |  | Concentration | mg/L | Y | (fl. conc.) |  | Chloromethane (fl. conc.) | 74-87-3 | Volatile Organic Compounds (VOCs); | Volatile Organic Compounds (VOCs); EMS code: C070 | Chemical |  | LAB
C070 | mg/L | Chloromethane | F007 | Static Headspace:HRGC/LRMS | Static headspace analysis is a method for the determination of volatile compounds in liquids and/or solids by analysing the vapour phase of a sample in a closed system; in this case a sealed vial. Static Headspace is sometimes referred to as Equilibrium Headspace because the vapour phase above the sample in a sealed vial is brought into thermodynamic equilibrium with the sample prior to analysis. Volatile organic compounds are thermally extracted from the sample matrix by Static Headspace sampling and introduced directly to a thick-film megabore GC column. Compounds eluting from the GC column are introduced into a Mass Spectrometer and target compounds are identified by a comparison of both retention times and mass spectra with those of reference standards stored in a computerized data system. Reference retention times and mass spectra for analytes are obtained by the measurement of calibration standards under precisely the same extraction conditions used for samples. The concentration of each identified component is measured by relating the response of the sample quantitation mass ions to the response of the quantitation mass ions of a series of external standards. A deuterated surrogate recovery standard is added to provide information on matrix effects. | 0.005 | Y | Y | Y | 329 | 4/15/2011 | 8/14/2024 | 8/16/2024 4:05:21 PM | 2012-03-28 |  | RDGIBSON |  | Concentration | mg/L | Y | (fl. conc.) |  | Chloromethane (fl. conc.) | 74-87-3 | Volatile Organic Compounds (VOCs); | Volatile Organic Compounds (VOCs); EMS code: C070 | Chemical |  | LAB
C070 | mg/L | Chloromethane | MS01 | Purge&Trap GC/MS (VOC-W) | Purge&Trap GC/MS (VOC-W) | 2.3 | Y | Y | Y | 288 | 4/15/2002 | 5/18/2011 1:13:00 PM | 7/9/2012 2:09:47 PM | 1996-04-12 | 3/18/2003 2:49:40 PM | EMS | UMULLER | Concentration | mg/L | Y | (fl. conc.) |  | Chloromethane (fl. conc.) | 74-87-3 | Volatile Organic Compounds (VOCs); | Volatile Organic Compounds (VOCs); EMS code: C070 | Chemical |  | LAB
C070 | ug/g | Chloromethane | MS02 | Purge&Trap GC/MS (VOC-S) | Purge&Trap GC/MS (VOC-S) | 0.012 | Y | Y | Y | 26 | 11/19/2004 | 2/4/2022 | 1/11/2024 1:07:07 PM | 1996-04-12 | 6/3/1998 6:51:45 PM | EMS |  | MassPerMass | mg/kg | Y | (mass conc.) |  | Chloromethane (mass conc.) | 74-87-3 | Volatile Organic Compounds (VOCs); | Volatile Organic Compounds (VOCs); EMS code: C070 | Chemical |  | LAB
C070 | mg/L | Chloromethane | MS04 | Static Headspace GC/MS (Water) | (Water) A static headspace procedure is used as a means of driving the analytes out of the matrix into a headspace. Once equilibrium is achieved, a fixed volume of headspeace is taken for subsequent determination by Gas Chromatography - Mass Spectroscopy (GC/MS) using Selected Ion Monitoring. | 0.001 | Y | Y | Y | 70 | 8/13/2009 | 6/28/2024 12:29:00 AM | 7/9/2024 3:31:02 PM | 2009-12-18 |  | RDGIBSON |  | Concentration | mg/L | Y | (fl. conc.) |  | Chloromethane (fl. conc.) | 74-87-3 | Volatile Organic Compounds (VOCs); | Volatile Organic Compounds (VOCs); EMS code: C070 | Chemical |  | LAB
C071 | ug/g | trans-Nonachlor | ZN03 | DCM+Ac.Extr:GC/MS | DCM+Ac.Extr:GC/MS EPA 8270E (mod) | 0.005 | Y | Y | Y | 2 | 11/14/2022 | 11/14/2022 | 12/15/2022 10:50:04 AM | 2022-06-23 |  | EMS |  | MassPerMass | mg/kg | Y | (mass conc.) |  | trans-Nonachlor (mass conc.) |  | Pesticides; | Pesticides; EMS code: C071 | Chemical |  | LAB
C071 | mg/L | trans-Nonachlor | MW03 | OCP1-LL-SF-ECD-VA | This analysis is carried out using procedures adapted from "Test Methods for Evaluating Solid Waste" SW-846, Methods 3510, 3630, 3660 and 8081, published by the United States Environmental Protection Agency (EPA). The procedure involves extraction of the entire water sample with dichloromethane. The extract then undergoes one or more of the following clean-up procedures (if required): silica gel clean-up and/or sulphur clean-up. The final extract is analyzed by dual capillary column gas chromatography with electron capture detection (GC/ECD). | 5e-05 | Y | Y | Y | 10 | 9/20/2016 | 9/21/2016 | 8/15/2019 2:42:34 PM | 2019-07-08 |  | KATFRASE |  | Concentration | mg/L | Y | (fl. conc.) |  | trans-Nonachlor (fl. conc.) |  | Pesticides; | Pesticides; EMS code: C071 | Chemical |  | LAB
C071 | mg/L | trans-Nonachlor | X268 | CH2Cl2 ext:GC/MS | CH2Cl2 ext:GC/MS |  | Y | Y | Y | 16 | 12/29/2021 | 9/2/2022 | 11/10/2022 9:54:06 AM | 2022-04-08 |  | EMS |  | Concentration | mg/L | Y | (fl. conc.) |  | trans-Nonachlor (fl. conc.) |  | Pesticides; | Pesticides; EMS code: C071 | Chemical |  | LAB
C071 | ug/g | trans-Nonachlor | E8081 | EPA 8081 | See methods EPA 3540, 3630, 3660, 8081 | 0.001 | Y | Y | Y | 9 | 3/18/2020 | 3/23/2020 | 2/22/2023 12:51:03 PM | 2022-07-29 |  | EMS |  | MassPerMass | mg/kg | Y | (mass conc.) |  | trans-Nonachlor (mass conc.) |  | Pesticides; | Pesticides; EMS code: C071 | Chemical |  | LAB
C080 | mg/L | 1;1;2;2-Tetrachloroethane | MS01 | Purge&Trap GC/MS (VOC-W) | Purge&Trap GC/MS (VOC-W) | 0.4 | Y | Y | Y | 544 | 4/15/2002 | 6/4/2024 10:24:00 AM | 6/13/2024 5:01:27 PM | 1996-04-12 | 3/18/2003 2:49:40 PM | EMS | UMULLER | Concentration | mg/L | Y | (fl. conc.) |  | 1;1;2;2-Tetrachloroethane (fl. conc.) |  | Volatile Organic Compounds (VOCs); | Volatile Organic Compounds (VOCs); EMS code: C080 | Chemical |  | LAB
C080 | mg/L | 1;1;2;2-Tetrachloroethane | X384 | Purge and trap GC/MS | The sample is analyzed by purge and trap gas chromatography with detection by mass spectrometry. The procedure involves purging the volatile compounds from the sample with an inert gas, and trapping them on a solid sorbent. When purging of the sample is complete, the trap is heated and the compounds are transferred to the gas chromatographic column. Analysis is then accomplished by separation of the components by gas chromatography with detection by mass spectrometry. | 0.001 | Y | Y | Y | 112 | 8/5/2004 10:31:00 AM | 7/26/2012 7:12:00 AM | 9/26/2012 1:07:08 PM | 1996-04-12 | 6/3/1998 6:51:34 PM | EMS |  | Concentration | mg/L | Y | (fl. conc.) |  | 1;1;2;2-Tetrachloroethane (fl. conc.) |  | Volatile Organic Compounds (VOCs); | Volatile Organic Compounds (VOCs); EMS code: C080 | Chemical |  | LAB
C080 | ug/g | 1;1;2;2-Tetrachloroethane | MS02 | Purge&Trap GC/MS (VOC-S) | Purge&Trap GC/MS (VOC-S) | 0.002 | Y | Y | Y | 27 | 11/19/2004 | 10/28/2022 5:04:00 PM | 1/11/2024 1:07:07 PM | 1996-04-12 | 6/3/1998 6:51:45 PM | EMS |  | MassPerMass | mg/kg | Y | (mass conc.) |  | 1;1;2;2-Tetrachloroethane (mass conc.) |  | Volatile Organic Compounds (VOCs); | Volatile Organic Compounds (VOCs); EMS code: C080 | Chemical |  | LAB
C080 | mg/L | 1;1;2;2-Tetrachloroethane | MS04 | Static Headspace GC/MS (Water) | (Water) A static headspace procedure is used as a means of driving the analytes out of the matrix into a headspace. Once equilibrium is achieved, a fixed volume of headspeace is taken for subsequent determination by Gas Chromatography - Mass Spectroscopy (GC/MS) using Selected Ion Monitoring. | 0.0005 | Y | Y | Y | 56 | 8/13/2009 | 6/28/2024 12:29:00 AM | 7/9/2024 3:31:02 PM | 2009-12-18 |  | RDGIBSON |  | Concentration | mg/L | Y | (fl. conc.) |  | 1;1;2;2-Tetrachloroethane (fl. conc.) |  | Volatile Organic Compounds (VOCs); | Volatile Organic Compounds (VOCs); EMS code: C080 | Chemical |  | LAB
C080 | mg/L | 1;1;2;2-Tetrachloroethane | F007 | Static Headspace:HRGC/LRMS | Static headspace analysis is a method for the determination of volatile compounds in liquids and/or solids by analysing the vapour phase of a sample in a closed system; in this case a sealed vial. Static Headspace is sometimes referred to as Equilibrium Headspace because the vapour phase above the sample in a sealed vial is brought into thermodynamic equilibrium with the sample prior to analysis. Volatile organic compounds are thermally extracted from the sample matrix by Static Headspace sampling and introduced directly to a thick-film megabore GC column. Compounds eluting from the GC column are introduced into a Mass Spectrometer and target compounds are identified by a comparison of both retention times and mass spectra with those of reference standards stored in a computerized data system. Reference retention times and mass spectra for analytes are obtained by the measurement of calibration standards under precisely the same extraction conditions used for samples. The concentration of each identified component is measured by relating the response of the sample quantitation mass ions to the response of the quantitation mass ions of a series of external standards. A deuterated surrogate recovery standard is added to provide information on matrix effects. | 0.005 | Y | Y | Y | 380 | 9/28/1998 3:08:00 PM | 8/14/2024 | 8/16/2024 4:05:21 PM | 1996-09-16 | 6/3/1998 6:51:45 PM | MPAK |  | Concentration | mg/L | Y | (fl. conc.) |  | 1;1;2;2-Tetrachloroethane (fl. conc.) |  | Volatile Organic Compounds (VOCs); | Volatile Organic Compounds (VOCs); EMS code: C080 | Chemical |  | LAB
C080 | ug/g | 1;1;2;2-Tetrachloroethane | FS04 | Static Headspace:HRGC/LRMS | Static Headspace anaylsis is a method for the determination of volatile compounds in liquids and/or solids by analysing the vapour phase of a sample in a closed system; in this case a sealed vial. Static Headspace is sometimes referred to as Equilibrium Headspace because the vapour phase above the sample in a sealed vial is brought into thermodynamic equilibrium with the sample prior to analysis. Volatile organic compounds are thermally extracted from the sample matrix by Static Headspace sampling and introduced directly to a thick-film megabore GC column. Compounds eluting from the GC column are introduced into a Mass Spectrometer and target compounds are identified by a comparison of both retention times and mass spectra with those of reference standards stored in a computerized data system. Reference retention times and mass spectra for analytes are obtained by the measurement of calibration standards under precisely the same extraction conditions used for samples. The concentration of each identified component is measured by relating the response of the sample quantitation mass ions to the response of the quantitation mass ions of a series of external standards. A deuterated surrogate recovery standard is added to provide information on matrix effects. | 0.05 | Y | Y | Y | 44 | 6/30/1999 4:41:00 PM | 5/22/2015 | 7/3/2015 3:19:04 PM | 1996-09-16 | 6/3/1998 6:51:45 PM | MPAK |  | MassPerMass | mg/kg | Y | (mass conc.) |  | 1;1;2;2-Tetrachloroethane (mass conc.) |  | Volatile Organic Compounds (VOCs); | Volatile Organic Compounds (VOCs); EMS code: C080 | Chemical |  | LAB
C081 | ug/g | 1;1;1;2-Tetrachloroethane | FS04 | Static Headspace:HRGC/LRMS | Static Headspace anaylsis is a method for the determination of volatile compounds in liquids and/or solids by analysing the vapour phase of a sample in a closed system; in this case a sealed vial. Static Headspace is sometimes referred to as Equilibrium Headspace because the vapour phase above the sample in a sealed vial is brought into thermodynamic equilibrium with the sample prior to analysis. Volatile organic compounds are thermally extracted from the sample matrix by Static Headspace sampling and introduced directly to a thick-film megabore GC column. Compounds eluting from the GC column are introduced into a Mass Spectrometer and target compounds are identified by a comparison of both retention times and mass spectra with those of reference standards stored in a computerized data system. Reference retention times and mass spectra for analytes are obtained by the measurement of calibration standards under precisely the same extraction conditions used for samples. The concentration of each identified component is measured by relating the response of the sample quantitation mass ions to the response of the quantitation mass ions of a series of external standards. A deuterated surrogate recovery standard is added to provide information on matrix effects. | 0.05 | Y | Y | Y | 43 | 5/21/2015 | 5/22/2015 | 7/3/2015 3:19:04 PM | 2015-07-03 |  | RDGIBSON |  | MassPerMass | mg/kg | Y | (mass conc.) |  | 1;1;1;2-Tetrachloroethane (mass conc.) |  | Volatile Organic Compounds (VOCs); | Volatile Organic Compounds (VOCs); EMS code: C081 | Chemical |  | LAB
C081 | ug/g | 1;1;1;2-Tetrachloroethane | MS02 | Purge&Trap GC/MS (VOC-S) | Purge&Trap GC/MS (VOC-S) | 0.05 | Y | Y | Y | 15 | 1/14/2022 | 2/4/2022 | 1/11/2024 1:07:07 PM | 2022-07-27 |  | EMS |  | MassPerMass | mg/kg | Y | (mass conc.) |  | 1;1;1;2-Tetrachloroethane (mass conc.) |  | Volatile Organic Compounds (VOCs); | Volatile Organic Compounds (VOCs); EMS code: C081 | Chemical |  | LAB
C081 | mg/L | 1;1;1;2-Tetrachloroethane | MS01 | Purge&Trap GC/MS (VOC-W) | Purge&Trap GC/MS (VOC-W) | 0.001 | Y | Y | Y | 89 | 4/15/2002 | 4/1/2011 | 7/9/2012 2:09:47 PM | 2003-03-18 | 6/17/2003 10:35:02 AM | UMULLER | UMULLER | Concentration | mg/L | Y | (fl. conc.) |  | 1;1;1;2-Tetrachloroethane (fl. conc.) |  | Volatile Organic Compounds (VOCs); | Volatile Organic Compounds (VOCs); EMS code: C081 | Chemical |  | LAB
C081 | mg/L | 1;1;1;2-Tetrachloroethane | F007 | Static Headspace:HRGC/LRMS | Static headspace analysis is a method for the determination of volatile compounds in liquids and/or solids by analysing the vapour phase of a sample in a closed system; in this case a sealed vial. Static Headspace is sometimes referred to as Equilibrium Headspace because the vapour phase above the sample in a sealed vial is brought into thermodynamic equilibrium with the sample prior to analysis. Volatile organic compounds are thermally extracted from the sample matrix by Static Headspace sampling and introduced directly to a thick-film megabore GC column. Compounds eluting from the GC column are introduced into a Mass Spectrometer and target compounds are identified by a comparison of both retention times and mass spectra with those of reference standards stored in a computerized data system. Reference retention times and mass spectra for analytes are obtained by the measurement of calibration standards under precisely the same extraction conditions used for samples. The concentration of each identified component is measured by relating the response of the sample quantitation mass ions to the response of the quantitation mass ions of a series of external standards. A deuterated surrogate recovery standard is added to provide information on matrix effects. | 0.001 | Y | Y | Y | 329 | 4/15/2011 | 8/14/2024 | 8/16/2024 4:05:21 PM | 2012-03-28 |  | RDGIBSON |  | Concentration | mg/L | Y | (fl. conc.) |  | 1;1;1;2-Tetrachloroethane (fl. conc.) |  | Volatile Organic Compounds (VOCs); | Volatile Organic Compounds (VOCs); EMS code: C081 | Chemical |  | LAB
C081 | mg/L | 1;1;1;2-Tetrachloroethane | MS04 | Static Headspace GC/MS (Water) | (Water) A static headspace procedure is used as a means of driving the analytes out of the matrix into a headspace. Once equilibrium is achieved, a fixed volume of headspeace is taken for subsequent determination by Gas Chromatography - Mass Spectroscopy (GC/MS) using Selected Ion Monitoring. | 0.0005 | Y | Y | Y | 70 | 8/13/2009 | 6/28/2024 12:29:00 AM | 7/9/2024 3:31:02 PM | 2009-12-18 |  | RDGIBSON |  | Concentration | mg/L | Y | (fl. conc.) |  | 1;1;1;2-Tetrachloroethane (fl. conc.) |  | Volatile Organic Compounds (VOCs); | Volatile Organic Compounds (VOCs); EMS code: C081 | Chemical |  | LAB
C081 | mg/L | 1;1;1;2-Tetrachloroethane | X384 | Purge and trap GC/MS | The sample is analyzed by purge and trap gas chromatography with detection by mass spectrometry. The procedure involves purging the volatile compounds from the sample with an inert gas, and trapping them on a solid sorbent. When purging of the sample is complete, the trap is heated and the compounds are transferred to the gas chromatographic column. Analysis is then accomplished by separation of the components by gas chromatography with detection by mass spectrometry. | 0.001 | Y | Y | Y | 93 | 4/27/2000 | 3/22/2007 | 3/31/2007 1:01:04 AM | 1996-04-12 | 6/3/1998 6:51:34 PM | EMS |  | Concentration | mg/L | Y | (fl. conc.) |  | 1;1;1;2-Tetrachloroethane (fl. conc.) |  | Volatile Organic Compounds (VOCs); | Volatile Organic Compounds (VOCs); EMS code: C081 | Chemical |  | LAB
C082 | mg/L | Tetrachloroethene | MS01 | Purge&Trap GC/MS (VOC-W) | Purge&Trap GC/MS (VOC-W) | 0.001 | Y | Y | Y | 233 | 1/15/2007 4:50:00 PM | 5/27/2016 6:28:00 PM | 12/30/2023 10:26:22 AM | 2003-06-17 |  | UMULLER |  | Concentration | mg/L | Y | (fl. conc.) |  | Tetrachloroethene (fl. conc.) | 127-18-4 | Volatile Organic Compounds (VOCs); | Volatile Organic Compounds (VOCs); EMS code: C082 | Chemical |  | LAB
C084 | mg/L | Carboxin | ZN01 | DCM+Ac.Extr:GC/MS | DCM+Ac.Extr:GC/MS | 0.0005 | Y | Y | Y | 1 | 7/29/2005 7:10:00 AM | 7/29/2005 7:10:00 AM | 6/2/2009 12:01:01 PM | 2003-06-16 |  | UMULLER |  | Concentration | mg/L | Y | (fl. conc.) |  | Carboxin (fl. conc.) | 5234-68-4 | Pesticides; | Pesticides; EMS code: C084 | Chemical |  | LAB
C085 | mg/L | Chlorbenzilate | ZN01 | DCM+Ac.Extr:GC/MS | DCM+Ac.Extr:GC/MS | 0.0005 | Y | Y | Y | 1 | 7/29/2005 7:10:00 AM | 7/29/2005 7:10:00 AM | 6/2/2009 12:01:01 PM | 2003-06-16 |  | UMULLER |  | Concentration | mg/L | Y | (fl. conc.) |  | Chlorbenzilate (fl. conc.) |  | Pesticides; | Pesticides; EMS code: C085 | Chemical |  | LAB
C086 | mg/L | Chlorbromuron | ZN01 | DCM+Ac.Extr:GC/MS | DCM+Ac.Extr:GC/MS | 0.0005 | Y | Y | Y | 1 | 7/29/2005 7:10:00 AM | 7/29/2005 7:10:00 AM | 6/2/2009 12:01:01 PM | 2003-06-16 |  | UMULLER |  | Concentration | mg/L | Y | (fl. conc.) |  | Chlorbromuron (fl. conc.) |  | Pesticides; | Pesticides; EMS code: C086 | Chemical |  | LAB
C090 | mg/L | Cyanophos | ZN01 | DCM+Ac.Extr:GC/MS | DCM+Ac.Extr:GC/MS | 0.0002 | Y | Y | Y | 28 | 7/29/2005 7:10:00 AM | 3/29/2012 10:15:00 PM | 5/3/2012 9:27:50 AM | 1996-04-12 | 11/16/1999 4:06:27 PM | EMS | NPEPPIN | Concentration | mg/L | Y | (fl. conc.) |  | Cyanophos (fl. conc.) |  | Pesticides; | Pesticides; EMS code: C090 | Chemical |  | LAB
C093 | mg/L | Crufomate | ZN01 | DCM+Ac.Extr:GC/MS | DCM+Ac.Extr:GC/MS | 0.0005 | Y | Y | Y | 1 | 7/29/2005 7:10:00 AM | 7/29/2005 7:10:00 AM | 6/2/2009 12:01:01 PM | 2003-06-16 |  | UMULLER |  | Concentration | mg/L | Y | (fl. conc.) |  | Crufomate (fl. conc.) |  | Pesticides; | Pesticides; EMS code: C093 | Chemical |  | LAB
C094 | mg/L | Cypermethrin | ZN01 | DCM+Ac.Extr:GC/MS | DCM+Ac.Extr:GC/MS | 0.0005 | Y | Y | Y | 1 | 7/29/2005 7:10:00 AM | 7/29/2005 7:10:00 AM | 6/2/2009 12:01:01 PM | 2003-06-16 |  | UMULLER |  | Concentration | mg/L | Y | (fl. conc.) |  | Cypermethrin (fl. conc.) | 52315-07-8 | Pesticides; | Pesticides; EMS code: C094 | Chemical |  | LAB
C095 | mg/L | Cyprazine | ZN01 | DCM+Ac.Extr:GC/MS | DCM+Ac.Extr:GC/MS | 0.0005 | Y | Y | Y | 1 | 7/29/2005 7:10:00 AM | 7/29/2005 7:10:00 AM | 6/2/2009 12:01:01 PM | 2003-06-16 |  | UMULLER |  | Concentration | mg/L | Y | (fl. conc.) |  | Cyprazine (fl. conc.) |  | Pesticides; | Pesticides; EMS code: C095 | Chemical |  | LAB
C097 | mg/L | Clomazone | ZN01 | DCM+Ac.Extr:GC/MS | DCM+Ac.Extr:GC/MS | 0.0005 | Y | Y | Y | 1 | 7/29/2005 7:10:00 AM | 7/29/2005 7:10:00 AM | 6/2/2009 12:01:01 PM | 2003-06-16 |  | UMULLER |  | Concentration | mg/L | Y | (fl. conc.) |  | Clomazone (fl. conc.) |  | Pesticides; | Pesticides; EMS code: C097 | Chemical |  | LAB
C1-2 | uS/cm | Conductivity (1:2 leachate) | E100 | CSSS Ch. 15 (mod)/APHA 2510 (mod) | CSSS Ch. 15 (mod)/APHA 2510 (mod) | 10000 | Y | Y | Y | 1 | 2/9/2023 | 2/9/2023 | 2/22/2023 12:19:05 PM | 2023-02-14 | 2/21/2023 4:20:20 PM | EMS | EMS | ElectricConductivity | µS/cm | Y | (elec. cond.) |  | Conductivity (1:2 leachate) (elec. cond.) |  |  | EMS code: C1-2 | Chemical |  | LAB
C100 | mg/L | Chinomethionate | ZN01 | DCM+Ac.Extr:GC/MS | DCM+Ac.Extr:GC/MS | 0.5 | Y | Y | Y | 1 | 7/29/2005 7:10:00 AM | 7/29/2005 7:10:00 AM | 6/2/2009 12:01:01 PM | 2009-02-20 |  | JCHUDAK |  | Concentration | mg/L | Y | (fl. conc.) |  | Chinomethionate (fl. conc.) |  | Pesticides; | Pesticides; EMS code: C100 | Chemical |  | LAB
C101 | mg/L | Chlordimeform | ZN01 | DCM+Ac.Extr:GC/MS | DCM+Ac.Extr:GC/MS | 0.5 | Y | Y | Y | 1 | 7/29/2005 7:10:00 AM | 7/29/2005 7:10:00 AM | 6/2/2009 12:01:01 PM | 2009-02-20 |  | JCHUDAK |  | Concentration | mg/L | Y | (fl. conc.) |  | Chlordimeform (fl. conc.) |  | Pesticides; | Pesticides; EMS code: C101 | Chemical |  | LAB
C113 | mg/L | 1;1;2-trichloro-1;2;2-trifluoroethane | MS04 | Static Headspace GC/MS (Water) | (Water) A static headspace procedure is used as a means of driving the analytes out of the matrix into a headspace. Once equilibrium is achieved, a fixed volume of headspeace is taken for subsequent determination by Gas Chromatography - Mass Spectroscopy (GC/MS) using Selected Ion Monitoring. | 0.002 | Y | Y | Y | 12 | 7/13/2011 | 6/28/2024 12:29:00 AM | 7/9/2024 3:31:02 PM | 2012-02-08 |  | RDGIBSON |  | Concentration | mg/L | Y | (fl. conc.) |  | 1;1;2-trichloro-1;2;2-trifluoroethane (fl. conc.) |  | Chlorofluorocarbon (CFCs); | Chlorofluorocarbon (CFCs); EMS code: C113 | Chemical |  | LAB
C122 | % | d4-1;2-dichloroethane | SURR | Surrogate Recovery | Surrogate Recovery | 0 | Y | Y | Y | 67 | 11/28/2003 | 11/29/2004 1:55:00 PM | 1/6/2005 1:01:55 AM | 1996-04-12 | 6/3/1998 6:51:34 PM | EMS |  | DimensionlessRatio | % | N | (ratio) |  | d4-1;2-dichloroethane (ratio) | 17060-07-0 |  | EMS code: C122 | Chemical |  | SURROGATE_RESULT
CA-D | mg/L | Calcium Dissolved | EQ01 | EQUIS (mg/L) | EQUIS (mg/L) | 0.01 | N | Y | Y | 12267 |  |  | 10/24/1996 3:16:55 PM | 1996-04-12 | 6/3/1998 6:51:34 PM | EMS |  | Concentration | mg/L | Y | (fl. conc.) |  | Calcium Dissolved (fl. conc.) | 7440-70-2 | Metals - Dissolved; | Metals - Dissolved; EMS code: CA-D | Chemical | Dissolved | LAB
CA-D | mg/L | Calcium Dissolved | F032 | Filt.,Acid,ICPMS | There is no acid digestion required for this analysis. A water sample is collected and preserved in 0.2 percent nitric acid. To analyse the extractable metals the sample is shaken, allowed to sit for at least twenty four hours, and a decanted portion is introduced to the spectrometer. When the dissolved metals are required, the samples must first be filtered through a 0.45 micron cellulose acetate filter paper and then followed by preservation in 0.2 percent nitric acid. It is important to note that different laboratories may have very different detection limits for some variables analyzed by ICPMS. | 0.02 | Y | Y | Y | 84269 | 8/16/1996 2:35:00 PM | 8/26/2024 | 8/26/2024 4:53:05 PM | 1997-01-16 | 4/29/2011 8:48:48 AM | MPAK | RDGIBSON | Concentration | mg/L | Y | (fl. conc.) |  | Calcium Dissolved (fl. conc.) | 7440-70-2 | Metals - Dissolved; | Metals - Dissolved; EMS code: CA-D | Chemical | Dissolved | LAB
CA-D | mg/L | Calcium Dissolved | F039 | Field filt + acid, ICP-AES | Dissolved and extractable water samples are ready for analysis once collected. Total samples are autoclaved with HNO3 and hydrochloric acid (HCI) in capped Teflon vials. Analysis is performed via Inductively Coupled Argon Plasma Atomic Emission Spectrometry (ICP-AES)(simultaneous multielement analysis), Graphite Furnace Atomic Absorption Spectrometry (GF-AAS)(low level analysis) and/or Flame Atomic Absorption Spectrometry (F-AAS). | 0.1 | Y | Y | Y | 14080 | 7/29/1996 6:03:00 PM | 11/22/2022 12:29:00 PM | 7/22/2024 3:17:26 PM | 1997-05-08 | 6/3/1998 6:51:46 PM | MPAK |  | Concentration | mg/L | Y | (fl. conc.) |  | Calcium Dissolved (fl. conc.) | 7440-70-2 | Metals - Dissolved; | Metals - Dissolved; EMS code: CA-D | Chemical | Dissolved | LAB
CA-D | mg/L | Calcium Dissolved | X053 | AA,Dir Asp. | AA;Dir Asp. | 0.02 | Y | Y | Y | 6754 |  |  | 3/29/2001 2:38:48 PM | 1996-04-12 | 6/3/1998 6:51:34 PM | EMS |  | Concentration | mg/L | Y | (fl. conc.) |  | Calcium Dissolved (fl. conc.) | 7440-70-2 | Metals - Dissolved; | Metals - Dissolved; EMS code: CA-D | Chemical | Dissolved | LAB
CA-D | mg/L | Calcium Dissolved | X170 | ICP | This code is a catch-all for a large variety of ICP methods including water and also solids digested into solutions as an intermediate result. Different procedures will have greatly different detection limits. (a) For air particulates, preliminary results are reported in mg/L and loading results as mg/dm2/d. (b) For Hi Vol samples preliminiary results are same as reported in mg/L and loading results as ug/m3. (c) Note that for Rain samples there is no filtration nor digestion steps. | 0.1 | Y | Y | Y | 19597 | 4/11/1996 6:01:00 PM | 11/6/2018 | 6/13/2024 3:58:21 PM | 1996-04-12 | 8/15/2019 11:53:28 AM | EMS | KATFRASE | Concentration | mg/L | Y | (fl. conc.) |  | Calcium Dissolved (fl. conc.) | 7440-70-2 | Metals - Dissolved; | Metals - Dissolved; EMS code: CA-D | Chemical | Dissolved | LAB
CA-D | mg/L | Calcium Dissolved | X562 | Dissolved metals: FF, FA, HNO3, HR-ICPMS | Ultra trace metals in water are analyzed by high resolution inductively coupled plasma mass spectrometry (HR-ICPMS), based on US EPA Method 200.8, (Oct. 1999). This procedure is intended for pristine field-filtered acid-preserved water samples. The detection limits (LOR) for this test are based on lab instrumental analysis only, not including filtration. ALS-supplied field filtration equipment does not support these LOR. Therefore, because of the high probability of false positives due to filtration, it is strongly recommended that a filtration blank be analyzed to aid in data interpretation. | 0.01 | Y | Y | Y | 1258 | 6/8/2015 | 10/19/2023 | 11/9/2023 7:41:27 AM | 2015-10-22 |  | RDGIBSON |  | Concentration | mg/L | Y | (fl. conc.) |  | Calcium Dissolved (fl. conc.) | 7440-70-2 | Metals - Dissolved; | Metals - Dissolved; EMS code: CA-D | Chemical | Dissolved | LAB
CA-D | mg/L | Calcium Dissolved | METC | MET-D-NP-CCMS-VA | Ultra trace metals in water are analyzed by CRC ICPMS, based on US EPA Method 6020B (July 2014). This procedure is intended for pristine field-filtered acid-preserved water samples. The detection limits (LOR) for this test are based on lab instrumental analysis only, not including filtration. ALS-supplied field filtration equipment does not support these LOR. Therefore, because of the high probability of false positives due to filtration, it is strongly recommended that a filtration blank be analysed to aid in data interpretation. | 0.01 | Y | Y | Y | 1156 | 11/13/2017 | 8/15/2024 | 8/23/2024 9:50:35 AM | 2019-08-20 |  | KATFRASE |  | Concentration | mg/L | Y | (fl. conc.) |  | Calcium Dissolved (fl. conc.) | 7440-70-2 | Metals - Dissolved; | Metals - Dissolved; EMS code: CA-D | Chemical | Dissolved | LAB
CA-D | mg/L | Calcium Dissolved | CS01 | System Calculation | System Calculation (mg/L) |  | N | Y | N | 112 |  |  | 6/7/2024 5:30:17 PM | 1996-04-12 | 6/3/1998 6:51:43 PM | EMS |  | Concentration | mg/L | Y | (fl. conc.) |  | Calcium Dissolved (fl. conc.) | 7440-70-2 | Metals - Dissolved; | Metals - Dissolved; EMS code: CA-D | Chemical | Dissolved | LAB
CA-D | mg/L | Calcium Dissolved | F049 | Filt,HNO3 dig,conc'n,nebul'ICP-OES | This method provides an automated procedure for the Inductively Coupled Argon Plasma Optical Emission Spectrophotometric (ICP-OES) determination and quantitation of trace amounts of metals in surface, ground and precipitation waters after a manual acid digestion and concentration procedure and ultrasonic nebulisation. Organically-bound metals and metal complexes are converted to metal ions by digestion with nitric acid. To increase the signal detected by the ICP, the samples are concentrated 10 times. The liquid sample is pumped onto the piezoelectric transducer, where the ultrasonic action efficiently converts it to a fine, dense aerosol. The nebulizer carrier gas flow then transports the wet aerosol through a heated U-tube, where the solvent is vaporized. Subsequently, most of the solvent vapour is removed from the aerosol stream by passing the stream through a chilled condenser. The result is a dry, analyte-laden aerosol which is introduced to the plasma. There, the high temperature desolvates the aerosol and dissociates the sample material into individual atoms and ions, which are excited to emit light at wave lengths characteristic of the elements present. Dissolved samples are field filtered through 0.45 micron cellulose acetate filter prior to acid preservation. Samples for extractable analysis are acidified with 2 mL concentrated HNO3 per litre and allowed to stand a minimum of 24 hours before analysis. Clients providing these samples are advised to preserve samples at the collection site. | 0.05 | Y | Y | Y | 82 | 3/10/1998 2:04:00 PM | 2/25/1999 2:07:00 PM | 10/20/1999 2:21:44 AM | 1998-11-17 |  | NPEPPIN |  | Concentration | mg/L | Y | (fl. conc.) |  | Calcium Dissolved (fl. conc.) | 7440-70-2 | Metals - Dissolved; | Metals - Dissolved; EMS code: CA-D | Chemical | Dissolved | LAB
CA-D | mg/dm2/d | Calcium Dissolved | X286 | ICP | Aliquot of prepared sample digested with nitric acid then with perchloric acid. Filtered digest analyzed by ICP. Intermediate results reported as mg/L. Loading results for air particulate reported as mg/dm2/d. | 0 | Y | Y | Y | 23 |  |  | 10/24/1996 11:21:52 AM | 1996-04-12 | 6/3/1998 6:51:41 PM | EMS |  | MassFlux | mg/dm²/d | Y | (mass flux) |  | Calcium Dissolved (mass flux) | 7440-70-2 | Metals - Dissolved; | Metals - Dissolved; EMS code: CA-D | Chemical | Dissolved | LAB
CA-D | mg/L | Calcium Dissolved | MSEA | Metals in Seawater (undigested) by (CRC) ICP-MS | This method describes the multi-elemental determination of trace elements by Collision/Reaction Cell (CRC) ICP-MS. The method measures ions produced by a radio-frequency inductively coupled plasma. Analyte species originating in a liquid are pneumatically nebulized and the resulting aerosol transported by argon gas into the plasma torch. The ions produced are entrained in the plasma gas and extracted, by means of a differentially pumped vacuum interface, into a mass spectrometer. The ions flow through a Collision/Reaction Cell (CRC) where depending on the mode of analysis either the ions passes through the cell unhindered (no gas - classical ICP-MS mode), or collisions/reactions chemistry takes place between the ions and a neutral gas such as hydrogen or helium. Any polyatomic ions which give interference on the analytes are removed. The ions are then sorted according to their mass-to-charge ratios by a quadrupole mass spectrometer. The ions are then sorted according to their mass-to-charge ratios by a quadrupole mass spectrometer. The ions transmitted through the quadrupole are quantified by a channel electron multiplier and the ion information processed by a data handling system. Although most polyatomic interference is removed, any additional interference must be assessed and valid corrections applied or the data flagged to indicate problems. Interference correction must include compensation for background ions contributed by the plasma gas, reagents, and constituents of the sample matrix. Instrument drift as well as suppressions or enhancements of instrument response caused by the sample matrix must be corrected for by the use of internal standards. | 1 | Y | Y | Y | 4 | 4/12/2023 | 4/29/2024 | 5/16/2024 2:00:14 PM | 2009-09-08 |  | RDGIBSON |  | Concentration | mg/L | Y | (fl. conc.) |  | Calcium Dissolved (fl. conc.) | 7440-70-2 | Metals - Dissolved; | Metals - Dissolved; EMS code: CA-D | Chemical | Dissolved | LAB
CA-D | mg/L | Calcium Dissolved | X085 | LF,LA,HNO3,AA | Samples are lab-filtered (0.45 um), and lab-acidified (HNO3); analysis by AA. | 0.02 | Y | Y | Y | 208 | 12/12/1996 | 9/18/2015 | 10/28/2015 9:54:16 AM | 1996-04-12 | 6/3/1998 6:51:34 PM | EMS |  | Concentration | mg/L | Y | (fl. conc.) |  | Calcium Dissolved (fl. conc.) | 7440-70-2 | Metals - Dissolved; | Metals - Dissolved; EMS code: CA-D | Chemical | Dissolved | LAB
CA-D | mg/L | Calcium Dissolved | X350 | FF,FA:HNO3:ICP | Samples are field-filtered (0.45 um); field-acidified with HNO3; analysis by ICP/AES. | 0.5 | Y | Y | Y | 6987 |  |  | 8/26/2024 2:41:32 PM | 1996-04-12 | 6/3/1998 6:51:34 PM | EMS |  | Concentration | mg/L | Y | (fl. conc.) |  | Calcium Dissolved (fl. conc.) | 7440-70-2 | Metals - Dissolved; | Metals - Dissolved; EMS code: CA-D | Chemical | Dissolved | LAB
CA-D | mg/L | Calcium Dissolved | MWID | Metals in Water by ICPMS (digested) | Metals in Water by ICPMS (digested) Based on EPA 200.8 | 1 | Y | Y | Y | 2283 |  |  | 7/31/2024 1:46:41 PM | 2009-12-24 |  | RDGIBSON |  | Concentration | mg/L | Y | (fl. conc.) |  | Calcium Dissolved (fl. conc.) | 7440-70-2 | Metals - Dissolved; | Metals - Dissolved; EMS code: CA-D | Chemical | Dissolved | LAB
CA-D | mg/L | Calcium Dissolved | X285 | HNO3 Dig:ICP Analysis | HNO3 Dig:ICP Analysis | 0.1 | Y | Y | Y | 767 | 3/1/2000 | 11/13/2002 | 3/18/2004 10:28:34 AM | 1998-05-22 | 6/3/1998 6:51:49 PM | NPEPPIN |  | Concentration | mg/L | Y | (fl. conc.) |  | Calcium Dissolved (fl. conc.) | 7440-70-2 | Metals - Dissolved; | Metals - Dissolved; EMS code: CA-D | Chemical | Dissolved | LAB
CA-D | mg/L | Calcium Dissolved | 0200 | 20%A.R.Dig:AA Anal. | 20%A.R.Dig:AA Anal. | 0.02 | Y | Y | Y | 8 | 9/29/2000 | 12/1/2010 | 12/9/2010 3:55:26 PM | 1996-04-12 | 6/3/1998 6:51:34 PM | EMS |  | Concentration | mg/L | Y | (fl. conc.) |  | Calcium Dissolved (fl. conc.) | 7440-70-2 | Metals - Dissolved; | Metals - Dissolved; EMS code: CA-D | Chemical | Dissolved | LAB
CA-D | mg/L | Calcium Dissolved | 5138 | AA Analysis | AA Analysis | 0.02 | Y | Y | Y | 29 |  |  | 10/24/1996 11:35:00 AM | 1996-04-12 | 6/3/1998 6:51:34 PM | EMS |  | Concentration | mg/L | Y | (fl. conc.) |  | Calcium Dissolved (fl. conc.) | 7440-70-2 | Metals - Dissolved; | Metals - Dissolved; EMS code: CA-D | Chemical | Dissolved | LAB
CA-E | mg/L | Calcium Extractable | F040 | Field Acid, lab filt, ICP-AES | Dissolved and extractable water samples are ready for analysis once collected. Total samples are autoclaved with HNO3 and hydrochloric acid (HCI) in capped Teflon vials. Analysis is performed via Inductively Coupled Argon Plasma Atomic Emission Spectrometry (ICP-AES)(simultaneous multielement analysis), Graphite Furnace Atomic Absorption Spectrometry (GF-AAS)(low level analysis) and/or Flame Atomic Absorption Spectrometry (F-AAS). | 0.1 | Y | Y | Y | 9030 | 8/5/1997 6:06:00 PM | 3/23/2012 3:11:00 AM | 9/11/2019 8:47:41 AM | 1997-05-08 | 6/3/1998 6:51:46 PM | MPAK |  | Concentration | mg/L | Y | (fl. conc.) |  | Calcium Extractable (fl. conc.) | 7440-70-2 | Metals - Extractable; | Metals - Extractable; EMS code: CA-E | Chemical | Extractable | LAB
CA-E | ug/g | Calcium Extractable | CAL2 | Calculation (ug/g) | Calculation (ug/g) | 0.3 | Y | Y | Y | 15 | 1/11/2022 | 2/7/2022 | 1/11/2024 1:07:07 PM | 2022-07-06 |  | EMS |  | MassPerMass | mg/kg | Y | (mass conc.) |  | Calcium Extractable (mass conc.) | 7440-70-2 | Metals - Extractable; | Metals - Extractable; EMS code: CA-E | Chemical | Extractable | LAB
CA-E | mg/L | Calcium Extractable | X553 | Metals in water, extractable, ICP-SFMS | The whole water sample is collected in a HDPE/Teflon bottle. In a Class 100 workstation it is preserved with concentrated Nitric Acid at concentration of 0.2 percent (v/v). The acid preservative extracts the metals from material in the sample and a decanted portion is introduced directly to an inductively coupled argon plasma-sector field mass spectrometer (ICP-SFMS) by pneumatic nebulization. Each respective element is measured at a specific mass to charge ratio, m/z value expressed in amu (atomic mass units), and at a specified instrument resolution. LR: Low Resolution ~300, MR: Mid Resolution ~4000 and HR: High Resolution ~10000. Ref: Method number 02-2061, National Laboratory for Environmental Testing, CCIW, Environment Canada, Burlington. Requested April 2003. | 2.5 | Y | Y | Y | 1 |  |  | 11/28/2023 2:11:53 PM | 2023-11-28 |  | EMS |  | Concentration | mg/L | Y | (fl. conc.) |  | Calcium Extractable (fl. conc.) | 7440-70-2 | Metals - Extractable; | Metals - Extractable; EMS code: CA-E | Chemical | Extractable | LAB
CA-E | ug/g | Calcium Extractable | ASCE | Ammonium Acetate soil colloid exchange | Available metals - Ammonium Acetate soil colloid exchange sites displacement |  | Y | Y | Y | 7 | 10/31/2022 | 10/31/2022 | 11/22/2022 2:40:39 PM | 2007-10-04 |  | RDGIBSON |  | MassPerMass | mg/kg | Y | (mass conc.) |  | Calcium Extractable (mass conc.) | 7440-70-2 | Metals - Extractable; | Metals - Extractable; EMS code: CA-E | Chemical | Extractable | LAB
CA-E | mg/L | Calcium Extractable | F055 | No Pres/No Filtr/No Dig,ICP | Shortcut method which excludes preservation, filtration and digestion steps. Results are used for estimation of water hardness. | 0.1 | Y | Y | Y | 2363 | 8/23/1996 6:04:00 PM | 12/10/2011 12:01:00 PM | 9/11/2019 8:46:38 AM | 1998-04-21 | 6/3/1998 6:51:48 PM | NPEPPIN |  | Concentration | mg/L | Y | (fl. conc.) |  | Calcium Extractable (fl. conc.) | 7440-70-2 | Metals - Extractable; | Metals - Extractable; EMS code: CA-E | Chemical | Extractable | LAB
CA-E | mg/L | Calcium Extractable | F069 | SEM extr,6N HCl,ICP/OES | 6.0 N Hydrochloric acid (HCI) is added to an aliquot of the sediment sample. The extract produced from the addition of the acid is then analyzed for simultaneously extracted metals (SEM) using inductively coupled plasma - optical emission spectrophotometry. The detection limits depend on the metal parameters in question and the instrumentation used. | 0.05 | Y | Y | Y | 12 |  |  | 2/28/2022 4:37:31 PM | 1999-07-09 |  | NPEPPIN |  | Concentration | mg/L | Y | (fl. conc.) |  | Calcium Extractable (fl. conc.) | 7440-70-2 | Metals - Extractable; | Metals - Extractable; EMS code: CA-E | Chemical | Extractable | LAB
CA-E | mg/L | Calcium Extractable | E6020B | EPA 6020B | See EPA 6020B and CSSS Ch. 15 | 0.6 | Y | Y | Y | 15 | 1/11/2022 | 2/4/2022 | 1/11/2024 1:07:07 PM | 2022-06-29 |  | EMS |  | Concentration | mg/L | Y | (fl. conc.) |  | Calcium Extractable (fl. conc.) | 7440-70-2 | Metals - Extractable; | Metals - Extractable; EMS code: CA-E | Chemical | Extractable | LAB
CA-E | ug/g | Calcium Extractable | X226 | Hom,HNO3/HCl,AA | Hom;HNO3/HCl;AA | 5 | Y | Y | Y | 70 |  |  | 10/24/1996 2:25:27 PM | 1996-04-12 | 6/3/1998 6:51:34 PM | EMS |  | MassPerMass | mg/kg | Y | (mass conc.) |  | Calcium Extractable (mass conc.) | 7440-70-2 | Metals - Extractable; | Metals - Extractable; EMS code: CA-E | Chemical | Extractable | LAB
CA-O | % | Calcium (as CaO) | ICP06 | ALS Minerals ME-ICP06 | ALS Minerals ME-ICP06. A prepared sample is fused with a lithium tetraborate/lithium metaborate flux and poured into a graphite crucible and fused in a furnace at 1000°C. The resulting melt is then cooled and dissolved in 100 mL of 4% HNO3/2% HCl followed by ICP-AES. | 0.01 | Y | Y | Y | 0 |  |  |  | 2024-01-29 |  | EMS |  | DimensionlessRatio | % | N | (ratio) |  | Calcium (as CaO) (ratio) |  |  | EMS code: CA-O | Chemical |  | LAB
CA-S | mg/L | Calcium - Soluble ion content | EQ01 | EQUIS (mg/L) | EQUIS (mg/L) | 0.01 | N | Y | Y | 106 |  |  | 10/22/1996 11:12:48 PM | 1996-04-12 | 6/3/1998 6:51:34 PM | EMS |  | Concentration | mg/L | Y | (fl. conc.) |  | Calcium - Soluble ion content (fl. conc.) | 7440-70-2 | Metals - Soluble; | Metals - Soluble; EMS code: CA-S | Chemical | Soluble | LAB
CA-S | mg/L | Calcium - Soluble ion content | E6010B | EPA 6010B/CSSS 15 | See CSSS CH15/EPA 6010B/AER D50. INDUCTIVELY COUPLED PLASMA-ATOMIC EMISSION SPECTROMETRY See EPA 6010B and CSSS Ch. 15. https://www.epa.gov/sites/default/files/documents/6010b.pdf | 5 | Y | Y | Y | 0 |  |  |  | 2024-01-30 |  | EMS |  | Concentration | mg/L | Y | (fl. conc.) |  | Calcium - Soluble ion content (fl. conc.) | 7440-70-2 | Metals - Soluble; | Metals - Soluble; EMS code: CA-S | Chemical | Soluble | LAB
CA-S | mg/L | Calcium - Soluble ion content | F032 | Filt.,Acid,ICPMS | There is no acid digestion required for this analysis. A water sample is collected and preserved in 0.2 percent nitric acid. To analyse the extractable metals the sample is shaken, allowed to sit for at least twenty four hours, and a decanted portion is introduced to the spectrometer. When the dissolved metals are required, the samples must first be filtered through a 0.45 micron cellulose acetate filter paper and then followed by preservation in 0.2 percent nitric acid. It is important to note that different laboratories may have very different detection limits for some variables analyzed by ICPMS. | 0.01 | Y | Y | Y | 337 | 3/1/2002 1:15:00 PM | 2/17/2005 | 3/17/2005 1:07:53 AM | 2004-02-10 |  | UMULLER |  | Concentration | mg/L | Y | (fl. conc.) |  | Calcium - Soluble ion content (fl. conc.) | 7440-70-2 | Metals - Soluble; | Metals - Soluble; EMS code: CA-S | Chemical | Soluble | LAB
CA-S | mg/L | Calcium - Soluble ion content | 5138 | AA Analysis | AA Analysis | 0.02 | Y | Y | Y | 338 |  |  | 10/24/1996 10:11:48 AM | 1996-04-12 | 6/3/1998 6:51:34 PM | EMS |  | Concentration | mg/L | Y | (fl. conc.) |  | Calcium - Soluble ion content (fl. conc.) | 7440-70-2 | Metals - Soluble; | Metals - Soluble; EMS code: CA-S | Chemical | Soluble | LAB
CA-S | mg/L | Calcium - Soluble ion content | X170 | ICP | This code is a catch-all for a large variety of ICP methods including water and also solids digested into solutions as an intermediate result. Different procedures will have greatly different detection limits. (a) For air particulates, preliminary results are reported in mg/L and loading results as mg/dm2/d. (b) For Hi Vol samples preliminiary results are same as reported in mg/L and loading results as ug/m3. (c) Note that for Rain samples there is no filtration nor digestion steps. | 0.02 | Y | Y | Y | 3143 | 4/6/2005 | 4/11/2008 | 5/28/2008 2:08:44 AM | 1996-04-12 | 6/3/1998 6:51:34 PM | EMS |  | Concentration | mg/L | Y | (fl. conc.) |  | Calcium - Soluble ion content (fl. conc.) | 7440-70-2 | Metals - Soluble; | Metals - Soluble; EMS code: CA-S | Chemical | Soluble | LAB
CA-S | ug/g | Calcium - Soluble ion content | CAL2 | Calculation (ug/g) | Calculation (ug/g) | 0.3 | Y | Y | Y | 0 |  |  |  | 2024-02-02 |  | EMS |  | MassPerMass | mg/kg | Y | (mass conc.) |  | Calcium - Soluble ion content (mass conc.) | 7440-70-2 | Metals - Soluble; | Metals - Soluble; EMS code: CA-S | Chemical | Soluble | LAB
CA-S | mg/L | Calcium - Soluble ion content | 5070 | Ion Chromatog-Cations II | (a) For the Soluble Particulates Pkg. a portion of filtrate from the Particulate - Total procedure made up to 500 mL is analyzed by IC. Intermediate results are reported in mg/L. Loading results are calculated and reported with units of mg/dm2/d.(b) For precipitation samples, all samples are filtered through a 0.22 um fritted glass filter prior to injection. The cations of interest are separated through a cation "guard" and cation "separator" column with a methane sulphonic acid eluent. After the cations are separated they exit at various times from the column in a background of eluent. A self regeneration membrane suppressor is attached to the end of the column to neutralize the acid before detection, and form the corresponding hydroxide. Note that ammonium is reported as NH4 not as N. (c) For HiVol samples, discs cut from the HiVol filter are extracted with DI water and the soluble anions are analyzed by IC. Intermediate results are reported in units of mg/L and HiVol loading results are reported as ug/m3. | 0.01 | Y | Y | Y | 1497 | 4/30/1997 3:50:00 PM | 1/7/2002 2:33:00 PM | 6/24/2003 10:43:20 AM | 1996-04-12 | 6/3/1998 6:51:41 PM | EMS |  | Concentration | mg/L | Y | (fl. conc.) |  | Calcium - Soluble ion content (fl. conc.) | 7440-70-2 | Metals - Soluble; | Metals - Soluble; EMS code: CA-S | Chemical | Soluble | LAB
CA-S | mg/L | Calcium - Soluble ion content | E6020B | EPA 6020B | See EPA 6020B and CSSS Ch. 15 | 0.6 | Y | Y | Y | 0 |  |  |  | 2024-02-02 |  | EMS |  | Concentration | mg/L | Y | (fl. conc.) |  | Calcium - Soluble ion content (fl. conc.) | 7440-70-2 | Metals - Soluble; | Metals - Soluble; EMS code: CA-S | Chemical | Soluble | LAB
CA-S | mg/L | Calcium - Soluble ion content | 5060 | Ion Chrom- Cations I | Ion Chrom- Cations | 0.02 | Y | Y | Y | 345 | 5/6/1997 4:36:00 PM | 5/7/1997 12:02:00 PM | 2/17/1999 8:58:06 AM | 1996-04-12 | 6/3/1998 6:51:42 PM | EMS |  | Concentration | mg/L | Y | (fl. conc.) |  | Calcium - Soluble ion content (fl. conc.) | 7440-70-2 | Metals - Soluble; | Metals - Soluble; EMS code: CA-S | Chemical | Soluble | LAB
CA-T | ug/g | Calcium Total | FB04 | Freeze Dry,Hom,Acid Dig,ICP | Sample moisture is determined by weighing samples before and after freeze-drying. Freeze-dried samples are ground in a Braun blender before being weighed into Teflon PFA liners and digested with concentrated HNO3, concentrated HCl, and 50% H2O2 in a microwave oven; which results in a complete sample dissolution. Samples are collected and volumized with Type I water. Analysis is performed via Inductively Coupled Argon Plasma Atomic Emission Spectrometry (ICP-AES) (simultaneous multielement analysis), Graphite Furnace Atomic Absorption Spectrometry (GF-AAS) (low level Cd and Pb analysis), Cold Vapor Atomic Fluorescence Spectrometry (CV-AFS) (Mercury analysis) and/or Flame Atomic Absorption Spectrometry (F-AAS) (K analysis). | 8 | Y | Y | Y | 241 | 8/22/1996 6:11:00 PM | 2/1/2002 2:30:00 PM | 6/23/2003 1:15:29 PM | 1997-02-27 | 6/3/1998 6:51:46 PM | MPAK |  | MassPerMass | mg/kg | Y | (mass conc.) |  | Calcium Total (mass conc.) | 7440-70-2 | Metals - Total; | Metals - Total; EMS code: CA-T | Chemical | Total | LAB
CA-T | ug/g | Calcium Total | ICPM | Metals in Soil by CRC ICPMS | This analysis is carried out using procedures from CSR Analytical Method: "Strong Acid Leachable Metals (SALM) in Soil", BC Ministry of Environment, 26 June 2009, and procedures adapted from EPA Method 200.2. The sample is manually homogenized, dried at 60 degrees Celsius, sieved through a 63 um (230 mesh) sieve, and a representative subsample of the dry material is weighed. The sample is then digested at 95 degrees Celsius for 2 hours by block digester using concentrated nitric and hydrochloric acids. Instrumental analysis is by collision cell inductively coupled plasma - mass spectrometry (modifed from EPA Method 6020A). Method Limitation: This method is not a total digestion technique. It is a very strong acid digestion that is intended to dissolve those metals that may be environmentally available. By design, elements bound in silicate structures are not normally dissolved by this procedure as they are not usually mobile in the environment. | 50 | Y | Y | Y | 103 | 9/19/2017 | 9/24/2023 | 1/12/2024 11:16:52 AM | 2018-01-09 |  | JKROGH |  | MassPerMass | mg/kg | Y | (mass conc.) |  | Calcium Total (mass conc.) | 7440-70-2 | Metals - Total; | Metals - Total; EMS code: CA-T | Chemical | Total | LAB
CA-T | mg/L | Calcium Total | F038 | HNO3/HCl dig,ICP-AES | Dissolved and extractable water samples are ready for analysis once collected. Total samples are autoclaved with HNO3 and hydrochloric acid (HCI) in capped Teflon vials. Analysis is performed via Inductively Coupled Argon Plasma Atomic Emission Spectrometry (ICP-AES)(simultaneous multielement analysis), Graphite Furnace Atomic Absorption Spectrometry (GF-AAS) (low level analysis) and/or Flame Atomic Absorption Spectrometry (F-AAS). | 0.1 | Y | Y | Y | 7047 | 8/1/1996 6:04:00 PM | 7/15/2008 1:00:00 PM | 9/11/2019 8:47:42 AM | 1997-05-08 | 6/3/1998 6:51:46 PM | MPAK |  | Concentration | mg/L | Y | (fl. conc.) |  | Calcium Total (fl. conc.) | 7440-70-2 | Metals - Total; | Metals - Total; EMS code: CA-T | Chemical | Total | LAB
CA-T | ug/g | Calcium Total | 0190 | D&G:N/P Dig:ICP Anal. | Sample is dried and grind-mixed before nitric perchloric digestion followed by ICP analysis. | 1 | Y | Y | Y | 949 |  |  | 10/24/1996 3:06:16 PM | 1996-04-12 | 6/3/1998 6:51:34 PM | EMS |  | MassPerMass | mg/kg | Y | (mass conc.) |  | Calcium Total (mass conc.) | 7440-70-2 | Metals - Total; | Metals - Total; EMS code: CA-T | Chemical | Total | LAB
CA-T | ug/g | Calcium Total | X248 | Dig:ICP | Dig:ICP | 1 | Y | Y | Y | 47 |  |  | 10/24/1996 12:04:28 PM | 1996-04-12 | 6/3/1998 6:51:34 PM | EMS |  | MassPerMass | mg/kg | Y | (mass conc.) |  | Calcium Total (mass conc.) | 7440-70-2 | Metals - Total; | Metals - Total; EMS code: CA-T | Chemical | Total | LAB
CA-T | ug/g | Calcium Total | PB02 | PBM,HNO3/HCl dig,dry60C,100mesh, | Performance Based Method. Samples are dried at 60 C and sieved using a .149 mm (100 mesh) sieve. Solid samples are digested with the combination of nitric and hydrochloric acids with a 1:1 ratio of acids. For environmental purposes, strong acid leaches are used routinely since they have the advantage of speed, low cost, and superio reproducibility. All metals trapped within the refractory matrix are unlikely to have environmental consequences. Instrumental analysis of extracts produced using this digestion method can be performed by a variety of techniques. |  | Y | Y | Y | 10 | 9/13/2001 | 9/13/2001 | 11/6/2002 3:52:34 PM | 2001-07-18 |  | NPEPPIN |  | MassPerMass | mg/kg | Y | (mass conc.) |  | Calcium Total (mass conc.) | 7440-70-2 | Metals - Total; | Metals - Total; EMS code: CA-T | Chemical | Total | LAB
CA-T | mg/L | Calcium Total | X170 | ICP | This code is a catch-all for a large variety of ICP methods including water and also solids digested into solutions as an intermediate result. Different procedures will have greatly different detection limits. (a) For air particulates, preliminary results are reported in mg/L and loading results as mg/dm2/d. (b) For Hi Vol samples preliminiary results are same as reported in mg/L and loading results as ug/m3. (c) Note that for Rain samples there is no filtration nor digestion steps. | 0.1 | Y | Y | Y | 10775 | 3/28/1996 6:03:00 PM | 11/7/2018 | 12/8/2023 4:30:26 PM | 1996-04-12 | 8/15/2019 11:54:00 AM | EMS | KATFRASE | Concentration | mg/L | Y | (fl. conc.) |  | Calcium Total (fl. conc.) | 7440-70-2 | Metals - Total; | Metals - Total; EMS code: CA-T | Chemical | Total | LAB
CA-T | mg/L | Calcium Total | MSEA | Metals in Seawater (undigested) by (CRC) ICP-MS | This method describes the multi-elemental determination of trace elements by Collision/Reaction Cell (CRC) ICP-MS. The method measures ions produced by a radio-frequency inductively coupled plasma. Analyte species originating in a liquid are pneumatically nebulized and the resulting aerosol transported by argon gas into the plasma torch. The ions produced are entrained in the plasma gas and extracted, by means of a differentially pumped vacuum interface, into a mass spectrometer. The ions flow through a Collision/Reaction Cell (CRC) where depending on the mode of analysis either the ions passes through the cell unhindered (no gas - classical ICP-MS mode), or collisions/reactions chemistry takes place between the ions and a neutral gas such as hydrogen or helium. Any polyatomic ions which give interference on the analytes are removed. The ions are then sorted according to their mass-to-charge ratios by a quadrupole mass spectrometer. The ions are then sorted according to their mass-to-charge ratios by a quadrupole mass spectrometer. The ions transmitted through the quadrupole are quantified by a channel electron multiplier and the ion information processed by a data handling system. Although most polyatomic interference is removed, any additional interference must be assessed and valid corrections applied or the data flagged to indicate problems. Interference correction must include compensation for background ions contributed by the plasma gas, reagents, and constituents of the sample matrix. Instrument drift as well as suppressions or enhancements of instrument response caused by the sample matrix must be corrected for by the use of internal standards. | 1 | Y | Y | Y | 441 | 8/11/2009 11:57:00 PM | 4/29/2024 | 6/7/2024 10:03:16 AM | 2009-09-08 |  | RDGIBSON |  | Concentration | mg/L | Y | (fl. conc.) |  | Calcium Total (fl. conc.) | 7440-70-2 | Metals - Total; | Metals - Total; EMS code: CA-T | Chemical | Total | LAB
CA-T | mg/L | Calcium Total | F033 | Acid Dig,Centrifuge,ICPMS | For the determination of total recoverable elements in water or waste water, take 100 ml aliquot from a well mixed, acid preserved, sample containing not more than 0.25% (w/v) total solids and transfer to a 250 ml Griffin beaker. (If total solids are greater than 0.25% reduce the size of the aliquot by a proportionate amount). Add 1 ml of conc. nitric acid and 0.5 ml conc. hydrochloric acid. Heat on a hot plate until the volume has been reduced to 15-20ml, ensuring that the sample does not boil. Allow to cool and transfer to a 50 ml volumetric flask. Dilute to volume with ASTM type I water. Centrifuge the sample or allow to stand overnight to separate insoluble material. Prior analysis, pipette 20 ml into a 50 ml volumetric flask and dilute to volume with ASTM type I water. Analysis should be performed within seventy two hours of the completed preparation. Note that different laboratories may have very different field preservation, digestion procedures and detection limits for this method, so these should be documented. | 0.02 | Y | Y | Y | 26752 | 10/15/1996 9:27:00 AM | 8/22/2024 3:24:00 AM | 8/25/2024 9:47:36 PM | 1997-01-16 | 6/3/1998 6:51:46 PM | MPAK |  | Concentration | mg/L | Y | (fl. conc.) |  | Calcium Total (fl. conc.) | 7440-70-2 | Metals - Total; | Metals - Total; EMS code: CA-T | Chemical | Total | LAB
CA-T | mg/L | Calcium Total | 0042 | GI HNO3 Dig: ICP | Sample aliquot taken from unfiltered unpreserved General Ions bottle; HNO3 digestion; ICP analysis. | 0.05 | Y | Y | Y | 12830 | 9/20/1996 | 7/5/2019 3:21:00 PM | 7/12/2019 6:49:13 PM | 1996-04-12 | 6/3/1998 6:51:34 PM | EMS |  | Concentration | mg/L | Y | (fl. conc.) |  | Calcium Total (fl. conc.) | 7440-70-2 | Metals - Total; | Metals - Total; EMS code: CA-T | Chemical | Total | LAB
CA-T | mg/L | Calcium Total | X285 | HNO3 Dig:ICP Analysis | HNO3 Dig:ICP Analysis | 0.1 | Y | Y | Y | 799 | 8/10/1993 | 7/25/2012 7:11:00 AM | 9/26/2012 1:06:57 PM | 1998-05-22 | 6/3/1998 6:51:49 PM | NPEPPIN |  | Concentration | mg/L | Y | (fl. conc.) |  | Calcium Total (fl. conc.) | 7440-70-2 | Metals - Total; | Metals - Total; EMS code: CA-T | Chemical | Total | LAB
CA-T | ug/g | Calcium Total | FS89 | Homo,HNO3 dig,ICP-AES | Homo;HNO3 dig;ICP-AES | 20 | Y | Y | Y | 4 | 1/5/2001 | 8/20/2001 12:00:00 PM | 10/23/2002 1:42:41 PM | 2001-03-08 |  | NPEPPIN |  | MassPerMass | mg/kg | Y | (mass conc.) |  | Calcium Total (mass conc.) | 7440-70-2 | Metals - Total; | Metals - Total; EMS code: CA-T | Chemical | Total | LAB
CA-T | mg/L | Calcium Total | F082 | Aliquot,HNO3/HCL dig,ICPMS | Aliquot microwave digested with 3:1 HNO3/HCL. Analysis by ICPMS. For HiVol samples preliminary result are reported in mg/L and loading results are reported in ug/m3 | 0.02 | Y | Y | Y | 24803 | 1/24/2011 | 8/23/2024 | 8/26/2024 4:53:05 PM | 2015-04-15 |  | RDGIBSON |  | Concentration | mg/L | Y | (fl. conc.) |  | Calcium Total (fl. conc.) | 7440-70-2 | Metals - Total; | Metals - Total; EMS code: CA-T | Chemical | Total | LAB
CA-T | mg/L | Calcium Total | EQ01 | EQUIS (mg/L) | EQUIS (mg/L) | 0.01 | N | Y | Y | 3599 |  |  | 10/24/1996 3:16:18 PM | 1996-04-12 | 6/3/1998 6:51:34 PM | EMS |  | Concentration | mg/L | Y | (fl. conc.) |  | Calcium Total (fl. conc.) | 7440-70-2 | Metals - Total; | Metals - Total; EMS code: CA-T | Chemical | Total | LAB
CA-T | mg/L | Calcium Total | X073 | FA,HNO3,Direct Flame,AA | Samples are field-acidified with HNO3; HNO3 digestion; analysis by AA. | 0.01 | Y | Y | Y | 159 |  |  | 1/25/2007 4:06:11 PM | 1996-04-12 | 6/3/1998 6:51:34 PM | EMS |  | Concentration | mg/L | Y | (fl. conc.) |  | Calcium Total (fl. conc.) | 7440-70-2 | Metals - Total; | Metals - Total; EMS code: CA-T | Chemical | Total | LAB
CA-T | mg/dm2/d | Calcium Total | DUST | HNO3 dig.(microwave or hot block),ICPMS | HNO3 dig.(microwave or hot block);ICPMS |  | Y | Y | Y | 267 | 7/19/2016 1:11:00 PM | 7/31/2024 | 7/31/2024 11:23:12 AM | 2008-02-07 |  | RDGIBSON |  | MassFlux | mg/dm²/d | Y | (mass flux) |  | Calcium Total (mass flux) | 7440-70-2 | Metals - Total; | Metals - Total; EMS code: CA-T | Chemical | Total | LAB
CA-T | mg/L | Calcium Total | MWID | Metals in Water by ICPMS (digested) | Metals in Water by ICPMS (digested) Based on EPA 200.8 | 1 | Y | Y | Y | 36642 |  |  | 7/31/2024 1:46:41 PM | 2009-12-24 |  | RDGIBSON |  | Concentration | mg/L | Y | (fl. conc.) |  | Calcium Total (fl. conc.) | 7440-70-2 | Metals - Total; | Metals - Total; EMS code: CA-T | Chemical | Total | LAB
CA-T | % | Calcium Total | 6041 | Aqua Regia by ICPMS | These were sediment/soils digested with Aqua-Regia (Nitric Acid & Hydrochloric Acids) and then analysis by ICPMS. Results would be reported in mg/kg and/or %. | 0.01 | Y | Y | Y | 92 | 10/13/2019 | 7/27/2024 | 7/29/2024 9:13:08 AM | 2019-09-06 |  | KATFRASE |  | DimensionlessRatio | % | N | (ratio) |  | Calcium Total (ratio) | 7440-70-2 | Metals - Total; | Metals - Total; EMS code: CA-T | Chemical | Total | LAB
CA-T | mg/L | Calcium Total | F048 | Acid dig,conc'n,nebul.,ICP-OES | This method provides an automated procedure for the Inductively Coupled Argon Plasma Optical Emission Spectrophotometric (ICP-OES) determination and quantitation of trace amounts of metals in surface, ground and precipitation waters after a manual acid digestion and concentration procedure and ultrasonic nebulisation. Organically-bound metals and metal complexes are converted to metal ions by digestion with nitric acid [method details says HCl]. To increase the signal detected by the ICP, the samples are concentrated 10 times. The liquid sample is pumped onto the piezoelectric transducer, where the ultrasonic action efficiently converts it to a fine, dense aerosol. The nebulizer carrier gas flow then transports the wet aerosol through a heated U-tube, where the solvent is vaporized. Subsequently, most of the solvent vapour is removed from the aerosol stream by passing the stream through a chilled condenser. The result is a dry, analyte-laden aerosol which is introduced to the plasma. There, the high temperature desolvates the aerosol and dissociates the sample material into individual atoms and ions, which are excited to emit light at wave lengths characteristic of the elements present. Dissolved samples are field filtered through 0.45 micron cellulose acetate filter prior to acid preservation. Samples for extractable analysis are acidified with 2 mL concentrated HNO3 per litre and allowed to stand a minimum of 24 hours before analysis. Clients providing these samples are advised to preserve samples at the collection site. | 0.05 | Y | Y | Y | 231 | 11/20/1997 9:31:00 AM | 7/17/2012 7:03:00 AM | 7/26/2012 12:28:00 PM | 1998-11-13 |  | NPEPPIN |  | Concentration | mg/L | Y | (fl. conc.) |  | Calcium Total (fl. conc.) | 7440-70-2 | Metals - Total; | Metals - Total; EMS code: CA-T | Chemical | Total | LAB
CA-T | ug/g | Calcium Total | 0192 | Hom, Nitric/Perchlor,ICP | Sample already in a dry state so does not require drying before homogenization and nitric perchlorid digestion followed by ICP analysis. | 1 | Y | Y | Y | 708 |  |  | 10/24/1996 12:50:25 PM | 1996-04-12 | 6/3/1998 6:51:34 PM | EMS |  | MassPerMass | mg/kg | Y | (mass conc.) |  | Calcium Total (mass conc.) | 7440-70-2 | Metals - Total; | Metals - Total; EMS code: CA-T | Chemical | Total | LAB
CA-T | ug/g | Calcium Total | X230 | Hom,HNO3/HCl,ICP | The sample is dried, ground to pass 100 mesh, and digested with a nitric-hydrochloric acid mixture to solubilize the solid matter and to remove the organic material by oxidation and volatilization. The completed digests are submitted to metals instrumentation personnel for metals analysis by Inductively Coupled Atomic Emission Spectrophotometry (ICP-AES), or hydride analysis by atomic absorption. | 10 | Y | Y | Y | 45 | 2/12/2002 | 12/30/2005 | 2/3/2006 1:13:12 AM | 1998-11-05 |  | NPEPPIN |  | MassPerMass | mg/kg | Y | (mass conc.) |  | Calcium Total (mass conc.) | 7440-70-2 | Metals - Total; | Metals - Total; EMS code: CA-T | Chemical | Total | LAB
CA-T | ug/g wet | Calcium Total | X560 | Metals in Biota (wet), total, acid digestion, ICP-MS | Five grams of wet, homogenized biota is weighed in a centrifuge tube, and 2.5 mL of Nitric Acid and 2.5 mL of Hydrochloric acid are added. The sample is digested at room temperature for one hour and then digested at 90-95 degrees C in a water bath or heat block for one more hour. The sample is cooled and the volume is made to 50 mL with DI. The digestate is then diluted at ten fold and introduced to a collision/reaction cell inductively coupled argon plasma mass spectrometer (CRC ICPMS) by pneumatic nebulization. Each respective element is measured at a specific mass to charge ratio, m/z value expressed in amu (atomic mass units). - Maxxam Analytics (BC). Added to EMS Dec 10, 2008. | 10 | Y | Y | Y | 541 | 11/15/2008 4:00:00 AM | 10/30/2023 | 11/28/2023 2:21:33 PM | 2008-12-11 | 3/11/2015 4:17:10 PM | RDGIBSON | EMS | MassPerMass | µg/g | Y | (mass conc.) | (wet) | Calcium Total (wet) (mass conc.) | 7440-70-2 | Metals - Total; | Metals - Total; EMS code: CA-T | Chemical | Total | LAB
CA-T | mg/L | Calcium Total | METN | MET-T-NP-CCMS-VA | Ultra trace metals in water are analyzed by CRC ICPMS, based on US EPA Method 6020B (July 2014). The detection limits provided can only be met for undigested samples. This procedure is intended for colorless, non-turbid, acid-preserved water samples (i.e. pristine water samples), having turbidity < 1 NTU and no odor. Where turbidity exceeds 1 NTU, and/or the sample is colored and has an odor, results may be biased low compared to true Total Metals concentrations. ALS recommends that turbidity analysis be requested on samples submitted for this test to aid with interpretation of results. Where turbidity is <1NTU, undigested metals are equivalent to total metals concentrations. | 0.01 | Y | Y | Y | 3028 | 11/14/2017 | 8/19/2024 | 8/26/2024 1:03:06 PM | 2019-08-20 |  | KATFRASE |  | Concentration | mg/L | Y | (fl. conc.) |  | Calcium Total (fl. conc.) | 7440-70-2 | Metals - Total; | Metals - Total; EMS code: CA-T | Chemical | Total | LAB
CA-T | % | Calcium Total | MS61 | ALS Minerals ME-MS61 | ALS Minerals ME-MS61. A prepared sample is digested with HNO3; HF; and HClO4 in a graphite heating block to near dryness. After cooling; the sample is boiled in diluted HCl. Bulk to volume with deionized water; mixed and analyzed by ICP-AES and ICP-MS for all elements. | 0.01 | Y | Y | Y | 0 |  |  |  | 2024-01-29 |  | EMS |  | DimensionlessRatio | % | N | (ratio) |  | Calcium Total (ratio) | 7440-70-2 | Metals - Total; | Metals - Total; EMS code: CA-T | Chemical | Total | LAB
CA-T | mg/L | Calcium Total | 0040 | HNO3 Dig: ICP Analysis | HNO3 Dig: ICP Analysis | 0.02 | Y | Y | Y | 12366 |  |  | 3/23/2010 5:08:08 PM | 1996-04-12 | 6/3/1998 6:51:34 PM | EMS |  | Concentration | mg/L | Y | (fl. conc.) |  | Calcium Total (fl. conc.) | 7440-70-2 | Metals - Total; | Metals - Total; EMS code: CA-T | Chemical | Total | LAB
CA-T | % | Calcium Total | MS41 | ALS Minerals ME-MS41 | ALS Minerals ME-MS41. A prepared sample is digested with aqua regia in a graphite heating block. After cooling; the resulting solution is diluted to with deionized water; mixed and analyzed by ICP-AES. Samples are then analyzed by ICP-MS for the remaining suite of elements. | 0.01 | Y | Y | Y | 0 |  |  |  | 2024-01-29 |  | EMS |  | DimensionlessRatio | % | N | (ratio) |  | Calcium Total (ratio) | 7440-70-2 | Metals - Total; | Metals - Total; EMS code: CA-T | Chemical | Total | LAB
CA-T | ug/g | Calcium Total | FS82 | PBM,dry60C,10mesh,SALM-ICP-MS | PBM;dry60C;10mesh;SALM-ICP-MS | 100 | Y | Y | Y | 3307 | 5/21/2002 | 7/24/2024 | 7/29/2024 5:17:09 PM | 2002-11-12 |  | UMULLER |  | MassPerMass | mg/kg | Y | (mass conc.) |  | Calcium Total (mass conc.) | 7440-70-2 | Metals - Total; | Metals - Total; EMS code: CA-T | Chemical | Total | LAB
CA-T | mg/L | Calcium Total | CS01 | System Calculation | System Calculation (mg/L) |  | N | Y | N | 214 |  |  | 9/2/2020 1:59:11 PM | 1996-04-12 | 6/3/1998 6:51:43 PM | EMS |  | Concentration | mg/L | Y | (fl. conc.) |  | Calcium Total (fl. conc.) | 7440-70-2 | Metals - Total; | Metals - Total; EMS code: CA-T | Chemical | Total | LAB
CA-T | ug/g | Calcium Total | FS01 | D&G,Nitric/HCl:ICP | Samples are dried at 60 C, then either sieved to -100mesh (150 um) and rolled to homogenize, or finely ground with a mortar and pestle or a shatter box laboratory mill and rolled (depending on the submitter's requirements). The samples are then weighed into Teflon digestion liners and digested with HNO3 and HCl in a microwave oven. Samples are cooled, volumized, and allowed to settle overnight before being decanted. Analysis is then performed via Inductively Coupled Argon Plasma Atomic Emission Spectrometry (ICP-AES) (Simultaneous multi element analysis). In ICP Spectrometry, the liquid sample is pumped to the nebulizer, where it is convertd into an aerosol. the nebulizer carrier gas flow than transports the wet aerosol past the impact bead, where the large droplets condense out and flow to the waste container. The remaining fine aerosol is carried into the plasma. There, the high temperature desolvates the aerosol and dissociates the sample material into individual atoms and ions, which are excited to emit light at wavelengths characteristic of the elements present. A diffraction grating in the spectrometer diffracts the various wavelengths of light through exit slits to the photo-multiplier tubes, which convert the light energy to electronic signals. | 20 | Y | Y | Y | 2642 | 8/30/1996 1:52:00 PM | 10/10/2007 | 5/24/2008 1:15:02 AM | 1996-11-20 | 6/3/1998 6:51:46 PM | MPAK |  | MassPerMass | mg/kg | Y | (mass conc.) |  | Calcium Total (mass conc.) | 7440-70-2 | Metals - Total; | Metals - Total; EMS code: CA-T | Chemical | Total | LAB
CA-T | mg/L | Calcium Total | 1349 | Tot Metals in Seawater by CRC ICPMS (BC) | Tot Metals in Seawater by CRC ICPMS (BC) | 1 | Y | Y | Y | 98 | 8/16/2019 | 12/5/2021 | 12/13/2021 12:34:16 PM | 2019-09-30 |  | KATFRASE |  | Concentration | mg/L | Y | (fl. conc.) |  | Calcium Total (fl. conc.) | 7440-70-2 | Metals - Total; | Metals - Total; EMS code: CA-T | Chemical | Total | LAB
CA-T | ug/g | Calcium Total | EQ02 | EQUIS (ug/g) | EQUIS (ug/g) | 1 | N | Y | Y | 780 |  |  | 10/23/1996 12:08:28 AM | 1996-04-12 | 6/3/1998 6:51:34 PM | EMS |  | MassPerMass | mg/kg | Y | (mass conc.) |  | Calcium Total (mass conc.) | 7440-70-2 | Metals - Total; | Metals - Total; EMS code: CA-T | Chemical | Total | LAB
CA-T | ug/g | Calcium Total | X551 | Metals in Biota (dry), total, acid digestion, ICP-SFMS | 0.1 gram of dried, homogenized biota is weighed in a microwave TFM vessel, and 8mL of Nitric Acid and 1mL of Hydrogen Peroxide are added. The sample is digested at 200 degrees C for 15 minutes in a closed vessel high pressure microwave oven. The sample is cooled, the volume adjusted to 100mL and introduced to an iductively coupled argon plasma-sector field mass spectrometer (ICP-SFMS) by pneumatic nebulization. Each respective element is measured at a specific mass to charge ratio, m/z value expressed in amu (atomic mass units), and at a specified instrument resolution. LR: Low Resolution ~300, MR: Mid Resolution ~4000 and HR: Hich Resolution ~10000. Ref: Method number 02-2705, National Laboratory for Environmental Testing, CCIW, Environment Canada, Burlington. Requested April 2003. | 20 | Y | Y | Y | 49 | 8/11/2015 | 10/30/2023 | 11/28/2023 2:21:33 PM | 2015-03-11 |  | RDGIBSON |  | MassPerMass | mg/kg | Y | (mass conc.) |  | Calcium Total (mass conc.) | 7440-70-2 | Metals - Total; | Metals - Total; EMS code: CA-T | Chemical | Total | LAB
CA-T | mg/L | Calcium Total | X349 | FA:HNO3:ICP | Samples are field-acidified with HNO3; HNO3 digestion; analysis by ICP/AES. | 0.5 | Y | Y | Y | 6492 |  |  | 8/26/2024 2:41:32 PM | 1996-04-12 | 6/3/1998 6:51:34 PM | EMS |  | Concentration | mg/L | Y | (fl. conc.) |  | Calcium Total (fl. conc.) | 7440-70-2 | Metals - Total; | Metals - Total; EMS code: CA-T | Chemical | Total | LAB
CA-T | mg/L | Calcium Total | F121 | Acid,ICPMS | Acidified, no digestion, ICPMS | 0.01 | Y | Y | Y | 10204 | 10/18/2007 | 8/16/2024 | 8/26/2024 8:06:03 AM | 2002-07-17 | 7/16/2012 3:33:24 PM | UMULLER | RDGIBSON | Concentration | mg/L | Y | (fl. conc.) |  | Calcium Total (fl. conc.) | 7440-70-2 | Metals - Total; | Metals - Total; EMS code: CA-T | Chemical | Total | LAB
CA-T | ug/g wet | Calcium Total | MTIW | MET-WET-CCMS-MID-VA | Metals in Tissue by CRC ICPMS (WET) Methodology: This method is conducted following British Columbia Lab Manual method "Metals in Animal Tissue and Vegetation (Biota) - Prescriptive". Tissue samples are homogenized and sub-sampled prior to hotblock digestion with nitric and hydrochloric acids, in combination with addition of hydrogen peroxide. Instrumental analysis is by collision cell inductively coupled plasma - mass spectrometry (modified from EPA Method 6020A). Method Limitation: This method employs a strong acid/peroxide digestion, and is intended to provide a conservative estimate of bio-available metals. Near complete recoveries are achieved for most toxicologically important metals, but elements associated with recalcitrant minerals may be only partially recovered. | 4 | Y | Y | Y | 24 | 11/25/2018 | 12/6/2018 | 7/21/2021 2:55:07 PM | 2018-12-24 |  | EMS |  | MassPerMass | µg/g | Y | (mass conc.) | (wet) | Calcium Total (wet) (mass conc.) | 7440-70-2 | Metals - Total; | Metals - Total; EMS code: CA-T | Chemical | Total | LAB
CA-T | mg/L | Calcium Total | X561 | Total metals: FA, HNO3, HR-ICPMS (No Prep)Prep) | Ultra trace metals in water are analyzed by high resolution inductively coupled plasma mass spectrometry (HR-ICPMS), modified from US EPA 200.8, (Revision 5.5). The detection limits provided can only be met for undigested samples. This procedure is intended for colorless, non-turbid, acid-preserved water samples (i.e. pristine water samples), having turbidity < 1 NTU and no odor. Where turbidity exceeds 1 NTU, and/or the sample is colored and has an odor, results may be biased low compared to true Total Metals concentrations. ALS recommends that turbidity analysis be requested on samples submitted for this test to aid with interpretation of results. | 0.01 | Y | Y | Y | 2756 | 6/8/2015 | 10/5/2023 | 1/31/2024 2:37:41 PM | 2015-10-22 |  | RDGIBSON |  | Concentration | mg/L | Y | (fl. conc.) |  | Calcium Total (fl. conc.) | 7440-70-2 | Metals - Total; | Metals - Total; EMS code: CA-T | Chemical | Total | LAB
CA-T | ug/g wet | Calcium Total | BPSY | MET-WET-BPSY-CCMS-VA | Metals in Animal Tissue and Vegetation (Biota) - Prescriptive. EPA 200.3/6020B This method is conducted following British Columbia Lab Manual method "Metals in Animal Tissue and Vegetation (Biota) - Prescriptive". Tissue samples are digested with nitric and hydrochloric acids; in combination with addition of hydrogen peroxide. Instrumental analysis is by collision cell inductively coupled plasma - mass spectrometry (modified from EPA Method 6020B). | 4 | Y | Y | Y | 2 | 11/7/2019 | 11/7/2019 | 6/10/2021 1:57:10 PM | 2021-06-04 |  | EMS |  | MassPerMass | µg/g | Y | (mass conc.) | (wet) | Calcium Total (wet) (mass conc.) | 7440-70-2 | Metals - Total; | Metals - Total; EMS code: CA-T | Chemical | Total | LAB
CA-T | ug/g | Calcium Total | MTID | MET-DRY-CCMS-VA | Metals in Tissue by CRC ICPMS (DRY) Methodology: This method is conducted following British Columbia Lab Manual method "Metals in Animal Tissue and Vegetation (Biota) - Prescriptive". Tissue samples are homogenized and sub-sampled prior to hotblock digestion with nitric and hydrochloric acids, in combination with addition of hydrogen peroxide. Instrumental analysis is by collision cell inductively coupled plasma - mass spectrometry (modified from EPA Method 6020A). Method Limitation: This method employs a strong acid/peroxide digestion, and is intended to provide a conservative estimate of bio-available metals. Near complete recoveries are achieved for most toxicologically important metals, but elements associated with recalcitrant minerals may be only partially recovered. | 20 | Y | Y | Y | 122 | 8/13/2014 | 9/13/2022 | 10/26/2022 7:14:14 AM | 2015-03-09 | 3/11/2015 4:12:05 PM | RDGIBSON | EMS | MassPerMass | mg/kg | Y | (mass conc.) |  | Calcium Total (mass conc.) | 7440-70-2 | Metals - Total; | Metals - Total; EMS code: CA-T | Chemical | Total | LAB
CA-T | mg/L | Calcium Total | 0020 | HNO3 Dig:AA Analysis | HNO3 Dig:AA Analysis | 0.02 | Y | Y | Y | 98 | 6/25/2010 3:15:00 PM | 12/1/2010 | 12/9/2010 3:55:26 PM | 1996-04-12 | 6/3/1998 6:51:41 PM | EMS |  | Concentration | mg/L | Y | (fl. conc.) |  | Calcium Total (fl. conc.) | 7440-70-2 | Metals - Total; | Metals - Total; EMS code: CA-T | Chemical | Total | LAB
CA-T | mg/dm2/d | Calcium Total | X286 | ICP | Aliquot of prepared sample digested with nitric acid then with perchloric acid. Filtered digest analyzed by ICP. Intermediate results reported as mg/L. Loading results for air particulate reported as mg/dm2/d. | 0 | Y | Y | Y | 23 |  |  | 10/24/1996 11:21:52 AM | 1996-04-12 | 6/3/1998 6:51:41 PM | EMS |  | MassFlux | mg/dm²/d | Y | (mass flux) |  | Calcium Total (mass flux) | 7440-70-2 | Metals - Total; | Metals - Total; EMS code: CA-T | Chemical | Total | LAB
CA-T | mg/L | Calcium Total | X145 | Dig, ICP | Dig; ICP | 0.02 | Y | Y | Y | 885 | 5/8/2003 9:40:00 AM | 9/22/2016 | 6/13/2024 3:58:21 PM | 1996-04-12 | 6/3/1998 6:51:41 PM | EMS |  | Concentration | mg/L | Y | (fl. conc.) |  | Calcium Total (fl. conc.) | 7440-70-2 | Metals - Total; | Metals - Total; EMS code: CA-T | Chemical | Total | LAB
CA-T | ug/g | Calcium Total | 0196 | D&G: LTA/HCl: ICP | D&G: LTA/HCl: ICP | 1 | Y | Y | Y | 27 |  |  | 10/24/1996 2:08:12 PM | 1996-04-12 | 6/3/1998 6:51:34 PM | EMS |  | MassPerMass | mg/kg | Y | (mass conc.) |  | Calcium Total (mass conc.) | 7440-70-2 | Metals - Total; | Metals - Total; EMS code: CA-T | Chemical | Total | LAB
CA-T | mg/L | Calcium Total | 0200 | 20%A.R.Dig:AA Anal. | 20%A.R.Dig:AA Anal. | 0.02 | Y | Y | Y | 148 |  |  | 10/24/1996 11:35:00 AM | 1996-04-12 | 6/3/1998 6:51:34 PM | EMS |  | Concentration | mg/L | Y | (fl. conc.) |  | Calcium Total (fl. conc.) | 7440-70-2 | Metals - Total; | Metals - Total; EMS code: CA-T | Chemical | Total | LAB
CA-T | ug/g wet | Calcium Total | X552 | Metals in Biota (wet), total, acid digestion, ICP-SFMS | One gram of wet, homogenized biota is weighed in a microwave TFM vessel, and 8mL of Nitric Acid and 1mL of Hydrogen Peroxide are added. The sample is digested at 200 degrees C for 15 minutes in a closed vessel high pressure microwave oven. The sample is cooled, the volume adjusted to 100mL and introduced to an inductively coupled argon plasma-sector field mass spectrometer (ICP-SFMS) by pneumatic nebulization. Each respective element is measured at a specific mass to charge ratio, m/z value expressed in amu (atomic mass units), and at a specified instrument resolution. LR: Low Resolution ~300, MR: Mid Resolution ~4000. Ref: 02-2704, National Laboratory for Environmental Testing, CCIW, Environment Canada, Burlington. Requested April 2003. | 4 | Y | Y | Y | 37 | 8/9/2016 | 10/16/2023 | 11/28/2023 2:21:32 PM | 2015-03-11 |  | RDGIBSON |  | MassPerMass | µg/g | Y | (mass conc.) | (wet) | Calcium Total (wet) (mass conc.) | 7440-70-2 | Metals - Total; | Metals - Total; EMS code: CA-T | Chemical | Total | LAB
CABR | % | Ion Balance Ratio % =(Cation sum / Anion sum])*100% | XCAL | Calculation (mg/L) | Calculation (mg/L) | 0.01 | Y | Y | Y | 72 | 1/9/2020 | 6/6/2024 | 6/10/2024 2:03:06 PM | 2020-01-21 | 1/21/2020 8:38:35 AM | RDGIBSON | RDGIBSON | DimensionlessRatio | % | N | (ratio) |  | Ion Balance Ratio % =(Cation sum / Anion sum])*100% (ratio) |  | Major Ions; | Major Ions; EMS code: CABR | Chemical |  | LAB
CAFE | mg/L | Caffeine | XMSP | HPLC and Tri. Quadpole MS/MS | EPA 1694 An aliquot of filtered sample containing 5% organic is injected directly. Each compound is separated by reversed phase HPLC. Analyzed by Electro Spray Ionization MS/MS detection using Triple Quadrupole MS/MS detector. Quantiitation with internal or external standard. | 5e-06 | Y | Y | Y | 32 | 12/14/2017 | 6/4/2024 | 8/21/2024 3:32:03 PM | 2018-01-09 |  | JKROGH |  | Concentration | mg/L | Y | (fl. conc.) |  | Caffeine (fl. conc.) |  | Pharmaceuticals and Personal Care Products (PPCP); | Pharmaceuticals and Personal Care Products (PPCP); EMS code: CAFE | Chemical |  | LAB
CAPA | mg/L | Caproic Acid | VFAD | Volatile Fatty Acids by GC/FID | Volatile Fatty Acids by GC/FID; Standard Methods for the Examination of Water and Wastewater; 21st Edition, Method 5560 D; pages 5-59 to 5-61. 5560 D. Gas Chromatographic Method Principle: This gas chromatographic procedure may be usedto determine the individual concentrations of the following fattyacids: acetic; propionic; butyric; isobutyric; valeric; andisovaleric acids in digester sludge samples. Caproic and heptanoicacids also may be determined; but with less certaintybecause of their reduced solubility in water. Volatile fatty acidsare generally soluble in digester sludge and exist in either theprotonated or unprotonated form; depending on the pH of thesludge sample. Acetic and propionic acids usually are in higherconcentrations than the other fatty acids in digester sludge. Allacids must be converted to the protonated (volatile) form beforeinjection into the gas chromatograph (GC) so vaporization mayoccur. The fatty acids must also be separated from the solidsmaterial in the sample before introduction into the GC to minimizedegradation of the GC column. Digester sludge samples areprepared by acidification; centrifugation; and filtration. The sampleis analyzed by direct injection into a gas chromatographequipped with a flame ionization detector after both centrifugationand filtration.b. Interference: The blank amount may be high relative tosample analyte. High blanks may be caused by a buildup ofcontaminants in the injector and guard column or by samplecarry-over. The analyst must be aware of the blank levels andcorrect unacceptably high blank levels that are above the reportinglimits.c. Minimum detectable concentration: The method detectionlevel has been determined by the USEPA method1 in reagentwater to be about 3 mg/L for acetic acid and about 1 mg/L for allother target compounds. | 2 | Y | Y | Y | 5 | 3/14/2017 10:30:00 AM | 11/19/2017 5:10:00 PM | 6/20/2018 9:02:11 AM | 2018-06-19 |  | RDGIBSON |  | Concentration | mg/L | Y | (fl. conc.) |  | Caproic Acid (fl. conc.) |  | Organics; | Organics; EMS code: CAPA | Chemical |  | LAB
CARB | mg/L | Carbamazepine CAS: 298-46-4 | XMSP | HPLC and Tri. Quadpole MS/MS | EPA 1694 An aliquot of filtered sample containing 5% organic is injected directly. Each compound is separated by reversed phase HPLC. Analyzed by Electro Spray Ionization MS/MS detection using Triple Quadrupole MS/MS detector. Quantiitation with internal or external standard. |  | Y | Y | Y | 12 | 5/8/2020 | 6/4/2024 | 8/21/2024 3:32:03 PM | 2022-05-04 |  | EMS |  | Concentration | mg/L | Y | (fl. conc.) |  | Carbamazepine (fl. conc.) | 298-46-4 | Pesticides; | Pesticides; EMS code: CARB | Chemical |  | LAB
CATP | mg/L | Chlorantraniliprole | E751 | E751D Pesticides in Water by LC-MS/MS (POS1 List) | E751D Pesticides in Water by LC-MS/MS (POS1 List) (MOECC E3501) | 5e-05 | Y | Y | Y | 39 | 12/9/2021 | 12/31/2021 | 2/15/2022 9:50:05 AM | 2021-12-14 |  | EMS |  | Concentration | mg/L | Y | (fl. conc.) |  | Chlorantraniliprole (fl. conc.) |  | Pesticides; | Pesticides; EMS code: CATP | Chemical |  | LAB
CBAL | % | Cation-Anion Balance %=100*(Cations-Anions)/(Cations+Anions) | XCAL | Calculation (mg/L) | Calculation (mg/L) |  | Y | Y | Y | 4512 | 6/3/2013 | 8/15/2024 | 8/26/2024 9:09:02 AM | 2016-04-05 |  | RDGIBSON |  | DimensionlessRatio | % | N | (ratio) |  | Cation-Anion Balance %=100*(Cations-Anions)/(Cations+Anions) (ratio) |  |  | EMS code: CBAL | Chemical |  | LAB
CBOD | mg/L | Carbonaceous Biochemical Oxygen Demand | X013 | Seeded Dil Nit Inh 5 Day | The biochemical oxygen demand is determined by measuring the amount of dissolved oxygen (DO) that is consumed in the sample under aerobic conditions during the incubation period (five days at 20 C) after seeding material has been added. A polarographic membrane electrode that has been calibrated in saturated air is used to acquire all dissolved oxygen measurements. Nitrification inhibitor added. A modified Winkler procedure may be substituted for the DO probe procedure. Performance Base Method published circa Dec. 2000. | 5 | Y | Y | Y | 5942 | 5/8/2003 9:40:00 AM | 8/21/2024 11:55:00 AM | 8/26/2024 2:48:03 PM | 2000-10-24 |  | NPEPPIN |  | Concentration | mg/L | Y | (fl. conc.) |  | Carbonaceous Biochemical Oxygen Demand (fl. conc.) |  | Aggregate Organics; | Aggregate Organics; EMS code: CBOD | Chemical |  | LAB
CBOD | mg/L | Carbonaceous Biochemical Oxygen Demand | X140 | Unseeded Dil Nit Inh 5 Days | The biochemical oxygen demand is determined by measuring the amount of dissolved oxygen (DO) that is consumed in the sample under aerobic conditions during the incubation period (five days at 20 C). A polarographic membrane electrode that has been calibrated in saturated air is used to acquire all dissolved oxygen measurements. Nitrification inhibitor added. No seed is added. A modified Winkler procedure may be substituted for the DO probe procedure. Performance Base Method published circa Dec. 2000. | 5 | Y | Y | Y | 3285 |  |  | 7/23/2024 11:41:51 AM | 2000-10-24 |  | NPEPPIN |  | Concentration | mg/L | Y | (fl. conc.) |  | Carbonaceous Biochemical Oxygen Demand (fl. conc.) |  | Aggregate Organics; | Aggregate Organics; EMS code: CBOD | Chemical |  | LAB
CD-D | mg/L | Cadmium Dissolved | X095 | FF,FA,HNO3,ICP | FF;FA;HNO3;ICP | 0.005 | Y | Y | Y | 191 |  |  | 10/24/1996 12:20:01 PM | 1996-04-12 | 6/3/1998 6:51:34 PM | EMS |  | Concentration | mg/L | Y | (fl. conc.) |  | Cadmium Dissolved (fl. conc.) | 7440-43-9 | Metals - Dissolved; | Metals - Dissolved; EMS code: CD-D | Chemical | Dissolved | LAB
CD-D | mg/L | Cadmium Dissolved | F049 | Filt,HNO3 dig,conc'n,nebul'ICP-OES | This method provides an automated procedure for the Inductively Coupled Argon Plasma Optical Emission Spectrophotometric (ICP-OES) determination and quantitation of trace amounts of metals in surface, ground and precipitation waters after a manual acid digestion and concentration procedure and ultrasonic nebulisation. Organically-bound metals and metal complexes are converted to metal ions by digestion with nitric acid. To increase the signal detected by the ICP, the samples are concentrated 10 times. The liquid sample is pumped onto the piezoelectric transducer, where the ultrasonic action efficiently converts it to a fine, dense aerosol. The nebulizer carrier gas flow then transports the wet aerosol through a heated U-tube, where the solvent is vaporized. Subsequently, most of the solvent vapour is removed from the aerosol stream by passing the stream through a chilled condenser. The result is a dry, analyte-laden aerosol which is introduced to the plasma. There, the high temperature desolvates the aerosol and dissociates the sample material into individual atoms and ions, which are excited to emit light at wave lengths characteristic of the elements present. Dissolved samples are field filtered through 0.45 micron cellulose acetate filter prior to acid preservation. Samples for extractable analysis are acidified with 2 mL concentrated HNO3 per litre and allowed to stand a minimum of 24 hours before analysis. Clients providing these samples are advised to preserve samples at the collection site. | 0.0001 | Y | Y | Y | 1455 |  |  | 5/19/2015 3:10:33 PM | 1997-10-20 | 6/3/1998 6:51:47 PM | NPEPPIN |  | Concentration | mg/L | Y | (fl. conc.) |  | Cadmium Dissolved (fl. conc.) | 7440-43-9 | Metals - Dissolved; | Metals - Dissolved; EMS code: CD-D | Chemical | Dissolved | LAB
CD-D | mg/L | Cadmium Dissolved | 0051 | Graphite Furnace AA | Graphite Furnace | 0.0001 | Y | Y | Y | 1435 | 5/9/2003 9:40:00 AM | 7/15/2003 12:03:00 PM | 8/14/2003 5:26:02 PM | 1996-04-12 | 6/3/1998 6:51:34 PM | EMS |  | Concentration | mg/L | Y | (fl. conc.) |  | Cadmium Dissolved (fl. conc.) | 7440-43-9 | Metals - Dissolved; | Metals - Dissolved; EMS code: CD-D | Chemical | Dissolved | LAB
CD-D | mg/L | Cadmium Dissolved | X544 | FF,FA,HNO3,ICPMS | Analytes nebulized followed by ion entrained in plasma gas and extracted by means of a differentially pumped vacuum interface into a mass spectrometer | 0.0001 | Y | Y | Y | 30 |  |  | 10/30/2014 8:31:17 AM | 2001-09-18 | 9/18/2001 1:36:23 PM | NPEPPIN | NPEPPIN | Concentration | mg/L | Y | (fl. conc.) |  | Cadmium Dissolved (fl. conc.) | 7440-43-9 | Metals - Dissolved; | Metals - Dissolved; EMS code: CD-D | Chemical | Dissolved | LAB
CD-D | mg/L | Cadmium Dissolved | X285 | HNO3 Dig:ICP Analysis | HNO3 Dig:ICP Analysis | 0.01 | Y | Y | Y | 951 | 1/30/2002 | 11/13/2002 | 3/18/2004 10:28:34 AM | 1996-04-12 | 6/3/1998 6:51:34 PM | EMS |  | Concentration | mg/L | Y | (fl. conc.) |  | Cadmium Dissolved (fl. conc.) | 7440-43-9 | Metals - Dissolved; | Metals - Dissolved; EMS code: CD-D | Chemical | Dissolved | LAB
CD-D | mg/L | Cadmium Dissolved | 0090 | Ext.(APDC):AA Anal. | Ext.(APDC):AA Anal. | 0.0005 | Y | Y | Y | 55 |  |  | 10/24/1996 11:35:00 AM | 1996-04-12 | 6/3/1998 6:51:34 PM | EMS |  | Concentration | mg/L | Y | (fl. conc.) |  | Cadmium Dissolved (fl. conc.) | 7440-43-9 | Metals - Dissolved; | Metals - Dissolved; EMS code: CD-D | Chemical | Dissolved | LAB
CD-D | mg/L | Cadmium Dissolved | X072 | FA,HNO3,Graph.Furnace,AA | Samples are field-acidified with HNO3; HNO3 digestion; analysis by GFAA. | 0.0005 | Y | Y | Y | 956 |  |  | 10/29/2001 11:34:41 AM | 1996-04-12 | 6/3/1998 6:51:41 PM | EMS |  | Concentration | mg/L | Y | (fl. conc.) |  | Cadmium Dissolved (fl. conc.) | 7440-43-9 | Metals - Dissolved; | Metals - Dissolved; EMS code: CD-D | Chemical | Dissolved | LAB
CD-D | mg/L | Cadmium Dissolved | X113 | HNO3/HCl/Perchlor,AA | HNO3/HCl/Perchlor;AA | 0.0005 | Y | Y | Y | 1216 |  |  | 12/9/1996 10:29:28 PM | 1996-04-12 | 6/3/1998 6:51:41 PM | EMS |  | Concentration | mg/L | Y | (fl. conc.) |  | Cadmium Dissolved (fl. conc.) | 7440-43-9 | Metals - Dissolved; | Metals - Dissolved; EMS code: CD-D | Chemical | Dissolved | LAB
CD-D | mg/L | Cadmium Dissolved | F023 | HNO3,GFAA | Dissolved and extractable water samples are ready for analysis once collected. Total samples are autoclaved with HNO3 and hydrochloric acid (HCl) in capped Teflon vials. Analysis is performed via Inductively Coupled Argon Plasma Atomic Emission Spectrometry (ICP-AES) (simultaneous multielement analysis), Graphite Furnace Atomic Absorption Spectrometry (GF-AAS) (low level analysis) and/or Flame Atomic Absorption Spectrometry (F-AAS). Note that extractable samples are acidified with 0.5 mL HNO3 per 100 mL of sample at time of sample collection. | 0.0002 | Y | Y | Y | 393 | 8/6/1996 7:16:00 PM | 9/16/2003 3:18:00 PM | 11/28/2003 3:22:10 PM | 1996-11-20 | 6/3/1998 6:51:46 PM | MPAK |  | Concentration | mg/L | Y | (fl. conc.) |  | Cadmium Dissolved (fl. conc.) | 7440-43-9 | Metals - Dissolved; | Metals - Dissolved; EMS code: CD-D | Chemical | Dissolved | LAB
CD-D | mg/L | Cadmium Dissolved | F068 | HNO3dig,TeflonTube/Autoclave,ICP-MS | This method is applicable to determiniation of dissolved, extractable, and total recoverable metals in ground waters, surface waters, drinking waters, and wastewaters. Dissolved, extractable, and total recoverable sample fractions are acidified to pH <1.5 with analytical grade nitric acid (HNO3) to prevent the metals from adsorbing to the container material or precipitating. Extractable and total recoverable water samples are collected in 250 ml acid washed polyethylene bottles and are preserved with Nitric acid (0.5 ml acid for 100 mLs of sample). All preserved samples are stored at room temperature and can be held for up to 6 months. Unfiltered, un-preserved samples are stored at 4C for no more than 30 days. For dissolved and extractable water fractions, no sample preparation is necessary. Samples are simply poured in autosampler tubes (i.e. 15 ml centrifuge tubes), and placed on the autosampler and analyzed. Samples for total recoverable metals are digested with Nitric acid in special Teflon tubes in an autoclave prior to analyses. | 1e-05 | Y | Y | Y | 3006 | 6/1/1999 1:20:00 PM | 4/17/2015 1:06:00 PM | 7/18/2024 8:07:20 AM | 1999-06-28 |  | NPEPPIN |  | Concentration | mg/L | Y | (fl. conc.) |  | Cadmium Dissolved (fl. conc.) | 7440-43-9 | Metals - Dissolved; | Metals - Dissolved; EMS code: CD-D | Chemical | Dissolved | LAB
CD-D | mg/L | Cadmium Dissolved | X287 | FA,HNO3,Graph.F.,AA(LL) | FA;HNO3;Graph.F.;AA(LL) | 0.0001 | Y | Y | Y | 90 |  |  | 10/24/1996 12:45:10 PM | 1996-04-12 | 6/3/1998 6:51:34 PM | EMS |  | Concentration | mg/L | Y | (fl. conc.) |  | Cadmium Dissolved (fl. conc.) | 7440-43-9 | Metals - Dissolved; | Metals - Dissolved; EMS code: CD-D | Chemical | Dissolved | LAB
CD-D | mg/L | Cadmium Dissolved | X170 | ICP | This code is a catch-all for a large variety of ICP methods including water and also solids digested into solutions as an intermediate result. Different procedures will have greatly different detection limits. (a) For air particulates, preliminary results are reported in mg/L and loading results as mg/dm2/d. (b) For Hi Vol samples preliminiary results are same as reported in mg/L and loading results as ug/m3. (c) Note that for Rain samples there is no filtration nor digestion steps. | 0.01 | Y | Y | Y | 13356 | 4/11/1996 6:01:00 PM | 11/6/2018 | 4/12/2024 1:44:29 PM | 1996-04-12 | 6/3/1998 6:51:34 PM | EMS |  | Concentration | mg/L | Y | (fl. conc.) |  | Cadmium Dissolved (fl. conc.) | 7440-43-9 | Metals - Dissolved; | Metals - Dissolved; EMS code: CD-D | Chemical | Dissolved | LAB
CD-D | mg/L | Cadmium Dissolved | X028 | FA,HNO3,AA Analysis | FA;HNO3;AA Analysis | 0.0005 | Y | Y | Y | 18 |  |  | 10/23/1996 11:14:48 PM | 1996-04-12 | 6/3/1998 6:51:34 PM | EMS |  | Concentration | mg/L | Y | (fl. conc.) |  | Cadmium Dissolved (fl. conc.) | 7440-43-9 | Metals - Dissolved; | Metals - Dissolved; EMS code: CD-D | Chemical | Dissolved | LAB
CD-D | mg/L | Cadmium Dissolved | METC | MET-D-NP-CCMS-VA | Ultra trace metals in water are analyzed by CRC ICPMS, based on US EPA Method 6020B (July 2014). This procedure is intended for pristine field-filtered acid-preserved water samples. The detection limits (LOR) for this test are based on lab instrumental analysis only, not including filtration. ALS-supplied field filtration equipment does not support these LOR. Therefore, because of the high probability of false positives due to filtration, it is strongly recommended that a filtration blank be analysed to aid in data interpretation. | 5e-06 | Y | Y | Y | 1776 | 4/3/2012 | 8/15/2024 | 8/23/2024 9:50:35 AM | 2019-08-20 |  | KATFRASE |  | Concentration | mg/L | Y | (fl. conc.) |  | Cadmium Dissolved (fl. conc.) | 7440-43-9 | Metals - Dissolved; | Metals - Dissolved; EMS code: CD-D | Chemical | Dissolved | LAB
CD-D | mg/L | Cadmium Dissolved | F032 | Filt.,Acid,ICPMS | There is no acid digestion required for this analysis. A water sample is collected and preserved in 0.2 percent nitric acid. To analyse the extractable metals the sample is shaken, allowed to sit for at least twenty four hours, and a decanted portion is introduced to the spectrometer. When the dissolved metals are required, the samples must first be filtered through a 0.45 micron cellulose acetate filter paper and then followed by preservation in 0.2 percent nitric acid. It is important to note that different laboratories may have very different detection limits for some variables analyzed by ICPMS. | 5e-06 | Y | Y | Y | 81982 | 8/16/1996 2:35:00 PM | 8/26/2024 | 8/26/2024 4:53:05 PM | 1997-01-16 | 8/15/2019 10:54:12 AM | MPAK | KATFRASE | Concentration | mg/L | Y | (fl. conc.) |  | Cadmium Dissolved (fl. conc.) | 7440-43-9 | Metals - Dissolved; | Metals - Dissolved; EMS code: CD-D | Chemical | Dissolved | LAB
CD-D | mg/L | Cadmium Dissolved | F039 | Field filt + acid, ICP-AES | Dissolved and extractable water samples are ready for analysis once collected. Total samples are autoclaved with HNO3 and hydrochloric acid (HCI) in capped Teflon vials. Analysis is performed via Inductively Coupled Argon Plasma Atomic Emission Spectrometry (ICP-AES)(simultaneous multielement analysis), Graphite Furnace Atomic Absorption Spectrometry (GF-AAS)(low level analysis) and/or Flame Atomic Absorption Spectrometry (F-AAS). | 0.005 | Y | Y | Y | 5937 | 7/29/1996 6:03:00 PM | 7/5/2019 1:21:00 PM | 9/11/2019 8:47:50 AM | 1997-05-08 | 6/3/1998 6:51:46 PM | MPAK |  | Concentration | mg/L | Y | (fl. conc.) |  | Cadmium Dissolved (fl. conc.) | 7440-43-9 | Metals - Dissolved; | Metals - Dissolved; EMS code: CD-D | Chemical | Dissolved | LAB
CD-D | mg/L | Cadmium Dissolved | X562 | Dissolved metals: FF, FA, HNO3, HR-ICPMS | Ultra trace metals in water are analyzed by high resolution inductively coupled plasma mass spectrometry (HR-ICPMS), based on US EPA Method 200.8, (Oct. 1999). This procedure is intended for pristine field-filtered acid-preserved water samples. The detection limits (LOR) for this test are based on lab instrumental analysis only, not including filtration. ALS-supplied field filtration equipment does not support these LOR. Therefore, because of the high probability of false positives due to filtration, it is strongly recommended that a filtration blank be analyzed to aid in data interpretation. | 5e-06 | Y | Y | Y | 1264 | 6/8/2015 | 10/19/2023 | 11/9/2023 7:41:27 AM | 2015-10-22 |  | RDGIBSON |  | Concentration | mg/L | Y | (fl. conc.) |  | Cadmium Dissolved (fl. conc.) | 7440-43-9 | Metals - Dissolved; | Metals - Dissolved; EMS code: CD-D | Chemical | Dissolved | LAB
CD-D | mg/L | Cadmium Dissolved | X356 | LF,LA:HNO3:ICP | Samples are lab-filtered (0.45 um); lab-acidified with HNO3; analysis by ICP/AES. | 0.01 | Y | Y | Y | 426 |  |  | 3/23/2010 5:08:08 PM | 1996-04-12 | 6/3/1998 6:51:34 PM | EMS |  | Concentration | mg/L | Y | (fl. conc.) |  | Cadmium Dissolved (fl. conc.) | 7440-43-9 | Metals - Dissolved; | Metals - Dissolved; EMS code: CD-D | Chemical | Dissolved | LAB
CD-D | mg/L | Cadmium Dissolved | X556 | Metals in water, dissolved, "in-bottle" digestion, ICP-SFMS | The whole water sample is filtered through a 0.45 micron cellulose acetate filter paper and the filtrate is stored in a HDPE/Teflon bottle. The container is de-capped in a Class 100 workstation, acidified to 2% (v/v) HNO3 and 1% (v/v) HCl and recapped. The sample is "in-bottle" digested in a convection oven, set at 60 degrees C for 16 hours. The sample is cooled and introduced to an inductively coupled argon plasma-sector field mass spectrometer (ICP-SFMS) by pneumatic nebulization. Each respective element is measured at a specific mass to charge ratio, m/z value expressed in amu (atomic mass units), and at a specified intrument resolution. LR: Low Resolution ~300, MR: Mid Resolution ~4000 and HR: High Resolution ~10000. Ref: Method number 02-2003, National Laboratory for Environmental Testing, CCIW, Environment Canada, Burlington, Requested April 2003. | 1e-06 | Y | Y | Y | 4241 | 6/11/2003 | 12/6/2023 2:04:00 AM | 2/14/2024 9:05:32 AM | 2004-01-19 |  | UMULLER |  | Concentration | mg/L | Y | (fl. conc.) |  | Cadmium Dissolved (fl. conc.) | 7440-43-9 | Metals - Dissolved; | Metals - Dissolved; EMS code: CD-D | Chemical | Dissolved | LAB
CD-D | mg/L | Cadmium Dissolved | X350 | FF,FA:HNO3:ICP | Samples are field-filtered (0.45 um); field-acidified with HNO3; analysis by ICP/AES. | 0.01 | Y | Y | Y | 5139 |  |  | 8/26/2024 2:41:32 PM | 1996-04-12 | 6/3/1998 6:51:41 PM | EMS |  | Concentration | mg/L | Y | (fl. conc.) |  | Cadmium Dissolved (fl. conc.) | 7440-43-9 | Metals - Dissolved; | Metals - Dissolved; EMS code: CD-D | Chemical | Dissolved | LAB
CD-D | mg/L | Cadmium Dissolved | X357 | LF,LA:HNO3:GFAA | Samples are lab-filtered (0.45 um) and lab-acidified with HNO3; analysis by GFAA. | 0.0001 | Y | Y | Y | 27 | 1/20/1998 | 9/23/1998 | 10/30/1998 9:32:50 AM | 1996-04-12 | 6/3/1998 6:51:34 PM | EMS |  | Concentration | mg/L | Y | (fl. conc.) |  | Cadmium Dissolved (fl. conc.) | 7440-43-9 | Metals - Dissolved; | Metals - Dissolved; EMS code: CD-D | Chemical | Dissolved | LAB
CD-D | mg/L | Cadmium Dissolved | MSEA | Metals in Seawater (undigested) by (CRC) ICP-MS | This method describes the multi-elemental determination of trace elements by Collision/Reaction Cell (CRC) ICP-MS. The method measures ions produced by a radio-frequency inductively coupled plasma. Analyte species originating in a liquid are pneumatically nebulized and the resulting aerosol transported by argon gas into the plasma torch. The ions produced are entrained in the plasma gas and extracted, by means of a differentially pumped vacuum interface, into a mass spectrometer. The ions flow through a Collision/Reaction Cell (CRC) where depending on the mode of analysis either the ions passes through the cell unhindered (no gas - classical ICP-MS mode), or collisions/reactions chemistry takes place between the ions and a neutral gas such as hydrogen or helium. Any polyatomic ions which give interference on the analytes are removed. The ions are then sorted according to their mass-to-charge ratios by a quadrupole mass spectrometer. The ions are then sorted according to their mass-to-charge ratios by a quadrupole mass spectrometer. The ions transmitted through the quadrupole are quantified by a channel electron multiplier and the ion information processed by a data handling system. Although most polyatomic interference is removed, any additional interference must be assessed and valid corrections applied or the data flagged to indicate problems. Interference correction must include compensation for background ions contributed by the plasma gas, reagents, and constituents of the sample matrix. Instrument drift as well as suppressions or enhancements of instrument response caused by the sample matrix must be corrected for by the use of internal standards. | 2e-05 | Y | Y | Y | 4 | 4/12/2023 | 4/29/2024 | 5/16/2024 2:00:14 PM | 2009-09-08 |  | RDGIBSON |  | Concentration | mg/L | Y | (fl. conc.) |  | Cadmium Dissolved (fl. conc.) | 7440-43-9 | Metals - Dissolved; | Metals - Dissolved; EMS code: CD-D | Chemical | Dissolved | LAB
CD-D | mg/L | Cadmium Dissolved | X116 | FF,FA,HNO3,Graph.Furn,AA | Samples are field-filtered (0.45 um), field-acidified with HNO3; analysis by GFAA. | 0.0002 | Y | Y | Y | 2122 |  |  | 10/1/2003 9:41:03 AM | 1996-04-12 | 6/3/1998 6:51:34 PM | EMS |  | Concentration | mg/L | Y | (fl. conc.) |  | Cadmium Dissolved (fl. conc.) | 7440-43-9 | Metals - Dissolved; | Metals - Dissolved; EMS code: CD-D | Chemical | Dissolved | LAB
CD-D | mg/L | Cadmium Dissolved | EQ01 | EQUIS (mg/L) | EQUIS (mg/L) | 0.0005 | N | Y | Y | 4114 |  |  | 10/24/1996 3:16:55 PM | 1996-04-12 | 6/3/1998 6:51:41 PM | EMS |  | Concentration | mg/L | Y | (fl. conc.) |  | Cadmium Dissolved (fl. conc.) | 7440-43-9 | Metals - Dissolved; | Metals - Dissolved; EMS code: CD-D | Chemical | Dissolved | LAB
CD-D | mg/L | Cadmium Dissolved | X085 | LF,LA,HNO3,AA | Samples are lab-filtered (0.45 um), and lab-acidified (HNO3); analysis by AA. | 0.0005 | Y | Y | Y | 739 | 12/12/1996 | 9/18/2015 | 10/28/2015 9:54:16 AM | 1996-04-12 | 6/3/1998 6:51:34 PM | EMS |  | Concentration | mg/L | Y | (fl. conc.) |  | Cadmium Dissolved (fl. conc.) | 7440-43-9 | Metals - Dissolved; | Metals - Dissolved; EMS code: CD-D | Chemical | Dissolved | LAB
CD-D | mg/L | Cadmium Dissolved | MWID | Metals in Water by ICPMS (digested) | Metals in Water by ICPMS (digested) Based on EPA 200.8 | 0.0001 | Y | Y | Y | 31896 | 10/26/2018 5:20:00 PM | 10/26/2018 5:20:00 PM | 7/31/2024 1:46:41 PM | 2009-12-24 |  | RDGIBSON |  | Concentration | mg/L | Y | (fl. conc.) |  | Cadmium Dissolved (fl. conc.) | 7440-43-9 | Metals - Dissolved; | Metals - Dissolved; EMS code: CD-D | Chemical | Dissolved | LAB
CD-D | t/d | Cadmium Dissolved | CALC | Calculation (t/d) | Calculation (t/d) | 0 | Y | Y | Y | 3755 |  |  | 7/3/2024 12:59:26 PM | 1996-04-12 | 6/3/1998 6:51:34 PM | EMS |  | MassFlow | t/d | Y | (mass flow) |  | Cadmium Dissolved (mass flow) | 7440-43-9 | Metals - Dissolved; | Metals - Dissolved; EMS code: CD-D | Chemical | Dissolved | LAB
CD-D | mg/L | Cadmium Dissolved | 0050 | Graphite Furnace AA | Graphite Furnace AA | 0.0005 | Y | Y | Y | 2201 | 9/29/2000 | 12/1/2010 | 12/9/2010 3:55:26 PM | 1996-04-12 | 6/3/1998 6:51:34 PM | EMS |  | Concentration | mg/L | Y | (fl. conc.) |  | Cadmium Dissolved (fl. conc.) | 7440-43-9 | Metals - Dissolved; | Metals - Dissolved; EMS code: CD-D | Chemical | Dissolved | LAB
CD-E | umol/g | Cadmium Extractable | FSEM | AVS extraction, SEM Analysis ICPMS | Determination of simultaneous extractable metals (SEM) in sediment after sample extracted with hydrochloric acid for acid volatile sulfide (AVS) | 0.01 | Y | Y | Y | 19 | 3/12/2020 | 3/24/2020 | 2/22/2023 12:51:03 PM | 2010-11-16 |  | RDGIBSON |  | ParticlesPerMass | µmol/g | Y | (part. /m) |  | Cadmium Extractable (part. /m) | 7440-43-9 | Metals - Extractable; | Metals - Extractable; EMS code: CD-E | Chemical | Extractable | LAB
CD-E | mg/L | Cadmium Extractable | F040 | Field Acid, lab filt, ICP-AES | Dissolved and extractable water samples are ready for analysis once collected. Total samples are autoclaved with HNO3 and hydrochloric acid (HCI) in capped Teflon vials. Analysis is performed via Inductively Coupled Argon Plasma Atomic Emission Spectrometry (ICP-AES)(simultaneous multielement analysis), Graphite Furnace Atomic Absorption Spectrometry (GF-AAS)(low level analysis) and/or Flame Atomic Absorption Spectrometry (F-AAS). | 0.005 | Y | Y | Y | 2361 | 8/5/1997 6:06:00 PM | 3/14/2008 | 9/11/2019 8:47:41 AM | 1997-05-08 | 6/3/1998 6:51:46 PM | MPAK |  | Concentration | mg/L | Y | (fl. conc.) |  | Cadmium Extractable (fl. conc.) | 7440-43-9 | Metals - Extractable; | Metals - Extractable; EMS code: CD-E | Chemical | Extractable | LAB
CD-E | mg/L | Cadmium Extractable | X512 | HNO3,pHto7.5,MIK extr'n,AA | HNO3 Digestion;pH to 7.5; MIK solvent extraction;AA | 0.001 | Y | Y | Y | 867 |  |  | 8/23/1999 3:43:46 PM | 1999-04-22 |  | NPEPPIN |  | Concentration | mg/L | Y | (fl. conc.) |  | Cadmium Extractable (fl. conc.) | 7440-43-9 | Metals - Extractable; | Metals - Extractable; EMS code: CD-E | Chemical | Extractable | LAB
CD-E | ug/g | Cadmium Extractable | X230 | Hom,HNO3/HCl,ICP | The sample is dried, ground to pass 100 mesh, and digested with a nitric-hydrochloric acid mixture to solubilize the solid matter and to remove the organic material by oxidation and volatilization. The completed digests are submitted to metals instrumentation personnel for metals analysis by Inductively Coupled Atomic Emission Spectrophotometry (ICP-AES), or hydride analysis by atomic absorption. | 0.5 | Y | Y | Y | 74 | 12/13/2011 12:02:00 PM | 12/13/2011 12:02:00 PM | 2/9/2012 12:15:56 PM | 1996-04-12 | 6/3/1998 6:51:34 PM | EMS |  | MassPerMass | mg/kg | Y | (mass conc.) |  | Cadmium Extractable (mass conc.) | 7440-43-9 | Metals - Extractable; | Metals - Extractable; EMS code: CD-E | Chemical | Extractable | LAB
CD-E | mg/L | Cadmium Extractable | F047 | 24hr HNO3 dil'n,ICP-MS | There is no acid digestion required for this analysis. A water sample is collected and preserved in 0.2 percent nitric acid. To analyse the extractable metals the sample is shaken, allowed to sit for at least twenty four hours, and a decanted portion is introduced to the spectrometer. When the dissolved metals are required, the samples must first be filtered through a 0.45 micron cellulose acetate filter paper and then followed by preservation in 0.2 percent nitric acid. It is important to note that different laboratories may have very different detection limits for some variables analyzed by ICP/MS. | 5e-06 | Y | Y | Y | 2676 | 4/16/1998 10:21:00 AM | 6/26/2003 | 9/11/2019 8:46:17 AM | 1997-10-20 | 6/3/1998 6:51:47 PM | NPEPPIN |  | Concentration | mg/L | Y | (fl. conc.) |  | Cadmium Extractable (fl. conc.) | 7440-43-9 | Metals - Extractable; | Metals - Extractable; EMS code: CD-E | Chemical | Extractable | LAB
CD-E | mg/L | Cadmium Extractable | X553 | Metals in water, extractable, ICP-SFMS | The whole water sample is collected in a HDPE/Teflon bottle. In a Class 100 workstation it is preserved with concentrated Nitric Acid at concentration of 0.2 percent (v/v). The acid preservative extracts the metals from material in the sample and a decanted portion is introduced directly to an inductively coupled argon plasma-sector field mass spectrometer (ICP-SFMS) by pneumatic nebulization. Each respective element is measured at a specific mass to charge ratio, m/z value expressed in amu (atomic mass units), and at a specified instrument resolution. LR: Low Resolution ~300, MR: Mid Resolution ~4000 and HR: High Resolution ~10000. Ref: Method number 02-2061, National Laboratory for Environmental Testing, CCIW, Environment Canada, Burlington. Requested April 2003. | 1e-06 | Y | Y | Y | 2597 | 5/22/2003 | 9/21/2007 | 11/28/2023 2:11:53 PM | 2004-01-19 |  | UMULLER |  | Concentration | mg/L | Y | (fl. conc.) |  | Cadmium Extractable (fl. conc.) | 7440-43-9 | Metals - Extractable; | Metals - Extractable; EMS code: CD-E | Chemical | Extractable | LAB
CD-E | mg/L | Cadmium Extractable | F069 | SEM extr,6N HCl,ICP/OES | 6.0 N Hydrochloric acid (HCI) is added to an aliquot of the sediment sample. The extract produced from the addition of the acid is then analyzed for simultaneously extracted metals (SEM) using inductively coupled plasma - optical emission spectrophotometry. The detection limits depend on the metal parameters in question and the instrumentation used. | 0.01 | Y | Y | Y | 10 | 3/26/2012 3:01:00 AM | 3/26/2012 3:01:00 AM | 11/25/2021 1:50:15 PM | 1999-07-09 |  | NPEPPIN |  | Concentration | mg/L | Y | (fl. conc.) |  | Cadmium Extractable (fl. conc.) | 7440-43-9 | Metals - Extractable; | Metals - Extractable; EMS code: CD-E | Chemical | Extractable | LAB
CD-I | mg/dm2/d | Cadmium-Insoluble | EQ15 | EQUIS (mg/dm2/d) | EQUIS (mg/dm2/d) | 0.0001 | N | Y | Y | 169 |  |  | 10/22/1996 10:18:48 PM | 1996-04-12 | 6/3/1998 6:51:34 PM | EMS |  | MassFlux | mg/dm²/d | Y | (mass flux) |  | Cadmium-Insoluble (mass flux) | 7440-43-9 | Metals - Insoluble; | Metals - Insoluble; EMS code: CD-I | Chemical | Insoluble | LAB
CD-I | mg/dm2/d | Cadmium-Insoluble | 5039 | N/P Dig, ICP | Aliquot of prepared sample is digested first with nitric acid then with perchloric acid. Analysis by ICP. (a) For air particulates, preliminary results are reported in mg/L and loading results as mg/dm2/d. (b) For Hi Vol samples preliminary results are same as reported in mg/L and loading results as ug/m3. | 0.0001 | Y | Y | Y | 176 | 12/5/1997 | 5/15/2000 | 11/18/2000 1:30:12 AM | 1996-04-12 | 6/3/1998 6:51:34 PM | EMS |  | MassFlux | mg/dm²/d | Y | (mass flux) |  | Cadmium-Insoluble (mass flux) | 7440-43-9 | Metals - Insoluble; | Metals - Insoluble; EMS code: CD-I | Chemical | Insoluble | LAB
CD-S | mg/dm2/d | Cadmium-Soluble | EQ15 | EQUIS (mg/dm2/d) | EQUIS (mg/dm2/d) | 0.0001 | N | Y | Y | 174 |  |  | 10/22/1996 10:18:48 PM | 1996-04-12 | 6/3/1998 6:51:34 PM | EMS |  | MassFlux | mg/dm²/d | Y | (mass flux) |  | Cadmium-Soluble (mass flux) | 7440-43-9 | Metals - Soluble; | Metals - Soluble; EMS code: CD-S | Chemical | Soluble | LAB
CD-S | umol/g | Cadmium-Soluble | FSEM | AVS extraction, SEM Analysis ICPMS | Determination of simultaneous extractable metals (SEM) in sediment after sample extracted with hydrochloric acid for acid volatile sulfide (AVS) | 0.01 | Y | Y | Y | 0 |  |  |  | 2024-02-02 |  | EMS |  | ParticlesPerMass | µmol/g | Y | (part. /m) |  | Cadmium-Soluble (part. /m) | 7440-43-9 | Metals - Soluble; | Metals - Soluble; EMS code: CD-S | Chemical | Soluble | LAB
CD-S | mg/dm2/d | Cadmium-Soluble | X170 | ICP | This code is a catch-all for a large variety of ICP methods including water and also solids digested into solutions as an intermediate result. Different procedures will have greatly different detection limits. (a) For air particulates, preliminary results are reported in mg/L and loading results as mg/dm2/d. (b) For Hi Vol samples preliminiary results are same as reported in mg/L and loading results as ug/m3. (c) Note that for Rain samples there is no filtration nor digestion steps. | 0.0001 | Y | Y | Y | 140 |  |  | 2/13/1998 2:28:22 AM | 1996-04-12 | 6/3/1998 6:51:41 PM | EMS |  | MassFlux | mg/dm²/d | Y | (mass flux) |  | Cadmium-Soluble (mass flux) | 7440-43-9 | Metals - Soluble; | Metals - Soluble; EMS code: CD-S | Chemical | Soluble | LAB
CD-S | mg/dm2/d | Cadmium-Soluble | 0050 | Graphite Furnace AA | Graphite Furnace AA | 1e-05 | Y | Y | Y | 7 |  |  | 10/23/1996 3:44:01 PM | 1996-04-12 | 6/3/1998 6:51:34 PM | EMS |  | MassFlux | mg/dm²/d | Y | (mass flux) |  | Cadmium-Soluble (mass flux) | 7440-43-9 | Metals - Soluble; | Metals - Soluble; EMS code: CD-S | Chemical | Soluble | LAB
CD-S | mg/dm2/d | Cadmium-Soluble | 5045 | ICP Analysis | ICP a) For Air particulates analyzed without an acid digestion step, preliminary results are reported in mg/L and loading results as mg/dm2/d. b) For HiVol samples, preliminary results are reported as mg/L and loading results as ug/m3. | 0.0001 | Y | Y | Y | 39 | 11/28/1997 | 5/15/2000 | 11/18/2000 1:30:12 AM | 1998-10-28 |  | NPEPPIN |  | MassFlux | mg/dm²/d | Y | (mass flux) |  | Cadmium-Soluble (mass flux) | 7440-43-9 | Metals - Soluble; | Metals - Soluble; EMS code: CD-S | Chemical | Soluble | LAB
CD-T | ug/g | Cadmium Total | PB02 | PBM,HNO3/HCl dig,dry60C,100mesh, | Performance Based Method. Samples are dried at 60 C and sieved using a .149 mm (100 mesh) sieve. Solid samples are digested with the combination of nitric and hydrochloric acids with a 1:1 ratio of acids. For environmental purposes, strong acid leaches are used routinely since they have the advantage of speed, low cost, and superio reproducibility. All metals trapped within the refractory matrix are unlikely to have environmental consequences. Instrumental analysis of extracts produced using this digestion method can be performed by a variety of techniques. |  | Y | Y | Y | 52 | 5/24/2001 3:42:00 PM | 11/26/2001 10:23:00 AM | 5/26/2003 2:27:35 PM | 2001-07-18 |  | NPEPPIN |  | MassPerMass | mg/kg | Y | (mass conc.) |  | Cadmium Total (mass conc.) | 7440-43-9 | Metals - Total; | Metals - Total; EMS code: CD-T | Chemical | Total | LAB
CD-T | mg/L | Cadmium Total | X528 | unknown method,plasma emission spec. | unknown method;plasma emission spec |  | Y | Y | Y | 7 |  |  | 8/17/1999 2:50:49 AM | 1999-07-21 |  | NPEPPIN |  | Concentration | mg/L | Y | (fl. conc.) |  | Cadmium Total (fl. conc.) | 7440-43-9 | Metals - Total; | Metals - Total; EMS code: CD-T | Chemical | Total | LAB
CD-T | mg/dm2/d | Cadmium Total | CS15 | System Calculation | System Calculation (mg/dm2/d) |  | N | Y | N | 44 |  |  | 11/18/2000 1:30:24 AM | 1996-04-12 | 6/3/1998 6:51:43 PM | EMS |  | MassFlux | mg/dm²/d | Y | (mass flux) |  | Cadmium Total (mass flux) | 7440-43-9 | Metals - Total; | Metals - Total; EMS code: CD-T | Chemical | Total | LAB
CD-T | mg/L | Cadmium Total | X561 | Total metals: FA, HNO3, HR-ICPMS (No Prep)Prep) | Ultra trace metals in water are analyzed by high resolution inductively coupled plasma mass spectrometry (HR-ICPMS), modified from US EPA 200.8, (Revision 5.5). The detection limits provided can only be met for undigested samples. This procedure is intended for colorless, non-turbid, acid-preserved water samples (i.e. pristine water samples), having turbidity < 1 NTU and no odor. Where turbidity exceeds 1 NTU, and/or the sample is colored and has an odor, results may be biased low compared to true Total Metals concentrations. ALS recommends that turbidity analysis be requested on samples submitted for this test to aid with interpretation of results. | 5e-06 | Y | Y | Y | 2756 | 6/8/2015 | 10/5/2023 | 1/31/2024 2:37:41 PM | 2015-10-22 |  | RDGIBSON |  | Concentration | mg/L | Y | (fl. conc.) |  | Cadmium Total (fl. conc.) | 7440-43-9 | Metals - Total; | Metals - Total; EMS code: CD-T | Chemical | Total | LAB
CD-T | mg/L | Cadmium Total | 0020 | HNO3 Dig:AA Analysis | HNO3 Dig:AA Analysis | 0.01 | Y | Y | Y | 93 | 6/25/2010 3:15:00 PM | 12/1/2010 | 12/9/2010 3:55:26 PM | 1996-04-12 | 6/3/1998 6:51:34 PM | EMS |  | Concentration | mg/L | Y | (fl. conc.) |  | Cadmium Total (fl. conc.) | 7440-43-9 | Metals - Total; | Metals - Total; EMS code: CD-T | Chemical | Total | LAB
CD-T | mg/L | Cadmium Total | METN | MET-T-NP-CCMS-VA | Ultra trace metals in water are analyzed by CRC ICPMS, based on US EPA Method 6020B (July 2014). The detection limits provided can only be met for undigested samples. This procedure is intended for colorless, non-turbid, acid-preserved water samples (i.e. pristine water samples), having turbidity < 1 NTU and no odor. Where turbidity exceeds 1 NTU, and/or the sample is colored and has an odor, results may be biased low compared to true Total Metals concentrations. ALS recommends that turbidity analysis be requested on samples submitted for this test to aid with interpretation of results. Where turbidity is <1NTU, undigested metals are equivalent to total metals concentrations. | 5e-06 | Y | Y | Y | 3648 | 4/3/2012 | 8/19/2024 | 8/26/2024 1:03:06 PM | 2019-08-20 |  | KATFRASE |  | Concentration | mg/L | Y | (fl. conc.) |  | Cadmium Total (fl. conc.) | 7440-43-9 | Metals - Total; | Metals - Total; EMS code: CD-T | Chemical | Total | LAB
CD-T | ug/g | Cadmium Total | 0031 | ICP | ICP | 2 | Y | Y | Y | 5 | 5/6/1997 4:38:00 PM | 5/6/1997 4:38:00 PM | 9/2/1997 6:04:07 PM | 1996-04-12 | 6/3/1998 6:51:41 PM | EMS |  | MassPerMass | mg/kg | Y | (mass conc.) |  | Cadmium Total (mass conc.) | 7440-43-9 | Metals - Total; | Metals - Total; EMS code: CD-T | Chemical | Total | LAB
CD-T | ug/m3 | Cadmium Total | 5097 | N/P Dig(Filt+Wash),ICP | N/P Dig(Filt+Wash);ICP | 0.001 | Y | Y | Y | 10475 |  |  | 7/3/2024 12:59:27 PM | 1996-04-12 | 6/3/1998 6:51:34 PM | EMS |  | AirConcentration | µg/m³ | Y | (air conc.) |  | Cadmium Total (air conc.) | 7440-43-9 | Metals - Total; | Metals - Total; EMS code: CD-T | Chemical | Total | LAB
CD-T | ug/g | Cadmium Total | FS02 | D&G,Nitric/HCl:GF | Samples are dried at 60 C, then either sieved to -100 mesh (150 um) and rolled to homogenize, or finely ground with a mortar and pestle or a shatter box laboratory mill and rolled (depending on the submitter's requirements). The samples are then weighed into Teflon digestion liners and digested with HNO3 and HCl in a microwave oven. Samples are cooled, volumized, and allowed to settle overnight before being decanted. Analysis is then performed via Graphite Furnace Atomic Absorption Spectrometry (GF-AAS) (low level analysis). GFAA Spectrometry involves heating the sample in three or more stages. First, a low current heats the graphite tube to dry the sample. The second, or charring stage, destroys the organic matter and volatilizes other components at an intermediate temperature. Finally, a high current heats the tube to incandescence and, in an inert atmosphere, atomizes the element being determined. Additional stages may be added to drying and charring, and to cool and clean the tubes in between samples. The resultant ground-state atomic vapour absorbs monochromatic radiation from the source. A photoelectric detector measures the intensity of the transmitted radiation characteristic of the element being determined, which is inversely proportional to the concentration of ground-state atoms in the optical path over a limited range. Low level Cadmium and Lead are determined using Deuterium lamp background correction system whereas Arsenic and Selenium are determined by using Zeeman background correction system. | 0.04 | Y | Y | Y | 23 | 2/20/1998 6:02:00 PM | 1/3/2001 12:00:00 PM | 5/30/2001 1:29:55 AM | 1996-06-19 | 6/3/1998 6:51:44 PM | MPAK |  | MassPerMass | mg/kg | Y | (mass conc.) |  | Cadmium Total (mass conc.) | 7440-43-9 | Metals - Total; | Metals - Total; EMS code: CD-T | Chemical | Total | LAB
CD-T | ug/g | Cadmium Total | FB02 | Freeze Dry,Hom,Acid Dig,GF | Sample moisture is determined by weighing samples before and after freeze-drying. Freeze-dried samples are ground in a Braun blender before being weighed into Teflon PFA liners and digested with concentrated HNO3, concentrated HCl, and 50% H2O2 in a microwave oven; which results in a complete sample dissolution. Samples are cooled and volumized with Type I water. Analysis is performed via Inductively Coupled Argon Plasma Atomic Emission Spectrometry (ICP-AES) (simultaneous multielement analysis), Graphite Furnace Atomic Absorption Spectrometry (GF-AAS) (low level Cd and Pb analysis), Cold Vapor Atomic Fluorescence Spectrometry (CV-AFS) (Mercury analysis) and/or Flame Atomic Absorption Spectrometry (F-AAS) (K analysis). | 0.02 | Y | Y | Y | 45 | 8/28/2001 4:46:00 PM | 8/28/2001 4:51:00 PM | 11/20/2002 1:41:04 PM | 1996-06-19 | 6/3/1998 6:51:44 PM | MPAK |  | MassPerMass | mg/kg | Y | (mass conc.) |  | Cadmium Total (mass conc.) | 7440-43-9 | Metals - Total; | Metals - Total; EMS code: CD-T | Chemical | Total | LAB
CD-T | mg/L | Cadmium Total | F121 | Acid,ICPMS | Acidified, no digestion, ICPMS | 1e-05 | Y | Y | Y | 8652 | 3/12/2010 7:57:00 PM | 8/15/2024 9:10:00 PM | 8/26/2024 8:06:03 AM | 2002-07-17 | 7/16/2012 3:27:22 PM | UMULLER | RDGIBSON | Concentration | mg/L | Y | (fl. conc.) |  | Cadmium Total (fl. conc.) | 7440-43-9 | Metals - Total; | Metals - Total; EMS code: CD-T | Chemical | Total | LAB
CD-T | ug/g wet | Cadmium Total | X215 | Hom,Nitr/Perchl,Grph Fur | Hom;Nitr/Perchl;Grph Fur | 0.005 | Y | Y | Y | 73 |  |  | 10/24/1996 10:01:25 AM | 1996-04-12 | 6/3/1998 6:51:41 PM | EMS |  | MassPerMass | µg/g | Y | (mass conc.) | (wet) | Cadmium Total (wet) (mass conc.) | 7440-43-9 | Metals - Total; | Metals - Total; EMS code: CD-T | Chemical | Total | LAB
CD-T | ppm (S) | Cadmium Total | MS61 | ALS Minerals ME-MS61 | ALS Minerals ME-MS61. A prepared sample is digested with HNO3; HF; and HClO4 in a graphite heating block to near dryness. After cooling; the sample is boiled in diluted HCl. Bulk to volume with deionized water; mixed and analyzed by ICP-AES and ICP-MS for all elements. | 0.02 | Y | Y | Y | 0 |  |  |  | 2024-01-29 |  | EMS |  | DimensionlessRatio | ppm | N | (ratio) |  | Cadmium Total (ratio) | 7440-43-9 | Metals - Total; | Metals - Total; EMS code: CD-T | Chemical | Total | LAB
CD-T | mg/L | Cadmium Total | 5078 | N/P Digestion/ ICP Anal. | N/P Digestion/ ICP Anal. | 0.01 | N | Y | Y | 32 |  |  | 10/24/1996 2:10:29 PM | 1996-10-15 | 6/3/1998 6:51:46 PM | EMS |  | Concentration | mg/L | Y | (fl. conc.) |  | Cadmium Total (fl. conc.) | 7440-43-9 | Metals - Total; | Metals - Total; EMS code: CD-T | Chemical | Total | LAB
CD-T | mg/L | Cadmium Total | ULTR | Total Metals in Water by ICPMS (Ultra) | Total Metals in Water by ICPMS (Ultra), EPA SW-846 3005A/6020A, MET-TOT-ULTRA-MS-VA This analysis is carried out using procedures adapted from "Standard Methods for the Examination of Water and Wastewater" published by the American Public Health Association, and with procedures adapted from "Test Methods for Evaluating Solid Waste" SW-846 published by the United States Environmental Protection Agency (EPA). The procedures may involve preliminary sample treatment by acid digestion, using either hotblock or microwave oven, or filtration (EPA Method 3005A). Instrumental analysis is by inductively coupled plasma - mass spectrometry (EPA Method 6020A). | 1.7e-05 | Y | Y | Y | 33 | 5/25/2007 | 7/13/2011 | 7/10/2012 2:14:47 PM | 2012-07-05 |  | RDGIBSON |  | Concentration | mg/L | Y | (fl. conc.) |  | Cadmium Total (fl. conc.) | 7440-43-9 | Metals - Total; | Metals - Total; EMS code: CD-T | Chemical | Total | LAB
CD-T | ug/m3 | Cadmium Total | 5012 | N/P Dig:HGA Analysis | N/P Dig:HGA Analysis | 0.01 | N | Y | Y | 723 |  |  | 10/24/1996 6:54:32 AM | 1996-04-12 | 6/3/1998 6:51:34 PM | EMS |  | AirConcentration | µg/m³ | Y | (air conc.) |  | Cadmium Total (air conc.) | 7440-43-9 | Metals - Total; | Metals - Total; EMS code: CD-T | Chemical | Total | LAB
CD-T | ug/g | Cadmium Total | FB04 | Freeze Dry,Hom,Acid Dig,ICP | Sample moisture is determined by weighing samples before and after freeze-drying. Freeze-dried samples are ground in a Braun blender before being weighed into Teflon PFA liners and digested with concentrated HNO3, concentrated HCl, and 50% H2O2 in a microwave oven; which results in a complete sample dissolution. Samples are collected and volumized with Type I water. Analysis is performed via Inductively Coupled Argon Plasma Atomic Emission Spectrometry (ICP-AES) (simultaneous multielement analysis), Graphite Furnace Atomic Absorption Spectrometry (GF-AAS) (low level Cd and Pb analysis), Cold Vapor Atomic Fluorescence Spectrometry (CV-AFS) (Mercury analysis) and/or Flame Atomic Absorption Spectrometry (F-AAS) (K analysis). | 0.4 | Y | Y | Y | 241 | 8/22/1996 6:11:00 PM | 2/1/2002 2:30:00 PM | 6/23/2003 1:15:29 PM | 1997-02-27 | 6/3/1998 6:51:46 PM | MPAK |  | MassPerMass | mg/kg | Y | (mass conc.) |  | Cadmium Total (mass conc.) | 7440-43-9 | Metals - Total; | Metals - Total; EMS code: CD-T | Chemical | Total | LAB
CD-T | ug/g | Cadmium Total | X230 | Hom,HNO3/HCl,ICP | The sample is dried, ground to pass 100 mesh, and digested with a nitric-hydrochloric acid mixture to solubilize the solid matter and to remove the organic material by oxidation and volatilization. The completed digests are submitted to metals instrumentation personnel for metals analysis by Inductively Coupled Atomic Emission Spectrophotometry (ICP-AES), or hydride analysis by atomic absorption. | 1 | Y | Y | Y | 5 | 12/30/2005 | 12/30/2005 | 2/3/2006 1:13:12 AM | 1998-11-05 |  | NPEPPIN |  | MassPerMass | mg/kg | Y | (mass conc.) |  | Cadmium Total (mass conc.) | 7440-43-9 | Metals - Total; | Metals - Total; EMS code: CD-T | Chemical | Total | LAB
CD-T | ug/g | Cadmium Total | X551 | Metals in Biota (dry), total, acid digestion, ICP-SFMS | 0.1 gram of dried, homogenized biota is weighed in a microwave TFM vessel, and 8mL of Nitric Acid and 1mL of Hydrogen Peroxide are added. The sample is digested at 200 degrees C for 15 minutes in a closed vessel high pressure microwave oven. The sample is cooled, the volume adjusted to 100mL and introduced to an iductively coupled argon plasma-sector field mass spectrometer (ICP-SFMS) by pneumatic nebulization. Each respective element is measured at a specific mass to charge ratio, m/z value expressed in amu (atomic mass units), and at a specified instrument resolution. LR: Low Resolution ~300, MR: Mid Resolution ~4000 and HR: Hich Resolution ~10000. Ref: Method number 02-2705, National Laboratory for Environmental Testing, CCIW, Environment Canada, Burlington. Requested April 2003. | 0.001 | Y | Y | Y | 49 | 8/11/2015 | 10/30/2023 | 11/28/2023 2:21:33 PM | 2004-01-19 |  | UMULLER |  | MassPerMass | mg/kg | Y | (mass conc.) |  | Cadmium Total (mass conc.) | 7440-43-9 | Metals - Total; | Metals - Total; EMS code: CD-T | Chemical | Total | LAB
CD-T | mg/L | Cadmium Total | MWID | Metals in Water by ICPMS (digested) | Metals in Water by ICPMS (digested) Based on EPA 200.8 | 0.0001 | Y | Y | Y | 36816 |  |  | 7/31/2024 1:46:41 PM | 2009-12-24 |  | RDGIBSON |  | Concentration | mg/L | Y | (fl. conc.) |  | Cadmium Total (fl. conc.) | 7440-43-9 | Metals - Total; | Metals - Total; EMS code: CD-T | Chemical | Total | LAB
CD-T | mg/L | Cadmium Total | F108 | SPR,IDA Chelation,ICPMS | SPR;IDA Chelation;ICPMS |  | Y | Y | Y | 4 | 11/30/2000 10:32:00 AM | 11/30/2000 10:32:00 AM | 9/10/2002 9:29:47 AM | 2001-03-08 |  | NPEPPIN |  | Concentration | mg/L | Y | (fl. conc.) |  | Cadmium Total (fl. conc.) | 7440-43-9 | Metals - Total; | Metals - Total; EMS code: CD-T | Chemical | Total | LAB
CD-T | mg/L | Cadmium Total | X167 | HNO3 Dig:Solvent Extr:AA | HNO3 Dig:Solvent Extr:AA | 0.0001 | Y | Y | Y | 220 |  |  | 10/24/1996 1:01:56 PM | 1996-04-12 | 6/3/1998 6:51:34 PM | EMS |  | Concentration | mg/L | Y | (fl. conc.) |  | Cadmium Total (fl. conc.) | 7440-43-9 | Metals - Total; | Metals - Total; EMS code: CD-T | Chemical | Total | LAB
CD-T | mg/L | Cadmium Total | 0100 | Dig:Ext.(APDC):AA Anal. | Dig:Ext.(APDC):AA Anal. | 0.0005 | Y | Y | Y | 920 |  |  | 6/16/2000 1:41:55 PM | 1996-04-12 | 6/3/1998 6:51:34 PM | EMS |  | Concentration | mg/L | Y | (fl. conc.) |  | Cadmium Total (fl. conc.) | 7440-43-9 | Metals - Total; | Metals - Total; EMS code: CD-T | Chemical | Total | LAB
CD-T | mg/L | Cadmium Total | F082 | Aliquot,HNO3/HCL dig,ICPMS | Aliquot microwave digested with 3:1 HNO3/HCL. Analysis by ICPMS. For HiVol samples preliminary result are reported in mg/L and loading results are reported in ug/m3 | 5e-06 | Y | Y | Y | 28466 | 7/26/2005 7:04:00 AM | 8/23/2024 | 8/26/2024 4:53:05 PM | 2000-06-13 | 8/15/2019 11:06:03 AM | NPEPPIN | KATFRASE | Concentration | mg/L | Y | (fl. conc.) |  | Cadmium Total (fl. conc.) | 7440-43-9 | Metals - Total; | Metals - Total; EMS code: CD-T | Chemical | Total | LAB
CD-T | mg/L | Cadmium Total | 0042 | GI HNO3 Dig: ICP | Sample aliquot taken from unfiltered unpreserved General Ions bottle; HNO3 digestion; ICP analysis. | 0.002 | Y | Y | Y | 5524 | 9/20/1996 | 7/5/2019 3:21:00 PM | 7/12/2019 6:49:13 PM | 1996-04-12 | 6/3/1998 6:51:34 PM | EMS |  | Concentration | mg/L | Y | (fl. conc.) |  | Cadmium Total (fl. conc.) | 7440-43-9 | Metals - Total; | Metals - Total; EMS code: CD-T | Chemical | Total | LAB
CD-T | ppm (S) | Cadmium Total | MS41 | ALS Minerals ME-MS41 | ALS Minerals ME-MS41. A prepared sample is digested with aqua regia in a graphite heating block. After cooling; the resulting solution is diluted to with deionized water; mixed and analyzed by ICP-AES. Samples are then analyzed by ICP-MS for the remaining suite of elements. | 0.01 | Y | Y | Y | 0 |  |  |  | 2024-01-29 |  | EMS |  | DimensionlessRatio | ppm | N | (ratio) |  | Cadmium Total (ratio) | 7440-43-9 | Metals - Total; | Metals - Total; EMS code: CD-T | Chemical | Total | LAB
CD-T | ug/g | Cadmium Total | FS82 | PBM,dry60C,10mesh,SALM-ICP-MS | PBM;dry60C;10mesh;SALM-ICP-MS | 0.05 | Y | Y | Y | 3486 | 5/21/2002 | 7/24/2024 | 7/30/2024 10:00:14 AM | 2002-11-12 |  | UMULLER |  | MassPerMass | mg/kg | Y | (mass conc.) |  | Cadmium Total (mass conc.) | 7440-43-9 | Metals - Total; | Metals - Total; EMS code: CD-T | Chemical | Total | LAB
CD-T | mg/L | Cadmium Total | F038 | HNO3/HCl dig,ICP-AES | Dissolved and extractable water samples are ready for analysis once collected. Total samples are autoclaved with HNO3 and hydrochloric acid (HCI) in capped Teflon vials. Analysis is performed via Inductively Coupled Argon Plasma Atomic Emission Spectrometry (ICP-AES)(simultaneous multielement analysis), Graphite Furnace Atomic Absorption Spectrometry (GF-AAS) (low level analysis) and/or Flame Atomic Absorption Spectrometry (F-AAS). | 0.006 | Y | Y | Y | 5429 | 8/1/1996 6:04:00 PM | 5/6/2008 | 9/11/2019 8:47:42 AM | 1997-05-08 | 6/3/1998 6:51:46 PM | MPAK |  | Concentration | mg/L | Y | (fl. conc.) |  | Cadmium Total (fl. conc.) | 7440-43-9 | Metals - Total; | Metals - Total; EMS code: CD-T | Chemical | Total | LAB
CD-T | ug/m3 | Cadmium Total | 5312 | Tef:N/P Dig:ICP Analysis | A nitric/perchloric acid digestion if used to bring the metals into solution prior to ICP analysis. For Hi-Vol samples preliminary results are reported as mg/L and 24 hour loading results as ug/m3. | 0.03 | Y | Y | Y | 1688 | 11/7/1997 | 3/29/2000 | 9/28/2000 1:34:41 PM | 1996-04-12 | 6/3/1998 6:51:34 PM | EMS |  | AirConcentration | µg/m³ | Y | (air conc.) |  | Cadmium Total (air conc.) | 7440-43-9 | Metals - Total; | Metals - Total; EMS code: CD-T | Chemical | Total | LAB
CD-T | mg/L | Cadmium Total | EQ01 | EQUIS (mg/L) | EQUIS (mg/L) | 0.0005 | N | Y | Y | 5199 |  |  | 10/24/1996 3:16:00 PM | 1996-04-12 | 6/3/1998 6:51:41 PM | EMS |  | Concentration | mg/L | Y | (fl. conc.) |  | Cadmium Total (fl. conc.) | 7440-43-9 | Metals - Total; | Metals - Total; EMS code: CD-T | Chemical | Total | LAB
CD-T | ug/g | Cadmium Total | 0192 | Hom, Nitric/Perchlor,ICP | Sample already in a dry state so does not require drying before homogenization and nitric perchlorid digestion followed by ICP analysis. | 1 | Y | Y | Y | 585 | 2/13/2002 | 5/29/2002 | 1/14/2004 1:00:43 AM | 1996-04-12 | 6/3/1998 6:51:34 PM | EMS |  | MassPerMass | mg/kg | Y | (mass conc.) |  | Cadmium Total (mass conc.) | 7440-43-9 | Metals - Total; | Metals - Total; EMS code: CD-T | Chemical | Total | LAB
CD-T | ug/g | Cadmium Total | X492 | Homo,HCl-HNO3 dig,ICP-OES,dry calc | Homogenize samples and take sub-samples. For Total Recoverable Metals Digest with Aqua Regia. Analyses by ICPOES (Inductively Coupled Plasma Optical Emission Spectroscopy). Perform moisture analysis on separate sub-sample and calculate dry value for trace elements. | 2 | Y | Y | Y | 1 |  |  | 8/24/2016 2:59:13 PM | 1998-11-13 |  | NPEPPIN |  | MassPerMass | mg/kg | Y | (mass conc.) |  | Cadmium Total (mass conc.) | 7440-43-9 | Metals - Total; | Metals - Total; EMS code: CD-T | Chemical | Total | LAB
CD-T | mg/L | Cadmium Total | X285 | HNO3 Dig:ICP Analysis | HNO3 Dig:ICP Analysis | 0.002 | Y | Y | Y | 1093 | 8/10/1993 | 10/5/1993 | 4/30/2004 1:59:45 PM | 1996-04-12 | 6/3/1998 6:51:41 PM | EMS |  | Concentration | mg/L | Y | (fl. conc.) |  | Cadmium Total (fl. conc.) | 7440-43-9 | Metals - Total; | Metals - Total; EMS code: CD-T | Chemical | Total | LAB
CD-T | ug/g | Cadmium Total | MTID | MET-DRY-CCMS-VA | Metals in Tissue by CRC ICPMS (DRY) Methodology: This method is conducted following British Columbia Lab Manual method "Metals in Animal Tissue and Vegetation (Biota) - Prescriptive". Tissue samples are homogenized and sub-sampled prior to hotblock digestion with nitric and hydrochloric acids, in combination with addition of hydrogen peroxide. Instrumental analysis is by collision cell inductively coupled plasma - mass spectrometry (modified from EPA Method 6020A). Method Limitation: This method employs a strong acid/peroxide digestion, and is intended to provide a conservative estimate of bio-available metals. Near complete recoveries are achieved for most toxicologically important metals, but elements associated with recalcitrant minerals may be only partially recovered. | 0.005 | Y | Y | Y | 122 | 8/13/2014 | 9/13/2022 | 10/26/2022 7:14:14 AM | 2015-03-09 | 3/11/2015 4:12:05 PM | RDGIBSON | EMS | MassPerMass | mg/kg | Y | (mass conc.) |  | Cadmium Total (mass conc.) | 7440-43-9 | Metals - Total; | Metals - Total; EMS code: CD-T | Chemical | Total | LAB
CD-T | mg/L | Cadmium Total | F024 | HNO3/HCl,GFAA | Dissolved and extractable water samples are ready for analysis once collected. Total samples are autoclaved with HNO3 and hydrochloric acid (HCl) in capped Teflon vials. Analysis is performed via Inductively Coupled Argon Plasma Atomic Emission Spectrometry (ICP-AES) (simultaneous multielement analysis), Graphite Furnace Atomic Absorption Spectrometry (GF-AAS) (low level analysis) and/or Flame Atomic Absorption Spectrometry (F-AAS). | 0.0002 | Y | Y | Y | 490 | 8/6/1996 7:16:00 PM | 7/5/2001 10:24:00 AM | 8/11/2001 2:02:37 AM | 1996-11-20 | 6/3/1998 6:51:46 PM | MPAK |  | Concentration | mg/L | Y | (fl. conc.) |  | Cadmium Total (fl. conc.) | 7440-43-9 | Metals - Total; | Metals - Total; EMS code: CD-T | Chemical | Total | LAB
CD-T | ug/g | Cadmium Total | X226 | Hom,HNO3/HCl,AA | Hom;HNO3/HCl;AA | 0.01 | Y | Y | Y | 182 |  |  | 3/19/1998 11:42:49 AM | 1996-04-12 | 6/3/1998 6:51:34 PM | EMS |  | MassPerMass | mg/kg | Y | (mass conc.) |  | Cadmium Total (mass conc.) | 7440-43-9 | Metals - Total; | Metals - Total; EMS code: CD-T | Chemical | Total | LAB
CD-T | mg/L | Cadmium Total | 0061 | HNO3 Dig: GFAA Analysis | HNO3 Dig: GFAA Analysis | 0.0001 | Y | Y | Y | 1926 | 12/6/1996 | 12/6/1996 | 3/8/2001 9:35:13 AM | 1996-04-12 | 6/3/1998 6:51:34 PM | EMS |  | Concentration | mg/L | Y | (fl. conc.) |  | Cadmium Total (fl. conc.) | 7440-43-9 | Metals - Total; | Metals - Total; EMS code: CD-T | Chemical | Total | LAB
CD-T | t/d | Cadmium Total | CALC | Calculation (t/d) | Calculation (t/d) | 0 | Y | Y | Y | 5775 |  |  | 7/3/2024 12:59:29 PM | 1996-04-12 | 6/3/1998 6:51:34 PM | EMS |  | MassFlow | t/d | Y | (mass flow) |  | Cadmium Total (mass flow) | 7440-43-9 | Metals - Total; | Metals - Total; EMS code: CD-T | Chemical | Total | LAB
CD-T | mg/L | Cadmium Total | 0050 | Graphite Furnace AA | Graphite Furnace AA | 0.0005 | Y | Y | Y | 49 | 12/27/2002 | 2/3/2005 | 5/16/2007 1:57:57 PM | 2002-09-10 |  | UMULLER |  | Concentration | mg/L | Y | (fl. conc.) |  | Cadmium Total (fl. conc.) | 7440-43-9 | Metals - Total; | Metals - Total; EMS code: CD-T | Chemical | Total | LAB
CD-T | mg/L | Cadmium Total | X072 | FA,HNO3,Graph.Furnace,AA | Samples are field-acidified with HNO3; HNO3 digestion; analysis by GFAA. | 0.0002 | Y | Y | Y | 7215 |  |  | 8/26/2024 2:41:32 PM | 1996-04-12 | 6/3/1998 6:51:34 PM | EMS |  | Concentration | mg/L | Y | (fl. conc.) |  | Cadmium Total (fl. conc.) | 7440-43-9 | Metals - Total; | Metals - Total; EMS code: CD-T | Chemical | Total | LAB
CD-T | mg/L | Cadmium Total | X512 | HNO3,pHto7.5,MIK extr'n,AA | HNO3 Digestion;pH to 7.5; MIK solvent extraction;AA | 0.001 | Y | Y | Y | 2981 |  |  | 8/23/1999 4:52:44 PM | 1999-04-22 |  | NPEPPIN |  | Concentration | mg/L | Y | (fl. conc.) |  | Cadmium Total (fl. conc.) | 7440-43-9 | Metals - Total; | Metals - Total; EMS code: CD-T | Chemical | Total | LAB
CD-T | ug/m3 | Cadmium Total | X180 | LVOL:Dig,AA | LVOL:Dig;AA | 0 | Y | Y | Y | 1826 |  |  | 3/9/1998 10:35:51 AM | 1996-04-12 | 6/3/1998 6:51:34 PM | EMS |  | AirConcentration | µg/m³ | Y | (air conc.) |  | Cadmium Total (air conc.) | 7440-43-9 | Metals - Total; | Metals - Total; EMS code: CD-T | Chemical | Total | LAB
CD-T | mg/L | Cadmium Total | 0051 | Graphite Furnace AA | Graphite Furnace | 0.0001 | Y | Y | Y | 72 | 8/26/2002 | 9/12/2003 3:16:00 PM | 12/1/2003 5:30:18 PM | 1996-04-12 | 6/3/1998 6:51:41 PM | EMS |  | Concentration | mg/L | Y | (fl. conc.) |  | Cadmium Total (fl. conc.) | 7440-43-9 | Metals - Total; | Metals - Total; EMS code: CD-T | Chemical | Total | LAB
CD-T | ug/g | Cadmium Total | ICPM | Metals in Soil by CRC ICPMS | This analysis is carried out using procedures from CSR Analytical Method: "Strong Acid Leachable Metals (SALM) in Soil", BC Ministry of Environment, 26 June 2009, and procedures adapted from EPA Method 200.2. The sample is manually homogenized, dried at 60 degrees Celsius, sieved through a 63 um (230 mesh) sieve, and a representative subsample of the dry material is weighed. The sample is then digested at 95 degrees Celsius for 2 hours by block digester using concentrated nitric and hydrochloric acids. Instrumental analysis is by collision cell inductively coupled plasma - mass spectrometry (modifed from EPA Method 6020A). Method Limitation: This method is not a total digestion technique. It is a very strong acid digestion that is intended to dissolve those metals that may be environmentally available. By design, elements bound in silicate structures are not normally dissolved by this procedure as they are not usually mobile in the environment. | 0.05 | Y | Y | Y | 103 | 9/19/2017 | 9/24/2023 | 1/12/2024 11:16:52 AM | 2018-01-09 |  | JKROGH |  | MassPerMass | mg/kg | Y | (mass conc.) |  | Cadmium Total (mass conc.) | 7440-43-9 | Metals - Total; | Metals - Total; EMS code: CD-T | Chemical | Total | LAB
CD-T | mg/L | Cadmium Total | 0105 | 5%AR Dig:APDC Ext:AA Anl | 5%AR Dig:APDC Ext:AA Anl | 0.0005 | Y | Y | Y | 520 |  |  | 10/24/1996 2:13:46 PM | 1996-04-12 | 6/3/1998 6:51:41 PM | EMS |  | Concentration | mg/L | Y | (fl. conc.) |  | Cadmium Total (fl. conc.) | 7440-43-9 | Metals - Total; | Metals - Total; EMS code: CD-T | Chemical | Total | LAB
CD-T | ug/m3 | Cadmium Total | EQ40 | EQUIS (mg/m3) | EQUIS (mg/m3) | 0.001 | N | Y | Y | 3039 |  |  | 10/23/1996 12:01:36 AM | 1996-04-12 | 6/3/1998 6:51:41 PM | EMS |  | AirConcentration | µg/m³ | Y | (air conc.) |  | Cadmium Total (air conc.) | 7440-43-9 | Metals - Total; | Metals - Total; EMS code: CD-T | Chemical | Total | LAB
CD-T | ppm (S) | Cadmium Total | 6041 | Aqua Regia by ICPMS | These were sediment/soils digested with Aqua-Regia (Nitric Acid & Hydrochloric Acids) and then analysis by ICPMS. Results would be reported in mg/kg and/or %. | 0.001 | Y | Y | Y | 21 | 10/13/2019 | 8/5/2020 | 8/10/2020 3:52:06 PM | 2019-09-06 | 10/21/2019 11:54:32 AM | KATFRASE | KATFRASE | DimensionlessRatio | ppm | N | (ratio) |  | Cadmium Total (ratio) | 7440-43-9 | Metals - Total; | Metals - Total; EMS code: CD-T | Chemical | Total | LAB
CD-T | ppm (S) | Cadmium Total | 6042 | Aqua Regia by ICPMS | These were sediment/soils digested with Aqua-Regia (Nitric Acid & Hydrochloric Acids) and then analysis by ICPMS. Results would be reported in ppm s. | 0.001 | Y | Y | Y | 71 | 5/14/2022 | 7/27/2024 | 7/29/2024 9:13:08 AM | 2022-06-28 |  | EMS |  | DimensionlessRatio | ppm | N | (ratio) |  | Cadmium Total (ratio) | 7440-43-9 | Metals - Total; | Metals - Total; EMS code: CD-T | Chemical | Total | LAB
CD-T | mg/L | Cadmium Total | 1349 | Tot Metals in Seawater by CRC ICPMS (BC) | Tot Metals in Seawater by CRC ICPMS (BC) | 1e-05 | Y | Y | Y | 98 | 8/16/2019 | 12/5/2021 | 12/13/2021 12:34:16 PM | 2019-09-30 |  | KATFRASE |  | Concentration | mg/L | Y | (fl. conc.) |  | Cadmium Total (fl. conc.) | 7440-43-9 | Metals - Total; | Metals - Total; EMS code: CD-T | Chemical | Total | LAB
CD-T | mg/dm2/d | Cadmium Total | EQ15 | EQUIS (mg/dm2/d) | EQUIS (mg/dm2/d) | 0.001 | N | Y | Y | 200 |  |  | 10/22/1996 10:18:48 PM | 1996-04-12 | 6/3/1998 6:51:34 PM | EMS |  | MassFlux | mg/dm²/d | Y | (mass flux) |  | Cadmium Total (mass flux) | 7440-43-9 | Metals - Total; | Metals - Total; EMS code: CD-T | Chemical | Total | LAB
CD-T | None | Cadmium Total | AISP | Special Analysis | Special Analysis | 0 | N | Y | N | 7 |  |  | 10/24/1996 11:08:37 AM | 1996-04-12 | 6/3/1998 6:51:34 PM | EMS |  | None | None | N | (unitless) |  | Cadmium Total (unitless) | 7440-43-9 | Metals - Total; | Metals - Total; EMS code: CD-T | Chemical | Total | LAB
CD-T | mg/L | Cadmium Total | 0060 | HNO3 Dig: HGA Analysis | HNO3 Dig: HGA Analysis | 0.0005 | Y | Y | Y | 3706 |  |  | 2/4/2000 9:38:26 AM | 1996-04-12 | 6/3/1998 6:51:34 PM | EMS |  | Concentration | mg/L | Y | (fl. conc.) |  | Cadmium Total (fl. conc.) | 7440-43-9 | Metals - Total; | Metals - Total; EMS code: CD-T | Chemical | Total | LAB
CD-T | mg/L | Cadmium Total | 0040 | HNO3 Dig: ICP Analysis | HNO3 Dig: ICP Analysis | 0.01 | Y | Y | Y | 11437 |  |  | 10/18/2023 8:55:31 AM | 1996-04-12 | 6/3/1998 6:51:34 PM | EMS |  | Concentration | mg/L | Y | (fl. conc.) |  | Cadmium Total (fl. conc.) | 7440-43-9 | Metals - Total; | Metals - Total; EMS code: CD-T | Chemical | Total | LAB
CD-T | mg/L | Cadmium Total | 0200 | 20%A.R.Dig:AA Anal. | 20%A.R.Dig:AA Anal. | 0.01 | Y | Y | Y | 7 |  |  | 10/24/1996 2:35:04 AM | 1996-04-12 | 6/3/1998 6:51:34 PM | EMS |  | Concentration | mg/L | Y | (fl. conc.) |  | Cadmium Total (fl. conc.) | 7440-43-9 | Metals - Total; | Metals - Total; EMS code: CD-T | Chemical | Total | LAB
CD-T | mg/L | Cadmium Total | X545 | FA,HNO3,ICPMS | Analytes nebulized followed by ion entrained in plasma gas and extracted by means of a differentially pumped vacuum interface into a mass spectrometer | 0.0001 | Y | Y | Y | 30 |  |  | 10/30/2014 8:31:17 AM | 2001-09-18 | 9/18/2001 1:36:23 PM | NPEPPIN | NPEPPIN | Concentration | mg/L | Y | (fl. conc.) |  | Cadmium Total (fl. conc.) | 7440-43-9 | Metals - Total; | Metals - Total; EMS code: CD-T | Chemical | Total | LAB
CD-T | mg/L | Cadmium Total | X113 | HNO3/HCl/Perchlor,AA | HNO3/HCl/Perchlor;AA | 0.0005 | Y | Y | Y | 835 |  |  | 12/9/1996 10:18:15 PM | 1996-04-12 | 6/3/1998 6:51:34 PM | EMS |  | Concentration | mg/L | Y | (fl. conc.) |  | Cadmium Total (fl. conc.) | 7440-43-9 | Metals - Total; | Metals - Total; EMS code: CD-T | Chemical | Total | LAB
CD-T | ug/g | Cadmium Total | FS89 | Homo,HNO3 dig,ICP-AES | Homo;HNO3 dig;ICP-AES | 0.8 | Y | Y | Y | 4 | 1/5/2001 | 8/20/2001 12:00:00 PM | 10/23/2002 1:42:43 PM | 2001-03-08 |  | NPEPPIN |  | MassPerMass | mg/kg | Y | (mass conc.) |  | Cadmium Total (mass conc.) | 7440-43-9 | Metals - Total; | Metals - Total; EMS code: CD-T | Chemical | Total | LAB
CD-T | ug/m3 | Cadmium Total | 6040 | Aqua Regia Dig:ICP | Aqua regia digestion is used to bring the metals into solution prior to ICP analysis. For Hi-Vol samples, preliminary results are reported as mg/L and24 hr loadings as ug/m3 | 0.03 | Y | Y | Y | 437 | 2/13/2002 | 2/9/2005 | 2/23/2005 1:13:47 AM | 1998-05-05 | 6/3/1998 6:51:48 PM | NPEPPIN |  | AirConcentration | µg/m³ | Y | (air conc.) |  | Cadmium Total (air conc.) | 7440-43-9 | Metals - Total; | Metals - Total; EMS code: CD-T | Chemical | Total | LAB
CD-T | mg/dm2/d | Cadmium Total | DUST | HNO3 dig.(microwave or hot block),ICPMS | HNO3 dig.(microwave or hot block);ICPMS |  | Y | Y | Y | 705 | 7/19/2016 1:11:00 PM | 7/31/2024 | 7/31/2024 11:23:12 AM | 2008-02-07 |  | RDGIBSON |  | MassFlux | mg/dm²/d | Y | (mass flux) |  | Cadmium Total (mass flux) | 7440-43-9 | Metals - Total; | Metals - Total; EMS code: CD-T | Chemical | Total | LAB
CD-T | ug/g | Cadmium Total | X208 | Hom,Nitr/Perchl,Grph Fur | Hom;Nitr/Perchl;Grph Fur | 0.025 | Y | Y | Y | 36 |  |  | 10/24/1996 2:25:59 PM | 1996-04-12 | 6/3/1998 6:51:41 PM | EMS |  | MassPerMass | mg/kg | Y | (mass conc.) |  | Cadmium Total (mass conc.) | 7440-43-9 | Metals - Total; | Metals - Total; EMS code: CD-T | Chemical | Total | LAB
CD-T | ug/g | Cadmium Total | 0196 | D&G: LTA/HCl: ICP | D&G: LTA/HCl: ICP | 1 | Y | Y | Y | 26 |  |  | 10/24/1996 12:50:32 PM | 1996-04-12 | 6/3/1998 6:51:41 PM | EMS |  | MassPerMass | mg/kg | Y | (mass conc.) |  | Cadmium Total (mass conc.) | 7440-43-9 | Metals - Total; | Metals - Total; EMS code: CD-T | Chemical | Total | LAB
CD-T | ug/g wet | Cadmium Total | X552 | Metals in Biota (wet), total, acid digestion, ICP-SFMS | One gram of wet, homogenized biota is weighed in a microwave TFM vessel, and 8mL of Nitric Acid and 1mL of Hydrogen Peroxide are added. The sample is digested at 200 degrees C for 15 minutes in a closed vessel high pressure microwave oven. The sample is cooled, the volume adjusted to 100mL and introduced to an inductively coupled argon plasma-sector field mass spectrometer (ICP-SFMS) by pneumatic nebulization. Each respective element is measured at a specific mass to charge ratio, m/z value expressed in amu (atomic mass units), and at a specified instrument resolution. LR: Low Resolution ~300, MR: Mid Resolution ~4000. Ref: 02-2704, National Laboratory for Environmental Testing, CCIW, Environment Canada, Burlington. Requested April 2003. | 0.0001 | Y | Y | Y | 37 | 8/9/2016 | 10/16/2023 | 11/28/2023 2:21:32 PM | 2004-01-19 | 4/19/2021 1:22:00 PM | UMULLER | EMS | MassPerMass | µg/g | Y | (mass conc.) | (wet) | Cadmium Total (wet) (mass conc.) | 7440-43-9 | Metals - Total; | Metals - Total; EMS code: CD-T | Chemical | Total | LAB
CD-T | ug/g wet | Cadmium Total | MTIW | MET-WET-CCMS-MID-VA | Metals in Tissue by CRC ICPMS (WET) Methodology: This method is conducted following British Columbia Lab Manual method "Metals in Animal Tissue and Vegetation (Biota) - Prescriptive". Tissue samples are homogenized and sub-sampled prior to hotblock digestion with nitric and hydrochloric acids, in combination with addition of hydrogen peroxide. Instrumental analysis is by collision cell inductively coupled plasma - mass spectrometry (modified from EPA Method 6020A). Method Limitation: This method employs a strong acid/peroxide digestion, and is intended to provide a conservative estimate of bio-available metals. Near complete recoveries are achieved for most toxicologically important metals, but elements associated with recalcitrant minerals may be only partially recovered. | 0.002 | Y | Y | Y | 24 | 11/25/2018 | 12/6/2018 | 7/21/2021 2:55:07 PM | 2018-12-24 |  | EMS |  | MassPerMass | µg/g | Y | (mass conc.) | (wet) | Cadmium Total (wet) (mass conc.) | 7440-43-9 | Metals - Total; | Metals - Total; EMS code: CD-T | Chemical | Total | LAB
CD-T | ug/g | Cadmium Total | EQ02 | EQUIS (ug/g) | EQUIS (ug/g) | 1 | N | Y | Y | 865 |  |  | 10/23/1996 12:08:28 AM | 1996-04-12 | 6/3/1998 6:51:34 PM | EMS |  | MassPerMass | mg/kg | Y | (mass conc.) |  | Cadmium Total (mass conc.) | 7440-43-9 | Metals - Total; | Metals - Total; EMS code: CD-T | Chemical | Total | LAB
CD-T | mg/L | Cadmium Total | X555 | Metals in water, total, "in-bottle" digestion, ICP-SFMS | The whole water sample collected in a HDPE/Teflon bottle, is de-capped in a Class 100 workstation, acidified to 2% (v/v) HNO3 and 1% (v/v) HCl and recapped. The sample is "in-bottle" digested in a convection oven, set at 60 degrees C for 16 hours. The sample is cooled and introduced to an inductively coupled argon plasma-sector field mass spectrometer (ICP-SFMS) by pneumatic nebulization. Each respective element is measured at a specific mass to charge ratio, m/z value expressed in amu (atomic mass units), and at resolution. LR: Low Resolution ~300, MR: Mid Resolution ~4000 and HR: High a specified instrument Resolution ~10000. Ref: Method number 02-2003, National Laboratory for Environmental Testing, CCIW, Environment Canada, Burlington. Requested April 2003. | 1e-06 | Y | Y | Y | 17342 | 5/20/2003 | 12/6/2023 2:08:00 AM | 2/14/2024 9:05:32 AM | 2004-01-14 |  | UMULLER |  | Concentration | mg/L | Y | (fl. conc.) |  | Cadmium Total (fl. conc.) | 7440-43-9 | Metals - Total; | Metals - Total; EMS code: CD-T | Chemical | Total | LAB
CD-T | ug/m3 | Cadmium Total | F088 | Aliquot,HNO3/HCl dig, ICPMS,loading | Aliquot microwave digested with 3:1 HNO3/HCL. Analysis by ICPMS. For HiVol samples preliminary result are reported in mg/L and loading results are reported in ug/m3 |  | Y | Y | Y | 565 | 8/17/2000 10:46:00 AM | 12/14/2006 | 12/30/2006 1:01:36 AM | 2000-08-16 |  | NPEPPIN |  | AirConcentration | µg/m³ | Y | (air conc.) |  | Cadmium Total (air conc.) | 7440-43-9 | Metals - Total; | Metals - Total; EMS code: CD-T | Chemical | Total | LAB
CD-T | mg/L | Cadmium Total | X179 | LA,HNO3,Graphite Furnace | Samples are lab-acidified with HNO3; HNO3 digestion; analysis by GFAA. | 0.0001 | Y | Y | Y | 185 | 2/11/1998 | 7/23/2002 | 8/6/2002 5:43:17 PM | 1996-04-12 | 6/3/1998 6:51:34 PM | EMS |  | Concentration | mg/L | Y | (fl. conc.) |  | Cadmium Total (fl. conc.) | 7440-43-9 | Metals - Total; | Metals - Total; EMS code: CD-T | Chemical | Total | LAB
CD-T | mg/L | Cadmium Total | F068 | HNO3dig,TeflonTube/Autoclave,ICP-MS | This method is applicable to determiniation of dissolved, extractable, and total recoverable metals in ground waters, surface waters, drinking waters, and wastewaters. Dissolved, extractable, and total recoverable sample fractions are acidified to pH <1.5 with analytical grade nitric acid (HNO3) to prevent the metals from adsorbing to the container material or precipitating. Extractable and total recoverable water samples are collected in 250 ml acid washed polyethylene bottles and are preserved with Nitric acid (0.5 ml acid for 100 mLs of sample). All preserved samples are stored at room temperature and can be held for up to 6 months. Unfiltered, un-preserved samples are stored at 4C for no more than 30 days. For dissolved and extractable water fractions, no sample preparation is necessary. Samples are simply poured in autosampler tubes (i.e. 15 ml centrifuge tubes), and placed on the autosampler and analyzed. Samples for total recoverable metals are digested with Nitric acid in special Teflon tubes in an autoclave prior to analyses. | 1e-05 | Y | Y | Y | 4073 | 6/1/1999 1:32:00 PM | 7/4/2008 9:48:00 AM | 7/30/2024 6:32:12 AM | 1999-06-28 |  | NPEPPIN |  | Concentration | mg/L | Y | (fl. conc.) |  | Cadmium Total (fl. conc.) | 7440-43-9 | Metals - Total; | Metals - Total; EMS code: CD-T | Chemical | Total | LAB
CD-T | mg/dm2/d | Cadmium Total | CAL1 | Calculation (mg/dm2/d) | Calculation (mg/dm2/d) | 0.0006 | Y | Y | Y | 3040 |  |  | 7/3/2024 12:59:27 PM | 1996-04-12 | 6/3/1998 6:51:41 PM | EMS |  | MassFlux | mg/dm²/d | Y | (mass flux) |  | Cadmium Total (mass flux) | 7440-43-9 | Metals - Total; | Metals - Total; EMS code: CD-T | Chemical | Total | LAB
CD-T | mg/L | Cadmium Total | X078 | Aqua Reg.Dig,Hydr.gen,AA | Aqua Reg.Dig;Hydr.gen;AA | 0.0005 | Y | Y | Y | 232 |  |  | 10/29/2001 11:34:41 AM | 1996-04-12 | 6/3/1998 6:51:34 PM | EMS |  | Concentration | mg/L | Y | (fl. conc.) |  | Cadmium Total (fl. conc.) | 7440-43-9 | Metals - Total; | Metals - Total; EMS code: CD-T | Chemical | Total | LAB
CD-T | ug/m3 | Cadmium Total | X136 | HNO3/HCl/Perchlor,AA | HNO3/HCl/Perchlor;AA | 0 | Y | Y | Y | 317 |  |  | 3/9/1998 8:54:23 AM | 1996-04-12 | 6/3/1998 6:51:42 PM | EMS |  | AirConcentration | µg/m³ | Y | (air conc.) |  | Cadmium Total (air conc.) | 7440-43-9 | Metals - Total; | Metals - Total; EMS code: CD-T | Chemical | Total | LAB
CD-T | mg/L | Cadmium Total | X170 | ICP | This code is a catch-all for a large variety of ICP methods including water and also solids digested into solutions as an intermediate result. Different procedures will have greatly different detection limits. (a) For air particulates, preliminary results are reported in mg/L and loading results as mg/dm2/d. (b) For Hi Vol samples preliminiary results are same as reported in mg/L and loading results as ug/m3. (c) Note that for Rain samples there is no filtration nor digestion steps. | 0.003 | Y | Y | Y | 10817 | 3/28/1996 6:03:00 PM | 11/7/2018 | 7/25/2024 11:22:07 AM | 1996-04-12 | 6/3/1998 6:51:34 PM | EMS |  | Concentration | mg/L | Y | (fl. conc.) |  | Cadmium Total (fl. conc.) | 7440-43-9 | Metals - Total; | Metals - Total; EMS code: CD-T | Chemical | Total | LAB
CD-T | mg/L | Cadmium Total | X352 | LA:HNO3:AA | Samples are lab-acidified with HNO3; HNO3 digestion; analysis by ICP/AES. | 0.01 | Y | Y | Y | 36 | 4/2/2003 | 10/31/2003 | 1/20/2004 12:08:07 PM | 1996-04-12 | 6/3/1998 6:51:34 PM | EMS |  | Concentration | mg/L | Y | (fl. conc.) |  | Cadmium Total (fl. conc.) | 7440-43-9 | Metals - Total; | Metals - Total; EMS code: CD-T | Chemical | Total | LAB
CD-T | ug/g wet | Cadmium Total | BPSY | MET-WET-BPSY-CCMS-VA | Metals in Animal Tissue and Vegetation (Biota) - Prescriptive. EPA 200.3/6020B This method is conducted following British Columbia Lab Manual method "Metals in Animal Tissue and Vegetation (Biota) - Prescriptive". Tissue samples are digested with nitric and hydrochloric acids; in combination with addition of hydrogen peroxide. Instrumental analysis is by collision cell inductively coupled plasma - mass spectrometry (modified from EPA Method 6020B). | 0.002 | Y | Y | Y | 2 | 11/7/2019 | 11/7/2019 | 6/10/2021 1:57:10 PM | 2021-06-04 |  | EMS |  | MassPerMass | µg/g | Y | (mass conc.) | (wet) | Cadmium Total (wet) (mass conc.) | 7440-43-9 | Metals - Total; | Metals - Total; EMS code: CD-T | Chemical | Total | LAB
CD-T | mg/L | Cadmium Total | F033 | Acid Dig,Centrifuge,ICPMS | For the determination of total recoverable elements in water or waste water, take 100 ml aliquot from a well mixed, acid preserved, sample containing not more than 0.25% (w/v) total solids and transfer to a 250 ml Griffin beaker. (If total solids are greater than 0.25% reduce the size of the aliquot by a proportionate amount). Add 1 ml of conc. nitric acid and 0.5 ml conc. hydrochloric acid. Heat on a hot plate until the volume has been reduced to 15-20ml, ensuring that the sample does not boil. Allow to cool and transfer to a 50 ml volumetric flask. Dilute to volume with ASTM type I water. Centrifuge the sample or allow to stand overnight to separate insoluble material. Prior analysis, pipette 20 ml into a 50 ml volumetric flask and dilute to volume with ASTM type I water. Analysis should be performed within seventy two hours of the completed preparation. Note that different laboratories may have very different field preservation, digestion procedures and detection limits for this method, so these should be documented. | 0.0001 | Y | Y | Y | 40254 | 10/15/1996 9:27:00 AM | 8/22/2024 3:24:00 AM | 8/25/2024 9:47:36 PM | 1997-01-16 | 6/3/1998 6:51:46 PM | MPAK |  | Concentration | mg/L | Y | (fl. conc.) |  | Cadmium Total (fl. conc.) | 7440-43-9 | Metals - Total; | Metals - Total; EMS code: CD-T | Chemical | Total | LAB
CD-T | mg/L | Cadmium Total | F048 | Acid dig,conc'n,nebul.,ICP-OES | This method provides an automated procedure for the Inductively Coupled Argon Plasma Optical Emission Spectrophotometric (ICP-OES) determination and quantitation of trace amounts of metals in surface, ground and precipitation waters after a manual acid digestion and concentration procedure and ultrasonic nebulisation. Organically-bound metals and metal complexes are converted to metal ions by digestion with nitric acid [method details says HCl]. To increase the signal detected by the ICP, the samples are concentrated 10 times. The liquid sample is pumped onto the piezoelectric transducer, where the ultrasonic action efficiently converts it to a fine, dense aerosol. The nebulizer carrier gas flow then transports the wet aerosol through a heated U-tube, where the solvent is vaporized. Subsequently, most of the solvent vapour is removed from the aerosol stream by passing the stream through a chilled condenser. The result is a dry, analyte-laden aerosol which is introduced to the plasma. There, the high temperature desolvates the aerosol and dissociates the sample material into individual atoms and ions, which are excited to emit light at wave lengths characteristic of the elements present. Dissolved samples are field filtered through 0.45 micron cellulose acetate filter prior to acid preservation. Samples for extractable analysis are acidified with 2 mL concentrated HNO3 per litre and allowed to stand a minimum of 24 hours before analysis. Clients providing these samples are advised to preserve samples at the collection site. | 0.0001 | Y | Y | Y | 6159 | 7/4/1997 2:35:00 PM | 4/23/2003 | 9/11/2019 8:46:38 AM | 1997-10-20 | 6/3/1998 6:51:47 PM | NPEPPIN |  | Concentration | mg/L | Y | (fl. conc.) |  | Cadmium Total (fl. conc.) | 7440-43-9 | Metals - Total; | Metals - Total; EMS code: CD-T | Chemical | Total | LAB
CD-T | ug/g | Cadmium Total | FS01 | D&G,Nitric/HCl:ICP | Samples are dried at 60 C, then either sieved to -100mesh (150 um) and rolled to homogenize, or finely ground with a mortar and pestle or a shatter box laboratory mill and rolled (depending on the submitter's requirements). The samples are then weighed into Teflon digestion liners and digested with HNO3 and HCl in a microwave oven. Samples are cooled, volumized, and allowed to settle overnight before being decanted. Analysis is then performed via Inductively Coupled Argon Plasma Atomic Emission Spectrometry (ICP-AES) (Simultaneous multi element analysis). In ICP Spectrometry, the liquid sample is pumped to the nebulizer, where it is convertd into an aerosol. the nebulizer carrier gas flow than transports the wet aerosol past the impact bead, where the large droplets condense out and flow to the waste container. The remaining fine aerosol is carried into the plasma. There, the high temperature desolvates the aerosol and dissociates the sample material into individual atoms and ions, which are excited to emit light at wavelengths characteristic of the elements present. A diffraction grating in the spectrometer diffracts the various wavelengths of light through exit slits to the photo-multiplier tubes, which convert the light energy to electronic signals. | 0.8 | Y | Y | Y | 2641 | 8/30/1996 1:52:00 PM | 10/10/2007 | 5/24/2008 1:15:02 AM | 1996-11-20 | 6/3/1998 6:51:46 PM | MPAK |  | MassPerMass | mg/kg | Y | (mass conc.) |  | Cadmium Total (mass conc.) | 7440-43-9 | Metals - Total; | Metals - Total; EMS code: CD-T | Chemical | Total | LAB
CD-T | ug/g wet | Cadmium Total | X560 | Metals in Biota (wet), total, acid digestion, ICP-MS | Five grams of wet, homogenized biota is weighed in a centrifuge tube, and 2.5 mL of Nitric Acid and 2.5 mL of Hydrochloric acid are added. The sample is digested at room temperature for one hour and then digested at 90-95 degrees C in a water bath or heat block for one more hour. The sample is cooled and the volume is made to 50 mL with DI. The digestate is then diluted at ten fold and introduced to a collision/reaction cell inductively coupled argon plasma mass spectrometer (CRC ICPMS) by pneumatic nebulization. Each respective element is measured at a specific mass to charge ratio, m/z value expressed in amu (atomic mass units). - Maxxam Analytics (BC). Added to EMS Dec 10, 2008. | 0.01 | Y | Y | Y | 541 | 11/15/2008 4:00:00 AM | 10/30/2023 | 11/28/2023 2:21:33 PM | 2008-12-11 | 3/11/2015 4:17:10 PM | RDGIBSON | EMS | MassPerMass | µg/g | Y | (mass conc.) | (wet) | Cadmium Total (wet) (mass conc.) | 7440-43-9 | Metals - Total; | Metals - Total; EMS code: CD-T | Chemical | Total | LAB
CD-T | mg/L | Cadmium Total | TCLP | Toxicity Characteristic Leaching Procedure | The TCLP procedure is designed to simulate the climatic leaching action | 0.1 | Y | Y | Y | 38 | 12/8/2009 7:49:00 PM | 9/28/2011 5:53:00 PM | 8/24/2016 2:59:13 PM | 2009-03-05 |  | JCHUDAK |  | Concentration | mg/L | Y | (fl. conc.) |  | Cadmium Total (fl. conc.) | 7440-43-9 | Metals - Total; | Metals - Total; EMS code: CD-T | Chemical | Total | LAB
CD-T | ug/g | Cadmium Total | 0190 | D&G:N/P Dig:ICP Anal. | Sample is dried and grind-mixed before nitric perchloric digestion followed by ICP analysis. | 1 | Y | Y | Y | 875 |  |  | 10/24/1996 3:05:58 PM | 1996-04-12 | 6/3/1998 6:51:34 PM | EMS |  | MassPerMass | mg/kg | Y | (mass conc.) |  | Cadmium Total (mass conc.) | 7440-43-9 | Metals - Total; | Metals - Total; EMS code: CD-T | Chemical | Total | LAB
CD-T | ug/g | Cadmium Total | X227 | Hom,HNO3/HCl,Grph Furn | Hom;HNO3/HCl;Grph Furn | 0.01 | Y | Y | Y | 4 |  |  | 9/23/2003 9:46:56 PM | 1996-04-12 | 6/3/1998 6:51:34 PM | EMS |  | MassPerMass | mg/kg | Y | (mass conc.) |  | Cadmium Total (mass conc.) | 7440-43-9 | Metals - Total; | Metals - Total; EMS code: CD-T | Chemical | Total | LAB
CD-T | mg/L | Cadmium Total | MSEA | Metals in Seawater (undigested) by (CRC) ICP-MS | This method describes the multi-elemental determination of trace elements by Collision/Reaction Cell (CRC) ICP-MS. The method measures ions produced by a radio-frequency inductively coupled plasma. Analyte species originating in a liquid are pneumatically nebulized and the resulting aerosol transported by argon gas into the plasma torch. The ions produced are entrained in the plasma gas and extracted, by means of a differentially pumped vacuum interface, into a mass spectrometer. The ions flow through a Collision/Reaction Cell (CRC) where depending on the mode of analysis either the ions passes through the cell unhindered (no gas - classical ICP-MS mode), or collisions/reactions chemistry takes place between the ions and a neutral gas such as hydrogen or helium. Any polyatomic ions which give interference on the analytes are removed. The ions are then sorted according to their mass-to-charge ratios by a quadrupole mass spectrometer. The ions are then sorted according to their mass-to-charge ratios by a quadrupole mass spectrometer. The ions transmitted through the quadrupole are quantified by a channel electron multiplier and the ion information processed by a data handling system. Although most polyatomic interference is removed, any additional interference must be assessed and valid corrections applied or the data flagged to indicate problems. Interference correction must include compensation for background ions contributed by the plasma gas, reagents, and constituents of the sample matrix. Instrument drift as well as suppressions or enhancements of instrument response caused by the sample matrix must be corrected for by the use of internal standards. | 2e-05 | Y | Y | Y | 432 | 8/12/2009 12:08:00 AM | 4/29/2024 | 6/7/2024 10:03:16 AM | 2009-09-08 |  | RDGIBSON |  | Concentration | mg/L | Y | (fl. conc.) |  | Cadmium Total (fl. conc.) | 7440-43-9 | Metals - Total; | Metals - Total; EMS code: CD-T | Chemical | Total | LAB
CE-D | mg/L | Cerium Dissolved | X556 | Metals in water, dissolved, "in-bottle" digestion, ICP-SFMS | The whole water sample is filtered through a 0.45 micron cellulose acetate filter paper and the filtrate is stored in a HDPE/Teflon bottle. The container is de-capped in a Class 100 workstation, acidified to 2% (v/v) HNO3 and 1% (v/v) HCl and recapped. The sample is "in-bottle" digested in a convection oven, set at 60 degrees C for 16 hours. The sample is cooled and introduced to an inductively coupled argon plasma-sector field mass spectrometer (ICP-SFMS) by pneumatic nebulization. Each respective element is measured at a specific mass to charge ratio, m/z value expressed in amu (atomic mass units), and at a specified intrument resolution. LR: Low Resolution ~300, MR: Mid Resolution ~4000 and HR: High Resolution ~10000. Ref: Method number 02-2003, National Laboratory for Environmental Testing, CCIW, Environment Canada, Burlington, Requested April 2003. | 5e-06 | Y | Y | Y | 4159 | 6/27/2004 | 12/6/2023 2:04:00 AM | 2/14/2024 9:05:32 AM | 2008-06-04 | 6/4/2008 4:51:31 PM | RDGIBSON | RDGIBSON | Concentration | mg/L | Y | (fl. conc.) |  | Cerium Dissolved (fl. conc.) | 7440-45-1 | Metals - Dissolved; | Metals - Dissolved; EMS code: CE-D | Chemical | Dissolved | LAB
CE-E | mg/L | Cerium Extractable | X553 | Metals in water, extractable, ICP-SFMS | The whole water sample is collected in a HDPE/Teflon bottle. In a Class 100 workstation it is preserved with concentrated Nitric Acid at concentration of 0.2 percent (v/v). The acid preservative extracts the metals from material in the sample and a decanted portion is introduced directly to an inductively coupled argon plasma-sector field mass spectrometer (ICP-SFMS) by pneumatic nebulization. Each respective element is measured at a specific mass to charge ratio, m/z value expressed in amu (atomic mass units), and at a specified instrument resolution. LR: Low Resolution ~300, MR: Mid Resolution ~4000 and HR: High Resolution ~10000. Ref: Method number 02-2061, National Laboratory for Environmental Testing, CCIW, Environment Canada, Burlington. Requested April 2003. | 0.002 | Y | Y | Y | 1716 | 6/27/2004 | 9/21/2007 | 4/18/2018 9:28:02 AM | 2009-05-14 | 5/15/2009 11:20:54 AM | RDGIBSON | RDGIBSON | Concentration | mg/L | Y | (fl. conc.) |  | Cerium Extractable (fl. conc.) | 7440-45-1 | Metals - Extractable; | Metals - Extractable; EMS code: CE-E | Chemical | Extractable | LAB
CE-T | ppm (S) | Cerium Total | MS41 | ALS Minerals ME-MS41 | ALS Minerals ME-MS41. A prepared sample is digested with aqua regia in a graphite heating block. After cooling; the resulting solution is diluted to with deionized water; mixed and analyzed by ICP-AES. Samples are then analyzed by ICP-MS for the remaining suite of elements. | 0.02 | Y | Y | Y | 0 |  |  |  | 2024-01-29 |  | EMS |  | DimensionlessRatio | ppm | N | (ratio) |  | Cerium Total (ratio) | 7440-45-1 | Metals - Total; | Metals - Total; EMS code: CE-T | Chemical | Total | LAB
CE-T | ppm (S) | Cerium Total | 6042 | Aqua Regia by ICPMS | These were sediment/soils digested with Aqua-Regia (Nitric Acid & Hydrochloric Acids) and then analysis by ICPMS. Results would be reported in ppm s. | 0.003 | Y | Y | Y | 71 | 5/14/2022 | 7/27/2024 | 7/29/2024 9:13:08 AM | 2022-06-28 |  | EMS |  | DimensionlessRatio | ppm | N | (ratio) |  | Cerium Total (ratio) | 7440-45-1 | Metals - Total; | Metals - Total; EMS code: CE-T | Chemical | Total | LAB
CE-T | ppm (S) | Cerium Total | MS61 | ALS Minerals ME-MS61 | ALS Minerals ME-MS61. A prepared sample is digested with HNO3; HF; and HClO4 in a graphite heating block to near dryness. After cooling; the sample is boiled in diluted HCl. Bulk to volume with deionized water; mixed and analyzed by ICP-AES and ICP-MS for all elements. | 0.01 | Y | Y | Y | 0 |  |  |  | 2024-01-29 |  | EMS |  | DimensionlessRatio | ppm | N | (ratio) |  | Cerium Total (ratio) | 7440-45-1 | Metals - Total; | Metals - Total; EMS code: CE-T | Chemical | Total | LAB
CE-T | mg/L | Cerium Total | X555 | Metals in water, total, "in-bottle" digestion, ICP-SFMS | The whole water sample collected in a HDPE/Teflon bottle, is de-capped in a Class 100 workstation, acidified to 2% (v/v) HNO3 and 1% (v/v) HCl and recapped. The sample is "in-bottle" digested in a convection oven, set at 60 degrees C for 16 hours. The sample is cooled and introduced to an inductively coupled argon plasma-sector field mass spectrometer (ICP-SFMS) by pneumatic nebulization. Each respective element is measured at a specific mass to charge ratio, m/z value expressed in amu (atomic mass units), and at resolution. LR: Low Resolution ~300, MR: Mid Resolution ~4000 and HR: High a specified instrument Resolution ~10000. Ref: Method number 02-2003, National Laboratory for Environmental Testing, CCIW, Environment Canada, Burlington. Requested April 2003. | 2e-06 | Y | Y | Y | 15777 | 6/27/2004 | 12/6/2023 2:08:00 AM | 2/14/2024 9:05:32 AM | 2009-03-31 |  | JCHUDAK |  | Concentration | mg/L | Y | (fl. conc.) |  | Cerium Total (fl. conc.) | 7440-45-1 | Metals - Total; | Metals - Total; EMS code: CE-T | Chemical | Total | LAB
CE-T | ppm (S) | Cerium Total | MS81 | ALS Minerals: ME-MS81 | ALS Minerals: ME-MS81. A prepared sample (0.100 g) is added to lithium metaborate/lithium tetraborate flux; mixed well and fused in a furnace at 1025°C. The resulting melt is then cooled and dissolved in an acid mixture containing nitric; hydrochloric and hydrofluoric acids. This solution is then analyzed by inductively coupled plasma - mass spectrometry. | 0.1 | Y | Y | Y | 0 |  |  |  | 2024-01-29 |  | EMS |  | DimensionlessRatio | ppm | N | (ratio) |  | Cerium Total (ratio) | 7440-45-1 | Metals - Total; | Metals - Total; EMS code: CE-T | Chemical | Total | LAB
CE-T | ppm (S) | Cerium Total | 6041 | Aqua Regia by ICPMS | These were sediment/soils digested with Aqua-Regia (Nitric Acid & Hydrochloric Acids) and then analysis by ICPMS. Results would be reported in mg/kg and/or %. | 0.003 | Y | Y | Y | 21 | 10/13/2019 | 8/5/2020 | 8/10/2020 3:52:06 PM | 2019-09-06 | 10/21/2019 11:54:37 AM | KATFRASE | KATFRASE | DimensionlessRatio | ppm | N | (ratio) |  | Cerium Total (ratio) | 7440-45-1 | Metals - Total; | Metals - Total; EMS code: CE-T | Chemical | Total | LAB
CGHT | m | Corrected Gauge Height | XM19 | Contin.-Electr.Sensors | Contin.-Electr.Sensors | 0 | Y | Y | Y | 161 | 1/6/2021 | 12/31/2023 | 3/18/2024 5:01:07 PM | 2000-01-20 |  | NPEPPIN |  | Length | m | Y | (len.) |  | Corrected Gauge Height (len.) |  |  | EMS code: CGHT | Chemical |  | LAB
CH4- | ppm W | Methane | MEED | Methane, Ethane and Ethene, Dissolved via GC-FID | Water samples are collected in 44 mL P&T vials without headspace and preserved with Sodium Bisulfate. Headspace is then generated using Nitrogen. The sample is then shaken to allow gas equilibration. A portion of the headspace vapour is withdrawn and injected into a gas chromatograph equipped with capillary columns and a flame ionization detector (FID). The concentration of the gas in the liquid is proportional to the partial pressure of the gas above the liquid, and is calculated using Henrys Law. Ref: EPA REGION 1, NATATTEN.WPD, REV. 1 | 4 | Y | Y | Y | 3 | 11/4/2021 | 11/4/2021 | 10/5/2022 3:23:07 PM | 2022-05-30 |  | EMS |  | DimensionlessRatio | ppm | N | (ratio) |  | Methane (ratio) | 74-82-8 | Petroleum Hydrocarbons; | Petroleum Hydrocarbons; EMS code: CH4- | Chemical |  | LAB
CH4- | mg/L | Methane | XCAL | Calculation (mg/L) | Calculation (mg/L) |  | Y | Y | Y | 3 | 11/8/2021 | 11/8/2021 | 10/5/2022 3:23:07 PM | 2022-06-01 |  | EMS |  | Concentration | mg/L | Y | (fl. conc.) |  | Methane (fl. conc.) | 74-82-8 | Petroleum Hydrocarbons; | Petroleum Hydrocarbons; EMS code: CH4- | Chemical |  | LAB
CH4- | % LEL | Methane | XM10 | Meter | Meter (April 9/97 - changed description from 'Flow Meter' to 'Meter'; this method code is associated with parameter 0011 - Specific Conductance therefore 'Meter' is more appropriate. Note that X330 is also called 'Meter' and is the preferred method code.) | 1 | Y | Y | Y | 259 |  |  | 6/9/2006 1:36:45 PM | 2002-09-24 |  | EMS |  | DimensionlessRatio | % | N | (ratio) | (LEL) | Methane (LEL) (ratio) | 74-82-8 | Petroleum Hydrocarbons; | Petroleum Hydrocarbons; EMS code: CH4- | Chemical |  | LAB
CHCL3 | mg/L | Total Trichloromethane | THMHAA | THM/HAA CALC | THM/HAA CALC | 0.01 | Y | Y | Y | 2 | 3/24/2023 10:24:00 PM | 5/31/2023 12:38:00 AM | 6/8/2023 12:14:57 PM | 2023-03-31 |  | EMS |  | Concentration | mg/L | Y | (fl. conc.) |  | Total Trichloromethane (fl. conc.) |  | Major Ions; | Major Ions; EMS code: CHCL3 | Chemical | Total | LAB
CHL-E | mg/L | Chloride Extractable | E300 | EPA 300.1/CSSS 15 | See EPA method 300.1 and Canadian Society of Soil Sciences (2008) Chapter 15 | 20 | Y | Y | Y | 16 | 1/11/2022 | 2/7/2023 | 1/11/2024 1:07:07 PM | 2022-06-29 |  | EMS |  | Concentration | mg/L | Y | (fl. conc.) |  | Chloride Extractable (fl. conc.) | 16887-00-6 | Major Ions; | Major Ions; EMS code: CHL-E | Chemical | Extractable | LAB
CHL-E | ug/g | Chloride Extractable | CLPST | Chloride in soil (paste) by IC | A soil extract produced by the saturated paste extraction procedure is analyzed for chloride by Ion Chromatography with conductivity detection. | 1 | Y | Y | Y | 16 | 1/13/2022 | 2/7/2023 | 1/11/2024 1:07:07 PM | 2022-07-07 |  | EMS |  | MassPerMass | mg/kg | Y | (mass conc.) |  | Chloride Extractable (mass conc.) | 16887-00-6 | Major Ions; | Major Ions; EMS code: CHL-E | Chemical | Extractable | LAB
CHL-S | mg/L | Chloride; soluble ion content | E6010B | EPA 6010B/CSSS 15 | See CSSS CH15/EPA 6010B/AER D50. INDUCTIVELY COUPLED PLASMA-ATOMIC EMISSION SPECTROMETRY See EPA 6010B and CSSS Ch. 15. https://www.epa.gov/sites/default/files/documents/6010b.pdf | 5 | Y | Y | Y | 0 |  |  |  | 2024-01-30 |  | EMS |  | Concentration | mg/L | Y | (fl. conc.) |  | Chloride; soluble ion content (fl. conc.) |  | Major Ions; | Major Ions; EMS code: CHL-S | Chemical | Soluble | LAB
CHL-S | mg/L | Chloride; soluble ion content | E300 | EPA 300.1/CSSS 15 | See EPA method 300.1 and Canadian Society of Soil Sciences (2008) Chapter 15 | 20 | Y | Y | Y | 0 |  |  |  | 2024-02-02 |  | EMS |  | Concentration | mg/L | Y | (fl. conc.) |  | Chloride; soluble ion content (fl. conc.) |  | Major Ions; | Major Ions; EMS code: CHL-S | Chemical | Soluble | LAB
CHL-S | ug/g | Chloride; soluble ion content | CLPST | Chloride in soil (paste) by IC | A soil extract produced by the saturated paste extraction procedure is analyzed for chloride by Ion Chromatography with conductivity detection. | 1 | Y | Y | Y | 0 |  |  |  | 2024-02-02 |  | EMS |  | MassPerMass | mg/kg | Y | (mass conc.) |  | Chloride; soluble ion content (mass conc.) |  | Major Ions; | Major Ions; EMS code: CHL-S | Chemical | Soluble | LAB
CHRY | mg/L | Chrysene d12 | MW02 | DCM extrn,GC/MS-SIM (PAH) | Water samples are liquid - liquid extracted with dichloromethane. The extract is then cleaned by silica gel column chromatography and analyzed by gas chromatography - mass spectroscopy. | 1e-05 | Y | Y | Y | 72 | 2/5/2002 | 8/16/2004 | 9/21/2004 1:03:54 AM | 1996-04-12 | 6/3/1998 6:51:41 PM | EMS |  | Concentration | mg/L | Y | (fl. conc.) |  | Chrysene d12 (fl. conc.) | 1719-03-5 | Polycyclic Aromatic Hydrocarbons (PAHs); | Polycyclic Aromatic Hydrocarbons (PAHs); EMS code: CHRY | Chemical |  | SURROGATE_RESULT
CHRY | % | Chrysene d12 | PAH1 | PUF:GC/MS | To understand the extent of human exposure to B[a]P and other PAHs, a reliable sampling and analytical method has been established which involves a combination quartz fibre filter/absorbent cartridge with subsequent extraction and analysis by gas chromatography (GC) with mass spectrometry (MS) detection (GC/MS). The analytical methods are a modification of EPA Test Method 625, Methods for Organic Chemical Analysis of Municipal and Industrial Wastewater, and Methods 8270, Test Methods for Evaluation of Solid Waste. PAH1 utilizes PUF as an adsorbant; PAH2 utilizes XAD-2 as an adsorbant. Note that isotopic surrogates may be added and are reported as % recovery. | 0 | Y | Y | Y | 100 | 10/14/1997 2:50:00 PM | 4/22/1998 5:40:00 PM | 1/19/1999 6:45:09 PM | 1996-04-12 | 6/3/1998 6:51:34 PM | EMS |  | DimensionlessRatio | % | N | (ratio) |  | Chrysene d12 (ratio) | 1719-03-5 | Polycyclic Aromatic Hydrocarbons (PAHs); | Polycyclic Aromatic Hydrocarbons (PAHs); EMS code: CHRY | Chemical |  | SURROGATE_RESULT
CHRY | % | Chrysene d12 | FS87 | DCM,Ac/Hex extr,PAM,HRGC/LRMS | DCM;Ac/Hex extr;PAM;HRGC/LRMS | 1 | Y | Y | Y | 43 | 5/20/2015 | 5/23/2015 | 7/3/2015 3:19:04 PM | 2015-05-27 |  | RDGIBSON |  | DimensionlessRatio | % | N | (ratio) |  | Chrysene d12 (ratio) | 1719-03-5 | Polycyclic Aromatic Hydrocarbons (PAHs); | Polycyclic Aromatic Hydrocarbons (PAHs); EMS code: CHRY | Chemical |  | SURROGATE_RESULT
CHRY | % | Chrysene d12 | X262 | CH2Cl2 Ext:Silica:GC/MS | The sample is placed in a separatory funnel, the pH of the sample is adjusted to alkaline and then acidic conditions, and extracted each time with dichloromethane to selectively remove compounds. The final extracts (basic and acidic) can be combined or run separately. The extract is cleaned by silica gel column chromatography, then analyzed by GC/MS. | 1 | Y | Y | Y | 1 | 8/28/2014 | 8/28/2014 | 3/11/2015 3:57:18 PM | 2015-03-11 |  | RDGIBSON |  | DimensionlessRatio | % | N | (ratio) |  | Chrysene d12 (ratio) | 1719-03-5 | Polycyclic Aromatic Hydrocarbons (PAHs); | Polycyclic Aromatic Hydrocarbons (PAHs); EMS code: CHRY | Chemical |  | SURROGATE_RESULT
CHRY | % | Chrysene d12 | SURR | Surrogate Recovery | Surrogate Recovery | 30 | Y | Y | Y | 1472 | 8/18/2003 | 8/15/2024 | 8/16/2024 1:36:03 PM | 1996-04-12 | 6/3/1998 6:51:41 PM | EMS |  | DimensionlessRatio | % | N | (ratio) |  | Chrysene d12 (ratio) | 1719-03-5 | Polycyclic Aromatic Hydrocarbons (PAHs); | Polycyclic Aromatic Hydrocarbons (PAHs); EMS code: CHRY | Chemical |  | SURROGATE_RESULT
CHRY | ug/g | Chrysene d12 | MP02 | GC/MS | GC/MS | 0.001 | Y | Y | Y | 101 | 3/28/2002 | 3/28/2011 | 5/9/2011 5:20:46 PM | 1996-04-12 | 6/3/1998 6:51:34 PM | EMS |  | MassPerMass | mg/kg | Y | (mass conc.) |  | Chrysene d12 (mass conc.) | 1719-03-5 | Polycyclic Aromatic Hydrocarbons (PAHs); | Polycyclic Aromatic Hydrocarbons (PAHs); EMS code: CHRY | Chemical |  | SURROGATE_RESULT
CIMT | mg/L | Cimetidine CAS: 51481-61-9 | XMSP | HPLC and Tri. Quadpole MS/MS | EPA 1694 An aliquot of filtered sample containing 5% organic is injected directly. Each compound is separated by reversed phase HPLC. Analyzed by Electro Spray Ionization MS/MS detection using Triple Quadrupole MS/MS detector. Quantiitation with internal or external standard. |  | Y | Y | Y | 12 | 5/8/2020 | 6/4/2024 | 8/21/2024 3:32:03 PM | 2022-05-04 |  | EMS |  | Concentration | mg/L | Y | (fl. conc.) |  | Cimetidine (fl. conc.) | 51481-61-9 | Pesticides; | Pesticides; EMS code: CIMT | Chemical |  | LAB
CL-L | mg/L | Chloride Leachable | MEND | MEND (mod) EPA 300.1 (mod) | MEND (mod) EPA 300.1 (mod). Inorganic anions are analyzed by obtaining an extract from a soil sample that has been added in a 3:1 ratio of deionized water to soil; then leached and the extract filtered prior to analysis; which is then analyzed by Ion Chromatography with conductivity and/or UV detection. | 0.5 | Y | Y | Y | 0 |  |  |  | 2024-01-24 |  | RDGIBSON |  | Concentration | mg/L | Y | (fl. conc.) |  | Chloride Leachable (fl. conc.) |  | Major Ions; | Major Ions; EMS code: CL-L | Chemical | Leachable | LAB
CL-S | ug/m3 | Chloride-Soluble | EQ40 | EQUIS (mg/m3) | EQUIS (mg/m3) | 0 | N | Y | Y | 2382 |  |  | 10/22/1996 10:21:01 PM | 1996-04-12 | 6/3/1998 6:51:34 PM | EMS |  | AirConcentration | µg/m³ | Y | (air conc.) |  | Chloride-Soluble (air conc.) | 16887-00-6 | Major Ions; | Major Ions; EMS code: CL-S | Chemical | Soluble | LAB
CL-S | mg/dm2/d | Chloride-Soluble | EQ15 | EQUIS (mg/dm2/d) | EQUIS (mg/dm2/d) | 0.002 | N | Y | Y | 526 |  |  | 10/22/1996 10:21:00 PM | 1996-04-12 | 6/3/1998 6:51:41 PM | EMS |  | MassFlux | mg/dm²/d | Y | (mass flux) |  | Chloride-Soluble (mass flux) | 16887-00-6 | Major Ions; | Major Ions; EMS code: CL-S | Chemical | Soluble | LAB
CL-S | mg/L | Chloride-Soluble | 5068 | Ion Chromatography-Anion | (a) For the Soluble Particulates Pkg., a portion of filtrate from the Particulate-Total procedure made up to 500 mL is analyzed by Ion Chromatography. Intermediate results are in mg/L. Loading results are then calculated with units of mg/dm2/d. (b) For precipitation samples, samples and standards are spiked with concentrated carbonate/bicarbonate solution to give the sample the same background as the eluent used (avoiding a "water dip" effect). The anions of interest are separated through an anion "guard" and anion "separator" column. An anion micro membrane suppressor following the separator columns is used to reduce the background eluent conductivity by converting the carbonate and bicarbonate species to carbonic acid, while enhancing the conductivity of the ions of interest by converting these ions to their corresponding acids. A conductivity detector senses the same species in direct proportion to their initial concentration. (c) For HiVol samples, discs are cut from the HiVol filter, extracted with DI water, and anions analyzed by ion chromatograph, Preliminary results are reported in mg/L and intermediate HiVol loading results are reported as ug/m3. (*) Note that the air-convention is to report on basis of ion weight not element, e.g. as NO3 for nitrate, and as PO4 for phosphate. | 0.01 | Y | Y | Y | 6600 | 4/30/1997 3:50:00 PM | 4/11/2008 | 5/28/2008 2:08:44 AM | 1996-04-12 | 6/3/1998 6:51:34 PM | EMS |  | Concentration | mg/L | Y | (fl. conc.) |  | Chloride-Soluble (fl. conc.) | 16887-00-6 | Major Ions; | Major Ions; EMS code: CL-S | Chemical | Soluble | LAB
CL-S | mg/dm2/d | Chloride-Soluble | 5049 | Soluble Particulates Pkg. Ion Chrom. Anions | For the Soluble Particulates Pkg., a portion of filtrate from the Particulate-Total procedure made up to 500 mL is analyzed by Ion Chromatography. Intermediate results are in mg/L. Loading results are then calculated with units of mg/dm2/d. Note that the air-convention is to report on basis of ion weight not element, e.g. as NO3 for nitrate, and as PO4 for phosphate. | 0.1 | Y | Y | Y | 136 |  |  | 4/1/2008 1:24:44 PM | 1996-04-12 | 6/3/1998 6:51:34 PM | EMS |  | MassFlux | mg/dm²/d | Y | (mass flux) |  | Chloride-Soluble (mass flux) | 16887-00-6 | Major Ions; | Major Ions; EMS code: CL-S | Chemical | Soluble | LAB
CL-S | mg/L | Chloride-Soluble | EQ01 | EQUIS (mg/L) | EQUIS (mg/L) | 0.01 | N | Y | Y | 130 |  |  | 10/22/1996 11:12:48 PM | 1996-04-12 | 6/3/1998 6:51:34 PM | EMS |  | Concentration | mg/L | Y | (fl. conc.) |  | Chloride-Soluble (fl. conc.) | 16887-00-6 | Major Ions; | Major Ions; EMS code: CL-S | Chemical | Soluble | LAB
CL-S | ug/m3 | Chloride-Soluble | 5022 | Water Extr,Ion Chr-Anion | Discs cut from the HiVol filter are extracted with deionized water and the resulting anions are separated an measured using an ion chromatograph. (See Anions-Ion Chromatography-Precipitation, methods NO3-5068, S04-5068, CL-S5068). Values are then converted to units of ug/m3. | 0.5 | Y | Y | Y | 533 |  |  | 10/24/1996 6:07:34 AM | 1996-04-12 | 6/3/1998 6:51:34 PM | EMS |  | AirConcentration | µg/m³ | Y | (air conc.) |  | Chloride-Soluble (air conc.) | 16887-00-6 | Major Ions; | Major Ions; EMS code: CL-S | Chemical | Soluble | LAB
CL2- | mg/L | Chlorine | X257 | Colorimetric | Colorimetric. General use for several parameters. E.g.Total Sulfide is determined by spectrophotometer using the methlyene blue colourimetric method. The H2S calculation converts Total Sulphide as (S2-) and reports it as Total Sulphide as (H2S). | 0.1 | Y | Y | Y | 8312 | 4/1/1999 8:00:00 AM | 6/30/2022 8:00:00 AM | 8/2/2022 1:10:53 PM | 1996-04-12 | 6/3/1998 6:51:34 PM | EMS |  | Concentration | mg/L | Y | (fl. conc.) |  | Chlorine (fl. conc.) | 7782-50-5 | Inorganics; | Inorganics; EMS code: CL2- | Chemical |  | LAB
CL2- | ug/m3 | Chlorine | X300 | Iodine Titration-Stack | Iodine Titration-Stack | 0.001 | Y | Y | Y | 264 | 2/18/2014 8:00:00 AM | 6/29/2024 8:00:00 AM | 7/30/2024 10:08:09 AM | 1996-04-12 | 6/3/1998 6:51:42 PM | EMS |  | AirConcentration | µg/m³ | Y | (air conc.) |  | Chlorine (air conc.) | 7782-50-5 | Inorganics; | Inorganics; EMS code: CL2- | Chemical |  | LAB
CL2F | mg/L | Total Chloramines (as Cl2) | XCAL | Calculation (mg/L) | Calculation (mg/L) | 0.2 | Y | Y | Y | 10 | 4/15/2015 | 4/24/2015 | 2/10/2016 3:32:39 PM | 2016-02-10 |  | RDGIBSON |  | Concentration | mg/L | Y | (fl. conc.) |  | Total Chloramines (as Cl2) (fl. conc.) | 10599-90-3, 10025-85-1 | Organics; | Organics; EMS code: CL2F | Chemical | Total | LAB
CLAS | % | Classification | GS01 | Gravimetric, balance, Separation by sieve >.075 mm | Gravimetric, balance; Separation by sieve >.075 mm. These are particle sizing determinations performed on soil. These two parameters reported will need the > and < associated with them because the way they have been reported is not based on a number of distinct and specific sieve sizes but instead more general size of everything that is either larger or greater than a size of 0.075mm. |  | Y | Y | Y | 7 | 10/29/2022 | 10/29/2022 | 1/7/2023 3:12:11 PM | 2022-12-21 |  | EMS |  | DimensionlessRatio | % | N | (ratio) |  | Classification (ratio) |  |  | EMS code: CLAS | Chemical |  | LAB
CLAY | % | Clay Content | HYDR | Hydrometer Analysis | Hydrometer Analysis | 2 | Y | Y | Y | 144 | 10/4/2013 3:54:00 PM | 11/1/2022 | 1/11/2024 1:07:07 PM | 2010-11-16 |  | RDGIBSON |  | DimensionlessRatio | % | N | (ratio) |  | Clay Content (ratio) |  | Sediment Size; | Sediment Size; EMS code: CLAY | Chemical |  | LAB
CLFF | mg/L | Free Chlorine, Final Concentration | 5710B | SM 5710 B (2021) | SM 5710 B (2021); Chlorination, 7-day Incubation | 0.1 | Y | Y | Y | 2 | 3/18/2023 4:37:00 PM | 5/23/2023 12:06:00 PM | 6/8/2023 12:14:57 PM | 2023-03-31 |  | EMS |  | Concentration | mg/L | Y | (fl. conc.) |  | Free Chlorine, Final Concentration (fl. conc.) | 7782-50-5 |  | EMS code: CLFF | Chemical |  | LAB
CLFI | mg/L | Free Chlorine, Initial Dose | 5710B | SM 5710 B (2021) | SM 5710 B (2021); Chlorination, 7-day Incubation | 0.1 | Y | Y | Y | 2 | 3/18/2023 4:37:00 PM | 5/23/2023 12:06:00 PM | 6/8/2023 12:14:57 PM | 2023-03-31 |  | EMS |  | Concentration | mg/L | Y | (fl. conc.) |  | Free Chlorine, Initial Dose (fl. conc.) | 7782-50-5 |  | EMS code: CLFI | Chemical |  | LAB
CLO2 | t/d | Chlorine Dioxide | X456 | Iodine Titration-Stack | Iodine Titration-Stack | 0 | Y | Y | Y | 390 |  |  | 8/22/2024 1:49:14 PM | 1996-04-12 | 6/3/1998 6:51:42 PM | EMS |  | MassFlow | t/d | Y | (mass flow) |  | Chlorine Dioxide (mass flow) | 10049-04-4 | Inorganics; | Inorganics; EMS code: CLO2 | Chemical |  | LAB
CLO2 | ppm A | Chlorine Dioxide | X300 | Iodine Titration-Stack | Iodine Titration-Stack | 0.001 | Y | Y | Y | 492 |  |  | 8/22/2024 1:49:14 PM | 1996-04-12 | 8/27/2020 11:15:32 AM | EMS | RDGIBSON | DimensionlessRatio | ppm | N | (ratio) |  | Chlorine Dioxide (ratio) | 10049-04-4 | Inorganics; | Inorganics; EMS code: CLO2 | Chemical |  | LAB
CLO3 | mg/L | Chlorate Dissolved | X044 | Ion Chromatography | A clear (Total) or filtered (Dissolved) water sample is injected into a stream of carbonate-bicarbonate eluent and passed through a series of ion exchanger. The anions of interested are separated on the basis of their relative affinities for a low capacity , strongly basic anion exchanger (guard and separate columns). The separated anions are directed through the anion self-regenerating suppressor (ASRS). Hydronium ions in the suppressor will neutralize the eluent to its weakly ionized form and convert the separated anions in their conductive acid form. The separated anions are measured by conductivity. They are identified on the basis of retention time as compared to standards. Quantitation is by measurement of peak area of peak height against a calibration curve generated from known standards. | 0.005 | Y | Y | Y | 33 | 12/19/1996 11:54:00 AM | 8/16/2009 | 1/13/2010 4:24:21 PM | 1996-04-12 | 6/3/1998 6:51:45 PM | EMS |  | Concentration | mg/L | Y | (fl. conc.) |  | Chlorate Dissolved (fl. conc.) | 14866-68-3 | Inorganics; | Inorganics; EMS code: CLO3 | Chemical | Dissolved | LAB
CLOR | mg/L | Clotrimazole CAS: 23593-75-1 | XMSP | HPLC and Tri. Quadpole MS/MS | EPA 1694 An aliquot of filtered sample containing 5% organic is injected directly. Each compound is separated by reversed phase HPLC. Analyzed by Electro Spray Ionization MS/MS detection using Triple Quadrupole MS/MS detector. Quantiitation with internal or external standard. |  | Y | Y | Y | 12 | 5/8/2020 | 6/4/2024 | 8/21/2024 3:32:03 PM | 2022-05-04 |  | EMS |  | Concentration | mg/L | Y | (fl. conc.) |  | Clotrimazole (fl. conc.) | 23593-75-1 | Pesticides; | Pesticides; EMS code: CLOR | Chemical |  | LAB
CLOT | mg/L | Clothianidin (Insecticide/Pesticide) | E751 | E751D Pesticides in Water by LC-MS/MS (POS1 List) | E751D Pesticides in Water by LC-MS/MS (POS1 List) (MOECC E3501) | 5e-05 | Y | Y | Y | 39 | 12/9/2021 | 12/31/2021 | 2/15/2022 9:50:05 AM | 2021-12-14 |  | EMS |  | Concentration | mg/L | Y | (fl. conc.) |  | Clothianidin (Insecticide/Pesticide) (fl. conc.) | 210880-92-5 | Pesticides; | Pesticides; EMS code: CLOT | Chemical |  | LAB
CLOT | mg/L | Clothianidin (Insecticide/Pesticide) | 3544 | MOECC E3544 NEONIC-DI-LCMS-WT | MOECC E3544 NEONIC-DI-LCMS-WT Water sample is filtered; spiked with internal standard and analyzed by direct injection using a HPLC Triple Quadrupole MS/MS with Electro Spray Ionization. | 5e-06 | Y | Y | Y | 29 | 6/5/2020 | 7/27/2020 | 7/21/2021 6:01:34 PM | 2020-06-30 |  | RDGIBSON |  | Concentration | mg/L | Y | (fl. conc.) |  | Clothianidin (Insecticide/Pesticide) (fl. conc.) | 210880-92-5 | Pesticides; | Pesticides; EMS code: CLOT | Chemical |  | LAB
CLTF | mg/L | Chlorine, total, field | FLD | Field sampling | Field sampling method | 0.01 | Y | Y | Y | 16 | 3/6/2023 | 8/14/2024 | 8/17/2024 8:30:10 AM | 2023-04-19 |  | EMS |  | Concentration | mg/L | Y | (fl. conc.) |  | Chlorine, total, field (fl. conc.) | 7782-50-5 | Major Ions; | Major Ions; EMS code: CLTF | Chemical | Total | FIELD_RESULT | Field sampling
CN-F | mg/L | Cyanide; Free | FCFI | Free Cyanide with FIA,GDS,AD | This analysis is carried out using procedures adapted from ASTM Method 7237 "Free Cyanide with Flow Injection Analysis (FIA) Utilizing Gas Diffusion Separation and Amperometric Detection". Free cyanide is determined by in-line gas diffusion at pH 6 with final determination by colourimetric analysis. | 0.005 | Y | Y | Y | 1452 | 5/9/2013 | 7/29/2024 | 8/12/2024 11:52:10 AM | 2016-04-21 |  | RDGIBSON |  | Concentration | mg/L | Y | (fl. conc.) |  | Cyanide; Free (fl. conc.) | 57-12-5 | Inorganics; | Inorganics; EMS code: CN-F | Chemical |  | LAB
CN-T | mg/L | Cyanide Total | X115 | NaOH,Dist,Col'metric | NaOH;Dist;Col'metric | 0.005 | Y | Y | Y | 2256 |  |  | 10/30/2014 8:31:17 AM | 1996-04-12 | 6/3/1998 6:51:34 PM | EMS |  | Concentration | mg/L | Y | (fl. conc.) |  | Cyanide Total (fl. conc.) | 57-12-5 | Inorganics; | Inorganics; EMS code: CN-T | Chemical | Total | LAB
CN-T | mg/L | Cyanide Total | X153 | UV:Auto Pyr-Barb Acid | UV:Auto Pyr-Barb Acid | 0.0005 | Y | Y | Y | 1615 | 10/17/2000 | 10/17/2000 | 2/11/2003 3:37:01 PM | 1996-04-12 | 6/3/1998 6:51:34 PM | EMS |  | Concentration | mg/L | Y | (fl. conc.) |  | Cyanide Total (fl. conc.) | 57-12-5 | Inorganics; | Inorganics; EMS code: CN-T | Chemical | Total | LAB
CN-T | ug/m3 | Cyanide Total | CAL6 | Calculation (mg/m3) | Calculation (mg/m3) | 1 | N | Y | Y | 32 |  |  | 10/24/1996 12:16:49 PM | 1997-01-16 | 7/9/1999 11:16:20 AM | MPAK | NPEPPIN | AirConcentration | µg/m³ | Y | (air conc.) |  | Cyanide Total (air conc.) | 57-12-5 | Inorganics; | Inorganics; EMS code: CN-T | Chemical | Total | LAB
CN-T | mg/L | Cyanide Total | TCFA | FIA-UV digestion, distil., isonicotinic colmetric | This analysis is carried out using procedures adapted from ISO Method 14403:2002 "Determination of Total Cyanide using Flow Analysis (FIA and CFA)". Total or strong acid dissociable (SAD) cyanide is determined by in-line UV digestion along with sample distillation and final determination by colourimetric analysis. Method Limitation: This method is susceptible to interference from thiocyanate (SCN). If SCN is present in the sample, there could be a positive interference with this method, but it would be less than 1% and could be as low as zero. | 0.005 | Y | Y | Y | 6661 | 6/12/2012 | 8/22/2024 | 8/26/2024 3:51:07 PM | 2012-07-12 |  | RDGIBSON |  | Concentration | mg/L | Y | (fl. conc.) |  | Cyanide Total (fl. conc.) | 57-12-5 | Inorganics; | Inorganics; EMS code: CN-T | Chemical | Total | LAB
CN-T | mg/L | Cyanide Total | X254 | EPA:H2SO4Dist:NaCl:Col'm | EPA:H2SO4Dist:NaCl:Col'm | 0.005 | Y | Y | Y | 820 | 1/12/2007 | 5/7/2012 | 2/15/2018 11:14:40 AM | 1996-04-12 | 6/3/1998 6:51:45 PM | EMS |  | Concentration | mg/L | Y | (fl. conc.) |  | Cyanide Total (fl. conc.) | 57-12-5 | Inorganics; | Inorganics; EMS code: CN-T | Chemical | Total | LAB
CNBBTF | ug/g | 4-Chloro-3-Nitrobenzotrifluoride | ZN03 | DCM+Ac.Extr:GC/MS | DCM+Ac.Extr:GC/MS EPA 8270E (mod) | 0.005 | Y | Y | Y | 2 | 11/2/2022 12:53:00 AM | 11/2/2022 4:41:00 AM | 1/7/2023 3:12:10 PM | 2022-12-29 |  | EMS |  | MassPerMass | mg/kg | Y | (mass conc.) |  | 4-Chloro-3-Nitrobenzotrifluoride (mass conc.) | 121-17-5 | Organics; | Organics; EMS code: CNBBTF | Chemical |  | LAB
CNR | None | C:N ratio | EC | Environment Canada | Environment Canada | 0.01 | Y | Y | Y | 1 | 2/9/2023 | 2/9/2023 | 2/22/2023 12:19:06 PM | 2023-02-14 |  | EMS |  | None | None | N | (unitless) |  | C:N ratio (unitless) |  |  | EMS code: CNR | Chemical |  | LAB
CO-- | ppm A | Carbon Monoxide | FLD | Field sampling | Field sampling method |  | Y | Y | Y | 210 | 5/8/2023 10:00:00 AM | 6/25/2024 10:00:00 AM | 7/4/2024 12:26:51 PM | 2023-08-25 |  | EMS |  | DimensionlessRatio | ppm | N | (ratio) |  | Carbon Monoxide (ratio) |  |  | EMS code: CO-- | Chemical |  | FIELD_RESULT | Field sampling
CO-D | mg/L | Cobalt Dissolved | F068 | HNO3dig,TeflonTube/Autoclave,ICP-MS | This method is applicable to determiniation of dissolved, extractable, and total recoverable metals in ground waters, surface waters, drinking waters, and wastewaters. Dissolved, extractable, and total recoverable sample fractions are acidified to pH <1.5 with analytical grade nitric acid (HNO3) to prevent the metals from adsorbing to the container material or precipitating. Extractable and total recoverable water samples are collected in 250 ml acid washed polyethylene bottles and are preserved with Nitric acid (0.5 ml acid for 100 mLs of sample). All preserved samples are stored at room temperature and can be held for up to 6 months. Unfiltered, un-preserved samples are stored at 4C for no more than 30 days. For dissolved and extractable water fractions, no sample preparation is necessary. Samples are simply poured in autosampler tubes (i.e. 15 ml centrifuge tubes), and placed on the autosampler and analyzed. Samples for total recoverable metals are digested with Nitric acid in special Teflon tubes in an autoclave prior to analyses. | 5e-06 | Y | Y | Y | 1398 | 6/1/1999 1:20:00 PM | 4/17/2015 1:06:00 PM | 8/11/2021 12:04:11 PM | 1999-06-28 |  | NPEPPIN |  | Concentration | mg/L | Y | (fl. conc.) |  | Cobalt Dissolved (fl. conc.) | 7440-48-4 | Metals - Dissolved; | Metals - Dissolved; EMS code: CO-D | Chemical | Dissolved | LAB
CO-D | mg/L | Cobalt Dissolved | F039 | Field filt + acid, ICP-AES | Dissolved and extractable water samples are ready for analysis once collected. Total samples are autoclaved with HNO3 and hydrochloric acid (HCI) in capped Teflon vials. Analysis is performed via Inductively Coupled Argon Plasma Atomic Emission Spectrometry (ICP-AES)(simultaneous multielement analysis), Graphite Furnace Atomic Absorption Spectrometry (GF-AAS)(low level analysis) and/or Flame Atomic Absorption Spectrometry (F-AAS). | 0.005 | Y | Y | Y | 5938 | 7/29/1996 6:03:00 PM | 7/5/2019 1:21:00 PM | 9/11/2019 8:47:50 AM | 1997-05-08 | 6/3/1998 6:51:46 PM | MPAK |  | Concentration | mg/L | Y | (fl. conc.) |  | Cobalt Dissolved (fl. conc.) | 7440-48-4 | Metals - Dissolved; | Metals - Dissolved; EMS code: CO-D | Chemical | Dissolved | LAB
CO-D | mg/L | Cobalt Dissolved | F032 | Filt.,Acid,ICPMS | There is no acid digestion required for this analysis. A water sample is collected and preserved in 0.2 percent nitric acid. To analyse the extractable metals the sample is shaken, allowed to sit for at least twenty four hours, and a decanted portion is introduced to the spectrometer. When the dissolved metals are required, the samples must first be filtered through a 0.45 micron cellulose acetate filter paper and then followed by preservation in 0.2 percent nitric acid. It is important to note that different laboratories may have very different detection limits for some variables analyzed by ICPMS. | 0.0003 | Y | Y | Y | 81117 | 3/10/1998 2:06:00 PM | 8/26/2024 | 8/26/2024 4:53:05 PM | 1997-10-20 | 8/15/2019 10:54:45 AM | NPEPPIN | KATFRASE | Concentration | mg/L | Y | (fl. conc.) |  | Cobalt Dissolved (fl. conc.) | 7440-48-4 | Metals - Dissolved; | Metals - Dissolved; EMS code: CO-D | Chemical | Dissolved | LAB
CO-D | mg/L | Cobalt Dissolved | X170 | ICP | This code is a catch-all for a large variety of ICP methods including water and also solids digested into solutions as an intermediate result. Different procedures will have greatly different detection limits. (a) For air particulates, preliminary results are reported in mg/L and loading results as mg/dm2/d. (b) For Hi Vol samples preliminiary results are same as reported in mg/L and loading results as ug/m3. (c) Note that for Rain samples there is no filtration nor digestion steps. | 0.1 | Y | Y | Y | 12204 | 4/11/1996 6:01:00 PM | 11/6/2018 | 6/12/2020 1:21:11 PM | 1996-04-12 | 6/3/1998 6:51:34 PM | EMS |  | Concentration | mg/L | Y | (fl. conc.) |  | Cobalt Dissolved (fl. conc.) | 7440-48-4 | Metals - Dissolved; | Metals - Dissolved; EMS code: CO-D | Chemical | Dissolved | LAB
CO-D | mg/L | Cobalt Dissolved | X085 | LF,LA,HNO3,AA | Samples are lab-filtered (0.45 um), and lab-acidified (HNO3); analysis by AA. | 0.005 | Y | Y | Y | 207 | 12/12/1996 | 9/18/2015 | 10/28/2015 9:54:16 AM | 1996-04-12 | 6/3/1998 6:51:41 PM | EMS |  | Concentration | mg/L | Y | (fl. conc.) |  | Cobalt Dissolved (fl. conc.) | 7440-48-4 | Metals - Dissolved; | Metals - Dissolved; EMS code: CO-D | Chemical | Dissolved | LAB
CO-D | mg/L | Cobalt Dissolved | F049 | Filt,HNO3 dig,conc'n,nebul'ICP-OES | This method provides an automated procedure for the Inductively Coupled Argon Plasma Optical Emission Spectrophotometric (ICP-OES) determination and quantitation of trace amounts of metals in surface, ground and precipitation waters after a manual acid digestion and concentration procedure and ultrasonic nebulisation. Organically-bound metals and metal complexes are converted to metal ions by digestion with nitric acid. To increase the signal detected by the ICP, the samples are concentrated 10 times. The liquid sample is pumped onto the piezoelectric transducer, where the ultrasonic action efficiently converts it to a fine, dense aerosol. The nebulizer carrier gas flow then transports the wet aerosol through a heated U-tube, where the solvent is vaporized. Subsequently, most of the solvent vapour is removed from the aerosol stream by passing the stream through a chilled condenser. The result is a dry, analyte-laden aerosol which is introduced to the plasma. There, the high temperature desolvates the aerosol and dissociates the sample material into individual atoms and ions, which are excited to emit light at wave lengths characteristic of the elements present. Dissolved samples are field filtered through 0.45 micron cellulose acetate filter prior to acid preservation. Samples for extractable analysis are acidified with 2 mL concentrated HNO3 per litre and allowed to stand a minimum of 24 hours before analysis. Clients providing these samples are advised to preserve samples at the collection site. | 0.0001 | Y | Y | Y | 24736 |  |  | 5/19/2015 3:10:36 PM | 1997-10-20 | 6/3/1998 6:51:47 PM | NPEPPIN |  | Concentration | mg/L | Y | (fl. conc.) |  | Cobalt Dissolved (fl. conc.) | 7440-48-4 | Metals - Dissolved; | Metals - Dissolved; EMS code: CO-D | Chemical | Dissolved | LAB
CO-D | mg/L | Cobalt Dissolved | X350 | FF,FA:HNO3:ICP | Samples are field-filtered (0.45 um); field-acidified with HNO3; analysis by ICP/AES. | 0.015 | Y | Y | Y | 5754 |  |  | 8/26/2024 2:41:32 PM | 1996-04-12 | 6/3/1998 6:51:41 PM | EMS |  | Concentration | mg/L | Y | (fl. conc.) |  | Cobalt Dissolved (fl. conc.) | 7440-48-4 | Metals - Dissolved; | Metals - Dissolved; EMS code: CO-D | Chemical | Dissolved | LAB
CO-D | mg/L | Cobalt Dissolved | X562 | Dissolved metals: FF, FA, HNO3, HR-ICPMS | Ultra trace metals in water are analyzed by high resolution inductively coupled plasma mass spectrometry (HR-ICPMS), based on US EPA Method 200.8, (Oct. 1999). This procedure is intended for pristine field-filtered acid-preserved water samples. The detection limits (LOR) for this test are based on lab instrumental analysis only, not including filtration. ALS-supplied field filtration equipment does not support these LOR. Therefore, because of the high probability of false positives due to filtration, it is strongly recommended that a filtration blank be analyzed to aid in data interpretation. | 5e-06 | Y | Y | Y | 1264 | 6/8/2015 | 10/19/2023 | 11/9/2023 7:41:27 AM | 2015-10-22 |  | RDGIBSON |  | Concentration | mg/L | Y | (fl. conc.) |  | Cobalt Dissolved (fl. conc.) | 7440-48-4 | Metals - Dissolved; | Metals - Dissolved; EMS code: CO-D | Chemical | Dissolved | LAB
CO-D | mg/L | Cobalt Dissolved | METC | MET-D-NP-CCMS-VA | Ultra trace metals in water are analyzed by CRC ICPMS, based on US EPA Method 6020B (July 2014). This procedure is intended for pristine field-filtered acid-preserved water samples. The detection limits (LOR) for this test are based on lab instrumental analysis only, not including filtration. ALS-supplied field filtration equipment does not support these LOR. Therefore, because of the high probability of false positives due to filtration, it is strongly recommended that a filtration blank be analysed to aid in data interpretation. | 5e-06 | Y | Y | Y | 1776 | 4/3/2012 | 8/15/2024 | 8/23/2024 9:50:35 AM | 2019-08-20 |  | KATFRASE |  | Concentration | mg/L | Y | (fl. conc.) |  | Cobalt Dissolved (fl. conc.) | 7440-48-4 | Metals - Dissolved; | Metals - Dissolved; EMS code: CO-D | Chemical | Dissolved | LAB
CO-D | mg/L | Cobalt Dissolved | MWID | Metals in Water by ICPMS (digested) | Metals in Water by ICPMS (digested) Based on EPA 200.8 | 0.0005 | Y | Y | Y | 31670 | 10/26/2018 5:20:00 PM | 10/26/2018 5:20:00 PM | 7/31/2024 1:46:41 PM | 2009-12-24 |  | RDGIBSON |  | Concentration | mg/L | Y | (fl. conc.) |  | Cobalt Dissolved (fl. conc.) | 7440-48-4 | Metals - Dissolved; | Metals - Dissolved; EMS code: CO-D | Chemical | Dissolved | LAB
CO-D | mg/L | Cobalt Dissolved | 0050 | Graphite Furnace AA | Graphite Furnace AA | 0.001 | Y | Y | Y | 103 | 9/29/2000 | 12/1/2010 | 12/9/2010 3:55:26 PM | 1996-04-12 | 6/3/1998 6:51:34 PM | EMS |  | Concentration | mg/L | Y | (fl. conc.) |  | Cobalt Dissolved (fl. conc.) | 7440-48-4 | Metals - Dissolved; | Metals - Dissolved; EMS code: CO-D | Chemical | Dissolved | LAB
CO-D | mg/L | Cobalt Dissolved | MSEA | Metals in Seawater (undigested) by (CRC) ICP-MS | This method describes the multi-elemental determination of trace elements by Collision/Reaction Cell (CRC) ICP-MS. The method measures ions produced by a radio-frequency inductively coupled plasma. Analyte species originating in a liquid are pneumatically nebulized and the resulting aerosol transported by argon gas into the plasma torch. The ions produced are entrained in the plasma gas and extracted, by means of a differentially pumped vacuum interface, into a mass spectrometer. The ions flow through a Collision/Reaction Cell (CRC) where depending on the mode of analysis either the ions passes through the cell unhindered (no gas - classical ICP-MS mode), or collisions/reactions chemistry takes place between the ions and a neutral gas such as hydrogen or helium. Any polyatomic ions which give interference on the analytes are removed. The ions are then sorted according to their mass-to-charge ratios by a quadrupole mass spectrometer. The ions are then sorted according to their mass-to-charge ratios by a quadrupole mass spectrometer. The ions transmitted through the quadrupole are quantified by a channel electron multiplier and the ion information processed by a data handling system. Although most polyatomic interference is removed, any additional interference must be assessed and valid corrections applied or the data flagged to indicate problems. Interference correction must include compensation for background ions contributed by the plasma gas, reagents, and constituents of the sample matrix. Instrument drift as well as suppressions or enhancements of instrument response caused by the sample matrix must be corrected for by the use of internal standards. | 3e-05 | Y | Y | Y | 4 | 4/12/2023 | 4/29/2024 | 5/16/2024 2:00:14 PM | 2009-09-08 |  | RDGIBSON |  | Concentration | mg/L | Y | (fl. conc.) |  | Cobalt Dissolved (fl. conc.) | 7440-48-4 | Metals - Dissolved; | Metals - Dissolved; EMS code: CO-D | Chemical | Dissolved | LAB
CO-D | mg/L | Cobalt Dissolved | X285 | HNO3 Dig:ICP Analysis | HNO3 Dig:ICP Analysis | 0.015 | Y | Y | Y | 740 | 1/30/2002 | 9/16/2003 3:18:00 PM | 3/18/2004 10:28:34 AM | 1996-04-12 | 6/3/1998 6:51:34 PM | EMS |  | Concentration | mg/L | Y | (fl. conc.) |  | Cobalt Dissolved (fl. conc.) | 7440-48-4 | Metals - Dissolved; | Metals - Dissolved; EMS code: CO-D | Chemical | Dissolved | LAB
CO-D | mg/L | Cobalt Dissolved | 0051 | Graphite Furnace AA | Graphite Furnace | 0.001 | Y | Y | Y | 10 | 12/12/1996 9:40:00 AM | 12/12/1996 9:48:00 AM | 1/30/1997 11:20:54 PM | 1996-04-12 | 6/3/1998 6:51:34 PM | EMS |  | Concentration | mg/L | Y | (fl. conc.) |  | Cobalt Dissolved (fl. conc.) | 7440-48-4 | Metals - Dissolved; | Metals - Dissolved; EMS code: CO-D | Chemical | Dissolved | LAB
CO-D | mg/L | Cobalt Dissolved | X556 | Metals in water, dissolved, "in-bottle" digestion, ICP-SFMS | The whole water sample is filtered through a 0.45 micron cellulose acetate filter paper and the filtrate is stored in a HDPE/Teflon bottle. The container is de-capped in a Class 100 workstation, acidified to 2% (v/v) HNO3 and 1% (v/v) HCl and recapped. The sample is "in-bottle" digested in a convection oven, set at 60 degrees C for 16 hours. The sample is cooled and introduced to an inductively coupled argon plasma-sector field mass spectrometer (ICP-SFMS) by pneumatic nebulization. Each respective element is measured at a specific mass to charge ratio, m/z value expressed in amu (atomic mass units), and at a specified intrument resolution. LR: Low Resolution ~300, MR: Mid Resolution ~4000 and HR: High Resolution ~10000. Ref: Method number 02-2003, National Laboratory for Environmental Testing, CCIW, Environment Canada, Burlington, Requested April 2003. | 2e-06 | Y | Y | Y | 4234 | 6/11/2003 | 12/6/2023 2:04:00 AM | 2/14/2024 9:05:32 AM | 2004-01-19 |  | UMULLER |  | Concentration | mg/L | Y | (fl. conc.) |  | Cobalt Dissolved (fl. conc.) | 7440-48-4 | Metals - Dissolved; | Metals - Dissolved; EMS code: CO-D | Chemical | Dissolved | LAB
CO-D | mg/L | Cobalt Dissolved | F023 | HNO3,GFAA | Dissolved and extractable water samples are ready for analysis once collected. Total samples are autoclaved with HNO3 and hydrochloric acid (HCl) in capped Teflon vials. Analysis is performed via Inductively Coupled Argon Plasma Atomic Emission Spectrometry (ICP-AES) (simultaneous multielement analysis), Graphite Furnace Atomic Absorption Spectrometry (GF-AAS) (low level analysis) and/or Flame Atomic Absorption Spectrometry (F-AAS). Note that extractable samples are acidified with 0.5 mL HNO3 per 100 mL of sample at time of sample collection. | 0.001 | Y | Y | Y | 5 | 3/6/1997 4:28:00 PM | 5/25/1998 4:49:00 PM | 3/11/1999 6:01:06 PM | 1997-04-28 | 6/3/1998 6:51:46 PM | MPAK |  | Concentration | mg/L | Y | (fl. conc.) |  | Cobalt Dissolved (fl. conc.) | 7440-48-4 | Metals - Dissolved; | Metals - Dissolved; EMS code: CO-D | Chemical | Dissolved | LAB
CO-D | mg/L | Cobalt Dissolved | EQ01 | EQUIS (mg/L) | EQUIS (mg/L) | 0.005 | N | Y | Y | 927 |  |  | 10/23/1996 12:07:04 AM | 1996-04-12 | 6/3/1998 6:51:34 PM | EMS |  | Concentration | mg/L | Y | (fl. conc.) |  | Cobalt Dissolved (fl. conc.) | 7440-48-4 | Metals - Dissolved; | Metals - Dissolved; EMS code: CO-D | Chemical | Dissolved | LAB
CO-E | mg/L | Cobalt Extractable | X553 | Metals in water, extractable, ICP-SFMS | The whole water sample is collected in a HDPE/Teflon bottle. In a Class 100 workstation it is preserved with concentrated Nitric Acid at concentration of 0.2 percent (v/v). The acid preservative extracts the metals from material in the sample and a decanted portion is introduced directly to an inductively coupled argon plasma-sector field mass spectrometer (ICP-SFMS) by pneumatic nebulization. Each respective element is measured at a specific mass to charge ratio, m/z value expressed in amu (atomic mass units), and at a specified instrument resolution. LR: Low Resolution ~300, MR: Mid Resolution ~4000 and HR: High Resolution ~10000. Ref: Method number 02-2061, National Laboratory for Environmental Testing, CCIW, Environment Canada, Burlington. Requested April 2003. | 0.002 | Y | Y | Y | 2100 | 9/25/2003 | 9/21/2007 | 11/28/2023 2:11:53 PM | 2004-01-19 | 5/15/2009 11:23:23 AM | UMULLER | RDGIBSON | Concentration | mg/L | Y | (fl. conc.) |  | Cobalt Extractable (fl. conc.) | 7440-48-4 | Metals - Extractable; | Metals - Extractable; EMS code: CO-E | Chemical | Extractable | LAB
CO-E | ug/g | Cobalt Extractable | X230 | Hom,HNO3/HCl,ICP | The sample is dried, ground to pass 100 mesh, and digested with a nitric-hydrochloric acid mixture to solubilize the solid matter and to remove the organic material by oxidation and volatilization. The completed digests are submitted to metals instrumentation personnel for metals analysis by Inductively Coupled Atomic Emission Spectrophotometry (ICP-AES), or hydride analysis by atomic absorption. | 0.5 | Y | Y | Y | 70 |  |  | 10/24/1996 2:25:27 PM | 1996-04-12 | 6/3/1998 6:51:34 PM | EMS |  | MassPerMass | mg/kg | Y | (mass conc.) |  | Cobalt Extractable (mass conc.) | 7440-48-4 | Metals - Extractable; | Metals - Extractable; EMS code: CO-E | Chemical | Extractable | LAB
CO-E | mg/L | Cobalt Extractable | F069 | SEM extr,6N HCl,ICP/OES | 6.0 N Hydrochloric acid (HCI) is added to an aliquot of the sediment sample. The extract produced from the addition of the acid is then analyzed for simultaneously extracted metals (SEM) using inductively coupled plasma - optical emission spectrophotometry. The detection limits depend on the metal parameters in question and the instrumentation used. | 0.01 | Y | Y | Y | 1 |  |  | 11/25/2021 1:50:15 PM | 1999-07-09 |  | NPEPPIN |  | Concentration | mg/L | Y | (fl. conc.) |  | Cobalt Extractable (fl. conc.) | 7440-48-4 | Metals - Extractable; | Metals - Extractable; EMS code: CO-E | Chemical | Extractable | LAB
CO-E | mg/L | Cobalt Extractable | F040 | Field Acid, lab filt, ICP-AES | Dissolved and extractable water samples are ready for analysis once collected. Total samples are autoclaved with HNO3 and hydrochloric acid (HCI) in capped Teflon vials. Analysis is performed via Inductively Coupled Argon Plasma Atomic Emission Spectrometry (ICP-AES)(simultaneous multielement analysis), Graphite Furnace Atomic Absorption Spectrometry (GF-AAS)(low level analysis) and/or Flame Atomic Absorption Spectrometry (F-AAS). | 0.005 | Y | Y | Y | 2350 | 8/5/1997 6:06:00 PM | 3/14/2008 | 9/11/2019 8:47:41 AM | 1997-05-08 | 6/3/1998 6:51:46 PM | MPAK |  | Concentration | mg/L | Y | (fl. conc.) |  | Cobalt Extractable (fl. conc.) | 7440-48-4 | Metals - Extractable; | Metals - Extractable; EMS code: CO-E | Chemical | Extractable | LAB
CO-E | mg/L | Cobalt Extractable | F047 | 24hr HNO3 dil'n,ICP-MS | There is no acid digestion required for this analysis. A water sample is collected and preserved in 0.2 percent nitric acid. To analyse the extractable metals the sample is shaken, allowed to sit for at least twenty four hours, and a decanted portion is introduced to the spectrometer. When the dissolved metals are required, the samples must first be filtered through a 0.45 micron cellulose acetate filter paper and then followed by preservation in 0.2 percent nitric acid. It is important to note that different laboratories may have very different detection limits for some variables analyzed by ICP/MS. | 2e-06 | Y | Y | Y | 2676 | 4/16/1998 10:21:00 AM | 6/26/2003 | 9/11/2019 8:46:17 AM | 1997-10-20 | 6/3/1998 6:51:47 PM | NPEPPIN |  | Concentration | mg/L | Y | (fl. conc.) |  | Cobalt Extractable (fl. conc.) | 7440-48-4 | Metals - Extractable; | Metals - Extractable; EMS code: CO-E | Chemical | Extractable | LAB
CO-T | ug/g wet | Cobalt Total | BPSY | MET-WET-BPSY-CCMS-VA | Metals in Animal Tissue and Vegetation (Biota) - Prescriptive. EPA 200.3/6020B This method is conducted following British Columbia Lab Manual method "Metals in Animal Tissue and Vegetation (Biota) - Prescriptive". Tissue samples are digested with nitric and hydrochloric acids; in combination with addition of hydrogen peroxide. Instrumental analysis is by collision cell inductively coupled plasma - mass spectrometry (modified from EPA Method 6020B). | 0.004 | Y | Y | Y | 2 | 11/7/2019 | 11/7/2019 | 6/10/2021 1:57:10 PM | 2021-06-04 |  | EMS |  | MassPerMass | µg/g | Y | (mass conc.) | (wet) | Cobalt Total (wet) (mass conc.) | 7440-48-4 | Metals - Total; | Metals - Total; EMS code: CO-T | Chemical | Total | LAB
CO-T | mg/L | Cobalt Total | X349 | FA:HNO3:ICP | Samples are field-acidified with HNO3; HNO3 digestion; analysis by ICP/AES. | 0.015 | Y | Y | Y | 4901 |  |  | 8/26/2024 2:41:32 PM | 1996-04-12 | 6/3/1998 6:51:41 PM | EMS |  | Concentration | mg/L | Y | (fl. conc.) |  | Cobalt Total (fl. conc.) | 7440-48-4 | Metals - Total; | Metals - Total; EMS code: CO-T | Chemical | Total | LAB
CO-T | ppm (S) | Cobalt Total | 6041 | Aqua Regia by ICPMS | These were sediment/soils digested with Aqua-Regia (Nitric Acid & Hydrochloric Acids) and then analysis by ICPMS. Results would be reported in mg/kg and/or %. | 0.001 | Y | Y | Y | 21 | 10/13/2019 | 8/5/2020 | 8/10/2020 3:52:06 PM | 2019-09-06 | 10/21/2019 11:54:42 AM | KATFRASE | KATFRASE | DimensionlessRatio | ppm | N | (ratio) |  | Cobalt Total (ratio) | 7440-48-4 | Metals - Total; | Metals - Total; EMS code: CO-T | Chemical | Total | LAB
CO-T | mg/L | Cobalt Total | F121 | Acid,ICPMS | Acidified, no digestion, ICPMS | 1e-05 | Y | Y | Y | 8440 | 3/12/2010 7:57:00 PM | 8/15/2024 9:10:00 PM | 8/26/2024 8:06:03 AM | 2002-07-17 | 7/16/2012 3:27:09 PM | UMULLER | RDGIBSON | Concentration | mg/L | Y | (fl. conc.) |  | Cobalt Total (fl. conc.) | 7440-48-4 | Metals - Total; | Metals - Total; EMS code: CO-T | Chemical | Total | LAB
CO-T | mg/L | Cobalt Total | MWID | Metals in Water by ICPMS (digested) | Metals in Water by ICPMS (digested) Based on EPA 200.8 | 0.0005 | Y | Y | Y | 36811 |  |  | 7/31/2024 1:46:41 PM | 2009-12-24 |  | RDGIBSON |  | Concentration | mg/L | Y | (fl. conc.) |  | Cobalt Total (fl. conc.) | 7440-48-4 | Metals - Total; | Metals - Total; EMS code: CO-T | Chemical | Total | LAB
CO-T | mg/L | Cobalt Total | F108 | SPR,IDA Chelation,ICPMS | SPR;IDA Chelation;ICPMS |  | Y | Y | Y | 4 | 11/30/2000 10:32:00 AM | 11/30/2000 10:32:00 AM | 9/10/2002 9:29:47 AM | 2001-03-08 |  | NPEPPIN |  | Concentration | mg/L | Y | (fl. conc.) |  | Cobalt Total (fl. conc.) | 7440-48-4 | Metals - Total; | Metals - Total; EMS code: CO-T | Chemical | Total | LAB
CO-T | ug/g | Cobalt Total | X551 | Metals in Biota (dry), total, acid digestion, ICP-SFMS | 0.1 gram of dried, homogenized biota is weighed in a microwave TFM vessel, and 8mL of Nitric Acid and 1mL of Hydrogen Peroxide are added. The sample is digested at 200 degrees C for 15 minutes in a closed vessel high pressure microwave oven. The sample is cooled, the volume adjusted to 100mL and introduced to an iductively coupled argon plasma-sector field mass spectrometer (ICP-SFMS) by pneumatic nebulization. Each respective element is measured at a specific mass to charge ratio, m/z value expressed in amu (atomic mass units), and at a specified instrument resolution. LR: Low Resolution ~300, MR: Mid Resolution ~4000 and HR: Hich Resolution ~10000. Ref: Method number 02-2705, National Laboratory for Environmental Testing, CCIW, Environment Canada, Burlington. Requested April 2003. | 0.002 | Y | Y | Y | 57 | 8/11/2015 | 10/30/2023 | 11/28/2023 2:21:33 PM | 2004-01-19 |  | UMULLER |  | MassPerMass | mg/kg | Y | (mass conc.) |  | Cobalt Total (mass conc.) | 7440-48-4 | Metals - Total; | Metals - Total; EMS code: CO-T | Chemical | Total | LAB
CO-T | ppm (S) | Cobalt Total | MS61 | ALS Minerals ME-MS61 | ALS Minerals ME-MS61. A prepared sample is digested with HNO3; HF; and HClO4 in a graphite heating block to near dryness. After cooling; the sample is boiled in diluted HCl. Bulk to volume with deionized water; mixed and analyzed by ICP-AES and ICP-MS for all elements. | 0.1 | Y | Y | Y | 0 |  |  |  | 2024-01-29 |  | EMS |  | DimensionlessRatio | ppm | N | (ratio) |  | Cobalt Total (ratio) | 7440-48-4 | Metals - Total; | Metals - Total; EMS code: CO-T | Chemical | Total | LAB
CO-T | mg/L | Cobalt Total | X072 | FA,HNO3,Graph.Furnace,AA | Samples are field-acidified with HNO3; HNO3 digestion; analysis by GFAA. | 0.001 | Y | Y | Y | 10 | 12/12/1996 9:40:00 AM | 12/12/1996 9:48:00 AM | 1/30/1997 11:20:55 PM | 1996-04-12 | 6/3/1998 6:51:34 PM | EMS |  | Concentration | mg/L | Y | (fl. conc.) |  | Cobalt Total (fl. conc.) | 7440-48-4 | Metals - Total; | Metals - Total; EMS code: CO-T | Chemical | Total | LAB
CO-T | ug/g | Cobalt Total | 0031 | ICP | ICP | 2 | Y | Y | Y | 4 |  |  | 10/23/1996 11:03:49 PM | 1996-04-12 | 6/3/1998 6:51:41 PM | EMS |  | MassPerMass | mg/kg | Y | (mass conc.) |  | Cobalt Total (mass conc.) | 7440-48-4 | Metals - Total; | Metals - Total; EMS code: CO-T | Chemical | Total | LAB
CO-T | ug/g | Cobalt Total | 0192 | Hom, Nitric/Perchlor,ICP | Sample already in a dry state so does not require drying before homogenization and nitric perchlorid digestion followed by ICP analysis. | 10 | Y | Y | Y | 708 |  |  | 10/24/1996 12:50:25 PM | 1996-04-12 | 6/3/1998 6:51:34 PM | EMS |  | MassPerMass | mg/kg | Y | (mass conc.) |  | Cobalt Total (mass conc.) | 7440-48-4 | Metals - Total; | Metals - Total; EMS code: CO-T | Chemical | Total | LAB
CO-T | ug/g | Cobalt Total | FS82 | PBM,dry60C,10mesh,SALM-ICP-MS | PBM;dry60C;10mesh;SALM-ICP-MS | 0.3 | Y | Y | Y | 3688 | 5/21/2002 | 7/24/2024 | 7/30/2024 10:00:14 AM | 2002-11-12 |  | UMULLER |  | MassPerMass | mg/kg | Y | (mass conc.) |  | Cobalt Total (mass conc.) | 7440-48-4 | Metals - Total; | Metals - Total; EMS code: CO-T | Chemical | Total | LAB
CO-T | mg/L | Cobalt Total | 0061 | HNO3 Dig: GFAA Analysis | HNO3 Dig: GFAA Analysis | 0.002 | Y | Y | Y | 86 |  |  | 8/26/2004 4:21:02 PM | 1996-04-12 | 6/3/1998 6:51:34 PM | EMS |  | Concentration | mg/L | Y | (fl. conc.) |  | Cobalt Total (fl. conc.) | 7440-48-4 | Metals - Total; | Metals - Total; EMS code: CO-T | Chemical | Total | LAB
CO-T | ug/g | Cobalt Total | FS01 | D&G,Nitric/HCl:ICP | Samples are dried at 60 C, then either sieved to -100mesh (150 um) and rolled to homogenize, or finely ground with a mortar and pestle or a shatter box laboratory mill and rolled (depending on the submitter's requirements). The samples are then weighed into Teflon digestion liners and digested with HNO3 and HCl in a microwave oven. Samples are cooled, volumized, and allowed to settle overnight before being decanted. Analysis is then performed via Inductively Coupled Argon Plasma Atomic Emission Spectrometry (ICP-AES) (Simultaneous multi element analysis). In ICP Spectrometry, the liquid sample is pumped to the nebulizer, where it is convertd into an aerosol. the nebulizer carrier gas flow than transports the wet aerosol past the impact bead, where the large droplets condense out and flow to the waste container. The remaining fine aerosol is carried into the plasma. There, the high temperature desolvates the aerosol and dissociates the sample material into individual atoms and ions, which are excited to emit light at wavelengths characteristic of the elements present. A diffraction grating in the spectrometer diffracts the various wavelengths of light through exit slits to the photo-multiplier tubes, which convert the light energy to electronic signals. | 0.8 | Y | Y | Y | 2642 | 8/30/1996 1:52:00 PM | 10/10/2007 | 5/24/2008 1:15:02 AM | 1997-03-19 | 6/3/1998 6:51:46 PM | MPAK |  | MassPerMass | mg/kg | Y | (mass conc.) |  | Cobalt Total (mass conc.) | 7440-48-4 | Metals - Total; | Metals - Total; EMS code: CO-T | Chemical | Total | LAB
CO-T | mg/L | Cobalt Total | F048 | Acid dig,conc'n,nebul.,ICP-OES | This method provides an automated procedure for the Inductively Coupled Argon Plasma Optical Emission Spectrophotometric (ICP-OES) determination and quantitation of trace amounts of metals in surface, ground and precipitation waters after a manual acid digestion and concentration procedure and ultrasonic nebulisation. Organically-bound metals and metal complexes are converted to metal ions by digestion with nitric acid [method details says HCl]. To increase the signal detected by the ICP, the samples are concentrated 10 times. The liquid sample is pumped onto the piezoelectric transducer, where the ultrasonic action efficiently converts it to a fine, dense aerosol. The nebulizer carrier gas flow then transports the wet aerosol through a heated U-tube, where the solvent is vaporized. Subsequently, most of the solvent vapour is removed from the aerosol stream by passing the stream through a chilled condenser. The result is a dry, analyte-laden aerosol which is introduced to the plasma. There, the high temperature desolvates the aerosol and dissociates the sample material into individual atoms and ions, which are excited to emit light at wave lengths characteristic of the elements present. Dissolved samples are field filtered through 0.45 micron cellulose acetate filter prior to acid preservation. Samples for extractable analysis are acidified with 2 mL concentrated HNO3 per litre and allowed to stand a minimum of 24 hours before analysis. Clients providing these samples are advised to preserve samples at the collection site. | 0.0001 | Y | Y | Y | 6159 | 7/4/1997 2:35:00 PM | 4/23/2003 | 9/11/2019 8:46:38 AM | 1997-10-20 | 6/3/1998 6:51:47 PM | NPEPPIN |  | Concentration | mg/L | Y | (fl. conc.) |  | Cobalt Total (fl. conc.) | 7440-48-4 | Metals - Total; | Metals - Total; EMS code: CO-T | Chemical | Total | LAB
CO-T | mg/L | Cobalt Total | METN | MET-T-NP-CCMS-VA | Ultra trace metals in water are analyzed by CRC ICPMS, based on US EPA Method 6020B (July 2014). The detection limits provided can only be met for undigested samples. This procedure is intended for colorless, non-turbid, acid-preserved water samples (i.e. pristine water samples), having turbidity < 1 NTU and no odor. Where turbidity exceeds 1 NTU, and/or the sample is colored and has an odor, results may be biased low compared to true Total Metals concentrations. ALS recommends that turbidity analysis be requested on samples submitted for this test to aid with interpretation of results. Where turbidity is <1NTU, undigested metals are equivalent to total metals concentrations. | 5e-06 | Y | Y | Y | 3648 | 4/3/2012 | 8/19/2024 | 8/26/2024 1:03:06 PM | 2019-08-20 |  | KATFRASE |  | Concentration | mg/L | Y | (fl. conc.) |  | Cobalt Total (fl. conc.) | 7440-48-4 | Metals - Total; | Metals - Total; EMS code: CO-T | Chemical | Total | LAB
CO-T | ug/g | Cobalt Total | EQ02 | EQUIS (ug/g) | EQUIS (ug/g) | 10 | N | Y | Y | 760 |  |  | 10/23/1996 12:08:28 AM | 1996-04-12 | 6/3/1998 6:51:34 PM | EMS |  | MassPerMass | mg/kg | Y | (mass conc.) |  | Cobalt Total (mass conc.) | 7440-48-4 | Metals - Total; | Metals - Total; EMS code: CO-T | Chemical | Total | LAB
CO-T | mg/L | Cobalt Total | F082 | Aliquot,HNO3/HCL dig,ICPMS | Aliquot microwave digested with 3:1 HNO3/HCL. Analysis by ICPMS. For HiVol samples preliminary result are reported in mg/L and loading results are reported in ug/m3 | 0.0003 | Y | Y | Y | 27836 | 7/26/2005 7:04:00 AM | 8/23/2024 | 8/26/2024 4:53:05 PM | 2000-06-13 | 8/15/2019 11:06:33 AM | NPEPPIN | KATFRASE | Concentration | mg/L | Y | (fl. conc.) |  | Cobalt Total (fl. conc.) | 7440-48-4 | Metals - Total; | Metals - Total; EMS code: CO-T | Chemical | Total | LAB
CO-T | ug/g wet | Cobalt Total | X560 | Metals in Biota (wet), total, acid digestion, ICP-MS | Five grams of wet, homogenized biota is weighed in a centrifuge tube, and 2.5 mL of Nitric Acid and 2.5 mL of Hydrochloric acid are added. The sample is digested at room temperature for one hour and then digested at 90-95 degrees C in a water bath or heat block for one more hour. The sample is cooled and the volume is made to 50 mL with DI. The digestate is then diluted at ten fold and introduced to a collision/reaction cell inductively coupled argon plasma mass spectrometer (CRC ICPMS) by pneumatic nebulization. Each respective element is measured at a specific mass to charge ratio, m/z value expressed in amu (atomic mass units). - Maxxam Analytics (BC). Added to EMS Dec 10, 2008. | 0.1 | Y | Y | Y | 541 | 11/15/2008 4:00:00 AM | 10/30/2023 | 11/28/2023 2:21:33 PM | 2008-12-11 | 3/11/2015 4:17:10 PM | RDGIBSON | EMS | MassPerMass | µg/g | Y | (mass conc.) | (wet) | Cobalt Total (wet) (mass conc.) | 7440-48-4 | Metals - Total; | Metals - Total; EMS code: CO-T | Chemical | Total | LAB
CO-T | ug/g | Cobalt Total | 0196 | D&G: LTA/HCl: ICP | D&G: LTA/HCl: ICP | 10 | Y | Y | Y | 27 |  |  | 10/24/1996 2:08:12 PM | 1996-04-12 | 6/3/1998 6:51:34 PM | EMS |  | MassPerMass | mg/kg | Y | (mass conc.) |  | Cobalt Total (mass conc.) | 7440-48-4 | Metals - Total; | Metals - Total; EMS code: CO-T | Chemical | Total | LAB
CO-T | ug/g | Cobalt Total | MTID | MET-DRY-CCMS-VA | Metals in Tissue by CRC ICPMS (DRY) Methodology: This method is conducted following British Columbia Lab Manual method "Metals in Animal Tissue and Vegetation (Biota) - Prescriptive". Tissue samples are homogenized and sub-sampled prior to hotblock digestion with nitric and hydrochloric acids, in combination with addition of hydrogen peroxide. Instrumental analysis is by collision cell inductively coupled plasma - mass spectrometry (modified from EPA Method 6020A). Method Limitation: This method employs a strong acid/peroxide digestion, and is intended to provide a conservative estimate of bio-available metals. Near complete recoveries are achieved for most toxicologically important metals, but elements associated with recalcitrant minerals may be only partially recovered. | 0.02 | Y | Y | Y | 122 | 8/13/2014 | 9/13/2022 | 10/26/2022 7:14:14 AM | 2015-03-09 | 3/11/2015 4:12:05 PM | RDGIBSON | EMS | MassPerMass | mg/kg | Y | (mass conc.) |  | Cobalt Total (mass conc.) | 7440-48-4 | Metals - Total; | Metals - Total; EMS code: CO-T | Chemical | Total | LAB
CO-T | mg/L | Cobalt Total | X285 | HNO3 Dig:ICP Analysis | HNO3 Dig:ICP Analysis | 0.015 | Y | Y | Y | 686 | 8/10/1993 | 10/5/1993 | 4/30/2004 1:54:39 PM | 1998-05-22 | 6/3/1998 6:51:49 PM | NPEPPIN |  | Concentration | mg/L | Y | (fl. conc.) |  | Cobalt Total (fl. conc.) | 7440-48-4 | Metals - Total; | Metals - Total; EMS code: CO-T | Chemical | Total | LAB
CO-T | ug/g wet | Cobalt Total | X552 | Metals in Biota (wet), total, acid digestion, ICP-SFMS | One gram of wet, homogenized biota is weighed in a microwave TFM vessel, and 8mL of Nitric Acid and 1mL of Hydrogen Peroxide are added. The sample is digested at 200 degrees C for 15 minutes in a closed vessel high pressure microwave oven. The sample is cooled, the volume adjusted to 100mL and introduced to an inductively coupled argon plasma-sector field mass spectrometer (ICP-SFMS) by pneumatic nebulization. Each respective element is measured at a specific mass to charge ratio, m/z value expressed in amu (atomic mass units), and at a specified instrument resolution. LR: Low Resolution ~300, MR: Mid Resolution ~4000. Ref: 02-2704, National Laboratory for Environmental Testing, CCIW, Environment Canada, Burlington. Requested April 2003. | 0.0002 | Y | Y | Y | 37 | 8/9/2016 | 10/16/2023 | 11/28/2023 2:21:32 PM | 2004-01-19 | 4/19/2021 1:22:00 PM | UMULLER | EMS | MassPerMass | µg/g | Y | (mass conc.) | (wet) | Cobalt Total (wet) (mass conc.) | 7440-48-4 | Metals - Total; | Metals - Total; EMS code: CO-T | Chemical | Total | LAB
CO-T | ug/m3 | Cobalt Total | 5312 | Tef:N/P Dig:ICP Analysis | A nitric/perchloric acid digestion if used to bring the metals into solution prior to ICP analysis. For Hi-Vol samples preliminary results are reported as mg/L and 24 hour loading results as ug/m3. | 0.002 | Y | Y | Y | 1185 | 11/7/1997 | 3/29/2000 | 9/28/2000 1:34:41 PM | 1996-10-13 | 6/3/1998 6:51:46 PM | EMS |  | AirConcentration | µg/m³ | Y | (air conc.) |  | Cobalt Total (air conc.) | 7440-48-4 | Metals - Total; | Metals - Total; EMS code: CO-T | Chemical | Total | LAB
CO-T | mg/L | Cobalt Total | X145 | Dig, ICP | Dig; ICP | 0.015 | Y | Y | Y | 91 | 5/8/2003 9:40:00 AM | 9/12/2003 3:16:00 PM | 12/1/2003 5:30:19 PM | 1996-04-12 | 6/3/1998 6:51:41 PM | EMS |  | Concentration | mg/L | Y | (fl. conc.) |  | Cobalt Total (fl. conc.) | 7440-48-4 | Metals - Total; | Metals - Total; EMS code: CO-T | Chemical | Total | LAB
CO-T | mg/L | Cobalt Total | F033 | Acid Dig,Centrifuge,ICPMS | For the determination of total recoverable elements in water or waste water, take 100 ml aliquot from a well mixed, acid preserved, sample containing not more than 0.25% (w/v) total solids and transfer to a 250 ml Griffin beaker. (If total solids are greater than 0.25% reduce the size of the aliquot by a proportionate amount). Add 1 ml of conc. nitric acid and 0.5 ml conc. hydrochloric acid. Heat on a hot plate until the volume has been reduced to 15-20ml, ensuring that the sample does not boil. Allow to cool and transfer to a 50 ml volumetric flask. Dilute to volume with ASTM type I water. Centrifuge the sample or allow to stand overnight to separate insoluble material. Prior analysis, pipette 20 ml into a 50 ml volumetric flask and dilute to volume with ASTM type I water. Analysis should be performed within seventy two hours of the completed preparation. Note that different laboratories may have very different field preservation, digestion procedures and detection limits for this method, so these should be documented. | 0.001 | Y | Y | Y | 40322 | 11/20/1997 9:31:00 AM | 8/22/2024 3:24:00 AM | 8/25/2024 9:47:36 PM | 1998-05-14 | 6/3/1998 6:51:49 PM | NPEPPIN |  | Concentration | mg/L | Y | (fl. conc.) |  | Cobalt Total (fl. conc.) | 7440-48-4 | Metals - Total; | Metals - Total; EMS code: CO-T | Chemical | Total | LAB
CO-T | ug/g | Cobalt Total | ICPM | Metals in Soil by CRC ICPMS | This analysis is carried out using procedures from CSR Analytical Method: "Strong Acid Leachable Metals (SALM) in Soil", BC Ministry of Environment, 26 June 2009, and procedures adapted from EPA Method 200.2. The sample is manually homogenized, dried at 60 degrees Celsius, sieved through a 63 um (230 mesh) sieve, and a representative subsample of the dry material is weighed. The sample is then digested at 95 degrees Celsius for 2 hours by block digester using concentrated nitric and hydrochloric acids. Instrumental analysis is by collision cell inductively coupled plasma - mass spectrometry (modifed from EPA Method 6020A). Method Limitation: This method is not a total digestion technique. It is a very strong acid digestion that is intended to dissolve those metals that may be environmentally available. By design, elements bound in silicate structures are not normally dissolved by this procedure as they are not usually mobile in the environment. | 0.1 | Y | Y | Y | 103 | 9/19/2017 | 9/24/2023 | 1/12/2024 11:16:52 AM | 2018-01-09 |  | JKROGH |  | MassPerMass | mg/kg | Y | (mass conc.) |  | Cobalt Total (mass conc.) | 7440-48-4 | Metals - Total; | Metals - Total; EMS code: CO-T | Chemical | Total | LAB
CO-T | mg/L | Cobalt Total | F038 | HNO3/HCl dig,ICP-AES | Dissolved and extractable water samples are ready for analysis once collected. Total samples are autoclaved with HNO3 and hydrochloric acid (HCI) in capped Teflon vials. Analysis is performed via Inductively Coupled Argon Plasma Atomic Emission Spectrometry (ICP-AES)(simultaneous multielement analysis), Graphite Furnace Atomic Absorption Spectrometry (GF-AAS) (low level analysis) and/or Flame Atomic Absorption Spectrometry (F-AAS). | 0.006 | Y | Y | Y | 5424 | 8/1/1996 6:04:00 PM | 5/6/2008 | 9/11/2019 8:47:42 AM | 1997-05-08 | 6/3/1998 6:51:46 PM | MPAK |  | Concentration | mg/L | Y | (fl. conc.) |  | Cobalt Total (fl. conc.) | 7440-48-4 | Metals - Total; | Metals - Total; EMS code: CO-T | Chemical | Total | LAB
CO-T | ug/g | Cobalt Total | 0190 | D&G:N/P Dig:ICP Anal. | Sample is dried and grind-mixed before nitric perchloric digestion followed by ICP analysis. | 10 | Y | Y | Y | 949 |  |  | 10/24/1996 3:06:16 PM | 1996-04-12 | 6/3/1998 6:51:41 PM | EMS |  | MassPerMass | mg/kg | Y | (mass conc.) |  | Cobalt Total (mass conc.) | 7440-48-4 | Metals - Total; | Metals - Total; EMS code: CO-T | Chemical | Total | LAB
CO-T | mg/L | Cobalt Total | 0060 | HNO3 Dig: HGA Analysis | HNO3 Dig: HGA Analysis | 0.001 | Y | Y | Y | 253 |  |  | 10/24/1996 12:22:28 PM | 1996-04-12 | 6/3/1998 6:51:34 PM | EMS |  | Concentration | mg/L | Y | (fl. conc.) |  | Cobalt Total (fl. conc.) | 7440-48-4 | Metals - Total; | Metals - Total; EMS code: CO-T | Chemical | Total | LAB
CO-T | mg/dm2/d | Cobalt Total | DUST | HNO3 dig.(microwave or hot block),ICPMS | HNO3 dig.(microwave or hot block);ICPMS |  | Y | Y | Y | 420 | 7/19/2016 1:11:00 PM | 7/31/2024 | 7/31/2024 11:23:12 AM | 2008-02-07 |  | RDGIBSON |  | MassFlux | mg/dm²/d | Y | (mass flux) |  | Cobalt Total (mass flux) | 7440-48-4 | Metals - Total; | Metals - Total; EMS code: CO-T | Chemical | Total | LAB
CO-T | ug/g | Cobalt Total | FB04 | Freeze Dry,Hom,Acid Dig,ICP | Sample moisture is determined by weighing samples before and after freeze-drying. Freeze-dried samples are ground in a Braun blender before being weighed into Teflon PFA liners and digested with concentrated HNO3, concentrated HCl, and 50% H2O2 in a microwave oven; which results in a complete sample dissolution. Samples are collected and volumized with Type I water. Analysis is performed via Inductively Coupled Argon Plasma Atomic Emission Spectrometry (ICP-AES) (simultaneous multielement analysis), Graphite Furnace Atomic Absorption Spectrometry (GF-AAS) (low level Cd and Pb analysis), Cold Vapor Atomic Fluorescence Spectrometry (CV-AFS) (Mercury analysis) and/or Flame Atomic Absorption Spectrometry (F-AAS) (K analysis). | 0.4 | Y | Y | Y | 241 | 8/22/1996 6:11:00 PM | 2/1/2002 2:30:00 PM | 6/23/2003 1:15:29 PM | 1997-02-27 | 6/3/1998 6:51:46 PM | MPAK |  | MassPerMass | mg/kg | Y | (mass conc.) |  | Cobalt Total (mass conc.) | 7440-48-4 | Metals - Total; | Metals - Total; EMS code: CO-T | Chemical | Total | LAB
CO-T | mg/L | Cobalt Total | MSEA | Metals in Seawater (undigested) by (CRC) ICP-MS | This method describes the multi-elemental determination of trace elements by Collision/Reaction Cell (CRC) ICP-MS. The method measures ions produced by a radio-frequency inductively coupled plasma. Analyte species originating in a liquid are pneumatically nebulized and the resulting aerosol transported by argon gas into the plasma torch. The ions produced are entrained in the plasma gas and extracted, by means of a differentially pumped vacuum interface, into a mass spectrometer. The ions flow through a Collision/Reaction Cell (CRC) where depending on the mode of analysis either the ions passes through the cell unhindered (no gas - classical ICP-MS mode), or collisions/reactions chemistry takes place between the ions and a neutral gas such as hydrogen or helium. Any polyatomic ions which give interference on the analytes are removed. The ions are then sorted according to their mass-to-charge ratios by a quadrupole mass spectrometer. The ions are then sorted according to their mass-to-charge ratios by a quadrupole mass spectrometer. The ions transmitted through the quadrupole are quantified by a channel electron multiplier and the ion information processed by a data handling system. Although most polyatomic interference is removed, any additional interference must be assessed and valid corrections applied or the data flagged to indicate problems. Interference correction must include compensation for background ions contributed by the plasma gas, reagents, and constituents of the sample matrix. Instrument drift as well as suppressions or enhancements of instrument response caused by the sample matrix must be corrected for by the use of internal standards. | 3e-05 | Y | Y | Y | 441 | 8/12/2009 12:08:00 AM | 4/29/2024 | 6/7/2024 10:03:16 AM | 2009-09-08 |  | RDGIBSON |  | Concentration | mg/L | Y | (fl. conc.) |  | Cobalt Total (fl. conc.) | 7440-48-4 | Metals - Total; | Metals - Total; EMS code: CO-T | Chemical | Total | LAB
CO-T | ug/g wet | Cobalt Total | MTIW | MET-WET-CCMS-MID-VA | Metals in Tissue by CRC ICPMS (WET) Methodology: This method is conducted following British Columbia Lab Manual method "Metals in Animal Tissue and Vegetation (Biota) - Prescriptive". Tissue samples are homogenized and sub-sampled prior to hotblock digestion with nitric and hydrochloric acids, in combination with addition of hydrogen peroxide. Instrumental analysis is by collision cell inductively coupled plasma - mass spectrometry (modified from EPA Method 6020A). Method Limitation: This method employs a strong acid/peroxide digestion, and is intended to provide a conservative estimate of bio-available metals. Near complete recoveries are achieved for most toxicologically important metals, but elements associated with recalcitrant minerals may be only partially recovered. | 0.004 | Y | Y | Y | 24 | 11/25/2018 | 12/6/2018 | 7/21/2021 2:55:07 PM | 2018-12-24 |  | EMS |  | MassPerMass | µg/g | Y | (mass conc.) | (wet) | Cobalt Total (wet) (mass conc.) | 7440-48-4 | Metals - Total; | Metals - Total; EMS code: CO-T | Chemical | Total | LAB
CO-T | ug/m3 | Cobalt Total | 6040 | Aqua Regia Dig:ICP | Aqua regia digestion is used to bring the metals into solution prior to ICP analysis. For Hi-Vol samples, preliminary results are reported as mg/L and24 hr loadings as ug/m3 | 0.002 | Y | Y | Y | 437 | 2/13/2002 | 2/9/2005 | 2/23/2005 1:13:47 AM | 1998-05-05 | 6/3/1998 6:51:48 PM | NPEPPIN |  | AirConcentration | µg/m³ | Y | (air conc.) |  | Cobalt Total (air conc.) | 7440-48-4 | Metals - Total; | Metals - Total; EMS code: CO-T | Chemical | Total | LAB
CO-T | mg/L | Cobalt Total | 0042 | GI HNO3 Dig: ICP | Sample aliquot taken from unfiltered unpreserved General Ions bottle; HNO3 digestion; ICP analysis. | 0.004 | Y | Y | Y | 6238 | 9/20/1996 | 7/5/2019 3:21:00 PM | 7/12/2019 6:49:13 PM | 1996-04-12 | 6/3/1998 6:51:41 PM | EMS |  | Concentration | mg/L | Y | (fl. conc.) |  | Cobalt Total (fl. conc.) | 7440-48-4 | Metals - Total; | Metals - Total; EMS code: CO-T | Chemical | Total | LAB
CO-T | ug/m3 | Cobalt Total | F088 | Aliquot,HNO3/HCl dig, ICPMS,loading | Aliquot microwave digested with 3:1 HNO3/HCL. Analysis by ICPMS. For HiVol samples preliminary result are reported in mg/L and loading results are reported in ug/m3 |  | Y | Y | Y | 565 | 8/17/2000 10:46:00 AM | 12/14/2006 | 12/30/2006 1:01:36 AM | 2000-08-16 |  | NPEPPIN |  | AirConcentration | µg/m³ | Y | (air conc.) |  | Cobalt Total (air conc.) | 7440-48-4 | Metals - Total; | Metals - Total; EMS code: CO-T | Chemical | Total | LAB
CO-T | ppm (S) | Cobalt Total | 6042 | Aqua Regia by ICPMS | These were sediment/soils digested with Aqua-Regia (Nitric Acid & Hydrochloric Acids) and then analysis by ICPMS. Results would be reported in ppm s. | 0.001 | Y | Y | Y | 71 | 5/14/2022 | 7/27/2024 | 7/29/2024 9:13:08 AM | 2022-06-28 |  | EMS |  | DimensionlessRatio | ppm | N | (ratio) |  | Cobalt Total (ratio) | 7440-48-4 | Metals - Total; | Metals - Total; EMS code: CO-T | Chemical | Total | LAB
CO-T | ug/g | Cobalt Total | X230 | Hom,HNO3/HCl,ICP | The sample is dried, ground to pass 100 mesh, and digested with a nitric-hydrochloric acid mixture to solubilize the solid matter and to remove the organic material by oxidation and volatilization. The completed digests are submitted to metals instrumentation personnel for metals analysis by Inductively Coupled Atomic Emission Spectrophotometry (ICP-AES), or hydride analysis by atomic absorption. | 10 | Y | Y | Y | 176 | 2/12/2002 | 12/30/2005 | 2/3/2006 1:13:12 AM | 1996-04-12 | 6/3/1998 6:51:34 PM | EMS |  | MassPerMass | mg/kg | Y | (mass conc.) |  | Cobalt Total (mass conc.) | 7440-48-4 | Metals - Total; | Metals - Total; EMS code: CO-T | Chemical | Total | LAB
CO-T | ug/g | Cobalt Total | X492 | Homo,HCl-HNO3 dig,ICP-OES,dry calc | Homogenize samples and take sub-samples. For Total Recoverable Metals Digest with Aqua Regia. Analyses by ICPOES (Inductively Coupled Plasma Optical Emission Spectroscopy). Perform moisture analysis on separate sub-sample and calculate dry value for trace elements. | 2 | Y | Y | Y | 7 | 8/16/2022 6:31:00 PM | 8/16/2022 6:42:00 PM | 2/25/2023 1:41:51 PM | 1998-11-13 |  | NPEPPIN |  | MassPerMass | mg/kg | Y | (mass conc.) |  | Cobalt Total (mass conc.) | 7440-48-4 | Metals - Total; | Metals - Total; EMS code: CO-T | Chemical | Total | LAB
CO-T | ug/g | Cobalt Total | PB02 | PBM,HNO3/HCl dig,dry60C,100mesh, | Performance Based Method. Samples are dried at 60 C and sieved using a .149 mm (100 mesh) sieve. Solid samples are digested with the combination of nitric and hydrochloric acids with a 1:1 ratio of acids. For environmental purposes, strong acid leaches are used routinely since they have the advantage of speed, low cost, and superio reproducibility. All metals trapped within the refractory matrix are unlikely to have environmental consequences. Instrumental analysis of extracts produced using this digestion method can be performed by a variety of techniques. |  | Y | Y | Y | 52 | 5/24/2001 3:42:00 PM | 11/26/2001 10:23:00 AM | 5/26/2003 2:27:41 PM | 2001-07-18 |  | NPEPPIN |  | MassPerMass | mg/kg | Y | (mass conc.) |  | Cobalt Total (mass conc.) | 7440-48-4 | Metals - Total; | Metals - Total; EMS code: CO-T | Chemical | Total | LAB
CO-T | mg/L | Cobalt Total | X170 | ICP | This code is a catch-all for a large variety of ICP methods including water and also solids digested into solutions as an intermediate result. Different procedures will have greatly different detection limits. (a) For air particulates, preliminary results are reported in mg/L and loading results as mg/dm2/d. (b) For Hi Vol samples preliminiary results are same as reported in mg/L and loading results as ug/m3. (c) Note that for Rain samples there is no filtration nor digestion steps. | 0.003 | Y | Y | Y | 5701 | 3/28/1996 6:03:00 PM | 11/7/2018 | 10/24/2022 10:58:25 AM | 1996-04-12 | 6/3/1998 6:51:34 PM | EMS |  | Concentration | mg/L | Y | (fl. conc.) |  | Cobalt Total (fl. conc.) | 7440-48-4 | Metals - Total; | Metals - Total; EMS code: CO-T | Chemical | Total | LAB
CO-T | mg/L | Cobalt Total | 0040 | HNO3 Dig: ICP Analysis | HNO3 Dig: ICP Analysis | 0.1 | Y | Y | Y | 15010 | 6/25/2010 3:15:00 PM | 12/1/2010 | 12/9/2010 3:55:26 PM | 1996-04-12 | 6/3/1998 6:51:34 PM | EMS |  | Concentration | mg/L | Y | (fl. conc.) |  | Cobalt Total (fl. conc.) | 7440-48-4 | Metals - Total; | Metals - Total; EMS code: CO-T | Chemical | Total | LAB
CO-T | mg/L | Cobalt Total | F068 | HNO3dig,TeflonTube/Autoclave,ICP-MS | This method is applicable to determiniation of dissolved, extractable, and total recoverable metals in ground waters, surface waters, drinking waters, and wastewaters. Dissolved, extractable, and total recoverable sample fractions are acidified to pH <1.5 with analytical grade nitric acid (HNO3) to prevent the metals from adsorbing to the container material or precipitating. Extractable and total recoverable water samples are collected in 250 ml acid washed polyethylene bottles and are preserved with Nitric acid (0.5 ml acid for 100 mLs of sample). All preserved samples are stored at room temperature and can be held for up to 6 months. Unfiltered, un-preserved samples are stored at 4C for no more than 30 days. For dissolved and extractable water fractions, no sample preparation is necessary. Samples are simply poured in autosampler tubes (i.e. 15 ml centrifuge tubes), and placed on the autosampler and analyzed. Samples for total recoverable metals are digested with Nitric acid in special Teflon tubes in an autoclave prior to analyses. | 5e-06 | Y | Y | Y | 2431 | 6/1/1999 1:32:00 PM | 7/4/2008 9:48:00 AM | 7/30/2024 6:32:12 AM | 1999-06-28 | 6/30/1999 1:05:24 PM | NPEPPIN | NPEPPIN | Concentration | mg/L | Y | (fl. conc.) |  | Cobalt Total (fl. conc.) | 7440-48-4 | Metals - Total; | Metals - Total; EMS code: CO-T | Chemical | Total | LAB
CO-T | mg/L | Cobalt Total | TCLP | Toxicity Characteristic Leaching Procedure | The TCLP procedure is designed to simulate the climatic leaching action | 0.1 | Y | Y | Y | 22 | 12/8/2009 7:49:00 PM | 9/28/2011 5:53:00 PM | 11/3/2011 11:10:18 PM | 2009-03-05 |  | JCHUDAK |  | Concentration | mg/L | Y | (fl. conc.) |  | Cobalt Total (fl. conc.) | 7440-48-4 | Metals - Total; | Metals - Total; EMS code: CO-T | Chemical | Total | LAB
CO-T | mg/L | Cobalt Total | EQ01 | EQUIS (mg/L) | EQUIS (mg/L) | 0.005 | N | Y | Y | 1792 |  |  | 10/23/1996 12:08:27 AM | 1996-04-12 | 6/3/1998 6:51:34 PM | EMS |  | Concentration | mg/L | Y | (fl. conc.) |  | Cobalt Total (fl. conc.) | 7440-48-4 | Metals - Total; | Metals - Total; EMS code: CO-T | Chemical | Total | LAB
CO-T | mg/L | Cobalt Total | ULTR | Total Metals in Water by ICPMS (Ultra) | Total Metals in Water by ICPMS (Ultra), EPA SW-846 3005A/6020A, MET-TOT-ULTRA-MS-VA This analysis is carried out using procedures adapted from "Standard Methods for the Examination of Water and Wastewater" published by the American Public Health Association, and with procedures adapted from "Test Methods for Evaluating Solid Waste" SW-846 published by the United States Environmental Protection Agency (EPA). The procedures may involve preliminary sample treatment by acid digestion, using either hotblock or microwave oven, or filtration (EPA Method 3005A). Instrumental analysis is by inductively coupled plasma - mass spectrometry (EPA Method 6020A). | 5e-05 | Y | Y | Y | 1 | 12/28/2017 | 12/28/2017 | 1/9/2018 11:06:10 AM | 2018-01-09 |  | JKROGH |  | Concentration | mg/L | Y | (fl. conc.) |  | Cobalt Total (fl. conc.) | 7440-48-4 | Metals - Total; | Metals - Total; EMS code: CO-T | Chemical | Total | LAB
CO-T | mg/L | Cobalt Total | X561 | Total metals: FA, HNO3, HR-ICPMS (No Prep)Prep) | Ultra trace metals in water are analyzed by high resolution inductively coupled plasma mass spectrometry (HR-ICPMS), modified from US EPA 200.8, (Revision 5.5). The detection limits provided can only be met for undigested samples. This procedure is intended for colorless, non-turbid, acid-preserved water samples (i.e. pristine water samples), having turbidity < 1 NTU and no odor. Where turbidity exceeds 1 NTU, and/or the sample is colored and has an odor, results may be biased low compared to true Total Metals concentrations. ALS recommends that turbidity analysis be requested on samples submitted for this test to aid with interpretation of results. | 5e-06 | Y | Y | Y | 2756 | 6/8/2015 | 10/5/2023 | 1/31/2024 2:37:41 PM | 2015-10-22 |  | RDGIBSON |  | Concentration | mg/L | Y | (fl. conc.) |  | Cobalt Total (fl. conc.) | 7440-48-4 | Metals - Total; | Metals - Total; EMS code: CO-T | Chemical | Total | LAB
CO-T | mg/L | Cobalt Total | X555 | Metals in water, total, "in-bottle" digestion, ICP-SFMS | The whole water sample collected in a HDPE/Teflon bottle, is de-capped in a Class 100 workstation, acidified to 2% (v/v) HNO3 and 1% (v/v) HCl and recapped. The sample is "in-bottle" digested in a convection oven, set at 60 degrees C for 16 hours. The sample is cooled and introduced to an inductively coupled argon plasma-sector field mass spectrometer (ICP-SFMS) by pneumatic nebulization. Each respective element is measured at a specific mass to charge ratio, m/z value expressed in amu (atomic mass units), and at resolution. LR: Low Resolution ~300, MR: Mid Resolution ~4000 and HR: High a specified instrument Resolution ~10000. Ref: Method number 02-2003, National Laboratory for Environmental Testing, CCIW, Environment Canada, Burlington. Requested April 2003. | 2e-06 | Y | Y | Y | 17337 | 5/20/2003 | 12/6/2023 2:08:00 AM | 2/14/2024 9:05:32 AM | 2004-01-19 |  | UMULLER |  | Concentration | mg/L | Y | (fl. conc.) |  | Cobalt Total (fl. conc.) | 7440-48-4 | Metals - Total; | Metals - Total; EMS code: CO-T | Chemical | Total | LAB
CO-T | ppm (S) | Cobalt Total | MS41 | ALS Minerals ME-MS41 | ALS Minerals ME-MS41. A prepared sample is digested with aqua regia in a graphite heating block. After cooling; the resulting solution is diluted to with deionized water; mixed and analyzed by ICP-AES. Samples are then analyzed by ICP-MS for the remaining suite of elements. | 0.1 | Y | Y | Y | 0 |  |  |  | 2024-01-29 |  | EMS |  | DimensionlessRatio | ppm | N | (ratio) |  | Cobalt Total (ratio) | 7440-48-4 | Metals - Total; | Metals - Total; EMS code: CO-T | Chemical | Total | LAB
CO-T | mg/L | Cobalt Total | 1349 | Tot Metals in Seawater by CRC ICPMS (BC) | Tot Metals in Seawater by CRC ICPMS (BC) | 5e-05 | Y | Y | Y | 98 | 8/16/2019 | 12/5/2021 | 12/13/2021 12:34:16 PM | 2019-09-30 |  | KATFRASE |  | Concentration | mg/L | Y | (fl. conc.) |  | Cobalt Total (fl. conc.) | 7440-48-4 | Metals - Total; | Metals - Total; EMS code: CO-T | Chemical | Total | LAB
CO-T | mg/L | Cobalt Total | 0200 | 20%A.R.Dig:AA Anal. | 20%A.R.Dig:AA Anal. | 0.1 | Y | Y | Y | 1 |  |  | 10/23/1996 11:21:16 PM | 1996-04-12 | 6/3/1998 6:51:41 PM | EMS |  | Concentration | mg/L | Y | (fl. conc.) |  | Cobalt Total (fl. conc.) | 7440-48-4 | Metals - Total; | Metals - Total; EMS code: CO-T | Chemical | Total | LAB
CO-T | mg/L | Cobalt Total | F024 | HNO3/HCl,GFAA | Dissolved and extractable water samples are ready for analysis once collected. Total samples are autoclaved with HNO3 and hydrochloric acid (HCl) in capped Teflon vials. Analysis is performed via Inductively Coupled Argon Plasma Atomic Emission Spectrometry (ICP-AES) (simultaneous multielement analysis), Graphite Furnace Atomic Absorption Spectrometry (GF-AAS) (low level analysis) and/or Flame Atomic Absorption Spectrometry (F-AAS). | 0.001 | Y | Y | Y | 8 | 2/3/1998 5:16:00 PM | 7/14/1998 4:48:00 PM | 3/11/1999 6:01:07 PM | 1997-04-28 | 6/3/1998 6:51:46 PM | MPAK |  | Concentration | mg/L | Y | (fl. conc.) |  | Cobalt Total (fl. conc.) | 7440-48-4 | Metals - Total; | Metals - Total; EMS code: CO-T | Chemical | Total | LAB
CO2- | % | Carbon Dioxide | CGAS05 | ALS Minerals C-GAS05 | ALS Minerals C-GAS05. Determine Total Inorganic Carbon (TIC) by using diluted perchloric acid to react with metal carbonate and yield carbon dioxide. Concentration of CO2 is determined by coulometer. | 0.02 | Y | Y | Y | 0 |  |  |  | 2024-01-29 |  | RDGIBSON |  | DimensionlessRatio | % | N | (ratio) |  | Carbon Dioxide (ratio) |  |  | EMS code: CO2- | Chemical |  | LAB
CO3- | mg/L | carbonate (CAS 3812-32-6) | X310 | Auto Potentiometric Titration | For alkalinity, the samples are (if pH >8.3) first titrated to a pH of 8.3, which corresponds to the equivalence point for the conversion of carbonate to bicarbonate ion (phenolphthalein alkalinity). The titration is continued to a pH of 4.5, which corresponds approx. to the equivalence point for the conversion of bicarbonate ion to carbonic acid (entire titration indicates total alkalinity). An extrapolation technique based on the near proportionality of concentration of H ions to excess of titrant beyond the equivalence pt. The amount of standard acid required to reduce pH exactly 0.30pH (ie. pH4.2) unit is measured carefully. Because this change in pH corresponds to an exact coupling of the H ion concentration, a simple extrapolation can be made to the equivalence point. The actual pH of the stoichiometric end point in alkalinity determinations is the pH at which the inflection in the titration curve occurs and is directly related to the amount of carbonate alkalinity originally present in the sample. A curve can be constructed by plotting pH vs mL of acid to find this point directly. *** For acidity, the samples are titrated to a pH of 8.3, which corresponds to stoichiometric neutralization of carbonic acid to bicarbonate. Because the colour change of phenolphthalein indicator is close to pH 8.3, this value generally is accepted as a standard end point for titration of total acidity, including CO2 and most weak acids an equivalence point for the conversion of carbonate to bicarbonate ion (phenolphthalein alkalinity). The actual pH of the stoichiometric end point in acidity determinations is the pH at which the inflection in the titration curve occurs and is directly related to the amount of titratable acids originally present in the sample. A curve can be constructed by plotting pH vs. mL of NaOH to find this point directly. (To obtain the actual substances contributing to acidity/alkalinity, a pH vs mL of acid/NaOH curve should be constructed.) | 1 | Y | Y | Y | 90 | 11/24/2019 | 8/11/2024 | 8/20/2024 9:51:05 AM | 2020-01-29 |  | RDGIBSON |  | Concentration | mg/L | Y | (fl. conc.) |  | carbonate(fl. conc.) | 3812-32-6 | Major Ions; | Major Ions; EMS code: CO3- | Chemical |  | LAB
CO60 | Bq/L | Cobalt-60 | X240 | Gamma Spectrum | Gamma Spectrum | 0.001 | Y | Y | Y | 123 |  |  | 10/24/1996 7:20:15 AM | 1996-04-12 | 6/3/1998 6:51:34 PM | EMS |  | Radioactivity | Bq/L | Y | (rad./vol.) |  | Cobalt-60 (rad./vol.) | 10198-40-0 | Radionuclide; | Radionuclide; EMS code: CO60 | Chemical |  | LAB
COD | mg/L | Codeine CAS: 76-57-3 | XMSP | HPLC and Tri. Quadpole MS/MS | EPA 1694 An aliquot of filtered sample containing 5% organic is injected directly. Each compound is separated by reversed phase HPLC. Analyzed by Electro Spray Ionization MS/MS detection using Triple Quadrupole MS/MS detector. Quantiitation with internal or external standard. |  | Y | Y | Y | 12 | 5/8/2020 | 6/4/2024 | 8/21/2024 3:32:03 PM | 2022-05-04 |  | EMS |  | Concentration | mg/L | Y | (fl. conc.) |  | Codeine (fl. conc.) | 76-57-3 | Organics; | Organics; EMS code: COD | Chemical |  | LAB
CODS | mg/L | Chemical Oxygen Demand - Soluble | X504 | Closed refux,K2Cr2O7 dig,col'm,600nm | Organic and oxidizable inorganic substances are oxidized by potassium dichromate in H2SO4 solution at reflux temperature for 2 hours in a closed reflux. Oxygen consumption is measured colorimetrically against standards at 600nm. | 10 | Y | Y | Y | 2881 | 6/5/2014 1:22:00 PM | 9/15/2016 11:23:00 AM | 7/25/2024 2:49:05 PM | 2014-05-29 | 5/30/2014 10:54:34 AM | RDGIBSON | RDGIBSON | Concentration | mg/L | Y | (fl. conc.) |  | Chemical Oxygen Demand - Soluble (fl. conc.) |  | Aggregate Organics; | Aggregate Organics; EMS code: CODS | Chemical | Soluble | LAB
CODS | mg/L | Chemical Oxygen Demand - Soluble | X315 | Open reflux,K2Cr2O7/H2SO4 dig, FAS Titr. | Organic and oxidizable inorganic substances are oxidized by potassium dichromate in H2SO4 solution at reflux temperature for 2 hours. Excess dichromate is titrated with standard ferrous ammonium sulfate (0.1M) using orthophenanthroline ferrous complex (ferroin) as indicator. Open Reflux. | 2 | Y | Y | Y | 265 | 5/11/2011 11:27:00 AM | 5/23/2014 1:45:00 PM | 5/28/2014 1:51:10 PM | 2011-05-25 |  | RDGIBSON |  | Concentration | mg/L | Y | (fl. conc.) |  | Chemical Oxygen Demand - Soluble (fl. conc.) |  | Aggregate Organics; | Aggregate Organics; EMS code: CODS | Chemical | Soluble | LAB
COND-L | uS/cm | Conductivity; leachable | E103 | MEND (mod)/APHA 2510 (mod) | MEND (mod)/APHA 2510 (mod): Conductivity; also known as Electrical Conductivity (EC) or Specific Conductance; is measured by immersion of a conductivity cell with platinum electrodes into an extract from a soil sample that has been added in a 3:1 ratio of deionized water to soil; then leached and the extract filtered prior to analysis. Conductivity measurements are temperature-compensated to 25C. | 2 | Y | Y | Y | 0 |  |  |  | 2024-02-27 |  | RDGIBSON |  | ElectricConductivity | µS/cm | Y | (elec. cond.) |  | Conductivity; leachable (elec. cond.) |  | Physical Properties; | Physical Properties; EMS code: COND-L | Physical | Leachable | LAB
COND-L | uS/cm | Conductivity; leachable | E101 | .Conductivity (1:10 Soil:Water Extraction) | Conductivity (1:10 Soil:Water Extraction). Conductivity, also known as Electrical Conductivity (EC) or Specific Conductance, is measured by immersion of a conductivity cell with platinum electrodes into a soil sample that has been added in a defined ratio of soil to deionized water, then leached prior to analysis. Conductance is measured in the fluid that is observed in the upper layer. | 10 | Y | Y | Y | 0 |  |  |  | 2024-05-24 |  | RDGIBSON |  | ElectricConductivity | µS/cm | Y | (elec. cond.) |  | Conductivity; leachable (elec. cond.) |  | Physical Properties; | Physical Properties; EMS code: COND-L | Physical | Leachable | LAB
COT | mg/L | Cotinine CAS: 486-56-6 | XMSP | HPLC and Tri. Quadpole MS/MS | EPA 1694 An aliquot of filtered sample containing 5% organic is injected directly. Each compound is separated by reversed phase HPLC. Analyzed by Electro Spray Ionization MS/MS detection using Triple Quadrupole MS/MS detector. Quantiitation with internal or external standard. |  | Y | Y | Y | 12 | 5/8/2020 | 6/4/2024 | 8/21/2024 3:32:03 PM | 2022-05-04 |  | EMS |  | Concentration | mg/L | Y | (fl. conc.) |  | Cotinine (fl. conc.) | 486-56-6 | Organics; | Organics; EMS code: COT | Chemical |  | LAB
CP-T | mg/dm2/d | Particul: Total Combustible | CAL1 | Calculation (mg/dm2/d) | Calculation (mg/dm2/d) | 0.1 | Y | Y | Y | 517 |  |  | 10/20/1997 9:36:18 AM | 1996-04-12 | 6/3/1998 6:51:41 PM | EMS |  | MassFlux | mg/dm²/d | Y | (mass flux) |  | Particul: Total Combustible (mass flux) |  | Particulates; | Particulates; EMS code: CP-T | Chemical | Total | LAB
CP-T | mg/dm2/d | Particul: Total Combustible | EQ15 | EQUIS (mg/dm2/d) | EQUIS (mg/dm2/d) | 0 | N | Y | Y | 2719 |  |  | 10/23/1996 12:06:46 AM | 1996-04-12 | 6/3/1998 6:51:34 PM | EMS |  | MassFlux | mg/dm²/d | Y | (mass flux) |  | Particul: Total Combustible (mass flux) |  | Particulates; | Particulates; EMS code: CP-T | Chemical | Total | LAB
CP-T | mg/dm2/d | Particul: Total Combustible | CS15 | System Calculation | System Calculation (mg/dm2/d) |  | N | Y | N | 8 |  |  | 2/4/1998 5:55:18 PM | 1996-04-12 | 6/3/1998 6:51:43 PM | EMS |  | MassFlux | mg/dm²/d | Y | (mass flux) |  | Particul: Total Combustible (mass flux) |  | Particulates; | Particulates; EMS code: CP-T | Chemical | Total | LAB
CP-T | ug/m3 | Particul: Total Combustible | EQ40 | EQUIS (mg/m3) | EQUIS (mg/m3) | 2 | N | Y | Y | 49 |  |  | 10/22/1996 9:14:54 PM | 1996-04-12 | 6/3/1998 6:51:34 PM | EMS |  | AirConcentration | µg/m³ | Y | (air conc.) |  | Particul: Total Combustible (air conc.) |  | Particulates; | Particulates; EMS code: CP-T | Chemical | Total | LAB
CP-T | mg/dm2/d | Particul: Total Combustible | X192 | Gravimetric (Sol+Insol) | Gravimetric (Sol+Insol) | 0.1 | Y | Y | Y | 2777 |  |  | 7/11/2005 5:41:46 PM | 1996-04-12 | 6/3/1998 6:51:34 PM | EMS |  | MassFlux | mg/dm²/d | Y | (mass flux) |  | Particul: Total Combustible (mass flux) |  | Particulates; | Particulates; EMS code: CP-T | Chemical | Total | LAB
CP07 | mg/L | 3;4-Dichlorophenol | F028 | Acid,Extr.MeCl,HRGC/LRMS:TIM,SIM | A measured volume of sample is acidified and extracted with Methylene Chloride. For samples which have a high organic content that potentially produce chromatographic interferences an initial pre-extraction under basic conditions is performed. The sample extract is then evaporated to small volume and analysed by High Resolution Gas Chromatography / Low Resolution Mass Spectrometry (HRGC/LRMS). LRMS data is acquired in Total Ion Mode (TIM) and quantitation is performed using a single mass ion. Positive Identification of the various target compounds is made by comparison of absolute retention times to those of standards combined with computer based mass spectral mass/intensity library searching and matching routines. | 0.0001 | Y | Y | Y | 28 | 8/20/2009 | 11/1/2022 4:39:00 AM | 4/21/2023 4:26:45 PM | 2009-12-23 |  | RDGIBSON |  | Concentration | mg/L | Y | (fl. conc.) |  | 3;4-Dichlorophenol (fl. conc.) |  | Organics; | Organics; EMS code: CP07 | Chemical |  | LAB
CP07 | ug/g | 3;4-Dichlorophenol | 1332 | Phenolics by Tumbler/GC-MS | Phenolics by Tumbler/GC-MS | 0.02 | Y | Y | Y | 30 | 9/9/2019 | 9/22/2022 8:16:00 PM | 1/11/2024 1:07:07 PM | 2019-09-30 |  | KATFRASE |  | MassPerMass | mg/kg | Y | (mass conc.) |  | 3;4-Dichlorophenol (mass conc.) |  | Organics; | Organics; EMS code: CP07 | Chemical |  | LAB
CP07 | mg/L | 3;4-Dichlorophenol | X268 | CH2Cl2 ext:GC/MS | CH2Cl2 ext:GC/MS | 0.0003 | Y | Y | Y | 26 | 12/28/2021 | 6/7/2024 | 6/26/2024 4:28:10 PM | 2012-06-25 |  | RDGIBSON |  | Concentration | mg/L | Y | (fl. conc.) |  | 3;4-Dichlorophenol (fl. conc.) |  | Organics; | Organics; EMS code: CP07 | Chemical |  | LAB
CP25 | ug/g | 3+4-Chlorophenol | 1332 | Phenolics by Tumbler/GC-MS | Phenolics by Tumbler/GC-MS | 0.02 | Y | Y | Y | 7 | 9/22/2022 5:44:00 PM | 9/22/2022 8:16:00 PM | 3/9/2023 4:55:51 PM | 2022-09-29 |  | EMS |  | MassPerMass | mg/kg | Y | (mass conc.) |  | 3+4-Chlorophenol (mass conc.) | 108-43-0, 106-48-9 | Organics; | Organics; EMS code: CP25 | Chemical |  | LAB
CP25 | mg/L | 3+4-Chlorophenol | F028 | Acid,Extr.MeCl,HRGC/LRMS:TIM,SIM | A measured volume of sample is acidified and extracted with Methylene Chloride. For samples which have a high organic content that potentially produce chromatographic interferences an initial pre-extraction under basic conditions is performed. The sample extract is then evaporated to small volume and analysed by High Resolution Gas Chromatography / Low Resolution Mass Spectrometry (HRGC/LRMS). LRMS data is acquired in Total Ion Mode (TIM) and quantitation is performed using a single mass ion. Positive Identification of the various target compounds is made by comparison of absolute retention times to those of standards combined with computer based mass spectral mass/intensity library searching and matching routines. | 0.0001 | Y | Y | Y | 28 | 8/20/2009 | 11/1/2022 4:39:00 AM | 4/21/2023 4:26:45 PM | 2009-12-23 |  | RDGIBSON |  | Concentration | mg/L | Y | (fl. conc.) |  | 3+4-Chlorophenol (fl. conc.) | 108-43-0, 106-48-9 | Organics; | Organics; EMS code: CP25 | Chemical |  | LAB
CP37 | mg/L | 2;4+2;5-Dichlorophenol | X268 | CH2Cl2 ext:GC/MS | CH2Cl2 ext:GC/MS | 0.0001 | Y | Y | Y | 26 | 12/28/2021 | 6/7/2024 | 6/26/2024 4:28:10 PM | 2012-06-25 |  | RDGIBSON |  | Concentration | mg/L | Y | (fl. conc.) |  | 2;4+2;5-Dichlorophenol (fl. conc.) | 120-83-2, 583-78-8 | Organics; | Organics; EMS code: CP37 | Chemical |  | LAB
CP37 | mg/L | 2;4+2;5-Dichlorophenol | F028 | Acid,Extr.MeCl,HRGC/LRMS:TIM,SIM | A measured volume of sample is acidified and extracted with Methylene Chloride. For samples which have a high organic content that potentially produce chromatographic interferences an initial pre-extraction under basic conditions is performed. The sample extract is then evaporated to small volume and analysed by High Resolution Gas Chromatography / Low Resolution Mass Spectrometry (HRGC/LRMS). LRMS data is acquired in Total Ion Mode (TIM) and quantitation is performed using a single mass ion. Positive Identification of the various target compounds is made by comparison of absolute retention times to those of standards combined with computer based mass spectral mass/intensity library searching and matching routines. | 0.0001 | Y | Y | Y | 28 | 8/20/2009 | 11/1/2022 4:39:00 AM | 4/21/2023 4:26:45 PM | 2009-12-23 |  | RDGIBSON |  | Concentration | mg/L | Y | (fl. conc.) |  | 2;4+2;5-Dichlorophenol (fl. conc.) | 120-83-2, 583-78-8 | Organics; | Organics; EMS code: CP37 | Chemical |  | LAB
CP37 | ug/g | 2;4+2;5-Dichlorophenol | 1332 | Phenolics by Tumbler/GC-MS | Phenolics by Tumbler/GC-MS | 0.02 | Y | Y | Y | 30 | 9/9/2019 | 9/22/2022 8:16:00 PM | 1/11/2024 1:07:07 PM | 2019-09-30 |  | KATFRASE |  | MassPerMass | mg/kg | Y | (mass conc.) |  | 2;4+2;5-Dichlorophenol (mass conc.) | 120-83-2, 583-78-8 | Organics; | Organics; EMS code: CP37 | Chemical |  | LAB
CPF | mg/L | Cyclophosphamide CAS: 50-18-0 | XMSP | HPLC and Tri. Quadpole MS/MS | EPA 1694 An aliquot of filtered sample containing 5% organic is injected directly. Each compound is separated by reversed phase HPLC. Analyzed by Electro Spray Ionization MS/MS detection using Triple Quadrupole MS/MS detector. Quantiitation with internal or external standard. |  | Y | Y | Y | 12 | 5/8/2020 | 6/4/2024 | 8/21/2024 3:32:03 PM | 2022-05-04 |  | EMS |  | Concentration | mg/L | Y | (fl. conc.) |  | Cyclophosphamide (fl. conc.) | 50-18-0 | Organics; | Organics; EMS code: CPF | Chemical |  | LAB
CR-3 | mg/L | Chromium; Trivalent | XCAL | Calculation (mg/L) | Calculation (mg/L) | 0.005 | Y | Y | Y | 863 | 4/10/2015 | 8/15/2024 | 8/21/2024 1:05:08 PM | 1998-04-29 | 6/3/1998 6:51:48 PM | NPEPPIN |  | Concentration | mg/L | Y | (fl. conc.) |  | Chromium; Trivalent (fl. conc.) |  | Metals - Total; | Metals - Total; EMS code: CR-3 | Chemical |  | LAB
CR-6 | mg/L | Chromium; Hexavalent | 2112 | Col'metric:diphenylcarb. | The chromium VI in the sample is reacted with diphenylcarbazide in an acid solution to produce a red-violet colour. The intensity of the colour is measured at 540 nm. Sample is lab or field filtered and not acid preserved. | 0.002 | Y | Y | Y | 1706 | 2/20/2002 2:25:00 PM | 8/20/2024 12:00:00 PM | 8/23/2024 5:01:25 PM | 1998-04-29 | 6/3/1998 6:51:48 PM | NPEPPIN |  | Concentration | mg/L | Y | (fl. conc.) |  | Chromium; Hexavalent (fl. conc.) |  | Metals - Total; | Metals - Total; EMS code: CR-6 | Chemical |  | LAB
CR-D | mg/L | Chromium Dissolved | X285 | HNO3 Dig:ICP Analysis | HNO3 Dig:ICP Analysis | 0.015 | Y | Y | Y | 673 | 1/30/2002 | 11/13/2002 | 3/18/2004 10:28:34 AM | 1998-05-22 | 6/3/1998 6:51:49 PM | NPEPPIN |  | Concentration | mg/L | Y | (fl. conc.) |  | Chromium Dissolved (fl. conc.) | 7440-47-3 | Metals - Dissolved; | Metals - Dissolved; EMS code: CR-D | Chemical | Dissolved | LAB
CR-D | mg/L | Chromium Dissolved | X356 | LF,LA:HNO3:ICP | Samples are lab-filtered (0.45 um); lab-acidified with HNO3; analysis by ICP/AES. | 0.01 | Y | Y | Y | 397 |  |  | 3/23/2010 5:08:08 PM | 1996-04-12 | 6/3/1998 6:51:34 PM | EMS |  | Concentration | mg/L | Y | (fl. conc.) |  | Chromium Dissolved (fl. conc.) | 7440-47-3 | Metals - Dissolved; | Metals - Dissolved; EMS code: CR-D | Chemical | Dissolved | LAB
CR-D | mg/L | Chromium Dissolved | X556 | Metals in water, dissolved, "in-bottle" digestion, ICP-SFMS | The whole water sample is filtered through a 0.45 micron cellulose acetate filter paper and the filtrate is stored in a HDPE/Teflon bottle. The container is de-capped in a Class 100 workstation, acidified to 2% (v/v) HNO3 and 1% (v/v) HCl and recapped. The sample is "in-bottle" digested in a convection oven, set at 60 degrees C for 16 hours. The sample is cooled and introduced to an inductively coupled argon plasma-sector field mass spectrometer (ICP-SFMS) by pneumatic nebulization. Each respective element is measured at a specific mass to charge ratio, m/z value expressed in amu (atomic mass units), and at a specified intrument resolution. LR: Low Resolution ~300, MR: Mid Resolution ~4000 and HR: High Resolution ~10000. Ref: Method number 02-2003, National Laboratory for Environmental Testing, CCIW, Environment Canada, Burlington, Requested April 2003. | 5e-06 | Y | Y | Y | 4233 | 6/11/2003 | 12/6/2023 2:04:00 AM | 2/14/2024 9:05:32 AM | 2004-01-19 |  | UMULLER |  | Concentration | mg/L | Y | (fl. conc.) |  | Chromium Dissolved (fl. conc.) | 7440-47-3 | Metals - Dissolved; | Metals - Dissolved; EMS code: CR-D | Chemical | Dissolved | LAB
CR-D | mg/L | Chromium Dissolved | METC | MET-D-NP-CCMS-VA | Ultra trace metals in water are analyzed by CRC ICPMS, based on US EPA Method 6020B (July 2014). This procedure is intended for pristine field-filtered acid-preserved water samples. The detection limits (LOR) for this test are based on lab instrumental analysis only, not including filtration. ALS-supplied field filtration equipment does not support these LOR. Therefore, because of the high probability of false positives due to filtration, it is strongly recommended that a filtration blank be analysed to aid in data interpretation. | 0.0001 | Y | Y | Y | 1776 | 4/3/2012 | 8/15/2024 | 8/23/2024 9:50:35 AM | 2019-08-20 |  | KATFRASE |  | Concentration | mg/L | Y | (fl. conc.) |  | Chromium Dissolved (fl. conc.) | 7440-47-3 | Metals - Dissolved; | Metals - Dissolved; EMS code: CR-D | Chemical | Dissolved | LAB
CR-D | mg/L | Chromium Dissolved | F023 | HNO3,GFAA | Dissolved and extractable water samples are ready for analysis once collected. Total samples are autoclaved with HNO3 and hydrochloric acid (HCl) in capped Teflon vials. Analysis is performed via Inductively Coupled Argon Plasma Atomic Emission Spectrometry (ICP-AES) (simultaneous multielement analysis), Graphite Furnace Atomic Absorption Spectrometry (GF-AAS) (low level analysis) and/or Flame Atomic Absorption Spectrometry (F-AAS). Note that extractable samples are acidified with 0.5 mL HNO3 per 100 mL of sample at time of sample collection. | 0.0005 | Y | Y | Y | 104 | 8/9/1996 1:38:00 PM | 11/22/2000 2:07:00 PM | 5/30/2001 1:16:39 AM | 1997-04-02 | 6/3/1998 6:51:46 PM | MPAK |  | Concentration | mg/L | Y | (fl. conc.) |  | Chromium Dissolved (fl. conc.) | 7440-47-3 | Metals - Dissolved; | Metals - Dissolved; EMS code: CR-D | Chemical | Dissolved | LAB
CR-D | mg/L | Chromium Dissolved | MWID | Metals in Water by ICPMS (digested) | Metals in Water by ICPMS (digested) Based on EPA 200.8 | 0.002 | Y | Y | Y | 31752 | 10/26/2018 5:20:00 PM | 10/26/2018 5:20:00 PM | 7/31/2024 1:46:41 PM | 2009-12-24 |  | RDGIBSON |  | Concentration | mg/L | Y | (fl. conc.) |  | Chromium Dissolved (fl. conc.) | 7440-47-3 | Metals - Dissolved; | Metals - Dissolved; EMS code: CR-D | Chemical | Dissolved | LAB
CR-D | mg/L | Chromium Dissolved | X562 | Dissolved metals: FF, FA, HNO3, HR-ICPMS | Ultra trace metals in water are analyzed by high resolution inductively coupled plasma mass spectrometry (HR-ICPMS), based on US EPA Method 200.8, (Oct. 1999). This procedure is intended for pristine field-filtered acid-preserved water samples. The detection limits (LOR) for this test are based on lab instrumental analysis only, not including filtration. ALS-supplied field filtration equipment does not support these LOR. Therefore, because of the high probability of false positives due to filtration, it is strongly recommended that a filtration blank be analyzed to aid in data interpretation. | 0.0001 | Y | Y | Y | 1264 | 6/8/2015 | 10/19/2023 | 11/9/2023 7:41:27 AM | 2015-10-22 |  | RDGIBSON |  | Concentration | mg/L | Y | (fl. conc.) |  | Chromium Dissolved (fl. conc.) | 7440-47-3 | Metals - Dissolved; | Metals - Dissolved; EMS code: CR-D | Chemical | Dissolved | LAB
CR-D | mg/L | Chromium Dissolved | X357 | LF,LA:HNO3:GFAA | Samples are lab-filtered (0.45 um) and lab-acidified with HNO3; analysis by GFAA. | 0.001 | Y | Y | Y | 25 | 1/22/1998 | 9/29/1998 | 10/30/1998 9:32:50 AM | 1996-04-12 | 6/3/1998 6:51:34 PM | EMS |  | Concentration | mg/L | Y | (fl. conc.) |  | Chromium Dissolved (fl. conc.) | 7440-47-3 | Metals - Dissolved; | Metals - Dissolved; EMS code: CR-D | Chemical | Dissolved | LAB
CR-D | mg/L | Chromium Dissolved | EQ01 | EQUIS (mg/L) | EQUIS (mg/L) | 0.005 | N | Y | Y | 2400 |  |  | 10/24/1996 3:16:55 PM | 1996-04-12 | 6/3/1998 6:51:34 PM | EMS |  | Concentration | mg/L | Y | (fl. conc.) |  | Chromium Dissolved (fl. conc.) | 7440-47-3 | Metals - Dissolved; | Metals - Dissolved; EMS code: CR-D | Chemical | Dissolved | LAB
CR-D | mg/L | Chromium Dissolved | F032 | Filt.,Acid,ICPMS | There is no acid digestion required for this analysis. A water sample is collected and preserved in 0.2 percent nitric acid. To analyse the extractable metals the sample is shaken, allowed to sit for at least twenty four hours, and a decanted portion is introduced to the spectrometer. When the dissolved metals are required, the samples must first be filtered through a 0.45 micron cellulose acetate filter paper and then followed by preservation in 0.2 percent nitric acid. It is important to note that different laboratories may have very different detection limits for some variables analyzed by ICPMS. | 0.001 | Y | Y | Y | 81337 | 10/15/1996 9:06:00 AM | 8/26/2024 | 8/26/2024 4:53:05 PM | 1997-01-16 | 8/15/2019 10:55:15 AM | MPAK | KATFRASE | Concentration | mg/L | Y | (fl. conc.) |  | Chromium Dissolved (fl. conc.) | 7440-47-3 | Metals - Dissolved; | Metals - Dissolved; EMS code: CR-D | Chemical | Dissolved | LAB
CR-D | mg/L | Chromium Dissolved | 0050 | Graphite Furnace AA | Graphite Furnace AA | 0.005 | Y | Y | Y | 724 | 8/21/2002 | 7/24/2003 | 8/6/2003 4:55:09 PM | 1996-04-12 | 6/3/1998 6:51:34 PM | EMS |  | Concentration | mg/L | Y | (fl. conc.) |  | Chromium Dissolved (fl. conc.) | 7440-47-3 | Metals - Dissolved; | Metals - Dissolved; EMS code: CR-D | Chemical | Dissolved | LAB
CR-D | mg/L | Chromium Dissolved | X170 | ICP | This code is a catch-all for a large variety of ICP methods including water and also solids digested into solutions as an intermediate result. Different procedures will have greatly different detection limits. (a) For air particulates, preliminary results are reported in mg/L and loading results as mg/dm2/d. (b) For Hi Vol samples preliminiary results are same as reported in mg/L and loading results as ug/m3. (c) Note that for Rain samples there is no filtration nor digestion steps. | 0.01 | Y | Y | Y | 12126 | 4/11/1996 6:01:00 PM | 11/30/2019 | 10/18/2023 8:55:31 AM | 1996-04-12 | 6/3/1998 6:51:34 PM | EMS |  | Concentration | mg/L | Y | (fl. conc.) |  | Chromium Dissolved (fl. conc.) | 7440-47-3 | Metals - Dissolved; | Metals - Dissolved; EMS code: CR-D | Chemical | Dissolved | LAB
CR-D | mg/L | Chromium Dissolved | 0040 | HNO3 Dig: ICP Analysis | HNO3 Dig: ICP Analysis | 0.01 | Y | Y | Y | 114 | 5/9/2003 9:40:00 AM | 9/16/2003 3:18:00 PM | 11/28/2003 3:22:11 PM | 1996-04-12 | 6/3/1998 6:51:34 PM | EMS |  | Concentration | mg/L | Y | (fl. conc.) |  | Chromium Dissolved (fl. conc.) | 7440-47-3 | Metals - Dissolved; | Metals - Dissolved; EMS code: CR-D | Chemical | Dissolved | LAB
CR-D | mg/L | Chromium Dissolved | F039 | Field filt + acid, ICP-AES | Dissolved and extractable water samples are ready for analysis once collected. Total samples are autoclaved with HNO3 and hydrochloric acid (HCI) in capped Teflon vials. Analysis is performed via Inductively Coupled Argon Plasma Atomic Emission Spectrometry (ICP-AES)(simultaneous multielement analysis), Graphite Furnace Atomic Absorption Spectrometry (GF-AAS)(low level analysis) and/or Flame Atomic Absorption Spectrometry (F-AAS). | 0.005 | Y | Y | Y | 5968 | 7/29/1996 6:03:00 PM | 7/5/2019 1:21:00 PM | 9/11/2019 8:47:50 AM | 1997-05-08 | 6/3/1998 6:51:46 PM | MPAK |  | Concentration | mg/L | Y | (fl. conc.) |  | Chromium Dissolved (fl. conc.) | 7440-47-3 | Metals - Dissolved; | Metals - Dissolved; EMS code: CR-D | Chemical | Dissolved | LAB
CR-D | mg/L | Chromium Dissolved | X085 | LF,LA,HNO3,AA | Samples are lab-filtered (0.45 um), and lab-acidified (HNO3); analysis by AA. | 0.005 | Y | Y | Y | 705 | 12/12/1996 | 9/18/2015 | 10/28/2015 9:54:16 AM | 1996-04-12 | 6/3/1998 6:51:41 PM | EMS |  | Concentration | mg/L | Y | (fl. conc.) |  | Chromium Dissolved (fl. conc.) | 7440-47-3 | Metals - Dissolved; | Metals - Dissolved; EMS code: CR-D | Chemical | Dissolved | LAB
CR-D | mg/L | Chromium Dissolved | X116 | FF,FA,HNO3,Graph.Furn,AA | Samples are field-filtered (0.45 um), field-acidified with HNO3; analysis by GFAA. | 0.0001 | Y | Y | Y | 6572 |  |  | 8/26/2024 2:41:32 PM | 1996-04-12 | 6/3/1998 6:51:34 PM | EMS |  | Concentration | mg/L | Y | (fl. conc.) |  | Chromium Dissolved (fl. conc.) | 7440-47-3 | Metals - Dissolved; | Metals - Dissolved; EMS code: CR-D | Chemical | Dissolved | LAB
CR-D | mg/L | Chromium Dissolved | X095 | FF,FA,HNO3,ICP | FF;FA;HNO3;ICP | 0.005 | Y | Y | Y | 190 |  |  | 10/24/1996 12:20:01 PM | 1996-04-12 | 6/3/1998 6:51:34 PM | EMS |  | Concentration | mg/L | Y | (fl. conc.) |  | Chromium Dissolved (fl. conc.) | 7440-47-3 | Metals - Dissolved; | Metals - Dissolved; EMS code: CR-D | Chemical | Dissolved | LAB
CR-D | mg/L | Chromium Dissolved | MSEA | Metals in Seawater (undigested) by (CRC) ICP-MS | This method describes the multi-elemental determination of trace elements by Collision/Reaction Cell (CRC) ICP-MS. The method measures ions produced by a radio-frequency inductively coupled plasma. Analyte species originating in a liquid are pneumatically nebulized and the resulting aerosol transported by argon gas into the plasma torch. The ions produced are entrained in the plasma gas and extracted, by means of a differentially pumped vacuum interface, into a mass spectrometer. The ions flow through a Collision/Reaction Cell (CRC) where depending on the mode of analysis either the ions passes through the cell unhindered (no gas - classical ICP-MS mode), or collisions/reactions chemistry takes place between the ions and a neutral gas such as hydrogen or helium. Any polyatomic ions which give interference on the analytes are removed. The ions are then sorted according to their mass-to-charge ratios by a quadrupole mass spectrometer. The ions are then sorted according to their mass-to-charge ratios by a quadrupole mass spectrometer. The ions transmitted through the quadrupole are quantified by a channel electron multiplier and the ion information processed by a data handling system. Although most polyatomic interference is removed, any additional interference must be assessed and valid corrections applied or the data flagged to indicate problems. Interference correction must include compensation for background ions contributed by the plasma gas, reagents, and constituents of the sample matrix. Instrument drift as well as suppressions or enhancements of instrument response caused by the sample matrix must be corrected for by the use of internal standards. | 0.0002 | Y | Y | Y | 4 | 4/12/2023 | 4/29/2024 | 5/16/2024 2:00:14 PM | 2009-09-08 |  | RDGIBSON |  | Concentration | mg/L | Y | (fl. conc.) |  | Chromium Dissolved (fl. conc.) | 7440-47-3 | Metals - Dissolved; | Metals - Dissolved; EMS code: CR-D | Chemical | Dissolved | LAB
CR-D | mg/L | Chromium Dissolved | X350 | FF,FA:HNO3:ICP | Samples are field-filtered (0.45 um); field-acidified with HNO3; analysis by ICP/AES. | 0.01 | Y | Y | Y | 30 |  |  | 10/30/2014 8:31:17 AM | 1996-04-12 | 6/3/1998 6:51:34 PM | EMS |  | Concentration | mg/L | Y | (fl. conc.) |  | Chromium Dissolved (fl. conc.) | 7440-47-3 | Metals - Dissolved; | Metals - Dissolved; EMS code: CR-D | Chemical | Dissolved | LAB
CR-D | mg/L | Chromium Dissolved | F068 | HNO3dig,TeflonTube/Autoclave,ICP-MS | This method is applicable to determiniation of dissolved, extractable, and total recoverable metals in ground waters, surface waters, drinking waters, and wastewaters. Dissolved, extractable, and total recoverable sample fractions are acidified to pH <1.5 with analytical grade nitric acid (HNO3) to prevent the metals from adsorbing to the container material or precipitating. Extractable and total recoverable water samples are collected in 250 ml acid washed polyethylene bottles and are preserved with Nitric acid (0.5 ml acid for 100 mLs of sample). All preserved samples are stored at room temperature and can be held for up to 6 months. Unfiltered, un-preserved samples are stored at 4C for no more than 30 days. For dissolved and extractable water fractions, no sample preparation is necessary. Samples are simply poured in autosampler tubes (i.e. 15 ml centrifuge tubes), and placed on the autosampler and analyzed. Samples for total recoverable metals are digested with Nitric acid in special Teflon tubes in an autoclave prior to analyses. | 0.0002 | Y | Y | Y | 1398 | 6/1/1999 1:20:00 PM | 4/17/2015 1:06:00 PM | 8/11/2021 12:04:11 PM | 1999-06-28 |  | NPEPPIN |  | Concentration | mg/L | Y | (fl. conc.) |  | Chromium Dissolved (fl. conc.) | 7440-47-3 | Metals - Dissolved; | Metals - Dissolved; EMS code: CR-D | Chemical | Dissolved | LAB
CR-D | mg/L | Chromium Dissolved | F049 | Filt,HNO3 dig,conc'n,nebul'ICP-OES | This method provides an automated procedure for the Inductively Coupled Argon Plasma Optical Emission Spectrophotometric (ICP-OES) determination and quantitation of trace amounts of metals in surface, ground and precipitation waters after a manual acid digestion and concentration procedure and ultrasonic nebulisation. Organically-bound metals and metal complexes are converted to metal ions by digestion with nitric acid. To increase the signal detected by the ICP, the samples are concentrated 10 times. The liquid sample is pumped onto the piezoelectric transducer, where the ultrasonic action efficiently converts it to a fine, dense aerosol. The nebulizer carrier gas flow then transports the wet aerosol through a heated U-tube, where the solvent is vaporized. Subsequently, most of the solvent vapour is removed from the aerosol stream by passing the stream through a chilled condenser. The result is a dry, analyte-laden aerosol which is introduced to the plasma. There, the high temperature desolvates the aerosol and dissociates the sample material into individual atoms and ions, which are excited to emit light at wave lengths characteristic of the elements present. Dissolved samples are field filtered through 0.45 micron cellulose acetate filter prior to acid preservation. Samples for extractable analysis are acidified with 2 mL concentrated HNO3 per litre and allowed to stand a minimum of 24 hours before analysis. Clients providing these samples are advised to preserve samples at the collection site. | 0.0002 | Y | Y | Y | 1420 |  |  | 5/19/2015 3:10:33 PM | 1997-10-20 | 6/3/1998 6:51:47 PM | NPEPPIN |  | Concentration | mg/L | Y | (fl. conc.) |  | Chromium Dissolved (fl. conc.) | 7440-47-3 | Metals - Dissolved; | Metals - Dissolved; EMS code: CR-D | Chemical | Dissolved | LAB
CR-E | mg/L | Chromium Extractable | F040 | Field Acid, lab filt, ICP-AES | Dissolved and extractable water samples are ready for analysis once collected. Total samples are autoclaved with HNO3 and hydrochloric acid (HCI) in capped Teflon vials. Analysis is performed via Inductively Coupled Argon Plasma Atomic Emission Spectrometry (ICP-AES)(simultaneous multielement analysis), Graphite Furnace Atomic Absorption Spectrometry (GF-AAS)(low level analysis) and/or Flame Atomic Absorption Spectrometry (F-AAS). | 0.005 | Y | Y | Y | 2361 | 8/5/1997 6:06:00 PM | 3/14/2008 | 9/11/2019 8:47:41 AM | 1997-05-08 | 6/3/1998 6:51:46 PM | MPAK |  | Concentration | mg/L | Y | (fl. conc.) |  | Chromium Extractable (fl. conc.) | 7440-47-3 | Metals - Extractable; | Metals - Extractable; EMS code: CR-E | Chemical | Extractable | LAB
CR-E | ug/g | Chromium Extractable | X230 | Hom,HNO3/HCl,ICP | The sample is dried, ground to pass 100 mesh, and digested with a nitric-hydrochloric acid mixture to solubilize the solid matter and to remove the organic material by oxidation and volatilization. The completed digests are submitted to metals instrumentation personnel for metals analysis by Inductively Coupled Atomic Emission Spectrophotometry (ICP-AES), or hydride analysis by atomic absorption. | 0.5 | Y | Y | Y | 74 | 12/13/2011 12:02:00 PM | 12/13/2011 12:02:00 PM | 2/9/2012 12:15:56 PM | 1996-04-12 | 6/3/1998 6:51:41 PM | EMS |  | MassPerMass | mg/kg | Y | (mass conc.) |  | Chromium Extractable (mass conc.) | 7440-47-3 | Metals - Extractable; | Metals - Extractable; EMS code: CR-E | Chemical | Extractable | LAB
CR-E | mg/L | Chromium Extractable | F069 | SEM extr,6N HCl,ICP/OES | 6.0 N Hydrochloric acid (HCI) is added to an aliquot of the sediment sample. The extract produced from the addition of the acid is then analyzed for simultaneously extracted metals (SEM) using inductively coupled plasma - optical emission spectrophotometry. The detection limits depend on the metal parameters in question and the instrumentation used. | 0.01 | Y | Y | Y | 10 | 3/26/2012 3:01:00 AM | 3/26/2012 3:01:00 AM | 11/25/2021 1:50:15 PM | 1999-07-09 |  | NPEPPIN |  | Concentration | mg/L | Y | (fl. conc.) |  | Chromium Extractable (fl. conc.) | 7440-47-3 | Metals - Extractable; | Metals - Extractable; EMS code: CR-E | Chemical | Extractable | LAB
CR-E | mg/L | Chromium Extractable | X553 | Metals in water, extractable, ICP-SFMS | The whole water sample is collected in a HDPE/Teflon bottle. In a Class 100 workstation it is preserved with concentrated Nitric Acid at concentration of 0.2 percent (v/v). The acid preservative extracts the metals from material in the sample and a decanted portion is introduced directly to an inductively coupled argon plasma-sector field mass spectrometer (ICP-SFMS) by pneumatic nebulization. Each respective element is measured at a specific mass to charge ratio, m/z value expressed in amu (atomic mass units), and at a specified instrument resolution. LR: Low Resolution ~300, MR: Mid Resolution ~4000 and HR: High Resolution ~10000. Ref: Method number 02-2061, National Laboratory for Environmental Testing, CCIW, Environment Canada, Burlington. Requested April 2003. | 5e-06 | Y | Y | Y | 2597 | 5/22/2003 | 9/21/2007 | 11/28/2023 2:11:53 PM | 2004-01-19 |  | UMULLER |  | Concentration | mg/L | Y | (fl. conc.) |  | Chromium Extractable (fl. conc.) | 7440-47-3 | Metals - Extractable; | Metals - Extractable; EMS code: CR-E | Chemical | Extractable | LAB
CR-E | mg/L | Chromium Extractable | F047 | 24hr HNO3 dil'n,ICP-MS | There is no acid digestion required for this analysis. A water sample is collected and preserved in 0.2 percent nitric acid. To analyse the extractable metals the sample is shaken, allowed to sit for at least twenty four hours, and a decanted portion is introduced to the spectrometer. When the dissolved metals are required, the samples must first be filtered through a 0.45 micron cellulose acetate filter paper and then followed by preservation in 0.2 percent nitric acid. It is important to note that different laboratories may have very different detection limits for some variables analyzed by ICP/MS. | 2e-05 | Y | Y | Y | 2676 | 4/16/1998 10:21:00 AM | 6/26/2003 | 9/11/2019 8:46:17 AM | 1997-10-20 | 6/3/1998 6:51:47 PM | NPEPPIN |  | Concentration | mg/L | Y | (fl. conc.) |  | Chromium Extractable (fl. conc.) | 7440-47-3 | Metals - Extractable; | Metals - Extractable; EMS code: CR-E | Chemical | Extractable | LAB
CR-T | mg/L | Chromium Total | MSEA | Metals in Seawater (undigested) by (CRC) ICP-MS | This method describes the multi-elemental determination of trace elements by Collision/Reaction Cell (CRC) ICP-MS. The method measures ions produced by a radio-frequency inductively coupled plasma. Analyte species originating in a liquid are pneumatically nebulized and the resulting aerosol transported by argon gas into the plasma torch. The ions produced are entrained in the plasma gas and extracted, by means of a differentially pumped vacuum interface, into a mass spectrometer. The ions flow through a Collision/Reaction Cell (CRC) where depending on the mode of analysis either the ions passes through the cell unhindered (no gas - classical ICP-MS mode), or collisions/reactions chemistry takes place between the ions and a neutral gas such as hydrogen or helium. Any polyatomic ions which give interference on the analytes are removed. The ions are then sorted according to their mass-to-charge ratios by a quadrupole mass spectrometer. The ions are then sorted according to their mass-to-charge ratios by a quadrupole mass spectrometer. The ions transmitted through the quadrupole are quantified by a channel electron multiplier and the ion information processed by a data handling system. Although most polyatomic interference is removed, any additional interference must be assessed and valid corrections applied or the data flagged to indicate problems. Interference correction must include compensation for background ions contributed by the plasma gas, reagents, and constituents of the sample matrix. Instrument drift as well as suppressions or enhancements of instrument response caused by the sample matrix must be corrected for by the use of internal standards. | 0.0002 | Y | Y | Y | 441 | 8/12/2009 12:08:00 AM | 4/29/2024 | 6/7/2024 10:03:16 AM | 2009-09-08 |  | RDGIBSON |  | Concentration | mg/L | Y | (fl. conc.) |  | Chromium Total (fl. conc.) | 7440-47-3 | Metals - Total; | Metals - Total; EMS code: CR-T | Chemical | Total | LAB
CR-T | mg/L | Chromium Total | X555 | Metals in water, total, "in-bottle" digestion, ICP-SFMS | The whole water sample collected in a HDPE/Teflon bottle, is de-capped in a Class 100 workstation, acidified to 2% (v/v) HNO3 and 1% (v/v) HCl and recapped. The sample is "in-bottle" digested in a convection oven, set at 60 degrees C for 16 hours. The sample is cooled and introduced to an inductively coupled argon plasma-sector field mass spectrometer (ICP-SFMS) by pneumatic nebulization. Each respective element is measured at a specific mass to charge ratio, m/z value expressed in amu (atomic mass units), and at resolution. LR: Low Resolution ~300, MR: Mid Resolution ~4000 and HR: High a specified instrument Resolution ~10000. Ref: Method number 02-2003, National Laboratory for Environmental Testing, CCIW, Environment Canada, Burlington. Requested April 2003. | 5e-06 | Y | Y | Y | 17343 | 5/20/2003 | 12/6/2023 2:08:00 AM | 2/14/2024 9:05:32 AM | 2004-01-19 |  | UMULLER |  | Concentration | mg/L | Y | (fl. conc.) |  | Chromium Total (fl. conc.) | 7440-47-3 | Metals - Total; | Metals - Total; EMS code: CR-T | Chemical | Total | LAB
CR-T | mg/L | Chromium Total | TCLP | Toxicity Characteristic Leaching Procedure | The TCLP procedure is designed to simulate the climatic leaching action | 0.1 | Y | Y | Y | 38 | 12/8/2009 7:49:00 PM | 9/28/2011 5:53:00 PM | 8/24/2016 2:59:13 PM | 2009-03-05 |  | JCHUDAK |  | Concentration | mg/L | Y | (fl. conc.) |  | Chromium Total (fl. conc.) | 7440-47-3 | Metals - Total; | Metals - Total; EMS code: CR-T | Chemical | Total | LAB
CR-T | mg/L | Chromium Total | 0095 | %5 A.R. Dig:HGA Analysis | %5 A.R. Dig:HGA Analysis | 0.005 | Y | Y | Y | 37 |  |  | 10/24/1996 12:38:16 PM | 1996-04-12 | 6/3/1998 6:51:41 PM | EMS |  | Concentration | mg/L | Y | (fl. conc.) |  | Chromium Total (fl. conc.) | 7440-47-3 | Metals - Total; | Metals - Total; EMS code: CR-T | Chemical | Total | LAB
CR-T | ug/g | Chromium Total | EQ02 | EQUIS (ug/g) | EQUIS (ug/g) | 1 | N | Y | Y | 799 |  |  | 10/23/1996 12:08:28 AM | 1996-04-12 | 6/3/1998 6:51:35 PM | EMS |  | MassPerMass | mg/kg | Y | (mass conc.) |  | Chromium Total (mass conc.) | 7440-47-3 | Metals - Total; | Metals - Total; EMS code: CR-T | Chemical | Total | LAB
CR-T | ug/g | Chromium Total | PB02 | PBM,HNO3/HCl dig,dry60C,100mesh, | Performance Based Method. Samples are dried at 60 C and sieved using a .149 mm (100 mesh) sieve. Solid samples are digested with the combination of nitric and hydrochloric acids with a 1:1 ratio of acids. For environmental purposes, strong acid leaches are used routinely since they have the advantage of speed, low cost, and superio reproducibility. All metals trapped within the refractory matrix are unlikely to have environmental consequences. Instrumental analysis of extracts produced using this digestion method can be performed by a variety of techniques. |  | Y | Y | Y | 52 | 5/24/2001 3:42:00 PM | 11/26/2001 10:23:00 AM | 5/26/2003 2:27:46 PM | 2001-07-18 |  | NPEPPIN |  | MassPerMass | mg/kg | Y | (mass conc.) |  | Chromium Total (mass conc.) | 7440-47-3 | Metals - Total; | Metals - Total; EMS code: CR-T | Chemical | Total | LAB
CR-T | mg/L | Chromium Total | F024 | HNO3/HCl,GFAA | Dissolved and extractable water samples are ready for analysis once collected. Total samples are autoclaved with HNO3 and hydrochloric acid (HCl) in capped Teflon vials. Analysis is performed via Inductively Coupled Argon Plasma Atomic Emission Spectrometry (ICP-AES) (simultaneous multielement analysis), Graphite Furnace Atomic Absorption Spectrometry (GF-AAS) (low level analysis) and/or Flame Atomic Absorption Spectrometry (F-AAS). | 0.0005 | Y | Y | Y | 165 | 9/3/1996 10:53:00 AM | 11/22/2000 2:07:00 PM | 5/30/2001 1:16:40 AM | 1997-04-02 | 6/3/1998 6:51:46 PM | MPAK |  | Concentration | mg/L | Y | (fl. conc.) |  | Chromium Total (fl. conc.) | 7440-47-3 | Metals - Total; | Metals - Total; EMS code: CR-T | Chemical | Total | LAB
CR-T | mg/L | Chromium Total | X285 | HNO3 Dig:ICP Analysis | HNO3 Dig:ICP Analysis | 0.015 | Y | Y | Y | 736 | 8/10/1993 | 10/5/1993 | 4/30/2004 1:59:45 PM | 1996-04-12 | 6/3/1998 6:51:35 PM | EMS |  | Concentration | mg/L | Y | (fl. conc.) |  | Chromium Total (fl. conc.) | 7440-47-3 | Metals - Total; | Metals - Total; EMS code: CR-T | Chemical | Total | LAB
CR-T | ug/g | Chromium Total | 0190 | D&G:N/P Dig:ICP Anal. | Sample is dried and grind-mixed before nitric perchloric digestion followed by ICP analysis. | 1 | Y | Y | Y | 949 |  |  | 10/24/1996 3:06:16 PM | 1996-04-12 | 6/3/1998 6:51:34 PM | EMS |  | MassPerMass | mg/kg | Y | (mass conc.) |  | Chromium Total (mass conc.) | 7440-47-3 | Metals - Total; | Metals - Total; EMS code: CR-T | Chemical | Total | LAB
CR-T | ppm (S) | Chromium Total | MS61 | ALS Minerals ME-MS61 | ALS Minerals ME-MS61. A prepared sample is digested with HNO3; HF; and HClO4 in a graphite heating block to near dryness. After cooling; the sample is boiled in diluted HCl. Bulk to volume with deionized water; mixed and analyzed by ICP-AES and ICP-MS for all elements. | 1 | Y | Y | Y | 0 |  |  |  | 2024-01-29 |  | EMS |  | DimensionlessRatio | ppm | N | (ratio) |  | Chromium Total (ratio) | 7440-47-3 | Metals - Total; | Metals - Total; EMS code: CR-T | Chemical | Total | LAB
CR-T | ug/m3 | Chromium Total | F088 | Aliquot,HNO3/HCl dig, ICPMS,loading | Aliquot microwave digested with 3:1 HNO3/HCL. Analysis by ICPMS. For HiVol samples preliminary result are reported in mg/L and loading results are reported in ug/m3 |  | Y | Y | Y | 565 | 8/17/2000 10:46:00 AM | 12/14/2006 | 12/30/2006 1:01:36 AM | 2000-08-16 |  | NPEPPIN |  | AirConcentration | µg/m³ | Y | (air conc.) |  | Chromium Total (air conc.) | 7440-47-3 | Metals - Total; | Metals - Total; EMS code: CR-T | Chemical | Total | LAB
CR-T | ug/g | Chromium Total | MTID | MET-DRY-CCMS-VA | Metals in Tissue by CRC ICPMS (DRY) Methodology: This method is conducted following British Columbia Lab Manual method "Metals in Animal Tissue and Vegetation (Biota) - Prescriptive". Tissue samples are homogenized and sub-sampled prior to hotblock digestion with nitric and hydrochloric acids, in combination with addition of hydrogen peroxide. Instrumental analysis is by collision cell inductively coupled plasma - mass spectrometry (modified from EPA Method 6020A). Method Limitation: This method employs a strong acid/peroxide digestion, and is intended to provide a conservative estimate of bio-available metals. Near complete recoveries are achieved for most toxicologically important metals, but elements associated with recalcitrant minerals may be only partially recovered. | 0.05 | Y | Y | Y | 122 | 8/13/2014 | 9/13/2022 | 10/26/2022 7:14:14 AM | 2015-03-09 | 3/11/2015 4:12:05 PM | RDGIBSON | EMS | MassPerMass | mg/kg | Y | (mass conc.) |  | Chromium Total (mass conc.) | 7440-47-3 | Metals - Total; | Metals - Total; EMS code: CR-T | Chemical | Total | LAB
CR-T | mg/L | Chromium Total | X170 | ICP | This code is a catch-all for a large variety of ICP methods including water and also solids digested into solutions as an intermediate result. Different procedures will have greatly different detection limits. (a) For air particulates, preliminary results are reported in mg/L and loading results as mg/dm2/d. (b) For Hi Vol samples preliminiary results are same as reported in mg/L and loading results as ug/m3. (c) Note that for Rain samples there is no filtration nor digestion steps. | 0.01 | Y | Y | Y | 6400 | 3/28/1996 6:03:00 PM | 11/7/2018 | 10/24/2022 10:58:25 AM | 1996-04-12 | 6/3/1998 6:51:35 PM | EMS |  | Concentration | mg/L | Y | (fl. conc.) |  | Chromium Total (fl. conc.) | 7440-47-3 | Metals - Total; | Metals - Total; EMS code: CR-T | Chemical | Total | LAB
CR-T | mg/L | Chromium Total | F038 | HNO3/HCl dig,ICP-AES | Dissolved and extractable water samples are ready for analysis once collected. Total samples are autoclaved with HNO3 and hydrochloric acid (HCI) in capped Teflon vials. Analysis is performed via Inductively Coupled Argon Plasma Atomic Emission Spectrometry (ICP-AES)(simultaneous multielement analysis), Graphite Furnace Atomic Absorption Spectrometry (GF-AAS) (low level analysis) and/or Flame Atomic Absorption Spectrometry (F-AAS). | 0.006 | Y | Y | Y | 5429 | 8/1/1996 6:04:00 PM | 5/6/2008 | 9/11/2019 8:47:42 AM | 1997-05-08 | 6/3/1998 6:51:46 PM | MPAK |  | Concentration | mg/L | Y | (fl. conc.) |  | Chromium Total (fl. conc.) | 7440-47-3 | Metals - Total; | Metals - Total; EMS code: CR-T | Chemical | Total | LAB
CR-T | mg/L | Chromium Total | X179 | LA,HNO3,Graphite Furnace | Samples are lab-acidified with HNO3; HNO3 digestion; analysis by GFAA. | 0.001 | Y | Y | Y | 1 |  |  | 6/16/2000 3:49:03 PM | 1996-04-12 | 6/3/1998 6:51:41 PM | EMS |  | Concentration | mg/L | Y | (fl. conc.) |  | Chromium Total (fl. conc.) | 7440-47-3 | Metals - Total; | Metals - Total; EMS code: CR-T | Chemical | Total | LAB
CR-T | ug/g wet | Chromium Total | X560 | Metals in Biota (wet), total, acid digestion, ICP-MS | Five grams of wet, homogenized biota is weighed in a centrifuge tube, and 2.5 mL of Nitric Acid and 2.5 mL of Hydrochloric acid are added. The sample is digested at room temperature for one hour and then digested at 90-95 degrees C in a water bath or heat block for one more hour. The sample is cooled and the volume is made to 50 mL with DI. The digestate is then diluted at ten fold and introduced to a collision/reaction cell inductively coupled argon plasma mass spectrometer (CRC ICPMS) by pneumatic nebulization. Each respective element is measured at a specific mass to charge ratio, m/z value expressed in amu (atomic mass units). - Maxxam Analytics (BC). Added to EMS Dec 10, 2008. | 0.5 | Y | Y | Y | 541 | 11/15/2008 4:00:00 AM | 10/30/2023 | 11/28/2023 2:21:33 PM | 2008-12-11 | 3/11/2015 4:17:10 PM | RDGIBSON | EMS | MassPerMass | µg/g | Y | (mass conc.) | (wet) | Chromium Total (wet) (mass conc.) | 7440-47-3 | Metals - Total; | Metals - Total; EMS code: CR-T | Chemical | Total | LAB
CR-T | mg/L | Chromium Total | 0061 | HNO3 Dig: GFAA Analysis | HNO3 Dig: GFAA Analysis | 0.001 | Y | Y | Y | 2 | 6/25/2010 3:15:00 PM | 6/25/2010 3:15:00 PM | 7/22/2010 2:19:37 PM | 1998-12-08 |  | NPEPPIN |  | Concentration | mg/L | Y | (fl. conc.) |  | Chromium Total (fl. conc.) | 7440-47-3 | Metals - Total; | Metals - Total; EMS code: CR-T | Chemical | Total | LAB
CR-T | mg/L | Chromium Total | 0040 | HNO3 Dig: ICP Analysis | HNO3 Dig: ICP Analysis | 0.01 | Y | Y | Y | 14271 | 5/8/2003 9:40:00 AM | 9/12/2003 3:16:00 PM | 10/18/2023 8:55:31 AM | 1996-04-12 | 6/3/1998 6:51:34 PM | EMS |  | Concentration | mg/L | Y | (fl. conc.) |  | Chromium Total (fl. conc.) | 7440-47-3 | Metals - Total; | Metals - Total; EMS code: CR-T | Chemical | Total | LAB
CR-T | mg/L | Chromium Total | F068 | HNO3dig,TeflonTube/Autoclave,ICP-MS | This method is applicable to determiniation of dissolved, extractable, and total recoverable metals in ground waters, surface waters, drinking waters, and wastewaters. Dissolved, extractable, and total recoverable sample fractions are acidified to pH <1.5 with analytical grade nitric acid (HNO3) to prevent the metals from adsorbing to the container material or precipitating. Extractable and total recoverable water samples are collected in 250 ml acid washed polyethylene bottles and are preserved with Nitric acid (0.5 ml acid for 100 mLs of sample). All preserved samples are stored at room temperature and can be held for up to 6 months. Unfiltered, un-preserved samples are stored at 4C for no more than 30 days. For dissolved and extractable water fractions, no sample preparation is necessary. Samples are simply poured in autosampler tubes (i.e. 15 ml centrifuge tubes), and placed on the autosampler and analyzed. Samples for total recoverable metals are digested with Nitric acid in special Teflon tubes in an autoclave prior to analyses. | 0.0002 | Y | Y | Y | 1778 | 6/1/1999 1:32:00 PM | 7/4/2008 9:48:00 AM | 7/30/2024 6:32:12 AM | 1999-06-28 |  | NPEPPIN |  | Concentration | mg/L | Y | (fl. conc.) |  | Chromium Total (fl. conc.) | 7440-47-3 | Metals - Total; | Metals - Total; EMS code: CR-T | Chemical | Total | LAB
CR-T | mg/L | Chromium Total | 0200 | 20%A.R.Dig:AA Anal. | 20%A.R.Dig:AA Anal. | 0.02 | Y | Y | Y | 6 |  |  | 10/24/1996 2:35:04 AM | 1996-04-12 | 6/3/1998 6:51:35 PM | EMS |  | Concentration | mg/L | Y | (fl. conc.) |  | Chromium Total (fl. conc.) | 7440-47-3 | Metals - Total; | Metals - Total; EMS code: CR-T | Chemical | Total | LAB
CR-T | mg/L | Chromium Total | X072 | FA,HNO3,Graph.Furnace,AA | Samples are field-acidified with HNO3; HNO3 digestion; analysis by GFAA. | 0.001 | Y | Y | Y | 5937 |  |  | 8/26/2024 2:41:32 PM | 1996-04-12 | 6/3/1998 6:51:35 PM | EMS |  | Concentration | mg/L | Y | (fl. conc.) |  | Chromium Total (fl. conc.) | 7440-47-3 | Metals - Total; | Metals - Total; EMS code: CR-T | Chemical | Total | LAB
CR-T | ug/g | Chromium Total | 0192 | Hom, Nitric/Perchlor,ICP | Sample already in a dry state so does not require drying before homogenization and nitric perchlorid digestion followed by ICP analysis. | 1 | Y | Y | Y | 708 |  |  | 10/24/1996 12:50:25 PM | 1996-04-12 | 6/3/1998 6:51:41 PM | EMS |  | MassPerMass | mg/kg | Y | (mass conc.) |  | Chromium Total (mass conc.) | 7440-47-3 | Metals - Total; | Metals - Total; EMS code: CR-T | Chemical | Total | LAB
CR-T | mg/L | Chromium Total | 0042 | GI HNO3 Dig: ICP | Sample aliquot taken from unfiltered unpreserved General Ions bottle; HNO3 digestion; ICP analysis. | 0.002 | Y | Y | Y | 6823 | 9/20/1996 | 7/5/2019 3:21:00 PM | 7/12/2019 6:49:13 PM | 1996-04-12 | 6/3/1998 6:51:34 PM | EMS |  | Concentration | mg/L | Y | (fl. conc.) |  | Chromium Total (fl. conc.) | 7440-47-3 | Metals - Total; | Metals - Total; EMS code: CR-T | Chemical | Total | LAB
CR-T | mg/L | Chromium Total | EQ01 | EQUIS (mg/L) | EQUIS (mg/L) | 0.005 | N | Y | Y | 3204 |  |  | 10/24/1996 3:16:00 PM | 1996-04-12 | 6/3/1998 6:51:41 PM | EMS |  | Concentration | mg/L | Y | (fl. conc.) |  | Chromium Total (fl. conc.) | 7440-47-3 | Metals - Total; | Metals - Total; EMS code: CR-T | Chemical | Total | LAB
CR-T | ug/g wet | Chromium Total | MTIW | MET-WET-CCMS-MID-VA | Metals in Tissue by CRC ICPMS (WET) Methodology: This method is conducted following British Columbia Lab Manual method "Metals in Animal Tissue and Vegetation (Biota) - Prescriptive". Tissue samples are homogenized and sub-sampled prior to hotblock digestion with nitric and hydrochloric acids, in combination with addition of hydrogen peroxide. Instrumental analysis is by collision cell inductively coupled plasma - mass spectrometry (modified from EPA Method 6020A). Method Limitation: This method employs a strong acid/peroxide digestion, and is intended to provide a conservative estimate of bio-available metals. Near complete recoveries are achieved for most toxicologically important metals, but elements associated with recalcitrant minerals may be only partially recovered. | 0.04 | Y | Y | Y | 24 | 11/25/2018 | 12/6/2018 | 7/21/2021 2:55:07 PM | 2018-12-24 |  | EMS |  | MassPerMass | µg/g | Y | (mass conc.) | (wet) | Chromium Total (wet) (mass conc.) | 7440-47-3 | Metals - Total; | Metals - Total; EMS code: CR-T | Chemical | Total | LAB
CR-T | ug/g | Chromium Total | 0031 | ICP | ICP | 2 | Y | Y | Y | 38 |  |  | 1/29/2002 8:45:12 AM | 1996-04-12 | 6/3/1998 6:51:41 PM | EMS |  | MassPerMass | mg/kg | Y | (mass conc.) |  | Chromium Total (mass conc.) | 7440-47-3 | Metals - Total; | Metals - Total; EMS code: CR-T | Chemical | Total | LAB
CR-T | ug/g | Chromium Total | FS01 | D&G,Nitric/HCl:ICP | Samples are dried at 60 C, then either sieved to -100mesh (150 um) and rolled to homogenize, or finely ground with a mortar and pestle or a shatter box laboratory mill and rolled (depending on the submitter's requirements). The samples are then weighed into Teflon digestion liners and digested with HNO3 and HCl in a microwave oven. Samples are cooled, volumized, and allowed to settle overnight before being decanted. Analysis is then performed via Inductively Coupled Argon Plasma Atomic Emission Spectrometry (ICP-AES) (Simultaneous multi element analysis). In ICP Spectrometry, the liquid sample is pumped to the nebulizer, where it is convertd into an aerosol. the nebulizer carrier gas flow than transports the wet aerosol past the impact bead, where the large droplets condense out and flow to the waste container. The remaining fine aerosol is carried into the plasma. There, the high temperature desolvates the aerosol and dissociates the sample material into individual atoms and ions, which are excited to emit light at wavelengths characteristic of the elements present. A diffraction grating in the spectrometer diffracts the various wavelengths of light through exit slits to the photo-multiplier tubes, which convert the light energy to electronic signals. | 0.8 | Y | Y | Y | 2642 | 8/30/1996 1:52:00 PM | 10/10/2007 | 5/24/2008 1:15:02 AM | 1996-11-20 | 6/3/1998 6:51:46 PM | MPAK |  | MassPerMass | mg/kg | Y | (mass conc.) |  | Chromium Total (mass conc.) | 7440-47-3 | Metals - Total; | Metals - Total; EMS code: CR-T | Chemical | Total | LAB
CR-T | ug/g | Chromium Total | 0196 | D&G: LTA/HCl: ICP | D&G: LTA/HCl: ICP | 1 | Y | Y | Y | 27 |  |  | 10/24/1996 2:08:12 PM | 1996-04-12 | 6/3/1998 6:51:34 PM | EMS |  | MassPerMass | mg/kg | Y | (mass conc.) |  | Chromium Total (mass conc.) | 7440-47-3 | Metals - Total; | Metals - Total; EMS code: CR-T | Chemical | Total | LAB
CR-T | mg/L | Chromium Total | 1349 | Tot Metals in Seawater by CRC ICPMS (BC) | Tot Metals in Seawater by CRC ICPMS (BC) | 0.0005 | Y | Y | Y | 98 | 8/16/2019 | 12/5/2021 | 12/13/2021 12:34:16 PM | 2019-09-30 |  | KATFRASE |  | Concentration | mg/L | Y | (fl. conc.) |  | Chromium Total (fl. conc.) | 7440-47-3 | Metals - Total; | Metals - Total; EMS code: CR-T | Chemical | Total | LAB
CR-T | mg/L | Chromium Total | 0100 | Dig:Ext.(APDC):AA Anal. | Dig:Ext.(APDC):AA Anal. | 0.005 | Y | Y | Y | 14 |  |  | 10/24/1996 9:47:47 AM | 1996-04-12 | 6/3/1998 6:51:34 PM | EMS |  | Concentration | mg/L | Y | (fl. conc.) |  | Chromium Total (fl. conc.) | 7440-47-3 | Metals - Total; | Metals - Total; EMS code: CR-T | Chemical | Total | LAB
CR-T | ug/g | Chromium Total | FS82 | PBM,dry60C,10mesh,SALM-ICP-MS | PBM;dry60C;10mesh;SALM-ICP-MS | 0.2 | Y | Y | Y | 3697 | 5/21/2002 | 7/24/2024 | 7/30/2024 10:00:14 AM | 2002-11-12 |  | UMULLER |  | MassPerMass | mg/kg | Y | (mass conc.) |  | Chromium Total (mass conc.) | 7440-47-3 | Metals - Total; | Metals - Total; EMS code: CR-T | Chemical | Total | LAB
CR-T | ug/g | Chromium Total | FS89 | Homo,HNO3 dig,ICP-AES | Homo;HNO3 dig;ICP-AES | 0.8 | Y | Y | Y | 4 | 1/5/2001 | 8/20/2001 12:00:00 PM | 10/23/2002 1:42:44 PM | 2001-03-08 |  | NPEPPIN |  | MassPerMass | mg/kg | Y | (mass conc.) |  | Chromium Total (mass conc.) | 7440-47-3 | Metals - Total; | Metals - Total; EMS code: CR-T | Chemical | Total | LAB
CR-T | mg/L | Chromium Total | MWID | Metals in Water by ICPMS (digested) | Metals in Water by ICPMS (digested) Based on EPA 200.8 | 0.002 | Y | Y | Y | 36766 |  |  | 7/31/2024 1:46:41 PM | 2009-12-24 |  | RDGIBSON |  | Concentration | mg/L | Y | (fl. conc.) |  | Chromium Total (fl. conc.) | 7440-47-3 | Metals - Total; | Metals - Total; EMS code: CR-T | Chemical | Total | LAB
CR-T | mg/L | Chromium Total | F033 | Acid Dig,Centrifuge,ICPMS | For the determination of total recoverable elements in water or waste water, take 100 ml aliquot from a well mixed, acid preserved, sample containing not more than 0.25% (w/v) total solids and transfer to a 250 ml Griffin beaker. (If total solids are greater than 0.25% reduce the size of the aliquot by a proportionate amount). Add 1 ml of conc. nitric acid and 0.5 ml conc. hydrochloric acid. Heat on a hot plate until the volume has been reduced to 15-20ml, ensuring that the sample does not boil. Allow to cool and transfer to a 50 ml volumetric flask. Dilute to volume with ASTM type I water. Centrifuge the sample or allow to stand overnight to separate insoluble material. Prior analysis, pipette 20 ml into a 50 ml volumetric flask and dilute to volume with ASTM type I water. Analysis should be performed within seventy two hours of the completed preparation. Note that different laboratories may have very different field preservation, digestion procedures and detection limits for this method, so these should be documented. | 0.0005 | Y | Y | Y | 40714 | 10/15/1996 9:27:00 AM | 8/22/2024 3:24:00 AM | 8/25/2024 9:47:36 PM | 1997-01-16 | 6/3/1998 6:51:46 PM | MPAK |  | Concentration | mg/L | Y | (fl. conc.) |  | Chromium Total (fl. conc.) | 7440-47-3 | Metals - Total; | Metals - Total; EMS code: CR-T | Chemical | Total | LAB
CR-T | ug/g | Chromium Total | FB04 | Freeze Dry,Hom,Acid Dig,ICP | Sample moisture is determined by weighing samples before and after freeze-drying. Freeze-dried samples are ground in a Braun blender before being weighed into Teflon PFA liners and digested with concentrated HNO3, concentrated HCl, and 50% H2O2 in a microwave oven; which results in a complete sample dissolution. Samples are collected and volumized with Type I water. Analysis is performed via Inductively Coupled Argon Plasma Atomic Emission Spectrometry (ICP-AES) (simultaneous multielement analysis), Graphite Furnace Atomic Absorption Spectrometry (GF-AAS) (low level Cd and Pb analysis), Cold Vapor Atomic Fluorescence Spectrometry (CV-AFS) (Mercury analysis) and/or Flame Atomic Absorption Spectrometry (F-AAS) (K analysis). | 0.4 | Y | Y | Y | 241 | 8/22/1996 6:11:00 PM | 2/1/2002 2:30:00 PM | 6/23/2003 1:15:29 PM | 1997-02-27 | 6/3/1998 6:51:46 PM | MPAK |  | MassPerMass | mg/kg | Y | (mass conc.) |  | Chromium Total (mass conc.) | 7440-47-3 | Metals - Total; | Metals - Total; EMS code: CR-T | Chemical | Total | LAB
CR-T | ug/g | Chromium Total | X551 | Metals in Biota (dry), total, acid digestion, ICP-SFMS | 0.1 gram of dried, homogenized biota is weighed in a microwave TFM vessel, and 8mL of Nitric Acid and 1mL of Hydrogen Peroxide are added. The sample is digested at 200 degrees C for 15 minutes in a closed vessel high pressure microwave oven. The sample is cooled, the volume adjusted to 100mL and introduced to an iductively coupled argon plasma-sector field mass spectrometer (ICP-SFMS) by pneumatic nebulization. Each respective element is measured at a specific mass to charge ratio, m/z value expressed in amu (atomic mass units), and at a specified instrument resolution. LR: Low Resolution ~300, MR: Mid Resolution ~4000 and HR: Hich Resolution ~10000. Ref: Method number 02-2705, National Laboratory for Environmental Testing, CCIW, Environment Canada, Burlington. Requested April 2003. | 0.01 | Y | Y | Y | 70 | 8/11/2015 | 10/30/2023 | 11/28/2023 2:21:33 PM | 2004-01-19 |  | UMULLER |  | MassPerMass | mg/kg | Y | (mass conc.) |  | Chromium Total (mass conc.) | 7440-47-3 | Metals - Total; | Metals - Total; EMS code: CR-T | Chemical | Total | LAB
CR-T | mg/L | Chromium Total | F121 | Acid,ICPMS | Acidified, no digestion, ICPMS | 0.0001 | Y | Y | Y | 8500 | 3/12/2010 7:57:00 PM | 8/15/2024 9:10:00 PM | 8/26/2024 8:06:03 AM | 2002-07-17 | 7/16/2012 3:26:07 PM | UMULLER | RDGIBSON | Concentration | mg/L | Y | (fl. conc.) |  | Chromium Total (fl. conc.) | 7440-47-3 | Metals - Total; | Metals - Total; EMS code: CR-T | Chemical | Total | LAB
CR-T | mg/L | Chromium Total | F048 | Acid dig,conc'n,nebul.,ICP-OES | This method provides an automated procedure for the Inductively Coupled Argon Plasma Optical Emission Spectrophotometric (ICP-OES) determination and quantitation of trace amounts of metals in surface, ground and precipitation waters after a manual acid digestion and concentration procedure and ultrasonic nebulisation. Organically-bound metals and metal complexes are converted to metal ions by digestion with nitric acid [method details says HCl]. To increase the signal detected by the ICP, the samples are concentrated 10 times. The liquid sample is pumped onto the piezoelectric transducer, where the ultrasonic action efficiently converts it to a fine, dense aerosol. The nebulizer carrier gas flow then transports the wet aerosol through a heated U-tube, where the solvent is vaporized. Subsequently, most of the solvent vapour is removed from the aerosol stream by passing the stream through a chilled condenser. The result is a dry, analyte-laden aerosol which is introduced to the plasma. There, the high temperature desolvates the aerosol and dissociates the sample material into individual atoms and ions, which are excited to emit light at wave lengths characteristic of the elements present. Dissolved samples are field filtered through 0.45 micron cellulose acetate filter prior to acid preservation. Samples for extractable analysis are acidified with 2 mL concentrated HNO3 per litre and allowed to stand a minimum of 24 hours before analysis. Clients providing these samples are advised to preserve samples at the collection site. | 0.0002 | Y | Y | Y | 6159 | 7/4/1997 2:36:00 PM | 4/23/2003 | 9/11/2019 8:46:38 AM | 1997-10-20 | 6/3/1998 6:51:47 PM | NPEPPIN |  | Concentration | mg/L | Y | (fl. conc.) |  | Chromium Total (fl. conc.) | 7440-47-3 | Metals - Total; | Metals - Total; EMS code: CR-T | Chemical | Total | LAB
CR-T | ppm (S) | Chromium Total | 6041 | Aqua Regia by ICPMS | These were sediment/soils digested with Aqua-Regia (Nitric Acid & Hydrochloric Acids) and then analysis by ICPMS. Results would be reported in mg/kg and/or %. | 0.01 | Y | Y | Y | 21 | 10/13/2019 | 8/5/2020 | 8/10/2020 3:52:06 PM | 2019-09-06 | 10/21/2019 11:54:47 AM | KATFRASE | KATFRASE | DimensionlessRatio | ppm | N | (ratio) |  | Chromium Total (ratio) | 7440-47-3 | Metals - Total; | Metals - Total; EMS code: CR-T | Chemical | Total | LAB
CR-T | mg/L | Chromium Total | F082 | Aliquot,HNO3/HCL dig,ICPMS | Aliquot microwave digested with 3:1 HNO3/HCL. Analysis by ICPMS. For HiVol samples preliminary result are reported in mg/L and loading results are reported in ug/m3 | 0.001 | Y | Y | Y | 27314 | 7/26/2005 7:04:00 AM | 8/23/2024 | 8/26/2024 4:53:05 PM | 2000-06-13 | 8/15/2019 11:07:03 AM | NPEPPIN | KATFRASE | Concentration | mg/L | Y | (fl. conc.) |  | Chromium Total (fl. conc.) | 7440-47-3 | Metals - Total; | Metals - Total; EMS code: CR-T | Chemical | Total | LAB
CR-T | mg/L | Chromium Total | METN | MET-T-NP-CCMS-VA | Ultra trace metals in water are analyzed by CRC ICPMS, based on US EPA Method 6020B (July 2014). The detection limits provided can only be met for undigested samples. This procedure is intended for colorless, non-turbid, acid-preserved water samples (i.e. pristine water samples), having turbidity < 1 NTU and no odor. Where turbidity exceeds 1 NTU, and/or the sample is colored and has an odor, results may be biased low compared to true Total Metals concentrations. ALS recommends that turbidity analysis be requested on samples submitted for this test to aid with interpretation of results. Where turbidity is <1NTU, undigested metals are equivalent to total metals concentrations. | 0.0001 | Y | Y | Y | 3648 | 4/3/2012 | 8/19/2024 | 8/26/2024 1:03:06 PM | 2019-08-20 |  | KATFRASE |  | Concentration | mg/L | Y | (fl. conc.) |  | Chromium Total (fl. conc.) | 7440-47-3 | Metals - Total; | Metals - Total; EMS code: CR-T | Chemical | Total | LAB
CR-T | ug/g wet | Chromium Total | X215 | Hom,Nitr/Perchl,Grph Fur | Hom;Nitr/Perchl;Grph Fur | 0.01 | Y | Y | Y | 73 |  |  | 10/24/1996 10:01:25 AM | 1996-04-12 | 6/3/1998 6:51:41 PM | EMS |  | MassPerMass | µg/g | Y | (mass conc.) | (wet) | Chromium Total (wet) (mass conc.) | 7440-47-3 | Metals - Total; | Metals - Total; EMS code: CR-T | Chemical | Total | LAB
CR-T | ug/g wet | Chromium Total | X552 | Metals in Biota (wet), total, acid digestion, ICP-SFMS | One gram of wet, homogenized biota is weighed in a microwave TFM vessel, and 8mL of Nitric Acid and 1mL of Hydrogen Peroxide are added. The sample is digested at 200 degrees C for 15 minutes in a closed vessel high pressure microwave oven. The sample is cooled, the volume adjusted to 100mL and introduced to an inductively coupled argon plasma-sector field mass spectrometer (ICP-SFMS) by pneumatic nebulization. Each respective element is measured at a specific mass to charge ratio, m/z value expressed in amu (atomic mass units), and at a specified instrument resolution. LR: Low Resolution ~300, MR: Mid Resolution ~4000. Ref: 02-2704, National Laboratory for Environmental Testing, CCIW, Environment Canada, Burlington. Requested April 2003. | 0.001 | Y | Y | Y | 37 | 8/9/2016 | 10/16/2023 | 11/28/2023 2:21:32 PM | 2004-01-19 | 4/19/2021 1:22:00 PM | UMULLER | EMS | MassPerMass | µg/g | Y | (mass conc.) | (wet) | Chromium Total (wet) (mass conc.) | 7440-47-3 | Metals - Total; | Metals - Total; EMS code: CR-T | Chemical | Total | LAB
CR-T | mg/L | Chromium Total | X561 | Total metals: FA, HNO3, HR-ICPMS (No Prep)Prep) | Ultra trace metals in water are analyzed by high resolution inductively coupled plasma mass spectrometry (HR-ICPMS), modified from US EPA 200.8, (Revision 5.5). The detection limits provided can only be met for undigested samples. This procedure is intended for colorless, non-turbid, acid-preserved water samples (i.e. pristine water samples), having turbidity < 1 NTU and no odor. Where turbidity exceeds 1 NTU, and/or the sample is colored and has an odor, results may be biased low compared to true Total Metals concentrations. ALS recommends that turbidity analysis be requested on samples submitted for this test to aid with interpretation of results. | 0.0001 | Y | Y | Y | 2756 | 6/8/2015 | 10/5/2023 | 1/31/2024 2:37:41 PM | 2015-10-22 |  | RDGIBSON |  | Concentration | mg/L | Y | (fl. conc.) |  | Chromium Total (fl. conc.) | 7440-47-3 | Metals - Total; | Metals - Total; EMS code: CR-T | Chemical | Total | LAB
CR-T | mg/L | Chromium Total | X349 | FA:HNO3:ICP | Samples are field-acidified with HNO3; HNO3 digestion; analysis by ICP/AES. | 0.01 | Y | Y | Y | 30 |  |  | 10/30/2014 8:31:17 AM | 1996-04-12 | 6/3/1998 6:51:35 PM | EMS |  | Concentration | mg/L | Y | (fl. conc.) |  | Chromium Total (fl. conc.) | 7440-47-3 | Metals - Total; | Metals - Total; EMS code: CR-T | Chemical | Total | LAB
CR-T | ppm (S) | Chromium Total | MS81 | ALS Minerals: ME-MS81 | ALS Minerals: ME-MS81. A prepared sample (0.100 g) is added to lithium metaborate/lithium tetraborate flux; mixed well and fused in a furnace at 1025°C. The resulting melt is then cooled and dissolved in an acid mixture containing nitric; hydrochloric and hydrofluoric acids. This solution is then analyzed by inductively coupled plasma - mass spectrometry. | 10 | Y | Y | Y | 0 |  |  |  | 2024-01-29 |  | EMS |  | DimensionlessRatio | ppm | N | (ratio) |  | Chromium Total (ratio) | 7440-47-3 | Metals - Total; | Metals - Total; EMS code: CR-T | Chemical | Total | LAB
CR-T | ug/m3 | Chromium Total | 5312 | Tef:N/P Dig:ICP Analysis | A nitric/perchloric acid digestion if used to bring the metals into solution prior to ICP analysis. For Hi-Vol samples preliminary results are reported as mg/L and 24 hour loading results as ug/m3. | 0.01 | Y | Y | Y | 1185 | 11/7/1997 | 5/15/2000 | 11/18/2000 1:29:34 AM | 1996-10-13 | 6/3/1998 6:51:46 PM | EMS |  | AirConcentration | µg/m³ | Y | (air conc.) |  | Chromium Total (air conc.) | 7440-47-3 | Metals - Total; | Metals - Total; EMS code: CR-T | Chemical | Total | LAB
CR-T | ug/m3 | Chromium Total | 6040 | Aqua Regia Dig:ICP | Aqua regia digestion is used to bring the metals into solution prior to ICP analysis. For Hi-Vol samples, preliminary results are reported as mg/L and24 hr loadings as ug/m3 | 0.01 | Y | Y | Y | 437 | 2/13/2002 | 2/9/2005 | 2/23/2005 1:13:47 AM | 1998-05-05 | 6/3/1998 6:51:48 PM | NPEPPIN |  | AirConcentration | µg/m³ | Y | (air conc.) |  | Chromium Total (air conc.) | 7440-47-3 | Metals - Total; | Metals - Total; EMS code: CR-T | Chemical | Total | LAB
CR-T | mg/dm2/d | Chromium Total | DUST | HNO3 dig.(microwave or hot block),ICPMS | HNO3 dig.(microwave or hot block);ICPMS |  | Y | Y | Y | 855 | 7/19/2016 1:11:00 PM | 7/31/2024 | 7/31/2024 11:23:12 AM | 2008-02-07 |  | RDGIBSON |  | MassFlux | mg/dm²/d | Y | (mass flux) |  | Chromium Total (mass flux) | 7440-47-3 | Metals - Total; | Metals - Total; EMS code: CR-T | Chemical | Total | LAB
CR-T | ppm (S) | Chromium Total | 6042 | Aqua Regia by ICPMS | These were sediment/soils digested with Aqua-Regia (Nitric Acid & Hydrochloric Acids) and then analysis by ICPMS. Results would be reported in ppm s. | 0.01 | Y | Y | Y | 71 | 5/14/2022 | 7/27/2024 | 7/29/2024 9:13:08 AM | 2022-06-28 |  | EMS |  | DimensionlessRatio | ppm | N | (ratio) |  | Chromium Total (ratio) | 7440-47-3 | Metals - Total; | Metals - Total; EMS code: CR-T | Chemical | Total | LAB
CR-T | mg/L | Chromium Total | ULTR | Total Metals in Water by ICPMS (Ultra) | Total Metals in Water by ICPMS (Ultra), EPA SW-846 3005A/6020A, MET-TOT-ULTRA-MS-VA This analysis is carried out using procedures adapted from "Standard Methods for the Examination of Water and Wastewater" published by the American Public Health Association, and with procedures adapted from "Test Methods for Evaluating Solid Waste" SW-846 published by the United States Environmental Protection Agency (EPA). The procedures may involve preliminary sample treatment by acid digestion, using either hotblock or microwave oven, or filtration (EPA Method 3005A). Instrumental analysis is by inductively coupled plasma - mass spectrometry (EPA Method 6020A). | 0.0001 | Y | Y | Y | 33 | 5/25/2007 | 7/13/2011 | 7/10/2012 2:14:48 PM | 2012-07-05 |  | RDGIBSON |  | Concentration | mg/L | Y | (fl. conc.) |  | Chromium Total (fl. conc.) | 7440-47-3 | Metals - Total; | Metals - Total; EMS code: CR-T | Chemical | Total | LAB
CR-T | ug/g | Chromium Total | X492 | Homo,HCl-HNO3 dig,ICP-OES,dry calc | Homogenize samples and take sub-samples. For Total Recoverable Metals Digest with Aqua Regia. Analyses by ICPOES (Inductively Coupled Plasma Optical Emission Spectroscopy). Perform moisture analysis on separate sub-sample and calculate dry value for trace elements. | 2 | Y | Y | Y | 8 | 8/16/2022 6:33:00 PM | 8/17/2022 12:34:00 PM | 2/25/2023 1:41:51 PM | 1998-11-13 |  | NPEPPIN |  | MassPerMass | mg/kg | Y | (mass conc.) |  | Chromium Total (mass conc.) | 7440-47-3 | Metals - Total; | Metals - Total; EMS code: CR-T | Chemical | Total | LAB
CR-T | ug/g | Chromium Total | X230 | Hom,HNO3/HCl,ICP | The sample is dried, ground to pass 100 mesh, and digested with a nitric-hydrochloric acid mixture to solubilize the solid matter and to remove the organic material by oxidation and volatilization. The completed digests are submitted to metals instrumentation personnel for metals analysis by Inductively Coupled Atomic Emission Spectrophotometry (ICP-AES), or hydride analysis by atomic absorption. | 0.5 | Y | Y | Y | 266 | 2/12/2002 | 12/30/2005 | 2/3/2006 1:13:12 AM | 1996-04-12 | 6/3/1998 6:51:35 PM | EMS |  | MassPerMass | mg/kg | Y | (mass conc.) |  | Chromium Total (mass conc.) | 7440-47-3 | Metals - Total; | Metals - Total; EMS code: CR-T | Chemical | Total | LAB
CR-T | mg/L | Chromium Total | X514 | HNO3dig,BrCl2etc,MIK extr'n,AA | The sample is digested with HNO3. The pH of the solution is adjusted to 1.6, Br2 water is added, and the solution warmed until the Br2 colour vanishes. The pH of the solution is then adjusted to 3.5 with a buffer solution, ammonium pyrrolidine dithiocarbarnate solution is added, and the solution is extracted. The solvent layer is aspirated. The absorbance is measured spectrophotometrically at 358.0 mu and compared with those of identically prepared standard Cr207 ion solution an acetylene-air reducing flame is used. | 0.002 | Y | Y | Y | 537 |  |  | 8/18/1999 4:40:42 PM | 1999-04-22 |  | NPEPPIN |  | Concentration | mg/L | Y | (fl. conc.) |  | Chromium Total (fl. conc.) | 7440-47-3 | Metals - Total; | Metals - Total; EMS code: CR-T | Chemical | Total | LAB
CR-T | ppm (S) | Chromium Total | MS41 | ALS Minerals ME-MS41 | ALS Minerals ME-MS41. A prepared sample is digested with aqua regia in a graphite heating block. After cooling; the resulting solution is diluted to with deionized water; mixed and analyzed by ICP-AES. Samples are then analyzed by ICP-MS for the remaining suite of elements. | 1 | Y | Y | Y | 0 |  |  |  | 2024-01-29 |  | EMS |  | DimensionlessRatio | ppm | N | (ratio) |  | Chromium Total (ratio) | 7440-47-3 | Metals - Total; | Metals - Total; EMS code: CR-T | Chemical | Total | LAB
CR-T | mg/L | Chromium Total | 0060 | HNO3 Dig: HGA Analysis | HNO3 Dig: HGA Analysis | 0.005 | Y | Y | Y | 1236 |  |  | 9/22/1997 1:23:08 PM | 1996-04-12 | 6/3/1998 6:51:34 PM | EMS |  | Concentration | mg/L | Y | (fl. conc.) |  | Chromium Total (fl. conc.) | 7440-47-3 | Metals - Total; | Metals - Total; EMS code: CR-T | Chemical | Total | LAB
CR-T | ug/g wet | Chromium Total | BPSY | MET-WET-BPSY-CCMS-VA | Metals in Animal Tissue and Vegetation (Biota) - Prescriptive. EPA 200.3/6020B This method is conducted following British Columbia Lab Manual method "Metals in Animal Tissue and Vegetation (Biota) - Prescriptive". Tissue samples are digested with nitric and hydrochloric acids; in combination with addition of hydrogen peroxide. Instrumental analysis is by collision cell inductively coupled plasma - mass spectrometry (modified from EPA Method 6020B). | 0.04 | Y | Y | Y | 2 | 11/7/2019 | 11/7/2019 | 6/10/2021 1:57:10 PM | 2021-06-04 |  | EMS |  | MassPerMass | µg/g | Y | (mass conc.) | (wet) | Chromium Total (wet) (mass conc.) | 7440-47-3 | Metals - Total; | Metals - Total; EMS code: CR-T | Chemical | Total | LAB
CR-T | ug/g | Chromium Total | ICPM | Metals in Soil by CRC ICPMS | This analysis is carried out using procedures from CSR Analytical Method: "Strong Acid Leachable Metals (SALM) in Soil", BC Ministry of Environment, 26 June 2009, and procedures adapted from EPA Method 200.2. The sample is manually homogenized, dried at 60 degrees Celsius, sieved through a 63 um (230 mesh) sieve, and a representative subsample of the dry material is weighed. The sample is then digested at 95 degrees Celsius for 2 hours by block digester using concentrated nitric and hydrochloric acids. Instrumental analysis is by collision cell inductively coupled plasma - mass spectrometry (modifed from EPA Method 6020A). Method Limitation: This method is not a total digestion technique. It is a very strong acid digestion that is intended to dissolve those metals that may be environmentally available. By design, elements bound in silicate structures are not normally dissolved by this procedure as they are not usually mobile in the environment. | 0.5 | Y | Y | Y | 103 | 9/19/2017 | 9/24/2023 | 1/12/2024 11:16:52 AM | 2018-01-09 |  | JKROGH |  | MassPerMass | mg/kg | Y | (mass conc.) |  | Chromium Total (mass conc.) | 7440-47-3 | Metals - Total; | Metals - Total; EMS code: CR-T | Chemical | Total | LAB
CR-T | mg/L | Chromium Total | X028 | FA,HNO3,AA Analysis | FA;HNO3;AA Analysis | 0.005 | Y | Y | Y | 87 |  |  | 8/26/2004 4:21:02 PM | 1996-04-12 | 6/3/1998 6:51:35 PM | EMS |  | Concentration | mg/L | Y | (fl. conc.) |  | Chromium Total (fl. conc.) | 7440-47-3 | Metals - Total; | Metals - Total; EMS code: CR-T | Chemical | Total | LAB
CR2-O3 | % | Chromium (as Cr2O3) | ICP06 | ALS Minerals ME-ICP06 | ALS Minerals ME-ICP06. A prepared sample is fused with a lithium tetraborate/lithium metaborate flux and poured into a graphite crucible and fused in a furnace at 1000°C. The resulting melt is then cooled and dissolved in 100 mL of 4% HNO3/2% HCl followed by ICP-AES. | 0.002 | Y | Y | Y | 0 |  |  |  | 2024-01-29 |  | EMS |  | DimensionlessRatio | % | N | (ratio) |  | Chromium (as Cr2O3) (ratio) |  | Physical Properties; | Physical Properties; EMS code: CR2-O3 | Chemical |  | LAB
CR3D | mg/L | Chromium; Trivalent; Dissolved | XCAL | Calculation (mg/L) | Calculation (mg/L) | 0.005 | Y | Y | Y | 576 | 11/21/2018 | 8/14/2024 | 8/21/2024 11:30:19 AM | 2018-11-16 |  | RDGIBSON |  | Concentration | mg/L | Y | (fl. conc.) |  | Chromium; Trivalent; Dissolved (fl. conc.) | 16065-83-1 | Metals - Dissolved; | Metals - Dissolved; EMS code: CR3D | Chemical | Dissolved | LAB
CR6D | mg/L | Chromium; Hexavalent; Dissolved | FCDC | Filtration: Col'metric: diphenylcarb. | The chromium VI in the sample is reacted with diphenylcarbazide in an acid solution to produce a red-violet colour. The intensity of the colour is measured at 540 nm. Sample is lab or field filtered and preserved with NaOH. This analysis is carried out using procedures adapted from method 3500-Cr C in "Standard Methods for the Examination of Water and Wastewater" published by the American Public Health Association, and with procedures adapted from Method 1636 published by the United States Environmental Protection Agency (EPA). The procedure involves analysis for chromium (VI) by ion chromatography using diphenylcarbazide in a sulphuric acid solution. | 0.001 | Y | Y | Y | 1798 | 10/21/2015 | 8/21/2024 10:20:00 AM | 8/23/2024 5:01:25 PM | 2016-02-18 |  | RDGIBSON |  | Concentration | mg/L | Y | (fl. conc.) |  | Chromium; Hexavalent; Dissolved (fl. conc.) | 18540-29-9 | Metals - Dissolved; | Metals - Dissolved; EMS code: CR6D | Chemical | Dissolved | LAB
CS-D | mg/L | Cesium Dissolved | MWID | Metals in Water by ICPMS (digested) | Metals in Water by ICPMS (digested) Based on EPA 200.8 | 1e-05 | Y | Y | Y | 0 |  |  |  | 2023-10-10 |  | EMS |  | Concentration | mg/L | Y | (fl. conc.) |  | Cesium Dissolved (fl. conc.) | 7440-46-2 | Metals - Dissolved; | Metals - Dissolved; EMS code: CS-D | Chemical | Dissolved | LAB
CS-D | mg/L | Cesium Dissolved | X556 | Metals in water, dissolved, "in-bottle" digestion, ICP-SFMS | The whole water sample is filtered through a 0.45 micron cellulose acetate filter paper and the filtrate is stored in a HDPE/Teflon bottle. The container is de-capped in a Class 100 workstation, acidified to 2% (v/v) HNO3 and 1% (v/v) HCl and recapped. The sample is "in-bottle" digested in a convection oven, set at 60 degrees C for 16 hours. The sample is cooled and introduced to an inductively coupled argon plasma-sector field mass spectrometer (ICP-SFMS) by pneumatic nebulization. Each respective element is measured at a specific mass to charge ratio, m/z value expressed in amu (atomic mass units), and at a specified intrument resolution. LR: Low Resolution ~300, MR: Mid Resolution ~4000 and HR: High Resolution ~10000. Ref: Method number 02-2003, National Laboratory for Environmental Testing, CCIW, Environment Canada, Burlington, Requested April 2003. | 5e-06 | Y | Y | Y | 4196 | 6/27/2004 | 12/6/2023 2:04:00 AM | 2/14/2024 9:05:32 AM | 2008-06-04 | 6/4/2008 4:50:05 PM | RDGIBSON | RDGIBSON | Concentration | mg/L | Y | (fl. conc.) |  | Cesium Dissolved (fl. conc.) | 7440-46-2 | Metals - Dissolved; | Metals - Dissolved; EMS code: CS-D | Chemical | Dissolved | LAB
CS-D | mg/L | Cesium Dissolved | F032 | Filt.,Acid,ICPMS | There is no acid digestion required for this analysis. A water sample is collected and preserved in 0.2 percent nitric acid. To analyse the extractable metals the sample is shaken, allowed to sit for at least twenty four hours, and a decanted portion is introduced to the spectrometer. When the dissolved metals are required, the samples must first be filtered through a 0.45 micron cellulose acetate filter paper and then followed by preservation in 0.2 percent nitric acid. It is important to note that different laboratories may have very different detection limits for some variables analyzed by ICPMS. | 0.2 | Y | Y | Y | 9377 | 4/3/2015 | 8/26/2024 | 8/26/2024 4:51:04 PM | 2010-11-16 |  | RDGIBSON |  | Concentration | mg/L | Y | (fl. conc.) |  | Cesium Dissolved (fl. conc.) | 7440-46-2 | Metals - Dissolved; | Metals - Dissolved; EMS code: CS-D | Chemical | Dissolved | LAB
CS-D | mg/L | Cesium Dissolved | METC | MET-D-NP-CCMS-VA | Ultra trace metals in water are analyzed by CRC ICPMS, based on US EPA Method 6020B (July 2014). This procedure is intended for pristine field-filtered acid-preserved water samples. The detection limits (LOR) for this test are based on lab instrumental analysis only, not including filtration. ALS-supplied field filtration equipment does not support these LOR. Therefore, because of the high probability of false positives due to filtration, it is strongly recommended that a filtration blank be analysed to aid in data interpretation. | 5e-06 | Y | Y | Y | 394 | 3/30/2020 | 4/9/2024 | 4/29/2024 4:53:06 PM | 2022-10-07 |  | EMS |  | Concentration | mg/L | Y | (fl. conc.) |  | Cesium Dissolved (fl. conc.) | 7440-46-2 | Metals - Dissolved; | Metals - Dissolved; EMS code: CS-D | Chemical | Dissolved | LAB
CS-D | mg/L | Cesium Dissolved | X562 | Dissolved metals: FF, FA, HNO3, HR-ICPMS | Ultra trace metals in water are analyzed by high resolution inductively coupled plasma mass spectrometry (HR-ICPMS), based on US EPA Method 200.8, (Oct. 1999). This procedure is intended for pristine field-filtered acid-preserved water samples. The detection limits (LOR) for this test are based on lab instrumental analysis only, not including filtration. ALS-supplied field filtration equipment does not support these LOR. Therefore, because of the high probability of false positives due to filtration, it is strongly recommended that a filtration blank be analyzed to aid in data interpretation. | 5e-06 | Y | Y | Y | 10 | 10/20/2015 | 10/24/2018 | 10/24/2018 12:26:13 PM | 2016-03-24 |  | RDGIBSON |  | Concentration | mg/L | Y | (fl. conc.) |  | Cesium Dissolved (fl. conc.) | 7440-46-2 | Metals - Dissolved; | Metals - Dissolved; EMS code: CS-D | Chemical | Dissolved | LAB
CS-D | mg/L | Cesium Dissolved | MSEA | Metals in Seawater (undigested) by (CRC) ICP-MS | This method describes the multi-elemental determination of trace elements by Collision/Reaction Cell (CRC) ICP-MS. The method measures ions produced by a radio-frequency inductively coupled plasma. Analyte species originating in a liquid are pneumatically nebulized and the resulting aerosol transported by argon gas into the plasma torch. The ions produced are entrained in the plasma gas and extracted, by means of a differentially pumped vacuum interface, into a mass spectrometer. The ions flow through a Collision/Reaction Cell (CRC) where depending on the mode of analysis either the ions passes through the cell unhindered (no gas - classical ICP-MS mode), or collisions/reactions chemistry takes place between the ions and a neutral gas such as hydrogen or helium. Any polyatomic ions which give interference on the analytes are removed. The ions are then sorted according to their mass-to-charge ratios by a quadrupole mass spectrometer. The ions are then sorted according to their mass-to-charge ratios by a quadrupole mass spectrometer. The ions transmitted through the quadrupole are quantified by a channel electron multiplier and the ion information processed by a data handling system. Although most polyatomic interference is removed, any additional interference must be assessed and valid corrections applied or the data flagged to indicate problems. Interference correction must include compensation for background ions contributed by the plasma gas, reagents, and constituents of the sample matrix. Instrument drift as well as suppressions or enhancements of instrument response caused by the sample matrix must be corrected for by the use of internal standards. | 0.0005 | Y | Y | Y | 4 | 4/12/2023 | 4/29/2024 | 5/16/2024 2:00:14 PM | 2023-05-26 |  | EMS |  | Concentration | mg/L | Y | (fl. conc.) |  | Cesium Dissolved (fl. conc.) | 7440-46-2 | Metals - Dissolved; | Metals - Dissolved; EMS code: CS-D | Chemical | Dissolved | LAB
CS-E | mg/L | Cesium Extractable | X553 | Metals in water, extractable, ICP-SFMS | The whole water sample is collected in a HDPE/Teflon bottle. In a Class 100 workstation it is preserved with concentrated Nitric Acid at concentration of 0.2 percent (v/v). The acid preservative extracts the metals from material in the sample and a decanted portion is introduced directly to an inductively coupled argon plasma-sector field mass spectrometer (ICP-SFMS) by pneumatic nebulization. Each respective element is measured at a specific mass to charge ratio, m/z value expressed in amu (atomic mass units), and at a specified instrument resolution. LR: Low Resolution ~300, MR: Mid Resolution ~4000 and HR: High Resolution ~10000. Ref: Method number 02-2061, National Laboratory for Environmental Testing, CCIW, Environment Canada, Burlington. Requested April 2003. | 0.005 | Y | Y | Y | 1715 | 6/27/2004 | 9/21/2007 | 4/18/2018 9:28:02 AM | 2009-05-14 | 5/15/2009 11:25:09 AM | RDGIBSON | RDGIBSON | Concentration | mg/L | Y | (fl. conc.) |  | Cesium Extractable (fl. conc.) | 7440-46-2 | Metals - Extractable; | Metals - Extractable; EMS code: CS-E | Chemical | Extractable | LAB
CS-T | ppm (S) | Cesium Total | MS61 | ALS Minerals ME-MS61 | ALS Minerals ME-MS61. A prepared sample is digested with HNO3; HF; and HClO4 in a graphite heating block to near dryness. After cooling; the sample is boiled in diluted HCl. Bulk to volume with deionized water; mixed and analyzed by ICP-AES and ICP-MS for all elements. | 0.05 | Y | Y | Y | 0 |  |  |  | 2024-01-29 |  | EMS |  | DimensionlessRatio | ppm | N | (ratio) |  | Cesium Total (ratio) | 7440-46-2 | Metals - Total; | Metals - Total; EMS code: CS-T | Chemical | Total | LAB
CS-T | mg/L | Cesium Total | X555 | Metals in water, total, "in-bottle" digestion, ICP-SFMS | The whole water sample collected in a HDPE/Teflon bottle, is de-capped in a Class 100 workstation, acidified to 2% (v/v) HNO3 and 1% (v/v) HCl and recapped. The sample is "in-bottle" digested in a convection oven, set at 60 degrees C for 16 hours. The sample is cooled and introduced to an inductively coupled argon plasma-sector field mass spectrometer (ICP-SFMS) by pneumatic nebulization. Each respective element is measured at a specific mass to charge ratio, m/z value expressed in amu (atomic mass units), and at resolution. LR: Low Resolution ~300, MR: Mid Resolution ~4000 and HR: High a specified instrument Resolution ~10000. Ref: Method number 02-2003, National Laboratory for Environmental Testing, CCIW, Environment Canada, Burlington. Requested April 2003. | 5e-06 | Y | Y | Y | 15778 | 6/27/2004 | 12/6/2023 2:08:00 AM | 2/14/2024 9:05:32 AM | 2009-03-31 |  | JCHUDAK |  | Concentration | mg/L | Y | (fl. conc.) |  | Cesium Total (fl. conc.) | 7440-46-2 | Metals - Total; | Metals - Total; EMS code: CS-T | Chemical | Total | LAB
CS-T | ppm (S) | Cesium Total | 6042 | Aqua Regia by ICPMS | These were sediment/soils digested with Aqua-Regia (Nitric Acid & Hydrochloric Acids) and then analysis by ICPMS. Results would be reported in ppm s. | 0.005 | Y | Y | Y | 71 | 5/14/2022 | 7/27/2024 | 7/29/2024 9:13:08 AM | 2022-06-28 |  | EMS |  | DimensionlessRatio | ppm | N | (ratio) |  | Cesium Total (ratio) | 7440-46-2 | Metals - Total; | Metals - Total; EMS code: CS-T | Chemical | Total | LAB
CS-T | ug/g | Cesium Total | X551 | Metals in Biota (dry), total, acid digestion, ICP-SFMS | 0.1 gram of dried, homogenized biota is weighed in a microwave TFM vessel, and 8mL of Nitric Acid and 1mL of Hydrogen Peroxide are added. The sample is digested at 200 degrees C for 15 minutes in a closed vessel high pressure microwave oven. The sample is cooled, the volume adjusted to 100mL and introduced to an iductively coupled argon plasma-sector field mass spectrometer (ICP-SFMS) by pneumatic nebulization. Each respective element is measured at a specific mass to charge ratio, m/z value expressed in amu (atomic mass units), and at a specified instrument resolution. LR: Low Resolution ~300, MR: Mid Resolution ~4000 and HR: Hich Resolution ~10000. Ref: Method number 02-2705, National Laboratory for Environmental Testing, CCIW, Environment Canada, Burlington. Requested April 2003. | 0.005 | Y | Y | Y | 43 | 8/11/2015 | 10/30/2023 | 11/28/2023 2:21:33 PM | 2015-03-11 |  | RDGIBSON |  | MassPerMass | mg/kg | Y | (mass conc.) |  | Cesium Total (mass conc.) | 7440-46-2 | Metals - Total; | Metals - Total; EMS code: CS-T | Chemical | Total | LAB
CS-T | mg/L | Cesium Total | MSEA | Metals in Seawater (undigested) by (CRC) ICP-MS | This method describes the multi-elemental determination of trace elements by Collision/Reaction Cell (CRC) ICP-MS. The method measures ions produced by a radio-frequency inductively coupled plasma. Analyte species originating in a liquid are pneumatically nebulized and the resulting aerosol transported by argon gas into the plasma torch. The ions produced are entrained in the plasma gas and extracted, by means of a differentially pumped vacuum interface, into a mass spectrometer. The ions flow through a Collision/Reaction Cell (CRC) where depending on the mode of analysis either the ions passes through the cell unhindered (no gas - classical ICP-MS mode), or collisions/reactions chemistry takes place between the ions and a neutral gas such as hydrogen or helium. Any polyatomic ions which give interference on the analytes are removed. The ions are then sorted according to their mass-to-charge ratios by a quadrupole mass spectrometer. The ions are then sorted according to their mass-to-charge ratios by a quadrupole mass spectrometer. The ions transmitted through the quadrupole are quantified by a channel electron multiplier and the ion information processed by a data handling system. Although most polyatomic interference is removed, any additional interference must be assessed and valid corrections applied or the data flagged to indicate problems. Interference correction must include compensation for background ions contributed by the plasma gas, reagents, and constituents of the sample matrix. Instrument drift as well as suppressions or enhancements of instrument response caused by the sample matrix must be corrected for by the use of internal standards. | 0.0005 | Y | Y | Y | 20 | 4/13/2023 | 4/29/2024 | 6/7/2024 10:03:16 AM | 2023-05-26 |  | EMS |  | Concentration | mg/L | Y | (fl. conc.) |  | Cesium Total (fl. conc.) | 7440-46-2 | Metals - Total; | Metals - Total; EMS code: CS-T | Chemical | Total | LAB
CS-T | ug/g wet | Cesium Total | X560 | Metals in Biota (wet), total, acid digestion, ICP-MS | Five grams of wet, homogenized biota is weighed in a centrifuge tube, and 2.5 mL of Nitric Acid and 2.5 mL of Hydrochloric acid are added. The sample is digested at room temperature for one hour and then digested at 90-95 degrees C in a water bath or heat block for one more hour. The sample is cooled and the volume is made to 50 mL with DI. The digestate is then diluted at ten fold and introduced to a collision/reaction cell inductively coupled argon plasma mass spectrometer (CRC ICPMS) by pneumatic nebulization. Each respective element is measured at a specific mass to charge ratio, m/z value expressed in amu (atomic mass units). - Maxxam Analytics (BC). Added to EMS Dec 10, 2008. | 0.001 | Y | Y | Y | 179 | 8/13/2014 | 10/30/2023 | 11/28/2023 2:21:33 PM | 2015-03-11 | 3/11/2015 4:17:10 PM | RDGIBSON | EMS | MassPerMass | µg/g | Y | (mass conc.) | (wet) | Cesium Total (wet) (mass conc.) | 7440-46-2 | Metals - Total; | Metals - Total; EMS code: CS-T | Chemical | Total | LAB
CS-T | mg/L | Cesium Total | X561 | Total metals: FA, HNO3, HR-ICPMS (No Prep)Prep) | Ultra trace metals in water are analyzed by high resolution inductively coupled plasma mass spectrometry (HR-ICPMS), modified from US EPA 200.8, (Revision 5.5). The detection limits provided can only be met for undigested samples. This procedure is intended for colorless, non-turbid, acid-preserved water samples (i.e. pristine water samples), having turbidity < 1 NTU and no odor. Where turbidity exceeds 1 NTU, and/or the sample is colored and has an odor, results may be biased low compared to true Total Metals concentrations. ALS recommends that turbidity analysis be requested on samples submitted for this test to aid with interpretation of results. | 5e-06 | Y | Y | Y | 19 | 10/20/2015 | 10/24/2018 | 10/24/2018 12:26:13 PM | 2016-03-24 |  | RDGIBSON |  | Concentration | mg/L | Y | (fl. conc.) |  | Cesium Total (fl. conc.) | 7440-46-2 | Metals - Total; | Metals - Total; EMS code: CS-T | Chemical | Total | LAB
CS-T | mg/L | Cesium Total | 1349 | Tot Metals in Seawater by CRC ICPMS (BC) | Tot Metals in Seawater by CRC ICPMS (BC) | 0.0005 | Y | Y | Y | 98 | 8/16/2019 | 12/5/2021 | 12/13/2021 12:34:16 PM | 2019-09-30 |  | KATFRASE |  | Concentration | mg/L | Y | (fl. conc.) |  | Cesium Total (fl. conc.) | 7440-46-2 | Metals - Total; | Metals - Total; EMS code: CS-T | Chemical | Total | LAB
CS-T | ppm (S) | Cesium Total | MS81 | ALS Minerals: ME-MS81 | ALS Minerals: ME-MS81. A prepared sample (0.100 g) is added to lithium metaborate/lithium tetraborate flux; mixed well and fused in a furnace at 1025°C. The resulting melt is then cooled and dissolved in an acid mixture containing nitric; hydrochloric and hydrofluoric acids. This solution is then analyzed by inductively coupled plasma - mass spectrometry. | 0.01 | Y | Y | Y | 0 |  |  |  | 2024-01-29 |  | EMS |  | DimensionlessRatio | ppm | N | (ratio) |  | Cesium Total (ratio) | 7440-46-2 | Metals - Total; | Metals - Total; EMS code: CS-T | Chemical | Total | LAB
CS-T | ug/g wet | Cesium Total | BPSY | MET-WET-BPSY-CCMS-VA | Metals in Animal Tissue and Vegetation (Biota) - Prescriptive. EPA 200.3/6020B This method is conducted following British Columbia Lab Manual method "Metals in Animal Tissue and Vegetation (Biota) - Prescriptive". Tissue samples are digested with nitric and hydrochloric acids; in combination with addition of hydrogen peroxide. Instrumental analysis is by collision cell inductively coupled plasma - mass spectrometry (modified from EPA Method 6020B). | 0.001 | Y | Y | Y | 2 | 11/7/2019 | 11/7/2019 | 6/10/2021 1:57:10 PM | 2021-06-04 |  | EMS |  | MassPerMass | µg/g | Y | (mass conc.) | (wet) | Cesium Total (wet) (mass conc.) | 7440-46-2 | Metals - Total; | Metals - Total; EMS code: CS-T | Chemical | Total | LAB
CS-T | ppm (S) | Cesium Total | 6041 | Aqua Regia by ICPMS | These were sediment/soils digested with Aqua-Regia (Nitric Acid & Hydrochloric Acids) and then analysis by ICPMS. Results would be reported in mg/kg and/or %. | 0.005 | Y | Y | Y | 21 | 10/13/2019 | 8/5/2020 | 8/10/2020 3:52:06 PM | 2019-09-06 | 10/21/2019 11:54:50 AM | KATFRASE | KATFRASE | DimensionlessRatio | ppm | N | (ratio) |  | Cesium Total (ratio) | 7440-46-2 | Metals - Total; | Metals - Total; EMS code: CS-T | Chemical | Total | LAB
CS-T | mg/L | Cesium Total | F082 | Aliquot,HNO3/HCL dig,ICPMS | Aliquot microwave digested with 3:1 HNO3/HCL. Analysis by ICPMS. For HiVol samples preliminary result are reported in mg/L and loading results are reported in ug/m3 | 1e-05 | Y | Y | Y | 9143 | 4/6/2015 | 8/23/2024 | 8/26/2024 3:51:07 PM | 2015-07-30 |  | RDGIBSON |  | Concentration | mg/L | Y | (fl. conc.) |  | Cesium Total (fl. conc.) | 7440-46-2 | Metals - Total; | Metals - Total; EMS code: CS-T | Chemical | Total | LAB
CS-T | ug/g | Cesium Total | MTID | MET-DRY-CCMS-VA | Metals in Tissue by CRC ICPMS (DRY) Methodology: This method is conducted following British Columbia Lab Manual method "Metals in Animal Tissue and Vegetation (Biota) - Prescriptive". Tissue samples are homogenized and sub-sampled prior to hotblock digestion with nitric and hydrochloric acids, in combination with addition of hydrogen peroxide. Instrumental analysis is by collision cell inductively coupled plasma - mass spectrometry (modified from EPA Method 6020A). Method Limitation: This method employs a strong acid/peroxide digestion, and is intended to provide a conservative estimate of bio-available metals. Near complete recoveries are achieved for most toxicologically important metals, but elements associated with recalcitrant minerals may be only partially recovered. | 0.005 | Y | Y | Y | 122 | 8/13/2014 | 9/8/2022 | 10/26/2022 7:14:14 AM | 2015-03-09 | 3/11/2015 4:12:05 PM | RDGIBSON | EMS | MassPerMass | mg/kg | Y | (mass conc.) |  | Cesium Total (mass conc.) | 7440-46-2 | Metals - Total; | Metals - Total; EMS code: CS-T | Chemical | Total | LAB
CS-T | mg/L | Cesium Total | METN | MET-T-NP-CCMS-VA | Ultra trace metals in water are analyzed by CRC ICPMS, based on US EPA Method 6020B (July 2014). The detection limits provided can only be met for undigested samples. This procedure is intended for colorless, non-turbid, acid-preserved water samples (i.e. pristine water samples), having turbidity < 1 NTU and no odor. Where turbidity exceeds 1 NTU, and/or the sample is colored and has an odor, results may be biased low compared to true Total Metals concentrations. ALS recommends that turbidity analysis be requested on samples submitted for this test to aid with interpretation of results. Where turbidity is <1NTU, undigested metals are equivalent to total metals concentrations. | 5e-06 | Y | Y | Y | 280 | 8/6/2019 | 4/4/2024 | 4/29/2024 4:53:06 PM | 2021-06-16 |  | EMS |  | Concentration | mg/L | Y | (fl. conc.) |  | Cesium Total (fl. conc.) | 7440-46-2 | Metals - Total; | Metals - Total; EMS code: CS-T | Chemical | Total | LAB
CS-T | ug/g wet | Cesium Total | X552 | Metals in Biota (wet), total, acid digestion, ICP-SFMS | One gram of wet, homogenized biota is weighed in a microwave TFM vessel, and 8mL of Nitric Acid and 1mL of Hydrogen Peroxide are added. The sample is digested at 200 degrees C for 15 minutes in a closed vessel high pressure microwave oven. The sample is cooled, the volume adjusted to 100mL and introduced to an inductively coupled argon plasma-sector field mass spectrometer (ICP-SFMS) by pneumatic nebulization. Each respective element is measured at a specific mass to charge ratio, m/z value expressed in amu (atomic mass units), and at a specified instrument resolution. LR: Low Resolution ~300, MR: Mid Resolution ~4000. Ref: 02-2704, National Laboratory for Environmental Testing, CCIW, Environment Canada, Burlington. Requested April 2003. | 0.001 | Y | Y | Y | 37 | 8/9/2016 | 10/16/2023 | 11/28/2023 2:21:32 PM | 2015-03-11 |  | RDGIBSON |  | MassPerMass | µg/g | Y | (mass conc.) | (wet) | Cesium Total (wet) (mass conc.) | 7440-46-2 | Metals - Total; | Metals - Total; EMS code: CS-T | Chemical | Total | LAB
CS-T | ug/g wet | Cesium Total | MTIW | MET-WET-CCMS-MID-VA | Metals in Tissue by CRC ICPMS (WET) Methodology: This method is conducted following British Columbia Lab Manual method "Metals in Animal Tissue and Vegetation (Biota) - Prescriptive". Tissue samples are homogenized and sub-sampled prior to hotblock digestion with nitric and hydrochloric acids, in combination with addition of hydrogen peroxide. Instrumental analysis is by collision cell inductively coupled plasma - mass spectrometry (modified from EPA Method 6020A). Method Limitation: This method employs a strong acid/peroxide digestion, and is intended to provide a conservative estimate of bio-available metals. Near complete recoveries are achieved for most toxicologically important metals, but elements associated with recalcitrant minerals may be only partially recovered. | 0.001 | Y | Y | Y | 24 | 11/25/2018 | 12/6/2018 | 7/21/2021 2:55:07 PM | 2018-12-24 |  | EMS |  | MassPerMass | µg/g | Y | (mass conc.) | (wet) | Cesium Total (wet) (mass conc.) | 7440-46-2 | Metals - Total; | Metals - Total; EMS code: CS-T | Chemical | Total | LAB
CS-T | ppm (S) | Cesium Total | MS41 | ALS Minerals ME-MS41 | ALS Minerals ME-MS41. A prepared sample is digested with aqua regia in a graphite heating block. After cooling; the resulting solution is diluted to with deionized water; mixed and analyzed by ICP-AES. Samples are then analyzed by ICP-MS for the remaining suite of elements. | 0.05 | Y | Y | Y | 0 |  |  |  | 2024-01-29 |  | EMS |  | DimensionlessRatio | ppm | N | (ratio) |  | Cesium Total (ratio) | 7440-46-2 | Metals - Total; | Metals - Total; EMS code: CS-T | Chemical | Total | LAB
CS34 | Bq/L | Cesium-134 | X240 | Gamma Spectrum | Gamma Spectrum | 0.001 | Y | Y | Y | 127 |  |  | 10/24/1996 7:20:16 AM | 1996-04-12 | 6/3/1998 6:51:35 PM | EMS |  | Radioactivity | Bq/L | Y | (rad./vol.) |  | Cesium-134 (rad./vol.) | 13967-70-9 | Radionuclide; | Radionuclide; EMS code: CS34 | Chemical |  | LAB
CS37 | Bq/L | Cesium-137 | X240 | Gamma Spectrum | Gamma Spectrum | 0.001 | Y | Y | Y | 128 |  |  | 10/24/1996 7:20:16 AM | 1996-04-12 | 6/3/1998 6:51:35 PM | EMS |  | Radioactivity | Bq/L | Y | (rad./vol.) |  | Cesium-137 (rad./vol.) | 10045-97-3 | Radionuclide; | Radionuclide; EMS code: CS37 | Chemical |  | LAB
CSUM | meq/L | Cation Sum | XCAL | Calculation (mg/L) | Calculation (mg/L) |  | Y | Y | Y | 4350 | 6/3/2013 | 8/15/2024 | 8/26/2024 9:09:02 AM | 2016-04-05 |  | RDGIBSON |  | Equivalents | mEq/L | Y | (equiv./vol.) |  | Cation Sum (equiv./vol.) |  | Major Ions; | Major Ions; EMS code: CSUM | Chemical |  | LAB
CSUM-E | meq/L | Cation Sum Extractable | E6020B | EPA 6020B | See EPA 6020B and CSSS Ch. 15 | 0.1 | Y | Y | Y | 62 | 12/17/2021 | 7/11/2022 | 7/20/2022 4:31:25 PM | 2022-06-29 |  | EMS |  | Equivalents | mEq/L | Y | (equiv./vol.) |  | Cation Sum Extractable (equiv./vol.) |  | Physical Properties; | Physical Properties; EMS code: CSUM-E | Physical | Extractable | LAB
CU-D | mg/L | Copper Dissolved | X113 | HNO3/HCl/Perchlor,AA | HNO3/HCl/Perchlor;AA | 0.001 | Y | Y | Y | 884 |  |  | 10/24/1996 1:40:59 AM | 1996-04-12 | 6/3/1998 6:51:41 PM | EMS |  | Concentration | mg/L | Y | (fl. conc.) |  | Copper Dissolved (fl. conc.) | 7440-50-8 | Metals - Dissolved; | Metals - Dissolved; EMS code: CU-D | Chemical | Dissolved | LAB
CU-D | mg/L | Copper Dissolved | X170 | ICP | This code is a catch-all for a large variety of ICP methods including water and also solids digested into solutions as an intermediate result. Different procedures will have greatly different detection limits. (a) For air particulates, preliminary results are reported in mg/L and loading results as mg/dm2/d. (b) For Hi Vol samples preliminiary results are same as reported in mg/L and loading results as ug/m3. (c) Note that for Rain samples there is no filtration nor digestion steps. | 0.001 | Y | Y | Y | 14928 | 4/11/1996 6:01:00 PM | 11/6/2018 | 4/12/2024 1:44:29 PM | 1996-04-12 | 6/3/1998 6:51:41 PM | EMS |  | Concentration | mg/L | Y | (fl. conc.) |  | Copper Dissolved (fl. conc.) | 7440-50-8 | Metals - Dissolved; | Metals - Dissolved; EMS code: CU-D | Chemical | Dissolved | LAB
CU-D | mg/L | Copper Dissolved | X357 | LF,LA:HNO3:GFAA | Samples are lab-filtered (0.45 um) and lab-acidified with HNO3; analysis by GFAA. | 0.001 | Y | Y | Y | 34 | 1/23/1998 | 9/25/1998 | 10/30/1998 9:32:50 AM | 1996-04-12 | 6/3/1998 6:51:35 PM | EMS |  | Concentration | mg/L | Y | (fl. conc.) |  | Copper Dissolved (fl. conc.) | 7440-50-8 | Metals - Dissolved; | Metals - Dissolved; EMS code: CU-D | Chemical | Dissolved | LAB
CU-D | mg/L | Copper Dissolved | X072 | FA,HNO3,Graph.Furnace,AA | Samples are field-acidified with HNO3; HNO3 digestion; analysis by GFAA. | 0.001 | Y | Y | Y | 4604 |  |  | 1/29/2010 9:45:12 AM | 1996-04-12 | 6/3/1998 6:51:35 PM | EMS |  | Concentration | mg/L | Y | (fl. conc.) |  | Copper Dissolved (fl. conc.) | 7440-50-8 | Metals - Dissolved; | Metals - Dissolved; EMS code: CU-D | Chemical | Dissolved | LAB
CU-D | mg/L | Copper Dissolved | X285 | HNO3 Dig:ICP Analysis | HNO3 Dig:ICP Analysis | 0.01 | Y | Y | Y | 798 | 1/30/2002 | 9/16/2003 3:18:00 PM | 3/18/2004 10:28:34 AM | 1996-04-12 | 6/3/1998 6:51:35 PM | EMS |  | Concentration | mg/L | Y | (fl. conc.) |  | Copper Dissolved (fl. conc.) | 7440-50-8 | Metals - Dissolved; | Metals - Dissolved; EMS code: CU-D | Chemical | Dissolved | LAB
CU-D | mg/L | Copper Dissolved | X099 | FF,FA,HNO3,AA | FF;FA;HNO3;AA | 0.001 | Y | Y | Y | 83 |  |  | 10/24/1996 12:16:49 PM | 1996-04-12 | 6/3/1998 6:51:35 PM | EMS |  | Concentration | mg/L | Y | (fl. conc.) |  | Copper Dissolved (fl. conc.) | 7440-50-8 | Metals - Dissolved; | Metals - Dissolved; EMS code: CU-D | Chemical | Dissolved | LAB
CU-D | mg/L | Copper Dissolved | F039 | Field filt + acid, ICP-AES | Dissolved and extractable water samples are ready for analysis once collected. Total samples are autoclaved with HNO3 and hydrochloric acid (HCI) in capped Teflon vials. Analysis is performed via Inductively Coupled Argon Plasma Atomic Emission Spectrometry (ICP-AES)(simultaneous multielement analysis), Graphite Furnace Atomic Absorption Spectrometry (GF-AAS)(low level analysis) and/or Flame Atomic Absorption Spectrometry (F-AAS). | 0.005 | Y | Y | Y | 6723 | 7/29/1996 6:03:00 PM | 7/5/2019 1:21:00 PM | 9/11/2019 8:47:50 AM | 1997-05-08 | 6/3/1998 6:51:46 PM | MPAK |  | Concentration | mg/L | Y | (fl. conc.) |  | Copper Dissolved (fl. conc.) | 7440-50-8 | Metals - Dissolved; | Metals - Dissolved; EMS code: CU-D | Chemical | Dissolved | LAB
CU-D | mg/L | Copper Dissolved | X109 | LF, AA | LF; AA | 0.01 | Y | Y | Y | 1304 |  |  | 10/24/1996 2:20:41 PM | 1996-04-12 | 6/3/1998 6:51:35 PM | EMS |  | Concentration | mg/L | Y | (fl. conc.) |  | Copper Dissolved (fl. conc.) | 7440-50-8 | Metals - Dissolved; | Metals - Dissolved; EMS code: CU-D | Chemical | Dissolved | LAB
CU-D | mg/L | Copper Dissolved | X356 | LF,LA:HNO3:ICP | Samples are lab-filtered (0.45 um); lab-acidified with HNO3; analysis by ICP/AES. | 0.005 | Y | Y | Y | 423 | 10/17/2000 | 10/17/2000 | 3/23/2010 5:08:08 PM | 1996-04-12 | 6/3/1998 6:51:35 PM | EMS |  | Concentration | mg/L | Y | (fl. conc.) |  | Copper Dissolved (fl. conc.) | 7440-50-8 | Metals - Dissolved; | Metals - Dissolved; EMS code: CU-D | Chemical | Dissolved | LAB
CU-D | mg/L | Copper Dissolved | 0090 | Ext.(APDC):AA Anal. | Ext.(APDC):AA Anal. | 0.001 | Y | Y | Y | 361 | 9/29/2000 | 12/1/2010 | 12/9/2010 3:55:26 PM | 1996-04-12 | 6/3/1998 6:51:41 PM | EMS |  | Concentration | mg/L | Y | (fl. conc.) |  | Copper Dissolved (fl. conc.) | 7440-50-8 | Metals - Dissolved; | Metals - Dissolved; EMS code: CU-D | Chemical | Dissolved | LAB
CU-D | mg/L | Copper Dissolved | F032 | Filt.,Acid,ICPMS | There is no acid digestion required for this analysis. A water sample is collected and preserved in 0.2 percent nitric acid. To analyse the extractable metals the sample is shaken, allowed to sit for at least twenty four hours, and a decanted portion is introduced to the spectrometer. When the dissolved metals are required, the samples must first be filtered through a 0.45 micron cellulose acetate filter paper and then followed by preservation in 0.2 percent nitric acid. It is important to note that different laboratories may have very different detection limits for some variables analyzed by ICPMS. | 0.001 | Y | Y | Y | 89841 | 8/16/1996 2:35:00 PM | 8/26/2024 | 8/26/2024 4:53:05 PM | 1997-01-16 | 8/15/2019 10:55:43 AM | MPAK | KATFRASE | Concentration | mg/L | Y | (fl. conc.) |  | Copper Dissolved (fl. conc.) | 7440-50-8 | Metals - Dissolved; | Metals - Dissolved; EMS code: CU-D | Chemical | Dissolved | LAB
CU-D | mg/L | Copper Dissolved | X085 | LF,LA,HNO3,AA | Samples are lab-filtered (0.45 um), and lab-acidified (HNO3); analysis by AA. | 0.001 | Y | Y | Y | 2394 | 12/12/1996 | 9/18/2015 | 10/28/2015 9:54:16 AM | 1996-04-12 | 6/3/1998 6:51:41 PM | EMS |  | Concentration | mg/L | Y | (fl. conc.) |  | Copper Dissolved (fl. conc.) | 7440-50-8 | Metals - Dissolved; | Metals - Dissolved; EMS code: CU-D | Chemical | Dissolved | LAB
CU-D | mg/L | Copper Dissolved | X095 | FF,FA,HNO3,ICP | FF;FA;HNO3;ICP | 0.001 | Y | Y | Y | 345 |  |  | 10/24/1996 2:19:31 PM | 1996-04-12 | 6/3/1998 6:51:35 PM | EMS |  | Concentration | mg/L | Y | (fl. conc.) |  | Copper Dissolved (fl. conc.) | 7440-50-8 | Metals - Dissolved; | Metals - Dissolved; EMS code: CU-D | Chemical | Dissolved | LAB
CU-D | mg/L | Copper Dissolved | EQ01 | EQUIS (mg/L) | EQUIS (mg/L) | 0.001 | N | Y | Y | 8478 |  |  | 10/24/1996 3:16:55 PM | 1996-04-12 | 6/3/1998 6:51:35 PM | EMS |  | Concentration | mg/L | Y | (fl. conc.) |  | Copper Dissolved (fl. conc.) | 7440-50-8 | Metals - Dissolved; | Metals - Dissolved; EMS code: CU-D | Chemical | Dissolved | LAB
CU-D | mg/L | Copper Dissolved | F023 | HNO3,GFAA | Dissolved and extractable water samples are ready for analysis once collected. Total samples are autoclaved with HNO3 and hydrochloric acid (HCl) in capped Teflon vials. Analysis is performed via Inductively Coupled Argon Plasma Atomic Emission Spectrometry (ICP-AES) (simultaneous multielement analysis), Graphite Furnace Atomic Absorption Spectrometry (GF-AAS) (low level analysis) and/or Flame Atomic Absorption Spectrometry (F-AAS). Note that extractable samples are acidified with 0.5 mL HNO3 per 100 mL of sample at time of sample collection. | 0.0005 | Y | Y | Y | 460 | 9/5/1996 12:04:00 PM | 7/4/2001 2:06:00 PM | 8/11/2001 2:02:37 AM | 1996-11-20 | 6/3/1998 6:51:46 PM | MPAK |  | Concentration | mg/L | Y | (fl. conc.) |  | Copper Dissolved (fl. conc.) | 7440-50-8 | Metals - Dissolved; | Metals - Dissolved; EMS code: CU-D | Chemical | Dissolved | LAB
CU-D | mg/L | Copper Dissolved | 0050 | Graphite Furnace AA | Graphite Furnace AA | 0.001 | Y | Y | Y | 3393 | 8/21/2002 | 7/23/2003 | 8/6/2003 4:55:09 PM | 1996-04-12 | 6/3/1998 6:51:35 PM | EMS |  | Concentration | mg/L | Y | (fl. conc.) |  | Copper Dissolved (fl. conc.) | 7440-50-8 | Metals - Dissolved; | Metals - Dissolved; EMS code: CU-D | Chemical | Dissolved | LAB
CU-D | mg/L | Copper Dissolved | F049 | Filt,HNO3 dig,conc'n,nebul'ICP-OES | This method provides an automated procedure for the Inductively Coupled Argon Plasma Optical Emission Spectrophotometric (ICP-OES) determination and quantitation of trace amounts of metals in surface, ground and precipitation waters after a manual acid digestion and concentration procedure and ultrasonic nebulisation. Organically-bound metals and metal complexes are converted to metal ions by digestion with nitric acid. To increase the signal detected by the ICP, the samples are concentrated 10 times. The liquid sample is pumped onto the piezoelectric transducer, where the ultrasonic action efficiently converts it to a fine, dense aerosol. The nebulizer carrier gas flow then transports the wet aerosol through a heated U-tube, where the solvent is vaporized. Subsequently, most of the solvent vapour is removed from the aerosol stream by passing the stream through a chilled condenser. The result is a dry, analyte-laden aerosol which is introduced to the plasma. There, the high temperature desolvates the aerosol and dissociates the sample material into individual atoms and ions, which are excited to emit light at wave lengths characteristic of the elements present. Dissolved samples are field filtered through 0.45 micron cellulose acetate filter prior to acid preservation. Samples for extractable analysis are acidified with 2 mL concentrated HNO3 per litre and allowed to stand a minimum of 24 hours before analysis. Clients providing these samples are advised to preserve samples at the collection site. | 0.0002 | Y | Y | Y | 17678 |  |  | 5/19/2015 3:10:36 PM | 1997-10-20 | 6/3/1998 6:51:47 PM | NPEPPIN |  | Concentration | mg/L | Y | (fl. conc.) |  | Copper Dissolved (fl. conc.) | 7440-50-8 | Metals - Dissolved; | Metals - Dissolved; EMS code: CU-D | Chemical | Dissolved | LAB
CU-D | mg/L | Copper Dissolved | X556 | Metals in water, dissolved, "in-bottle" digestion, ICP-SFMS | The whole water sample is filtered through a 0.45 micron cellulose acetate filter paper and the filtrate is stored in a HDPE/Teflon bottle. The container is de-capped in a Class 100 workstation, acidified to 2% (v/v) HNO3 and 1% (v/v) HCl and recapped. The sample is "in-bottle" digested in a convection oven, set at 60 degrees C for 16 hours. The sample is cooled and introduced to an inductively coupled argon plasma-sector field mass spectrometer (ICP-SFMS) by pneumatic nebulization. Each respective element is measured at a specific mass to charge ratio, m/z value expressed in amu (atomic mass units), and at a specified intrument resolution. LR: Low Resolution ~300, MR: Mid Resolution ~4000 and HR: High Resolution ~10000. Ref: Method number 02-2003, National Laboratory for Environmental Testing, CCIW, Environment Canada, Burlington, Requested April 2003. | 2e-05 | Y | Y | Y | 4234 | 6/11/2003 | 12/6/2023 2:04:00 AM | 2/14/2024 9:05:32 AM | 2004-01-19 | 1/19/2004 2:00:32 PM | UMULLER | UMULLER | Concentration | mg/L | Y | (fl. conc.) |  | Copper Dissolved (fl. conc.) | 7440-50-8 | Metals - Dissolved; | Metals - Dissolved; EMS code: CU-D | Chemical | Dissolved | LAB
CU-D | mg/L | Copper Dissolved | METC | MET-D-NP-CCMS-VA | Ultra trace metals in water are analyzed by CRC ICPMS, based on US EPA Method 6020B (July 2014). This procedure is intended for pristine field-filtered acid-preserved water samples. The detection limits (LOR) for this test are based on lab instrumental analysis only, not including filtration. ALS-supplied field filtration equipment does not support these LOR. Therefore, because of the high probability of false positives due to filtration, it is strongly recommended that a filtration blank be analysed to aid in data interpretation. | 5e-05 | Y | Y | Y | 1776 | 4/3/2012 | 8/15/2024 | 8/23/2024 9:50:35 AM | 2019-08-20 |  | KATFRASE |  | Concentration | mg/L | Y | (fl. conc.) |  | Copper Dissolved (fl. conc.) | 7440-50-8 | Metals - Dissolved; | Metals - Dissolved; EMS code: CU-D | Chemical | Dissolved | LAB
CU-D | mg/L | Copper Dissolved | MSEA | Metals in Seawater (undigested) by (CRC) ICP-MS | This method describes the multi-elemental determination of trace elements by Collision/Reaction Cell (CRC) ICP-MS. The method measures ions produced by a radio-frequency inductively coupled plasma. Analyte species originating in a liquid are pneumatically nebulized and the resulting aerosol transported by argon gas into the plasma torch. The ions produced are entrained in the plasma gas and extracted, by means of a differentially pumped vacuum interface, into a mass spectrometer. The ions flow through a Collision/Reaction Cell (CRC) where depending on the mode of analysis either the ions passes through the cell unhindered (no gas - classical ICP-MS mode), or collisions/reactions chemistry takes place between the ions and a neutral gas such as hydrogen or helium. Any polyatomic ions which give interference on the analytes are removed. The ions are then sorted according to their mass-to-charge ratios by a quadrupole mass spectrometer. The ions are then sorted according to their mass-to-charge ratios by a quadrupole mass spectrometer. The ions transmitted through the quadrupole are quantified by a channel electron multiplier and the ion information processed by a data handling system. Although most polyatomic interference is removed, any additional interference must be assessed and valid corrections applied or the data flagged to indicate problems. Interference correction must include compensation for background ions contributed by the plasma gas, reagents, and constituents of the sample matrix. Instrument drift as well as suppressions or enhancements of instrument response caused by the sample matrix must be corrected for by the use of internal standards. | 0.0001 | Y | Y | Y | 4 | 4/12/2023 | 4/29/2024 | 5/16/2024 2:00:14 PM | 2009-09-08 |  | RDGIBSON |  | Concentration | mg/L | Y | (fl. conc.) |  | Copper Dissolved (fl. conc.) | 7440-50-8 | Metals - Dissolved; | Metals - Dissolved; EMS code: CU-D | Chemical | Dissolved | LAB
CU-D | mg/L | Copper Dissolved | 5138 | AA Analysis | AA Analysis | 0.01 | Y | Y | Y | 2 |  |  | 10/23/1996 4:28:23 PM | 1996-04-12 | 6/3/1998 6:51:35 PM | EMS |  | Concentration | mg/L | Y | (fl. conc.) |  | Copper Dissolved (fl. conc.) | 7440-50-8 | Metals - Dissolved; | Metals - Dissolved; EMS code: CU-D | Chemical | Dissolved | LAB
CU-D | mg/L | Copper Dissolved | X205 | LF,AA,Dir Asp | LF;AA;Dir Asp | 0.1 | Y | Y | Y | 153 |  |  | 10/24/1996 9:16:37 AM | 1996-04-12 | 6/3/1998 6:51:35 PM | EMS |  | Concentration | mg/L | Y | (fl. conc.) |  | Copper Dissolved (fl. conc.) | 7440-50-8 | Metals - Dissolved; | Metals - Dissolved; EMS code: CU-D | Chemical | Dissolved | LAB
CU-D | mg/L | Copper Dissolved | 0020 | HNO3 Dig:AA Analysis | HNO3 Dig:AA Analysis | 0.01 | Y | Y | Y | 6 |  |  | 3/4/1997 7:02:36 PM | 1996-04-12 | 6/3/1998 6:51:35 PM | EMS |  | Concentration | mg/L | Y | (fl. conc.) |  | Copper Dissolved (fl. conc.) | 7440-50-8 | Metals - Dissolved; | Metals - Dissolved; EMS code: CU-D | Chemical | Dissolved | LAB
CU-D | mg/L | Copper Dissolved | X116 | FF,FA,HNO3,Graph.Furn,AA | Samples are field-filtered (0.45 um), field-acidified with HNO3; analysis by GFAA. | 0.001 | Y | Y | Y | 9362 |  |  | 8/28/2018 3:14:03 PM | 1996-04-12 | 6/3/1998 6:51:35 PM | EMS |  | Concentration | mg/L | Y | (fl. conc.) |  | Copper Dissolved (fl. conc.) | 7440-50-8 | Metals - Dissolved; | Metals - Dissolved; EMS code: CU-D | Chemical | Dissolved | LAB
CU-D | mg/L | Copper Dissolved | X562 | Dissolved metals: FF, FA, HNO3, HR-ICPMS | Ultra trace metals in water are analyzed by high resolution inductively coupled plasma mass spectrometry (HR-ICPMS), based on US EPA Method 200.8, (Oct. 1999). This procedure is intended for pristine field-filtered acid-preserved water samples. The detection limits (LOR) for this test are based on lab instrumental analysis only, not including filtration. ALS-supplied field filtration equipment does not support these LOR. Therefore, because of the high probability of false positives due to filtration, it is strongly recommended that a filtration blank be analyzed to aid in data interpretation. | 5e-05 | Y | Y | Y | 1264 | 6/8/2015 | 10/19/2023 | 11/9/2023 7:41:27 AM | 2015-10-22 |  | RDGIBSON |  | Concentration | mg/L | Y | (fl. conc.) |  | Copper Dissolved (fl. conc.) | 7440-50-8 | Metals - Dissolved; | Metals - Dissolved; EMS code: CU-D | Chemical | Dissolved | LAB
CU-D | mg/L | Copper Dissolved | MWID | Metals in Water by ICPMS (digested) | Metals in Water by ICPMS (digested) Based on EPA 200.8 | 0.001 | Y | Y | Y | 31787 | 10/26/2018 5:20:00 PM | 10/26/2018 5:20:00 PM | 7/31/2024 1:46:41 PM | 2009-12-24 |  | RDGIBSON |  | Concentration | mg/L | Y | (fl. conc.) |  | Copper Dissolved (fl. conc.) | 7440-50-8 | Metals - Dissolved; | Metals - Dissolved; EMS code: CU-D | Chemical | Dissolved | LAB
CU-D | mg/L | Copper Dissolved | 0060 | HNO3 Dig: HGA Analysis | HNO3 Dig: HGA Analysis | 0.001 | Y | Y | Y | 1 |  |  | 10/24/1996 11:33:42 AM | 1996-04-12 | 6/3/1998 6:51:35 PM | EMS |  | Concentration | mg/L | Y | (fl. conc.) |  | Copper Dissolved (fl. conc.) | 7440-50-8 | Metals - Dissolved; | Metals - Dissolved; EMS code: CU-D | Chemical | Dissolved | LAB
CU-D | t/d | Copper Dissolved | CALC | Calculation (t/d) | Calculation (t/d) | 0 | Y | Y | Y | 3755 |  |  | 7/3/2024 12:59:26 PM | 1996-04-12 | 6/3/1998 6:51:35 PM | EMS |  | MassFlow | t/d | Y | (mass flow) |  | Copper Dissolved (mass flow) | 7440-50-8 | Metals - Dissolved; | Metals - Dissolved; EMS code: CU-D | Chemical | Dissolved | LAB
CU-D | mg/L | Copper Dissolved | X544 | FF,FA,HNO3,ICPMS | Analytes nebulized followed by ion entrained in plasma gas and extracted by means of a differentially pumped vacuum interface into a mass spectrometer | 0.0002 | Y | Y | Y | 90 |  |  | 10/30/2014 8:31:17 AM | 2001-09-18 | 9/18/2001 1:36:23 PM | NPEPPIN | NPEPPIN | Concentration | mg/L | Y | (fl. conc.) |  | Copper Dissolved (fl. conc.) | 7440-50-8 | Metals - Dissolved; | Metals - Dissolved; EMS code: CU-D | Chemical | Dissolved | LAB
CU-D | mg/L | Copper Dissolved | X028 | FA,HNO3,AA Analysis | FA;HNO3;AA Analysis | 0.001 | Y | Y | Y | 245 |  |  | 11/15/2000 1:52:13 PM | 1996-04-12 | 6/3/1998 6:51:35 PM | EMS |  | Concentration | mg/L | Y | (fl. conc.) |  | Copper Dissolved (fl. conc.) | 7440-50-8 | Metals - Dissolved; | Metals - Dissolved; EMS code: CU-D | Chemical | Dissolved | LAB
CU-D | mg/L | Copper Dissolved | F068 | HNO3dig,TeflonTube/Autoclave,ICP-MS | This method is applicable to determiniation of dissolved, extractable, and total recoverable metals in ground waters, surface waters, drinking waters, and wastewaters. Dissolved, extractable, and total recoverable sample fractions are acidified to pH <1.5 with analytical grade nitric acid (HNO3) to prevent the metals from adsorbing to the container material or precipitating. Extractable and total recoverable water samples are collected in 250 ml acid washed polyethylene bottles and are preserved with Nitric acid (0.5 ml acid for 100 mLs of sample). All preserved samples are stored at room temperature and can be held for up to 6 months. Unfiltered, un-preserved samples are stored at 4C for no more than 30 days. For dissolved and extractable water fractions, no sample preparation is necessary. Samples are simply poured in autosampler tubes (i.e. 15 ml centrifuge tubes), and placed on the autosampler and analyzed. Samples for total recoverable metals are digested with Nitric acid in special Teflon tubes in an autoclave prior to analyses. | 5e-05 | Y | Y | Y | 1398 | 6/1/1999 1:21:00 PM | 4/17/2015 1:06:00 PM | 8/11/2021 12:04:11 PM | 1999-06-28 |  | NPEPPIN |  | Concentration | mg/L | Y | (fl. conc.) |  | Copper Dissolved (fl. conc.) | 7440-50-8 | Metals - Dissolved; | Metals - Dissolved; EMS code: CU-D | Chemical | Dissolved | LAB
CU-E | mg/L | Copper Extractable | X553 | Metals in water, extractable, ICP-SFMS | The whole water sample is collected in a HDPE/Teflon bottle. In a Class 100 workstation it is preserved with concentrated Nitric Acid at concentration of 0.2 percent (v/v). The acid preservative extracts the metals from material in the sample and a decanted portion is introduced directly to an inductively coupled argon plasma-sector field mass spectrometer (ICP-SFMS) by pneumatic nebulization. Each respective element is measured at a specific mass to charge ratio, m/z value expressed in amu (atomic mass units), and at a specified instrument resolution. LR: Low Resolution ~300, MR: Mid Resolution ~4000 and HR: High Resolution ~10000. Ref: Method number 02-2061, National Laboratory for Environmental Testing, CCIW, Environment Canada, Burlington. Requested April 2003. | 2e-05 | Y | Y | Y | 2597 | 5/22/2003 | 9/21/2007 | 11/28/2023 2:11:53 PM | 2004-01-19 |  | UMULLER |  | Concentration | mg/L | Y | (fl. conc.) |  | Copper Extractable (fl. conc.) | 7440-50-8 | Metals - Extractable; | Metals - Extractable; EMS code: CU-E | Chemical | Extractable | LAB
CU-E | mg/L | Copper Extractable | F040 | Field Acid, lab filt, ICP-AES | Dissolved and extractable water samples are ready for analysis once collected. Total samples are autoclaved with HNO3 and hydrochloric acid (HCI) in capped Teflon vials. Analysis is performed via Inductively Coupled Argon Plasma Atomic Emission Spectrometry (ICP-AES)(simultaneous multielement analysis), Graphite Furnace Atomic Absorption Spectrometry (GF-AAS)(low level analysis) and/or Flame Atomic Absorption Spectrometry (F-AAS). | 0.005 | Y | Y | Y | 2361 | 8/5/1997 6:06:00 PM | 3/14/2008 | 9/11/2019 8:47:41 AM | 1997-05-08 | 6/3/1998 6:51:46 PM | MPAK |  | Concentration | mg/L | Y | (fl. conc.) |  | Copper Extractable (fl. conc.) | 7440-50-8 | Metals - Extractable; | Metals - Extractable; EMS code: CU-E | Chemical | Extractable | LAB
CU-E | umol/g | Copper Extractable | FSEM | AVS extraction, SEM Analysis ICPMS | Determination of simultaneous extractable metals (SEM) in sediment after sample extracted with hydrochloric acid for acid volatile sulfide (AVS) | 0.2 | Y | Y | Y | 19 | 3/12/2020 | 3/24/2020 | 2/22/2023 12:51:03 PM | 2010-11-16 |  | RDGIBSON |  | ParticlesPerMass | µmol/g | Y | (part. /m) |  | Copper Extractable (part. /m) | 7440-50-8 | Metals - Extractable; | Metals - Extractable; EMS code: CU-E | Chemical | Extractable | LAB
CU-E | mg/L | Copper Extractable | F069 | SEM extr,6N HCl,ICP/OES | 6.0 N Hydrochloric acid (HCI) is added to an aliquot of the sediment sample. The extract produced from the addition of the acid is then analyzed for simultaneously extracted metals (SEM) using inductively coupled plasma - optical emission spectrophotometry. The detection limits depend on the metal parameters in question and the instrumentation used. | 0.01 | Y | Y | Y | 10 | 3/26/2012 3:01:00 AM | 3/26/2012 3:01:00 AM | 11/25/2021 1:50:15 PM | 1999-07-09 |  | NPEPPIN |  | Concentration | mg/L | Y | (fl. conc.) |  | Copper Extractable (fl. conc.) | 7440-50-8 | Metals - Extractable; | Metals - Extractable; EMS code: CU-E | Chemical | Extractable | LAB
CU-E | ug/g | Copper Extractable | X230 | Hom,HNO3/HCl,ICP | The sample is dried, ground to pass 100 mesh, and digested with a nitric-hydrochloric acid mixture to solubilize the solid matter and to remove the organic material by oxidation and volatilization. The completed digests are submitted to metals instrumentation personnel for metals analysis by Inductively Coupled Atomic Emission Spectrophotometry (ICP-AES), or hydride analysis by atomic absorption. | 0.5 | Y | Y | Y | 74 | 12/13/2011 12:02:00 PM | 12/13/2011 12:02:00 PM | 2/9/2012 12:15:56 PM | 1996-04-12 | 6/3/1998 6:51:35 PM | EMS |  | MassPerMass | mg/kg | Y | (mass conc.) |  | Copper Extractable (mass conc.) | 7440-50-8 | Metals - Extractable; | Metals - Extractable; EMS code: CU-E | Chemical | Extractable | LAB
CU-E | mg/L | Copper Extractable | F047 | 24hr HNO3 dil'n,ICP-MS | There is no acid digestion required for this analysis. A water sample is collected and preserved in 0.2 percent nitric acid. To analyse the extractable metals the sample is shaken, allowed to sit for at least twenty four hours, and a decanted portion is introduced to the spectrometer. When the dissolved metals are required, the samples must first be filtered through a 0.45 micron cellulose acetate filter paper and then followed by preservation in 0.2 percent nitric acid. It is important to note that different laboratories may have very different detection limits for some variables analyzed by ICP/MS. | 4e-05 | Y | Y | Y | 2676 | 4/16/1998 10:21:00 AM | 6/26/2003 | 9/11/2019 8:46:18 AM | 1997-10-20 | 6/3/1998 6:51:47 PM | NPEPPIN |  | Concentration | mg/L | Y | (fl. conc.) |  | Copper Extractable (fl. conc.) | 7440-50-8 | Metals - Extractable; | Metals - Extractable; EMS code: CU-E | Chemical | Extractable | LAB
CU-E | mg/L | Copper Extractable | X512 | HNO3,pHto7.5,MIK extr'n,AA | HNO3 Digestion;pH to 7.5; MIK solvent extraction;AA | 0.001 | Y | Y | Y | 874 |  |  | 8/23/1999 3:43:46 PM | 1999-04-22 |  | NPEPPIN |  | Concentration | mg/L | Y | (fl. conc.) |  | Copper Extractable (fl. conc.) | 7440-50-8 | Metals - Extractable; | Metals - Extractable; EMS code: CU-E | Chemical | Extractable | LAB
CU-I | mg/dm2/d | Copper-Insoluble | 5039 | N/P Dig, ICP | Aliquot of prepared sample is digested first with nitric acid then with perchloric acid. Analysis by ICP. (a) For air particulates, preliminary results are reported in mg/L and loading results as mg/dm2/d. (b) For Hi Vol samples preliminary results are same as reported in mg/L and loading results as ug/m3. | 0.0005 | Y | Y | Y | 176 | 12/5/1997 | 5/15/2000 | 11/18/2000 1:30:12 AM | 1996-04-12 | 6/3/1998 6:51:35 PM | EMS |  | MassFlux | mg/dm²/d | Y | (mass flux) |  | Copper-Insoluble (mass flux) | 7440-50-8 | Metals - Insoluble; | Metals - Insoluble; EMS code: CU-I | Chemical | Insoluble | LAB
CU-I | mg/dm2/d | Copper-Insoluble | EQ15 | EQUIS (mg/dm2/d) | EQUIS (mg/dm2/d) | 0.001 | N | Y | Y | 115 |  |  | 10/22/1996 10:18:48 PM | 1996-04-12 | 6/3/1998 6:51:35 PM | EMS |  | MassFlux | mg/dm²/d | Y | (mass flux) |  | Copper-Insoluble (mass flux) | 7440-50-8 | Metals - Insoluble; | Metals - Insoluble; EMS code: CU-I | Chemical | Insoluble | LAB
CU-S | mg/dm2/d | Copper-Soluble | X286 | ICP | Aliquot of prepared sample digested with nitric acid then with perchloric acid. Filtered digest analyzed by ICP. Intermediate results reported as mg/L. Loading results for air particulate reported as mg/dm2/d. | 0.0005 | Y | Y | Y | 147 |  |  | 2/13/1998 2:28:22 AM | 1996-04-12 | 6/3/1998 6:51:42 PM | EMS |  | MassFlux | mg/dm²/d | Y | (mass flux) |  | Copper-Soluble (mass flux) | 7440-50-8 | Metals - Soluble; | Metals - Soluble; EMS code: CU-S | Chemical | Soluble | LAB
CU-S | mg/dm2/d | Copper-Soluble | 5045 | ICP Analysis | ICP a) For Air particulates analyzed without an acid digestion step, preliminary results are reported in mg/L and loading results as mg/dm2/d. b) For HiVol samples, preliminary results are reported as mg/L and loading results as ug/m3. | 0.0005 | Y | Y | Y | 39 | 11/28/1997 | 5/15/2000 | 11/18/2000 1:30:13 AM | 1998-10-28 |  | NPEPPIN |  | MassFlux | mg/dm²/d | Y | (mass flux) |  | Copper-Soluble (mass flux) | 7440-50-8 | Metals - Soluble; | Metals - Soluble; EMS code: CU-S | Chemical | Soluble | LAB
CU-S | mg/dm2/d | Copper-Soluble | EQ15 | EQUIS (mg/dm2/d) | EQUIS (mg/dm2/d) | 0.001 | N | Y | Y | 119 |  |  | 10/22/1996 10:18:48 PM | 1996-04-12 | 6/3/1998 6:51:35 PM | EMS |  | MassFlux | mg/dm²/d | Y | (mass flux) |  | Copper-Soluble (mass flux) | 7440-50-8 | Metals - Soluble; | Metals - Soluble; EMS code: CU-S | Chemical | Soluble | LAB
CU-S | umol/g | Copper-Soluble | FSEM | AVS extraction, SEM Analysis ICPMS | Determination of simultaneous extractable metals (SEM) in sediment after sample extracted with hydrochloric acid for acid volatile sulfide (AVS) | 0.2 | Y | Y | Y | 0 |  |  |  | 2024-02-02 |  | EMS |  | ParticlesPerMass | µmol/g | Y | (part. /m) |  | Copper-Soluble (part. /m) | 7440-50-8 | Metals - Soluble; | Metals - Soluble; EMS code: CU-S | Chemical | Soluble | LAB
CU-T | mg/L | Copper Total | X561 | Total metals: FA, HNO3, HR-ICPMS (No Prep)Prep) | Ultra trace metals in water are analyzed by high resolution inductively coupled plasma mass spectrometry (HR-ICPMS), modified from US EPA 200.8, (Revision 5.5). The detection limits provided can only be met for undigested samples. This procedure is intended for colorless, non-turbid, acid-preserved water samples (i.e. pristine water samples), having turbidity < 1 NTU and no odor. Where turbidity exceeds 1 NTU, and/or the sample is colored and has an odor, results may be biased low compared to true Total Metals concentrations. ALS recommends that turbidity analysis be requested on samples submitted for this test to aid with interpretation of results. | 5e-05 | Y | Y | Y | 2756 | 6/8/2015 | 10/5/2023 | 1/31/2024 2:37:41 PM | 2015-10-22 |  | RDGIBSON |  | Concentration | mg/L | Y | (fl. conc.) |  | Copper Total (fl. conc.) | 7440-50-8 | Metals - Total; | Metals - Total; EMS code: CU-T | Chemical | Total | LAB
CU-T | mg/dm2/d | Copper Total | DUST | HNO3 dig.(microwave or hot block),ICPMS | HNO3 dig.(microwave or hot block);ICPMS |  | Y | Y | Y | 781 | 7/19/2016 1:11:00 PM | 7/31/2024 | 7/31/2024 11:23:12 AM | 2008-02-07 |  | RDGIBSON |  | MassFlux | mg/dm²/d | Y | (mass flux) |  | Copper Total (mass flux) | 7440-50-8 | Metals - Total; | Metals - Total; EMS code: CU-T | Chemical | Total | LAB
CU-T | mg/L | Copper Total | X073 | FA,HNO3,Direct Flame,AA | Samples are field-acidified with HNO3; HNO3 digestion; analysis by AA. | 0.01 | Y | Y | Y | 4140 |  |  | 8/26/2024 2:41:32 PM | 1996-04-12 | 6/3/1998 6:51:41 PM | EMS |  | Concentration | mg/L | Y | (fl. conc.) |  | Copper Total (fl. conc.) | 7440-50-8 | Metals - Total; | Metals - Total; EMS code: CU-T | Chemical | Total | LAB
CU-T | ppm (S) | Copper Total | 6041 | Aqua Regia by ICPMS | These were sediment/soils digested with Aqua-Regia (Nitric Acid & Hydrochloric Acids) and then analysis by ICPMS. Results would be reported in mg/kg and/or %. | 0.01 | Y | Y | Y | 21 | 10/13/2019 | 8/5/2020 | 8/10/2020 3:52:06 PM | 2019-09-06 | 10/21/2019 11:54:54 AM | KATFRASE | KATFRASE | DimensionlessRatio | ppm | N | (ratio) |  | Copper Total (ratio) | 7440-50-8 | Metals - Total; | Metals - Total; EMS code: CU-T | Chemical | Total | LAB
CU-T | ug/m3 | Copper Total | F088 | Aliquot,HNO3/HCl dig, ICPMS,loading | Aliquot microwave digested with 3:1 HNO3/HCL. Analysis by ICPMS. For HiVol samples preliminary result are reported in mg/L and loading results are reported in ug/m3 |  | Y | Y | Y | 564 | 8/17/2000 10:46:00 AM | 12/14/2006 | 12/30/2006 1:01:36 AM | 2000-08-16 |  | NPEPPIN |  | AirConcentration | µg/m³ | Y | (air conc.) |  | Copper Total (air conc.) | 7440-50-8 | Metals - Total; | Metals - Total; EMS code: CU-T | Chemical | Total | LAB
CU-T | ug/g | Copper Total | PB02 | PBM,HNO3/HCl dig,dry60C,100mesh, | Performance Based Method. Samples are dried at 60 C and sieved using a .149 mm (100 mesh) sieve. Solid samples are digested with the combination of nitric and hydrochloric acids with a 1:1 ratio of acids. For environmental purposes, strong acid leaches are used routinely since they have the advantage of speed, low cost, and superio reproducibility. All metals trapped within the refractory matrix are unlikely to have environmental consequences. Instrumental analysis of extracts produced using this digestion method can be performed by a variety of techniques. |  | Y | Y | Y | 52 | 5/24/2001 3:42:00 PM | 11/26/2001 10:23:00 AM | 5/26/2003 2:27:50 PM | 2001-07-18 |  | NPEPPIN |  | MassPerMass | mg/kg | Y | (mass conc.) |  | Copper Total (mass conc.) | 7440-50-8 | Metals - Total; | Metals - Total; EMS code: CU-T | Chemical | Total | LAB
CU-T | mg/L | Copper Total | F108 | SPR,IDA Chelation,ICPMS | SPR;IDA Chelation;ICPMS |  | Y | Y | Y | 4 | 11/30/2000 10:32:00 AM | 11/30/2000 10:32:00 AM | 9/10/2002 9:29:48 AM | 2001-03-08 |  | NPEPPIN |  | Concentration | mg/L | Y | (fl. conc.) |  | Copper Total (fl. conc.) | 7440-50-8 | Metals - Total; | Metals - Total; EMS code: CU-T | Chemical | Total | LAB
CU-T | mg/L | Copper Total | X109 | LF, AA | LF; AA | 0.001 | Y | Y | Y | 18 |  |  | 10/24/1996 12:25:42 PM | 1996-04-12 | 6/3/1998 6:51:35 PM | EMS |  | Concentration | mg/L | Y | (fl. conc.) |  | Copper Total (fl. conc.) | 7440-50-8 | Metals - Total; | Metals - Total; EMS code: CU-T | Chemical | Total | LAB
CU-T | mg/dm2/d | Copper Total | CS15 | System Calculation | System Calculation (mg/dm2/d) |  | N | Y | N | 44 |  |  | 11/18/2000 1:30:24 AM | 1996-04-12 | 6/3/1998 6:51:44 PM | EMS |  | MassFlux | mg/dm²/d | Y | (mass flux) |  | Copper Total (mass flux) | 7440-50-8 | Metals - Total; | Metals - Total; EMS code: CU-T | Chemical | Total | LAB
CU-T | mg/L | Copper Total | MWID | Metals in Water by ICPMS (digested) | Metals in Water by ICPMS (digested) Based on EPA 200.8 | 0.001 | Y | Y | Y | 36959 |  |  | 7/31/2024 1:46:41 PM | 2009-12-24 |  | RDGIBSON |  | Concentration | mg/L | Y | (fl. conc.) |  | Copper Total (fl. conc.) | 7440-50-8 | Metals - Total; | Metals - Total; EMS code: CU-T | Chemical | Total | LAB
CU-T | ppm (S) | Copper Total | MS61 | ALS Minerals ME-MS61 | ALS Minerals ME-MS61. A prepared sample is digested with HNO3; HF; and HClO4 in a graphite heating block to near dryness. After cooling; the sample is boiled in diluted HCl. Bulk to volume with deionized water; mixed and analyzed by ICP-AES and ICP-MS for all elements. | 0.2 | Y | Y | Y | 0 |  |  |  | 2024-01-29 |  | EMS |  | DimensionlessRatio | ppm | N | (ratio) |  | Copper Total (ratio) | 7440-50-8 | Metals - Total; | Metals - Total; EMS code: CU-T | Chemical | Total | LAB
CU-T | mg/L | Copper Total | 1349 | Tot Metals in Seawater by CRC ICPMS (BC) | Tot Metals in Seawater by CRC ICPMS (BC) | 0.0005 | Y | Y | Y | 98 | 8/16/2019 | 12/5/2021 | 12/13/2021 12:34:16 PM | 2019-09-30 |  | KATFRASE |  | Concentration | mg/L | Y | (fl. conc.) |  | Copper Total (fl. conc.) | 7440-50-8 | Metals - Total; | Metals - Total; EMS code: CU-T | Chemical | Total | LAB
CU-T | ug/g | Copper Total | FS89 | Homo,HNO3 dig,ICP-AES | Homo;HNO3 dig;ICP-AES | 0.8 | Y | Y | Y | 11 | 1/5/2001 | 10/31/2022 | 11/22/2022 2:40:39 PM | 2001-03-08 |  | NPEPPIN |  | MassPerMass | mg/kg | Y | (mass conc.) |  | Copper Total (mass conc.) | 7440-50-8 | Metals - Total; | Metals - Total; EMS code: CU-T | Chemical | Total | LAB
CU-T | ug/g wet | Copper Total | X552 | Metals in Biota (wet), total, acid digestion, ICP-SFMS | One gram of wet, homogenized biota is weighed in a microwave TFM vessel, and 8mL of Nitric Acid and 1mL of Hydrogen Peroxide are added. The sample is digested at 200 degrees C for 15 minutes in a closed vessel high pressure microwave oven. The sample is cooled, the volume adjusted to 100mL and introduced to an inductively coupled argon plasma-sector field mass spectrometer (ICP-SFMS) by pneumatic nebulization. Each respective element is measured at a specific mass to charge ratio, m/z value expressed in amu (atomic mass units), and at a specified instrument resolution. LR: Low Resolution ~300, MR: Mid Resolution ~4000. Ref: 02-2704, National Laboratory for Environmental Testing, CCIW, Environment Canada, Burlington. Requested April 2003. | 0.01 | Y | Y | Y | 37 | 8/9/2016 | 10/16/2023 | 11/28/2023 2:21:32 PM | 2004-01-19 | 4/19/2021 1:22:00 PM | UMULLER | EMS | MassPerMass | µg/g | Y | (mass conc.) | (wet) | Copper Total (wet) (mass conc.) | 7440-50-8 | Metals - Total; | Metals - Total; EMS code: CU-T | Chemical | Total | LAB
CU-T | mg/L | Copper Total | X351 | LA:HNO3:AA | Samples are lab-acidified (HNO3); digested HNO3; analysis by AA. | 0.01 | Y | Y | Y | 23 |  |  | 1/19/2007 8:30:58 AM | 1996-04-12 | 6/3/1998 6:51:35 PM | EMS |  | Concentration | mg/L | Y | (fl. conc.) |  | Copper Total (fl. conc.) | 7440-50-8 | Metals - Total; | Metals - Total; EMS code: CU-T | Chemical | Total | LAB
CU-T | mg/L | Copper Total | 0020 | HNO3 Dig:AA Analysis | HNO3 Dig:AA Analysis | 0.01 | Y | Y | Y | 112 |  |  | 7/6/1998 11:15:38 AM | 1996-04-12 | 6/3/1998 6:51:35 PM | EMS |  | Concentration | mg/L | Y | (fl. conc.) |  | Copper Total (fl. conc.) | 7440-50-8 | Metals - Total; | Metals - Total; EMS code: CU-T | Chemical | Total | LAB
CU-T | ug/g | Copper Total | MTID | MET-DRY-CCMS-VA | Metals in Tissue by CRC ICPMS (DRY) Methodology: This method is conducted following British Columbia Lab Manual method "Metals in Animal Tissue and Vegetation (Biota) - Prescriptive". Tissue samples are homogenized and sub-sampled prior to hotblock digestion with nitric and hydrochloric acids, in combination with addition of hydrogen peroxide. Instrumental analysis is by collision cell inductively coupled plasma - mass spectrometry (modified from EPA Method 6020A). Method Limitation: This method employs a strong acid/peroxide digestion, and is intended to provide a conservative estimate of bio-available metals. Near complete recoveries are achieved for most toxicologically important metals, but elements associated with recalcitrant minerals may be only partially recovered. | 0.1 | Y | Y | Y | 122 | 8/13/2014 | 9/8/2022 | 10/26/2022 7:14:14 AM | 2015-03-09 | 3/11/2015 4:12:05 PM | RDGIBSON | EMS | MassPerMass | mg/kg | Y | (mass conc.) |  | Copper Total (mass conc.) | 7440-50-8 | Metals - Total; | Metals - Total; EMS code: CU-T | Chemical | Total | LAB
CU-T | mg/L | Copper Total | TCLP | Toxicity Characteristic Leaching Procedure | The TCLP procedure is designed to simulate the climatic leaching action | 0.1 | Y | Y | Y | 38 | 12/8/2009 7:49:00 PM | 9/28/2011 5:53:00 PM | 8/24/2016 2:59:13 PM | 2009-03-06 |  | JCHUDAK |  | Concentration | mg/L | Y | (fl. conc.) |  | Copper Total (fl. conc.) | 7440-50-8 | Metals - Total; | Metals - Total; EMS code: CU-T | Chemical | Total | LAB
CU-T | ug/g | Copper Total | 0192 | Hom, Nitric/Perchlor,ICP | Sample already in a dry state so does not require drying before homogenization and nitric perchlorid digestion followed by ICP analysis. | 1 | Y | Y | Y | 688 |  |  | 10/24/1996 12:50:25 PM | 1996-04-12 | 6/3/1998 6:51:35 PM | EMS |  | MassPerMass | mg/kg | Y | (mass conc.) |  | Copper Total (mass conc.) | 7440-50-8 | Metals - Total; | Metals - Total; EMS code: CU-T | Chemical | Total | LAB
CU-T | ug/g | Copper Total | E496 | CSSS (1993) 11 | CSSS (1993) 11: Plant available micro nutrients (Cu; Fe; Mn; Zn) are determined by extraction using 0.005 M DTPA followed by analysis using ICPOES. | 0.02 | Y | Y | Y | 0 |  |  |  | 2024-02-27 |  | RDGIBSON |  | MassPerMass | mg/kg | Y | (mass conc.) |  | Copper Total (mass conc.) | 7440-50-8 | Metals - Total; | Metals - Total; EMS code: CU-T | Chemical | Total | LAB
CU-T | mg/L | Copper Total | X349 | FA:HNO3:ICP | Samples are field-acidified with HNO3; HNO3 digestion; analysis by ICP/AES. | 0.005 | Y | Y | Y | 202 |  |  | 12/4/2001 11:18:16 AM | 1996-04-12 | 6/3/1998 6:51:35 PM | EMS |  | Concentration | mg/L | Y | (fl. conc.) |  | Copper Total (fl. conc.) | 7440-50-8 | Metals - Total; | Metals - Total; EMS code: CU-T | Chemical | Total | LAB
CU-T | ug/g | Copper Total | FS01 | D&G,Nitric/HCl:ICP | Samples are dried at 60 C, then either sieved to -100mesh (150 um) and rolled to homogenize, or finely ground with a mortar and pestle or a shatter box laboratory mill and rolled (depending on the submitter's requirements). The samples are then weighed into Teflon digestion liners and digested with HNO3 and HCl in a microwave oven. Samples are cooled, volumized, and allowed to settle overnight before being decanted. Analysis is then performed via Inductively Coupled Argon Plasma Atomic Emission Spectrometry (ICP-AES) (Simultaneous multi element analysis). In ICP Spectrometry, the liquid sample is pumped to the nebulizer, where it is convertd into an aerosol. the nebulizer carrier gas flow than transports the wet aerosol past the impact bead, where the large droplets condense out and flow to the waste container. The remaining fine aerosol is carried into the plasma. There, the high temperature desolvates the aerosol and dissociates the sample material into individual atoms and ions, which are excited to emit light at wavelengths characteristic of the elements present. A diffraction grating in the spectrometer diffracts the various wavelengths of light through exit slits to the photo-multiplier tubes, which convert the light energy to electronic signals. | 0.8 | Y | Y | Y | 2642 | 8/30/1996 1:52:00 PM | 10/10/2007 | 5/24/2008 1:15:02 AM | 1996-11-20 | 6/3/1998 6:51:46 PM | MPAK |  | MassPerMass | mg/kg | Y | (mass conc.) |  | Copper Total (mass conc.) | 7440-50-8 | Metals - Total; | Metals - Total; EMS code: CU-T | Chemical | Total | LAB
CU-T | ug/g | Copper Total | 0190 | D&G:N/P Dig:ICP Anal. | Sample is dried and grind-mixed before nitric perchloric digestion followed by ICP analysis. | 1 | Y | Y | Y | 949 |  |  | 10/24/1996 3:06:16 PM | 1996-04-12 | 6/3/1998 6:51:41 PM | EMS |  | MassPerMass | mg/kg | Y | (mass conc.) |  | Copper Total (mass conc.) | 7440-50-8 | Metals - Total; | Metals - Total; EMS code: CU-T | Chemical | Total | LAB
CU-T | ug/m3 | Copper Total | 5097 | N/P Dig(Filt+Wash),ICP | N/P Dig(Filt+Wash);ICP | 0.001 | Y | Y | Y | 212 |  |  | 7/3/2024 12:59:21 PM | 1996-04-12 | 6/3/1998 6:51:35 PM | EMS |  | AirConcentration | µg/m³ | Y | (air conc.) |  | Copper Total (air conc.) | 7440-50-8 | Metals - Total; | Metals - Total; EMS code: CU-T | Chemical | Total | LAB
CU-T | mg/L | Copper Total | X145 | Dig, ICP | Dig; ICP | 0.01 | Y | Y | Y | 81 | 5/8/2003 9:40:00 AM | 9/12/2003 3:16:00 PM | 12/1/2003 5:30:19 PM | 1996-04-12 | 6/3/1998 6:51:35 PM | EMS |  | Concentration | mg/L | Y | (fl. conc.) |  | Copper Total (fl. conc.) | 7440-50-8 | Metals - Total; | Metals - Total; EMS code: CU-T | Chemical | Total | LAB
CU-T | t/d | Copper Total | CALC | Calculation (t/d) | Calculation (t/d) | 0 | Y | Y | Y | 5777 |  |  | 7/3/2024 12:59:29 PM | 1996-04-12 | 6/3/1998 6:51:35 PM | EMS |  | MassFlow | t/d | Y | (mass flow) |  | Copper Total (mass flow) | 7440-50-8 | Metals - Total; | Metals - Total; EMS code: CU-T | Chemical | Total | LAB
CU-T | mg/L | Copper Total | F038 | HNO3/HCl dig,ICP-AES | Dissolved and extractable water samples are ready for analysis once collected. Total samples are autoclaved with HNO3 and hydrochloric acid (HCI) in capped Teflon vials. Analysis is performed via Inductively Coupled Argon Plasma Atomic Emission Spectrometry (ICP-AES)(simultaneous multielement analysis), Graphite Furnace Atomic Absorption Spectrometry (GF-AAS) (low level analysis) and/or Flame Atomic Absorption Spectrometry (F-AAS). | 0.006 | Y | Y | Y | 5939 | 8/1/1996 6:04:00 PM | 5/6/2008 | 9/11/2019 8:47:42 AM | 1997-05-08 | 6/3/1998 6:51:46 PM | MPAK |  | Concentration | mg/L | Y | (fl. conc.) |  | Copper Total (fl. conc.) | 7440-50-8 | Metals - Total; | Metals - Total; EMS code: CU-T | Chemical | Total | LAB
CU-T | ug/m3 | Copper Total | EQ40 | EQUIS (mg/m3) | EQUIS (mg/m3) | 0.01 | N | Y | Y | 1105 |  |  | 10/22/1996 10:18:49 PM | 1996-04-12 | 6/3/1998 6:51:35 PM | EMS |  | AirConcentration | µg/m³ | Y | (air conc.) |  | Copper Total (air conc.) | 7440-50-8 | Metals - Total; | Metals - Total; EMS code: CU-T | Chemical | Total | LAB
CU-T | mg/L | Copper Total | F068 | HNO3dig,TeflonTube/Autoclave,ICP-MS | This method is applicable to determiniation of dissolved, extractable, and total recoverable metals in ground waters, surface waters, drinking waters, and wastewaters. Dissolved, extractable, and total recoverable sample fractions are acidified to pH <1.5 with analytical grade nitric acid (HNO3) to prevent the metals from adsorbing to the container material or precipitating. Extractable and total recoverable water samples are collected in 250 ml acid washed polyethylene bottles and are preserved with Nitric acid (0.5 ml acid for 100 mLs of sample). All preserved samples are stored at room temperature and can be held for up to 6 months. Unfiltered, un-preserved samples are stored at 4C for no more than 30 days. For dissolved and extractable water fractions, no sample preparation is necessary. Samples are simply poured in autosampler tubes (i.e. 15 ml centrifuge tubes), and placed on the autosampler and analyzed. Samples for total recoverable metals are digested with Nitric acid in special Teflon tubes in an autoclave prior to analyses. | 5e-05 | Y | Y | Y | 3947 | 6/1/1999 1:32:00 PM | 7/4/2008 9:48:00 AM | 7/30/2024 6:32:12 AM | 1999-06-28 |  | NPEPPIN |  | Concentration | mg/L | Y | (fl. conc.) |  | Copper Total (fl. conc.) | 7440-50-8 | Metals - Total; | Metals - Total; EMS code: CU-T | Chemical | Total | LAB
CU-T | mg/L | Copper Total | F033 | Acid Dig,Centrifuge,ICPMS | For the determination of total recoverable elements in water or waste water, take 100 ml aliquot from a well mixed, acid preserved, sample containing not more than 0.25% (w/v) total solids and transfer to a 250 ml Griffin beaker. (If total solids are greater than 0.25% reduce the size of the aliquot by a proportionate amount). Add 1 ml of conc. nitric acid and 0.5 ml conc. hydrochloric acid. Heat on a hot plate until the volume has been reduced to 15-20ml, ensuring that the sample does not boil. Allow to cool and transfer to a 50 ml volumetric flask. Dilute to volume with ASTM type I water. Centrifuge the sample or allow to stand overnight to separate insoluble material. Prior analysis, pipette 20 ml into a 50 ml volumetric flask and dilute to volume with ASTM type I water. Analysis should be performed within seventy two hours of the completed preparation. Note that different laboratories may have very different field preservation, digestion procedures and detection limits for this method, so these should be documented. | 0.0005 | Y | Y | Y | 38834 | 10/15/1996 9:27:00 AM | 8/22/2024 3:24:00 AM | 8/25/2024 9:47:36 PM | 1997-01-16 | 6/3/1998 6:51:46 PM | MPAK |  | Concentration | mg/L | Y | (fl. conc.) |  | Copper Total (fl. conc.) | 7440-50-8 | Metals - Total; | Metals - Total; EMS code: CU-T | Chemical | Total | LAB
CU-T | ug/m3 | Copper Total | X136 | HNO3/HCl/Perchlor,AA | HNO3/HCl/Perchlor;AA | 0 | Y | Y | Y | 54 |  |  | 12/23/1997 1:42:11 PM | 1996-04-12 | 6/3/1998 6:51:35 PM | EMS |  | AirConcentration | µg/m³ | Y | (air conc.) |  | Copper Total (air conc.) | 7440-50-8 | Metals - Total; | Metals - Total; EMS code: CU-T | Chemical | Total | LAB
CU-T | ug/g wet | Copper Total | X217 | Hom,Nitric/Perchl,ICP | Sample is wet but not dried before homogenization and nitric perchloric digestion followed by ICP analysis. | 0.01 | Y | Y | Y | 72 |  |  | 10/24/1996 10:01:25 AM | 1996-04-12 | 6/3/1998 6:51:41 PM | EMS |  | MassPerMass | µg/g | Y | (mass conc.) | (wet) | Copper Total (wet) (mass conc.) | 7440-50-8 | Metals - Total; | Metals - Total; EMS code: CU-T | Chemical | Total | LAB
CU-T | ug/g wet | Copper Total | BPSY | MET-WET-BPSY-CCMS-VA | Metals in Animal Tissue and Vegetation (Biota) - Prescriptive. EPA 200.3/6020B This method is conducted following British Columbia Lab Manual method "Metals in Animal Tissue and Vegetation (Biota) - Prescriptive". Tissue samples are digested with nitric and hydrochloric acids; in combination with addition of hydrogen peroxide. Instrumental analysis is by collision cell inductively coupled plasma - mass spectrometry (modified from EPA Method 6020B). | 0.04 | Y | Y | Y | 2 | 11/7/2019 | 11/7/2019 | 6/10/2021 1:57:10 PM | 2021-06-04 |  | EMS |  | MassPerMass | µg/g | Y | (mass conc.) | (wet) | Copper Total (wet) (mass conc.) | 7440-50-8 | Metals - Total; | Metals - Total; EMS code: CU-T | Chemical | Total | LAB
CU-T | ug/g | Copper Total | 0196 | D&G: LTA/HCl: ICP | D&G: LTA/HCl: ICP | 1 | Y | Y | Y | 27 |  |  | 10/24/1996 2:08:12 PM | 1996-04-12 | 6/3/1998 6:51:35 PM | EMS |  | MassPerMass | mg/kg | Y | (mass conc.) |  | Copper Total (mass conc.) | 7440-50-8 | Metals - Total; | Metals - Total; EMS code: CU-T | Chemical | Total | LAB
CU-T | mg/L | Copper Total | X170 | ICP | This code is a catch-all for a large variety of ICP methods including water and also solids digested into solutions as an intermediate result. Different procedures will have greatly different detection limits. (a) For air particulates, preliminary results are reported in mg/L and loading results as mg/dm2/d. (b) For Hi Vol samples preliminiary results are same as reported in mg/L and loading results as ug/m3. (c) Note that for Rain samples there is no filtration nor digestion steps. | 0.01 | Y | Y | Y | 9655 | 3/28/1996 6:03:00 PM | 5/7/2024 | 7/25/2024 11:22:07 AM | 1996-04-12 | 6/3/1998 6:51:35 PM | EMS |  | Concentration | mg/L | Y | (fl. conc.) |  | Copper Total (fl. conc.) | 7440-50-8 | Metals - Total; | Metals - Total; EMS code: CU-T | Chemical | Total | LAB
CU-T | mg/L | Copper Total | X545 | FA,HNO3,ICPMS | Analytes nebulized followed by ion entrained in plasma gas and extracted by means of a differentially pumped vacuum interface into a mass spectrometer | 0.0002 | Y | Y | Y | 30 |  |  | 10/30/2014 8:31:17 AM | 2001-09-18 | 9/18/2001 1:36:23 PM | NPEPPIN | NPEPPIN | Concentration | mg/L | Y | (fl. conc.) |  | Copper Total (fl. conc.) | 7440-50-8 | Metals - Total; | Metals - Total; EMS code: CU-T | Chemical | Total | LAB
CU-T | mg/L | Copper Total | 0050 | Graphite Furnace AA | Graphite Furnace AA | 0.0005 | Y | Y | Y | 11 | 12/24/2002 | 7/23/2003 | 8/6/2003 4:51:04 PM | 2003-01-13 |  | UMULLER |  | Concentration | mg/L | Y | (fl. conc.) |  | Copper Total (fl. conc.) | 7440-50-8 | Metals - Total; | Metals - Total; EMS code: CU-T | Chemical | Total | LAB
CU-T | ug/g wet | Copper Total | X560 | Metals in Biota (wet), total, acid digestion, ICP-MS | Five grams of wet, homogenized biota is weighed in a centrifuge tube, and 2.5 mL of Nitric Acid and 2.5 mL of Hydrochloric acid are added. The sample is digested at room temperature for one hour and then digested at 90-95 degrees C in a water bath or heat block for one more hour. The sample is cooled and the volume is made to 50 mL with DI. The digestate is then diluted at ten fold and introduced to a collision/reaction cell inductively coupled argon plasma mass spectrometer (CRC ICPMS) by pneumatic nebulization. Each respective element is measured at a specific mass to charge ratio, m/z value expressed in amu (atomic mass units). - Maxxam Analytics (BC). Added to EMS Dec 10, 2008. | 0.5 | Y | Y | Y | 541 | 11/15/2008 4:00:00 AM | 10/30/2023 | 11/28/2023 2:21:33 PM | 2008-12-11 | 3/11/2015 4:17:10 PM | RDGIBSON | EMS | MassPerMass | µg/g | Y | (mass conc.) | (wet) | Copper Total (wet) (mass conc.) | 7440-50-8 | Metals - Total; | Metals - Total; EMS code: CU-T | Chemical | Total | LAB
CU-T | mg/L | Copper Total | X234 | Aqua Reg. Dig,AA | Aqua Reg. Dig;AA | 0.0001 | Y | Y | Y | 335 |  |  | 10/24/1996 12:12:10 PM | 1996-04-12 | 6/3/1998 6:51:35 PM | EMS |  | Concentration | mg/L | Y | (fl. conc.) |  | Copper Total (fl. conc.) | 7440-50-8 | Metals - Total; | Metals - Total; EMS code: CU-T | Chemical | Total | LAB
CU-T | mg/L | Copper Total | 0200 | 20%A.R.Dig:AA Anal. | 20%A.R.Dig:AA Anal. | 0.01 | Y | Y | Y | 10 |  |  | 10/24/1996 4:16:44 AM | 1996-04-12 | 6/3/1998 6:51:41 PM | EMS |  | Concentration | mg/L | Y | (fl. conc.) |  | Copper Total (fl. conc.) | 7440-50-8 | Metals - Total; | Metals - Total; EMS code: CU-T | Chemical | Total | LAB
CU-T | mg/L | Copper Total | EQ01 | EQUIS (mg/L) | EQUIS (mg/L) | 0.001 | N | Y | Y | 7456 |  |  | 10/24/1996 3:16:18 PM | 1996-04-12 | 6/3/1998 6:51:35 PM | EMS |  | Concentration | mg/L | Y | (fl. conc.) |  | Copper Total (fl. conc.) | 7440-50-8 | Metals - Total; | Metals - Total; EMS code: CU-T | Chemical | Total | LAB
CU-T | ug/g | Copper Total | EQ02 | EQUIS (ug/g) | EQUIS (ug/g) | 1 | N | Y | Y | 949 |  |  | 10/23/1996 12:08:28 AM | 1996-04-12 | 6/3/1998 6:51:41 PM | EMS |  | MassPerMass | mg/kg | Y | (mass conc.) |  | Copper Total (mass conc.) | 7440-50-8 | Metals - Total; | Metals - Total; EMS code: CU-T | Chemical | Total | LAB
CU-T | mg/L | Copper Total | X088 | Nitric/Perchlor Dig,AA | Nitric/Perchlor Dig;AA | 0.001 | Y | Y | Y | 182 |  |  | 11/15/2000 1:51:32 PM | 1996-04-12 | 6/3/1998 6:51:35 PM | EMS |  | Concentration | mg/L | Y | (fl. conc.) |  | Copper Total (fl. conc.) | 7440-50-8 | Metals - Total; | Metals - Total; EMS code: CU-T | Chemical | Total | LAB
CU-T | mg/L | Copper Total | 0040 | HNO3 Dig: ICP Analysis | HNO3 Dig: ICP Analysis | 0.01 | Y | Y | Y | 11211 | 10/17/2000 | 10/17/2000 | 10/18/2023 8:55:31 AM | 1996-04-12 | 6/3/1998 6:51:35 PM | EMS |  | Concentration | mg/L | Y | (fl. conc.) |  | Copper Total (fl. conc.) | 7440-50-8 | Metals - Total; | Metals - Total; EMS code: CU-T | Chemical | Total | LAB
CU-T | ug/g | Copper Total | X492 | Homo,HCl-HNO3 dig,ICP-OES,dry calc | Homogenize samples and take sub-samples. For Total Recoverable Metals Digest with Aqua Regia. Analyses by ICPOES (Inductively Coupled Plasma Optical Emission Spectroscopy). Perform moisture analysis on separate sub-sample and calculate dry value for trace elements. | 1 | Y | Y | Y | 1 |  |  | 8/24/2016 2:59:13 PM | 1998-11-13 |  | NPEPPIN |  | MassPerMass | mg/kg | Y | (mass conc.) |  | Copper Total (mass conc.) | 7440-50-8 | Metals - Total; | Metals - Total; EMS code: CU-T | Chemical | Total | LAB
CU-T | ug/g | Copper Total | FB04 | Freeze Dry,Hom,Acid Dig,ICP | Sample moisture is determined by weighing samples before and after freeze-drying. Freeze-dried samples are ground in a Braun blender before being weighed into Teflon PFA liners and digested with concentrated HNO3, concentrated HCl, and 50% H2O2 in a microwave oven; which results in a complete sample dissolution. Samples are collected and volumized with Type I water. Analysis is performed via Inductively Coupled Argon Plasma Atomic Emission Spectrometry (ICP-AES) (simultaneous multielement analysis), Graphite Furnace Atomic Absorption Spectrometry (GF-AAS) (low level Cd and Pb analysis), Cold Vapor Atomic Fluorescence Spectrometry (CV-AFS) (Mercury analysis) and/or Flame Atomic Absorption Spectrometry (F-AAS) (K analysis). | 0.4 | Y | Y | Y | 241 | 8/22/1996 6:11:00 PM | 2/1/2002 2:30:00 PM | 6/23/2003 1:15:30 PM | 1997-02-27 | 6/3/1998 6:51:46 PM | MPAK |  | MassPerMass | mg/kg | Y | (mass conc.) |  | Copper Total (mass conc.) | 7440-50-8 | Metals - Total; | Metals - Total; EMS code: CU-T | Chemical | Total | LAB
CU-T | ug/g | Copper Total | X230 | Hom,HNO3/HCl,ICP | The sample is dried, ground to pass 100 mesh, and digested with a nitric-hydrochloric acid mixture to solubilize the solid matter and to remove the organic material by oxidation and volatilization. The completed digests are submitted to metals instrumentation personnel for metals analysis by Inductively Coupled Atomic Emission Spectrophotometry (ICP-AES), or hydride analysis by atomic absorption. | 0.5 | Y | Y | Y | 259 | 2/12/2002 | 12/30/2005 | 2/3/2006 1:13:12 AM | 1996-04-12 | 6/3/1998 6:51:35 PM | EMS |  | MassPerMass | mg/kg | Y | (mass conc.) |  | Copper Total (mass conc.) | 7440-50-8 | Metals - Total; | Metals - Total; EMS code: CU-T | Chemical | Total | LAB
CU-T | mg/L | Copper Total | 0042 | GI HNO3 Dig: ICP | Sample aliquot taken from unfiltered unpreserved General Ions bottle; HNO3 digestion; ICP analysis. | 0.002 | Y | Y | Y | 6851 | 9/20/1996 | 7/5/2019 3:21:00 PM | 7/12/2019 6:49:13 PM | 1996-04-12 | 6/3/1998 6:51:35 PM | EMS |  | Concentration | mg/L | Y | (fl. conc.) |  | Copper Total (fl. conc.) | 7440-50-8 | Metals - Total; | Metals - Total; EMS code: CU-T | Chemical | Total | LAB
CU-T | mg/L | Copper Total | 0060 | HNO3 Dig: HGA Analysis | HNO3 Dig: HGA Analysis | 0.001 | Y | Y | Y | 4941 | 6/25/2010 3:15:00 PM | 12/1/2010 | 12/9/2010 3:55:26 PM | 1996-04-12 | 6/3/1998 6:51:41 PM | EMS |  | Concentration | mg/L | Y | (fl. conc.) |  | Copper Total (fl. conc.) | 7440-50-8 | Metals - Total; | Metals - Total; EMS code: CU-T | Chemical | Total | LAB
CU-T | mg/L | Copper Total | X285 | HNO3 Dig:ICP Analysis | HNO3 Dig:ICP Analysis | 0.001 | Y | Y | Y | 1017 | 8/10/1993 | 10/5/1993 | 4/30/2004 1:59:45 PM | 1996-04-12 | 6/3/1998 6:51:41 PM | EMS |  | Concentration | mg/L | Y | (fl. conc.) |  | Copper Total (fl. conc.) | 7440-50-8 | Metals - Total; | Metals - Total; EMS code: CU-T | Chemical | Total | LAB
CU-T | mg/L | Copper Total | 0105 | 5%AR Dig:APDC Ext:AA Anl | 5%AR Dig:APDC Ext:AA Anl | 0.001 | Y | Y | Y | 515 |  |  | 10/24/1996 2:13:46 PM | 1996-04-12 | 6/3/1998 6:51:35 PM | EMS |  | Concentration | mg/L | Y | (fl. conc.) |  | Copper Total (fl. conc.) | 7440-50-8 | Metals - Total; | Metals - Total; EMS code: CU-T | Chemical | Total | LAB
CU-T | mg/L | Copper Total | X179 | LA,HNO3,Graphite Furnace | Samples are lab-acidified with HNO3; HNO3 digestion; analysis by GFAA. | 0.001 | Y | Y | Y | 154 | 2/12/1998 | 6/18/2002 | 7/11/2002 3:18:05 PM | 1996-04-12 | 6/3/1998 6:51:35 PM | EMS |  | Concentration | mg/L | Y | (fl. conc.) |  | Copper Total (fl. conc.) | 7440-50-8 | Metals - Total; | Metals - Total; EMS code: CU-T | Chemical | Total | LAB
CU-T | ug/m3 | Copper Total | 5012 | N/P Dig:HGA Analysis | N/P Dig:HGA Analysis | 0.01 | N | Y | Y | 737 |  |  | 10/24/1996 6:54:32 AM | 1996-04-12 | 6/3/1998 6:51:41 PM | EMS |  | AirConcentration | µg/m³ | Y | (air conc.) |  | Copper Total (air conc.) | 7440-50-8 | Metals - Total; | Metals - Total; EMS code: CU-T | Chemical | Total | LAB
CU-T | mg/L | Copper Total | X512 | HNO3,pHto7.5,MIK extr'n,AA | HNO3 Digestion;pH to 7.5; MIK solvent extraction;AA | 0.001 | Y | Y | Y | 2980 |  |  | 8/23/1999 4:52:44 PM | 1999-04-22 |  | NPEPPIN |  | Concentration | mg/L | Y | (fl. conc.) |  | Copper Total (fl. conc.) | 7440-50-8 | Metals - Total; | Metals - Total; EMS code: CU-T | Chemical | Total | LAB
CU-T | mg/L | Copper Total | F024 | HNO3/HCl,GFAA | Dissolved and extractable water samples are ready for analysis once collected. Total samples are autoclaved with HNO3 and hydrochloric acid (HCl) in capped Teflon vials. Analysis is performed via Inductively Coupled Argon Plasma Atomic Emission Spectrometry (ICP-AES) (simultaneous multielement analysis), Graphite Furnace Atomic Absorption Spectrometry (GF-AAS) (low level analysis) and/or Flame Atomic Absorption Spectrometry (F-AAS). | 0.0005 | Y | Y | Y | 874 | 8/6/1996 7:16:00 PM | 12/7/2001 7:26:00 AM | 2/14/2002 1:17:10 AM | 1996-11-20 | 6/3/1998 6:51:46 PM | MPAK |  | Concentration | mg/L | Y | (fl. conc.) |  | Copper Total (fl. conc.) | 7440-50-8 | Metals - Total; | Metals - Total; EMS code: CU-T | Chemical | Total | LAB
CU-T | ug/g | Copper Total | FS82 | PBM,dry60C,10mesh,SALM-ICP-MS | PBM;dry60C;10mesh;SALM-ICP-MS | 0.5 | Y | Y | Y | 3690 | 5/21/2002 | 7/24/2024 | 7/30/2024 10:00:14 AM | 2002-11-12 |  | UMULLER |  | MassPerMass | mg/kg | Y | (mass conc.) |  | Copper Total (mass conc.) | 7440-50-8 | Metals - Total; | Metals - Total; EMS code: CU-T | Chemical | Total | LAB
CU-T | mg/L | Copper Total | MSEA | Metals in Seawater (undigested) by (CRC) ICP-MS | This method describes the multi-elemental determination of trace elements by Collision/Reaction Cell (CRC) ICP-MS. The method measures ions produced by a radio-frequency inductively coupled plasma. Analyte species originating in a liquid are pneumatically nebulized and the resulting aerosol transported by argon gas into the plasma torch. The ions produced are entrained in the plasma gas and extracted, by means of a differentially pumped vacuum interface, into a mass spectrometer. The ions flow through a Collision/Reaction Cell (CRC) where depending on the mode of analysis either the ions passes through the cell unhindered (no gas - classical ICP-MS mode), or collisions/reactions chemistry takes place between the ions and a neutral gas such as hydrogen or helium. Any polyatomic ions which give interference on the analytes are removed. The ions are then sorted according to their mass-to-charge ratios by a quadrupole mass spectrometer. The ions are then sorted according to their mass-to-charge ratios by a quadrupole mass spectrometer. The ions transmitted through the quadrupole are quantified by a channel electron multiplier and the ion information processed by a data handling system. Although most polyatomic interference is removed, any additional interference must be assessed and valid corrections applied or the data flagged to indicate problems. Interference correction must include compensation for background ions contributed by the plasma gas, reagents, and constituents of the sample matrix. Instrument drift as well as suppressions or enhancements of instrument response caused by the sample matrix must be corrected for by the use of internal standards. | 0.0001 | Y | Y | Y | 432 | 8/12/2009 12:08:00 AM | 4/29/2024 | 6/7/2024 10:03:16 AM | 2009-09-08 |  | RDGIBSON |  | Concentration | mg/L | Y | (fl. conc.) |  | Copper Total (fl. conc.) | 7440-50-8 | Metals - Total; | Metals - Total; EMS code: CU-T | Chemical | Total | LAB
CU-T | ug/m3 | Copper Total | 6040 | Aqua Regia Dig:ICP | Aqua regia digestion is used to bring the metals into solution prior to ICP analysis. For Hi-Vol samples, preliminary results are reported as mg/L and24 hr loadings as ug/m3 | 0.4 | Y | Y | Y | 437 | 2/13/2002 | 2/9/2005 | 2/23/2005 1:13:47 AM | 1998-05-05 | 6/3/1998 6:51:48 PM | NPEPPIN |  | AirConcentration | µg/m³ | Y | (air conc.) |  | Copper Total (air conc.) | 7440-50-8 | Metals - Total; | Metals - Total; EMS code: CU-T | Chemical | Total | LAB
CU-T | mg/dm2/d | Copper Total | EQ15 | EQUIS (mg/dm2/d) | EQUIS (mg/dm2/d) | 0.001 | N | Y | Y | 151 |  |  | 10/22/1996 10:18:48 PM | 1996-04-12 | 6/3/1998 6:51:35 PM | EMS |  | MassFlux | mg/dm²/d | Y | (mass flux) |  | Copper Total (mass flux) | 7440-50-8 | Metals - Total; | Metals - Total; EMS code: CU-T | Chemical | Total | LAB
CU-T | mg/L | Copper Total | X028 | FA,HNO3,AA Analysis | FA;HNO3;AA Analysis | 0.001 | Y | Y | Y | 259 |  |  | 8/26/2004 4:21:02 PM | 1996-04-12 | 6/3/1998 6:51:35 PM | EMS |  | Concentration | mg/L | Y | (fl. conc.) |  | Copper Total (fl. conc.) | 7440-50-8 | Metals - Total; | Metals - Total; EMS code: CU-T | Chemical | Total | LAB
CU-T | ppm (S) | Copper Total | MS41 | ALS Minerals ME-MS41 | ALS Minerals ME-MS41. A prepared sample is digested with aqua regia in a graphite heating block. After cooling; the resulting solution is diluted to with deionized water; mixed and analyzed by ICP-AES. Samples are then analyzed by ICP-MS for the remaining suite of elements. | 0.2 | Y | Y | Y | 0 |  |  |  | 2024-01-29 |  | EMS |  | DimensionlessRatio | ppm | N | (ratio) |  | Copper Total (ratio) | 7440-50-8 | Metals - Total; | Metals - Total; EMS code: CU-T | Chemical | Total | LAB
CU-T | mg/L | Copper Total | 5078 | N/P Digestion/ ICP Anal. | N/P Digestion/ ICP Anal. | 0.01 | N | Y | Y | 32 |  |  | 10/24/1996 2:10:29 PM | 1996-10-15 | 6/3/1998 6:51:46 PM | EMS |  | Concentration | mg/L | Y | (fl. conc.) |  | Copper Total (fl. conc.) | 7440-50-8 | Metals - Total; | Metals - Total; EMS code: CU-T | Chemical | Total | LAB
CU-T | mg/L | Copper Total | X113 | HNO3/HCl/Perchlor,AA | HNO3/HCl/Perchlor;AA | 0.001 | Y | Y | Y | 734 |  |  | 10/24/1996 1:40:59 AM | 1996-04-12 | 6/3/1998 6:51:35 PM | EMS |  | Concentration | mg/L | Y | (fl. conc.) |  | Copper Total (fl. conc.) | 7440-50-8 | Metals - Total; | Metals - Total; EMS code: CU-T | Chemical | Total | LAB
CU-T | ug/g | Copper Total | X551 | Metals in Biota (dry), total, acid digestion, ICP-SFMS | 0.1 gram of dried, homogenized biota is weighed in a microwave TFM vessel, and 8mL of Nitric Acid and 1mL of Hydrogen Peroxide are added. The sample is digested at 200 degrees C for 15 minutes in a closed vessel high pressure microwave oven. The sample is cooled, the volume adjusted to 100mL and introduced to an iductively coupled argon plasma-sector field mass spectrometer (ICP-SFMS) by pneumatic nebulization. Each respective element is measured at a specific mass to charge ratio, m/z value expressed in amu (atomic mass units), and at a specified instrument resolution. LR: Low Resolution ~300, MR: Mid Resolution ~4000 and HR: Hich Resolution ~10000. Ref: Method number 02-2705, National Laboratory for Environmental Testing, CCIW, Environment Canada, Burlington. Requested April 2003. | 0.1 | Y | Y | Y | 49 | 8/11/2015 | 10/30/2023 | 11/28/2023 2:21:33 PM | 2004-01-19 |  | UMULLER |  | MassPerMass | mg/kg | Y | (mass conc.) |  | Copper Total (mass conc.) | 7440-50-8 | Metals - Total; | Metals - Total; EMS code: CU-T | Chemical | Total | LAB
CU-T | ug/g | Copper Total | ICPM | Metals in Soil by CRC ICPMS | This analysis is carried out using procedures from CSR Analytical Method: "Strong Acid Leachable Metals (SALM) in Soil", BC Ministry of Environment, 26 June 2009, and procedures adapted from EPA Method 200.2. The sample is manually homogenized, dried at 60 degrees Celsius, sieved through a 63 um (230 mesh) sieve, and a representative subsample of the dry material is weighed. The sample is then digested at 95 degrees Celsius for 2 hours by block digester using concentrated nitric and hydrochloric acids. Instrumental analysis is by collision cell inductively coupled plasma - mass spectrometry (modifed from EPA Method 6020A). Method Limitation: This method is not a total digestion technique. It is a very strong acid digestion that is intended to dissolve those metals that may be environmentally available. By design, elements bound in silicate structures are not normally dissolved by this procedure as they are not usually mobile in the environment. | 0.5 | Y | Y | Y | 103 | 9/19/2017 | 9/24/2023 | 1/12/2024 11:16:52 AM | 2018-01-09 |  | JKROGH |  | MassPerMass | mg/kg | Y | (mass conc.) |  | Copper Total (mass conc.) | 7440-50-8 | Metals - Total; | Metals - Total; EMS code: CU-T | Chemical | Total | LAB
CU-T | mg/L | Copper Total | ULTR | Total Metals in Water by ICPMS (Ultra) | Total Metals in Water by ICPMS (Ultra), EPA SW-846 3005A/6020A, MET-TOT-ULTRA-MS-VA This analysis is carried out using procedures adapted from "Standard Methods for the Examination of Water and Wastewater" published by the American Public Health Association, and with procedures adapted from "Test Methods for Evaluating Solid Waste" SW-846 published by the United States Environmental Protection Agency (EPA). The procedures may involve preliminary sample treatment by acid digestion, using either hotblock or microwave oven, or filtration (EPA Method 3005A). Instrumental analysis is by inductively coupled plasma - mass spectrometry (EPA Method 6020A). | 0.00015 | Y | Y | Y | 1 | 12/28/2017 | 12/28/2017 | 1/9/2018 11:06:10 AM | 2018-01-09 |  | JKROGH |  | Concentration | mg/L | Y | (fl. conc.) |  | Copper Total (fl. conc.) | 7440-50-8 | Metals - Total; | Metals - Total; EMS code: CU-T | Chemical | Total | LAB
CU-T | ug/g wet | Copper Total | MTIW | MET-WET-CCMS-MID-VA | Metals in Tissue by CRC ICPMS (WET) Methodology: This method is conducted following British Columbia Lab Manual method "Metals in Animal Tissue and Vegetation (Biota) - Prescriptive". Tissue samples are homogenized and sub-sampled prior to hotblock digestion with nitric and hydrochloric acids, in combination with addition of hydrogen peroxide. Instrumental analysis is by collision cell inductively coupled plasma - mass spectrometry (modified from EPA Method 6020A). Method Limitation: This method employs a strong acid/peroxide digestion, and is intended to provide a conservative estimate of bio-available metals. Near complete recoveries are achieved for most toxicologically important metals, but elements associated with recalcitrant minerals may be only partially recovered. | 0.04 | Y | Y | Y | 24 | 11/25/2018 | 12/6/2018 | 7/21/2021 2:55:07 PM | 2018-12-24 |  | EMS |  | MassPerMass | µg/g | Y | (mass conc.) | (wet) | Copper Total (wet) (mass conc.) | 7440-50-8 | Metals - Total; | Metals - Total; EMS code: CU-T | Chemical | Total | LAB
CU-T | mg/L | Copper Total | F048 | Acid dig,conc'n,nebul.,ICP-OES | This method provides an automated procedure for the Inductively Coupled Argon Plasma Optical Emission Spectrophotometric (ICP-OES) determination and quantitation of trace amounts of metals in surface, ground and precipitation waters after a manual acid digestion and concentration procedure and ultrasonic nebulisation. Organically-bound metals and metal complexes are converted to metal ions by digestion with nitric acid [method details says HCl]. To increase the signal detected by the ICP, the samples are concentrated 10 times. The liquid sample is pumped onto the piezoelectric transducer, where the ultrasonic action efficiently converts it to a fine, dense aerosol. The nebulizer carrier gas flow then transports the wet aerosol through a heated U-tube, where the solvent is vaporized. Subsequently, most of the solvent vapour is removed from the aerosol stream by passing the stream through a chilled condenser. The result is a dry, analyte-laden aerosol which is introduced to the plasma. There, the high temperature desolvates the aerosol and dissociates the sample material into individual atoms and ions, which are excited to emit light at wave lengths characteristic of the elements present. Dissolved samples are field filtered through 0.45 micron cellulose acetate filter prior to acid preservation. Samples for extractable analysis are acidified with 2 mL concentrated HNO3 per litre and allowed to stand a minimum of 24 hours before analysis. Clients providing these samples are advised to preserve samples at the collection site. | 0.0002 | Y | Y | Y | 17153 | 7/4/1997 2:36:00 PM | 4/23/2003 | 9/11/2019 8:46:38 AM | 1997-10-20 | 6/3/1998 6:51:47 PM | NPEPPIN |  | Concentration | mg/L | Y | (fl. conc.) |  | Copper Total (fl. conc.) | 7440-50-8 | Metals - Total; | Metals - Total; EMS code: CU-T | Chemical | Total | LAB
CU-T | mg/L | Copper Total | F082 | Aliquot,HNO3/HCL dig,ICPMS | Aliquot microwave digested with 3:1 HNO3/HCL. Analysis by ICPMS. For HiVol samples preliminary result are reported in mg/L and loading results are reported in ug/m3 | 0.001 | Y | Y | Y | 28830 | 7/26/2005 7:04:00 AM | 8/23/2024 | 8/26/2024 4:53:05 PM | 2000-06-13 | 8/15/2019 11:07:36 AM | NPEPPIN | KATFRASE | Concentration | mg/L | Y | (fl. conc.) |  | Copper Total (fl. conc.) | 7440-50-8 | Metals - Total; | Metals - Total; EMS code: CU-T | Chemical | Total | LAB
CU-T | mg/L | Copper Total | 0100 | Dig:Ext.(APDC):AA Anal. | Dig:Ext.(APDC):AA Anal. | 0.001 | Y | Y | Y | 1213 |  |  | 6/16/2000 4:19:25 PM | 1996-04-12 | 6/3/1998 6:51:35 PM | EMS |  | Concentration | mg/L | Y | (fl. conc.) |  | Copper Total (fl. conc.) | 7440-50-8 | Metals - Total; | Metals - Total; EMS code: CU-T | Chemical | Total | LAB
CU-T | mg/L | Copper Total | F121 | Acid,ICPMS | Acidified, no digestion, ICPMS | 5e-05 | Y | Y | Y | 8554 | 3/12/2010 7:57:00 PM | 8/15/2024 9:10:00 PM | 8/26/2024 8:06:03 AM | 2002-07-17 | 7/16/2012 3:26:54 PM | UMULLER | RDGIBSON | Concentration | mg/L | Y | (fl. conc.) |  | Copper Total (fl. conc.) | 7440-50-8 | Metals - Total; | Metals - Total; EMS code: CU-T | Chemical | Total | LAB
CU-T | ug/m3 | Copper Total | 5312 | Tef:N/P Dig:ICP Analysis | A nitric/perchloric acid digestion if used to bring the metals into solution prior to ICP analysis. For Hi-Vol samples preliminary results are reported as mg/L and 24 hour loading results as ug/m3. | 0.4 | Y | Y | Y | 1688 | 11/7/1997 | 5/15/2000 | 11/18/2000 1:29:35 AM | 1996-04-12 | 6/3/1998 6:51:41 PM | EMS |  | AirConcentration | µg/m³ | Y | (air conc.) |  | Copper Total (air conc.) | 7440-50-8 | Metals - Total; | Metals - Total; EMS code: CU-T | Chemical | Total | LAB
CU-T | mg/L | Copper Total | METN | MET-T-NP-CCMS-VA | Ultra trace metals in water are analyzed by CRC ICPMS, based on US EPA Method 6020B (July 2014). The detection limits provided can only be met for undigested samples. This procedure is intended for colorless, non-turbid, acid-preserved water samples (i.e. pristine water samples), having turbidity < 1 NTU and no odor. Where turbidity exceeds 1 NTU, and/or the sample is colored and has an odor, results may be biased low compared to true Total Metals concentrations. ALS recommends that turbidity analysis be requested on samples submitted for this test to aid with interpretation of results. Where turbidity is <1NTU, undigested metals are equivalent to total metals concentrations. | 5e-05 | Y | Y | Y | 3648 | 4/3/2012 | 8/19/2024 | 8/26/2024 1:03:06 PM | 2019-08-20 |  | KATFRASE |  | Concentration | mg/L | Y | (fl. conc.) |  | Copper Total (fl. conc.) | 7440-50-8 | Metals - Total; | Metals - Total; EMS code: CU-T | Chemical | Total | LAB
CU-T | mg/L | Copper Total | X555 | Metals in water, total, "in-bottle" digestion, ICP-SFMS | The whole water sample collected in a HDPE/Teflon bottle, is de-capped in a Class 100 workstation, acidified to 2% (v/v) HNO3 and 1% (v/v) HCl and recapped. The sample is "in-bottle" digested in a convection oven, set at 60 degrees C for 16 hours. The sample is cooled and introduced to an inductively coupled argon plasma-sector field mass spectrometer (ICP-SFMS) by pneumatic nebulization. Each respective element is measured at a specific mass to charge ratio, m/z value expressed in amu (atomic mass units), and at resolution. LR: Low Resolution ~300, MR: Mid Resolution ~4000 and HR: High a specified instrument Resolution ~10000. Ref: Method number 02-2003, National Laboratory for Environmental Testing, CCIW, Environment Canada, Burlington. Requested April 2003. | 2e-05 | Y | Y | Y | 17336 | 5/20/2003 | 12/6/2023 2:08:00 AM | 2/14/2024 9:05:32 AM | 2004-01-19 |  | UMULLER |  | Concentration | mg/L | Y | (fl. conc.) |  | Copper Total (fl. conc.) | 7440-50-8 | Metals - Total; | Metals - Total; EMS code: CU-T | Chemical | Total | LAB
CU-T | ppm (S) | Copper Total | 6042 | Aqua Regia by ICPMS | These were sediment/soils digested with Aqua-Regia (Nitric Acid & Hydrochloric Acids) and then analysis by ICPMS. Results would be reported in ppm s. | 0.01 | Y | Y | Y | 71 | 5/14/2022 | 7/27/2024 | 7/29/2024 9:13:08 AM | 2022-06-28 |  | EMS |  | DimensionlessRatio | ppm | N | (ratio) |  | Copper Total (ratio) | 7440-50-8 | Metals - Total; | Metals - Total; EMS code: CU-T | Chemical | Total | LAB
CU-T | mg/L | Copper Total | X072 | FA,HNO3,Graph.Furnace,AA | Samples are field-acidified with HNO3; HNO3 digestion; analysis by GFAA. | 0.001 | Y | Y | Y | 11771 |  |  | 8/28/2018 3:14:03 PM | 1996-04-12 | 6/3/1998 6:51:35 PM | EMS |  | Concentration | mg/L | Y | (fl. conc.) |  | Copper Total (fl. conc.) | 7440-50-8 | Metals - Total; | Metals - Total; EMS code: CU-T | Chemical | Total | LAB
CYAN | mg/L | Cyanate Dissolved | X044 | Ion Chromatography | A clear (Total) or filtered (Dissolved) water sample is injected into a stream of carbonate-bicarbonate eluent and passed through a series of ion exchanger. The anions of interested are separated on the basis of their relative affinities for a low capacity , strongly basic anion exchanger (guard and separate columns). The separated anions are directed through the anion self-regenerating suppressor (ASRS). Hydronium ions in the suppressor will neutralize the eluent to its weakly ionized form and convert the separated anions in their conductive acid form. The separated anions are measured by conductivity. They are identified on the basis of retention time as compared to standards. Quantitation is by measurement of peak area of peak height against a calibration curve generated from known standards. | 0.01 | Y | Y | Y | 1447 | 5/18/2000 7:12:00 AM | 8/9/2024 | 8/26/2024 3:51:07 PM | 1996-04-12 | 6/3/1998 6:51:34 PM | EMS |  | Concentration | mg/L | Y | (fl. conc.) |  | Cyanate Dissolved (fl. conc.) | 661-20-1 | Inorganics; | Inorganics; EMS code: CYAN | Chemical | Dissolved | LAB
CYNZ | mg/L | Cyanazine | E660 | EPA 8270E (mod) | EPA 8270E (mod) Miscellaneous Pesticides by GC-MS | 0.0001 | Y | Y | Y | 68 | 12/7/2021 | 9/12/2022 | 12/8/2022 12:57:05 PM | 2021-12-09 | 12/9/2021 3:22:12 PM | EMS | EMS | Concentration | mg/L | Y | (fl. conc.) |  | Cyanazine (fl. conc.) | 21725-46-2 | Pesticides; | Pesticides; EMS code: CYNZ | Chemical |  | LAB
CYNZ | ug/g | Cyanazine | ZN03 | DCM+Ac.Extr:GC/MS | DCM+Ac.Extr:GC/MS EPA 8270E (mod) | 0.05 | Y | Y | Y | 16 | 1/25/2022 | 11/2/2022 4:41:00 AM | 1/7/2023 3:12:10 PM | 1996-04-12 | 11/16/1999 2:27:40 PM | EMS | NPEPPIN | MassPerMass | mg/kg | Y | (mass conc.) |  | Cyanazine (mass conc.) | 21725-46-2 | Pesticides; | Pesticides; EMS code: CYNZ | Chemical |  | LAB
CYNZ | mg/L | Cyanazine | ZN01 | DCM+Ac.Extr:GC/MS | DCM+Ac.Extr:GC/MS | 0.0005 | Y | Y | Y | 28 | 7/28/2005 7:03:00 AM | 3/29/2012 10:15:00 PM | 5/3/2012 9:27:50 AM | 1996-04-12 | 11/16/1999 4:06:27 PM | EMS | NPEPPIN | Concentration | mg/L | Y | (fl. conc.) |  | Cyanazine (fl. conc.) | 21725-46-2 | Pesticides; | Pesticides; EMS code: CYNZ | Chemical |  | LAB
D001 | ug/g | Dimethoate | ZN03 | DCM+Ac.Extr:GC/MS | DCM+Ac.Extr:GC/MS EPA 8270E (mod) | 0.05 | Y | Y | Y | 16 | 1/25/2022 | 11/2/2022 4:41:00 AM | 1/7/2023 3:12:10 PM | 1996-04-12 | 11/16/1999 2:27:40 PM | EMS | NPEPPIN | MassPerMass | mg/kg | Y | (mass conc.) |  | Dimethoate (mass conc.) | 60-51-5 | Pesticides; | Pesticides; EMS code: D001 | Chemical |  | LAB
D001 | mg/L | Dimethoate | ZN01 | DCM+Ac.Extr:GC/MS | DCM+Ac.Extr:GC/MS | 0.0005 | Y | Y | Y | 52 | 10/3/2002 | 3/29/2012 10:15:00 PM | 5/3/2012 9:27:50 AM | 1996-04-12 | 11/16/1999 4:06:27 PM | EMS | NPEPPIN | Concentration | mg/L | Y | (fl. conc.) |  | Dimethoate (fl. conc.) | 60-51-5 | Pesticides; | Pesticides; EMS code: D001 | Chemical |  | LAB
D001 | ug/g | Dimethoate | X260 | Sox-Ac/CH2Cl2 ext:GC/NPD | Sox-Ac/CH2Cl2 ext:GC/NPD | 0.02 | Y | Y | Y | 117 | 4/20/2002 10:25:00 AM | 10/16/2002 3:20:00 PM | 2/26/2004 1:05:27 AM | 1996-04-12 | 6/3/1998 6:51:35 PM | EMS |  | MassPerMass | mg/kg | Y | (mass conc.) |  | Dimethoate (mass conc.) | 60-51-5 | Pesticides; | Pesticides; EMS code: D001 | Chemical |  | LAB
D001 | mg/L | Dimethoate | F053 | DCM extr'n,conc'n,HGRC-PFPD/N,P det'n | Effluent samples are extracted with Dichloromethane. The extract is then concentrated and analysed using High Resolution Gas Chromatography with Pulsed Flame Photometric Dectection. Confirmation, if required, is by High Resolution Gas Chromatography with Nitrogen Phosphorus Detection. | 5e-05 | Y | Y | Y | 31 | 3/3/2000 7:36:00 AM | 1/22/2002 10:26:00 AM | 6/18/2003 2:53:38 PM | 1997-10-20 | 6/3/1998 6:51:47 PM | NPEPPIN |  | Concentration | mg/L | Y | (fl. conc.) |  | Dimethoate (fl. conc.) | 60-51-5 | Pesticides; | Pesticides; EMS code: D001 | Chemical |  | LAB
D001 | mg/L | Dimethoate | E660 | EPA 8270E (mod) | EPA 8270E (mod) Miscellaneous Pesticides by GC-MS | 0.0001 | Y | Y | Y | 68 | 12/7/2021 | 9/12/2022 | 12/8/2022 12:57:05 PM | 2021-12-09 | 12/9/2021 3:22:12 PM | EMS | EMS | Concentration | mg/L | Y | (fl. conc.) |  | Dimethoate (fl. conc.) | 60-51-5 | Pesticides; | Pesticides; EMS code: D001 | Chemical |  | LAB
D002 | mg/L | delta-BHC | E660 | EPA 8270E (mod) | EPA 8270E (mod) Miscellaneous Pesticides by GC-MS | 5e-06 | Y | Y | Y | 2 | 9/14/2022 | 9/14/2022 | 12/8/2022 9:48:13 AM | 2022-12-06 |  | EMS |  | Concentration | mg/L | Y | (fl. conc.) |  | delta-BHC (fl. conc.) | 319-86-8 | Pesticides; | Pesticides; EMS code: D002 | Chemical |  | LAB
D002 | mg/L | delta-BHC | X268 | CH2Cl2 ext:GC/MS | CH2Cl2 ext:GC/MS |  | Y | Y | Y | 16 | 12/29/2021 | 9/2/2022 | 11/10/2022 9:54:06 AM | 2022-04-08 |  | EMS |  | Concentration | mg/L | Y | (fl. conc.) |  | delta-BHC (fl. conc.) | 319-86-8 | Pesticides; | Pesticides; EMS code: D002 | Chemical |  | LAB
D002 | ug/g | delta-BHC | E8081 | EPA 8081 | See methods EPA 3540, 3630, 3660, 8081 | 0.001 | Y | Y | Y | 9 | 3/18/2020 | 3/23/2020 | 2/22/2023 12:51:03 PM | 2022-07-29 |  | EMS |  | MassPerMass | mg/kg | Y | (mass conc.) |  | delta-BHC (mass conc.) | 319-86-8 | Pesticides; | Pesticides; EMS code: D002 | Chemical |  | LAB
D002 | mg/L | delta-BHC | MW03 | OCP1-LL-SF-ECD-VA | This analysis is carried out using procedures adapted from "Test Methods for Evaluating Solid Waste" SW-846, Methods 3510, 3630, 3660 and 8081, published by the United States Environmental Protection Agency (EPA). The procedure involves extraction of the entire water sample with dichloromethane. The extract then undergoes one or more of the following clean-up procedures (if required): silica gel clean-up and/or sulphur clean-up. The final extract is analyzed by dual capillary column gas chromatography with electron capture detection (GC/ECD). | 5e-05 | Y | Y | Y | 10 | 9/20/2016 | 9/21/2016 | 8/15/2019 2:42:34 PM | 2019-07-08 |  | KATFRASE |  | Concentration | mg/L | Y | (fl. conc.) |  | delta-BHC (fl. conc.) | 319-86-8 | Pesticides; | Pesticides; EMS code: D002 | Chemical |  | LAB
D002 | ug/g | delta-BHC | ZN03 | DCM+Ac.Extr:GC/MS | DCM+Ac.Extr:GC/MS EPA 8270E (mod) | 0.02 | Y | Y | Y | 17 | 1/17/2022 | 11/14/2022 | 1/11/2024 1:07:07 PM | 2022-06-23 |  | EMS |  | MassPerMass | mg/kg | Y | (mass conc.) |  | delta-BHC (mass conc.) | 319-86-8 | Pesticides; | Pesticides; EMS code: D002 | Chemical |  | LAB
D007 | mg/L | Dicloran | ZN01 | DCM+Ac.Extr:GC/MS | DCM+Ac.Extr:GC/MS | 0.0001 | Y | Y | Y | 28 | 7/29/2005 7:10:00 AM | 3/29/2012 10:15:00 PM | 5/3/2012 9:27:50 AM | 1996-04-12 | 11/16/1999 4:06:27 PM | EMS | NPEPPIN | Concentration | mg/L | Y | (fl. conc.) |  | Dicloran (fl. conc.) |  | Pesticides; | Pesticides; EMS code: D007 | Chemical |  | LAB
D009 | mg/L | Diclofop | AEWA | H2SO4 acid,CHCl2 extr.florisil. GC/MS | Effluent and water samples are acidified with Sulphuric Acid and extracted with Dichloromethane. After concentration, the extract is methylated using ethereal Diazomethane. Extracts are cleaned up as required using Florisil. The extracts are analyzed for Acid Extractable Herbicides using GC/MS. | 0.0005 | Y | Y | Y | 1 | 7/28/2005 7:03:00 AM | 7/28/2005 7:03:00 AM | 6/2/2009 12:01:02 PM | 2003-06-17 |  | UMULLER |  | Concentration | mg/L | Y | (fl. conc.) |  | Diclofop (fl. conc.) |  | Pesticides; | Pesticides; EMS code: D009 | Chemical |  | LAB
D010 | mg/L | Diazinon | F053 | DCM extr'n,conc'n,HGRC-PFPD/N,P det'n | Effluent samples are extracted with Dichloromethane. The extract is then concentrated and analysed using High Resolution Gas Chromatography with Pulsed Flame Photometric Dectection. Confirmation, if required, is by High Resolution Gas Chromatography with Nitrogen Phosphorus Detection. | 5e-05 | Y | Y | Y | 31 | 3/3/2000 7:36:00 AM | 1/22/2002 10:26:00 AM | 6/18/2003 2:53:38 PM | 1997-10-20 | 6/3/1998 6:51:47 PM | NPEPPIN |  | Concentration | mg/L | Y | (fl. conc.) |  | Diazinon (fl. conc.) | 333-41-5 | Pesticides; | Pesticides; EMS code: D010 | Chemical |  | LAB
D010 | mg/L | Diazinon | E660 | EPA 8270E (mod) | EPA 8270E (mod) Miscellaneous Pesticides by GC-MS | 0.0001 | Y | Y | Y | 68 | 12/7/2021 | 9/12/2022 | 12/8/2022 12:57:05 PM | 2021-12-09 | 12/9/2021 3:22:12 PM | EMS | EMS | Concentration | mg/L | Y | (fl. conc.) |  | Diazinon (fl. conc.) | 333-41-5 | Pesticides; | Pesticides; EMS code: D010 | Chemical |  | LAB
D010 | ug/g | Diazinon | ZN03 | DCM+Ac.Extr:GC/MS | DCM+Ac.Extr:GC/MS EPA 8270E (mod) | 0.03 | Y | Y | Y | 16 | 1/25/2022 | 11/2/2022 4:41:00 AM | 1/7/2023 3:12:10 PM | 1996-04-12 | 11/16/1999 2:27:40 PM | EMS | NPEPPIN | MassPerMass | mg/kg | Y | (mass conc.) |  | Diazinon (mass conc.) | 333-41-5 | Pesticides; | Pesticides; EMS code: D010 | Chemical |  | LAB
D010 | ug/g | Diazinon | X260 | Sox-Ac/CH2Cl2 ext:GC/NPD | Sox-Ac/CH2Cl2 ext:GC/NPD | 0.02 | Y | Y | Y | 117 | 4/20/2002 10:25:00 AM | 10/16/2002 3:20:00 PM | 2/26/2004 1:05:27 AM | 1996-04-12 | 6/3/1998 6:51:35 PM | EMS |  | MassPerMass | mg/kg | Y | (mass conc.) |  | Diazinon (mass conc.) | 333-41-5 | Pesticides; | Pesticides; EMS code: D010 | Chemical |  | LAB
D010 | mg/L | Diazinon | EQ01 | EQUIS (mg/L) | EQUIS (mg/L) | 0.001 | N | Y | Y | 2 |  |  | 10/22/1996 5:23:43 PM | 1996-04-12 | 6/3/1998 6:51:35 PM | EMS |  | Concentration | mg/L | Y | (fl. conc.) |  | Diazinon (fl. conc.) | 333-41-5 | Pesticides; | Pesticides; EMS code: D010 | Chemical |  | LAB
D010 | mg/L | Diazinon | ZN01 | DCM+Ac.Extr:GC/MS | DCM+Ac.Extr:GC/MS | 0.0003 | Y | Y | Y | 52 | 10/3/2002 | 3/29/2012 10:15:00 PM | 5/3/2012 9:27:50 AM | 1996-04-12 | 11/16/1999 4:06:27 PM | EMS | NPEPPIN | Concentration | mg/L | Y | (fl. conc.) |  | Diazinon (fl. conc.) | 333-41-5 | Pesticides; | Pesticides; EMS code: D010 | Chemical |  | LAB
D017 | ug/g | Dichlorvos | ZN03 | DCM+Ac.Extr:GC/MS | DCM+Ac.Extr:GC/MS EPA 8270E (mod) | 0.01 | Y | Y | Y | 21 | 4/20/2002 10:25:00 AM | 11/2/2022 4:41:00 AM | 1/7/2023 3:12:10 PM | 1996-04-12 | 11/17/1999 2:20:19 PM | EMS | NPEPPIN | MassPerMass | mg/kg | Y | (mass conc.) |  | Dichlorvos (mass conc.) | 62-73-7 | Pesticides; | Pesticides; EMS code: D017 | Chemical |  | LAB
D017 | mg/L | Dichlorvos | ZN01 | DCM+Ac.Extr:GC/MS | DCM+Ac.Extr:GC/MS | 0.0005 | Y | Y | Y | 5 | 10/3/2002 | 7/29/2005 7:10:00 AM | 6/2/2009 12:01:02 PM | 1996-04-12 | 6/3/1998 6:51:42 PM | EMS |  | Concentration | mg/L | Y | (fl. conc.) |  | Dichlorvos (fl. conc.) | 62-73-7 | Pesticides; | Pesticides; EMS code: D017 | Chemical |  | LAB
D019 | mg/L | p;p'-DDD | F027 | Hexane Extr,GC/ECD | Effluent samples are solvent extracted with Hexane and the extract dried over Sodium Sulfate, concentrated and cleaned up, as required, using Florisil, activated Copper metal, concentrated Sulfuric Acid and Mercury, to remove interferences. The extracts are then analysed using megabore column Gas Chromatography with Electron Capture detection. | 5e-05 | Y | Y | Y | 20 | 3/3/2000 7:36:00 AM | 1/22/2002 10:58:00 AM | 6/18/2003 2:53:39 PM | 1997-01-14 | 6/3/1998 6:51:46 PM | MPAK |  | Concentration | mg/L | Y | (fl. conc.) |  | p;p'-DDD (fl. conc.) |  | Pesticides; | Pesticides; EMS code: D019 | Chemical |  | LAB
D019 | ug/g | p;p'-DDD | FS10 | Acetone Extr,GC/ECD | Sediment and homogenized biota samples are Acetone extracted using a wrist-action shaker, Hexane partitioned using a Sodium Chloride solution, then treated in the manner outlined for effluent samples. The extracts are analysed using megabore column Gas Chromatography with Electron Capture Detection. | 0.002 | Y | Y | Y | 26 | 4/20/2002 10:10:00 AM | 10/13/2002 10:30:00 AM | 2/26/2004 1:05:27 AM | 1997-01-14 | 6/3/1998 6:51:46 PM | MPAK |  | MassPerMass | mg/kg | Y | (mass conc.) |  | p;p'-DDD (mass conc.) |  | Pesticides; | Pesticides; EMS code: D019 | Chemical |  | LAB
D021 | mg/L | p;p'-DDT | F027 | Hexane Extr,GC/ECD | Effluent samples are solvent extracted with Hexane and the extract dried over Sodium Sulfate, concentrated and cleaned up, as required, using Florisil, activated Copper metal, concentrated Sulfuric Acid and Mercury, to remove interferences. The extracts are then analysed using megabore column Gas Chromatography with Electron Capture detection. | 5e-05 | Y | Y | Y | 20 | 3/3/2000 7:36:00 AM | 1/22/2002 10:58:00 AM | 6/18/2003 2:53:39 PM | 1997-01-14 | 6/3/1998 6:51:46 PM | MPAK |  | Concentration | mg/L | Y | (fl. conc.) |  | p;p'-DDT (fl. conc.) |  | Pesticides; | Pesticides; EMS code: D021 | Chemical |  | LAB
D021 | mg/L | p;p'-DDT | ZN01 | DCM+Ac.Extr:GC/MS | DCM+Ac.Extr:GC/MS | 0.0002 | Y | Y | Y | 52 | 10/3/2002 | 3/29/2012 10:15:00 PM | 5/3/2012 9:27:50 AM | 1996-04-12 | 11/16/1999 4:06:27 PM | EMS | NPEPPIN | Concentration | mg/L | Y | (fl. conc.) |  | p;p'-DDT (fl. conc.) |  | Pesticides; | Pesticides; EMS code: D021 | Chemical |  | LAB
D021 | ug/g | p;p'-DDT | ZN03 | DCM+Ac.Extr:GC/MS | DCM+Ac.Extr:GC/MS EPA 8270E (mod) | 0.02 | Y | Y | Y | 27 | 4/20/2002 10:10:00 AM | 10/13/2002 10:30:00 AM | 2/26/2004 1:05:27 AM | 1996-04-12 | 6/3/1998 6:51:35 PM | EMS |  | MassPerMass | mg/kg | Y | (mass conc.) |  | p;p'-DDT (mass conc.) |  | Pesticides; | Pesticides; EMS code: D021 | Chemical |  | LAB
D022 | mg/L | Dieldrin | E660 | EPA 8270E (mod) | EPA 8270E (mod) Miscellaneous Pesticides by GC-MS | 2e-06 | Y | Y | Y | 2 | 9/14/2022 | 9/14/2022 | 12/8/2022 9:48:13 AM | 2022-12-06 |  | EMS |  | Concentration | mg/L | Y | (fl. conc.) |  | Dieldrin (fl. conc.) | 60-57-1 | Pesticides; | Pesticides; EMS code: D022 | Chemical |  | LAB
D022 | mg/L | Dieldrin | MW03 | OCP1-LL-SF-ECD-VA | This analysis is carried out using procedures adapted from "Test Methods for Evaluating Solid Waste" SW-846, Methods 3510, 3630, 3660 and 8081, published by the United States Environmental Protection Agency (EPA). The procedure involves extraction of the entire water sample with dichloromethane. The extract then undergoes one or more of the following clean-up procedures (if required): silica gel clean-up and/or sulphur clean-up. The final extract is analyzed by dual capillary column gas chromatography with electron capture detection (GC/ECD). | 5e-05 | Y | Y | Y | 10 | 9/20/2016 | 9/21/2016 | 8/15/2019 2:42:34 PM | 2019-08-15 |  | KATFRASE |  | Concentration | mg/L | Y | (fl. conc.) |  | Dieldrin (fl. conc.) | 60-57-1 | Pesticides; | Pesticides; EMS code: D022 | Chemical |  | LAB
D022 | ug/g | Dieldrin | E8081 | EPA 8081 | See methods EPA 3540, 3630, 3660, 8081 | 0.001 | Y | Y | Y | 9 | 3/18/2020 | 3/23/2020 | 2/22/2023 12:51:03 PM | 2022-07-29 |  | EMS |  | MassPerMass | mg/kg | Y | (mass conc.) |  | Dieldrin (mass conc.) | 60-57-1 | Pesticides; | Pesticides; EMS code: D022 | Chemical |  | LAB
D022 | ug/g | Dieldrin | ZN03 | DCM+Ac.Extr:GC/MS | DCM+Ac.Extr:GC/MS EPA 8270E (mod) | 0.05 | Y | Y | Y | 19 | 1/17/2022 | 11/14/2022 | 1/11/2024 1:07:07 PM | 1996-04-12 | 11/16/1999 2:27:40 PM | EMS | NPEPPIN | MassPerMass | mg/kg | Y | (mass conc.) |  | Dieldrin (mass conc.) | 60-57-1 | Pesticides; | Pesticides; EMS code: D022 | Chemical |  | LAB
D022 | mg/L | Dieldrin | X268 | CH2Cl2 ext:GC/MS | CH2Cl2 ext:GC/MS |  | Y | Y | Y | 16 | 12/29/2021 | 9/2/2022 | 11/10/2022 9:54:06 AM | 2022-04-08 |  | EMS |  | Concentration | mg/L | Y | (fl. conc.) |  | Dieldrin (fl. conc.) | 60-57-1 | Pesticides; | Pesticides; EMS code: D022 | Chemical |  | LAB
D022 | ug/g | Dieldrin | X213 | Hex/AcExt,Part,Chrom,GC | Hex/AcExt;Part;Chrom;GC | 0.001 | Y | Y | Y | 187 | 4/20/2002 10:10:00 AM | 10/13/2002 10:30:00 AM | 2/26/2004 1:05:27 AM | 1996-04-12 | 6/3/1998 6:51:42 PM | EMS |  | MassPerMass | mg/kg | Y | (mass conc.) |  | Dieldrin (mass conc.) | 60-57-1 | Pesticides; | Pesticides; EMS code: D022 | Chemical |  | LAB
D022 | mg/L | Dieldrin | F027 | Hexane Extr,GC/ECD | Effluent samples are solvent extracted with Hexane and the extract dried over Sodium Sulfate, concentrated and cleaned up, as required, using Florisil, activated Copper metal, concentrated Sulfuric Acid and Mercury, to remove interferences. The extracts are then analysed using megabore column Gas Chromatography with Electron Capture detection. | 5e-05 | Y | Y | Y | 20 | 3/3/2000 7:36:00 AM | 1/22/2002 10:58:00 AM | 6/18/2003 2:53:39 PM | 2002-06-05 |  | UMULLER |  | Concentration | mg/L | Y | (fl. conc.) |  | Dieldrin (fl. conc.) | 60-57-1 | Pesticides; | Pesticides; EMS code: D022 | Chemical |  | LAB
D022 | ug/g wet | Dieldrin | X220 | Hex/AcExt,Part,Chrom,GC | Hex/AcExt;Part;Chrom;GC | 0.0002 | Y | Y | Y | 73 |  |  | 10/24/1996 10:01:25 AM | 1996-04-12 | 6/3/1998 6:51:45 PM | EMS |  | MassPerMass | µg/g | Y | (mass conc.) | (wet) | Dieldrin (wet) (mass conc.) | 60-57-1 | Pesticides; | Pesticides; EMS code: D022 | Chemical |  | LAB
D022 | mg/L | Dieldrin | EQ01 | EQUIS (mg/L) | EQUIS (mg/L) | 5e-05 | N | Y | Y | 17 |  |  | 10/22/1996 11:26:53 PM | 1996-04-12 | 6/3/1998 6:51:35 PM | EMS |  | Concentration | mg/L | Y | (fl. conc.) |  | Dieldrin (fl. conc.) | 60-57-1 | Pesticides; | Pesticides; EMS code: D022 | Chemical |  | LAB
D022 | mg/L | Dieldrin | ZN01 | DCM+Ac.Extr:GC/MS | DCM+Ac.Extr:GC/MS | 0.0005 | Y | Y | Y | 52 | 10/3/2002 | 3/29/2012 10:15:00 PM | 5/3/2012 9:27:50 AM | 1996-04-12 | 11/16/1999 4:06:27 PM | EMS | NPEPPIN | Concentration | mg/L | Y | (fl. conc.) |  | Dieldrin (fl. conc.) | 60-57-1 | Pesticides; | Pesticides; EMS code: D022 | Chemical |  | LAB
D023 | ug/g | DDE | X213 | Hex/AcExt,Part,Chrom,GC | Hex/AcExt;Part;Chrom;GC | 0.0005 | Y | Y | Y | 161 |  |  | 10/24/1996 2:26:15 PM | 1996-04-12 | 6/3/1998 6:51:35 PM | EMS |  | MassPerMass | mg/kg | Y | (mass conc.) |  | DDE (mass conc.) |  | Pesticides; | Pesticides; EMS code: D023 | Chemical |  | LAB
D023 | mg/L | DDE | E660 | EPA 8270E (mod) | EPA 8270E (mod) Miscellaneous Pesticides by GC-MS | 4e-07 | Y | Y | Y | 2 | 9/14/2022 | 9/14/2022 | 12/8/2022 9:48:13 AM | 2022-12-06 |  | EMS |  | Concentration | mg/L | Y | (fl. conc.) |  | DDE (fl. conc.) |  | Pesticides; | Pesticides; EMS code: D023 | Chemical |  | LAB
D023 | ug/g | DDE | E8081 | EPA 8081 | See methods EPA 3540, 3630, 3660, 8081 | 0.001 | Y | Y | Y | 9 | 3/18/2020 | 3/23/2020 | 2/22/2023 12:51:03 PM | 2022-07-29 |  | EMS |  | MassPerMass | mg/kg | Y | (mass conc.) |  | DDE (mass conc.) |  | Pesticides; | Pesticides; EMS code: D023 | Chemical |  | LAB
D023 | ug/g wet | DDE | X220 | Hex/AcExt,Part,Chrom,GC | Hex/AcExt;Part;Chrom;GC | 0.0001 | Y | Y | Y | 72 |  |  | 10/24/1996 10:01:25 AM | 1996-04-12 | 6/3/1998 6:51:42 PM | EMS |  | MassPerMass | µg/g | Y | (mass conc.) | (wet) | DDE (wet) (mass conc.) |  | Pesticides; | Pesticides; EMS code: D023 | Chemical |  | LAB
D023 | mg/L | DDE | X268 | CH2Cl2 ext:GC/MS | CH2Cl2 ext:GC/MS |  | Y | Y | Y | 16 | 12/29/2021 | 9/2/2022 | 11/10/2022 9:54:06 AM | 2022-04-08 |  | EMS |  | Concentration | mg/L | Y | (fl. conc.) |  | DDE (fl. conc.) |  | Pesticides; | Pesticides; EMS code: D023 | Chemical |  | LAB
D023 | mg/L | DDE | MW03 | OCP1-LL-SF-ECD-VA | This analysis is carried out using procedures adapted from "Test Methods for Evaluating Solid Waste" SW-846, Methods 3510, 3630, 3660 and 8081, published by the United States Environmental Protection Agency (EPA). The procedure involves extraction of the entire water sample with dichloromethane. The extract then undergoes one or more of the following clean-up procedures (if required): silica gel clean-up and/or sulphur clean-up. The final extract is analyzed by dual capillary column gas chromatography with electron capture detection (GC/ECD). | 5e-06 | Y | Y | Y | 10 | 9/20/2016 | 9/21/2016 | 8/15/2019 2:42:34 PM | 2019-08-15 |  | KATFRASE |  | Concentration | mg/L | Y | (fl. conc.) |  | DDE (fl. conc.) |  | Pesticides; | Pesticides; EMS code: D023 | Chemical |  | LAB
D023 | mg/L | DDE | EQ01 | EQUIS (mg/L) | EQUIS (mg/L) | 0.0001 | N | Y | Y | 17 |  |  | 10/22/1996 11:26:53 PM | 1996-04-12 | 6/3/1998 6:51:35 PM | EMS |  | Concentration | mg/L | Y | (fl. conc.) |  | DDE (fl. conc.) |  | Pesticides; | Pesticides; EMS code: D023 | Chemical |  | LAB
D023 | ug/g | DDE | ZN03 | DCM+Ac.Extr:GC/MS | DCM+Ac.Extr:GC/MS EPA 8270E (mod) | 0.01 | Y | Y | Y | 17 | 1/17/2022 | 11/14/2022 | 1/11/2024 1:07:07 PM | 1999-11-16 |  | NPEPPIN |  | MassPerMass | mg/kg | Y | (mass conc.) |  | DDE (mass conc.) |  | Pesticides; | Pesticides; EMS code: D023 | Chemical |  | LAB
D024 | mg/L | o;p'-DDT | ZN01 | DCM+Ac.Extr:GC/MS | DCM+Ac.Extr:GC/MS | 0.0002 | Y | Y | Y | 10 | 10/3/2002 | 11/6/2006 | 6/2/2009 12:01:03 PM | 1996-04-12 | 6/3/1998 6:51:42 PM | EMS |  | Concentration | mg/L | Y | (fl. conc.) |  | o;p'-DDT (fl. conc.) | 789-02-6 | Pesticides; | Pesticides; EMS code: D024 | Chemical |  | LAB
D024 | ug/g | o;p'-DDT | ZN03 | DCM+Ac.Extr:GC/MS | DCM+Ac.Extr:GC/MS EPA 8270E (mod) | 0.02 | Y | Y | Y | 27 | 4/20/2002 10:10:00 AM | 10/13/2002 10:30:00 AM | 2/26/2004 1:05:27 AM | 1996-04-12 | 11/16/1999 2:27:40 PM | EMS | NPEPPIN | MassPerMass | mg/kg | Y | (mass conc.) |  | o;p'-DDT (mass conc.) | 789-02-6 | Pesticides; | Pesticides; EMS code: D024 | Chemical |  | LAB
D025 | ug/g | DDD | E8081 | EPA 8081 | See methods EPA 3540, 3630, 3660, 8081 | 0.001 | Y | Y | Y | 9 | 3/18/2020 | 3/23/2020 | 2/22/2023 12:51:03 PM | 2022-07-29 |  | EMS |  | MassPerMass | mg/kg | Y | (mass conc.) |  | DDD (mass conc.) |  | Pesticides; | Pesticides; EMS code: D025 | Chemical |  | LAB
D025 | mg/L | DDD | EQ01 | EQUIS (mg/L) | EQUIS (mg/L) | 0.0001 | N | Y | Y | 17 |  |  | 10/22/1996 11:26:54 PM | 1996-04-12 | 6/3/1998 6:51:35 PM | EMS |  | Concentration | mg/L | Y | (fl. conc.) |  | DDD (fl. conc.) |  | Pesticides; | Pesticides; EMS code: D025 | Chemical |  | LAB
D025 | mg/L | DDD | X268 | CH2Cl2 ext:GC/MS | CH2Cl2 ext:GC/MS |  | Y | Y | Y | 16 | 12/29/2021 | 9/2/2022 | 11/10/2022 9:54:06 AM | 2022-04-08 |  | EMS |  | Concentration | mg/L | Y | (fl. conc.) |  | DDD (fl. conc.) |  | Pesticides; | Pesticides; EMS code: D025 | Chemical |  | LAB
D025 | mg/L | DDD | E660 | EPA 8270E (mod) | EPA 8270E (mod) Miscellaneous Pesticides by GC-MS | 4e-07 | Y | Y | Y | 2 | 9/14/2022 | 9/14/2022 | 12/8/2022 9:48:13 AM | 2022-12-06 |  | EMS |  | Concentration | mg/L | Y | (fl. conc.) |  | DDD (fl. conc.) |  | Pesticides; | Pesticides; EMS code: D025 | Chemical |  | LAB
D025 | ug/g | DDD | ZN03 | DCM+Ac.Extr:GC/MS | DCM+Ac.Extr:GC/MS EPA 8270E (mod) | 0.02 | Y | Y | Y | 17 | 1/17/2022 | 11/14/2022 | 1/11/2024 1:07:07 PM | 2022-06-21 |  | EMS |  | MassPerMass | mg/kg | Y | (mass conc.) |  | DDD (mass conc.) |  | Pesticides; | Pesticides; EMS code: D025 | Chemical |  | LAB
D025 | ug/g | DDD | X213 | Hex/AcExt,Part,Chrom,GC | Hex/AcExt;Part;Chrom;GC | 0.001 | Y | Y | Y | 161 |  |  | 10/24/1996 2:26:15 PM | 1996-04-12 | 6/3/1998 6:51:45 PM | EMS |  | MassPerMass | mg/kg | Y | (mass conc.) |  | DDD (mass conc.) |  | Pesticides; | Pesticides; EMS code: D025 | Chemical |  | LAB
D025 | mg/L | DDD | MW03 | OCP1-LL-SF-ECD-VA | This analysis is carried out using procedures adapted from "Test Methods for Evaluating Solid Waste" SW-846, Methods 3510, 3630, 3660 and 8081, published by the United States Environmental Protection Agency (EPA). The procedure involves extraction of the entire water sample with dichloromethane. The extract then undergoes one or more of the following clean-up procedures (if required): silica gel clean-up and/or sulphur clean-up. The final extract is analyzed by dual capillary column gas chromatography with electron capture detection (GC/ECD). | 5e-06 | Y | Y | Y | 10 | 9/20/2016 | 9/21/2016 | 8/15/2019 2:42:34 PM | 2019-08-15 |  | KATFRASE |  | Concentration | mg/L | Y | (fl. conc.) |  | DDD (fl. conc.) |  | Pesticides; | Pesticides; EMS code: D025 | Chemical |  | LAB
D025 | ug/g wet | DDD | X220 | Hex/AcExt,Part,Chrom,GC | Hex/AcExt;Part;Chrom;GC | 0.0002 | Y | Y | Y | 73 |  |  | 10/24/1996 10:01:25 AM | 1996-04-12 | 6/3/1998 6:51:42 PM | EMS |  | MassPerMass | µg/g | Y | (mass conc.) | (wet) | DDD (wet) (mass conc.) |  | Pesticides; | Pesticides; EMS code: D025 | Chemical |  | LAB
D026 | ug/g | DDT | X213 | Hex/AcExt,Part,Chrom,GC | Hex/AcExt;Part;Chrom;GC | 0.001 | Y | Y | Y | 161 |  |  | 10/24/1996 2:26:15 PM | 1996-04-12 | 6/3/1998 6:51:45 PM | EMS |  | MassPerMass | mg/kg | Y | (mass conc.) |  | DDT (mass conc.) |  | Pesticides; | Pesticides; EMS code: D026 | Chemical |  | LAB
D026 | ug/g | DDT | ZN03 | DCM+Ac.Extr:GC/MS | DCM+Ac.Extr:GC/MS EPA 8270E (mod) | 0.02 | Y | Y | Y | 17 | 1/17/2022 | 11/14/2022 | 1/11/2024 1:07:07 PM | 1999-11-16 |  | NPEPPIN |  | MassPerMass | mg/kg | Y | (mass conc.) |  | DDT (mass conc.) |  | Pesticides; | Pesticides; EMS code: D026 | Chemical |  | LAB
D026 | mg/L | DDT | MW03 | OCP1-LL-SF-ECD-VA | This analysis is carried out using procedures adapted from "Test Methods for Evaluating Solid Waste" SW-846, Methods 3510, 3630, 3660 and 8081, published by the United States Environmental Protection Agency (EPA). The procedure involves extraction of the entire water sample with dichloromethane. The extract then undergoes one or more of the following clean-up procedures (if required): silica gel clean-up and/or sulphur clean-up. The final extract is analyzed by dual capillary column gas chromatography with electron capture detection (GC/ECD). | 5e-06 | Y | Y | Y | 10 | 9/20/2016 | 9/21/2016 | 8/15/2019 2:42:34 PM | 2019-08-15 |  | KATFRASE |  | Concentration | mg/L | Y | (fl. conc.) |  | DDT (fl. conc.) |  | Pesticides; | Pesticides; EMS code: D026 | Chemical |  | LAB
D026 | mg/L | DDT | EQ01 | EQUIS (mg/L) | EQUIS (mg/L) | 0.0001 | N | Y | Y | 17 |  |  | 10/22/1996 11:26:54 PM | 1996-04-12 | 6/3/1998 6:51:35 PM | EMS |  | Concentration | mg/L | Y | (fl. conc.) |  | DDT (fl. conc.) |  | Pesticides; | Pesticides; EMS code: D026 | Chemical |  | LAB
D026 | ug/g | DDT | E8081 | EPA 8081 | See methods EPA 3540, 3630, 3660, 8081 | 0.25 | Y | Y | Y | 9 | 3/18/2020 | 3/23/2020 | 2/22/2023 12:51:03 PM | 2023-02-01 |  | EMS |  | MassPerMass | mg/kg | Y | (mass conc.) |  | DDT (mass conc.) |  | Pesticides; | Pesticides; EMS code: D026 | Chemical |  | LAB
D026 | ug/g wet | DDT | X220 | Hex/AcExt,Part,Chrom,GC | Hex/AcExt;Part;Chrom;GC | 0.0002 | Y | Y | Y | 73 |  |  | 10/24/1996 10:01:25 AM | 1996-04-12 | 6/3/1998 6:51:42 PM | EMS |  | MassPerMass | µg/g | Y | (mass conc.) | (wet) | DDT (wet) (mass conc.) |  | Pesticides; | Pesticides; EMS code: D026 | Chemical |  | LAB
D026 | mg/L | DDT | X268 | CH2Cl2 ext:GC/MS | CH2Cl2 ext:GC/MS |  | Y | Y | Y | 16 | 12/29/2021 | 9/2/2022 | 11/10/2022 9:54:06 AM | 2022-04-08 |  | EMS |  | Concentration | mg/L | Y | (fl. conc.) |  | DDT (fl. conc.) |  | Pesticides; | Pesticides; EMS code: D026 | Chemical |  | LAB
D026 | mg/L | DDT | E660 | EPA 8270E (mod) | EPA 8270E (mod) Miscellaneous Pesticides by GC-MS | 4e-07 | Y | Y | Y | 2 | 9/14/2022 | 9/14/2022 | 12/8/2022 9:48:13 AM | 2022-12-06 |  | EMS |  | Concentration | mg/L | Y | (fl. conc.) |  | DDT (fl. conc.) |  | Pesticides; | Pesticides; EMS code: D026 | Chemical |  | LAB
D027 | mg/L | Dicofol | ZN01 | DCM+Ac.Extr:GC/MS | DCM+Ac.Extr:GC/MS | 0.0002 | Y | Y | Y | 28 | 7/28/2005 7:03:00 AM | 3/29/2012 10:15:00 PM | 5/3/2012 9:27:50 AM | 1996-04-12 | 11/16/1999 4:06:27 PM | EMS | NPEPPIN | Concentration | mg/L | Y | (fl. conc.) |  | Dicofol (fl. conc.) |  | Pesticides; | Pesticides; EMS code: D027 | Chemical |  | LAB
D028 | mg/L | p;p'-DDE | ZN01 | DCM+Ac.Extr:GC/MS | DCM+Ac.Extr:GC/MS | 0.0001 | Y | Y | Y | 52 | 10/3/2002 | 3/29/2012 10:15:00 PM | 5/3/2012 9:27:50 AM | 1996-04-12 | 6/3/1998 6:51:42 PM | EMS |  | Concentration | mg/L | Y | (fl. conc.) |  | p;p'-DDE (fl. conc.) |  | Pesticides; | Pesticides; EMS code: D028 | Chemical |  | LAB
D028 | mg/L | p;p'-DDE | F027 | Hexane Extr,GC/ECD | Effluent samples are solvent extracted with Hexane and the extract dried over Sodium Sulfate, concentrated and cleaned up, as required, using Florisil, activated Copper metal, concentrated Sulfuric Acid and Mercury, to remove interferences. The extracts are then analysed using megabore column Gas Chromatography with Electron Capture detection. | 5e-05 | Y | Y | Y | 20 | 3/3/2000 7:36:00 AM | 1/22/2002 10:58:00 AM | 6/18/2003 2:53:39 PM | 1997-01-14 | 6/3/1998 6:51:46 PM | MPAK |  | Concentration | mg/L | Y | (fl. conc.) |  | p;p'-DDE (fl. conc.) |  | Pesticides; | Pesticides; EMS code: D028 | Chemical |  | LAB
D028 | ug/g | p;p'-DDE | ZN03 | DCM+Ac.Extr:GC/MS | DCM+Ac.Extr:GC/MS EPA 8270E (mod) | 0.01 | Y | Y | Y | 26 | 4/20/2002 10:10:00 AM | 10/13/2002 10:30:00 AM | 2/26/2004 1:05:27 AM | 1996-04-12 | 11/16/1999 2:27:40 PM | EMS | NPEPPIN | MassPerMass | mg/kg | Y | (mass conc.) |  | p;p'-DDE (mass conc.) |  | Pesticides; | Pesticides; EMS code: D028 | Chemical |  | LAB
D029 | ug/g | o;p'-DDE | ZN03 | DCM+Ac.Extr:GC/MS | DCM+Ac.Extr:GC/MS EPA 8270E (mod) | 0.02 | Y | Y | Y | 1 | 7/28/2005 7:03:00 AM | 7/28/2005 7:03:00 AM | 6/2/2009 12:01:03 PM | 1999-11-16 |  | NPEPPIN |  | MassPerMass | mg/kg | Y | (mass conc.) |  | o;p'-DDE (mass conc.) | 3424-82-6 | Pesticides; | Pesticides; EMS code: D029 | Chemical |  | LAB
D031 | ug/g | Dinoseb(DNBP) | FS15 | DCM/MEOH acid extr'n,DCM,methn,Florisil,HRGC-EC | Sediment samples are shaken with a mixture of Dichloromethane, Methanol and Sulphuric Acid, filtered, then extracted with Dichloromethane. After concentration, the extract is methylated using ethereal Diazomethane. Extracts are cleaned up as required using Florisil. The extracts are analysed for Acid Extractable Herbicides using High Resolution Gas Chromatography with Electron Capture detection. Positive results may be verified by use of a secondary confirmation GC column. | 0.05 | Y | Y | Y | 35 | 4/24/1997 4:15:00 PM | 6/24/2003 | 2/26/2004 1:05:27 AM | 1997-10-20 | 6/3/1998 6:51:47 PM | NPEPPIN |  | MassPerMass | mg/kg | Y | (mass conc.) |  | Dinoseb(DNBP) (mass conc.) | 88-85-7 | Pesticides; | Pesticides; EMS code: D031 | Chemical |  | LAB
D031 | mg/L | Dinoseb(DNBP) | AEWA | H2SO4 acid,CHCl2 extr.florisil. GC/MS | Effluent and water samples are acidified with Sulphuric Acid and extracted with Dichloromethane. After concentration, the extract is methylated using ethereal Diazomethane. Extracts are cleaned up as required using Florisil. The extracts are analyzed for Acid Extractable Herbicides using GC/MS. | 0.0002 | Y | Y | Y | 6 | 6/19/2003 10:35:00 AM | 7/29/2005 7:10:00 AM | 6/2/2009 12:01:03 PM | 1996-04-12 | 6/3/1998 6:51:42 PM | EMS |  | Concentration | mg/L | Y | (fl. conc.) |  | Dinoseb(DNBP) (fl. conc.) | 88-85-7 | Pesticides; | Pesticides; EMS code: D031 | Chemical |  | LAB
D031 | mg/L | Dinoseb(DNBP) | F052 | Acid,DCM extr'n,methn,Florisil,HRGC-EC | Effluent and water samples are acidified with Sulphuric Acid and extracted with Dichloromethane. After concentration, the extract is methylated using ethereal Diazomethane. Extracts are cleaned up as required using Florisil. The extracts are analysed for Acid Extractable Herbicides using High Resolution Gas Chromatography with Electron Capture detection. Positive results may be verified by use of a secondary confirmation GC column. | 5e-05 | Y | Y | Y | 33 | 12/8/1998 3:39:00 PM | 12/31/2001 12:00:00 PM | 6/18/2003 2:53:40 PM | 1997-10-20 | 6/3/1998 6:51:47 PM | NPEPPIN |  | Concentration | mg/L | Y | (fl. conc.) |  | Dinoseb(DNBP) (fl. conc.) | 88-85-7 | Pesticides; | Pesticides; EMS code: D031 | Chemical |  | LAB
D031 | mg/L | Dinoseb(DNBP) | P015 | Phenoxyacid Herbicides by LC-MS/MS | Phenoxyacid Herbicides by LC-MS/MS | 0.5 | Y | Y | Y | 71 | 10/9/2019 | 9/9/2022 | 12/8/2022 12:57:05 PM | 2019-10-17 |  | KATFRASE |  | Concentration | mg/L | Y | (fl. conc.) |  | Dinoseb(DNBP) (fl. conc.) | 88-85-7 | Pesticides; | Pesticides; EMS code: D031 | Chemical |  | LAB
D031 | ug/g | Dinoseb(DNBP) | ZN03 | DCM+Ac.Extr:GC/MS | DCM+Ac.Extr:GC/MS EPA 8270E (mod) | 0.0005 | Y | Y | Y | 15 | 1/14/2022 | 2/2/2022 | 1/11/2024 1:07:07 PM | 2022-07-27 |  | EMS |  | MassPerMass | mg/kg | Y | (mass conc.) |  | Dinoseb(DNBP) (mass conc.) | 88-85-7 | Pesticides; | Pesticides; EMS code: D031 | Chemical |  | LAB
D032 | ug/g | Dicamba | FS15 | DCM/MEOH acid extr'n,DCM,methn,Florisil,HRGC-EC | Sediment samples are shaken with a mixture of Dichloromethane, Methanol and Sulphuric Acid, filtered, then extracted with Dichloromethane. After concentration, the extract is methylated using ethereal Diazomethane. Extracts are cleaned up as required using Florisil. The extracts are analysed for Acid Extractable Herbicides using High Resolution Gas Chromatography with Electron Capture detection. Positive results may be verified by use of a secondary confirmation GC column. | 0.05 | Y | Y | Y | 35 | 4/24/1997 4:14:00 PM | 6/24/2003 | 2/26/2004 1:05:27 AM | 1997-10-20 | 6/3/1998 6:51:47 PM | NPEPPIN |  | MassPerMass | mg/kg | Y | (mass conc.) |  | Dicamba (mass conc.) | 1918-00-9 | Pesticides; | Pesticides; EMS code: D032 | Chemical |  | LAB
D032 | mg/L | Dicamba | P015 | Phenoxyacid Herbicides by LC-MS/MS | Phenoxyacid Herbicides by LC-MS/MS | 0.5 | Y | Y | Y | 71 | 10/9/2019 | 9/9/2022 | 12/8/2022 12:57:05 PM | 2019-10-17 |  | KATFRASE |  | Concentration | mg/L | Y | (fl. conc.) |  | Dicamba (fl. conc.) | 1918-00-9 | Pesticides; | Pesticides; EMS code: D032 | Chemical |  | LAB
D032 | mg/L | Dicamba | AEWA | H2SO4 acid,CHCl2 extr.florisil. GC/MS | Effluent and water samples are acidified with Sulphuric Acid and extracted with Dichloromethane. After concentration, the extract is methylated using ethereal Diazomethane. Extracts are cleaned up as required using Florisil. The extracts are analyzed for Acid Extractable Herbicides using GC/MS. | 0.0001 | Y | Y | Y | 6 | 6/16/2003 | 7/29/2005 7:10:00 AM | 6/2/2009 12:01:03 PM | 1996-04-12 | 6/3/1998 6:51:42 PM | EMS |  | Concentration | mg/L | Y | (fl. conc.) |  | Dicamba (fl. conc.) | 1918-00-9 | Pesticides; | Pesticides; EMS code: D032 | Chemical |  | LAB
D032 | ug/g | Dicamba | ZN03 | DCM+Ac.Extr:GC/MS | DCM+Ac.Extr:GC/MS EPA 8270E (mod) | 0.001 | Y | Y | Y | 15 | 1/14/2022 | 2/2/2022 | 1/11/2024 1:07:07 PM | 2022-07-27 |  | EMS |  | MassPerMass | mg/kg | Y | (mass conc.) |  | Dicamba (mass conc.) | 1918-00-9 | Pesticides; | Pesticides; EMS code: D032 | Chemical |  | LAB
D032 | mg/L | Dicamba | F052 | Acid,DCM extr'n,methn,Florisil,HRGC-EC | Effluent and water samples are acidified with Sulphuric Acid and extracted with Dichloromethane. After concentration, the extract is methylated using ethereal Diazomethane. Extracts are cleaned up as required using Florisil. The extracts are analysed for Acid Extractable Herbicides using High Resolution Gas Chromatography with Electron Capture detection. Positive results may be verified by use of a secondary confirmation GC column. | 5e-05 | Y | Y | Y | 33 | 12/8/1998 3:39:00 PM | 12/31/2001 12:00:00 PM | 6/18/2003 2:53:40 PM | 1997-10-20 | 6/3/1998 6:51:47 PM | NPEPPIN |  | Concentration | mg/L | Y | (fl. conc.) |  | Dicamba (fl. conc.) | 1918-00-9 | Pesticides; | Pesticides; EMS code: D032 | Chemical |  | LAB
D034 | mg/L | Diquat | F123 | ph10,Solid Phase Extr. HPLC-UV | Primary Reference - USEPA Method 549.1 "Determination of Diquat and Paraquat in Drinking Water by Liquid-Solid Extraction and HPLC-UV Detection". Approximately 200mL of water is adjusted to pH 10.5. Target organics are extracted using a prepared C18 solid sorbant cartridge by reverse phase/Ion Pair SPE. Compounds are then extracted from the cartridge by eluting with an acidic aqueous solvent. The eluate is injected into a HPLC for separation and quantification is achieved by measuring absorbance at 308nm for Diquat and 258nm for Paraquat. | 0.01 | Y | Y | Y | 3 | 10/2/2002 1:55:00 PM | 10/2/2002 1:55:00 PM | 3/24/2004 1:04:12 AM | 2002-11-05 |  | UMULLER |  | Concentration | mg/L | Y | (fl. conc.) |  | Diquat (fl. conc.) |  | Pesticides; | Pesticides; EMS code: D034 | Chemical |  | LAB
D035 | mg/L | Dicrotophos | ZN01 | DCM+Ac.Extr:GC/MS | DCM+Ac.Extr:GC/MS | 0.0005 | Y | Y | Y | 28 | 7/29/2005 7:10:00 AM | 3/29/2012 10:15:00 PM | 5/3/2012 9:27:50 AM | 1996-04-12 | 11/16/1999 4:06:27 PM | EMS | NPEPPIN | Concentration | mg/L | Y | (fl. conc.) |  | Dicrotophos (fl. conc.) | 141-66-2 | Pesticides; | Pesticides; EMS code: D035 | Chemical |  | LAB
D036 | mg/L | Dioxathion | ZN01 | DCM+Ac.Extr:GC/MS | DCM+Ac.Extr:GC/MS | 0.0005 | Y | Y | Y | 27 | 11/1/2008 | 3/29/2012 10:15:00 PM | 5/3/2012 9:27:50 AM | 1996-04-12 | 11/16/1999 4:06:27 PM | EMS | NPEPPIN | Concentration | mg/L | Y | (fl. conc.) |  | Dioxathion (fl. conc.) |  | Pesticides; | Pesticides; EMS code: D036 | Chemical |  | LAB
D037 | ug/g | Disulfoton | ZN03 | DCM+Ac.Extr:GC/MS | DCM+Ac.Extr:GC/MS EPA 8270E (mod) | 0.05 | Y | Y | Y | 2 | 11/2/2022 12:53:00 AM | 11/2/2022 4:41:00 AM | 1/7/2023 3:12:10 PM | 1996-04-12 | 11/16/1999 2:27:40 PM | EMS | NPEPPIN | MassPerMass | mg/kg | Y | (mass conc.) |  | Disulfoton (mass conc.) | 298-04-4 | Pesticides; | Pesticides; EMS code: D037 | Chemical |  | LAB
D037 | mg/L | Disulfoton | ZN01 | DCM+Ac.Extr:GC/MS | DCM+Ac.Extr:GC/MS | 0.0005 | Y | Y | Y | 28 | 7/29/2005 7:10:00 AM | 3/29/2012 10:15:00 PM | 5/3/2012 9:27:50 AM | 1996-04-12 | 11/16/1999 4:06:27 PM | EMS | NPEPPIN | Concentration | mg/L | Y | (fl. conc.) |  | Disulfoton (fl. conc.) | 298-04-4 | Pesticides; | Pesticides; EMS code: D037 | Chemical |  | LAB
D039 | ug/g | Demeton | ZN03 | DCM+Ac.Extr:GC/MS | DCM+Ac.Extr:GC/MS EPA 8270E (mod) | 0.05 | Y | Y | Y | 39 | 4/20/2002 10:25:00 AM | 10/16/2002 3:20:00 PM | 2/26/2004 1:05:27 AM | 1996-04-12 | 11/16/1999 2:27:40 PM | EMS | NPEPPIN | MassPerMass | mg/kg | Y | (mass conc.) |  | Demeton (mass conc.) | 8065-48-3 | Pesticides; | Pesticides; EMS code: D039 | Chemical |  | LAB
D039 | mg/L | Demeton | ZN01 | DCM+Ac.Extr:GC/MS | DCM+Ac.Extr:GC/MS | 0.0005 | Y | Y | Y | 52 | 10/3/2002 | 3/29/2012 10:15:00 PM | 5/3/2012 9:27:50 AM | 1996-04-12 | 11/16/1999 4:06:27 PM | EMS | NPEPPIN | Concentration | mg/L | Y | (fl. conc.) |  | Demeton (fl. conc.) | 8065-48-3 | Pesticides; | Pesticides; EMS code: D039 | Chemical |  | LAB
D041 | mg/L | 2;4-D | EQ01 | EQUIS (mg/L) | EQUIS (mg/L) | 0 | N | Y | Y | 2 |  |  | 10/22/1996 5:23:44 PM | 1996-04-12 | 6/3/1998 6:51:35 PM | EMS |  | Concentration | mg/L | Y | (fl. conc.) |  | 2;4-D (fl. conc.) | 94-75-7 | Pesticides; | Pesticides; EMS code: D041 | Chemical |  | LAB
D041 | mg/L | 2;4-D | P015 | Phenoxyacid Herbicides by LC-MS/MS | Phenoxyacid Herbicides by LC-MS/MS | 0.5 | Y | Y | Y | 71 | 10/9/2019 | 9/9/2022 | 12/8/2022 12:57:05 PM | 2019-10-17 |  | KATFRASE |  | Concentration | mg/L | Y | (fl. conc.) |  | 2;4-D (fl. conc.) | 94-75-7 | Pesticides; | Pesticides; EMS code: D041 | Chemical |  | LAB
D041 | mg/L | 2;4-D | F052 | Acid,DCM extr'n,methn,Florisil,HRGC-EC | Effluent and water samples are acidified with Sulphuric Acid and extracted with Dichloromethane. After concentration, the extract is methylated using ethereal Diazomethane. Extracts are cleaned up as required using Florisil. The extracts are analysed for Acid Extractable Herbicides using High Resolution Gas Chromatography with Electron Capture detection. Positive results may be verified by use of a secondary confirmation GC column. | 5e-05 | Y | Y | Y | 33 | 12/8/1998 3:39:00 PM | 12/31/2001 12:00:00 PM | 6/18/2003 2:53:41 PM | 1997-10-20 | 6/3/1998 6:51:47 PM | NPEPPIN |  | Concentration | mg/L | Y | (fl. conc.) |  | 2;4-D (fl. conc.) | 94-75-7 | Pesticides; | Pesticides; EMS code: D041 | Chemical |  | LAB
D041 | ug/g | 2;4-D | ZN03 | DCM+Ac.Extr:GC/MS | DCM+Ac.Extr:GC/MS EPA 8270E (mod) | 0.0005 | Y | Y | Y | 15 | 1/14/2022 | 2/2/2022 | 1/11/2024 1:07:07 PM | 2022-07-27 |  | EMS |  | MassPerMass | mg/kg | Y | (mass conc.) |  | 2;4-D (mass conc.) | 94-75-7 | Pesticides; | Pesticides; EMS code: D041 | Chemical |  | LAB
D041 | mg/L | 2;4-D | AEWA | H2SO4 acid,CHCl2 extr.florisil. GC/MS | Effluent and water samples are acidified with Sulphuric Acid and extracted with Dichloromethane. After concentration, the extract is methylated using ethereal Diazomethane. Extracts are cleaned up as required using Florisil. The extracts are analyzed for Acid Extractable Herbicides using GC/MS. | 0.0002 | Y | Y | Y | 6 | 6/16/2003 | 7/29/2005 7:10:00 AM | 6/2/2009 12:01:03 PM | 1996-04-12 | 6/3/1998 6:51:42 PM | EMS |  | Concentration | mg/L | Y | (fl. conc.) |  | 2;4-D (fl. conc.) | 94-75-7 | Pesticides; | Pesticides; EMS code: D041 | Chemical |  | LAB
D041 | ug/g | 2;4-D | FS15 | DCM/MEOH acid extr'n,DCM,methn,Florisil,HRGC-EC | Sediment samples are shaken with a mixture of Dichloromethane, Methanol and Sulphuric Acid, filtered, then extracted with Dichloromethane. After concentration, the extract is methylated using ethereal Diazomethane. Extracts are cleaned up as required using Florisil. The extracts are analysed for Acid Extractable Herbicides using High Resolution Gas Chromatography with Electron Capture detection. Positive results may be verified by use of a secondary confirmation GC column. | 0.05 | Y | Y | Y | 35 | 4/24/1997 4:15:00 PM | 6/24/2003 | 2/26/2004 1:05:27 AM | 1997-10-20 | 6/3/1998 6:51:47 PM | NPEPPIN |  | MassPerMass | mg/kg | Y | (mass conc.) |  | 2;4-D (mass conc.) | 94-75-7 | Pesticides; | Pesticides; EMS code: D041 | Chemical |  | LAB
D043 | mg/L | Demeton S | F053 | DCM extr'n,conc'n,HGRC-PFPD/N,P det'n | Effluent samples are extracted with Dichloromethane. The extract is then concentrated and analysed using High Resolution Gas Chromatography with Pulsed Flame Photometric Dectection. Confirmation, if required, is by High Resolution Gas Chromatography with Nitrogen Phosphorus Detection. | 0.0001 | Y | Y | Y | 31 | 3/3/2000 7:36:00 AM | 1/22/2002 10:26:00 AM | 6/18/2003 2:53:41 PM | 1997-10-20 | 6/3/1998 6:51:47 PM | NPEPPIN |  | Concentration | mg/L | Y | (fl. conc.) |  | Demeton S (fl. conc.) |  | Pesticides; | Pesticides; EMS code: D043 | Chemical |  | LAB
D044 | mg/L | 2;4-D B | AEWA | H2SO4 acid,CHCl2 extr.florisil. GC/MS | Effluent and water samples are acidified with Sulphuric Acid and extracted with Dichloromethane. After concentration, the extract is methylated using ethereal Diazomethane. Extracts are cleaned up as required using Florisil. The extracts are analyzed for Acid Extractable Herbicides using GC/MS. | 0.0002 | Y | Y | Y | 1 | 7/29/2005 7:10:00 AM | 7/29/2005 7:10:00 AM | 6/2/2009 12:01:03 PM | 1998-11-17 |  | NPEPPIN |  | Concentration | mg/L | Y | (fl. conc.) |  | 2;4-D B (fl. conc.) |  | Pesticides; | Pesticides; EMS code: D044 | Chemical |  | LAB
D045 | mg/L | Dichlorprop | AEWA | H2SO4 acid,CHCl2 extr.florisil. GC/MS | Effluent and water samples are acidified with Sulphuric Acid and extracted with Dichloromethane. After concentration, the extract is methylated using ethereal Diazomethane. Extracts are cleaned up as required using Florisil. The extracts are analyzed for Acid Extractable Herbicides using GC/MS. | 0.0002 | Y | Y | Y | 6 | 6/16/2003 | 7/29/2005 7:10:00 AM | 6/2/2009 12:01:03 PM | 1996-04-12 | 6/3/1998 6:51:42 PM | EMS |  | Concentration | mg/L | Y | (fl. conc.) |  | Dichlorprop (fl. conc.) |  | Pesticides; | Pesticides; EMS code: D045 | Chemical |  | LAB
D045 | mg/L | Dichlorprop | F052 | Acid,DCM extr'n,methn,Florisil,HRGC-EC | Effluent and water samples are acidified with Sulphuric Acid and extracted with Dichloromethane. After concentration, the extract is methylated using ethereal Diazomethane. Extracts are cleaned up as required using Florisil. The extracts are analysed for Acid Extractable Herbicides using High Resolution Gas Chromatography with Electron Capture detection. Positive results may be verified by use of a secondary confirmation GC column. | 5e-05 | Y | Y | Y | 33 | 12/8/1998 3:39:00 PM | 12/31/2001 12:00:00 PM | 6/18/2003 2:53:41 PM | 1997-10-20 | 6/3/1998 6:51:47 PM | NPEPPIN |  | Concentration | mg/L | Y | (fl. conc.) |  | Dichlorprop (fl. conc.) |  | Pesticides; | Pesticides; EMS code: D045 | Chemical |  | LAB
D045 | ug/g | Dichlorprop | FS15 | DCM/MEOH acid extr'n,DCM,methn,Florisil,HRGC-EC | Sediment samples are shaken with a mixture of Dichloromethane, Methanol and Sulphuric Acid, filtered, then extracted with Dichloromethane. After concentration, the extract is methylated using ethereal Diazomethane. Extracts are cleaned up as required using Florisil. The extracts are analysed for Acid Extractable Herbicides using High Resolution Gas Chromatography with Electron Capture detection. Positive results may be verified by use of a secondary confirmation GC column. | 0.05 | Y | Y | Y | 35 | 4/24/1997 4:14:00 PM | 6/24/2003 | 2/26/2004 1:05:28 AM | 1997-10-20 | 6/3/1998 6:51:47 PM | NPEPPIN |  | MassPerMass | mg/kg | Y | (mass conc.) |  | Dichlorprop (mass conc.) |  | Pesticides; | Pesticides; EMS code: D045 | Chemical |  | LAB
D047 | mg/L | 2-Methyl-4;6-Dinitrophenol | E-8041 | EPA 8041A | EPA 8041A Phenols by Gas Chromatography |  | Y | Y | Y | 37 | 12/29/2021 | 6/10/2024 | 6/26/2024 4:28:10 PM | 2022-04-20 |  | EMS |  | Concentration | mg/L | Y | (fl. conc.) |  | 2-Methyl-4;6-Dinitrophenol (fl. conc.) |  | Organics; | Organics; EMS code: D047 | Chemical |  | LAB
D047 | mg/L | 2-Methyl-4;6-Dinitrophenol | F028 | Acid,Extr.MeCl,HRGC/LRMS:TIM,SIM | A measured volume of sample is acidified and extracted with Methylene Chloride. For samples which have a high organic content that potentially produce chromatographic interferences an initial pre-extraction under basic conditions is performed. The sample extract is then evaporated to small volume and analysed by High Resolution Gas Chromatography / Low Resolution Mass Spectrometry (HRGC/LRMS). LRMS data is acquired in Total Ion Mode (TIM) and quantitation is performed using a single mass ion. Positive Identification of the various target compounds is made by comparison of absolute retention times to those of standards combined with computer based mass spectral mass/intensity library searching and matching routines. | 0.02 | Y | Y | Y | 99 | 10/28/1997 4:33:00 PM | 4/14/2015 12:26:00 AM | 4/16/2015 6:00:10 PM | 1997-01-14 | 6/3/1998 6:51:46 PM | MPAK |  | Concentration | mg/L | Y | (fl. conc.) |  | 2-Methyl-4;6-Dinitrophenol (fl. conc.) |  | Organics; | Organics; EMS code: D047 | Chemical |  | LAB
D047 | ug/g | 2-Methyl-4;6-Dinitrophenol | E8041S | EPA 8041A for soil | See EPA 8041A | 0.05 | Y | Y | Y | 15 | 1/14/2022 | 2/5/2022 | 1/11/2024 1:07:07 PM | 2022-06-23 |  | EMS |  | MassPerMass | mg/kg | Y | (mass conc.) |  | 2-Methyl-4;6-Dinitrophenol (mass conc.) |  | Organics; | Organics; EMS code: D047 | Chemical |  | LAB
D048 | mg/L | Dimethylphenol-2;4 | F028 | Acid,Extr.MeCl,HRGC/LRMS:TIM,SIM | A measured volume of sample is acidified and extracted with Methylene Chloride. For samples which have a high organic content that potentially produce chromatographic interferences an initial pre-extraction under basic conditions is performed. The sample extract is then evaporated to small volume and analysed by High Resolution Gas Chromatography / Low Resolution Mass Spectrometry (HRGC/LRMS). LRMS data is acquired in Total Ion Mode (TIM) and quantitation is performed using a single mass ion. Positive Identification of the various target compounds is made by comparison of absolute retention times to those of standards combined with computer based mass spectral mass/intensity library searching and matching routines. | 0.005 | Y | Y | Y | 100 | 4/10/1997 3:18:00 PM | 4/14/2015 12:26:00 AM | 4/16/2015 6:00:10 PM | 1997-01-14 | 6/3/1998 6:51:46 PM | MPAK |  | Concentration | mg/L | Y | (fl. conc.) |  | Dimethylphenol-2;4 (fl. conc.) |  | Organics; | Organics; EMS code: D048 | Chemical |  | LAB
D048 | ug/g | Dimethylphenol-2;4 | 1332 | Phenolics by Tumbler/GC-MS | Phenolics by Tumbler/GC-MS | 0.02 | Y | Y | Y | 15 | 1/17/2022 | 2/2/2022 | 1/11/2024 1:07:07 PM | 2022-07-07 |  | EMS |  | MassPerMass | mg/kg | Y | (mass conc.) |  | Dimethylphenol-2;4 (mass conc.) |  | Organics; | Organics; EMS code: D048 | Chemical |  | LAB
D048 | mg/L | Dimethylphenol-2;4 | X268 | CH2Cl2 ext:GC/MS | CH2Cl2 ext:GC/MS | 0.0005 | Y | Y | Y | 859 | 3/6/2007 | 8/14/2024 | 8/16/2024 4:05:21 PM | 2012-03-27 |  | RDGIBSON |  | Concentration | mg/L | Y | (fl. conc.) |  | Dimethylphenol-2;4 (fl. conc.) |  | Organics; | Organics; EMS code: D048 | Chemical |  | LAB
D049 | mg/L | Dinitrophenol-2;4 | E-8041 | EPA 8041A | EPA 8041A Phenols by Gas Chromatography |  | Y | Y | Y | 37 | 12/29/2021 | 6/10/2024 | 6/26/2024 4:28:10 PM | 2022-04-08 |  | EMS |  | Concentration | mg/L | Y | (fl. conc.) |  | Dinitrophenol-2;4 (fl. conc.) |  | Organics; | Organics; EMS code: D049 | Chemical |  | LAB
D049 | ug/g | Dinitrophenol-2;4 | E8041S | EPA 8041A for soil | See EPA 8041A | 0.05 | Y | Y | Y | 15 | 1/14/2022 | 2/5/2022 | 1/11/2024 1:07:07 PM | 2022-06-23 |  | EMS |  | MassPerMass | mg/kg | Y | (mass conc.) |  | Dinitrophenol-2;4 (mass conc.) |  | Organics; | Organics; EMS code: D049 | Chemical |  | LAB
D049 | mg/L | Dinitrophenol-2;4 | F028 | Acid,Extr.MeCl,HRGC/LRMS:TIM,SIM | A measured volume of sample is acidified and extracted with Methylene Chloride. For samples which have a high organic content that potentially produce chromatographic interferences an initial pre-extraction under basic conditions is performed. The sample extract is then evaporated to small volume and analysed by High Resolution Gas Chromatography / Low Resolution Mass Spectrometry (HRGC/LRMS). LRMS data is acquired in Total Ion Mode (TIM) and quantitation is performed using a single mass ion. Positive Identification of the various target compounds is made by comparison of absolute retention times to those of standards combined with computer based mass spectral mass/intensity library searching and matching routines. | 0.2 | Y | Y | Y | 100 | 4/10/1997 3:19:00 PM | 4/14/2015 12:26:00 AM | 4/16/2015 6:00:10 PM | 1997-01-14 | 6/3/1998 6:51:46 PM | MPAK |  | Concentration | mg/L | Y | (fl. conc.) |  | Dinitrophenol-2;4 (fl. conc.) |  | Organics; | Organics; EMS code: D049 | Chemical |  | LAB
D050 | mg/L | 2;4-Dichlorophenol | F028 | Acid,Extr.MeCl,HRGC/LRMS:TIM,SIM | A measured volume of sample is acidified and extracted with Methylene Chloride. For samples which have a high organic content that potentially produce chromatographic interferences an initial pre-extraction under basic conditions is performed. The sample extract is then evaporated to small volume and analysed by High Resolution Gas Chromatography / Low Resolution Mass Spectrometry (HRGC/LRMS). LRMS data is acquired in Total Ion Mode (TIM) and quantitation is performed using a single mass ion. Positive Identification of the various target compounds is made by comparison of absolute retention times to those of standards combined with computer based mass spectral mass/intensity library searching and matching routines. | 0.005 | Y | Y | Y | 76 | 4/10/1997 3:18:00 PM | 3/11/2003 | 4/1/2004 1:13:23 AM | 1997-01-14 | 6/3/1998 6:51:46 PM | MPAK |  | Concentration | mg/L | Y | (fl. conc.) |  | 2;4-Dichlorophenol (fl. conc.) |  | Organics; | Organics; EMS code: D050 | Chemical |  | LAB
D051 | mg/L | Dihydroisopimaric | F009 | pH 9,Diethyl Ether:HRGC/LRMS | Effluent samples are first adjusted to pH 9 and then extracted with Diethyl Ether. The extracts are rotary evaporated and the resulting residues, containing the Resin Acids and other solvent soluble organics are then redissolved in Methanol, transferred to small glass vials, methylated with Diazomethane then evaporated to dryness under Nitrogen. [Methylation Reaction R-COOH(Resin/Fatty Acids) + CH2N2(Diazomethane) = R-COOCH3(Methyl ester) + N2) The residue is re-dissolved in Methyl t-Butyl Ether (MTBE) and injected on a High Resolution Gas Chromatograph/Low Resolution Mass Spectrometer (HRGC/LRMS) system for analysis. LRMS data is acquired in Total Ion Mode (TIM) and quantitation is performed using selected mass ions. Positive identification of the various Resin Acids is made on the basis of a comparison of absolute retention times to those of the external standards and computer based mass spectral mass/intensity library matching routines. | 0.005 | Y | Y | Y | 131 | 12/27/1996 4:16:00 PM | 1/8/2002 12:49:00 PM | 5/28/2003 11:38:56 AM | 1996-09-16 | 6/3/1998 6:51:44 PM | MPAK |  | Concentration | mg/L | Y | (fl. conc.) |  | Dihydroisopimaric (fl. conc.) |  | Organics; | Organics; EMS code: D051 | Chemical |  | LAB
D051 | ug/g | Dihydroisopimaric | FS05 | MeOH,DCM extr:HRGC/LRMS | Sediment samples are extracted by shaking with Methanol in a screw cap centrifuge bottle. The extracts are rotary evaporated and the resulting residues, containing the Resin Acids and other solvent soluble organics, are then redissolved in Methanol, transferred to small glass vials, methylated with Diazomethane then evaporated to dryness under Nitrogen. [Methylation Reaction R-COOH(Resin/Fatty Acids) + CH2N2 (Diazomethane) = R-COOCH3 (Methyl ester) + N2] The residue is re-dissolved in Methyl t-Butyl Ether (MTBE) and injected on a High Resolution Gas Chromatograph / Low Resolution Mass Spectrometer (HRGC/LRMS) system for analysis. LRMS data is acquired in Total Ion Mode (TIM) and quantitation is performed using selected mass ions. Positive identification of the various Resin Acids is made on the basis of a comparison of absolute retention times to those of the external standards and computer based mass spectral mass/intensity library matching routines. | 0.2 | Y | Y | Y | 1 | 9/14/2001 3:29:00 PM | 9/14/2001 3:29:00 PM | 2/17/2003 3:39:56 PM | 1996-09-16 | 6/3/1998 6:51:44 PM | MPAK |  | MassPerMass | mg/kg | Y | (mass conc.) |  | Dihydroisopimaric (mass conc.) |  | Organics; | Organics; EMS code: D051 | Chemical |  | LAB
D052 | mg/L | Dehydroabietic Acid | P030 | Resin Acids | The water sample is acidified and extracted with dichloro-methane. The raw extract is concentrated and derivatized with diazomethane to produce the corresponding methyl ester derivatives. (Suitable substitute derivatives may be produced by other available techniques.) If required, the extracts are cleaned up by Florisil column chromatography. The derivatives are analyzed by gas chromatography with flame ionization detection. | 0.001 | Y | Y | Y | 480 | 8/26/2014 | 8/15/2017 | 8/15/2017 4:37:09 PM | 1996-04-12 | 6/3/1998 6:51:35 PM | EMS |  | Concentration | mg/L | Y | (fl. conc.) |  | Dehydroabietic Acid (fl. conc.) |  | Organics; | Organics; EMS code: D052 | Chemical |  | LAB
D052 | mg/L | Dehydroabietic Acid | P033 | Resin&Fatty Acids,GC/MS | Resin&Fatty Acids;GC/MS | 0.001 | Y | Y | Y | 31 |  |  | 6/7/2024 5:30:18 PM | 1996-04-12 | 6/3/1998 6:51:35 PM | EMS |  | Concentration | mg/L | Y | (fl. conc.) |  | Dehydroabietic Acid (fl. conc.) |  | Organics; | Organics; EMS code: D052 | Chemical |  | LAB
D052 | mg/L | Dehydroabietic Acid | X461 | Liguid Chromatography | Liguid Chromatography | 0.05 | Y | Y | Y | 3496 | 5/4/2000 8:00:00 AM | 10/6/2020 | 7/29/2024 11:17:51 AM | 1996-04-12 | 6/3/1998 6:51:35 PM | EMS |  | Concentration | mg/L | Y | (fl. conc.) |  | Dehydroabietic Acid (fl. conc.) |  | Organics; | Organics; EMS code: D052 | Chemical |  | LAB
D052 | mg/L | Dehydroabietic Acid | F009 | pH 9,Diethyl Ether:HRGC/LRMS | Effluent samples are first adjusted to pH 9 and then extracted with Diethyl Ether. The extracts are rotary evaporated and the resulting residues, containing the Resin Acids and other solvent soluble organics are then redissolved in Methanol, transferred to small glass vials, methylated with Diazomethane then evaporated to dryness under Nitrogen. [Methylation Reaction R-COOH(Resin/Fatty Acids) + CH2N2(Diazomethane) = R-COOCH3(Methyl ester) + N2) The residue is re-dissolved in Methyl t-Butyl Ether (MTBE) and injected on a High Resolution Gas Chromatograph/Low Resolution Mass Spectrometer (HRGC/LRMS) system for analysis. LRMS data is acquired in Total Ion Mode (TIM) and quantitation is performed using selected mass ions. Positive identification of the various Resin Acids is made on the basis of a comparison of absolute retention times to those of the external standards and computer based mass spectral mass/intensity library matching routines. | 0.005 | Y | Y | Y | 198 | 12/27/1996 4:15:00 PM | 5/13/2003 | 4/1/2004 1:08:45 AM | 1996-09-16 | 6/3/1998 6:51:44 PM | MPAK |  | Concentration | mg/L | Y | (fl. conc.) |  | Dehydroabietic Acid (fl. conc.) |  | Organics; | Organics; EMS code: D052 | Chemical |  | LAB
D052 | mg/L | Dehydroabietic Acid | RFAS | Separation of resin and fatty acids (RFAs) in water by GC/MS | Scientific Principle The separation of resin and fatty acids (RFAs) in water samples is performed using gas chromatography/mass spectrometry (GC/MS) in the Selected Ion Monitoring (SIM) mode. The separation is based on compound size; boiling point and polarity. Analyte concentration is calculated using the internal standard method. RFA compounds are extracted from water samples using a liquid-liquid solvent extraction with DCM. Extraction occurs after the sample is brought to a pH of ~ 2 and placed in a rotary tumbler (at approx. 40 rpm) for 1 hour. The resulting extract is transferred to a vial; dried with acidified baked sodium sulfate; methylated with diazomethane and concentrated to 1mL under nitrogen prior to analysis. The resulting extracts are vialed and fortified with labeled internal standard compounds prior to analysis by gas chromatography with a (50%-Phenyl)-methylpolysiloxane (DB-17) column and a mass spectrometer detector in SIM Mode. | 0.006 | Y | Y | Y | 1 | 9/17/2021 7:49:00 PM | 9/17/2021 7:49:00 PM | 2/1/2022 2:38:14 PM | 2022-01-19 |  | EMS |  | Concentration | mg/L | Y | (fl. conc.) |  | Dehydroabietic Acid (fl. conc.) |  | Organics; | Organics; EMS code: D052 | Chemical |  | LAB
D052 | mg/L | Dehydroabietic Acid | P035 | DCM-DIAZO-FLOR | DCM-DIAZO-FLOR | 0.0005 | Y | Y | Y | 22 |  |  | 10/24/1996 12:44:23 PM | 1996-04-12 | 6/3/1998 6:51:35 PM | EMS |  | Concentration | mg/L | Y | (fl. conc.) |  | Dehydroabietic Acid (fl. conc.) |  | Organics; | Organics; EMS code: D052 | Chemical |  | LAB
D052 | mg/L | Dehydroabietic Acid | E677 | Resins & Fatty Acids are methylated then analyzed by GCMS | EPA 8270E; AB Env. 129; NCASI RA/FA-85.02 (mod) Resins & Fatty Acids are methylated then analyzed by GCMS. Total fatty acids is the sum of arachidic acid; behenic acid; lauric acid; lignoceric acid; linoleic acid; linolenic acid; myristic acid; oleic acid; palmitic acid; and stearic acid. Total resin acids is the sum of abietic acid; 12-chlorodehydroabietic acid; 14-chlorodehydroabietic acid; dehydroabietic acid; 12;14-dichlorodehydroabietic acid; isopimaric acid; levopimaric acid; neoabietic acid; palustric acid; pimaric acid and sandaracopimaric acid. Total resin acids (BC WWQG) is the sum of abietic acid; neoabietic acid; pimaric acid; isopimaric acid; and sandaracopimaric acid. | 0.001 | Y | Y | Y | 12 | 11/28/2022 | 6/15/2024 | 7/8/2024 11:10:02 AM | 2021-12-30 |  | EMS |  | Concentration | mg/L | Y | (fl. conc.) |  | Dehydroabietic Acid (fl. conc.) |  | Organics; | Organics; EMS code: D052 | Chemical |  | LAB
D052 | mg/L | Dehydroabietic Acid | X239 | Gas Chromatography | Gas Chromatography | 0.001 | Y | Y | Y | 12 |  |  | 10/23/1996 11:56:36 PM | 1996-04-12 | 6/3/1998 6:51:35 PM | EMS |  | Concentration | mg/L | Y | (fl. conc.) |  | Dehydroabietic Acid (fl. conc.) |  | Organics; | Organics; EMS code: D052 | Chemical |  | LAB
D052 | ug/g | Dehydroabietic Acid | FS05 | MeOH,DCM extr:HRGC/LRMS | Sediment samples are extracted by shaking with Methanol in a screw cap centrifuge bottle. The extracts are rotary evaporated and the resulting residues, containing the Resin Acids and other solvent soluble organics, are then redissolved in Methanol, transferred to small glass vials, methylated with Diazomethane then evaporated to dryness under Nitrogen. [Methylation Reaction R-COOH(Resin/Fatty Acids) + CH2N2 (Diazomethane) = R-COOCH3 (Methyl ester) + N2] The residue is re-dissolved in Methyl t-Butyl Ether (MTBE) and injected on a High Resolution Gas Chromatograph / Low Resolution Mass Spectrometer (HRGC/LRMS) system for analysis. LRMS data is acquired in Total Ion Mode (TIM) and quantitation is performed using selected mass ions. Positive identification of the various Resin Acids is made on the basis of a comparison of absolute retention times to those of the external standards and computer based mass spectral mass/intensity library matching routines. | 0.2 | Y | Y | Y | 5 | 9/14/2001 3:29:00 PM | 12/3/2002 | 3/27/2004 1:09:20 AM | 1996-09-16 | 6/3/1998 6:51:44 PM | MPAK |  | MassPerMass | mg/kg | Y | (mass conc.) |  | Dehydroabietic Acid (mass conc.) |  | Organics; | Organics; EMS code: D052 | Chemical |  | LAB
D053 | mg/L | Dichlorodehydroabietic Acid | E677 | Resins & Fatty Acids are methylated then analyzed by GCMS | EPA 8270E; AB Env. 129; NCASI RA/FA-85.02 (mod) Resins & Fatty Acids are methylated then analyzed by GCMS. Total fatty acids is the sum of arachidic acid; behenic acid; lauric acid; lignoceric acid; linoleic acid; linolenic acid; myristic acid; oleic acid; palmitic acid; and stearic acid. Total resin acids is the sum of abietic acid; 12-chlorodehydroabietic acid; 14-chlorodehydroabietic acid; dehydroabietic acid; 12;14-dichlorodehydroabietic acid; isopimaric acid; levopimaric acid; neoabietic acid; palustric acid; pimaric acid and sandaracopimaric acid. Total resin acids (BC WWQG) is the sum of abietic acid; neoabietic acid; pimaric acid; isopimaric acid; and sandaracopimaric acid. | 0.001 | Y | Y | Y | 12 | 11/28/2022 | 6/15/2024 | 7/8/2024 11:10:02 AM | 2021-12-30 |  | EMS |  | Concentration | mg/L | Y | (fl. conc.) |  | Dichlorodehydroabietic Acid (fl. conc.) |  | Organics; | Organics; EMS code: D053 | Chemical |  | LAB
D053 | ug/g | Dichlorodehydroabietic Acid | FS05 | MeOH,DCM extr:HRGC/LRMS | Sediment samples are extracted by shaking with Methanol in a screw cap centrifuge bottle. The extracts are rotary evaporated and the resulting residues, containing the Resin Acids and other solvent soluble organics, are then redissolved in Methanol, transferred to small glass vials, methylated with Diazomethane then evaporated to dryness under Nitrogen. [Methylation Reaction R-COOH(Resin/Fatty Acids) + CH2N2 (Diazomethane) = R-COOCH3 (Methyl ester) + N2] The residue is re-dissolved in Methyl t-Butyl Ether (MTBE) and injected on a High Resolution Gas Chromatograph / Low Resolution Mass Spectrometer (HRGC/LRMS) system for analysis. LRMS data is acquired in Total Ion Mode (TIM) and quantitation is performed using selected mass ions. Positive identification of the various Resin Acids is made on the basis of a comparison of absolute retention times to those of the external standards and computer based mass spectral mass/intensity library matching routines. | 0.2 | Y | Y | Y | 5 | 9/14/2001 3:29:00 PM | 12/3/2002 | 3/27/2004 1:09:20 AM | 1997-01-13 | 6/3/1998 6:51:46 PM | MPAK |  | MassPerMass | mg/kg | Y | (mass conc.) |  | Dichlorodehydroabietic Acid (mass conc.) |  | Organics; | Organics; EMS code: D053 | Chemical |  | LAB
D053 | mg/L | Dichlorodehydroabietic Acid | F009 | pH 9,Diethyl Ether:HRGC/LRMS | Effluent samples are first adjusted to pH 9 and then extracted with Diethyl Ether. The extracts are rotary evaporated and the resulting residues, containing the Resin Acids and other solvent soluble organics are then redissolved in Methanol, transferred to small glass vials, methylated with Diazomethane then evaporated to dryness under Nitrogen. [Methylation Reaction R-COOH(Resin/Fatty Acids) + CH2N2(Diazomethane) = R-COOCH3(Methyl ester) + N2) The residue is re-dissolved in Methyl t-Butyl Ether (MTBE) and injected on a High Resolution Gas Chromatograph/Low Resolution Mass Spectrometer (HRGC/LRMS) system for analysis. LRMS data is acquired in Total Ion Mode (TIM) and quantitation is performed using selected mass ions. Positive identification of the various Resin Acids is made on the basis of a comparison of absolute retention times to those of the external standards and computer based mass spectral mass/intensity library matching routines. | 0.005 | Y | Y | Y | 198 | 12/27/1996 4:16:00 PM | 5/13/2003 | 4/1/2004 1:08:45 AM | 1997-01-13 | 6/3/1998 6:51:46 PM | MPAK |  | Concentration | mg/L | Y | (fl. conc.) |  | Dichlorodehydroabietic Acid (fl. conc.) |  | Organics; | Organics; EMS code: D053 | Chemical |  | LAB
D053 | mg/L | Dichlorodehydroabietic Acid | P035 | DCM-DIAZO-FLOR | DCM-DIAZO-FLOR | 0.0005 | Y | Y | Y | 22 |  |  | 10/24/1996 12:44:23 PM | 1996-04-12 | 6/3/1998 6:51:45 PM | EMS |  | Concentration | mg/L | Y | (fl. conc.) |  | Dichlorodehydroabietic Acid (fl. conc.) |  | Organics; | Organics; EMS code: D053 | Chemical |  | LAB
D053 | mg/L | Dichlorodehydroabietic Acid | X239 | Gas Chromatography | Gas Chromatography | 0.001 | Y | Y | Y | 12 |  |  | 10/23/1996 11:56:36 PM | 1996-04-12 | 6/3/1998 6:51:35 PM | EMS |  | Concentration | mg/L | Y | (fl. conc.) |  | Dichlorodehydroabietic Acid (fl. conc.) |  | Organics; | Organics; EMS code: D053 | Chemical |  | LAB
D053 | mg/L | Dichlorodehydroabietic Acid | RFAS | Separation of resin and fatty acids (RFAs) in water by GC/MS | Scientific Principle The separation of resin and fatty acids (RFAs) in water samples is performed using gas chromatography/mass spectrometry (GC/MS) in the Selected Ion Monitoring (SIM) mode. The separation is based on compound size; boiling point and polarity. Analyte concentration is calculated using the internal standard method. RFA compounds are extracted from water samples using a liquid-liquid solvent extraction with DCM. Extraction occurs after the sample is brought to a pH of ~ 2 and placed in a rotary tumbler (at approx. 40 rpm) for 1 hour. The resulting extract is transferred to a vial; dried with acidified baked sodium sulfate; methylated with diazomethane and concentrated to 1mL under nitrogen prior to analysis. The resulting extracts are vialed and fortified with labeled internal standard compounds prior to analysis by gas chromatography with a (50%-Phenyl)-methylpolysiloxane (DB-17) column and a mass spectrometer detector in SIM Mode. | 0.006 | Y | Y | Y | 1 | 9/17/2021 7:49:00 PM | 9/17/2021 7:49:00 PM | 2/1/2022 2:38:14 PM | 2022-01-19 |  | EMS |  | Concentration | mg/L | Y | (fl. conc.) |  | Dichlorodehydroabietic Acid (fl. conc.) |  | Organics; | Organics; EMS code: D053 | Chemical |  | LAB
D053 | mg/L | Dichlorodehydroabietic Acid | P030 | Resin Acids | The water sample is acidified and extracted with dichloro-methane. The raw extract is concentrated and derivatized with diazomethane to produce the corresponding methyl ester derivatives. (Suitable substitute derivatives may be produced by other available techniques.) If required, the extracts are cleaned up by Florisil column chromatography. The derivatives are analyzed by gas chromatography with flame ionization detection. | 0.001 | Y | Y | Y | 81 | 8/26/2014 | 8/15/2017 | 6/7/2024 5:30:17 PM | 1996-04-12 | 6/3/1998 6:51:35 PM | EMS |  | Concentration | mg/L | Y | (fl. conc.) |  | Dichlorodehydroabietic Acid (fl. conc.) |  | Organics; | Organics; EMS code: D053 | Chemical |  | LAB
D053 | mg/L | Dichlorodehydroabietic Acid | X266 | CH2Cl2Ext:PEBrT:GC/ECD | CH2Cl2Ext:PEBrT:GC/ECD | 0.01 | Y | Y | Y | 1 |  |  | 6/7/2024 5:30:18 PM | 1996-04-12 | 6/3/1998 6:51:35 PM | EMS |  | Concentration | mg/L | Y | (fl. conc.) |  | Dichlorodehydroabietic Acid (fl. conc.) |  | Organics; | Organics; EMS code: D053 | Chemical |  | LAB
D054 | ug/g | 2;4'-DDD | ZN03 | DCM+Ac.Extr:GC/MS | DCM+Ac.Extr:GC/MS EPA 8270E (mod) | 0.02 | Y | Y | Y | 17 | 1/17/2022 | 11/14/2022 | 1/11/2024 1:07:07 PM | 2022-06-21 |  | EMS |  | MassPerMass | mg/kg | Y | (mass conc.) |  | 2;4'-DDD (mass conc.) | 53-19-0 | Pesticides; | Pesticides; EMS code: D054 | Chemical |  | LAB
D054 | mg/L | 2;4'-DDD | X268 | CH2Cl2 ext:GC/MS | CH2Cl2 ext:GC/MS |  | Y | Y | Y | 16 | 12/29/2021 | 9/2/2022 | 11/10/2022 9:54:06 AM | 2022-04-08 |  | EMS |  | Concentration | mg/L | Y | (fl. conc.) |  | 2;4'-DDD (fl. conc.) | 53-19-0 | Pesticides; | Pesticides; EMS code: D054 | Chemical |  | LAB
D054 | mg/L | 2;4'-DDD | E660 | EPA 8270E (mod) | EPA 8270E (mod) Miscellaneous Pesticides by GC-MS | 4e-07 | Y | Y | Y | 2 | 9/14/2022 | 9/14/2022 | 12/8/2022 9:48:13 AM | 2022-12-06 |  | EMS |  | Concentration | mg/L | Y | (fl. conc.) |  | 2;4'-DDD (fl. conc.) | 53-19-0 | Pesticides; | Pesticides; EMS code: D054 | Chemical |  | LAB
D054 | mg/L | 2;4'-DDD | MW03 | OCP1-LL-SF-ECD-VA | This analysis is carried out using procedures adapted from "Test Methods for Evaluating Solid Waste" SW-846, Methods 3510, 3630, 3660 and 8081, published by the United States Environmental Protection Agency (EPA). The procedure involves extraction of the entire water sample with dichloromethane. The extract then undergoes one or more of the following clean-up procedures (if required): silica gel clean-up and/or sulphur clean-up. The final extract is analyzed by dual capillary column gas chromatography with electron capture detection (GC/ECD). | 5e-06 | Y | Y | Y | 10 | 9/20/2016 | 9/21/2016 | 8/15/2019 2:42:34 PM | 2019-07-08 | 8/15/2019 11:39:30 AM | KATFRASE | KATFRASE | Concentration | mg/L | Y | (fl. conc.) |  | 2;4'-DDD (fl. conc.) | 53-19-0 | Pesticides; | Pesticides; EMS code: D054 | Chemical |  | LAB
D054 | ug/g | 2;4'-DDD | E8081 | EPA 8081 | See methods EPA 3540, 3630, 3660, 8081 | 0.001 | Y | Y | Y | 9 | 3/18/2020 | 3/23/2020 | 2/22/2023 12:51:03 PM | 2022-07-29 |  | EMS |  | MassPerMass | mg/kg | Y | (mass conc.) |  | 2;4'-DDD (mass conc.) | 53-19-0 | Pesticides; | Pesticides; EMS code: D054 | Chemical |  | LAB
D055 | mg/L | Dichlofenthion | ZN01 | DCM+Ac.Extr:GC/MS | DCM+Ac.Extr:GC/MS | 0.0005 | Y | Y | Y | 28 | 7/29/2005 7:10:00 AM | 3/29/2012 10:15:00 PM | 5/3/2012 9:27:50 AM | 1996-04-12 | 11/16/1999 4:06:27 PM | EMS | NPEPPIN | Concentration | mg/L | Y | (fl. conc.) |  | Dichlofenthion (fl. conc.) |  | Pesticides; | Pesticides; EMS code: D055 | Chemical |  | LAB
D060 | mg/L | Dichlorophenols | PP11 | Ac. anhydride ext:GC/MS | Ac. anhydride ext:GC/MS | 5e-05 | Y | Y | Y | 232 |  |  | 10/24/1996 5:30:13 AM | 1996-04-12 | 6/3/1998 6:51:35 PM | EMS |  | Concentration | mg/L | Y | (fl. conc.) |  | Dichlorophenols (fl. conc.) |  | Organics; | Organics; EMS code: D060 | Chemical |  | LAB
D061 | mg/L | Dichloroguaiacol | PP11 | Ac. anhydride ext:GC/MS | Ac. anhydride ext:GC/MS | 5e-05 | Y | Y | Y | 227 |  |  | 10/24/1996 5:30:13 AM | 1996-04-12 | 6/3/1998 6:51:35 PM | EMS |  | Concentration | mg/L | Y | (fl. conc.) |  | Dichloroguaiacol (fl. conc.) |  | Organics; | Organics; EMS code: D061 | Chemical |  | LAB
D062 | mg/L | Dichlorocatechols | PP11 | Ac. anhydride ext:GC/MS | Ac. anhydride ext:GC/MS | 5e-05 | Y | Y | Y | 241 |  |  | 10/24/1996 5:30:13 AM | 1996-04-12 | 6/3/1998 6:51:46 PM | EMS |  | Concentration | mg/L | Y | (fl. conc.) |  | Dichlorocatechols (fl. conc.) |  | Organics; | Organics; EMS code: D062 | Chemical |  | LAB
D063 | mg/L | Dichloroveretrols | PP11 | Ac. anhydride ext:GC/MS | Ac. anhydride ext:GC/MS | 5e-05 | Y | Y | Y | 240 |  |  | 10/24/1996 5:30:20 AM | 1996-04-12 | 6/3/1998 6:51:35 PM | EMS |  | Concentration | mg/L | Y | (fl. conc.) |  | Dichloroveretrols (fl. conc.) |  | Organics; | Organics; EMS code: D063 | Chemical |  | LAB
D064 | mg/L | Dichlorovanillins | PP11 | Ac. anhydride ext:GC/MS | Ac. anhydride ext:GC/MS | 5e-05 | Y | Y | Y | 240 |  |  | 10/24/1996 5:30:20 AM | 1996-04-12 | 6/3/1998 6:51:35 PM | EMS |  | Concentration | mg/L | Y | (fl. conc.) |  | Dichlorovanillins (fl. conc.) |  | Organics; | Organics; EMS code: D064 | Chemical |  | LAB
D065 | mg/L | Dichlorbenil | ZN01 | DCM+Ac.Extr:GC/MS | DCM+Ac.Extr:GC/MS | 0.0002 | Y | Y | Y | 28 | 7/28/2005 7:03:00 AM | 3/29/2012 10:15:00 PM | 5/3/2012 9:27:50 AM | 1996-04-12 | 11/16/1999 4:06:27 PM | EMS | NPEPPIN | Concentration | mg/L | Y | (fl. conc.) |  | Dichlorbenil (fl. conc.) |  | Pesticides; | Pesticides; EMS code: D065 | Chemical |  | LAB
D066 | mg/L | Dichlorfluanid | ZN01 | DCM+Ac.Extr:GC/MS | DCM+Ac.Extr:GC/MS | 0.0002 | Y | Y | Y | 28 | 7/28/2005 7:03:00 AM | 3/29/2012 10:15:00 PM | 5/3/2012 9:27:50 AM | 1996-04-12 | 11/16/1999 4:06:27 PM | EMS | NPEPPIN | Concentration | mg/L | Y | (fl. conc.) |  | Dichlorfluanid (fl. conc.) |  | Pesticides; | Pesticides; EMS code: D066 | Chemical |  | LAB
D069 | mg/L | 2;4'-DDE | MW03 | OCP1-LL-SF-ECD-VA | This analysis is carried out using procedures adapted from "Test Methods for Evaluating Solid Waste" SW-846, Methods 3510, 3630, 3660 and 8081, published by the United States Environmental Protection Agency (EPA). The procedure involves extraction of the entire water sample with dichloromethane. The extract then undergoes one or more of the following clean-up procedures (if required): silica gel clean-up and/or sulphur clean-up. The final extract is analyzed by dual capillary column gas chromatography with electron capture detection (GC/ECD). | 5e-06 | Y | Y | Y | 10 | 9/20/2016 | 9/21/2016 | 8/15/2019 2:42:34 PM | 2019-07-08 | 8/15/2019 11:39:52 AM | KATFRASE | KATFRASE | Concentration | mg/L | Y | (fl. conc.) |  | 2;4'-DDE (fl. conc.) | 3424-82-6 | Pesticides; | Pesticides; EMS code: D069 | Chemical |  | LAB
D069 | mg/L | 2;4'-DDE | E660 | EPA 8270E (mod) | EPA 8270E (mod) Miscellaneous Pesticides by GC-MS | 4e-07 | Y | Y | Y | 2 | 9/14/2022 | 9/14/2022 | 12/8/2022 9:48:13 AM | 2022-12-06 |  | EMS |  | Concentration | mg/L | Y | (fl. conc.) |  | 2;4'-DDE (fl. conc.) | 3424-82-6 | Pesticides; | Pesticides; EMS code: D069 | Chemical |  | LAB
D069 | ug/g | 2;4'-DDE | E8081 | EPA 8081 | See methods EPA 3540, 3630, 3660, 8081 | 0.001 | Y | Y | Y | 9 | 3/18/2020 | 3/23/2020 | 2/22/2023 12:51:03 PM | 2022-07-29 |  | EMS |  | MassPerMass | mg/kg | Y | (mass conc.) |  | 2;4'-DDE (mass conc.) | 3424-82-6 | Pesticides; | Pesticides; EMS code: D069 | Chemical |  | LAB
D069 | ug/g | 2;4'-DDE | ZN03 | DCM+Ac.Extr:GC/MS | DCM+Ac.Extr:GC/MS EPA 8270E (mod) | 0.02 | Y | Y | Y | 17 | 1/17/2022 | 11/14/2022 | 1/11/2024 1:07:07 PM | 2022-06-23 |  | EMS |  | MassPerMass | mg/kg | Y | (mass conc.) |  | 2;4'-DDE (mass conc.) | 3424-82-6 | Pesticides; | Pesticides; EMS code: D069 | Chemical |  | LAB
D069 | mg/L | 2;4'-DDE | X268 | CH2Cl2 ext:GC/MS | CH2Cl2 ext:GC/MS |  | Y | Y | Y | 16 | 12/29/2021 | 9/2/2022 | 11/10/2022 9:54:06 AM | 2022-04-08 |  | EMS |  | Concentration | mg/L | Y | (fl. conc.) |  | 2;4'-DDE (fl. conc.) | 3424-82-6 | Pesticides; | Pesticides; EMS code: D069 | Chemical |  | LAB
D072 | mg/L | Diallate | ZN01 | DCM+Ac.Extr:GC/MS | DCM+Ac.Extr:GC/MS | 0.0005 | Y | Y | Y | 31 | 7/28/2005 7:03:00 AM | 3/29/2012 10:15:00 PM | 5/3/2012 9:27:50 AM | 1996-04-12 | 11/16/1999 4:06:27 PM | EMS | NPEPPIN | Concentration | mg/L | Y | (fl. conc.) |  | Diallate (fl. conc.) | 2303-16-4 | Pesticides; | Pesticides; EMS code: D072 | Chemical |  | LAB
D072 | ug/g | Diallate | ZN03 | DCM+Ac.Extr:GC/MS | DCM+Ac.Extr:GC/MS EPA 8270E (mod) | 0.05 | Y | Y | Y | 6 |  |  | 10/24/1996 3:05:57 PM | 1996-04-12 | 6/3/1998 6:51:35 PM | EMS |  | MassPerMass | mg/kg | Y | (mass conc.) |  | Diallate (mass conc.) | 2303-16-4 | Pesticides; | Pesticides; EMS code: D072 | Chemical |  | LAB
D073 | mg/L | 2;3-Dichlorophenol | X268 | CH2Cl2 ext:GC/MS | CH2Cl2 ext:GC/MS | 0.0003 | Y | Y | Y | 26 | 12/28/2021 | 6/7/2024 | 6/26/2024 4:28:10 PM | 2012-06-25 |  | RDGIBSON |  | Concentration | mg/L | Y | (fl. conc.) |  | 2;3-Dichlorophenol (fl. conc.) |  | Organics; | Organics; EMS code: D073 | Chemical |  | LAB
D073 | ug/g | 2;3-Dichlorophenol | 1332 | Phenolics by Tumbler/GC-MS | Phenolics by Tumbler/GC-MS | 0.02 | Y | Y | Y | 30 | 9/9/2019 | 9/22/2022 8:16:00 PM | 1/11/2024 1:07:07 PM | 2019-09-30 |  | KATFRASE |  | MassPerMass | mg/kg | Y | (mass conc.) |  | 2;3-Dichlorophenol (mass conc.) |  | Organics; | Organics; EMS code: D073 | Chemical |  | LAB
D073 | mg/L | 2;3-Dichlorophenol | F028 | Acid,Extr.MeCl,HRGC/LRMS:TIM,SIM | A measured volume of sample is acidified and extracted with Methylene Chloride. For samples which have a high organic content that potentially produce chromatographic interferences an initial pre-extraction under basic conditions is performed. The sample extract is then evaporated to small volume and analysed by High Resolution Gas Chromatography / Low Resolution Mass Spectrometry (HRGC/LRMS). LRMS data is acquired in Total Ion Mode (TIM) and quantitation is performed using a single mass ion. Positive Identification of the various target compounds is made by comparison of absolute retention times to those of standards combined with computer based mass spectral mass/intensity library searching and matching routines. | 0.0001 | Y | Y | Y | 28 | 8/20/2009 | 11/1/2022 4:39:00 AM | 4/21/2023 4:26:45 PM | 2009-12-23 |  | RDGIBSON |  | Concentration | mg/L | Y | (fl. conc.) |  | 2;3-Dichlorophenol (fl. conc.) |  | Organics; | Organics; EMS code: D073 | Chemical |  | LAB
D076 | mg/L | 2;6-Dichlorophenol | F028 | Acid,Extr.MeCl,HRGC/LRMS:TIM,SIM | A measured volume of sample is acidified and extracted with Methylene Chloride. For samples which have a high organic content that potentially produce chromatographic interferences an initial pre-extraction under basic conditions is performed. The sample extract is then evaporated to small volume and analysed by High Resolution Gas Chromatography / Low Resolution Mass Spectrometry (HRGC/LRMS). LRMS data is acquired in Total Ion Mode (TIM) and quantitation is performed using a single mass ion. Positive Identification of the various target compounds is made by comparison of absolute retention times to those of standards combined with computer based mass spectral mass/intensity library searching and matching routines. | 0.0001 | Y | Y | Y | 28 | 8/20/2009 | 11/1/2022 4:39:00 AM | 4/21/2023 4:26:45 PM | 2009-12-23 |  | RDGIBSON |  | Concentration | mg/L | Y | (fl. conc.) |  | 2;6-Dichlorophenol (fl. conc.) |  | Organics; | Organics; EMS code: D076 | Chemical |  | LAB
D076 | mg/L | 2;6-Dichlorophenol | X268 | CH2Cl2 ext:GC/MS | CH2Cl2 ext:GC/MS | 0.0003 | Y | Y | Y | 26 | 12/28/2021 | 6/7/2024 | 6/26/2024 4:28:10 PM | 2012-06-25 |  | RDGIBSON |  | Concentration | mg/L | Y | (fl. conc.) |  | 2;6-Dichlorophenol (fl. conc.) |  | Organics; | Organics; EMS code: D076 | Chemical |  | LAB
D076 | ug/g | 2;6-Dichlorophenol | 1332 | Phenolics by Tumbler/GC-MS | Phenolics by Tumbler/GC-MS | 0.04 | Y | Y | Y | 30 | 9/9/2019 | 9/22/2022 8:16:00 PM | 1/11/2024 1:07:07 PM | 2019-09-30 |  | KATFRASE |  | MassPerMass | mg/kg | Y | (mass conc.) |  | 2;6-Dichlorophenol (mass conc.) |  | Organics; | Organics; EMS code: D076 | Chemical |  | LAB
D077 | ug/g | 3;5-Dichlorophenol | 1332 | Phenolics by Tumbler/GC-MS | Phenolics by Tumbler/GC-MS | 0.04 | Y | Y | Y | 30 | 9/9/2019 | 9/22/2022 8:16:00 PM | 1/11/2024 1:07:07 PM | 2019-09-30 |  | KATFRASE |  | MassPerMass | mg/kg | Y | (mass conc.) |  | 3;5-Dichlorophenol (mass conc.) |  | Organics; | Organics; EMS code: D077 | Chemical |  | LAB
D077 | mg/L | 3;5-Dichlorophenol | F028 | Acid,Extr.MeCl,HRGC/LRMS:TIM,SIM | A measured volume of sample is acidified and extracted with Methylene Chloride. For samples which have a high organic content that potentially produce chromatographic interferences an initial pre-extraction under basic conditions is performed. The sample extract is then evaporated to small volume and analysed by High Resolution Gas Chromatography / Low Resolution Mass Spectrometry (HRGC/LRMS). LRMS data is acquired in Total Ion Mode (TIM) and quantitation is performed using a single mass ion. Positive Identification of the various target compounds is made by comparison of absolute retention times to those of standards combined with computer based mass spectral mass/intensity library searching and matching routines. | 0.0001 | Y | Y | Y | 28 | 8/20/2009 | 11/1/2022 4:39:00 AM | 4/21/2023 4:26:45 PM | 2009-12-23 |  | RDGIBSON |  | Concentration | mg/L | Y | (fl. conc.) |  | 3;5-Dichlorophenol (fl. conc.) |  | Organics; | Organics; EMS code: D077 | Chemical |  | LAB
D077 | mg/L | 3;5-Dichlorophenol | X268 | CH2Cl2 ext:GC/MS | CH2Cl2 ext:GC/MS | 0.0003 | Y | Y | Y | 26 | 12/28/2021 | 6/7/2024 | 6/26/2024 4:28:10 PM | 2012-06-25 |  | RDGIBSON |  | Concentration | mg/L | Y | (fl. conc.) |  | 3;5-Dichlorophenol (fl. conc.) |  | Organics; | Organics; EMS code: D077 | Chemical |  | LAB
D078 | mg/L | Diuron | ZN01 | DCM+Ac.Extr:GC/MS | DCM+Ac.Extr:GC/MS | 0.0005 | Y | Y | Y | 1 | 7/29/2005 7:11:00 AM | 7/29/2005 7:11:00 AM | 6/2/2009 12:01:03 PM | 2003-06-16 |  | UMULLER |  | Concentration | mg/L | Y | (fl. conc.) |  | Diuron (fl. conc.) | 330-54-1 | Pesticides; | Pesticides; EMS code: D078 | Chemical |  | LAB
D078 | ug/g | Diuron | ZN03 | DCM+Ac.Extr:GC/MS | DCM+Ac.Extr:GC/MS EPA 8270E (mod) | 0.05 | Y | Y | Y | 2 | 11/2/2022 12:53:00 AM | 11/2/2022 4:41:00 AM | 1/7/2023 3:12:10 PM | 2003-06-16 |  | UMULLER |  | MassPerMass | mg/kg | Y | (mass conc.) |  | Diuron (mass conc.) | 330-54-1 | Pesticides; | Pesticides; EMS code: D078 | Chemical |  | LAB
D078 | mg/L | Diuron | E751 | E751D Pesticides in Water by LC-MS/MS (POS1 List) | E751D Pesticides in Water by LC-MS/MS (POS1 List) (MOECC E3501) | 5e-05 | Y | Y | Y | 39 | 12/9/2021 | 12/31/2021 | 2/15/2022 9:50:05 AM | 2021-12-14 |  | EMS |  | Concentration | mg/L | Y | (fl. conc.) |  | Diuron (fl. conc.) | 330-54-1 | Pesticides; | Pesticides; EMS code: D078 | Chemical |  | LAB
D079 | ug/g | Deltamethring | ZN03 | DCM+Ac.Extr:GC/MS | DCM+Ac.Extr:GC/MS EPA 8270E (mod) | 0.05 | Y | Y | Y | 2 | 11/2/2022 12:53:00 AM | 11/2/2022 4:41:00 AM | 1/7/2023 3:12:10 PM | 2003-06-16 |  | UMULLER |  | MassPerMass | mg/kg | Y | (mass conc.) |  | Deltamethring (mass conc.) |  | Pesticides; | Pesticides; EMS code: D079 | Chemical |  | LAB
D079 | mg/L | Deltamethring | ZN01 | DCM+Ac.Extr:GC/MS | DCM+Ac.Extr:GC/MS | 0.0005 | Y | Y | Y | 1 | 7/29/2005 7:10:00 AM | 7/29/2005 7:10:00 AM | 6/2/2009 12:01:03 PM | 2003-06-16 |  | UMULLER |  | Concentration | mg/L | Y | (fl. conc.) |  | Deltamethring (fl. conc.) |  | Pesticides; | Pesticides; EMS code: D079 | Chemical |  | LAB
D080 | mg/L | Dialifos | ZN01 | DCM+Ac.Extr:GC/MS | DCM+Ac.Extr:GC/MS | 0.0005 | Y | Y | Y | 1 | 7/29/2005 7:10:00 AM | 7/29/2005 7:10:00 AM | 6/2/2009 12:01:03 PM | 2003-06-16 |  | UMULLER |  | Concentration | mg/L | Y | (fl. conc.) |  | Dialifos (fl. conc.) |  | Pesticides; | Pesticides; EMS code: D080 | Chemical |  | LAB
D081 | mg/L | Dimethachlor | ZN01 | DCM+Ac.Extr:GC/MS | DCM+Ac.Extr:GC/MS | 0.0005 | Y | Y | Y | 1 | 7/29/2005 7:10:00 AM | 7/29/2005 7:10:00 AM | 6/2/2009 12:01:04 PM | 2003-06-16 |  | UMULLER |  | Concentration | mg/L | Y | (fl. conc.) |  | Dimethachlor (fl. conc.) |  | Pesticides; | Pesticides; EMS code: D081 | Chemical |  | LAB
D082 | mg/L | Dinitramine | ZN01 | DCM+Ac.Extr:GC/MS | DCM+Ac.Extr:GC/MS | 0.0005 | Y | Y | Y | 1 | 7/29/2005 7:10:00 AM | 7/29/2005 7:10:00 AM | 6/2/2009 12:01:04 PM | 2003-06-16 |  | UMULLER |  | Concentration | mg/L | Y | (fl. conc.) |  | Dinitramine (fl. conc.) |  | Pesticides; | Pesticides; EMS code: D082 | Chemical |  | LAB
D083 | mg/L | Diphenamid | ZN01 | DCM+Ac.Extr:GC/MS | DCM+Ac.Extr:GC/MS | 0.0005 | Y | Y | Y | 1 | 7/29/2005 7:10:00 AM | 7/29/2005 7:10:00 AM | 6/2/2009 12:01:04 PM | 2003-06-16 |  | UMULLER |  | Concentration | mg/L | Y | (fl. conc.) |  | Diphenamid (fl. conc.) | 957-51-7 | Pesticides; | Pesticides; EMS code: D083 | Chemical |  | LAB
D084 | mg/L | o-Cresol | X268 | CH2Cl2 ext:GC/MS | CH2Cl2 ext:GC/MS | 0.0005 | Y | Y | Y | 859 | 3/6/2007 | 8/14/2024 | 8/16/2024 4:05:21 PM | 2012-03-27 |  | RDGIBSON |  | Concentration | mg/L | Y | (fl. conc.) |  | o-Cresol (fl. conc.) |  | Organics; | Organics; EMS code: D084 | Chemical |  | LAB
D084 | ug/g | o-Cresol | 1332 | Phenolics by Tumbler/GC-MS | Phenolics by Tumbler/GC-MS | 0.02 | Y | Y | Y | 15 | 1/17/2022 | 2/2/2022 | 1/11/2024 1:07:07 PM | 2022-07-07 |  | EMS |  | MassPerMass | mg/kg | Y | (mass conc.) |  | o-Cresol (mass conc.) |  | Organics; | Organics; EMS code: D084 | Chemical |  | LAB
D085 | ug/g | m-Cresol | 1332 | Phenolics by Tumbler/GC-MS | Phenolics by Tumbler/GC-MS | 0.02 | Y | Y | Y | 15 | 1/17/2022 | 2/2/2022 | 1/11/2024 1:07:07 PM | 2022-07-07 |  | EMS |  | MassPerMass | mg/kg | Y | (mass conc.) |  | m-Cresol (mass conc.) |  | Organics; | Organics; EMS code: D085 | Chemical |  | LAB
D085 | mg/L | m-Cresol | X268 | CH2Cl2 ext:GC/MS | CH2Cl2 ext:GC/MS | 0.0005 | Y | Y | Y | 859 | 3/6/2007 | 8/14/2024 | 8/16/2024 4:05:21 PM | 2012-03-27 |  | RDGIBSON |  | Concentration | mg/L | Y | (fl. conc.) |  | m-Cresol (fl. conc.) |  | Organics; | Organics; EMS code: D085 | Chemical |  | LAB
D086 | mg/L | p-Cresol | X268 | CH2Cl2 ext:GC/MS | CH2Cl2 ext:GC/MS | 0.0005 | Y | Y | Y | 859 | 3/6/2007 | 8/14/2024 | 8/16/2024 4:05:21 PM | 2012-03-27 |  | RDGIBSON |  | Concentration | mg/L | Y | (fl. conc.) |  | p-Cresol (fl. conc.) |  | Organics; | Organics; EMS code: D086 | Chemical |  | LAB
D086 | ug/g | p-Cresol | 1332 | Phenolics by Tumbler/GC-MS | Phenolics by Tumbler/GC-MS | 0.02 | Y | Y | Y | 15 | 1/17/2022 | 2/2/2022 | 1/11/2024 1:07:07 PM | 2022-07-07 |  | EMS |  | MassPerMass | mg/kg | Y | (mass conc.) |  | p-Cresol (mass conc.) |  | Organics; | Organics; EMS code: D086 | Chemical |  | LAB
D088 | ug/g | 2;4'-DDT | E8081 | EPA 8081 | See methods EPA 3540, 3630, 3660, 8081 | 0.002 | Y | Y | Y | 9 | 3/18/2020 | 3/23/2020 | 2/22/2023 12:51:03 PM | 2022-07-29 |  | EMS |  | MassPerMass | mg/kg | Y | (mass conc.) |  | 2;4'-DDT (mass conc.) | 789-02-6 | Pesticides; | Pesticides; EMS code: D088 | Chemical |  | LAB
D088 | mg/L | 2;4'-DDT | E660 | EPA 8270E (mod) | EPA 8270E (mod) Miscellaneous Pesticides by GC-MS | 4e-07 | Y | Y | Y | 2 | 9/14/2022 | 9/14/2022 | 12/8/2022 9:48:13 AM | 2022-12-06 |  | EMS |  | Concentration | mg/L | Y | (fl. conc.) |  | 2;4'-DDT (fl. conc.) | 789-02-6 | Pesticides; | Pesticides; EMS code: D088 | Chemical |  | LAB
D088 | ug/g | 2;4'-DDT | ZN03 | DCM+Ac.Extr:GC/MS | DCM+Ac.Extr:GC/MS EPA 8270E (mod) | 0.02 | Y | Y | Y | 17 | 1/17/2022 | 11/14/2022 | 1/11/2024 1:07:07 PM | 2022-06-23 |  | EMS |  | MassPerMass | mg/kg | Y | (mass conc.) |  | 2;4'-DDT (mass conc.) | 789-02-6 | Pesticides; | Pesticides; EMS code: D088 | Chemical |  | LAB
D088 | mg/L | 2;4'-DDT | X268 | CH2Cl2 ext:GC/MS | CH2Cl2 ext:GC/MS |  | Y | Y | Y | 16 | 12/29/2021 | 9/2/2022 | 11/10/2022 9:54:06 AM | 2022-04-08 |  | EMS |  | Concentration | mg/L | Y | (fl. conc.) |  | 2;4'-DDT (fl. conc.) | 789-02-6 | Pesticides; | Pesticides; EMS code: D088 | Chemical |  | LAB
D088 | mg/L | 2;4'-DDT | MW03 | OCP1-LL-SF-ECD-VA | This analysis is carried out using procedures adapted from "Test Methods for Evaluating Solid Waste" SW-846, Methods 3510, 3630, 3660 and 8081, published by the United States Environmental Protection Agency (EPA). The procedure involves extraction of the entire water sample with dichloromethane. The extract then undergoes one or more of the following clean-up procedures (if required): silica gel clean-up and/or sulphur clean-up. The final extract is analyzed by dual capillary column gas chromatography with electron capture detection (GC/ECD). | 5e-06 | Y | Y | Y | 10 | 9/20/2016 | 9/21/2016 | 8/15/2019 2:42:34 PM | 2019-07-08 | 8/15/2019 11:40:15 AM | KATFRASE | KATFRASE | Concentration | mg/L | Y | (fl. conc.) |  | 2;4'-DDT (fl. conc.) | 789-02-6 | Pesticides; | Pesticides; EMS code: D088 | Chemical |  | LAB
D101 | mg/L | DDD-o;p | ZN01 | DCM+Ac.Extr:GC/MS | DCM+Ac.Extr:GC/MS | 0.1 | Y | Y | Y | 1 | 7/28/2005 7:03:00 AM | 7/28/2005 7:03:00 AM | 6/2/2009 12:01:04 PM | 2009-02-20 |  | JCHUDAK |  | Concentration | mg/L | Y | (fl. conc.) |  | DDD-o;p (fl. conc.) | 53-19-0 | Pesticides; | Pesticides; EMS code: D101 | Chemical |  | LAB
D102 | mg/L | Alkalinity pH 4.5/4.2 | 1212 | Pot. Tit. SA pH 4.5/4.2 | Alkalinity is determined by potentiometric titration with sulfuric acid which has been standardized with sodium carbonate. The sample (if pH>8.3) is first titrated to a pH=8.3, which corresponds to the equivalence point for the conversion of carbonate to bicarbonate ion (phenolphthalein alkalinity). The titration is continued to a pH of 4.5, which corresponds approximately to the equivalence point for the conversion of bicarbonate ion to carbonic acid (entire titration indicates total alkalinity). An extrapolation technique based on the near proportionality of concentration of hydrogen ions to excess of titrant beyond the equivalence point. The amount of standard acid required to reduce pH exactly 0.30 pH (i.e. pH 4.2) unit is measured carefully. Because this change in pH corresponds to an exact coupling of the hydrogen ion concentration, a simple extrapolation can be made to the equivalence point. The actual pH of the stoichiometric end point in alkalinity determinations is the pH at which the inflection in the titration curve occurs and is directly related to the amount of carbonate alkalinity orginally present in the sample. A curve can be constructed by plotting pH versus mL of acid to find this point directly. To obtain the actual substances contributing to alkalinity, such a curve should be constructed. | 0.5 | Y | Y | Y | 5255 | 8/1/1996 4:17:00 PM | 12/22/2015 2:06:00 PM | 9/11/2019 8:47:13 AM | 1996-04-12 | 6/3/1998 6:51:36 PM | EMS |  | Concentration | mg/L | Y | (fl. conc.) |  | Alkalinity pH 4.5/4.2 (fl. conc.) |  | Major Ions; | Major Ions; EMS code: D102 | Chemical |  | LAB
DBAA | mg/L | Dibromoacetic Acid | HAAW | Haloacetic Acids,MTBE ex | Haloacetic Acids;MTBE ex | 0.4 | Y | Y | Y | 2 | 3/27/2023 9:58:00 AM | 5/31/2023 12:14:00 PM | 6/8/2023 12:14:57 PM | 1996-10-13 | 6/3/1998 6:51:46 PM | EMS |  | Concentration | mg/L | Y | (fl. conc.) |  | Dibromoacetic Acid (fl. conc.) | 79-43-6 | Organics; | Organics; EMS code: DBAA | Chemical |  | LAB
DBRM | ug/g | Dibromomethane | MS02 | Purge&Trap GC/MS (VOC-S) | Purge&Trap GC/MS (VOC-S) |  | Y | Y | Y | 1 | 10/28/2022 5:04:00 PM | 10/28/2022 5:04:00 PM | 1/7/2023 3:12:09 PM | 2007-10-04 |  | RDGIBSON |  | MassPerMass | mg/kg | Y | (mass conc.) |  | Dibromomethane (mass conc.) | 74-95-3 | Volatile Organic Compounds (VOCs); | Volatile Organic Compounds (VOCs); EMS code: DBRM | Chemical |  | LAB
DBRM | mg/L | Dibromomethane | MS01 | Purge&Trap GC/MS (VOC-W) | Purge&Trap GC/MS (VOC-W) |  | Y | Y | Y | 212 | 6/3/2014 9:09:00 PM | 6/4/2024 10:24:00 AM | 6/13/2024 5:01:27 PM | 2007-10-04 |  | RDGIBSON |  | Concentration | mg/L | Y | (fl. conc.) |  | Dibromomethane (fl. conc.) | 74-95-3 | Volatile Organic Compounds (VOCs); | Volatile Organic Compounds (VOCs); EMS code: DBRM | Chemical |  | LAB
DC01 | mg/L | Decanoic Acid | RFAS | Separation of resin and fatty acids (RFAs) in water by GC/MS | Scientific Principle The separation of resin and fatty acids (RFAs) in water samples is performed using gas chromatography/mass spectrometry (GC/MS) in the Selected Ion Monitoring (SIM) mode. The separation is based on compound size; boiling point and polarity. Analyte concentration is calculated using the internal standard method. RFA compounds are extracted from water samples using a liquid-liquid solvent extraction with DCM. Extraction occurs after the sample is brought to a pH of ~ 2 and placed in a rotary tumbler (at approx. 40 rpm) for 1 hour. The resulting extract is transferred to a vial; dried with acidified baked sodium sulfate; methylated with diazomethane and concentrated to 1mL under nitrogen prior to analysis. The resulting extracts are vialed and fortified with labeled internal standard compounds prior to analysis by gas chromatography with a (50%-Phenyl)-methylpolysiloxane (DB-17) column and a mass spectrometer detector in SIM Mode. | 0.006 | Y | Y | Y | 1 | 9/17/2021 7:49:00 PM | 9/17/2021 7:49:00 PM | 2/1/2022 2:38:14 PM | 2022-01-19 |  | EMS |  | Concentration | mg/L | Y | (fl. conc.) |  | Decanoic Acid (fl. conc.) |  | Organics; | Organics; EMS code: DC01 | Chemical |  | LAB
DC02 | mg/L | Docosanoic Acid | RFAS | Separation of resin and fatty acids (RFAs) in water by GC/MS | Scientific Principle The separation of resin and fatty acids (RFAs) in water samples is performed using gas chromatography/mass spectrometry (GC/MS) in the Selected Ion Monitoring (SIM) mode. The separation is based on compound size; boiling point and polarity. Analyte concentration is calculated using the internal standard method. RFA compounds are extracted from water samples using a liquid-liquid solvent extraction with DCM. Extraction occurs after the sample is brought to a pH of ~ 2 and placed in a rotary tumbler (at approx. 40 rpm) for 1 hour. The resulting extract is transferred to a vial; dried with acidified baked sodium sulfate; methylated with diazomethane and concentrated to 1mL under nitrogen prior to analysis. The resulting extracts are vialed and fortified with labeled internal standard compounds prior to analysis by gas chromatography with a (50%-Phenyl)-methylpolysiloxane (DB-17) column and a mass spectrometer detector in SIM Mode. | 0.006 | Y | Y | Y | 1 | 9/17/2021 7:49:00 PM | 9/17/2021 7:49:00 PM | 2/1/2022 2:38:14 PM | 2022-01-19 |  | EMS |  | Concentration | mg/L | Y | (fl. conc.) |  | Docosanoic Acid (fl. conc.) |  | Organics; | Organics; EMS code: DC02 | Chemical |  | LAB
DC03 | mg/L | Dodecanoic Acid | RFAS | Separation of resin and fatty acids (RFAs) in water by GC/MS | Scientific Principle The separation of resin and fatty acids (RFAs) in water samples is performed using gas chromatography/mass spectrometry (GC/MS) in the Selected Ion Monitoring (SIM) mode. The separation is based on compound size; boiling point and polarity. Analyte concentration is calculated using the internal standard method. RFA compounds are extracted from water samples using a liquid-liquid solvent extraction with DCM. Extraction occurs after the sample is brought to a pH of ~ 2 and placed in a rotary tumbler (at approx. 40 rpm) for 1 hour. The resulting extract is transferred to a vial; dried with acidified baked sodium sulfate; methylated with diazomethane and concentrated to 1mL under nitrogen prior to analysis. The resulting extracts are vialed and fortified with labeled internal standard compounds prior to analysis by gas chromatography with a (50%-Phenyl)-methylpolysiloxane (DB-17) column and a mass spectrometer detector in SIM Mode. | 0.006 | Y | Y | Y | 1 | 9/17/2021 7:49:00 PM | 9/17/2021 7:49:00 PM | 2/1/2022 2:38:14 PM | 2022-01-19 |  | EMS |  | Concentration | mg/L | Y | (fl. conc.) |  | Dodecanoic Acid (fl. conc.) |  | Organics; | Organics; EMS code: DC03 | Chemical |  | LAB
DCAA | mg/L | Dichloroacetic Acid | HAAW | Haloacetic Acids,MTBE ex | Haloacetic Acids;MTBE ex | 0.4 | Y | Y | Y | 2 | 3/27/2023 9:58:00 AM | 5/31/2023 12:14:00 PM | 6/8/2023 12:14:57 PM | 1996-10-13 | 6/3/1998 6:51:46 PM | EMS |  | Concentration | mg/L | Y | (fl. conc.) |  | Dichloroacetic Acid (fl. conc.) | 79-43-6 | Organics; | Organics; EMS code: DCAA | Chemical |  | LAB
DCBP | % | Decachlorobiphenyl CAS: 2051-24-3 | SURR | Surrogate Recovery | Surrogate Recovery |  | Y | Y | Y | 71 | 12/28/2021 | 11/1/2023 | 1/11/2024 1:07:07 PM | 2022-04-21 | 4/21/2022 1:43:21 PM | EMS | EMS | DimensionlessRatio | % | N | (ratio) |  | Decachlorobiphenyl (ratio) | 2051-24-3 | Polychlorinated biphenylds (PCBs); | Polychlorinated biphenylds (PCBs); EMS code: DCBP | Chemical |  | SURROGATE_RESULT
DCL3 | % | Surrogate standard used for VH analysis | F083 | He purge+trap,GC/FID,nC6-nC10 | Volatile organic compounds (nC6-nC10) are purged from water samples with helium, adsorbed onto a sorbent trap, and then thermally desorbed to a capillary column gas chromatograph equipped with a flame ionization detector. | 1 | Y | Y | Y | 43 | 5/21/2015 | 5/22/2015 | 7/3/2015 3:19:04 PM | 2015-05-27 |  | RDGIBSON |  | DimensionlessRatio | % | N | (ratio) |  | Surrogate standard used for VH analysis (ratio) |  | Petroleum Hydrocarbons; | Petroleum Hydrocarbons; EMS code: DCL3 | Chemical |  | SURROGATE_RESULT
DCL3 | mg/L | Surrogate standard used for VH analysis | E611 | E611S: EPA 8260D (mod) | E611S: Dimethyl selenide and dimethyl diselenide are analyzed by static headspace GC-MS. Samples are prepared in headspace vials and are heated and agitated on the headspace autosampler; causing VOCs to partition between the aqueous phase and the headspace in accordance with Henry`s law. | 0.001 | Y | Y | Y | 0 |  |  |  | 2024-01-15 | 1/15/2024 2:17:10 PM | EMS | EMS | Concentration | mg/L | Y | (fl. conc.) |  | Surrogate standard used for VH analysis (fl. conc.) |  | Petroleum Hydrocarbons; | Petroleum Hydrocarbons; EMS code: DCL3 | Chemical |  | SURROGATE_RESULT
DCL3 | % | Surrogate standard used for VH analysis | SURR | Surrogate Recovery | Surrogate Recovery | 1 | Y | Y | Y | 602 | 5/13/2016 | 8/12/2024 | 8/15/2024 11:32:04 AM | 2016-07-19 |  | RRENTON |  | DimensionlessRatio | % | N | (ratio) |  | Surrogate standard used for VH analysis (ratio) |  | Petroleum Hydrocarbons; | Petroleum Hydrocarbons; EMS code: DCL3 | Chemical |  | SURROGATE_RESULT
DCLT | % | Surrogate standard used for F1(C6-C10) analysis. | F083 | He purge+trap,GC/FID,nC6-nC10 | Volatile organic compounds (nC6-nC10) are purged from water samples with helium, adsorbed onto a sorbent trap, and then thermally desorbed to a capillary column gas chromatograph equipped with a flame ionization detector. | 1 | Y | Y | Y | 1 | 6/22/2007 | 6/22/2007 | 7/9/2012 2:09:47 PM | 2012-06-25 |  | RDGIBSON |  | DimensionlessRatio | % | N | (ratio) |  | Surrogate standard used for F1(C6-C10) analysis. (ratio) |  | Petroleum Hydrocarbons; | Petroleum Hydrocarbons; EMS code: DCLT | Chemical |  | SURROGATE_RESULT
DCPA | mg/L | DACTHAL (DCPA) | ZN01 | DCM+Ac.Extr:GC/MS | DCM+Ac.Extr:GC/MS | 0.0001 | Y | Y | Y | 27 | 11/1/2008 | 3/29/2012 10:15:00 PM | 5/3/2012 9:27:50 AM | 1999-11-17 |  | NPEPPIN |  | Concentration | mg/L | Y | (fl. conc.) |  | DACTHAL (DCPA) (fl. conc.) | 1861-32-1 | Pesticides; | Pesticides; EMS code: DCPA | Chemical |  | LAB
DCPM | ug/g | Diclofop-methyl | ZN03 | DCM+Ac.Extr:GC/MS | DCM+Ac.Extr:GC/MS EPA 8270E (mod) | 0.05 | Y | Y | Y | 16 | 1/25/2022 | 11/2/2022 4:41:00 AM | 1/7/2023 3:12:10 PM | 2022-06-13 |  | EMS |  | MassPerMass | mg/kg | Y | (mass conc.) |  | Diclofop-methyl (mass conc.) | 51338-27-3 | Pesticides; | Pesticides; EMS code: DCPM | Chemical |  | LAB
DCPM | mg/L | Diclofop-methyl | E660 | EPA 8270E (mod) | EPA 8270E (mod) Miscellaneous Pesticides by GC-MS | 0.0001 | Y | Y | Y | 68 | 12/7/2021 | 9/12/2022 | 12/8/2022 12:57:05 PM | 2021-12-09 | 12/9/2021 3:22:12 PM | EMS | EMS | Concentration | mg/L | Y | (fl. conc.) |  | Diclofop-methyl (fl. conc.) | 51338-27-3 | Pesticides; | Pesticides; EMS code: DCPM | Chemical |  | LAB
DCPR | mg/L | 1;3-Dichloropropene (cis & trans) | F007 | Static Headspace:HRGC/LRMS | Static headspace analysis is a method for the determination of volatile compounds in liquids and/or solids by analysing the vapour phase of a sample in a closed system; in this case a sealed vial. Static Headspace is sometimes referred to as Equilibrium Headspace because the vapour phase above the sample in a sealed vial is brought into thermodynamic equilibrium with the sample prior to analysis. Volatile organic compounds are thermally extracted from the sample matrix by Static Headspace sampling and introduced directly to a thick-film megabore GC column. Compounds eluting from the GC column are introduced into a Mass Spectrometer and target compounds are identified by a comparison of both retention times and mass spectra with those of reference standards stored in a computerized data system. Reference retention times and mass spectra for analytes are obtained by the measurement of calibration standards under precisely the same extraction conditions used for samples. The concentration of each identified component is measured by relating the response of the sample quantitation mass ions to the response of the quantitation mass ions of a series of external standards. A deuterated surrogate recovery standard is added to provide information on matrix effects. | 0.00141 | Y | Y | Y | 322 | 11/10/2011 | 8/14/2024 | 8/16/2024 4:05:21 PM | 2012-05-24 |  | RDGIBSON |  | Concentration | mg/L | Y | (fl. conc.) |  | 1;3-Dichloropropene (cis & trans) (fl. conc.) |  | Volatile Organic Compounds (VOCs); | Volatile Organic Compounds (VOCs); EMS code: DCPR | Chemical |  | LAB
DCPR | ug/g | 1;3-Dichloropropene (cis & trans) | MS02 | Purge&Trap GC/MS (VOC-S) | Purge&Trap GC/MS (VOC-S) | 0.075 | Y | Y | Y | 16 | 1/14/2022 | 10/28/2022 5:04:00 PM | 1/11/2024 1:07:07 PM | 2022-07-27 |  | EMS |  | MassPerMass | mg/kg | Y | (mass conc.) |  | 1;3-Dichloropropene (cis & trans) (mass conc.) |  | Volatile Organic Compounds (VOCs); | Volatile Organic Compounds (VOCs); EMS code: DCPR | Chemical |  | LAB
DCPR | mg/L | 1;3-Dichloropropene (cis & trans) | MS01 | Purge&Trap GC/MS (VOC-W) | Purge&Trap GC/MS (VOC-W) | 1e-06 | Y | Y | Y | 164 | 5/12/2016 7:09:00 AM | 6/4/2024 10:24:00 AM | 6/13/2024 5:01:27 PM | 2016-05-19 |  | RDGIBSON |  | Concentration | mg/L | Y | (fl. conc.) |  | 1;3-Dichloropropene (cis & trans) (fl. conc.) |  | Volatile Organic Compounds (VOCs); | Volatile Organic Compounds (VOCs); EMS code: DCPR | Chemical |  | LAB
DDD-T | ug/g | Total DDD | ZN03 | DCM+Ac.Extr:GC/MS | DCM+Ac.Extr:GC/MS EPA 8270E (mod) | 0.03 | Y | Y | Y | 17 | 1/17/2022 | 11/14/2022 | 1/11/2024 1:07:07 PM | 2022-06-23 |  | EMS |  | MassPerMass | mg/kg | Y | (mass conc.) |  | Total DDD (mass conc.) |  | Pesticides; | Pesticides; EMS code: DDD-T | Chemical |  | LAB
DDD-T | mg/L | Total DDD | X268 | CH2Cl2 ext:GC/MS | CH2Cl2 ext:GC/MS |  | Y | Y | Y | 16 | 12/29/2021 | 9/2/2022 | 11/10/2022 9:54:06 AM | 2022-04-04 |  | EMS |  | Concentration | mg/L | Y | (fl. conc.) |  | Total DDD (fl. conc.) |  | Pesticides; | Pesticides; EMS code: DDD-T | Chemical |  | LAB
DDD-T | mg/L | Total DDD | E660 | EPA 8270E (mod) | EPA 8270E (mod) Miscellaneous Pesticides by GC-MS | 6e-07 | Y | Y | Y | 2 | 9/14/2022 | 9/14/2022 | 12/8/2022 9:48:13 AM | 2022-12-06 |  | EMS |  | Concentration | mg/L | Y | (fl. conc.) |  | Total DDD (fl. conc.) |  | Pesticides; | Pesticides; EMS code: DDD-T | Chemical |  | LAB
DDE-T | mg/L | Total DDE CAS: 72-55-9 | X268 | CH2Cl2 ext:GC/MS | CH2Cl2 ext:GC/MS |  | Y | Y | Y | 16 | 12/29/2021 | 9/2/2022 | 11/10/2022 9:54:07 AM | 2022-04-04 |  | EMS |  | Concentration | mg/L | Y | (fl. conc.) |  | Total DDE (fl. conc.) | 72-55-9 | Pesticides; | Pesticides; EMS code: DDE-T | Chemical |  | LAB
DDE-T | ug/g | Total DDE CAS: 72-55-9 | ZN03 | DCM+Ac.Extr:GC/MS | DCM+Ac.Extr:GC/MS EPA 8270E (mod) | 0.03 | Y | Y | Y | 17 | 1/17/2022 | 11/14/2022 | 1/11/2024 1:07:07 PM | 2022-06-23 |  | EMS |  | MassPerMass | mg/kg | Y | (mass conc.) |  | Total DDE (mass conc.) | 72-55-9 | Pesticides; | Pesticides; EMS code: DDE-T | Chemical |  | LAB
DDE-T | mg/L | Total DDE CAS: 72-55-9 | E660 | EPA 8270E (mod) | EPA 8270E (mod) Miscellaneous Pesticides by GC-MS | 6e-07 | Y | Y | Y | 2 | 9/14/2022 | 9/14/2022 | 12/8/2022 9:48:13 AM | 2022-12-06 |  | EMS |  | Concentration | mg/L | Y | (fl. conc.) |  | Total DDE (fl. conc.) | 72-55-9 | Pesticides; | Pesticides; EMS code: DDE-T | Chemical |  | LAB
DDT-T | ug/g | Total of DDT | ZN03 | DCM+Ac.Extr:GC/MS | DCM+Ac.Extr:GC/MS EPA 8270E (mod) | 0.03 | Y | Y | Y | 19 | 1/17/2022 | 11/14/2022 | 1/11/2024 1:07:07 PM | 2022-06-23 |  | EMS |  | MassPerMass | mg/kg | Y | (mass conc.) |  | Total of DDT (mass conc.) |  | Pesticides; | Pesticides; EMS code: DDT-T | Chemical |  | LAB
DDT-T | mg/L | Total of DDT | E660 | EPA 8270E (mod) | EPA 8270E (mod) Miscellaneous Pesticides by GC-MS | 6e-07 | Y | Y | Y | 2 | 9/14/2022 | 9/14/2022 | 12/8/2022 9:48:13 AM | 2022-12-06 |  | EMS |  | Concentration | mg/L | Y | (fl. conc.) |  | Total of DDT (fl. conc.) |  | Pesticides; | Pesticides; EMS code: DDT-T | Chemical |  | LAB
DDT-T | mg/L | Total of DDT | X268 | CH2Cl2 ext:GC/MS | CH2Cl2 ext:GC/MS |  | Y | Y | Y | 16 | 12/29/2021 | 9/2/2022 | 11/10/2022 9:54:07 AM | 2022-04-04 |  | EMS |  | Concentration | mg/L | Y | (fl. conc.) |  | Total of DDT (fl. conc.) |  | Pesticides; | Pesticides; EMS code: DDT-T | Chemical |  | LAB
DDT-TM | ug/g | DDT + metabolites, total | ZN03 | DCM+Ac.Extr:GC/MS | DCM+Ac.Extr:GC/MS EPA 8270E (mod) | 0.05 | Y | Y | Y | 2 | 11/14/2022 | 11/14/2022 | 12/15/2022 10:50:04 AM | 2022-06-23 |  | EMS |  | MassPerMass | mg/kg | Y | (mass conc.) |  | DDT + metabolites, total (mass conc.) |  | Pesticides; | Pesticides; EMS code: DDT-TM | Chemical | Total | LAB
DDT-TM | mg/L | DDT + metabolites, total | E660 | EPA 8270E (mod) | EPA 8270E (mod) Miscellaneous Pesticides by GC-MS | 1e-06 | Y | Y | Y | 2 | 9/14/2022 | 9/14/2022 | 12/8/2022 9:48:13 AM | 2022-12-06 |  | EMS |  | Concentration | mg/L | Y | (fl. conc.) |  | DDT + metabolites, total (fl. conc.) |  | Pesticides; | Pesticides; EMS code: DDT-TM | Chemical | Total | LAB
DDT-TM | mg/L | DDT + metabolites, total | X268 | CH2Cl2 ext:GC/MS | CH2Cl2 ext:GC/MS |  | Y | Y | Y | 16 | 12/29/2021 | 9/2/2022 | 11/10/2022 9:54:07 AM | 2022-04-04 |  | EMS |  | Concentration | mg/L | Y | (fl. conc.) |  | DDT + metabolites, total (fl. conc.) |  | Pesticides; | Pesticides; EMS code: DDT-TM | Chemical | Total | LAB
DDTP | mg/L | DDD-p;p | ZN01 | DCM+Ac.Extr:GC/MS | DCM+Ac.Extr:GC/MS | 0.1 | Y | Y | Y | 1 | 7/28/2005 7:03:00 AM | 7/28/2005 7:03:00 AM | 6/2/2009 12:01:04 PM | 2009-02-20 |  | JCHUDAK |  | Concentration | mg/L | Y | (fl. conc.) |  | DDD-p;p (fl. conc.) |  | Pesticides; | Pesticides; EMS code: DDTP | Chemical |  | LAB
DENS | mg/L | Bulk Density | FS88 | core,oven dried@105C | Core; oven dried @105C | 0 | Y | Y | Y | 16 | 9/7/2000 9:46:00 AM | 9/7/2000 9:52:00 AM | 4/26/2001 1:04:05 AM | 2001-02-23 | 3/8/2001 9:05:42 AM | NPEPPIN | NPEPPIN | Concentration | mg/L | Y | (fl. conc.) |  | Bulk Density (fl. conc.) |  | Physical Properties; | Physical Properties; EMS code: DENS | Chemical |  | LAB
DESM | mg/L | Desmetryn | ZN01 | DCM+Ac.Extr:GC/MS | DCM+Ac.Extr:GC/MS | 0.0003 | Y | Y | Y | 28 | 7/28/2005 7:03:00 AM | 3/29/2012 10:15:00 PM | 5/3/2012 9:27:50 AM | 1996-04-12 | 11/16/1999 4:06:27 PM | EMS | NPEPPIN | Concentration | mg/L | Y | (fl. conc.) |  | Desmetryn (fl. conc.) |  | Pesticides; | Pesticides; EMS code: DESM | Chemical |  | LAB
DFLF | mg/L | Diflufenican | P015 | Phenoxyacid Herbicides by LC-MS/MS | Phenoxyacid Herbicides by LC-MS/MS | 1e-05 | Y | Y | Y | 68 | 12/6/2021 | 9/9/2022 | 12/8/2022 12:57:05 PM | 2021-12-09 | 12/9/2021 3:22:12 PM | EMS | EMS | Concentration | mg/L | Y | (fl. conc.) |  | Diflufenican (fl. conc.) | 83164-33-4 | Pesticides; | Pesticides; EMS code: DFLF | Chemical |  | LAB
DHFP | mg/L | Dehydronifedipine CAS: 67035-22-7 | XMSP | HPLC and Tri. Quadpole MS/MS | EPA 1694 An aliquot of filtered sample containing 5% organic is injected directly. Each compound is separated by reversed phase HPLC. Analyzed by Electro Spray Ionization MS/MS detection using Triple Quadrupole MS/MS detector. Quantiitation with internal or external standard. |  | Y | Y | Y | 12 | 5/8/2020 | 6/4/2024 | 8/21/2024 3:32:03 PM | 2022-05-04 |  | EMS |  | Concentration | mg/L | Y | (fl. conc.) |  | Dehydronifedipine (fl. conc.) | 67035-22-7 | Pesticides; | Pesticides; EMS code: DHFP | Chemical |  | LAB
DISN | mg/L | Dibutyl Tin | SN01 | Water/Grignard/GCMS | Water/Grignard/GCMS | 0.5 | Y | Y | Y | 11 |  |  | 10/24/1996 1:49:08 PM | 1996-04-12 | 6/3/1998 6:51:36 PM | EMS |  | Concentration | mg/L | Y | (fl. conc.) |  | Dibutyl Tin (fl. conc.) | 1002-53-5 | Organometallics; | Organometallics; EMS code: DISN | Chemical |  | LAB
DISN | mg/L | Dibutyl Tin | X450 | Tropo/Grignard/GCMS | Tropo/Grignard/GCMS | 0.5 | Y | Y | Y | 15 |  |  | 10/24/1996 1:46:01 PM | 1996-04-12 | 6/3/1998 6:51:36 PM | EMS |  | Concentration | mg/L | Y | (fl. conc.) |  | Dibutyl Tin (fl. conc.) | 1002-53-5 | Organometallics; | Organometallics; EMS code: DISN | Chemical |  | LAB
DISN | ug/g | Dibutyl Tin | SN02 | Soil/Grignard/GC/MS | Soil/Grignard/GC/MS | 0.005 | N | Y | Y | 6 |  |  | 10/24/1996 1:45:43 PM | 1996-04-12 | 6/3/1998 6:51:36 PM | EMS |  | MassPerMass | mg/kg | Y | (mass conc.) |  | Dibutyl Tin (mass conc.) | 1002-53-5 | Organometallics; | Organometallics; EMS code: DISN | Chemical |  | LAB
DMDS | mg/L | Dimethyl Diselenide [DMDSe]; as Se | X564 | IC-ICP-CRC-MS | Ion chromatography inductively coupled plasma collision reaction cell mass spectrometry | 1e-05 | Y | Y | Y | 121 |  |  | 11/15/2022 11:41:12 AM | 2022-11-01 |  | EMS |  | Concentration | mg/L | Y | (fl. conc.) |  | Dimethyl Diselenide [DMDSe]; as Se (fl. conc.) | 7101-31-7 | Organometallics; | Organometallics; EMS code: DMDS | Chemical |  | LAB
DMDS | mg/L | Dimethyl Diselenide [DMDSe]; as Se | E611 | E611S: EPA 8260D (mod) | E611S: Dimethyl selenide and dimethyl diselenide are analyzed by static headspace GC-MS. Samples are prepared in headspace vials and are heated and agitated on the headspace autosampler; causing VOCs to partition between the aqueous phase and the headspace in accordance with Henry`s law. | 0.005 | Y | Y | Y | 178 | 12/9/2021 | 8/10/2024 | 8/19/2024 12:19:03 PM | 2020-04-21 |  | RDGIBSON |  | Concentration | mg/L | Y | (fl. conc.) |  | Dimethyl Diselenide [DMDSe]; as Se (fl. conc.) | 7101-31-7 | Organometallics; | Organometallics; EMS code: DMDS | Chemical |  | LAB
DMGC | %(V/V) | Bio. LC50:Daphnia Magna | EQ11 | EQUIS (%V/V) | EQUIS (%V/V) | 0 | N | Y | Y | 38 |  |  | 10/22/1996 10:26:42 PM | 1996-04-12 | 6/3/1998 6:51:36 PM | EMS |  | DimensionlessRatio | % | N | (ratio) | (by vol.) | Bio. LC50:Daphnia Magna (by vol.) (ratio) |  | Toxicity; | Toxicity; EMS code: DMGC | Chemical |  | LAB
DMGC | % (Mortality) | Bio. LC50:Daphnia Magna | FA10 | Daph.Magna/48hr(%mortality)/100% conc. | Daph.Magna/48 hr(%mortality)/100% concentration | 0 | Y | Y | Y | 27280 | 1/4/1997 | 7/31/2024 | 8/23/2024 9:49:37 AM | 1997-02-26 | 6/3/1998 6:51:46 PM | MPAK |  | DimensionlessRatio | % | N | (ratio) | (mortality) | Bio. LC50:Daphnia Magna (mortality) (ratio) |  | Toxicity; | Toxicity; EMS code: DMGC | Chemical |  | LAB
DMGC | %(V/V) | Bio. LC50:Daphnia Magna | X296 | Daph. Magna/48hour/Vol. | This method measures the concentration (expressed as volume/volume) of sample at which 50% mortality occurs. | 0 | Y | Y | Y | 19487 | 1/14/1996 | 8/2/2024 | 8/23/2024 9:49:37 AM | 1996-04-12 | 6/3/1998 6:51:36 PM | EMS |  | DimensionlessRatio | % | N | (ratio) | (by vol.) | Bio. LC50:Daphnia Magna (by vol.) (ratio) |  | Toxicity; | Toxicity; EMS code: DMGC | Chemical |  | LAB
DMGT | h | Bio. LT50:Daphnia Magna | X541 | Daphn.Magna:mortality/100%conc. | This method measures the time (in hours) to 50% mortality in 100% concentration. | 0 | Y | Y | Y | 8 | 6/30/2000 8:48:00 AM | 2/14/2002 10:57:00 AM | 3/28/2002 1:10:56 AM | 2000-04-27 |  | NPEPPIN |  | Time | hr | Y | (time) |  | Bio. LT50:Daphnia Magna (time) |  | Toxicity; | Toxicity; EMS code: DMGT | Chemical |  | LAB
DMPF | % (Mortality) | Bioassay(Pass/Fail):Daphnia Magna | FA10 | Daph.Magna/48hr(%mortality)/100% conc. | Daph.Magna/48 hr(%mortality)/100% concentration |  | Y | Y | Y | 4056 |  |  | 7/29/2024 4:35:58 PM | 2015-06-24 |  | RDGIBSON |  | DimensionlessRatio | % | N | (ratio) | (mortality) | Bioassay(Pass/Fail):Daphnia Magna (mortality) (ratio) |  | Toxicity; | Toxicity; EMS code: DMPF | Chemical |  | LAB
DMSE | mg/L | Dimethyl Selenide [DMSe]; as Se | X564 | IC-ICP-CRC-MS | Ion chromatography inductively coupled plasma collision reaction cell mass spectrometry | 1e-05 | Y | Y | Y | 121 |  |  | 11/15/2022 11:41:12 AM | 2022-11-01 |  | EMS |  | Concentration | mg/L | Y | (fl. conc.) |  | Dimethyl Selenide [DMSe]; as Se (fl. conc.) | 593-79-3 | Organometallics; | Organometallics; EMS code: DMSE | Chemical |  | LAB
DMSE | mg/L | Dimethyl Selenide [DMSe]; as Se | E611 | E611S: EPA 8260D (mod) | E611S: Dimethyl selenide and dimethyl diselenide are analyzed by static headspace GC-MS. Samples are prepared in headspace vials and are heated and agitated on the headspace autosampler; causing VOCs to partition between the aqueous phase and the headspace in accordance with Henry`s law. | 0.0005 | Y | Y | Y | 178 | 12/9/2021 | 8/10/2024 | 8/19/2024 12:19:03 PM | 2020-04-21 |  | RDGIBSON |  | Concentration | mg/L | Y | (fl. conc.) |  | Dimethyl Selenide [DMSe]; as Se (fl. conc.) | 593-79-3 | Organometallics; | Organometallics; EMS code: DMSE | Chemical |  | LAB
DMSF | % | Dimethyl Sulfone | SURR | Surrogate Recovery | Surrogate Recovery | 1 | Y | Y | Y | 16 | 6/5/2018 | 9/4/2018 | 12/10/2019 3:45:08 PM | 2018-09-17 | 9/17/2018 3:54:55 PM | RDGIBSON | EMS | DimensionlessRatio | % | N | (ratio) |  | Dimethyl Sulfone (ratio) | 67-71-0 | Organics; | Organics; EMS code: DMSF | Chemical |  | SURROGATE_RESULT
DMSO | mg/L | Dimethylseleneoxide | X564 | IC-ICP-CRC-MS | Ion chromatography inductively coupled plasma collision reaction cell mass spectrometry | 5e-06 | Y | Y | Y | 6327 |  |  | 7/30/2024 12:34:59 PM | 2017-03-23 |  | RRENTON |  | Concentration | mg/L | Y | (fl. conc.) |  | Dimethylseleneoxide (fl. conc.) | 4371-90-8 | Organometallics; | Organometallics; EMS code: DMSO | Chemical |  | LAB
DO-F | mg/L | Dissolved Oxygen-Field | FLD | Field sampling | Field sampling method | 0.01 | Y | Y | Y | 7122 | 2/16/2016 3:00:00 PM | 8/22/2024 | 8/26/2024 4:51:04 PM | 2019-07-10 | 6/30/2020 4:44:43 PM | KATFRASE | RDGIBSON | Concentration | mg/L | Y | (fl. conc.) |  | Dissolved Oxygen-Field (fl. conc.) | 7782-44-7 | Field; | Field; EMS code: DO-F | Chemical | Dissolved | FIELD_RESULT | Field sampling
DO-F | mg/L | Dissolved Oxygen-Field | XM01 | DO Meter | Following the manufacturer's instructions, the probe with oxygen permeable membrane is calibrated against air or samples of known DO concentration (e.g. from Winkler titrations). The samples are then measured for DO, again following all precautions recommended by the manufacturer to insure acceptable results. | 0 | Y | Y | Y | 76433 | 8/8/2007 12:50:00 PM | 7/26/2024 3:58:00 PM | 8/20/2024 7:35:41 AM | 2012-10-12 | 8/12/2020 9:04:53 AM | RDGIBSON | RDGIBSON | Concentration | mg/L | Y | (fl. conc.) |  | Dissolved Oxygen-Field (fl. conc.) | 7782-44-7 | Field; | Field; EMS code: DO-F | Chemical | Dissolved | FIELD_RESULT | DO Meter
DO-F | % Saturation | Dissolved Oxygen-Field | DOFS | Field Dissolved Oxygen | Following the manufacturer's instructions, the probe with oxygen permeable membrane is calibrated against air or samples of known DO concentration (e.g. from Winkler titrations). The samples are then measured for DO, again following all precautions recommended by the manufacturer to insure acceptable results. | 1 | Y | Y | Y | 346 | 8/22/2017 1:48:00 PM | 10/24/2023 9:35:00 AM | 7/5/2024 3:39:54 PM | 2018-01-23 | 8/12/2020 9:04:36 AM | JKROGH | RDGIBSON | DimensionlessRatio | % | N | (ratio) | (saturation) | Dissolved Oxygen-Field (saturation) (ratio) | 7782-44-7 |  | EMS code: DO-F | Chemical | Dissolved | FIELD_RESULT | Field Dissolved Oxygen
DOCS | % | Dissolved Oxygen (%), Client Supplied | EXOSO | EXO Sonde YSI | EXO Sonde YSI |  | Y | Y | Y | 204 |  |  | 10/17/2023 2:53:35 PM | 2023-06-12 |  | EMS |  | DimensionlessRatio | % | N | (ratio) |  | Dissolved Oxygen (%), Client Supplied (ratio) | 7782-44-7 | Field; | Field; EMS code: DOCS | Chemical | Dissolved | FIELD_RESULT | EXO Sonde YSI
DOCS | % | Dissolved Oxygen (%), Client Supplied | DOFS | Field Dissolved Oxygen | Following the manufacturer's instructions, the probe with oxygen permeable membrane is calibrated against air or samples of known DO concentration (e.g. from Winkler titrations). The samples are then measured for DO, again following all precautions recommended by the manufacturer to insure acceptable results. | 1 | Y | Y | Y | 328 | 7/26/2019 | 5/15/2024 3:33:00 PM | 6/21/2024 2:59:14 PM | 2019-09-06 |  | KATFRASE |  | DimensionlessRatio | % | N | (ratio) |  | Dissolved Oxygen (%), Client Supplied (ratio) | 7782-44-7 | Field; | Field; EMS code: DOCS | Chemical | Dissolved | FIELD_RESULT | Field Dissolved Oxygen
DOSA-F | % | Oxygen, dissolved saturation %, field | FLD | Field sampling | Field sampling method | 0.1 | Y | Y | Y | 234 | 8/12/2020 | 8/16/2024 | 8/23/2024 9:50:35 AM | 2023-01-30 |  | EMS |  | DimensionlessRatio | % | N | (ratio) |  | Oxygen, dissolved saturation %, field (ratio) | 7782-44-7 | Field; | Field; EMS code: DOSA-F | Chemical | Dissolved | FIELD_RESULT | Field sampling
DPAM | mg/L | Diphenylamine | ZN01 | DCM+Ac.Extr:GC/MS | DCM+Ac.Extr:GC/MS | 0.0001 | Y | Y | Y | 28 | 7/28/2005 7:03:00 AM | 3/29/2012 10:15:00 PM | 5/3/2012 9:27:50 AM | 1996-04-12 | 11/16/1999 4:06:27 PM | EMS | NPEPPIN | Concentration | mg/L | Y | (fl. conc.) |  | Diphenylamine (fl. conc.) |  | Pesticides; | Pesticides; EMS code: DPAM | Chemical |  | LAB
DPHM | mg/L | Diphenhydramine CAS: 58-73-1 | XMSP | HPLC and Tri. Quadpole MS/MS | EPA 1694 An aliquot of filtered sample containing 5% organic is injected directly. Each compound is separated by reversed phase HPLC. Analyzed by Electro Spray Ionization MS/MS detection using Triple Quadrupole MS/MS detector. Quantiitation with internal or external standard. |  | Y | Y | Y | 12 | 5/8/2020 | 6/4/2024 | 8/21/2024 3:32:03 PM | 2022-05-04 |  | EMS |  | Concentration | mg/L | Y | (fl. conc.) |  | Diphenhydramine (fl. conc.) | 58-73-1 | Pesticides; | Pesticides; EMS code: DPHM | Chemical |  | LAB
DTZ | mg/L | Diltiazem CAS: 56209-45-1 | XMSP | HPLC and Tri. Quadpole MS/MS | EPA 1694 An aliquot of filtered sample containing 5% organic is injected directly. Each compound is separated by reversed phase HPLC. Analyzed by Electro Spray Ionization MS/MS detection using Triple Quadrupole MS/MS detector. Quantiitation with internal or external standard. |  | Y | Y | Y | 12 | 5/8/2020 | 6/4/2024 | 8/21/2024 3:32:03 PM | 2022-05-04 |  | EMS |  | Concentration | mg/L | Y | (fl. conc.) |  | Diltiazem (fl. conc.) | 56209-45-1 | Pesticides; | Pesticides; EMS code: DTZ | Chemical |  | LAB
DY-T | ppm (S) | Dysprosium | MS81 | ALS Minerals: ME-MS81 | ALS Minerals: ME-MS81. A prepared sample (0.100 g) is added to lithium metaborate/lithium tetraborate flux; mixed well and fused in a furnace at 1025°C. The resulting melt is then cooled and dissolved in an acid mixture containing nitric; hydrochloric and hydrofluoric acids. This solution is then analyzed by inductively coupled plasma - mass spectrometry. | 0.05 | Y | Y | Y | 0 |  |  |  | 2024-01-29 |  | EMS |  | DimensionlessRatio | ppm | N | (ratio) |  | Dysprosium (ratio) |  | Metals; | Metals; EMS code: DY-T | Chemical |  | LAB
DZP | mg/L | Diazepam CAS: 439-14-5 | XMSP | HPLC and Tri. Quadpole MS/MS | EPA 1694 An aliquot of filtered sample containing 5% organic is injected directly. Each compound is separated by reversed phase HPLC. Analyzed by Electro Spray Ionization MS/MS detection using Triple Quadrupole MS/MS detector. Quantiitation with internal or external standard. |  | Y | Y | Y | 12 | 5/8/2020 | 6/4/2024 | 8/21/2024 3:32:03 PM | 2022-05-04 |  | EMS |  | Concentration | mg/L | Y | (fl. conc.) |  | Diazepam (fl. conc.) | 439-14-5 | Pesticides; | Pesticides; EMS code: DZP | Chemical |  | LAB
E001 | ug/g | Ethyl Alcohol | E8015S | EPA 8015 D in soil | See EPA 8015D | 10 | Y | Y | Y | 15 | 1/12/2022 | 2/1/2022 | 1/11/2024 1:07:07 PM | 2022-06-23 |  | EMS |  | MassPerMass | mg/kg | Y | (mass conc.) |  | Ethyl Alcohol (mass conc.) | 64-17-5 | Volatile Organic Compounds (VOCs); | Volatile Organic Compounds (VOCs); EMS code: E001 | Chemical |  | LAB
E001 | mg/L | Ethyl Alcohol | E-8015 | EPA 8015D | EPA Method 8015D (SW-846): Nonhalogenated Organics Using GC/FID |  | Y | Y | Y | 14 | 12/29/2021 | 1/12/2022 | 9/2/2022 6:21:53 PM | 2022-04-05 |  | EMS |  | Concentration | mg/L | Y | (fl. conc.) |  | Ethyl Alcohol (fl. conc.) | 64-17-5 | Volatile Organic Compounds (VOCs); | Volatile Organic Compounds (VOCs); EMS code: E001 | Chemical |  | LAB
E006 | ug/g | Endrin Ketone | ZN03 | DCM+Ac.Extr:GC/MS | DCM+Ac.Extr:GC/MS EPA 8270E (mod) | 0.005 | Y | Y | Y | 2 | 11/2/2022 12:53:00 AM | 11/2/2022 4:41:00 AM | 1/7/2023 3:12:10 PM | 2022-12-23 |  | EMS |  | MassPerMass | mg/kg | Y | (mass conc.) |  | Endrin Ketone (mass conc.) | 53494-70-5 | Pesticides; | Pesticides; EMS code: E006 | Chemical |  | LAB
E007 | mg/L | Endrin | F027 | Hexane Extr,GC/ECD | Effluent samples are solvent extracted with Hexane and the extract dried over Sodium Sulfate, concentrated and cleaned up, as required, using Florisil, activated Copper metal, concentrated Sulfuric Acid and Mercury, to remove interferences. The extracts are then analysed using megabore column Gas Chromatography with Electron Capture detection. | 5e-05 | Y | Y | Y | 20 | 3/3/2000 7:36:00 AM | 1/22/2002 10:58:00 AM | 6/18/2003 2:53:41 PM | 2002-06-05 |  | UMULLER |  | Concentration | mg/L | Y | (fl. conc.) |  | Endrin (fl. conc.) | 72-20-8 | Pesticides; | Pesticides; EMS code: E007 | Chemical |  | LAB
E007 | ug/g | Endrin | X213 | Hex/AcExt,Part,Chrom,GC | Hex/AcExt;Part;Chrom;GC | 0.0005 | Y | Y | Y | 187 | 4/20/2002 10:10:00 AM | 10/13/2002 10:30:00 AM | 2/26/2004 1:05:28 AM | 1996-04-12 | 6/3/1998 6:51:36 PM | EMS |  | MassPerMass | mg/kg | Y | (mass conc.) |  | Endrin (mass conc.) | 72-20-8 | Pesticides; | Pesticides; EMS code: E007 | Chemical |  | LAB
E007 | mg/L | Endrin | E660 | EPA 8270E (mod) | EPA 8270E (mod) Miscellaneous Pesticides by GC-MS | 5e-06 | Y | Y | Y | 2 | 9/14/2022 | 9/14/2022 | 12/8/2022 9:48:14 AM | 2022-12-06 |  | EMS |  | Concentration | mg/L | Y | (fl. conc.) |  | Endrin (fl. conc.) | 72-20-8 | Pesticides; | Pesticides; EMS code: E007 | Chemical |  | LAB
E007 | mg/L | Endrin | X268 | CH2Cl2 ext:GC/MS | CH2Cl2 ext:GC/MS |  | Y | Y | Y | 16 | 12/29/2021 | 9/2/2022 | 11/10/2022 9:54:07 AM | 2022-04-08 |  | EMS |  | Concentration | mg/L | Y | (fl. conc.) |  | Endrin (fl. conc.) | 72-20-8 | Pesticides; | Pesticides; EMS code: E007 | Chemical |  | LAB
E007 | mg/L | Endrin | MW03 | OCP1-LL-SF-ECD-VA | This analysis is carried out using procedures adapted from "Test Methods for Evaluating Solid Waste" SW-846, Methods 3510, 3630, 3660 and 8081, published by the United States Environmental Protection Agency (EPA). The procedure involves extraction of the entire water sample with dichloromethane. The extract then undergoes one or more of the following clean-up procedures (if required): silica gel clean-up and/or sulphur clean-up. The final extract is analyzed by dual capillary column gas chromatography with electron capture detection (GC/ECD). | 2e-05 | Y | Y | Y | 10 | 9/20/2016 | 9/21/2016 | 8/15/2019 2:42:34 PM | 2019-08-15 |  | KATFRASE |  | Concentration | mg/L | Y | (fl. conc.) |  | Endrin (fl. conc.) | 72-20-8 | Pesticides; | Pesticides; EMS code: E007 | Chemical |  | LAB
E007 | mg/L | Endrin | ZN01 | DCM+Ac.Extr:GC/MS | DCM+Ac.Extr:GC/MS | 0.0005 | Y | Y | Y | 52 | 10/3/2002 | 3/29/2012 10:15:00 PM | 5/3/2012 9:27:50 AM | 1996-04-12 | 11/16/1999 4:06:27 PM | EMS | NPEPPIN | Concentration | mg/L | Y | (fl. conc.) |  | Endrin (fl. conc.) | 72-20-8 | Pesticides; | Pesticides; EMS code: E007 | Chemical |  | LAB
E007 | ug/g wet | Endrin | X220 | Hex/AcExt,Part,Chrom,GC | Hex/AcExt;Part;Chrom;GC | 0.0002 | Y | Y | Y | 71 |  |  | 10/24/1996 10:01:25 AM | 1996-04-12 | 6/3/1998 6:51:42 PM | EMS |  | MassPerMass | µg/g | Y | (mass conc.) | (wet) | Endrin (wet) (mass conc.) | 72-20-8 | Pesticides; | Pesticides; EMS code: E007 | Chemical |  | LAB
E007 | mg/L | Endrin | EQ01 | EQUIS (mg/L) | EQUIS (mg/L) | 0.0008 | N | Y | Y | 16 |  |  | 10/22/1996 11:26:54 PM | 1996-04-12 | 6/3/1998 6:51:36 PM | EMS |  | Concentration | mg/L | Y | (fl. conc.) |  | Endrin (fl. conc.) | 72-20-8 | Pesticides; | Pesticides; EMS code: E007 | Chemical |  | LAB
E007 | ug/g | Endrin | E8081 | EPA 8081 | See methods EPA 3540, 3630, 3660, 8081 | 0.005 | Y | Y | Y | 9 | 3/18/2020 | 3/23/2020 | 2/22/2023 12:51:03 PM | 2022-07-29 |  | EMS |  | MassPerMass | mg/kg | Y | (mass conc.) |  | Endrin (mass conc.) | 72-20-8 | Pesticides; | Pesticides; EMS code: E007 | Chemical |  | LAB
E007 | ug/g | Endrin | ZN03 | DCM+Ac.Extr:GC/MS | DCM+Ac.Extr:GC/MS EPA 8270E (mod) | 0.05 | Y | Y | Y | 19 | 1/17/2022 | 11/14/2022 | 1/11/2024 1:07:07 PM | 1996-04-12 | 11/16/1999 2:27:40 PM | EMS | NPEPPIN | MassPerMass | mg/kg | Y | (mass conc.) |  | Endrin (mass conc.) | 72-20-8 | Pesticides; | Pesticides; EMS code: E007 | Chemical |  | LAB
E008 | ug/g | Endrin Aldehyde | ZN03 | DCM+Ac.Extr:GC/MS | DCM+Ac.Extr:GC/MS EPA 8270E (mod) | 0.01 | Y | Y | Y | 4 | 11/2/2022 12:53:00 AM | 11/14/2022 | 1/7/2023 3:12:10 PM | 2022-06-23 |  | EMS |  | MassPerMass | mg/kg | Y | (mass conc.) |  | Endrin Aldehyde (mass conc.) |  | Pesticides; | Pesticides; EMS code: E008 | Chemical |  | LAB
E008 | mg/L | Endrin Aldehyde | E660 | EPA 8270E (mod) | EPA 8270E (mod) Miscellaneous Pesticides by GC-MS | 5e-06 | Y | Y | Y | 2 | 9/14/2022 | 9/14/2022 | 12/8/2022 9:48:14 AM | 2022-12-06 |  | EMS |  | Concentration | mg/L | Y | (fl. conc.) |  | Endrin Aldehyde (fl. conc.) |  | Pesticides; | Pesticides; EMS code: E008 | Chemical |  | LAB
E008 | mg/L | Endrin Aldehyde | X268 | CH2Cl2 ext:GC/MS | CH2Cl2 ext:GC/MS |  | Y | Y | Y | 16 | 12/29/2021 | 9/2/2022 | 11/10/2022 9:54:07 AM | 2022-04-08 |  | EMS |  | Concentration | mg/L | Y | (fl. conc.) |  | Endrin Aldehyde (fl. conc.) |  | Pesticides; | Pesticides; EMS code: E008 | Chemical |  | LAB
E009 | ug/g | Ethion | ZN03 | DCM+Ac.Extr:GC/MS | DCM+Ac.Extr:GC/MS EPA 8270E (mod) | 0.02 | Y | Y | Y | 38 | 4/20/2002 10:25:00 AM | 10/16/2002 3:20:00 PM | 2/26/2004 1:05:28 AM | 1996-04-12 | 11/16/1999 2:27:40 PM | EMS | NPEPPIN | MassPerMass | mg/kg | Y | (mass conc.) |  | Ethion (mass conc.) | 563-12-2 | Pesticides; | Pesticides; EMS code: E009 | Chemical |  | LAB
E009 | mg/L | Ethion | ZN01 | DCM+Ac.Extr:GC/MS | DCM+Ac.Extr:GC/MS | 0.0002 | Y | Y | Y | 52 | 10/3/2002 | 3/29/2012 10:15:00 PM | 5/3/2012 9:27:50 AM | 1996-04-12 | 11/16/1999 4:06:27 PM | EMS | NPEPPIN | Concentration | mg/L | Y | (fl. conc.) |  | Ethion (fl. conc.) | 563-12-2 | Pesticides; | Pesticides; EMS code: E009 | Chemical |  | LAB
E009 | mg/L | Ethion | F053 | DCM extr'n,conc'n,HGRC-PFPD/N,P det'n | Effluent samples are extracted with Dichloromethane. The extract is then concentrated and analysed using High Resolution Gas Chromatography with Pulsed Flame Photometric Dectection. Confirmation, if required, is by High Resolution Gas Chromatography with Nitrogen Phosphorus Detection. | 5e-05 | Y | Y | Y | 31 | 3/3/2000 7:36:00 AM | 1/22/2002 10:26:00 AM | 6/18/2003 2:53:42 PM | 1997-10-20 | 6/3/1998 6:51:47 PM | NPEPPIN |  | Concentration | mg/L | Y | (fl. conc.) |  | Ethion (fl. conc.) | 563-12-2 | Pesticides; | Pesticides; EMS code: E009 | Chemical |  | LAB
E009 | ug/g | Ethion | FS16 | DCM/MEOH extr'n,DCM,conc'n,HRGC-PFPD/N,P det'n | Sediment samples are shaken with a mixture of Dichloromethane and Methanol, filtered, then extracted with Dichloromethane. The extract is then concentrated and analysed using High Resolution Gas Chromatography with Pulsed Flame Photometric Detection. Confirmation, if required, is by High Resolution Gas Chromatography with Nitrogen Phosphorus Detection. | 0.05 | Y | Y | Y | 1 | 4/24/2002 4:00:00 PM | 4/24/2002 4:00:00 PM | 2/12/2004 1:02:35 AM | 1997-10-20 | 6/3/1998 6:51:47 PM | NPEPPIN |  | MassPerMass | mg/kg | Y | (mass conc.) |  | Ethion (mass conc.) | 563-12-2 | Pesticides; | Pesticides; EMS code: E009 | Chemical |  | LAB
E012 | mg/L | EPN | ZN01 | DCM+Ac.Extr:GC/MS | DCM+Ac.Extr:GC/MS | 0.0005 | Y | Y | Y | 28 | 7/29/2005 7:10:00 AM | 3/29/2012 10:15:00 PM | 5/3/2012 9:27:50 AM | 1996-04-12 | 11/16/1999 4:06:27 PM | EMS | NPEPPIN | Concentration | mg/L | Y | (fl. conc.) |  | EPN (fl. conc.) | 2104-64-5 | Pesticides; | Pesticides; EMS code: E012 | Chemical |  | LAB
E020 | mg/L | Ethylene Glycol | GLYC | Glycols in Soil and Water by GC/FID | GLYCOLS IN SOIL AND WATER PURPOSE/PRINCIPLE OF METHOD: Soils are extracted with water and liquids are injected directly or diluted with water. The final extract is analyzed by direct injection gas chromatography with a flame ionization detector (GC/FID) due to the high solubility of glycols. SCOPE: Glycols are commonly used for many applications such as a de-icing agent in aircraft and runway, as an anti-freeze in heating and cooling systems, as an industrial solvent, and in paint and plastics. The end data is used for various purposes including contamination assessment, environmental regulatory compliance and land reclamation. REFERENCES: Nonhalogenated Organics Using GC/FID. US EPA. Test Methods for Evaluating Solid Waste, Physical/Chemical Methods -- SW-846, Method 8015B Revision 2, December 1996 | 10 | Y | Y | Y | 23 | 10/7/2013 10:15:00 PM | 9/8/2018 | 12/10/2019 3:45:08 PM | 2012-05-30 | 6/4/2012 12:03:55 PM | RDGIBSON | RDGIBSON | Concentration | mg/L | Y | (fl. conc.) |  | Ethylene Glycol (fl. conc.) | 107-21-1 | Organics; | Organics; EMS code: E020 | Chemical |  | LAB
E020 | mg/L | Ethylene Glycol | F026 | Acetone Dilution, HRGC-FID | Effluent samples are diluted with Acetone and analysed using direct injection into a High Resolutioin Gas Chromatograph using Flame Ionization Detection (HRGC-FID). | 0.05 | Y | Y | Y | 4 | 12/8/1998 12:28:00 PM | 3/27/2000 10:38:00 AM | 8/19/2000 2:23:06 AM | 1997-01-14 | 6/3/1998 6:51:46 PM | MPAK |  | Concentration | mg/L | Y | (fl. conc.) |  | Ethylene Glycol (fl. conc.) | 107-21-1 | Organics; | Organics; EMS code: E020 | Chemical |  | LAB
E040 | ug/g | Endosulfan I | X213 | Hex/AcExt,Part,Chrom,GC | Hex/AcExt;Part;Chrom;GC | 0.0005 | Y | Y | Y | 187 | 4/20/2002 10:10:00 AM | 10/13/2002 10:30:00 AM | 2/26/2004 1:05:28 AM | 1996-04-12 | 6/3/1998 6:51:42 PM | EMS |  | MassPerMass | mg/kg | Y | (mass conc.) |  | Endosulfan I (mass conc.) |  | Pesticides; | Pesticides; EMS code: E040 | Chemical |  | LAB
E040 | mg/L | Endosulfan I | X268 | CH2Cl2 ext:GC/MS | CH2Cl2 ext:GC/MS |  | Y | Y | Y | 16 | 12/29/2021 | 9/2/2022 | 11/10/2022 9:54:07 AM | 2022-04-08 |  | EMS |  | Concentration | mg/L | Y | (fl. conc.) |  | Endosulfan I (fl. conc.) |  | Pesticides; | Pesticides; EMS code: E040 | Chemical |  | LAB
E040 | mg/L | Endosulfan I | E660 | EPA 8270E (mod) | EPA 8270E (mod) Miscellaneous Pesticides by GC-MS | 1.4e-06 | Y | Y | Y | 2 | 9/14/2022 | 9/14/2022 | 12/8/2022 9:48:14 AM | 2022-12-06 |  | EMS |  | Concentration | mg/L | Y | (fl. conc.) |  | Endosulfan I (fl. conc.) |  | Pesticides; | Pesticides; EMS code: E040 | Chemical |  | LAB
E040 | ug/g wet | Endosulfan I | X220 | Hex/AcExt,Part,Chrom,GC | Hex/AcExt;Part;Chrom;GC | 0.0001 | Y | Y | Y | 73 |  |  | 10/24/1996 10:01:25 AM | 1996-04-12 | 6/3/1998 6:51:42 PM | EMS |  | MassPerMass | µg/g | Y | (mass conc.) | (wet) | Endosulfan I (wet) (mass conc.) |  | Pesticides; | Pesticides; EMS code: E040 | Chemical |  | LAB
E040 | mg/L | Endosulfan I | F027 | Hexane Extr,GC/ECD | Effluent samples are solvent extracted with Hexane and the extract dried over Sodium Sulfate, concentrated and cleaned up, as required, using Florisil, activated Copper metal, concentrated Sulfuric Acid and Mercury, to remove interferences. The extracts are then analysed using megabore column Gas Chromatography with Electron Capture detection. | 5e-05 | Y | Y | Y | 20 | 3/3/2000 7:36:00 AM | 1/22/2002 10:58:00 AM | 6/18/2003 2:53:42 PM | 2002-06-05 |  | UMULLER |  | Concentration | mg/L | Y | (fl. conc.) |  | Endosulfan I (fl. conc.) |  | Pesticides; | Pesticides; EMS code: E040 | Chemical |  | LAB
E040 | mg/L | Endosulfan I | MW03 | OCP1-LL-SF-ECD-VA | This analysis is carried out using procedures adapted from "Test Methods for Evaluating Solid Waste" SW-846, Methods 3510, 3630, 3660 and 8081, published by the United States Environmental Protection Agency (EPA). The procedure involves extraction of the entire water sample with dichloromethane. The extract then undergoes one or more of the following clean-up procedures (if required): silica gel clean-up and/or sulphur clean-up. The final extract is analyzed by dual capillary column gas chromatography with electron capture detection (GC/ECD). | 2e-06 | Y | Y | Y | 10 | 9/20/2016 | 9/21/2016 | 8/15/2019 2:42:34 PM | 2019-08-15 |  | KATFRASE |  | Concentration | mg/L | Y | (fl. conc.) |  | Endosulfan I (fl. conc.) |  | Pesticides; | Pesticides; EMS code: E040 | Chemical |  | LAB
E040 | ug/g | Endosulfan I | E8081 | EPA 8081 | See methods EPA 3540, 3630, 3660, 8081 | 0.001 | Y | Y | Y | 9 | 3/18/2020 | 3/23/2020 | 2/22/2023 12:51:03 PM | 2022-07-29 |  | EMS |  | MassPerMass | mg/kg | Y | (mass conc.) |  | Endosulfan I (mass conc.) |  | Pesticides; | Pesticides; EMS code: E040 | Chemical |  | LAB
E040 | mg/L | Endosulfan I | ZN01 | DCM+Ac.Extr:GC/MS | DCM+Ac.Extr:GC/MS | 0.0005 | Y | Y | Y | 52 | 10/3/2002 | 3/29/2012 10:15:00 PM | 5/3/2012 9:27:50 AM | 1996-04-12 | 11/16/1999 4:06:27 PM | EMS | NPEPPIN | Concentration | mg/L | Y | (fl. conc.) |  | Endosulfan I (fl. conc.) |  | Pesticides; | Pesticides; EMS code: E040 | Chemical |  | LAB
E040 | ug/g | Endosulfan I | ZN03 | DCM+Ac.Extr:GC/MS | DCM+Ac.Extr:GC/MS EPA 8270E (mod) | 0.05 | Y | Y | Y | 17 | 1/17/2022 | 11/14/2022 | 1/11/2024 1:07:07 PM | 1996-04-12 | 11/16/1999 2:27:40 PM | EMS | NPEPPIN | MassPerMass | mg/kg | Y | (mass conc.) |  | Endosulfan I (mass conc.) |  | Pesticides; | Pesticides; EMS code: E040 | Chemical |  | LAB
E041 | ug/g | Endosulfan II | E8081 | EPA 8081 | See methods EPA 3540, 3630, 3660, 8081 | 0.001 | Y | Y | Y | 9 | 3/18/2020 | 3/23/2020 | 2/22/2023 12:51:03 PM | 2022-07-29 |  | EMS |  | MassPerMass | mg/kg | Y | (mass conc.) |  | Endosulfan II (mass conc.) |  | Pesticides; | Pesticides; EMS code: E041 | Chemical |  | LAB
E041 | mg/L | Endosulfan II | X268 | CH2Cl2 ext:GC/MS | CH2Cl2 ext:GC/MS |  | Y | Y | Y | 16 | 12/29/2021 | 9/2/2022 | 11/10/2022 9:54:07 AM | 2022-04-08 |  | EMS |  | Concentration | mg/L | Y | (fl. conc.) |  | Endosulfan II (fl. conc.) |  | Pesticides; | Pesticides; EMS code: E041 | Chemical |  | LAB
E041 | mg/L | Endosulfan II | F027 | Hexane Extr,GC/ECD | Effluent samples are solvent extracted with Hexane and the extract dried over Sodium Sulfate, concentrated and cleaned up, as required, using Florisil, activated Copper metal, concentrated Sulfuric Acid and Mercury, to remove interferences. The extracts are then analysed using megabore column Gas Chromatography with Electron Capture detection. | 5e-05 | Y | Y | Y | 20 | 3/3/2000 7:36:00 AM | 1/22/2002 10:58:00 AM | 6/18/2003 2:53:42 PM | 2002-06-05 |  | UMULLER |  | Concentration | mg/L | Y | (fl. conc.) |  | Endosulfan II (fl. conc.) |  | Pesticides; | Pesticides; EMS code: E041 | Chemical |  | LAB
E041 | mg/L | Endosulfan II | E660 | EPA 8270E (mod) | EPA 8270E (mod) Miscellaneous Pesticides by GC-MS | 1.4e-06 | Y | Y | Y | 2 | 9/14/2022 | 9/14/2022 | 12/8/2022 9:48:14 AM | 2022-12-06 |  | EMS |  | Concentration | mg/L | Y | (fl. conc.) |  | Endosulfan II (fl. conc.) |  | Pesticides; | Pesticides; EMS code: E041 | Chemical |  | LAB
E041 | ug/g | Endosulfan II | X213 | Hex/AcExt,Part,Chrom,GC | Hex/AcExt;Part;Chrom;GC | 0.001 | Y | Y | Y | 187 | 4/20/2002 10:10:00 AM | 10/13/2002 10:30:00 AM | 2/26/2004 1:05:28 AM | 1996-04-12 | 6/3/1998 6:51:45 PM | EMS |  | MassPerMass | mg/kg | Y | (mass conc.) |  | Endosulfan II (mass conc.) |  | Pesticides; | Pesticides; EMS code: E041 | Chemical |  | LAB
E041 | ug/g wet | Endosulfan II | X220 | Hex/AcExt,Part,Chrom,GC | Hex/AcExt;Part;Chrom;GC | 0.0004 | Y | Y | Y | 73 |  |  | 10/24/1996 10:01:25 AM | 1996-04-12 | 6/3/1998 6:51:42 PM | EMS |  | MassPerMass | µg/g | Y | (mass conc.) | (wet) | Endosulfan II (wet) (mass conc.) |  | Pesticides; | Pesticides; EMS code: E041 | Chemical |  | LAB
E041 | mg/L | Endosulfan II | MW03 | OCP1-LL-SF-ECD-VA | This analysis is carried out using procedures adapted from "Test Methods for Evaluating Solid Waste" SW-846, Methods 3510, 3630, 3660 and 8081, published by the United States Environmental Protection Agency (EPA). The procedure involves extraction of the entire water sample with dichloromethane. The extract then undergoes one or more of the following clean-up procedures (if required): silica gel clean-up and/or sulphur clean-up. The final extract is analyzed by dual capillary column gas chromatography with electron capture detection (GC/ECD). | 2e-06 | Y | Y | Y | 10 | 9/20/2016 | 9/21/2016 | 8/15/2019 2:42:34 PM | 2019-08-15 |  | KATFRASE |  | Concentration | mg/L | Y | (fl. conc.) |  | Endosulfan II (fl. conc.) |  | Pesticides; | Pesticides; EMS code: E041 | Chemical |  | LAB
E041 | ug/g | Endosulfan II | ZN03 | DCM+Ac.Extr:GC/MS | DCM+Ac.Extr:GC/MS EPA 8270E (mod) | 0.05 | Y | Y | Y | 17 | 1/17/2022 | 11/14/2022 | 1/11/2024 1:07:07 PM | 1996-04-12 | 11/16/1999 2:27:40 PM | EMS | NPEPPIN | MassPerMass | mg/kg | Y | (mass conc.) |  | Endosulfan II (mass conc.) |  | Pesticides; | Pesticides; EMS code: E041 | Chemical |  | LAB
E041 | mg/L | Endosulfan II | ZN01 | DCM+Ac.Extr:GC/MS | DCM+Ac.Extr:GC/MS | 0.0005 | Y | Y | Y | 52 | 10/3/2002 | 3/29/2012 10:15:00 PM | 5/3/2012 9:27:50 AM | 1996-04-12 | 11/16/1999 4:06:27 PM | EMS | NPEPPIN | Concentration | mg/L | Y | (fl. conc.) |  | Endosulfan II (fl. conc.) |  | Pesticides; | Pesticides; EMS code: E041 | Chemical |  | LAB
E042 | mg/L | Endosulfan Sulphate | X268 | CH2Cl2 ext:GC/MS | CH2Cl2 ext:GC/MS |  | Y | Y | Y | 16 | 12/29/2021 | 9/2/2022 | 11/10/2022 9:54:07 AM | 2022-04-08 |  | EMS |  | Concentration | mg/L | Y | (fl. conc.) |  | Endosulfan Sulphate (fl. conc.) |  | Pesticides; | Pesticides; EMS code: E042 | Chemical |  | LAB
E042 | mg/L | Endosulfan Sulphate | E660 | EPA 8270E (mod) | EPA 8270E (mod) Miscellaneous Pesticides by GC-MS | 2e-06 | Y | Y | Y | 2 | 9/14/2022 | 9/14/2022 | 12/8/2022 9:48:14 AM | 2022-12-06 |  | EMS |  | Concentration | mg/L | Y | (fl. conc.) |  | Endosulfan Sulphate (fl. conc.) |  | Pesticides; | Pesticides; EMS code: E042 | Chemical |  | LAB
E042 | ug/g | Endosulfan Sulphate | ZN03 | DCM+Ac.Extr:GC/MS | DCM+Ac.Extr:GC/MS EPA 8270E (mod) | 0.05 | Y | Y | Y | 4 | 11/2/2022 12:53:00 AM | 11/14/2022 | 1/7/2023 3:12:10 PM | 1996-04-12 | 11/16/1999 2:27:40 PM | EMS | NPEPPIN | MassPerMass | mg/kg | Y | (mass conc.) |  | Endosulfan Sulphate (mass conc.) |  | Pesticides; | Pesticides; EMS code: E042 | Chemical |  | LAB
E042 | ug/g wet | Endosulfan Sulphate | X220 | Hex/AcExt,Part,Chrom,GC | Hex/AcExt;Part;Chrom;GC | 0.002 | Y | Y | Y | 73 |  |  | 10/24/1996 10:01:25 AM | 1996-04-12 | 6/3/1998 6:51:42 PM | EMS |  | MassPerMass | µg/g | Y | (mass conc.) | (wet) | Endosulfan Sulphate (wet) (mass conc.) |  | Pesticides; | Pesticides; EMS code: E042 | Chemical |  | LAB
E042 | ug/g | Endosulfan Sulphate | X213 | Hex/AcExt,Part,Chrom,GC | Hex/AcExt;Part;Chrom;GC | 0.01 | Y | Y | Y | 187 | 4/20/2002 10:10:00 AM | 10/13/2002 10:30:00 AM | 2/26/2004 1:05:28 AM | 1996-04-12 | 6/3/1998 6:51:36 PM | EMS |  | MassPerMass | mg/kg | Y | (mass conc.) |  | Endosulfan Sulphate (mass conc.) |  | Pesticides; | Pesticides; EMS code: E042 | Chemical |  | LAB
E042 | mg/L | Endosulfan Sulphate | F027 | Hexane Extr,GC/ECD | Effluent samples are solvent extracted with Hexane and the extract dried over Sodium Sulfate, concentrated and cleaned up, as required, using Florisil, activated Copper metal, concentrated Sulfuric Acid and Mercury, to remove interferences. The extracts are then analysed using megabore column Gas Chromatography with Electron Capture detection. | 5e-05 | Y | Y | Y | 20 | 3/3/2000 7:36:00 AM | 1/22/2002 10:58:00 AM | 6/18/2003 2:53:43 PM | 2002-06-05 |  | UMULLER |  | Concentration | mg/L | Y | (fl. conc.) |  | Endosulfan Sulphate (fl. conc.) |  | Pesticides; | Pesticides; EMS code: E042 | Chemical |  | LAB
E042 | ug/g | Endosulfan Sulphate | E8081 | EPA 8081 | See methods EPA 3540, 3630, 3660, 8081 | 0.001 | Y | Y | Y | 9 | 3/18/2020 | 3/23/2020 | 2/22/2023 12:51:03 PM | 2022-07-29 |  | EMS |  | MassPerMass | mg/kg | Y | (mass conc.) |  | Endosulfan Sulphate (mass conc.) |  | Pesticides; | Pesticides; EMS code: E042 | Chemical |  | LAB
E042 | mg/L | Endosulfan Sulphate | ZN01 | DCM+Ac.Extr:GC/MS | DCM+Ac.Extr:GC/MS | 0.0005 | Y | Y | Y | 52 | 10/3/2002 | 3/29/2012 10:15:00 PM | 5/3/2012 9:27:50 AM | 1996-04-12 | 11/16/1999 4:06:27 PM | EMS | NPEPPIN | Concentration | mg/L | Y | (fl. conc.) |  | Endosulfan Sulphate (fl. conc.) |  | Pesticides; | Pesticides; EMS code: E042 | Chemical |  | LAB
E042 | mg/L | Endosulfan Sulphate | MW03 | OCP1-LL-SF-ECD-VA | This analysis is carried out using procedures adapted from "Test Methods for Evaluating Solid Waste" SW-846, Methods 3510, 3630, 3660 and 8081, published by the United States Environmental Protection Agency (EPA). The procedure involves extraction of the entire water sample with dichloromethane. The extract then undergoes one or more of the following clean-up procedures (if required): silica gel clean-up and/or sulphur clean-up. The final extract is analyzed by dual capillary column gas chromatography with electron capture detection (GC/ECD). | 2e-06 | Y | Y | Y | 10 | 9/20/2016 | 9/21/2016 | 8/15/2019 2:42:34 PM | 2019-08-15 |  | KATFRASE |  | Concentration | mg/L | Y | (fl. conc.) |  | Endosulfan Sulphate (fl. conc.) |  | Pesticides; | Pesticides; EMS code: E042 | Chemical |  | LAB
E043 | mg/L | Endosulfan-Total | X268 | CH2Cl2 ext:GC/MS | CH2Cl2 ext:GC/MS |  | Y | Y | Y | 16 | 12/29/2021 | 9/2/2022 | 11/10/2022 9:54:07 AM | 2022-04-08 |  | EMS |  | Concentration | mg/L | Y | (fl. conc.) |  | Endosulfan-Total (fl. conc.) |  | Pesticides; | Pesticides; EMS code: E043 | Chemical | Total | LAB
E043 | ug/g | Endosulfan-Total | ZN03 | DCM+Ac.Extr:GC/MS | DCM+Ac.Extr:GC/MS EPA 8270E (mod) | 0.05 | Y | Y | Y | 19 | 1/17/2022 | 11/14/2022 | 1/11/2024 1:07:07 PM | 2003-06-16 |  | UMULLER |  | MassPerMass | mg/kg | Y | (mass conc.) |  | Endosulfan-Total (mass conc.) |  | Pesticides; | Pesticides; EMS code: E043 | Chemical | Total | LAB
E043 | mg/L | Endosulfan-Total | E660 | EPA 8270E (mod) | EPA 8270E (mod) Miscellaneous Pesticides by GC-MS | 2e-06 | Y | Y | Y | 2 | 9/14/2022 | 9/14/2022 | 12/8/2022 9:48:14 AM | 2022-12-06 |  | EMS |  | Concentration | mg/L | Y | (fl. conc.) |  | Endosulfan-Total (fl. conc.) |  | Pesticides; | Pesticides; EMS code: E043 | Chemical | Total | LAB
E046 | mg/L | Etrimfos | ZN01 | DCM+Ac.Extr:GC/MS | DCM+Ac.Extr:GC/MS | 0.0005 | Y | Y | Y | 1 | 7/29/2005 7:10:00 AM | 7/29/2005 7:10:00 AM | 6/2/2009 12:01:04 PM | 2003-06-16 |  | UMULLER |  | Concentration | mg/L | Y | (fl. conc.) |  | Etrimfos (fl. conc.) |  | Pesticides; | Pesticides; EMS code: E046 | Chemical |  | LAB
E100 | mg/L | Ethofumesate | ZN01 | DCM+Ac.Extr:GC/MS | DCM+Ac.Extr:GC/MS | 0.5 | Y | Y | Y | 1 | 7/29/2005 7:10:00 AM | 7/29/2005 7:10:00 AM | 6/2/2009 12:01:05 PM | 2009-02-20 |  | JCHUDAK |  | Concentration | mg/L | Y | (fl. conc.) |  | Ethofumesate (fl. conc.) |  | Pesticides; | Pesticides; EMS code: E100 | Chemical |  | LAB
E220 | mg/L | Diethylene Glycol | F026 | Acetone Dilution, HRGC-FID | Effluent samples are diluted with Acetone and analysed using direct injection into a High Resolutioin Gas Chromatograph using Flame Ionization Detection (HRGC-FID). | 0.1 | Y | Y | Y | 4 | 12/8/1998 12:28:00 PM | 3/27/2000 10:38:00 AM | 8/19/2000 2:23:06 AM | 1997-01-14 | 6/3/1998 6:51:46 PM | MPAK |  | Concentration | mg/L | Y | (fl. conc.) |  | Diethylene Glycol (fl. conc.) |  | Organics; | Organics; EMS code: E220 | Chemical |  | LAB
E220 | mg/L | Diethylene Glycol | GLYC | Glycols in Soil and Water by GC/FID | GLYCOLS IN SOIL AND WATER PURPOSE/PRINCIPLE OF METHOD: Soils are extracted with water and liquids are injected directly or diluted with water. The final extract is analyzed by direct injection gas chromatography with a flame ionization detector (GC/FID) due to the high solubility of glycols. SCOPE: Glycols are commonly used for many applications such as a de-icing agent in aircraft and runway, as an anti-freeze in heating and cooling systems, as an industrial solvent, and in paint and plastics. The end data is used for various purposes including contamination assessment, environmental regulatory compliance and land reclamation. REFERENCES: Nonhalogenated Organics Using GC/FID. US EPA. Test Methods for Evaluating Solid Waste, Physical/Chemical Methods -- SW-846, Method 8015B Revision 2, December 1996 | 10 | Y | Y | Y | 17 | 4/14/2015 10:15:00 PM | 9/8/2018 | 12/10/2019 3:45:08 PM | 2012-05-30 | 6/4/2012 12:03:55 PM | RDGIBSON | RDGIBSON | Concentration | mg/L | Y | (fl. conc.) |  | Diethylene Glycol (fl. conc.) |  | Organics; | Organics; EMS code: E220 | Chemical |  | LAB
E320 | mg/L | Triethylene Glycol | GLYC | Glycols in Soil and Water by GC/FID | GLYCOLS IN SOIL AND WATER PURPOSE/PRINCIPLE OF METHOD: Soils are extracted with water and liquids are injected directly or diluted with water. The final extract is analyzed by direct injection gas chromatography with a flame ionization detector (GC/FID) due to the high solubility of glycols. SCOPE: Glycols are commonly used for many applications such as a de-icing agent in aircraft and runway, as an anti-freeze in heating and cooling systems, as an industrial solvent, and in paint and plastics. The end data is used for various purposes including contamination assessment, environmental regulatory compliance and land reclamation. REFERENCES: Nonhalogenated Organics Using GC/FID. US EPA. Test Methods for Evaluating Solid Waste, Physical/Chemical Methods -- SW-846, Method 8015B Revision 2, December 1996 | 10 | Y | Y | Y | 17 | 4/14/2015 10:15:00 PM | 9/8/2018 | 12/10/2019 3:45:08 PM | 2012-05-30 | 6/4/2012 12:03:55 PM | RDGIBSON | RDGIBSON | Concentration | mg/L | Y | (fl. conc.) |  | Triethylene Glycol (fl. conc.) | 112-27-6 | Organics; | Organics; EMS code: E320 | Chemical |  | LAB
EA | ug/g | Ethyl Acetate CAS: 141-78-6 | MS02 | Purge&Trap GC/MS (VOC-S) | Purge&Trap GC/MS (VOC-S) | 0.05 | Y | Y | Y | 15 | 1/14/2022 | 2/4/2022 | 1/11/2024 1:07:07 PM | 2022-07-27 |  | EMS |  | MassPerMass | mg/kg | Y | (mass conc.) |  | Ethyl Acetate (mass conc.) | 141-78-6 | Organics; | Organics; EMS code: EA | Chemical |  | LAB
EA | mg/L | Ethyl Acetate CAS: 141-78-6 | MS04 | Static Headspace GC/MS (Water) | (Water) A static headspace procedure is used as a means of driving the analytes out of the matrix into a headspace. Once equilibrium is achieved, a fixed volume of headspeace is taken for subsequent determination by Gas Chromatography - Mass Spectroscopy (GC/MS) using Selected Ion Monitoring. |  | Y | Y | Y | 14 | 12/30/2021 | 1/12/2022 | 9/2/2022 6:21:53 PM | 2022-03-17 |  | EMS |  | Concentration | mg/L | Y | (fl. conc.) |  | Ethyl Acetate (fl. conc.) | 141-78-6 | Organics; | Organics; EMS code: EA | Chemical |  | LAB
EC-F | uS/cm | Electrical Conductivity - Field | COND | Conductivity APHA 2510B | Conductivity APHA 2510B. Conductance of a solution measured between two spatially fixed and chemically inert electrodes. |  | Y | Y | Y | 32 | 1/11/2022 | 8/16/2024 9:04:00 AM | 8/26/2024 8:06:03 AM | 2022-07-07 |  | EMS |  | ElectricConductivity | µS/cm | Y | (elec. cond.) |  | Electrical Conductivity - Field (elec. cond.) |  | Physical Properties; | Physical Properties; EMS code: EC-F | Physical |  | FIELD_RESULT | Conductivity APHA 2510B
EC-F | uS/cm | Electrical Conductivity - Field | EXOSO | EXO Sonde YSI | EXO Sonde YSI |  | Y | Y | Y | 111 |  |  | 8/7/2024 10:45:52 AM | 2023-06-12 |  | EMS |  | ElectricConductivity | µS/cm | Y | (elec. cond.) |  | Electrical Conductivity - Field (elec. cond.) |  | Physical Properties; | Physical Properties; EMS code: EC-F | Physical |  | FIELD_RESULT | EXO Sonde YSI
EC01 | mg/L | Eicosanoic Acid | RFAS | Separation of resin and fatty acids (RFAs) in water by GC/MS | Scientific Principle The separation of resin and fatty acids (RFAs) in water samples is performed using gas chromatography/mass spectrometry (GC/MS) in the Selected Ion Monitoring (SIM) mode. The separation is based on compound size; boiling point and polarity. Analyte concentration is calculated using the internal standard method. RFA compounds are extracted from water samples using a liquid-liquid solvent extraction with DCM. Extraction occurs after the sample is brought to a pH of ~ 2 and placed in a rotary tumbler (at approx. 40 rpm) for 1 hour. The resulting extract is transferred to a vial; dried with acidified baked sodium sulfate; methylated with diazomethane and concentrated to 1mL under nitrogen prior to analysis. The resulting extracts are vialed and fortified with labeled internal standard compounds prior to analysis by gas chromatography with a (50%-Phenyl)-methylpolysiloxane (DB-17) column and a mass spectrometer detector in SIM Mode. | 0.006 | Y | Y | Y | 1 | 9/17/2021 7:49:00 PM | 9/17/2021 7:49:00 PM | 2/1/2022 2:38:14 PM | 2022-01-19 |  | EMS |  | Concentration | mg/L | Y | (fl. conc.) |  | Eicosanoic Acid (fl. conc.) |  | Organics; | Organics; EMS code: EC01 | Chemical |  | LAB
EFOS1 | mg/L | Ethyl perfluorooctanesulfonamide, n- (NEtFOSA) CAS:4151-50-2 | E3533 | MECP E3533 | PFAs in water by direct injection LC/MS/MS | 2e-05 | Y | Y | Y | 2 | 6/3/2024 | 6/3/2024 | 8/21/2024 3:32:03 PM | 2023-11-15 | 6/27/2024 5:06:31 PM | EMS | EMS | Concentration | mg/L | Y | (fl. conc.) |  | Ethyl perfluorooctanesulfonamide, n- (NEtFOSA) (fl. conc.) | 4151-50-2 | Per- and polyfluoroalkyl substances (PFAS); | Per- and polyfluoroalkyl substances (PFAS); EMS code: EFOS1 | Chemical |  | LAB
EFOS2 | mg/L | Ethyl perfluorooctanesulfonamidoacetic acid, n- (NEtFOSAA) | E3533 | MECP E3533 | PFAs in water by direct injection LC/MS/MS | 2e-05 | Y | Y | Y | 2 | 6/3/2024 | 6/3/2024 | 8/21/2024 3:32:03 PM | 2023-11-15 | 6/27/2024 5:06:31 PM | EMS | EMS | Concentration | mg/L | Y | (fl. conc.) |  | Ethyl perfluorooctanesulfonamidoacetic acid, n- (NEtFOSAA) (fl. conc.) |  | Per- and polyfluoroalkyl substances (PFAS); | Per- and polyfluoroalkyl substances (PFAS); EMS code: EFOS2 | Chemical |  | LAB
EFOS3 | mg/L | Ethyl perfluorooctanesulfonamidoethanol, n- (NEtFOSE) | E3533 | MECP E3533 | PFAs in water by direct injection LC/MS/MS | 2e-05 | Y | Y | Y | 2 | 6/3/2024 | 6/3/2024 | 8/21/2024 3:32:03 PM | 2023-11-15 | 6/27/2024 5:06:31 PM | EMS | EMS | Concentration | mg/L | Y | (fl. conc.) |  | Ethyl perfluorooctanesulfonamidoethanol, n- (NEtFOSE) (fl. conc.) |  | Per- and polyfluoroalkyl substances (PFAS); | Per- and polyfluoroalkyl substances (PFAS); EMS code: EFOS3 | Chemical |  | LAB
EH1032 | mg/L | Extractable Petroleum Hydrocarbons C10 - C32 | F065 | EPH:DCM,liq-liq extr'n,GC/FID,(PAH not sub.) | Samples are extracted with dichloromethane (DCM) in a separatory funnel. Extracts are dried, concentrated and analyzed by capillary gas chromatography with flame ionization detector. Results are not corrected for specified PAH concentrations. LEPH is defined as the sum of components that elute on gc between nC10 and nC19; HEPH is defined as the sum of components that elute on gc between nC19 and nC32. Performance Base Method published circa Dec. 31, 2000, but distributed 1999. | 0.4 | Y | Y | Y | 15 | 2/16/2024 | 7/8/2024 | 7/9/2024 12:14:06 PM | 2022-05-09 |  | EMS |  | Concentration | mg/L | Y | (fl. conc.) |  | Extractable Petroleum Hydrocarbons C10 - C32 (fl. conc.) |  | Petroleum Hydrocarbons; | Petroleum Hydrocarbons; EMS code: EH1032 | Chemical | Extractable | LAB
EOX- | ug/g | ExtOrg.Halide Solids | DM12 | Extr.Org.Halides/Dohrman | Extr.Org.Halides/Dohrman | 10 | Y | Y | Y | 2 |  |  | 10/24/1996 12:41:35 PM | 1996-04-12 | 6/3/1998 6:51:36 PM | EMS |  | MassPerMass | mg/kg | Y | (mass conc.) |  | ExtOrg.Halide Solids (mass conc.) |  | Aggregate Organics; | Aggregate Organics; EMS code: EOX- | Chemical |  | LAB
EPH1 | mg/L | EPHw10-19 | E3511 | EPA 3511/BCMOE EPHw | EPA 3511/BCMOE EPHw; Hexane MicroExtraction (Base/Neutral) / Gas Chromatography (GC-FID). Sample extracts are subjected to column silica gel treatment prior to analysis by GC-FID for BC hydrocarbon fractions (EPHsg). | 0.25 | Y | Y | Y | 115 | 10/22/2022 2:30:00 PM | 5/10/2024 | 5/28/2024 1:16:02 PM | 2022-12-21 |  | EMS |  | Concentration | mg/L | Y | (fl. conc.) |  | EPHw10-19 (fl. conc.) |  | Petroleum Hydrocarbons; | Petroleum Hydrocarbons; EMS code: EPH1 | Chemical | Extractable | LAB
EPH2 | mg/L | EPHw19-32 | E3511 | EPA 3511/BCMOE EPHw | EPA 3511/BCMOE EPHw; Hexane MicroExtraction (Base/Neutral) / Gas Chromatography (GC-FID). Sample extracts are subjected to column silica gel treatment prior to analysis by GC-FID for BC hydrocarbon fractions (EPHsg). | 0.25 | Y | Y | Y | 115 | 10/22/2022 2:30:00 PM | 5/10/2024 | 5/28/2024 1:16:02 PM | 2022-12-21 |  | EMS |  | Concentration | mg/L | Y | (fl. conc.) |  | EPHw19-32 (fl. conc.) |  | Petroleum Hydrocarbons; | Petroleum Hydrocarbons; EMS code: EPH2 | Chemical | Extractable | LAB
EPH3 | ug/g | EPHs10-19 | E3570 | EPA 3570/BCMOE EPHs | EPA 3570/BCMOE EPHs; Shaker Extraction (Hexane-Acetone 1:1) / Gas Chromatography (GC-FID) | 50 | Y | Y | Y | 5 | 10/29/2022 8:51:00 AM | 10/29/2022 9:40:00 AM | 1/7/2023 3:12:10 PM | 2022-12-21 |  | EMS |  | MassPerMass | mg/kg | Y | (mass conc.) |  | EPHs10-19 (mass conc.) |  | Petroleum Hydrocarbons; | Petroleum Hydrocarbons; EMS code: EPH3 | Chemical |  | LAB
EPH4 | ug/g | EPHs19-32 | E3570 | EPA 3570/BCMOE EPHs | EPA 3570/BCMOE EPHs; Shaker Extraction (Hexane-Acetone 1:1) / Gas Chromatography (GC-FID) | 50 | Y | Y | Y | 5 | 10/29/2022 8:51:00 AM | 10/29/2022 9:40:00 AM | 1/7/2023 3:12:10 PM | 2022-12-21 |  | EMS |  | MassPerMass | mg/kg | Y | (mass conc.) |  | EPHs19-32 (mass conc.) |  | Petroleum Hydrocarbons; | Petroleum Hydrocarbons; EMS code: EPH4 | Chemical | Extractable | LAB
EPTM | mg/L | Eptam | ZN01 | DCM+Ac.Extr:GC/MS | DCM+Ac.Extr:GC/MS | 0.0005 | Y | Y | Y | 28 | 7/28/2005 7:03:00 AM | 3/29/2012 10:15:00 PM | 5/3/2012 9:27:51 AM | 1996-04-12 | 11/16/1999 4:06:27 PM | EMS | NPEPPIN | Concentration | mg/L | Y | (fl. conc.) |  | Eptam (fl. conc.) |  | Pesticides; | Pesticides; EMS code: EPTM | Chemical |  | LAB
ER-T | ppm (S) | Erbium | MS81 | ALS Minerals: ME-MS81 | ALS Minerals: ME-MS81. A prepared sample (0.100 g) is added to lithium metaborate/lithium tetraborate flux; mixed well and fused in a furnace at 1025°C. The resulting melt is then cooled and dissolved in an acid mixture containing nitric; hydrochloric and hydrofluoric acids. This solution is then analyzed by inductively coupled plasma - mass spectrometry. | 0.03 | Y | Y | Y | 0 |  |  |  | 2024-01-29 |  | EMS |  | DimensionlessRatio | ppm | N | (ratio) |  | Erbium (ratio) |  | Metals; | Metals; EMS code: ER-T | Chemical |  | LAB
ETFR | mg/L | Ethalfluralin | E660 | EPA 8270E (mod) | EPA 8270E (mod) Miscellaneous Pesticides by GC-MS | 0.0001 | Y | Y | Y | 4 | 9/12/2022 | 9/12/2022 | 12/8/2022 12:57:05 PM | 2022-12-06 |  | EMS |  | Concentration | mg/L | Y | (fl. conc.) |  | Ethalfluralin (fl. conc.) |  | Pesticides; | Pesticides; EMS code: ETFR | Chemical |  | LAB
ETFR | ug/g | Ethalfluralin | ZN03 | DCM+Ac.Extr:GC/MS | DCM+Ac.Extr:GC/MS EPA 8270E (mod) | 0.05 | Y | Y | Y | 2 | 9/8/2022 | 9/8/2022 | 12/15/2022 10:50:07 AM | 1996-04-12 | 11/16/1999 2:27:40 PM | EMS | NPEPPIN | MassPerMass | mg/kg | Y | (mass conc.) |  | Ethalfluralin (mass conc.) |  | Pesticides; | Pesticides; EMS code: ETFR | Chemical |  | LAB
ETFR | mg/L | Ethalfluralin | ZN01 | DCM+Ac.Extr:GC/MS | DCM+Ac.Extr:GC/MS | 0.0005 | Y | Y | Y | 28 | 7/28/2005 7:03:00 AM | 3/29/2012 10:15:00 PM | 5/3/2012 9:27:51 AM | 1996-04-12 | 11/16/1999 4:06:27 PM | EMS | NPEPPIN | Concentration | mg/L | Y | (fl. conc.) |  | Ethalfluralin (fl. conc.) |  | Pesticides; | Pesticides; EMS code: ETFR | Chemical |  | LAB
ETHA | mg/L | Ethane; Dissolved | MEED | Methane, Ethane and Ethene, Dissolved via GC-FID | Water samples are collected in 44 mL P&T vials without headspace and preserved with Sodium Bisulfate. Headspace is then generated using Nitrogen. The sample is then shaken to allow gas equilibration. A portion of the headspace vapour is withdrawn and injected into a gas chromatograph equipped with capillary columns and a flame ionization detector (FID). The concentration of the gas in the liquid is proportional to the partial pressure of the gas above the liquid, and is calculated using Henrys Law. Ref: EPA REGION 1, NATATTEN.WPD, REV. 1 | 0.005 | Y | Y | Y | 27 | 2/20/2014 | 8/1/2014 | 12/22/2014 5:16:41 PM | 2014-05-30 |  | RDGIBSON |  | Concentration | mg/L | Y | (fl. conc.) |  | Ethane; Dissolved (fl. conc.) |  | Volatile Organic Compounds (VOCs); | Volatile Organic Compounds (VOCs); EMS code: ETHA | Chemical | Dissolved | LAB
ETHE | mg/L | Ethene; Dissolved | MEED | Methane, Ethane and Ethene, Dissolved via GC-FID | Water samples are collected in 44 mL P&T vials without headspace and preserved with Sodium Bisulfate. Headspace is then generated using Nitrogen. The sample is then shaken to allow gas equilibration. A portion of the headspace vapour is withdrawn and injected into a gas chromatograph equipped with capillary columns and a flame ionization detector (FID). The concentration of the gas in the liquid is proportional to the partial pressure of the gas above the liquid, and is calculated using Henrys Law. Ref: EPA REGION 1, NATATTEN.WPD, REV. 1 | 0.005 | Y | Y | Y | 27 | 2/20/2014 | 8/1/2014 | 12/22/2014 5:16:41 PM | 2014-05-30 |  | RDGIBSON |  | Concentration | mg/L | Y | (fl. conc.) |  | Ethene; Dissolved (fl. conc.) |  | Volatile Organic Compounds (VOCs); | Volatile Organic Compounds (VOCs); EMS code: ETHE | Chemical | Dissolved | LAB
EU-T | ppm (S) | Europium | MS81 | ALS Minerals: ME-MS81 | ALS Minerals: ME-MS81. A prepared sample (0.100 g) is added to lithium metaborate/lithium tetraborate flux; mixed well and fused in a furnace at 1025°C. The resulting melt is then cooled and dissolved in an acid mixture containing nitric; hydrochloric and hydrofluoric acids. This solution is then analyzed by inductively coupled plasma - mass spectrometry. | 0.02 | Y | Y | Y | 0 |  |  |  | 2024-01-29 |  | EMS |  | DimensionlessRatio | ppm | N | (ratio) |  | Europium (ratio) |  | Metals; | Metals; EMS code: EU-T | Chemical |  | LAB
F--L | mg/L | Fluoride Leachable | MEND | MEND (mod) EPA 300.1 (mod) | MEND (mod) EPA 300.1 (mod). Inorganic anions are analyzed by obtaining an extract from a soil sample that has been added in a 3:1 ratio of deionized water to soil; then leached and the extract filtered prior to analysis; which is then analyzed by Ion Chromatography with conductivity and/or UV detection. | 0.02 | Y | Y | Y | 0 |  |  |  | 2024-01-24 |  | RDGIBSON |  | Concentration | mg/L | Y | (fl. conc.) |  | Fluoride Leachable (fl. conc.) |  |  | EMS code: F--L | Chemical |  | LAB
F--S | mg/L | Fluoride-Soluble | EQ01 | EQUIS (mg/L) | EQUIS (mg/L) | 0.01 | N | Y | Y | 38 |  |  | 10/22/1996 10:26:38 PM | 1996-04-12 | 6/3/1998 6:51:36 PM | EMS |  | Concentration | mg/L | Y | (fl. conc.) |  | Fluoride-Soluble (fl. conc.) |  |  | EMS code: F--S | Chemical | Soluble | LAB
F--S | mg/L | Fluoride-Soluble | 5068 | Ion Chromatography-Anion | (a) For the Soluble Particulates Pkg., a portion of filtrate from the Particulate-Total procedure made up to 500 mL is analyzed by Ion Chromatography. Intermediate results are in mg/L. Loading results are then calculated with units of mg/dm2/d. (b) For precipitation samples, samples and standards are spiked with concentrated carbonate/bicarbonate solution to give the sample the same background as the eluent used (avoiding a "water dip" effect). The anions of interest are separated through an anion "guard" and anion "separator" column. An anion micro membrane suppressor following the separator columns is used to reduce the background eluent conductivity by converting the carbonate and bicarbonate species to carbonic acid, while enhancing the conductivity of the ions of interest by converting these ions to their corresponding acids. A conductivity detector senses the same species in direct proportion to their initial concentration. (c) For HiVol samples, discs are cut from the HiVol filter, extracted with DI water, and anions analyzed by ion chromatograph, Preliminary results are reported in mg/L and intermediate HiVol loading results are reported as ug/m3. (*) Note that the air-convention is to report on basis of ion weight not element, e.g. as NO3 for nitrate, and as PO4 for phosphate. | 0.01 | Y | Y | Y | 151 | 12/13/2011 12:03:00 PM | 12/13/2011 12:03:00 PM | 2/9/2012 12:15:56 PM | 1996-04-12 | 6/3/1998 6:51:45 PM | EMS |  | Concentration | mg/L | Y | (fl. conc.) |  | Fluoride-Soluble (fl. conc.) |  |  | EMS code: F--S | Chemical | Soluble | LAB
F--S | mg/L | Fluoride-Soluble | X143 | Selective Ion Electrode | The analyte is determined potentiometrically by direct measurement using a selective ion electrode with standard reference electrode and either a meter with a direct concentration scale or a pH meter with an expanded mV scale. | 0.04 | Y | Y | Y | 654 |  |  | 1/18/1999 9:39:22 AM | 1996-04-12 | 6/3/1998 6:51:42 PM | EMS |  | Concentration | mg/L | Y | (fl. conc.) |  | Fluoride-Soluble (fl. conc.) |  |  | EMS code: F--S | Chemical | Soluble | LAB
F--S | ug/m3 | Fluoride-Soluble | EQ40 | EQUIS (mg/m3) | EQUIS (mg/m3) | 0.05 | N | Y | Y | 2857 |  |  | 10/22/1996 8:19:52 PM | 1996-04-12 | 6/3/1998 6:51:36 PM | EMS |  | AirConcentration | µg/m³ | Y | (air conc.) |  | Fluoride-Soluble (air conc.) |  |  | EMS code: F--S | Chemical | Soluble | LAB
F--S | ug/m3 | Fluoride-Soluble | X194 | Specific Ion Electr- Stack | Specific Ion Electr- Stack | 0.001 | Y | Y | Y | 419 |  |  | 3/9/1998 8:54:23 AM | 1996-04-12 | 6/3/1998 6:51:36 PM | EMS |  | AirConcentration | µg/m³ | Y | (air conc.) |  | Fluoride-Soluble (air conc.) |  |  | EMS code: F--S | Chemical | Soluble | LAB
F--S | ug/m3 | Fluoride-Soluble | 5087 | Specific Ion Electrode | Specific Ion Electrode | 0.5 | Y | Y | Y | 747 |  |  | 7/3/2024 12:59:18 PM | 1996-04-12 | 6/3/1998 6:51:42 PM | EMS |  | AirConcentration | µg/m³ | Y | (air conc.) |  | Fluoride-Soluble (air conc.) |  |  | EMS code: F--S | Chemical | Soluble | LAB
F009 | ug/g | Fenitrothion | ZN03 | DCM+Ac.Extr:GC/MS | DCM+Ac.Extr:GC/MS EPA 8270E (mod) | 0.05 | Y | Y | Y | 19 | 4/20/2002 10:25:00 AM | 10/16/2002 3:20:00 PM | 2/25/2004 1:09:45 AM | 1996-04-12 | 11/16/1999 2:27:40 PM | EMS | NPEPPIN | MassPerMass | mg/kg | Y | (mass conc.) |  | Fenitrothion (mass conc.) | 122-14-5 | Pesticides; | Pesticides; EMS code: F009 | Chemical |  | LAB
F009 | mg/L | Fenitrothion | ZN01 | DCM+Ac.Extr:GC/MS | DCM+Ac.Extr:GC/MS | 0.0005 | Y | Y | Y | 52 | 10/3/2002 | 3/29/2012 10:15:00 PM | 5/3/2012 9:27:51 AM | 1996-04-12 | 11/16/1999 4:06:27 PM | EMS | NPEPPIN | Concentration | mg/L | Y | (fl. conc.) |  | Fenitrothion (fl. conc.) | 122-14-5 | Pesticides; | Pesticides; EMS code: F009 | Chemical |  | LAB
F010 | ug/m3 | Formaldehyde CAS: 50-00-0 | DNPA | DNPH + HPLC | Ambient air is drawn through a pre-packed silica gel cartridge coated with acidified 2,4-dinitrophenylhydrazine (DNPH) reagent, at a sampling rate of 500 - 1200 mL/min. Aldehydes and ketones readily form a stable derivative with DNPH reagent. These derivatives are analyzed using HPLC. This method uses a coated adsorbent for sampling formaldehyde. The reaction of organic carbonyl compounds (aldehydes and ketones) with DNPH-coated cartridges in the presence of acid, forms a stable derivative. The sampling method gives a time weighted average and can be used for a 1 - 24 hours ambient air sampling time where the concentration of formaldehyde is in the low ppb (1 - 20 v/v) range or for short term (5 - 60 min) of source-impacted atmosphere where the concentration is in the ppm range. Sampling flow rate is limited to 1.5 L/min because of the high pressure drop across the DNPH-coated silica gel cartridges. This procedure is therefore not compatible with pumps used in personal sampling equipment. | 0.03 | Y | Y | Y | 113 |  |  | 3/3/1998 11:23:38 AM | 1996-04-12 | 6/3/1998 6:51:36 PM | EMS |  | AirConcentration | µg/m³ | Y | (air conc.) |  | Formaldehyde (air conc.) | 50-00-0 | Volatile Organic Compounds (VOCs); | Volatile Organic Compounds (VOCs); EMS code: F010 | Chemical |  | LAB
F010 | mg/L | Formaldehyde CAS: 50-00-0 | E-556 | EPA 556.1 | See referance for EPA mehod 556.1 DETERMINATION OF CARBONYL COMPOUNDS IN |  | Y | Y | Y | 14 | 12/29/2021 | 1/12/2022 | 9/2/2022 6:21:53 PM | 2022-04-06 |  | EMS |  | Concentration | mg/L | Y | (fl. conc.) |  | Formaldehyde (fl. conc.) | 50-00-0 | Volatile Organic Compounds (VOCs); | Volatile Organic Compounds (VOCs); EMS code: F010 | Chemical |  | LAB
F010 | ug/g | Formaldehyde CAS: 50-00-0 | E-556S | EPA 556.1 in Soil | See EPA method 556.1 | 0.04 | Y | Y | Y | 15 | 1/13/2022 | 2/2/2022 | 1/11/2024 1:07:07 PM | 2022-06-27 |  | EMS |  | MassPerMass | mg/kg | Y | (mass conc.) |  | Formaldehyde (mass conc.) | 50-00-0 | Volatile Organic Compounds (VOCs); | Volatile Organic Compounds (VOCs); EMS code: F010 | Chemical |  | LAB
F018 | mg/L | Fluazifop-p-butyl | AEWA | H2SO4 acid,CHCl2 extr.florisil. GC/MS | Effluent and water samples are acidified with Sulphuric Acid and extracted with Dichloromethane. After concentration, the extract is methylated using ethereal Diazomethane. Extracts are cleaned up as required using Florisil. The extracts are analyzed for Acid Extractable Herbicides using GC/MS. | 0.0005 | Y | Y | Y | 1 | 7/28/2005 7:03:00 AM | 7/28/2005 7:03:00 AM | 6/2/2009 12:01:05 PM | 2003-06-17 |  | UMULLER |  | Concentration | mg/L | Y | (fl. conc.) |  | Fluazifop-p-butyl (fl. conc.) |  | Pesticides; | Pesticides; EMS code: F018 | Chemical |  | LAB
F018 | ug/g | Fluazifop-p-butyl | ZN03 | DCM+Ac.Extr:GC/MS | DCM+Ac.Extr:GC/MS EPA 8270E (mod) | 0.05 | Y | Y | Y | 2 | 9/8/2022 | 9/8/2022 | 12/15/2022 10:50:07 AM | 2022-12-14 |  | EMS |  | MassPerMass | mg/kg | Y | (mass conc.) |  | Fluazifop-p-butyl (mass conc.) |  | Pesticides; | Pesticides; EMS code: F018 | Chemical |  | LAB
F018 | mg/L | Fluazifop-p-butyl | E660 | EPA 8270E (mod) | EPA 8270E (mod) Miscellaneous Pesticides by GC-MS | 0.0001 | Y | Y | Y | 68 | 12/7/2021 | 9/12/2022 | 12/8/2022 12:57:05 PM | 2021-12-09 | 12/9/2021 3:22:12 PM | EMS | EMS | Concentration | mg/L | Y | (fl. conc.) |  | Fluazifop-p-butyl (fl. conc.) |  | Pesticides; | Pesticides; EMS code: F018 | Chemical |  | LAB
F030 | ug/g | Fensulfothion | X260 | Sox-Ac/CH2Cl2 ext:GC/NPD | Sox-Ac/CH2Cl2 ext:GC/NPD | 0.01 | Y | Y | Y | 97 | 4/20/2002 10:25:00 AM | 10/16/2002 3:20:00 PM | 2/25/2004 1:09:45 AM | 1996-04-12 | 6/3/1998 6:51:36 PM | EMS |  | MassPerMass | mg/kg | Y | (mass conc.) |  | Fensulfothion (mass conc.) |  | Pesticides; | Pesticides; EMS code: F030 | Chemical |  | LAB
F030 | mg/L | Fensulfothion | ZN01 | DCM+Ac.Extr:GC/MS | DCM+Ac.Extr:GC/MS | 0.0002 | Y | Y | Y | 27 | 11/1/2008 | 3/29/2012 10:15:00 PM | 5/3/2012 9:27:51 AM | 1996-04-12 | 11/16/1999 4:06:27 PM | EMS | NPEPPIN | Concentration | mg/L | Y | (fl. conc.) |  | Fensulfothion (fl. conc.) |  | Pesticides; | Pesticides; EMS code: F030 | Chemical |  | LAB
F031 | mg/L | Fenthion | ZN01 | DCM+Ac.Extr:GC/MS | DCM+Ac.Extr:GC/MS | 0.0001 | Y | Y | Y | 51 | 10/3/2002 | 3/29/2012 10:15:00 PM | 5/3/2012 9:27:51 AM | 1996-04-12 | 11/16/1999 4:06:27 PM | EMS | NPEPPIN | Concentration | mg/L | Y | (fl. conc.) |  | Fenthion (fl. conc.) |  | Pesticides; | Pesticides; EMS code: F031 | Chemical |  | LAB
F031 | ug/g | Fenthion | X260 | Sox-Ac/CH2Cl2 ext:GC/NPD | Sox-Ac/CH2Cl2 ext:GC/NPD | 0.02 | Y | Y | Y | 97 | 4/20/2002 10:25:00 AM | 10/16/2002 3:20:00 PM | 2/25/2004 1:09:45 AM | 1996-04-12 | 6/3/1998 6:51:36 PM | EMS |  | MassPerMass | mg/kg | Y | (mass conc.) |  | Fenthion (mass conc.) |  | Pesticides; | Pesticides; EMS code: F031 | Chemical |  | LAB
F032 | mg/L | Fonofos | ZN01 | DCM+Ac.Extr:GC/MS | DCM+Ac.Extr:GC/MS | 0.0001 | Y | Y | Y | 52 | 10/3/2002 | 3/29/2012 10:15:00 PM | 5/3/2012 9:27:51 AM | 1996-04-12 | 11/16/1999 4:06:27 PM | EMS | NPEPPIN | Concentration | mg/L | Y | (fl. conc.) |  | Fonofos (fl. conc.) | 944-22-9 | Pesticides; | Pesticides; EMS code: F032 | Chemical |  | LAB
F032 | ug/g | Fonofos | X260 | Sox-Ac/CH2Cl2 ext:GC/NPD | Sox-Ac/CH2Cl2 ext:GC/NPD | 0.02 | Y | Y | Y | 97 | 4/20/2002 10:25:00 AM | 10/16/2002 3:20:00 PM | 2/25/2004 1:09:45 AM | 1996-04-12 | 6/3/1998 6:51:45 PM | EMS |  | MassPerMass | mg/kg | Y | (mass conc.) |  | Fonofos (mass conc.) | 944-22-9 | Pesticides; | Pesticides; EMS code: F032 | Chemical |  | LAB
F033 | mg/L | Folpet | ZN01 | DCM+Ac.Extr:GC/MS | DCM+Ac.Extr:GC/MS | 0.001 | Y | Y | Y | 28 | 7/28/2005 7:03:00 AM | 3/29/2012 10:15:00 PM | 5/3/2012 9:27:51 AM | 1996-04-12 | 11/16/1999 4:06:27 PM | EMS | NPEPPIN | Concentration | mg/L | Y | (fl. conc.) |  | Folpet (fl. conc.) | 133-07-3 | Pesticides; | Pesticides; EMS code: F033 | Chemical |  | LAB
F034 | ug/g | Fenchlorophos | ZN03 | DCM+Ac.Extr:GC/MS | DCM+Ac.Extr:GC/MS EPA 8270E (mod) | 0.01 | Y | Y | Y | 2 | 11/2/2022 12:53:00 AM | 11/2/2022 4:41:00 AM | 1/7/2023 3:12:10 PM | 1996-04-12 | 11/16/1999 2:27:40 PM | EMS | NPEPPIN | MassPerMass | mg/kg | Y | (mass conc.) |  | Fenchlorophos (mass conc.) |  | Pesticides; | Pesticides; EMS code: F034 | Chemical |  | LAB
F034 | mg/L | Fenchlorophos | ZN01 | DCM+Ac.Extr:GC/MS | DCM+Ac.Extr:GC/MS | 0.0001 | Y | Y | Y | 1 | 7/29/2005 7:10:00 AM | 7/29/2005 7:10:00 AM | 6/2/2009 12:01:05 PM | 1996-04-12 | 11/16/1999 4:06:27 PM | EMS | NPEPPIN | Concentration | mg/L | Y | (fl. conc.) |  | Fenchlorophos (fl. conc.) |  | Pesticides; | Pesticides; EMS code: F034 | Chemical |  | LAB
F036 | mg/L | Fenson | ZN01 | DCM+Ac.Extr:GC/MS | DCM+Ac.Extr:GC/MS | 0.0005 | Y | Y | Y | 1 | 7/29/2005 7:10:00 AM | 7/29/2005 7:10:00 AM | 6/2/2009 12:01:05 PM | 2003-06-16 |  | UMULLER |  | Concentration | mg/L | Y | (fl. conc.) |  | Fenson (fl. conc.) |  | Pesticides; | Pesticides; EMS code: F036 | Chemical |  | LAB
F037 | mg/L | Fenvalerate | ZN01 | DCM+Ac.Extr:GC/MS | DCM+Ac.Extr:GC/MS | 0.0005 | Y | Y | Y | 1 | 7/29/2005 7:10:00 AM | 7/29/2005 7:10:00 AM | 6/2/2009 12:01:05 PM | 2003-06-17 |  | UMULLER |  | Concentration | mg/L | Y | (fl. conc.) |  | Fenvalerate (fl. conc.) | 51630-58-1 | Pesticides; | Pesticides; EMS code: F037 | Chemical |  | LAB
F039 | mg/L | Fluchloralin | ZN01 | DCM+Ac.Extr:GC/MS | DCM+Ac.Extr:GC/MS | 0.0005 | Y | Y | Y | 1 | 7/29/2005 7:10:00 AM | 7/29/2005 7:10:00 AM | 6/2/2009 12:01:05 PM | 2003-06-17 |  | UMULLER |  | Concentration | mg/L | Y | (fl. conc.) |  | Fluchloralin (fl. conc.) |  | Pesticides; | Pesticides; EMS code: F039 | Chemical |  | LAB
F100 | ml/L | Fenamiphos | ZN01 | DCM+Ac.Extr:GC/MS | DCM+Ac.Extr:GC/MS | 0.5 | Y | Y | Y | 1 | 7/29/2005 7:10:00 AM | 7/29/2005 7:10:00 AM | 6/2/2009 12:01:05 PM | 2009-02-20 |  | JCHUDAK |  | VolumePerVolume | mL/L | Y | (vol./vol.) |  | Fenamiphos (vol./vol.) | 22224-92-6 | Pesticides; | Pesticides; EMS code: F100 | Chemical |  | LAB
F101 | mg/L | Fenoxapropethyl | ZN01 | DCM+Ac.Extr:GC/MS | DCM+Ac.Extr:GC/MS | 0.5 | Y | Y | Y | 1 | 7/28/2005 7:03:00 AM | 7/28/2005 7:03:00 AM | 6/2/2009 12:01:05 PM | 2009-02-20 |  | JCHUDAK |  | Concentration | mg/L | Y | (fl. conc.) |  | Fenoxapropethyl (fl. conc.) |  | Pesticides; | Pesticides; EMS code: F101 | Chemical |  | LAB
F1BT | mg/L | F1-BTEX | F1BX | F1-BTX-CALC-VA | This analysis is based on the "Reference Method for the Canada-Wide Standard for Petroleum Hydrocarbons in Soil - Tier 1 Method, Canadian Council of Ministers of the Environment, December 2000." For F1 (C6-C10), the sample undergoes a purge and trap extraction prior to analysis by GC/FID. The F1-BTEX result is calculated as follows: F1-BTEX: F1 (C6-C10) minus benzene, toluene, ethylbenzene and xylenes (BTEX). | 0.1 | Y | Y | Y | 183 | 7/11/2019 | 4/30/2024 | 5/2/2024 1:21:06 PM | 2019-08-20 |  | KATFRASE |  | Concentration | mg/L | Y | (fl. conc.) |  | F1-BTEX (fl. conc.) |  | Petroleum Hydrocarbons; | Petroleum Hydrocarbons; EMS code: F1BT | Chemical |  | LAB
F1BT | ug/g | F1-BTEX | F1BXS | CCCME PHC Soil | CCME PHC in Soil - Tier 1 | 5 | Y | Y | Y | 15 | 1/14/2022 | 2/7/2022 | 1/11/2024 1:07:07 PM | 2022-06-29 |  | EMS |  | MassPerMass | mg/kg | Y | (mass conc.) |  | F1-BTEX (mass conc.) |  | Petroleum Hydrocarbons; | Petroleum Hydrocarbons; EMS code: F1BT | Chemical |  | LAB
F1C6 | ug/g | F1 (C6-C10) | MS03 | Static Headspace GC/MS (Soil) | (Soil) A static headspace procedure is used as a means of driving the analytes out of the matrix into a headspace. Once equilibrium is achieved, a fixed volume of headspeace is taken for subsequent determination by Gas Chromatography - Mass Spectroscopy (GC/MS) using Selected Ion Monitoring. | 5 | Y | Y | Y | 14 | 1/14/2022 | 2/4/2022 | 1/11/2024 1:07:07 PM | 2022-07-27 |  | EMS |  | MassPerMass | mg/kg | Y | (mass conc.) |  | F1 (C6-C10) (mass conc.) |  | Petroleum Hydrocarbons; | Petroleum Hydrocarbons; EMS code: F1C6 | Chemical |  | LAB
F1C6 | mg/L | F1 (C6-C10) | F1HS | F1-HSFID-VA | This analysis is based on the "Reference Method for the Canada-Wide Standard for Petroleum Hydrocarbons in Soil - Tier 1 Method, Canadian Council of Ministers of the Environment, December 2000." For F1 (C6-C10), the sample undergoes a headspace purge prior to analysis by GC/FID. F1 (C6-C10): Sum of all hydrocarbons that elute between nC6 and nC10. | 0.1 | Y | Y | Y | 340 | 7/11/2019 | 7/31/2024 | 8/2/2024 12:05:19 PM | 2019-08-20 |  | KATFRASE |  | Concentration | mg/L | Y | (fl. conc.) |  | F1 (C6-C10) (fl. conc.) |  | Petroleum Hydrocarbons; | Petroleum Hydrocarbons; EMS code: F1C6 | Chemical |  | LAB
F1CD | ug/g | F1 corrected (C6-C10) | PHCS | Petroleum Hydrocarbon in Soil | CCME - Canada Wide Standard for Petroleum Hydrocarbon in Soil |  | Y | Y | Y | 4 | 9/4/2009 | 8/16/2010 | 8/27/2010 4:03:15 PM | 2007-10-04 |  | RDGIBSON |  | MassPerMass | mg/kg | Y | (mass conc.) |  | F1 corrected (C6-C10) (mass conc.) |  | Petroleum Hydrocarbons; | Petroleum Hydrocarbons; EMS code: F1CD | Chemical |  | LAB
F1UC | mg/L | F1 uncorrected (C6-C10) | PFCS | Petroleum Fraction | CCME - Canada Wide Standard Petroleum Fraction |  | Y | Y | Y | 255 | 1/17/2011 | 7/6/2024 9:57:00 PM | 7/22/2024 3:17:33 PM | 2007-10-04 |  | RDGIBSON |  | Concentration | mg/L | Y | (fl. conc.) |  | F1 uncorrected (C6-C10) (fl. conc.) |  | Petroleum Hydrocarbons; | Petroleum Hydrocarbons; EMS code: F1UC | Chemical |  | LAB
F1UC | ug/g | F1 uncorrected (C6-C10) | PHCS | Petroleum Hydrocarbon in Soil | CCME - Canada Wide Standard for Petroleum Hydrocarbon in Soil |  | Y | Y | Y | 18 | 3/4/2013 6:06:00 PM | 3/7/2013 3:50:00 AM | 3/19/2013 1:15:54 PM | 2007-10-04 |  | RDGIBSON |  | MassPerMass | mg/kg | Y | (mass conc.) |  | F1 uncorrected (C6-C10) (mass conc.) |  | Petroleum Hydrocarbons; | Petroleum Hydrocarbons; EMS code: F1UC | Chemical |  | LAB
F2 | ug/g | F2 (C10-C16) | PHCS | Petroleum Hydrocarbon in Soil | CCME - Canada Wide Standard for Petroleum Hydrocarbon in Soil | 25 | Y | Y | Y | 15 | 1/10/2022 | 2/7/2022 | 1/11/2024 1:07:07 PM | 2022-07-27 |  | EMS |  | MassPerMass | mg/kg | Y | (mass conc.) |  | F2 (C10-C16) (mass conc.) |  | Petroleum Hydrocarbons; | Petroleum Hydrocarbons; EMS code: F2 | Chemical |  | LAB
F2 | mg/L | F2 (C10-C16) | PHC | PHC in Soil | CCME PHC in Soil - Tier 1 |  | Y | Y | Y | 14 | 12/30/2021 | 1/13/2022 | 9/2/2022 6:21:53 PM | 2022-04-20 |  | EMS |  | Concentration | mg/L | Y | (fl. conc.) |  | F2 (C10-C16) (fl. conc.) |  | Petroleum Hydrocarbons; | Petroleum Hydrocarbons; EMS code: F2 | Chemical |  | LAB
F2UC | ug/g | F2 uncorrected (C10-C16) | PHCS | Petroleum Hydrocarbon in Soil | CCME - Canada Wide Standard for Petroleum Hydrocarbon in Soil |  | Y | Y | Y | 2 | 9/2/2009 | 9/2/2009 | 9/28/2009 4:14:43 PM | 2007-10-04 |  | RDGIBSON |  | MassPerMass | mg/kg | Y | (mass conc.) |  | F2 uncorrected (C10-C16) (mass conc.) |  | Petroleum Hydrocarbons; | Petroleum Hydrocarbons; EMS code: F2UC | Chemical |  | LAB
F3 | mg/L | F3 (C16-C34) | PHC | PHC in Soil | CCME PHC in Soil - Tier 1 |  | Y | Y | Y | 14 | 12/30/2021 | 1/13/2022 | 9/2/2022 6:21:53 PM | 2022-04-20 |  | EMS |  | Concentration | mg/L | Y | (fl. conc.) |  | F3 (C16-C34) (fl. conc.) |  | Petroleum Hydrocarbons; | Petroleum Hydrocarbons; EMS code: F3 | Chemical |  | LAB
F3 | ug/g | F3 (C16-C34) | PHCS | Petroleum Hydrocarbon in Soil | CCME - Canada Wide Standard for Petroleum Hydrocarbon in Soil | 50 | Y | Y | Y | 15 | 1/10/2022 | 2/7/2022 | 1/11/2024 1:07:07 PM | 2022-07-27 |  | EMS |  | MassPerMass | mg/kg | Y | (mass conc.) |  | F3 (C16-C34) (mass conc.) |  | Petroleum Hydrocarbons; | Petroleum Hydrocarbons; EMS code: F3 | Chemical |  | LAB
F3UC | ug/g | F3 uncorrected (C16-C34) | PHCS | Petroleum Hydrocarbon in Soil | CCME - Canada Wide Standard for Petroleum Hydrocarbon in Soil |  | Y | Y | Y | 2 | 9/2/2009 | 9/2/2009 | 9/28/2009 4:14:43 PM | 2007-10-04 |  | RDGIBSON |  | MassPerMass | mg/kg | Y | (mass conc.) |  | F3 uncorrected (C16-C34) (mass conc.) |  | Petroleum Hydrocarbons; | Petroleum Hydrocarbons; EMS code: F3UC | Chemical |  | LAB
F4 | mg/L | F4 (C34-C50) | PHC | PHC in Soil | CCME PHC in Soil - Tier 1 |  | Y | Y | Y | 14 | 12/30/2021 | 1/13/2022 | 9/2/2022 6:21:53 PM | 2022-04-20 |  | EMS |  | Concentration | mg/L | Y | (fl. conc.) |  | F4 (C34-C50) (fl. conc.) |  | Petroleum Hydrocarbons; | Petroleum Hydrocarbons; EMS code: F4 | Chemical |  | LAB
F4 | ug/g | F4 (C34-C50) | PHCS | Petroleum Hydrocarbon in Soil | CCME - Canada Wide Standard for Petroleum Hydrocarbon in Soil | 50 | Y | Y | Y | 15 | 1/10/2022 | 2/7/2022 | 1/11/2024 1:07:07 PM | 2022-07-27 |  | EMS |  | MassPerMass | mg/kg | Y | (mass conc.) |  | F4 (C34-C50) (mass conc.) |  | Petroleum Hydrocarbons; | Petroleum Hydrocarbons; EMS code: F4 | Chemical |  | LAB
FA01 | ug/g | Lauric Acid | MP02 | GC/MS | GC/MS | 0.05 | Y | Y | Y | 2 | 6/21/2002 12:00:00 PM | 6/21/2002 12:25:00 PM | 3/27/2004 1:07:19 AM | 1996-04-12 | 6/3/1998 6:51:42 PM | EMS |  | MassPerMass | mg/kg | Y | (mass conc.) |  | Lauric Acid (mass conc.) |  | Organics; | Organics; EMS code: FA01 | Chemical |  | LAB
FA01 | mg/L | Lauric Acid | F045 | pH2,DEE extr'n,conc,MTBEextr'n,methyl'n,HRGC/LMRS | Effluent samples are first adjusted to pH 2 and then extracted with Diethyl Ether. The extracts are rotary evaporated and the resulting residues, containing the Fatty Acids and other solvent soluble organics, are then redissolved in Methyl-t-butyl Ether (MTBE), transferred to small glass vials, methylated with Diazomethane, concentrated and cleaned up. The extract is injected on a High Resolution Gas Chromatograph/Low Resolution Mass Spectrometer (HRGC/LRMS) system for analysis. LRMS data is acquired in Total Ion Mode (TIM) and quantitation is performed using selected mass ions. Positive identification of the various Fatty Acids is made on the basis of a comparison of absolute retention times to those of the external standards and computer based mass spectral mass/intensity library matching routines. | 0.001 | Y | Y | Y | 71 | 3/18/1999 2:44:00 PM | 5/13/2003 | 4/1/2004 1:08:45 AM | 1997-10-20 | 6/3/1998 6:51:47 PM | NPEPPIN |  | Concentration | mg/L | Y | (fl. conc.) |  | Lauric Acid (fl. conc.) |  | Organics; | Organics; EMS code: FA01 | Chemical |  | LAB
FA01 | mg/L | Lauric Acid | E677 | Resins & Fatty Acids are methylated then analyzed by GCMS | EPA 8270E; AB Env. 129; NCASI RA/FA-85.02 (mod) Resins & Fatty Acids are methylated then analyzed by GCMS. Total fatty acids is the sum of arachidic acid; behenic acid; lauric acid; lignoceric acid; linoleic acid; linolenic acid; myristic acid; oleic acid; palmitic acid; and stearic acid. Total resin acids is the sum of abietic acid; 12-chlorodehydroabietic acid; 14-chlorodehydroabietic acid; dehydroabietic acid; 12;14-dichlorodehydroabietic acid; isopimaric acid; levopimaric acid; neoabietic acid; palustric acid; pimaric acid and sandaracopimaric acid. Total resin acids (BC WWQG) is the sum of abietic acid; neoabietic acid; pimaric acid; isopimaric acid; and sandaracopimaric acid. | 0.05 | Y | Y | Y | 12 | 11/28/2022 | 6/15/2024 | 7/8/2024 11:10:02 AM | 2021-12-30 |  | EMS |  | Concentration | mg/L | Y | (fl. conc.) |  | Lauric Acid (fl. conc.) |  | Organics; | Organics; EMS code: FA01 | Chemical |  | LAB
FA01 | ug/g | Lauric Acid | P035 | DCM-DIAZO-FLOR | DCM-DIAZO-FLOR | 0.05 | Y | Y | Y | 22 |  |  | 10/24/1996 12:44:23 PM | 1996-04-12 | 3/8/2001 11:15:17 AM | EMS | NPEPPIN | MassPerMass | mg/kg | Y | (mass conc.) |  | Lauric Acid (mass conc.) |  | Organics; | Organics; EMS code: FA01 | Chemical |  | LAB
FA01 | mg/L | Lauric Acid | X239 | Gas Chromatography | Gas Chromatography | 0.001 | Y | Y | Y | 12 |  |  | 10/23/1996 11:56:36 PM | 1996-04-12 | 6/3/1998 6:51:36 PM | EMS |  | Concentration | mg/L | Y | (fl. conc.) |  | Lauric Acid (fl. conc.) |  | Organics; | Organics; EMS code: FA01 | Chemical |  | LAB
FA01 | mg/L | Lauric Acid | P030 | Resin Acids | The water sample is acidified and extracted with dichloro-methane. The raw extract is concentrated and derivatized with diazomethane to produce the corresponding methyl ester derivatives. (Suitable substitute derivatives may be produced by other available techniques.) If required, the extracts are cleaned up by Florisil column chromatography. The derivatives are analyzed by gas chromatography with flame ionization detection. | 0.001 | Y | Y | Y | 75 | 8/28/2014 | 8/15/2017 | 8/15/2017 4:37:09 PM | 1996-04-12 | 6/3/1998 6:51:36 PM | EMS |  | Concentration | mg/L | Y | (fl. conc.) |  | Lauric Acid (fl. conc.) |  | Organics; | Organics; EMS code: FA01 | Chemical |  | LAB
FA02 | mg/L | Myristic | X239 | Gas Chromatography | Gas Chromatography | 0.001 | Y | Y | Y | 12 |  |  | 10/23/1996 11:56:36 PM | 1996-04-12 | 6/3/1998 6:51:36 PM | EMS |  | Concentration | mg/L | Y | (fl. conc.) |  | Myristic (fl. conc.) |  | Organics; | Organics; EMS code: FA02 | Chemical |  | LAB
FA02 | mg/L | Myristic | E677 | Resins & Fatty Acids are methylated then analyzed by GCMS | EPA 8270E; AB Env. 129; NCASI RA/FA-85.02 (mod) Resins & Fatty Acids are methylated then analyzed by GCMS. Total fatty acids is the sum of arachidic acid; behenic acid; lauric acid; lignoceric acid; linoleic acid; linolenic acid; myristic acid; oleic acid; palmitic acid; and stearic acid. Total resin acids is the sum of abietic acid; 12-chlorodehydroabietic acid; 14-chlorodehydroabietic acid; dehydroabietic acid; 12;14-dichlorodehydroabietic acid; isopimaric acid; levopimaric acid; neoabietic acid; palustric acid; pimaric acid and sandaracopimaric acid. Total resin acids (BC WWQG) is the sum of abietic acid; neoabietic acid; pimaric acid; isopimaric acid; and sandaracopimaric acid. | 0.005 | Y | Y | Y | 12 | 11/28/2022 | 6/15/2024 | 7/8/2024 11:10:02 AM | 2021-12-30 |  | EMS |  | Concentration | mg/L | Y | (fl. conc.) |  | Myristic (fl. conc.) |  | Organics; | Organics; EMS code: FA02 | Chemical |  | LAB
FA02 | mg/L | Myristic | P030 | Resin Acids | The water sample is acidified and extracted with dichloro-methane. The raw extract is concentrated and derivatized with diazomethane to produce the corresponding methyl ester derivatives. (Suitable substitute derivatives may be produced by other available techniques.) If required, the extracts are cleaned up by Florisil column chromatography. The derivatives are analyzed by gas chromatography with flame ionization detection. | 0.001 | Y | Y | Y | 73 | 8/28/2014 | 8/15/2017 | 8/15/2017 4:37:09 PM | 1996-04-12 | 6/3/1998 6:51:36 PM | EMS |  | Concentration | mg/L | Y | (fl. conc.) |  | Myristic (fl. conc.) |  | Organics; | Organics; EMS code: FA02 | Chemical |  | LAB
FA02 | ug/g | Myristic | MP02 | GC/MS | GC/MS | 0.05 | Y | Y | Y | 2 | 6/21/2002 12:00:00 PM | 6/21/2002 12:25:00 PM | 3/27/2004 1:07:19 AM | 1996-04-12 | 6/3/1998 6:51:42 PM | EMS |  | MassPerMass | mg/kg | Y | (mass conc.) |  | Myristic (mass conc.) |  | Organics; | Organics; EMS code: FA02 | Chemical |  | LAB
FA02 | mg/L | Myristic | P035 | DCM-DIAZO-FLOR | DCM-DIAZO-FLOR | 0.0005 | Y | Y | Y | 22 |  |  | 10/24/1996 12:44:23 PM | 1996-04-12 | 6/3/1998 6:51:36 PM | EMS |  | Concentration | mg/L | Y | (fl. conc.) |  | Myristic (fl. conc.) |  | Organics; | Organics; EMS code: FA02 | Chemical |  | LAB
FA02 | mg/L | Myristic | F045 | pH2,DEE extr'n,conc,MTBEextr'n,methyl'n,HRGC/LMRS | Effluent samples are first adjusted to pH 2 and then extracted with Diethyl Ether. The extracts are rotary evaporated and the resulting residues, containing the Fatty Acids and other solvent soluble organics, are then redissolved in Methyl-t-butyl Ether (MTBE), transferred to small glass vials, methylated with Diazomethane, concentrated and cleaned up. The extract is injected on a High Resolution Gas Chromatograph/Low Resolution Mass Spectrometer (HRGC/LRMS) system for analysis. LRMS data is acquired in Total Ion Mode (TIM) and quantitation is performed using selected mass ions. Positive identification of the various Fatty Acids is made on the basis of a comparison of absolute retention times to those of the external standards and computer based mass spectral mass/intensity library matching routines. | 0.001 | Y | Y | Y | 71 | 3/18/1999 2:44:00 PM | 5/13/2003 | 4/1/2004 1:08:45 AM | 1997-10-20 | 6/3/1998 6:51:47 PM | NPEPPIN |  | Concentration | mg/L | Y | (fl. conc.) |  | Myristic (fl. conc.) |  | Organics; | Organics; EMS code: FA02 | Chemical |  | LAB
FA03 | mg/L | Palmitic | P035 | DCM-DIAZO-FLOR | DCM-DIAZO-FLOR | 0.0005 | Y | Y | Y | 22 |  |  | 10/24/1996 12:44:23 PM | 1996-04-12 | 6/3/1998 6:51:36 PM | EMS |  | Concentration | mg/L | Y | (fl. conc.) |  | Palmitic (fl. conc.) |  | Organics; | Organics; EMS code: FA03 | Chemical |  | LAB
FA03 | mg/L | Palmitic | E677 | Resins & Fatty Acids are methylated then analyzed by GCMS | EPA 8270E; AB Env. 129; NCASI RA/FA-85.02 (mod) Resins & Fatty Acids are methylated then analyzed by GCMS. Total fatty acids is the sum of arachidic acid; behenic acid; lauric acid; lignoceric acid; linoleic acid; linolenic acid; myristic acid; oleic acid; palmitic acid; and stearic acid. Total resin acids is the sum of abietic acid; 12-chlorodehydroabietic acid; 14-chlorodehydroabietic acid; dehydroabietic acid; 12;14-dichlorodehydroabietic acid; isopimaric acid; levopimaric acid; neoabietic acid; palustric acid; pimaric acid and sandaracopimaric acid. Total resin acids (BC WWQG) is the sum of abietic acid; neoabietic acid; pimaric acid; isopimaric acid; and sandaracopimaric acid. | 0.05 | Y | Y | Y | 12 | 11/28/2022 | 6/15/2024 | 7/8/2024 11:10:02 AM | 2021-12-30 |  | EMS |  | Concentration | mg/L | Y | (fl. conc.) |  | Palmitic (fl. conc.) |  | Organics; | Organics; EMS code: FA03 | Chemical |  | LAB
FA03 | mg/L | Palmitic | X239 | Gas Chromatography | Gas Chromatography | 0.001 | Y | Y | Y | 12 |  |  | 10/23/1996 11:56:36 PM | 1996-04-12 | 6/3/1998 6:51:36 PM | EMS |  | Concentration | mg/L | Y | (fl. conc.) |  | Palmitic (fl. conc.) |  | Organics; | Organics; EMS code: FA03 | Chemical |  | LAB
FA03 | ug/g | Palmitic | MP02 | GC/MS | GC/MS | 0.05 | Y | Y | Y | 2 | 6/21/2002 12:00:00 PM | 6/21/2002 12:25:00 PM | 3/27/2004 1:07:19 AM | 1996-04-12 | 6/3/1998 6:51:42 PM | EMS |  | MassPerMass | mg/kg | Y | (mass conc.) |  | Palmitic (mass conc.) |  | Organics; | Organics; EMS code: FA03 | Chemical |  | LAB
FA03 | mg/L | Palmitic | F045 | pH2,DEE extr'n,conc,MTBEextr'n,methyl'n,HRGC/LMRS | Effluent samples are first adjusted to pH 2 and then extracted with Diethyl Ether. The extracts are rotary evaporated and the resulting residues, containing the Fatty Acids and other solvent soluble organics, are then redissolved in Methyl-t-butyl Ether (MTBE), transferred to small glass vials, methylated with Diazomethane, concentrated and cleaned up. The extract is injected on a High Resolution Gas Chromatograph/Low Resolution Mass Spectrometer (HRGC/LRMS) system for analysis. LRMS data is acquired in Total Ion Mode (TIM) and quantitation is performed using selected mass ions. Positive identification of the various Fatty Acids is made on the basis of a comparison of absolute retention times to those of the external standards and computer based mass spectral mass/intensity library matching routines. | 0.001 | Y | Y | Y | 71 | 3/18/1999 2:45:00 PM | 5/13/2003 | 4/1/2004 1:08:46 AM | 1997-10-20 | 6/3/1998 6:51:47 PM | NPEPPIN |  | Concentration | mg/L | Y | (fl. conc.) |  | Palmitic (fl. conc.) |  | Organics; | Organics; EMS code: FA03 | Chemical |  | LAB
FA03 | mg/L | Palmitic | P030 | Resin Acids | The water sample is acidified and extracted with dichloro-methane. The raw extract is concentrated and derivatized with diazomethane to produce the corresponding methyl ester derivatives. (Suitable substitute derivatives may be produced by other available techniques.) If required, the extracts are cleaned up by Florisil column chromatography. The derivatives are analyzed by gas chromatography with flame ionization detection. | 0.001 | Y | Y | Y | 78 | 8/28/2014 | 8/15/2017 | 8/15/2017 4:37:09 PM | 1996-04-12 | 6/3/1998 6:51:36 PM | EMS |  | Concentration | mg/L | Y | (fl. conc.) |  | Palmitic (fl. conc.) |  | Organics; | Organics; EMS code: FA03 | Chemical |  | LAB
FA04 | mg/L | Oleic/Linoleic | X239 | Gas Chromatography | Gas Chromatography | 0.001 | Y | Y | Y | 12 |  |  | 10/23/1996 11:56:36 PM | 1996-04-12 | 6/3/1998 6:51:36 PM | EMS |  | Concentration | mg/L | Y | (fl. conc.) |  | Oleic/Linoleic (fl. conc.) |  | Organics; | Organics; EMS code: FA04 | Chemical |  | LAB
FA05 | mg/L | Linoleic | P035 | DCM-DIAZO-FLOR | DCM-DIAZO-FLOR | 0.0005 | Y | Y | Y | 22 |  |  | 10/24/1996 12:44:23 PM | 1996-04-12 | 6/3/1998 6:51:36 PM | EMS |  | Concentration | mg/L | Y | (fl. conc.) |  | Linoleic (fl. conc.) |  | Organics; | Organics; EMS code: FA05 | Chemical |  | LAB
FA05 | mg/L | Linoleic | X239 | Gas Chromatography | Gas Chromatography | 0.001 | Y | Y | Y | 12 |  |  | 10/23/1996 11:56:36 PM | 1996-04-12 | 6/3/1998 6:51:36 PM | EMS |  | Concentration | mg/L | Y | (fl. conc.) |  | Linoleic (fl. conc.) |  | Organics; | Organics; EMS code: FA05 | Chemical |  | LAB
FA05 | mg/L | Linoleic | P030 | Resin Acids | The water sample is acidified and extracted with dichloro-methane. The raw extract is concentrated and derivatized with diazomethane to produce the corresponding methyl ester derivatives. (Suitable substitute derivatives may be produced by other available techniques.) If required, the extracts are cleaned up by Florisil column chromatography. The derivatives are analyzed by gas chromatography with flame ionization detection. | 0.001 | Y | Y | Y | 79 | 8/28/2014 | 8/15/2017 | 6/7/2024 5:30:18 PM | 1996-04-12 | 6/3/1998 6:51:45 PM | EMS |  | Concentration | mg/L | Y | (fl. conc.) |  | Linoleic (fl. conc.) |  | Organics; | Organics; EMS code: FA05 | Chemical |  | LAB
FA05 | mg/L | Linoleic | E677 | Resins & Fatty Acids are methylated then analyzed by GCMS | EPA 8270E; AB Env. 129; NCASI RA/FA-85.02 (mod) Resins & Fatty Acids are methylated then analyzed by GCMS. Total fatty acids is the sum of arachidic acid; behenic acid; lauric acid; lignoceric acid; linoleic acid; linolenic acid; myristic acid; oleic acid; palmitic acid; and stearic acid. Total resin acids is the sum of abietic acid; 12-chlorodehydroabietic acid; 14-chlorodehydroabietic acid; dehydroabietic acid; 12;14-dichlorodehydroabietic acid; isopimaric acid; levopimaric acid; neoabietic acid; palustric acid; pimaric acid and sandaracopimaric acid. Total resin acids (BC WWQG) is the sum of abietic acid; neoabietic acid; pimaric acid; isopimaric acid; and sandaracopimaric acid. | 0.005 | Y | Y | Y | 12 | 11/28/2022 | 6/15/2024 | 7/8/2024 11:10:02 AM | 2021-12-30 |  | EMS |  | Concentration | mg/L | Y | (fl. conc.) |  | Linoleic (fl. conc.) |  | Organics; | Organics; EMS code: FA05 | Chemical |  | LAB
FA05 | mg/L | Linoleic | RFAS | Separation of resin and fatty acids (RFAs) in water by GC/MS | Scientific Principle The separation of resin and fatty acids (RFAs) in water samples is performed using gas chromatography/mass spectrometry (GC/MS) in the Selected Ion Monitoring (SIM) mode. The separation is based on compound size; boiling point and polarity. Analyte concentration is calculated using the internal standard method. RFA compounds are extracted from water samples using a liquid-liquid solvent extraction with DCM. Extraction occurs after the sample is brought to a pH of ~ 2 and placed in a rotary tumbler (at approx. 40 rpm) for 1 hour. The resulting extract is transferred to a vial; dried with acidified baked sodium sulfate; methylated with diazomethane and concentrated to 1mL under nitrogen prior to analysis. The resulting extracts are vialed and fortified with labeled internal standard compounds prior to analysis by gas chromatography with a (50%-Phenyl)-methylpolysiloxane (DB-17) column and a mass spectrometer detector in SIM Mode. | 0.006 | Y | Y | Y | 1 | 9/17/2021 7:49:00 PM | 9/17/2021 7:49:00 PM | 2/1/2022 2:38:14 PM | 2022-01-19 |  | EMS |  | Concentration | mg/L | Y | (fl. conc.) |  | Linoleic (fl. conc.) |  | Organics; | Organics; EMS code: FA05 | Chemical |  | LAB
FA05 | mg/L | Linoleic | F045 | pH2,DEE extr'n,conc,MTBEextr'n,methyl'n,HRGC/LMRS | Effluent samples are first adjusted to pH 2 and then extracted with Diethyl Ether. The extracts are rotary evaporated and the resulting residues, containing the Fatty Acids and other solvent soluble organics, are then redissolved in Methyl-t-butyl Ether (MTBE), transferred to small glass vials, methylated with Diazomethane, concentrated and cleaned up. The extract is injected on a High Resolution Gas Chromatograph/Low Resolution Mass Spectrometer (HRGC/LRMS) system for analysis. LRMS data is acquired in Total Ion Mode (TIM) and quantitation is performed using selected mass ions. Positive identification of the various Fatty Acids is made on the basis of a comparison of absolute retention times to those of the external standards and computer based mass spectral mass/intensity library matching routines. | 0.001 | Y | Y | Y | 71 | 3/18/1999 2:45:00 PM | 5/13/2003 | 4/1/2004 1:08:46 AM | 1997-10-20 | 6/3/1998 6:51:47 PM | NPEPPIN |  | Concentration | mg/L | Y | (fl. conc.) |  | Linoleic (fl. conc.) |  | Organics; | Organics; EMS code: FA05 | Chemical |  | LAB
FA05 | ug/g | Linoleic | MP02 | GC/MS | GC/MS | 0.05 | Y | Y | Y | 2 | 6/21/2002 12:00:00 PM | 6/21/2002 12:25:00 PM | 3/27/2004 1:07:20 AM | 1996-04-12 | 6/3/1998 6:51:42 PM | EMS |  | MassPerMass | mg/kg | Y | (mass conc.) |  | Linoleic (mass conc.) |  | Organics; | Organics; EMS code: FA05 | Chemical |  | LAB
FA06 | mg/L | Stearic Acid | E677 | Resins & Fatty Acids are methylated then analyzed by GCMS | EPA 8270E; AB Env. 129; NCASI RA/FA-85.02 (mod) Resins & Fatty Acids are methylated then analyzed by GCMS. Total fatty acids is the sum of arachidic acid; behenic acid; lauric acid; lignoceric acid; linoleic acid; linolenic acid; myristic acid; oleic acid; palmitic acid; and stearic acid. Total resin acids is the sum of abietic acid; 12-chlorodehydroabietic acid; 14-chlorodehydroabietic acid; dehydroabietic acid; 12;14-dichlorodehydroabietic acid; isopimaric acid; levopimaric acid; neoabietic acid; palustric acid; pimaric acid and sandaracopimaric acid. Total resin acids (BC WWQG) is the sum of abietic acid; neoabietic acid; pimaric acid; isopimaric acid; and sandaracopimaric acid. | 0.05 | Y | Y | Y | 12 | 11/28/2022 | 6/15/2024 | 7/8/2024 11:10:02 AM | 2021-12-30 |  | EMS |  | Concentration | mg/L | Y | (fl. conc.) |  | Stearic Acid (fl. conc.) |  | Organics; | Organics; EMS code: FA06 | Chemical |  | LAB
FA06 | mg/L | Stearic Acid | X239 | Gas Chromatography | Gas Chromatography | 0.001 | Y | Y | Y | 12 |  |  | 10/23/1996 11:56:36 PM | 1996-04-12 | 6/3/1998 6:51:36 PM | EMS |  | Concentration | mg/L | Y | (fl. conc.) |  | Stearic Acid (fl. conc.) |  | Organics; | Organics; EMS code: FA06 | Chemical |  | LAB
FA06 | mg/L | Stearic Acid | P035 | DCM-DIAZO-FLOR | DCM-DIAZO-FLOR | 0.0005 | Y | Y | Y | 22 |  |  | 10/24/1996 12:44:23 PM | 1996-04-12 | 6/3/1998 6:51:45 PM | EMS |  | Concentration | mg/L | Y | (fl. conc.) |  | Stearic Acid (fl. conc.) |  | Organics; | Organics; EMS code: FA06 | Chemical |  | LAB
FA06 | ug/g | Stearic Acid | MP02 | GC/MS | GC/MS | 0.05 | Y | Y | Y | 2 | 6/21/2002 12:00:00 PM | 6/21/2002 12:25:00 PM | 3/27/2004 1:07:20 AM | 1996-04-12 | 6/3/1998 6:51:42 PM | EMS |  | MassPerMass | mg/kg | Y | (mass conc.) |  | Stearic Acid (mass conc.) |  | Organics; | Organics; EMS code: FA06 | Chemical |  | LAB
FA06 | mg/L | Stearic Acid | F045 | pH2,DEE extr'n,conc,MTBEextr'n,methyl'n,HRGC/LMRS | Effluent samples are first adjusted to pH 2 and then extracted with Diethyl Ether. The extracts are rotary evaporated and the resulting residues, containing the Fatty Acids and other solvent soluble organics, are then redissolved in Methyl-t-butyl Ether (MTBE), transferred to small glass vials, methylated with Diazomethane, concentrated and cleaned up. The extract is injected on a High Resolution Gas Chromatograph/Low Resolution Mass Spectrometer (HRGC/LRMS) system for analysis. LRMS data is acquired in Total Ion Mode (TIM) and quantitation is performed using selected mass ions. Positive identification of the various Fatty Acids is made on the basis of a comparison of absolute retention times to those of the external standards and computer based mass spectral mass/intensity library matching routines. | 0.001 | Y | Y | Y | 71 | 3/18/1999 2:46:00 PM | 5/13/2003 | 4/1/2004 1:08:46 AM | 1997-10-20 | 6/3/1998 6:51:47 PM | NPEPPIN |  | Concentration | mg/L | Y | (fl. conc.) |  | Stearic Acid (fl. conc.) |  | Organics; | Organics; EMS code: FA06 | Chemical |  | LAB
FA06 | mg/L | Stearic Acid | P030 | Resin Acids | The water sample is acidified and extracted with dichloro-methane. The raw extract is concentrated and derivatized with diazomethane to produce the corresponding methyl ester derivatives. (Suitable substitute derivatives may be produced by other available techniques.) If required, the extracts are cleaned up by Florisil column chromatography. The derivatives are analyzed by gas chromatography with flame ionization detection. | 0.001 | Y | Y | Y | 83 | 8/28/2014 | 8/15/2017 | 6/7/2024 5:30:18 PM | 1996-04-12 | 6/3/1998 6:51:36 PM | EMS |  | Concentration | mg/L | Y | (fl. conc.) |  | Stearic Acid (fl. conc.) |  | Organics; | Organics; EMS code: FA06 | Chemical |  | LAB
FA07 | mg/L | Arachidic | F045 | pH2,DEE extr'n,conc,MTBEextr'n,methyl'n,HRGC/LMRS | Effluent samples are first adjusted to pH 2 and then extracted with Diethyl Ether. The extracts are rotary evaporated and the resulting residues, containing the Fatty Acids and other solvent soluble organics, are then redissolved in Methyl-t-butyl Ether (MTBE), transferred to small glass vials, methylated with Diazomethane, concentrated and cleaned up. The extract is injected on a High Resolution Gas Chromatograph/Low Resolution Mass Spectrometer (HRGC/LRMS) system for analysis. LRMS data is acquired in Total Ion Mode (TIM) and quantitation is performed using selected mass ions. Positive identification of the various Fatty Acids is made on the basis of a comparison of absolute retention times to those of the external standards and computer based mass spectral mass/intensity library matching routines. | 0.001 | Y | Y | Y | 71 | 3/18/1999 2:46:00 PM | 5/13/2003 | 4/1/2004 1:08:46 AM | 1997-10-20 | 6/3/1998 6:51:47 PM | NPEPPIN |  | Concentration | mg/L | Y | (fl. conc.) |  | Arachidic (fl. conc.) |  | Organics; | Organics; EMS code: FA07 | Chemical |  | LAB
FA07 | mg/L | Arachidic | E677 | Resins & Fatty Acids are methylated then analyzed by GCMS | EPA 8270E; AB Env. 129; NCASI RA/FA-85.02 (mod) Resins & Fatty Acids are methylated then analyzed by GCMS. Total fatty acids is the sum of arachidic acid; behenic acid; lauric acid; lignoceric acid; linoleic acid; linolenic acid; myristic acid; oleic acid; palmitic acid; and stearic acid. Total resin acids is the sum of abietic acid; 12-chlorodehydroabietic acid; 14-chlorodehydroabietic acid; dehydroabietic acid; 12;14-dichlorodehydroabietic acid; isopimaric acid; levopimaric acid; neoabietic acid; palustric acid; pimaric acid and sandaracopimaric acid. Total resin acids (BC WWQG) is the sum of abietic acid; neoabietic acid; pimaric acid; isopimaric acid; and sandaracopimaric acid. | 0.005 | Y | Y | Y | 12 | 11/28/2022 | 6/15/2024 | 7/8/2024 11:10:02 AM | 2021-12-30 |  | EMS |  | Concentration | mg/L | Y | (fl. conc.) |  | Arachidic (fl. conc.) |  | Organics; | Organics; EMS code: FA07 | Chemical |  | LAB
FA07 | ug/g | Arachidic | MP02 | GC/MS | GC/MS | 0.05 | Y | Y | Y | 2 | 6/21/2002 12:00:00 PM | 6/21/2002 12:25:00 PM | 3/27/2004 1:07:20 AM | 1996-04-12 | 6/3/1998 6:51:42 PM | EMS |  | MassPerMass | mg/kg | Y | (mass conc.) |  | Arachidic (mass conc.) |  | Organics; | Organics; EMS code: FA07 | Chemical |  | LAB
FA07 | mg/L | Arachidic | X239 | Gas Chromatography | Gas Chromatography | 0.001 | Y | Y | Y | 12 |  |  | 10/23/1996 11:56:36 PM | 1996-04-12 | 6/3/1998 6:51:36 PM | EMS |  | Concentration | mg/L | Y | (fl. conc.) |  | Arachidic (fl. conc.) |  | Organics; | Organics; EMS code: FA07 | Chemical |  | LAB
FA07 | mg/L | Arachidic | P030 | Resin Acids | The water sample is acidified and extracted with dichloro-methane. The raw extract is concentrated and derivatized with diazomethane to produce the corresponding methyl ester derivatives. (Suitable substitute derivatives may be produced by other available techniques.) If required, the extracts are cleaned up by Florisil column chromatography. The derivatives are analyzed by gas chromatography with flame ionization detection. | 0.001 | Y | Y | Y | 75 | 8/28/2014 | 8/15/2017 | 8/15/2017 4:37:09 PM | 1996-04-12 | 6/3/1998 6:51:36 PM | EMS |  | Concentration | mg/L | Y | (fl. conc.) |  | Arachidic (fl. conc.) |  | Organics; | Organics; EMS code: FA07 | Chemical |  | LAB
FA07 | mg/L | Arachidic | P035 | DCM-DIAZO-FLOR | DCM-DIAZO-FLOR | 0.0005 | Y | Y | Y | 22 |  |  | 10/24/1996 12:44:23 PM | 1996-04-12 | 6/3/1998 6:51:45 PM | EMS |  | Concentration | mg/L | Y | (fl. conc.) |  | Arachidic (fl. conc.) |  | Organics; | Organics; EMS code: FA07 | Chemical |  | LAB
FA08 | mg/L | Behenic | F045 | pH2,DEE extr'n,conc,MTBEextr'n,methyl'n,HRGC/LMRS | Effluent samples are first adjusted to pH 2 and then extracted with Diethyl Ether. The extracts are rotary evaporated and the resulting residues, containing the Fatty Acids and other solvent soluble organics, are then redissolved in Methyl-t-butyl Ether (MTBE), transferred to small glass vials, methylated with Diazomethane, concentrated and cleaned up. The extract is injected on a High Resolution Gas Chromatograph/Low Resolution Mass Spectrometer (HRGC/LRMS) system for analysis. LRMS data is acquired in Total Ion Mode (TIM) and quantitation is performed using selected mass ions. Positive identification of the various Fatty Acids is made on the basis of a comparison of absolute retention times to those of the external standards and computer based mass spectral mass/intensity library matching routines. | 0.001 | Y | Y | Y | 71 | 3/18/1999 2:46:00 PM | 5/13/2003 | 4/1/2004 1:08:46 AM | 1997-10-20 | 6/3/1998 6:51:47 PM | NPEPPIN |  | Concentration | mg/L | Y | (fl. conc.) |  | Behenic (fl. conc.) |  | Organics; | Organics; EMS code: FA08 | Chemical |  | LAB
FA08 | mg/L | Behenic | P030 | Resin Acids | The water sample is acidified and extracted with dichloro-methane. The raw extract is concentrated and derivatized with diazomethane to produce the corresponding methyl ester derivatives. (Suitable substitute derivatives may be produced by other available techniques.) If required, the extracts are cleaned up by Florisil column chromatography. The derivatives are analyzed by gas chromatography with flame ionization detection. | 0.001 | Y | Y | Y | 76 | 8/28/2014 | 8/15/2017 | 8/15/2017 4:37:09 PM | 1996-04-12 | 6/3/1998 6:51:36 PM | EMS |  | Concentration | mg/L | Y | (fl. conc.) |  | Behenic (fl. conc.) |  | Organics; | Organics; EMS code: FA08 | Chemical |  | LAB
FA08 | mg/L | Behenic | P035 | DCM-DIAZO-FLOR | DCM-DIAZO-FLOR | 0.0005 | Y | Y | Y | 22 |  |  | 10/24/1996 12:44:23 PM | 1996-04-12 | 6/3/1998 6:51:36 PM | EMS |  | Concentration | mg/L | Y | (fl. conc.) |  | Behenic (fl. conc.) |  | Organics; | Organics; EMS code: FA08 | Chemical |  | LAB
FA08 | mg/L | Behenic | X239 | Gas Chromatography | Gas Chromatography | 0.001 | Y | Y | Y | 12 |  |  | 10/23/1996 11:56:36 PM | 1996-04-12 | 6/3/1998 6:51:45 PM | EMS |  | Concentration | mg/L | Y | (fl. conc.) |  | Behenic (fl. conc.) |  | Organics; | Organics; EMS code: FA08 | Chemical |  | LAB
FA08 | ug/g | Behenic | MP02 | GC/MS | GC/MS | 0.05 | Y | Y | Y | 2 | 6/21/2002 12:00:00 PM | 6/21/2002 12:25:00 PM | 3/27/2004 1:07:20 AM | 1996-04-12 | 6/3/1998 6:51:42 PM | EMS |  | MassPerMass | mg/kg | Y | (mass conc.) |  | Behenic (mass conc.) |  | Organics; | Organics; EMS code: FA08 | Chemical |  | LAB
FA08 | mg/L | Behenic | E677 | Resins & Fatty Acids are methylated then analyzed by GCMS | EPA 8270E; AB Env. 129; NCASI RA/FA-85.02 (mod) Resins & Fatty Acids are methylated then analyzed by GCMS. Total fatty acids is the sum of arachidic acid; behenic acid; lauric acid; lignoceric acid; linoleic acid; linolenic acid; myristic acid; oleic acid; palmitic acid; and stearic acid. Total resin acids is the sum of abietic acid; 12-chlorodehydroabietic acid; 14-chlorodehydroabietic acid; dehydroabietic acid; 12;14-dichlorodehydroabietic acid; isopimaric acid; levopimaric acid; neoabietic acid; palustric acid; pimaric acid and sandaracopimaric acid. Total resin acids (BC WWQG) is the sum of abietic acid; neoabietic acid; pimaric acid; isopimaric acid; and sandaracopimaric acid. | 0.005 | Y | Y | Y | 12 | 11/28/2022 | 6/15/2024 | 7/8/2024 11:10:02 AM | 2021-12-30 |  | EMS |  | Concentration | mg/L | Y | (fl. conc.) |  | Behenic (fl. conc.) |  | Organics; | Organics; EMS code: FA08 | Chemical |  | LAB
FA09 | mg/L | Lignoceric | E677 | Resins & Fatty Acids are methylated then analyzed by GCMS | EPA 8270E; AB Env. 129; NCASI RA/FA-85.02 (mod) Resins & Fatty Acids are methylated then analyzed by GCMS. Total fatty acids is the sum of arachidic acid; behenic acid; lauric acid; lignoceric acid; linoleic acid; linolenic acid; myristic acid; oleic acid; palmitic acid; and stearic acid. Total resin acids is the sum of abietic acid; 12-chlorodehydroabietic acid; 14-chlorodehydroabietic acid; dehydroabietic acid; 12;14-dichlorodehydroabietic acid; isopimaric acid; levopimaric acid; neoabietic acid; palustric acid; pimaric acid and sandaracopimaric acid. Total resin acids (BC WWQG) is the sum of abietic acid; neoabietic acid; pimaric acid; isopimaric acid; and sandaracopimaric acid. | 0.005 | Y | Y | Y | 12 | 11/28/2022 | 6/15/2024 | 7/8/2024 11:10:02 AM | 2021-12-30 |  | EMS |  | Concentration | mg/L | Y | (fl. conc.) |  | Lignoceric (fl. conc.) |  | Organics; | Organics; EMS code: FA09 | Chemical |  | LAB
FA09 | mg/L | Lignoceric | X239 | Gas Chromatography | Gas Chromatography | 0.001 | Y | Y | Y | 12 |  |  | 10/23/1996 11:56:36 PM | 1996-04-12 | 6/3/1998 6:51:45 PM | EMS |  | Concentration | mg/L | Y | (fl. conc.) |  | Lignoceric (fl. conc.) |  | Organics; | Organics; EMS code: FA09 | Chemical |  | LAB
FA09 | mg/L | Lignoceric | P030 | Resin Acids | The water sample is acidified and extracted with dichloro-methane. The raw extract is concentrated and derivatized with diazomethane to produce the corresponding methyl ester derivatives. (Suitable substitute derivatives may be produced by other available techniques.) If required, the extracts are cleaned up by Florisil column chromatography. The derivatives are analyzed by gas chromatography with flame ionization detection. | 0.001 | Y | Y | Y | 80 | 8/28/2014 | 8/15/2017 | 8/15/2017 4:37:09 PM | 1996-04-12 | 6/3/1998 6:51:36 PM | EMS |  | Concentration | mg/L | Y | (fl. conc.) |  | Lignoceric (fl. conc.) |  | Organics; | Organics; EMS code: FA09 | Chemical |  | LAB
FA09 | ug/g | Lignoceric | MP02 | GC/MS | GC/MS | 0.05 | Y | Y | Y | 2 | 6/21/2002 12:00:00 PM | 6/21/2002 12:25:00 PM | 3/27/2004 1:07:20 AM | 1996-04-12 | 6/3/1998 6:51:42 PM | EMS |  | MassPerMass | mg/kg | Y | (mass conc.) |  | Lignoceric (mass conc.) |  | Organics; | Organics; EMS code: FA09 | Chemical |  | LAB
FA09 | mg/L | Lignoceric | P035 | DCM-DIAZO-FLOR | DCM-DIAZO-FLOR | 0.0005 | Y | Y | Y | 22 |  |  | 10/24/1996 12:44:23 PM | 1996-04-12 | 6/3/1998 6:51:36 PM | EMS |  | Concentration | mg/L | Y | (fl. conc.) |  | Lignoceric (fl. conc.) |  | Organics; | Organics; EMS code: FA09 | Chemical |  | LAB
FA09 | mg/L | Lignoceric | F045 | pH2,DEE extr'n,conc,MTBEextr'n,methyl'n,HRGC/LMRS | Effluent samples are first adjusted to pH 2 and then extracted with Diethyl Ether. The extracts are rotary evaporated and the resulting residues, containing the Fatty Acids and other solvent soluble organics, are then redissolved in Methyl-t-butyl Ether (MTBE), transferred to small glass vials, methylated with Diazomethane, concentrated and cleaned up. The extract is injected on a High Resolution Gas Chromatograph/Low Resolution Mass Spectrometer (HRGC/LRMS) system for analysis. LRMS data is acquired in Total Ion Mode (TIM) and quantitation is performed using selected mass ions. Positive identification of the various Fatty Acids is made on the basis of a comparison of absolute retention times to those of the external standards and computer based mass spectral mass/intensity library matching routines. | 0.001 | Y | Y | Y | 71 | 3/18/1999 2:46:00 PM | 5/13/2003 | 4/1/2004 1:08:46 AM | 1997-10-20 | 6/3/1998 6:51:47 PM | NPEPPIN |  | Concentration | mg/L | Y | (fl. conc.) |  | Lignoceric (fl. conc.) |  | Organics; | Organics; EMS code: FA09 | Chemical |  | LAB
FA10 | mg/L | Linolenic Acid | F045 | pH2,DEE extr'n,conc,MTBEextr'n,methyl'n,HRGC/LMRS | Effluent samples are first adjusted to pH 2 and then extracted with Diethyl Ether. The extracts are rotary evaporated and the resulting residues, containing the Fatty Acids and other solvent soluble organics, are then redissolved in Methyl-t-butyl Ether (MTBE), transferred to small glass vials, methylated with Diazomethane, concentrated and cleaned up. The extract is injected on a High Resolution Gas Chromatograph/Low Resolution Mass Spectrometer (HRGC/LRMS) system for analysis. LRMS data is acquired in Total Ion Mode (TIM) and quantitation is performed using selected mass ions. Positive identification of the various Fatty Acids is made on the basis of a comparison of absolute retention times to those of the external standards and computer based mass spectral mass/intensity library matching routines. | 0.001 | Y | Y | Y | 69 | 1/31/2002 | 5/13/2003 | 4/1/2004 1:08:46 AM | 1997-10-20 | 6/3/1998 6:51:47 PM | NPEPPIN |  | Concentration | mg/L | Y | (fl. conc.) |  | Linolenic Acid (fl. conc.) |  | Organics; | Organics; EMS code: FA10 | Chemical |  | LAB
FA10 | mg/L | Linolenic Acid | P030 | Resin Acids | The water sample is acidified and extracted with dichloro-methane. The raw extract is concentrated and derivatized with diazomethane to produce the corresponding methyl ester derivatives. (Suitable substitute derivatives may be produced by other available techniques.) If required, the extracts are cleaned up by Florisil column chromatography. The derivatives are analyzed by gas chromatography with flame ionization detection. | 0.001 | Y | Y | Y | 79 | 8/28/2014 | 8/15/2017 | 6/7/2024 5:30:18 PM | 1996-04-12 | 6/3/1998 6:51:36 PM | EMS |  | Concentration | mg/L | Y | (fl. conc.) |  | Linolenic Acid (fl. conc.) |  | Organics; | Organics; EMS code: FA10 | Chemical |  | LAB
FA10 | mg/L | Linolenic Acid | P035 | DCM-DIAZO-FLOR | DCM-DIAZO-FLOR | 0.0005 | Y | Y | Y | 22 |  |  | 10/24/1996 12:44:24 PM | 1996-04-12 | 6/3/1998 6:51:36 PM | EMS |  | Concentration | mg/L | Y | (fl. conc.) |  | Linolenic Acid (fl. conc.) |  | Organics; | Organics; EMS code: FA10 | Chemical |  | LAB
FA10 | mg/L | Linolenic Acid | RFAS | Separation of resin and fatty acids (RFAs) in water by GC/MS | Scientific Principle The separation of resin and fatty acids (RFAs) in water samples is performed using gas chromatography/mass spectrometry (GC/MS) in the Selected Ion Monitoring (SIM) mode. The separation is based on compound size; boiling point and polarity. Analyte concentration is calculated using the internal standard method. RFA compounds are extracted from water samples using a liquid-liquid solvent extraction with DCM. Extraction occurs after the sample is brought to a pH of ~ 2 and placed in a rotary tumbler (at approx. 40 rpm) for 1 hour. The resulting extract is transferred to a vial; dried with acidified baked sodium sulfate; methylated with diazomethane and concentrated to 1mL under nitrogen prior to analysis. The resulting extracts are vialed and fortified with labeled internal standard compounds prior to analysis by gas chromatography with a (50%-Phenyl)-methylpolysiloxane (DB-17) column and a mass spectrometer detector in SIM Mode. | 0.006 | Y | Y | Y | 1 | 9/17/2021 7:49:00 PM | 9/17/2021 7:49:00 PM | 2/1/2022 2:38:14 PM | 2022-01-19 |  | EMS |  | Concentration | mg/L | Y | (fl. conc.) |  | Linolenic Acid (fl. conc.) |  | Organics; | Organics; EMS code: FA10 | Chemical |  | LAB
FA10 | ug/g | Linolenic Acid | MP02 | GC/MS | GC/MS | 0.05 | Y | Y | Y | 2 | 6/21/2002 12:00:00 PM | 6/21/2002 12:25:00 PM | 3/27/2004 1:07:20 AM | 1996-04-12 | 6/3/1998 6:51:42 PM | EMS |  | MassPerMass | mg/kg | Y | (mass conc.) |  | Linolenic Acid (mass conc.) |  | Organics; | Organics; EMS code: FA10 | Chemical |  | LAB
FA10 | mg/L | Linolenic Acid | E677 | Resins & Fatty Acids are methylated then analyzed by GCMS | EPA 8270E; AB Env. 129; NCASI RA/FA-85.02 (mod) Resins & Fatty Acids are methylated then analyzed by GCMS. Total fatty acids is the sum of arachidic acid; behenic acid; lauric acid; lignoceric acid; linoleic acid; linolenic acid; myristic acid; oleic acid; palmitic acid; and stearic acid. Total resin acids is the sum of abietic acid; 12-chlorodehydroabietic acid; 14-chlorodehydroabietic acid; dehydroabietic acid; 12;14-dichlorodehydroabietic acid; isopimaric acid; levopimaric acid; neoabietic acid; palustric acid; pimaric acid and sandaracopimaric acid. Total resin acids (BC WWQG) is the sum of abietic acid; neoabietic acid; pimaric acid; isopimaric acid; and sandaracopimaric acid. | 0.005 | Y | Y | Y | 12 | 11/28/2022 | 6/15/2024 | 7/8/2024 11:10:02 AM | 2021-12-30 |  | EMS |  | Concentration | mg/L | Y | (fl. conc.) |  | Linolenic Acid (fl. conc.) |  | Organics; | Organics; EMS code: FA10 | Chemical |  | LAB
FA11 | mg/L | Oleic Acid | P035 | DCM-DIAZO-FLOR | DCM-DIAZO-FLOR | 0.0005 | Y | Y | Y | 22 |  |  | 10/24/1996 12:44:24 PM | 1996-04-12 | 6/3/1998 6:51:45 PM | EMS |  | Concentration | mg/L | Y | (fl. conc.) |  | Oleic Acid (fl. conc.) |  | Organics; | Organics; EMS code: FA11 | Chemical |  | LAB
FA11 | mg/L | Oleic Acid | F045 | pH2,DEE extr'n,conc,MTBEextr'n,methyl'n,HRGC/LMRS | Effluent samples are first adjusted to pH 2 and then extracted with Diethyl Ether. The extracts are rotary evaporated and the resulting residues, containing the Fatty Acids and other solvent soluble organics, are then redissolved in Methyl-t-butyl Ether (MTBE), transferred to small glass vials, methylated with Diazomethane, concentrated and cleaned up. The extract is injected on a High Resolution Gas Chromatograph/Low Resolution Mass Spectrometer (HRGC/LRMS) system for analysis. LRMS data is acquired in Total Ion Mode (TIM) and quantitation is performed using selected mass ions. Positive identification of the various Fatty Acids is made on the basis of a comparison of absolute retention times to those of the external standards and computer based mass spectral mass/intensity library matching routines. | 0.001 | Y | Y | Y | 71 | 3/18/1999 2:46:00 PM | 5/13/2003 | 4/1/2004 1:08:46 AM | 1997-10-20 | 6/3/1998 6:51:47 PM | NPEPPIN |  | Concentration | mg/L | Y | (fl. conc.) |  | Oleic Acid (fl. conc.) |  | Organics; | Organics; EMS code: FA11 | Chemical |  | LAB
FA11 | ug/g | Oleic Acid | MP02 | GC/MS | GC/MS | 0.05 | Y | Y | Y | 2 | 6/21/2002 12:00:00 PM | 6/21/2002 12:25:00 PM | 3/27/2004 1:07:20 AM | 1996-04-12 | 6/3/1998 6:51:42 PM | EMS |  | MassPerMass | mg/kg | Y | (mass conc.) |  | Oleic Acid (mass conc.) |  | Organics; | Organics; EMS code: FA11 | Chemical |  | LAB
FA11 | mg/L | Oleic Acid | P030 | Resin Acids | The water sample is acidified and extracted with dichloro-methane. The raw extract is concentrated and derivatized with diazomethane to produce the corresponding methyl ester derivatives. (Suitable substitute derivatives may be produced by other available techniques.) If required, the extracts are cleaned up by Florisil column chromatography. The derivatives are analyzed by gas chromatography with flame ionization detection. | 0.001 | Y | Y | Y | 86 | 8/28/2014 | 9/17/2021 7:49:00 PM | 6/7/2024 5:30:18 PM | 1996-04-12 | 6/3/1998 6:51:36 PM | EMS |  | Concentration | mg/L | Y | (fl. conc.) |  | Oleic Acid (fl. conc.) |  | Organics; | Organics; EMS code: FA11 | Chemical |  | LAB
FA11 | mg/L | Oleic Acid | E677 | Resins & Fatty Acids are methylated then analyzed by GCMS | EPA 8270E; AB Env. 129; NCASI RA/FA-85.02 (mod) Resins & Fatty Acids are methylated then analyzed by GCMS. Total fatty acids is the sum of arachidic acid; behenic acid; lauric acid; lignoceric acid; linoleic acid; linolenic acid; myristic acid; oleic acid; palmitic acid; and stearic acid. Total resin acids is the sum of abietic acid; 12-chlorodehydroabietic acid; 14-chlorodehydroabietic acid; dehydroabietic acid; 12;14-dichlorodehydroabietic acid; isopimaric acid; levopimaric acid; neoabietic acid; palustric acid; pimaric acid and sandaracopimaric acid. Total resin acids (BC WWQG) is the sum of abietic acid; neoabietic acid; pimaric acid; isopimaric acid; and sandaracopimaric acid. | 0.005 | Y | Y | Y | 12 | 11/28/2022 | 6/15/2024 | 7/8/2024 11:10:02 AM | 2021-12-30 |  | EMS |  | Concentration | mg/L | Y | (fl. conc.) |  | Oleic Acid (fl. conc.) |  | Organics; | Organics; EMS code: FA11 | Chemical |  | LAB
FA12 | mg/L | Palustric Acid | RFAS | Separation of resin and fatty acids (RFAs) in water by GC/MS | Scientific Principle The separation of resin and fatty acids (RFAs) in water samples is performed using gas chromatography/mass spectrometry (GC/MS) in the Selected Ion Monitoring (SIM) mode. The separation is based on compound size; boiling point and polarity. Analyte concentration is calculated using the internal standard method. RFA compounds are extracted from water samples using a liquid-liquid solvent extraction with DCM. Extraction occurs after the sample is brought to a pH of ~ 2 and placed in a rotary tumbler (at approx. 40 rpm) for 1 hour. The resulting extract is transferred to a vial; dried with acidified baked sodium sulfate; methylated with diazomethane and concentrated to 1mL under nitrogen prior to analysis. The resulting extracts are vialed and fortified with labeled internal standard compounds prior to analysis by gas chromatography with a (50%-Phenyl)-methylpolysiloxane (DB-17) column and a mass spectrometer detector in SIM Mode. | 0.006 | Y | Y | Y | 1 | 9/17/2021 7:49:00 PM | 9/17/2021 7:49:00 PM | 2/1/2022 2:38:14 PM | 2022-01-19 |  | EMS |  | Concentration | mg/L | Y | (fl. conc.) |  | Palustric Acid (fl. conc.) |  | Organics; | Organics; EMS code: FA12 | Chemical |  | LAB
FA12 | mg/L | Palustric Acid | E677 | Resins & Fatty Acids are methylated then analyzed by GCMS | EPA 8270E; AB Env. 129; NCASI RA/FA-85.02 (mod) Resins & Fatty Acids are methylated then analyzed by GCMS. Total fatty acids is the sum of arachidic acid; behenic acid; lauric acid; lignoceric acid; linoleic acid; linolenic acid; myristic acid; oleic acid; palmitic acid; and stearic acid. Total resin acids is the sum of abietic acid; 12-chlorodehydroabietic acid; 14-chlorodehydroabietic acid; dehydroabietic acid; 12;14-dichlorodehydroabietic acid; isopimaric acid; levopimaric acid; neoabietic acid; palustric acid; pimaric acid and sandaracopimaric acid. Total resin acids (BC WWQG) is the sum of abietic acid; neoabietic acid; pimaric acid; isopimaric acid; and sandaracopimaric acid. | 0.001 | Y | Y | Y | 12 | 11/28/2022 | 6/15/2024 | 7/8/2024 11:10:02 AM | 2021-12-30 |  | EMS |  | Concentration | mg/L | Y | (fl. conc.) |  | Palustric Acid (fl. conc.) |  | Organics; | Organics; EMS code: FA12 | Chemical |  | LAB
FA12 | mg/L | Palustric Acid | P030 | Resin Acids | The water sample is acidified and extracted with dichloro-methane. The raw extract is concentrated and derivatized with diazomethane to produce the corresponding methyl ester derivatives. (Suitable substitute derivatives may be produced by other available techniques.) If required, the extracts are cleaned up by Florisil column chromatography. The derivatives are analyzed by gas chromatography with flame ionization detection. | 0.001 | Y | Y | Y | 68 |  |  | 6/7/2024 5:30:18 PM | 1996-04-12 | 6/3/1998 6:51:36 PM | EMS |  | Concentration | mg/L | Y | (fl. conc.) |  | Palustric Acid (fl. conc.) |  | Organics; | Organics; EMS code: FA12 | Chemical |  | LAB
FA12 | mg/L | Palustric Acid | P035 | DCM-DIAZO-FLOR | DCM-DIAZO-FLOR | 0.0005 | Y | Y | Y | 22 |  |  | 10/24/1996 12:44:24 PM | 1996-04-12 | 6/3/1998 6:51:36 PM | EMS |  | Concentration | mg/L | Y | (fl. conc.) |  | Palustric Acid (fl. conc.) |  | Organics; | Organics; EMS code: FA12 | Chemical |  | LAB
FA12 | ug/g | Palustric Acid | FS05 | MeOH,DCM extr:HRGC/LRMS | Sediment samples are extracted by shaking with Methanol in a screw cap centrifuge bottle. The extracts are rotary evaporated and the resulting residues, containing the Resin Acids and other solvent soluble organics, are then redissolved in Methanol, transferred to small glass vials, methylated with Diazomethane then evaporated to dryness under Nitrogen. [Methylation Reaction R-COOH(Resin/Fatty Acids) + CH2N2 (Diazomethane) = R-COOCH3 (Methyl ester) + N2] The residue is re-dissolved in Methyl t-Butyl Ether (MTBE) and injected on a High Resolution Gas Chromatograph / Low Resolution Mass Spectrometer (HRGC/LRMS) system for analysis. LRMS data is acquired in Total Ion Mode (TIM) and quantitation is performed using selected mass ions. Positive identification of the various Resin Acids is made on the basis of a comparison of absolute retention times to those of the external standards and computer based mass spectral mass/intensity library matching routines. | 0.2 | Y | Y | Y | 3 | 9/14/2001 3:29:00 PM | 12/13/2002 12:55:00 PM | 3/27/2004 1:09:20 AM | 1996-09-16 | 6/3/1998 6:51:44 PM | MPAK |  | MassPerMass | mg/kg | Y | (mass conc.) |  | Palustric Acid (mass conc.) |  | Organics; | Organics; EMS code: FA12 | Chemical |  | LAB
FA12 | mg/L | Palustric Acid | F009 | pH 9,Diethyl Ether:HRGC/LRMS | Effluent samples are first adjusted to pH 9 and then extracted with Diethyl Ether. The extracts are rotary evaporated and the resulting residues, containing the Resin Acids and other solvent soluble organics are then redissolved in Methanol, transferred to small glass vials, methylated with Diazomethane then evaporated to dryness under Nitrogen. [Methylation Reaction R-COOH(Resin/Fatty Acids) + CH2N2(Diazomethane) = R-COOCH3(Methyl ester) + N2) The residue is re-dissolved in Methyl t-Butyl Ether (MTBE) and injected on a High Resolution Gas Chromatograph/Low Resolution Mass Spectrometer (HRGC/LRMS) system for analysis. LRMS data is acquired in Total Ion Mode (TIM) and quantitation is performed using selected mass ions. Positive identification of the various Resin Acids is made on the basis of a comparison of absolute retention times to those of the external standards and computer based mass spectral mass/intensity library matching routines. | 0.005 | Y | Y | Y | 193 | 12/27/1996 4:16:00 PM | 5/13/2003 | 4/1/2004 1:08:46 AM | 1996-09-16 | 6/3/1998 6:51:44 PM | MPAK |  | Concentration | mg/L | Y | (fl. conc.) |  | Palustric Acid (fl. conc.) |  | Organics; | Organics; EMS code: FA12 | Chemical |  | LAB
FA56 | ug/g | dichlorostearic acid; 9;10- | MP02 | GC/MS | GC/MS | 0.05 | Y | Y | Y | 76 | 1/31/2002 | 5/13/2003 | 4/1/2004 1:08:46 AM | 1996-04-12 | 6/3/1998 6:51:42 PM | EMS |  | MassPerMass | mg/kg | Y | (mass conc.) |  | dichlorostearic acid; 9;10- (mass conc.) |  | Organics; | Organics; EMS code: FA56 | Chemical |  | LAB
FA56 | mg/L | dichlorostearic acid; 9;10- | E677 | Resins & Fatty Acids are methylated then analyzed by GCMS | EPA 8270E; AB Env. 129; NCASI RA/FA-85.02 (mod) Resins & Fatty Acids are methylated then analyzed by GCMS. Total fatty acids is the sum of arachidic acid; behenic acid; lauric acid; lignoceric acid; linoleic acid; linolenic acid; myristic acid; oleic acid; palmitic acid; and stearic acid. Total resin acids is the sum of abietic acid; 12-chlorodehydroabietic acid; 14-chlorodehydroabietic acid; dehydroabietic acid; 12;14-dichlorodehydroabietic acid; isopimaric acid; levopimaric acid; neoabietic acid; palustric acid; pimaric acid and sandaracopimaric acid. Total resin acids (BC WWQG) is the sum of abietic acid; neoabietic acid; pimaric acid; isopimaric acid; and sandaracopimaric acid. | 0.001 | Y | Y | Y | 12 | 11/28/2022 | 6/15/2024 | 7/8/2024 11:10:02 AM | 2021-12-30 |  | EMS |  | Concentration | mg/L | Y | (fl. conc.) |  | dichlorostearic acid; 9;10- (fl. conc.) |  | Organics; | Organics; EMS code: FA56 | Chemical |  | LAB
FA56 | mg/L | dichlorostearic acid; 9;10- | RFAS | Separation of resin and fatty acids (RFAs) in water by GC/MS | Scientific Principle The separation of resin and fatty acids (RFAs) in water samples is performed using gas chromatography/mass spectrometry (GC/MS) in the Selected Ion Monitoring (SIM) mode. The separation is based on compound size; boiling point and polarity. Analyte concentration is calculated using the internal standard method. RFA compounds are extracted from water samples using a liquid-liquid solvent extraction with DCM. Extraction occurs after the sample is brought to a pH of ~ 2 and placed in a rotary tumbler (at approx. 40 rpm) for 1 hour. The resulting extract is transferred to a vial; dried with acidified baked sodium sulfate; methylated with diazomethane and concentrated to 1mL under nitrogen prior to analysis. The resulting extracts are vialed and fortified with labeled internal standard compounds prior to analysis by gas chromatography with a (50%-Phenyl)-methylpolysiloxane (DB-17) column and a mass spectrometer detector in SIM Mode. | 0.006 | Y | Y | Y | 1 | 9/17/2021 7:49:00 PM | 9/17/2021 7:49:00 PM | 2/1/2022 2:38:14 PM | 2022-01-19 |  | EMS |  | Concentration | mg/L | Y | (fl. conc.) |  | dichlorostearic acid; 9;10- (fl. conc.) |  | Organics; | Organics; EMS code: FA56 | Chemical |  | LAB
FBEN | % | Fluorobenzene (SS) | F007 | Static Headspace:HRGC/LRMS | Static headspace analysis is a method for the determination of volatile compounds in liquids and/or solids by analysing the vapour phase of a sample in a closed system; in this case a sealed vial. Static Headspace is sometimes referred to as Equilibrium Headspace because the vapour phase above the sample in a sealed vial is brought into thermodynamic equilibrium with the sample prior to analysis. Volatile organic compounds are thermally extracted from the sample matrix by Static Headspace sampling and introduced directly to a thick-film megabore GC column. Compounds eluting from the GC column are introduced into a Mass Spectrometer and target compounds are identified by a comparison of both retention times and mass spectra with those of reference standards stored in a computerized data system. Reference retention times and mass spectra for analytes are obtained by the measurement of calibration standards under precisely the same extraction conditions used for samples. The concentration of each identified component is measured by relating the response of the sample quantitation mass ions to the response of the quantitation mass ions of a series of external standards. A deuterated surrogate recovery standard is added to provide information on matrix effects. | 1 | Y | Y | Y | 35 | 3/20/2007 | 6/7/2010 | 7/9/2012 2:09:48 PM | 2012-06-25 |  | RDGIBSON |  | DimensionlessRatio | % | N | (ratio) |  | Fluorobenzene (SS) (ratio) |  |  | EMS code: FBEN | Chemical |  | LAB
FBPH | % | Fluorobiphenyl; 2- | SURR | Surrogate Recovery | Surrogate Recovery | 0.1 | Y | Y | Y | 164 | 12/7/2021 | 6/4/2024 | 7/4/2024 11:06:09 AM | 2021-12-14 | 12/30/2021 5:18:52 PM | RDGIBSON | RDGIBSON | DimensionlessRatio | % | N | (ratio) |  | Fluorobiphenyl; 2- (ratio) |  |  | EMS code: FBPH | Chemical |  | SURROGATE_RESULT
FDST | mg/dm2/d | Fixed Dustfall | X192 | Gravimetric (Sol+Insol) | Gravimetric (Sol+Insol) |  | Y | Y | Y | 1421 | 9/15/2015 4:14:00 PM | 8/19/2024 | 8/19/2024 4:02:08 PM | 2008-06-04 |  | RDGIBSON |  | MassFlux | mg/dm²/d | Y | (mass flux) |  | Fixed Dustfall (mass flux) |  | Particulates; | Particulates; EMS code: FDST | Chemical |  | LAB
FDST | mg/dm2/d | Fixed Dustfall | DUST | HNO3 dig.(microwave or hot block),ICPMS | HNO3 dig.(microwave or hot block);ICPMS |  | Y | Y | Y | 13 | 1/31/2020 | 1/31/2020 | 1/16/2023 3:51:10 PM | 2008-02-07 |  | RDGIBSON |  | MassFlux | mg/dm²/d | Y | (mass flux) |  | Fixed Dustfall (mass flux) |  | Particulates; | Particulates; EMS code: FDST | Chemical |  | LAB
FE-D | mg/L | Iron Dissolved | F032 | Filt.,Acid,ICPMS | There is no acid digestion required for this analysis. A water sample is collected and preserved in 0.2 percent nitric acid. To analyse the extractable metals the sample is shaken, allowed to sit for at least twenty four hours, and a decanted portion is introduced to the spectrometer. When the dissolved metals are required, the samples must first be filtered through a 0.45 micron cellulose acetate filter paper and then followed by preservation in 0.2 percent nitric acid. It is important to note that different laboratories may have very different detection limits for some variables analyzed by ICPMS. | 0.01 | Y | Y | Y | 69669 | 8/16/1996 2:35:00 PM | 8/26/2024 | 8/26/2024 4:53:05 PM | 1997-01-16 | 6/3/1998 6:51:46 PM | MPAK |  | Concentration | mg/L | Y | (fl. conc.) |  | Iron Dissolved (fl. conc.) |  | Metals - Dissolved; | Metals - Dissolved; EMS code: FE-D | Chemical | Dissolved | LAB
FE-D | mg/L | Iron Dissolved | X356 | LF,LA:HNO3:ICP | Samples are lab-filtered (0.45 um); lab-acidified with HNO3; analysis by ICP/AES. | 0.001 | Y | Y | Y | 125 | 10/17/2000 | 5/1/2002 | 1/20/2004 12:07:44 PM | 2001-02-12 |  | NPEPPIN |  | Concentration | mg/L | Y | (fl. conc.) |  | Iron Dissolved (fl. conc.) |  | Metals - Dissolved; | Metals - Dissolved; EMS code: FE-D | Chemical | Dissolved | LAB
FE-D | mg/L | Iron Dissolved | F023 | HNO3,GFAA | Dissolved and extractable water samples are ready for analysis once collected. Total samples are autoclaved with HNO3 and hydrochloric acid (HCl) in capped Teflon vials. Analysis is performed via Inductively Coupled Argon Plasma Atomic Emission Spectrometry (ICP-AES) (simultaneous multielement analysis), Graphite Furnace Atomic Absorption Spectrometry (GF-AAS) (low level analysis) and/or Flame Atomic Absorption Spectrometry (F-AAS). Note that extractable samples are acidified with 0.5 mL HNO3 per 100 mL of sample at time of sample collection. | 0.001 | Y | Y | Y | 78 |  |  | 2/19/2002 10:02:55 AM | 2001-02-12 |  | NPEPPIN |  | Concentration | mg/L | Y | (fl. conc.) |  | Iron Dissolved (fl. conc.) |  | Metals - Dissolved; | Metals - Dissolved; EMS code: FE-D | Chemical | Dissolved | LAB
FE-D | mg/L | Iron Dissolved | X085 | LF,LA,HNO3,AA | Samples are lab-filtered (0.45 um), and lab-acidified (HNO3); analysis by AA. | 0.005 | Y | Y | Y | 1331 | 12/12/1996 | 9/18/2015 | 10/28/2015 9:54:16 AM | 1996-04-12 | 6/3/1998 6:51:36 PM | EMS |  | Concentration | mg/L | Y | (fl. conc.) |  | Iron Dissolved (fl. conc.) |  | Metals - Dissolved; | Metals - Dissolved; EMS code: FE-D | Chemical | Dissolved | LAB
FE-D | mg/L | Iron Dissolved | 0020 | HNO3 Dig:AA Analysis | HNO3 Dig:AA Analysis | 0.1 | Y | Y | Y | 5 |  |  | 10/24/1996 11:30:41 AM | 1996-04-12 | 6/3/1998 6:51:36 PM | EMS |  | Concentration | mg/L | Y | (fl. conc.) |  | Iron Dissolved (fl. conc.) |  | Metals - Dissolved; | Metals - Dissolved; EMS code: FE-D | Chemical | Dissolved | LAB
FE-D | mg/L | Iron Dissolved | 0050 | Graphite Furnace AA | Graphite Furnace AA | 0.005 | Y | Y | Y | 243 |  |  | 10/24/1996 12:07:14 PM | 1996-04-12 | 6/3/1998 6:51:36 PM | EMS |  | Concentration | mg/L | Y | (fl. conc.) |  | Iron Dissolved (fl. conc.) |  | Metals - Dissolved; | Metals - Dissolved; EMS code: FE-D | Chemical | Dissolved | LAB
FE-D | mg/L | Iron Dissolved | X357 | LF,LA:HNO3:GFAA | Samples are lab-filtered (0.45 um) and lab-acidified with HNO3; analysis by GFAA. | 0.001 | Y | Y | Y | 9 |  |  | 10/23/1996 9:26:37 PM | 1996-04-12 | 6/3/1998 6:51:36 PM | EMS |  | Concentration | mg/L | Y | (fl. conc.) |  | Iron Dissolved (fl. conc.) |  | Metals - Dissolved; | Metals - Dissolved; EMS code: FE-D | Chemical | Dissolved | LAB
FE-D | mg/L | Iron Dissolved | 5138 | AA Analysis | AA Analysis | 0.1 | Y | Y | Y | 25 |  |  | 10/24/1996 9:37:08 AM | 1996-04-12 | 6/3/1998 6:51:36 PM | EMS |  | Concentration | mg/L | Y | (fl. conc.) |  | Iron Dissolved (fl. conc.) |  | Metals - Dissolved; | Metals - Dissolved; EMS code: FE-D | Chemical | Dissolved | LAB
FE-D | mg/L | Iron Dissolved | X072 | FA,HNO3,Graph.Furnace,AA | Samples are field-acidified with HNO3; HNO3 digestion; analysis by GFAA. | 0.005 | Y | Y | Y | 880 | 9/29/2000 | 12/1/2010 | 12/9/2010 3:55:26 PM | 1996-04-12 | 6/3/1998 6:51:41 PM | EMS |  | Concentration | mg/L | Y | (fl. conc.) |  | Iron Dissolved (fl. conc.) |  | Metals - Dissolved; | Metals - Dissolved; EMS code: FE-D | Chemical | Dissolved | LAB
FE-D | mg/L | Iron Dissolved | X203 | FF:FA,HNO3,Dir. Flame,AA | Samples are field-filtered (0.45 um), and field-acidified (HNO3); analysis by AA. | 0.005 | Y | Y | Y | 337 |  |  | 10/24/1996 12:04:45 PM | 1996-04-12 | 6/3/1998 6:51:36 PM | EMS |  | Concentration | mg/L | Y | (fl. conc.) |  | Iron Dissolved (fl. conc.) |  | Metals - Dissolved; | Metals - Dissolved; EMS code: FE-D | Chemical | Dissolved | LAB
FE-D | mg/L | Iron Dissolved | METC | MET-D-NP-CCMS-VA | Ultra trace metals in water are analyzed by CRC ICPMS, based on US EPA Method 6020B (July 2014). This procedure is intended for pristine field-filtered acid-preserved water samples. The detection limits (LOR) for this test are based on lab instrumental analysis only, not including filtration. ALS-supplied field filtration equipment does not support these LOR. Therefore, because of the high probability of false positives due to filtration, it is strongly recommended that a filtration blank be analysed to aid in data interpretation. | 0.001 | Y | Y | Y | 1776 | 4/3/2012 | 8/15/2024 | 8/23/2024 9:50:35 AM | 2019-08-20 |  | KATFRASE |  | Concentration | mg/L | Y | (fl. conc.) |  | Iron Dissolved (fl. conc.) |  | Metals - Dissolved; | Metals - Dissolved; EMS code: FE-D | Chemical | Dissolved | LAB
FE-D | mg/L | Iron Dissolved | X028 | FA,HNO3,AA Analysis | FA;HNO3;AA Analysis | 0.005 | Y | Y | Y | 1376 |  |  | 10/29/2001 11:34:42 AM | 1996-04-12 | 6/3/1998 6:51:36 PM | EMS |  | Concentration | mg/L | Y | (fl. conc.) |  | Iron Dissolved (fl. conc.) |  | Metals - Dissolved; | Metals - Dissolved; EMS code: FE-D | Chemical | Dissolved | LAB
FE-D | mg/L | Iron Dissolved | X116 | FF,FA,HNO3,Graph.Furn,AA | Samples are field-filtered (0.45 um), field-acidified with HNO3; analysis by GFAA. | 0.01 | Y | Y | Y | 7424 |  |  | 8/26/2024 2:41:32 PM | 1996-04-12 | 6/3/1998 6:51:36 PM | EMS |  | Concentration | mg/L | Y | (fl. conc.) |  | Iron Dissolved (fl. conc.) |  | Metals - Dissolved; | Metals - Dissolved; EMS code: FE-D | Chemical | Dissolved | LAB
FE-D | mg/L | Iron Dissolved | F039 | Field filt + acid, ICP-AES | Dissolved and extractable water samples are ready for analysis once collected. Total samples are autoclaved with HNO3 and hydrochloric acid (HCI) in capped Teflon vials. Analysis is performed via Inductively Coupled Argon Plasma Atomic Emission Spectrometry (ICP-AES)(simultaneous multielement analysis), Graphite Furnace Atomic Absorption Spectrometry (GF-AAS)(low level analysis) and/or Flame Atomic Absorption Spectrometry (F-AAS). | 0.005 | Y | Y | Y | 8587 | 7/29/1996 6:03:00 PM | 7/5/2019 1:21:00 PM | 5/16/2022 4:21:33 PM | 1997-05-08 | 6/3/1998 6:51:46 PM | MPAK |  | Concentration | mg/L | Y | (fl. conc.) |  | Iron Dissolved (fl. conc.) |  | Metals - Dissolved; | Metals - Dissolved; EMS code: FE-D | Chemical | Dissolved | LAB
FE-D | mg/L | Iron Dissolved | X074 | FA,HNO3,Plasma Spec. | FA;HNO3;Plasma Spec. | 0.005 | Y | Y | Y | 7528 |  |  | 8/28/2018 3:14:04 PM | 1996-04-12 | 6/3/1998 6:51:36 PM | EMS |  | Concentration | mg/L | Y | (fl. conc.) |  | Iron Dissolved (fl. conc.) |  | Metals - Dissolved; | Metals - Dissolved; EMS code: FE-D | Chemical | Dissolved | LAB
FE-D | mg/L | Iron Dissolved | X095 | FF,FA,HNO3,ICP | FF;FA;HNO3;ICP | 0.005 | Y | Y | Y | 786 |  |  | 10/30/2014 8:31:17 AM | 1996-04-12 | 6/3/1998 6:51:41 PM | EMS |  | Concentration | mg/L | Y | (fl. conc.) |  | Iron Dissolved (fl. conc.) |  | Metals - Dissolved; | Metals - Dissolved; EMS code: FE-D | Chemical | Dissolved | LAB
FE-D | mg/L | Iron Dissolved | EQ01 | EQUIS (mg/L) | EQUIS (mg/L) | 0.005 | N | Y | Y | 7222 |  |  | 10/24/1996 3:16:55 PM | 1996-04-12 | 6/3/1998 6:51:36 PM | EMS |  | Concentration | mg/L | Y | (fl. conc.) |  | Iron Dissolved (fl. conc.) |  | Metals - Dissolved; | Metals - Dissolved; EMS code: FE-D | Chemical | Dissolved | LAB
FE-D | mg/L | Iron Dissolved | X562 | Dissolved metals: FF, FA, HNO3, HR-ICPMS | Ultra trace metals in water are analyzed by high resolution inductively coupled plasma mass spectrometry (HR-ICPMS), based on US EPA Method 200.8, (Oct. 1999). This procedure is intended for pristine field-filtered acid-preserved water samples. The detection limits (LOR) for this test are based on lab instrumental analysis only, not including filtration. ALS-supplied field filtration equipment does not support these LOR. Therefore, because of the high probability of false positives due to filtration, it is strongly recommended that a filtration blank be analyzed to aid in data interpretation. | 0.001 | Y | Y | Y | 1270 | 6/8/2015 | 10/19/2023 | 11/9/2023 7:41:27 AM | 2015-10-22 |  | RDGIBSON |  | Concentration | mg/L | Y | (fl. conc.) |  | Iron Dissolved (fl. conc.) |  | Metals - Dissolved; | Metals - Dissolved; EMS code: FE-D | Chemical | Dissolved | LAB
FE-D | mg/L | Iron Dissolved | 0040 | HNO3 Dig: ICP Analysis | HNO3 Dig: ICP Analysis | 0.01 | Y | Y | Y | 726 | 5/9/2003 9:40:00 AM | 9/16/2003 3:18:00 PM | 7/18/2024 7:48:39 AM | 1996-04-12 | 6/3/1998 6:51:36 PM | EMS |  | Concentration | mg/L | Y | (fl. conc.) |  | Iron Dissolved (fl. conc.) |  | Metals - Dissolved; | Metals - Dissolved; EMS code: FE-D | Chemical | Dissolved | LAB
FE-D | mg/L | Iron Dissolved | 0090 | Ext.(APDC):AA Anal. | Ext.(APDC):AA Anal. | 0.005 | Y | Y | Y | 232 |  |  | 10/24/1996 11:35:00 AM | 1996-04-12 | 6/3/1998 6:51:36 PM | EMS |  | Concentration | mg/L | Y | (fl. conc.) |  | Iron Dissolved (fl. conc.) |  | Metals - Dissolved; | Metals - Dissolved; EMS code: FE-D | Chemical | Dissolved | LAB
FE-D | mg/L | Iron Dissolved | MWID | Metals in Water by ICPMS (digested) | Metals in Water by ICPMS (digested) Based on EPA 200.8 | 0.02 | Y | Y | Y | 31768 | 10/26/2018 5:20:00 PM | 10/26/2018 5:20:00 PM | 7/31/2024 1:46:41 PM | 2009-12-24 |  | RDGIBSON |  | Concentration | mg/L | Y | (fl. conc.) |  | Iron Dissolved (fl. conc.) |  | Metals - Dissolved; | Metals - Dissolved; EMS code: FE-D | Chemical | Dissolved | LAB
FE-D | mg/L | Iron Dissolved | X099 | FF,FA,HNO3,AA | FF;FA;HNO3;AA | 0.005 | Y | Y | Y | 676 |  |  | 6/5/2002 1:53:32 PM | 1996-04-12 | 6/3/1998 6:51:36 PM | EMS |  | Concentration | mg/L | Y | (fl. conc.) |  | Iron Dissolved (fl. conc.) |  | Metals - Dissolved; | Metals - Dissolved; EMS code: FE-D | Chemical | Dissolved | LAB
FE-D | mg/L | Iron Dissolved | MSEA | Metals in Seawater (undigested) by (CRC) ICP-MS | This method describes the multi-elemental determination of trace elements by Collision/Reaction Cell (CRC) ICP-MS. The method measures ions produced by a radio-frequency inductively coupled plasma. Analyte species originating in a liquid are pneumatically nebulized and the resulting aerosol transported by argon gas into the plasma torch. The ions produced are entrained in the plasma gas and extracted, by means of a differentially pumped vacuum interface, into a mass spectrometer. The ions flow through a Collision/Reaction Cell (CRC) where depending on the mode of analysis either the ions passes through the cell unhindered (no gas - classical ICP-MS mode), or collisions/reactions chemistry takes place between the ions and a neutral gas such as hydrogen or helium. Any polyatomic ions which give interference on the analytes are removed. The ions are then sorted according to their mass-to-charge ratios by a quadrupole mass spectrometer. The ions are then sorted according to their mass-to-charge ratios by a quadrupole mass spectrometer. The ions transmitted through the quadrupole are quantified by a channel electron multiplier and the ion information processed by a data handling system. Although most polyatomic interference is removed, any additional interference must be assessed and valid corrections applied or the data flagged to indicate problems. Interference correction must include compensation for background ions contributed by the plasma gas, reagents, and constituents of the sample matrix. Instrument drift as well as suppressions or enhancements of instrument response caused by the sample matrix must be corrected for by the use of internal standards. | 0.0004 | Y | Y | Y | 4 | 4/12/2023 | 4/29/2024 | 5/16/2024 2:00:14 PM | 2009-09-08 |  | RDGIBSON |  | Concentration | mg/L | Y | (fl. conc.) |  | Iron Dissolved (fl. conc.) |  | Metals - Dissolved; | Metals - Dissolved; EMS code: FE-D | Chemical | Dissolved | LAB
FE-D | mg/L | Iron Dissolved | X556 | Metals in water, dissolved, "in-bottle" digestion, ICP-SFMS | The whole water sample is filtered through a 0.45 micron cellulose acetate filter paper and the filtrate is stored in a HDPE/Teflon bottle. The container is de-capped in a Class 100 workstation, acidified to 2% (v/v) HNO3 and 1% (v/v) HCl and recapped. The sample is "in-bottle" digested in a convection oven, set at 60 degrees C for 16 hours. The sample is cooled and introduced to an inductively coupled argon plasma-sector field mass spectrometer (ICP-SFMS) by pneumatic nebulization. Each respective element is measured at a specific mass to charge ratio, m/z value expressed in amu (atomic mass units), and at a specified intrument resolution. LR: Low Resolution ~300, MR: Mid Resolution ~4000 and HR: High Resolution ~10000. Ref: Method number 02-2003, National Laboratory for Environmental Testing, CCIW, Environment Canada, Burlington, Requested April 2003. | 0.0005 | Y | Y | Y | 4234 | 6/11/2003 | 12/6/2023 2:04:00 AM | 2/14/2024 9:05:32 AM | 2004-01-19 |  | UMULLER |  | Concentration | mg/L | Y | (fl. conc.) |  | Iron Dissolved (fl. conc.) |  | Metals - Dissolved; | Metals - Dissolved; EMS code: FE-D | Chemical | Dissolved | LAB
FE-D | mg/L | Iron Dissolved | X205 | LF,AA,Dir Asp | LF;AA;Dir Asp | 0.1 | Y | Y | Y | 165 |  |  | 10/24/1996 9:16:37 AM | 1996-04-12 | 6/3/1998 6:51:41 PM | EMS |  | Concentration | mg/L | Y | (fl. conc.) |  | Iron Dissolved (fl. conc.) |  | Metals - Dissolved; | Metals - Dissolved; EMS code: FE-D | Chemical | Dissolved | LAB
FE-D | mg/L | Iron Dissolved | X285 | HNO3 Dig:ICP Analysis | HNO3 Dig:ICP Analysis | 0.03 | Y | Y | Y | 765 | 3/1/2000 | 11/13/2002 | 3/18/2004 10:28:34 AM | 1996-04-12 | 6/3/1998 6:51:36 PM | EMS |  | Concentration | mg/L | Y | (fl. conc.) |  | Iron Dissolved (fl. conc.) |  | Metals - Dissolved; | Metals - Dissolved; EMS code: FE-D | Chemical | Dissolved | LAB
FE-D | mg/L | Iron Dissolved | X170 | ICP | This code is a catch-all for a large variety of ICP methods including water and also solids digested into solutions as an intermediate result. Different procedures will have greatly different detection limits. (a) For air particulates, preliminary results are reported in mg/L and loading results as mg/dm2/d. (b) For Hi Vol samples preliminiary results are same as reported in mg/L and loading results as ug/m3. (c) Note that for Rain samples there is no filtration nor digestion steps. | 0.03 | Y | Y | Y | 20422 | 4/11/1996 6:01:00 PM | 9/23/2020 | 6/13/2024 3:58:21 PM | 1996-04-12 | 8/15/2019 1:18:05 PM | EMS | KATFRASE | Concentration | mg/L | Y | (fl. conc.) |  | Iron Dissolved (fl. conc.) |  | Metals - Dissolved; | Metals - Dissolved; EMS code: FE-D | Chemical | Dissolved | LAB
FE-D | mg/L | Iron Dissolved | X109 | LF, AA | LF; AA | 0.02 | Y | Y | Y | 1438 |  |  | 10/24/1996 2:20:41 PM | 1996-04-12 | 6/3/1998 6:51:36 PM | EMS |  | Concentration | mg/L | Y | (fl. conc.) |  | Iron Dissolved (fl. conc.) |  | Metals - Dissolved; | Metals - Dissolved; EMS code: FE-D | Chemical | Dissolved | LAB
FE-D | mg/L | Iron Dissolved | F049 | Filt,HNO3 dig,conc'n,nebul'ICP-OES | This method provides an automated procedure for the Inductively Coupled Argon Plasma Optical Emission Spectrophotometric (ICP-OES) determination and quantitation of trace amounts of metals in surface, ground and precipitation waters after a manual acid digestion and concentration procedure and ultrasonic nebulisation. Organically-bound metals and metal complexes are converted to metal ions by digestion with nitric acid. To increase the signal detected by the ICP, the samples are concentrated 10 times. The liquid sample is pumped onto the piezoelectric transducer, where the ultrasonic action efficiently converts it to a fine, dense aerosol. The nebulizer carrier gas flow then transports the wet aerosol through a heated U-tube, where the solvent is vaporized. Subsequently, most of the solvent vapour is removed from the aerosol stream by passing the stream through a chilled condenser. The result is a dry, analyte-laden aerosol which is introduced to the plasma. There, the high temperature desolvates the aerosol and dissociates the sample material into individual atoms and ions, which are excited to emit light at wave lengths characteristic of the elements present. Dissolved samples are field filtered through 0.45 micron cellulose acetate filter prior to acid preservation. Samples for extractable analysis are acidified with 2 mL concentrated HNO3 per litre and allowed to stand a minimum of 24 hours before analysis. Clients providing these samples are advised to preserve samples at the collection site. | 0.0004 | Y | Y | Y | 21524 | 3/10/1998 2:06:00 PM | 2/25/1999 2:07:00 PM | 7/23/2024 11:41:49 AM | 1997-10-20 | 6/3/1998 6:51:47 PM | NPEPPIN |  | Concentration | mg/L | Y | (fl. conc.) |  | Iron Dissolved (fl. conc.) |  | Metals - Dissolved; | Metals - Dissolved; EMS code: FE-D | Chemical | Dissolved | LAB
FE-E | mg/L | Iron Extractable | F069 | SEM extr,6N HCl,ICP/OES | 6.0 N Hydrochloric acid (HCI) is added to an aliquot of the sediment sample. The extract produced from the addition of the acid is then analyzed for simultaneously extracted metals (SEM) using inductively coupled plasma - optical emission spectrophotometry. The detection limits depend on the metal parameters in question and the instrumentation used. | 0.03 | Y | Y | Y | 1 |  |  | 11/25/2021 1:50:15 PM | 1999-07-09 |  | NPEPPIN |  | Concentration | mg/L | Y | (fl. conc.) |  | Iron Extractable (fl. conc.) |  | Metals - Extractable; | Metals - Extractable; EMS code: FE-E | Chemical | Extractable | LAB
FE-E | mg/L | Iron Extractable | X130 | Acid Extr,AA Dir. Asp. | Acid Extr;AA Dir. Asp. | 0.02 | Y | Y | Y | 876 |  |  | 8/23/1999 3:43:46 PM | 1996-04-12 | 6/3/1998 6:51:36 PM | EMS |  | Concentration | mg/L | Y | (fl. conc.) |  | Iron Extractable (fl. conc.) |  | Metals - Extractable; | Metals - Extractable; EMS code: FE-E | Chemical | Extractable | LAB
FE-E | mg/L | Iron Extractable | X553 | Metals in water, extractable, ICP-SFMS | The whole water sample is collected in a HDPE/Teflon bottle. In a Class 100 workstation it is preserved with concentrated Nitric Acid at concentration of 0.2 percent (v/v). The acid preservative extracts the metals from material in the sample and a decanted portion is introduced directly to an inductively coupled argon plasma-sector field mass spectrometer (ICP-SFMS) by pneumatic nebulization. Each respective element is measured at a specific mass to charge ratio, m/z value expressed in amu (atomic mass units), and at a specified instrument resolution. LR: Low Resolution ~300, MR: Mid Resolution ~4000 and HR: High Resolution ~10000. Ref: Method number 02-2061, National Laboratory for Environmental Testing, CCIW, Environment Canada, Burlington. Requested April 2003. | 0.0005 | Y | Y | Y | 2596 | 5/22/2003 | 9/21/2007 | 11/28/2023 2:11:53 PM | 2004-01-19 |  | UMULLER |  | Concentration | mg/L | Y | (fl. conc.) |  | Iron Extractable (fl. conc.) |  | Metals - Extractable; | Metals - Extractable; EMS code: FE-E | Chemical | Extractable | LAB
FE-E | mg/L | Iron Extractable | F040 | Field Acid, lab filt, ICP-AES | Dissolved and extractable water samples are ready for analysis once collected. Total samples are autoclaved with HNO3 and hydrochloric acid (HCI) in capped Teflon vials. Analysis is performed via Inductively Coupled Argon Plasma Atomic Emission Spectrometry (ICP-AES)(simultaneous multielement analysis), Graphite Furnace Atomic Absorption Spectrometry (GF-AAS)(low level analysis) and/or Flame Atomic Absorption Spectrometry (F-AAS). | 0.005 | Y | Y | Y | 2370 | 8/5/1997 6:06:00 PM | 3/23/2012 3:11:00 AM | 9/11/2019 8:47:41 AM | 1997-05-08 | 6/3/1998 6:51:46 PM | MPAK |  | Concentration | mg/L | Y | (fl. conc.) |  | Iron Extractable (fl. conc.) |  | Metals - Extractable; | Metals - Extractable; EMS code: FE-E | Chemical | Extractable | LAB
FE-E | mg/L | Iron Extractable | F050 | 24hr acid,filt,acid dig,conc'n,nebul.,ICP-OES | This method provides an automated procedure for the Inductively Coupled Argon Plasma Optical Emission Spectrophotometric (ICP-OES) determination and quantitation of trace amounts of metals in surface, ground and precipitation waters after a manual acid digestion and concentration procedure and ultrasonic nebulisation. Organically-bound metals and metal complexes are converted to metal ions by digestion with nitric acid. To increase the signal detected by the ICP, the samples are concentrated 10 times. The liquid sample is pumped onto the piezoelectric transducer, where the ultrasonic action efficiently converts it to a fine, dense aerosol. The nebulizer carrier gas flow then transports the wet aerosol through a heated U-tube, where the solvent is vaporized. Subsequently, most of the solvent vapour is removed from the aerosol stream by passing the stream through a chilled condenser. The result is a dry, analyte-laden aerosol which is introduced to the plasma. There, the high temperature desolvates the aerosol and dissociates the sample material into individual atoms and ions, which are excited to emit light at wave lengths characteristic of the elements present. Dissolved samples are field filtered through 0.45 micron cellulose acetate filter prior to acid preservation. Samples for extractable analysis are acidified with 2 mL concentrated HNO3 per litre and allowed to stand a minimum of 24 hours before analysis. Clients providing these samples are advised to preserve samples at the collection site. | 0.0004 | Y | Y | Y | 1 | 7/31/2004 10:58:00 AM | 7/31/2004 10:58:00 AM | 11/28/2005 10:12:20 AM | 1997-10-20 | 6/3/1998 6:51:47 PM | NPEPPIN |  | Concentration | mg/L | Y | (fl. conc.) |  | Iron Extractable (fl. conc.) |  | Metals - Extractable; | Metals - Extractable; EMS code: FE-E | Chemical | Extractable | LAB
FE-E | ug/g | Iron Extractable | X230 | Hom,HNO3/HCl,ICP | The sample is dried, ground to pass 100 mesh, and digested with a nitric-hydrochloric acid mixture to solubilize the solid matter and to remove the organic material by oxidation and volatilization. The completed digests are submitted to metals instrumentation personnel for metals analysis by Inductively Coupled Atomic Emission Spectrophotometry (ICP-AES), or hydride analysis by atomic absorption. | 5 | Y | Y | Y | 70 |  |  | 10/24/1996 2:25:27 PM | 1996-04-12 | 6/3/1998 6:51:36 PM | EMS |  | MassPerMass | mg/kg | Y | (mass conc.) |  | Iron Extractable (mass conc.) |  | Metals - Extractable; | Metals - Extractable; EMS code: FE-E | Chemical | Extractable | LAB
FE-I | % | Iron-Insoluble | EQ34 | EQUIS (%) | EQUIS (%) | 0.01 | N | Y | Y | 14 |  |  | 10/22/1996 7:35:43 PM | 1996-04-12 | 6/3/1998 6:51:36 PM | EMS |  | DimensionlessRatio | % | N | (ratio) |  | Iron-Insoluble (ratio) |  | Metals - Insoluble; | Metals - Insoluble; EMS code: FE-I | Chemical | Insoluble | LAB
FE-S | mg/L | Iron-Soluble | EQ01 | EQUIS (mg/L) | EQUIS (mg/L) | 0.01 | N | Y | Y | 4 |  |  | 10/22/1996 2:40:23 PM | 1996-04-12 | 6/3/1998 6:51:36 PM | EMS |  | Concentration | mg/L | Y | (fl. conc.) |  | Iron-Soluble (fl. conc.) |  | Metals - Soluble; | Metals - Soluble; EMS code: FE-S | Chemical | Soluble | LAB
FE-S | mg/dm2/d | Iron-Soluble | EQ15 | EQUIS (mg/dm2/d) | EQUIS (mg/dm2/d) | 0.01 | N | Y | Y | 2 |  |  | 10/22/1996 6:13:42 PM | 1996-04-12 | 6/3/1998 6:51:36 PM | EMS |  | MassFlux | mg/dm²/d | Y | (mass flux) |  | Iron-Soluble (mass flux) |  | Metals - Soluble; | Metals - Soluble; EMS code: FE-S | Chemical | Soluble | LAB
FE-T | mg/L | Iron Total | X351 | LA:HNO3:AA | Samples are lab-acidified (HNO3); digested HNO3; analysis by AA. | 0.03 | Y | Y | Y | 7 |  |  | 5/30/2006 8:05:04 AM | 1996-04-12 | 6/3/1998 6:51:36 PM | EMS |  | Concentration | mg/L | Y | (fl. conc.) |  | Iron Total (fl. conc.) |  | Metals - Total; | Metals - Total; EMS code: FE-T | Chemical | Total | LAB
FE-T | ug/g | Iron Total | FS89 | Homo,HNO3 dig,ICP-AES | Homo;HNO3 dig;ICP-AES | 0.8 | Y | Y | Y | 11 | 1/5/2001 | 10/31/2022 | 11/22/2022 2:40:39 PM | 2001-03-08 |  | NPEPPIN |  | MassPerMass | mg/kg | Y | (mass conc.) |  | Iron Total (mass conc.) |  | Metals - Total; | Metals - Total; EMS code: FE-T | Chemical | Total | LAB
FE-T | mg/L | Iron Total | MSEA | Metals in Seawater (undigested) by (CRC) ICP-MS | This method describes the multi-elemental determination of trace elements by Collision/Reaction Cell (CRC) ICP-MS. The method measures ions produced by a radio-frequency inductively coupled plasma. Analyte species originating in a liquid are pneumatically nebulized and the resulting aerosol transported by argon gas into the plasma torch. The ions produced are entrained in the plasma gas and extracted, by means of a differentially pumped vacuum interface, into a mass spectrometer. The ions flow through a Collision/Reaction Cell (CRC) where depending on the mode of analysis either the ions passes through the cell unhindered (no gas - classical ICP-MS mode), or collisions/reactions chemistry takes place between the ions and a neutral gas such as hydrogen or helium. Any polyatomic ions which give interference on the analytes are removed. The ions are then sorted according to their mass-to-charge ratios by a quadrupole mass spectrometer. The ions are then sorted according to their mass-to-charge ratios by a quadrupole mass spectrometer. The ions transmitted through the quadrupole are quantified by a channel electron multiplier and the ion information processed by a data handling system. Although most polyatomic interference is removed, any additional interference must be assessed and valid corrections applied or the data flagged to indicate problems. Interference correction must include compensation for background ions contributed by the plasma gas, reagents, and constituents of the sample matrix. Instrument drift as well as suppressions or enhancements of instrument response caused by the sample matrix must be corrected for by the use of internal standards. | 0.0004 | Y | Y | Y | 432 | 8/12/2009 12:08:00 AM | 4/29/2024 | 6/7/2024 10:03:16 AM | 2009-09-08 |  | RDGIBSON |  | Concentration | mg/L | Y | (fl. conc.) |  | Iron Total (fl. conc.) |  | Metals - Total; | Metals - Total; EMS code: FE-T | Chemical | Total | LAB
FE-T | mg/L | Iron Total | 1349 | Tot Metals in Seawater by CRC ICPMS (BC) | Tot Metals in Seawater by CRC ICPMS (BC) | 0.01 | Y | Y | Y | 98 | 8/16/2019 | 12/5/2021 | 12/13/2021 12:34:16 PM | 2019-09-30 |  | KATFRASE |  | Concentration | mg/L | Y | (fl. conc.) |  | Iron Total (fl. conc.) |  | Metals - Total; | Metals - Total; EMS code: FE-T | Chemical | Total | LAB
FE-T | mg/L | Iron Total | METN | MET-T-NP-CCMS-VA | Ultra trace metals in water are analyzed by CRC ICPMS, based on US EPA Method 6020B (July 2014). The detection limits provided can only be met for undigested samples. This procedure is intended for colorless, non-turbid, acid-preserved water samples (i.e. pristine water samples), having turbidity < 1 NTU and no odor. Where turbidity exceeds 1 NTU, and/or the sample is colored and has an odor, results may be biased low compared to true Total Metals concentrations. ALS recommends that turbidity analysis be requested on samples submitted for this test to aid with interpretation of results. Where turbidity is <1NTU, undigested metals are equivalent to total metals concentrations. | 0.001 | Y | Y | Y | 3648 | 4/3/2012 | 8/19/2024 | 8/26/2024 1:03:06 PM | 2019-08-20 |  | KATFRASE |  | Concentration | mg/L | Y | (fl. conc.) |  | Iron Total (fl. conc.) |  | Metals - Total; | Metals - Total; EMS code: FE-T | Chemical | Total | LAB
FE-T | ug/g | Iron Total | MTID | MET-DRY-CCMS-VA | Metals in Tissue by CRC ICPMS (DRY) Methodology: This method is conducted following British Columbia Lab Manual method "Metals in Animal Tissue and Vegetation (Biota) - Prescriptive". Tissue samples are homogenized and sub-sampled prior to hotblock digestion with nitric and hydrochloric acids, in combination with addition of hydrogen peroxide. Instrumental analysis is by collision cell inductively coupled plasma - mass spectrometry (modified from EPA Method 6020A). Method Limitation: This method employs a strong acid/peroxide digestion, and is intended to provide a conservative estimate of bio-available metals. Near complete recoveries are achieved for most toxicologically important metals, but elements associated with recalcitrant minerals may be only partially recovered. | 3 | Y | Y | Y | 122 | 8/13/2014 | 9/8/2022 | 10/26/2022 7:14:15 AM | 2015-03-09 | 3/11/2015 4:12:05 PM | RDGIBSON | EMS | MassPerMass | mg/kg | Y | (mass conc.) |  | Iron Total (mass conc.) |  | Metals - Total; | Metals - Total; EMS code: FE-T | Chemical | Total | LAB
FE-T | mg/L | Iron Total | X349 | FA:HNO3:ICP | Samples are field-acidified with HNO3; HNO3 digestion; analysis by ICP/AES. | 0.002 | Y | Y | Y | 30 |  |  | 10/30/2014 8:31:17 AM | 1998-05-22 | 6/3/1998 6:51:49 PM | NPEPPIN |  | Concentration | mg/L | Y | (fl. conc.) |  | Iron Total (fl. conc.) |  | Metals - Total; | Metals - Total; EMS code: FE-T | Chemical | Total | LAB
FE-T | ug/g wet | Iron Total | X552 | Metals in Biota (wet), total, acid digestion, ICP-SFMS | One gram of wet, homogenized biota is weighed in a microwave TFM vessel, and 8mL of Nitric Acid and 1mL of Hydrogen Peroxide are added. The sample is digested at 200 degrees C for 15 minutes in a closed vessel high pressure microwave oven. The sample is cooled, the volume adjusted to 100mL and introduced to an inductively coupled argon plasma-sector field mass spectrometer (ICP-SFMS) by pneumatic nebulization. Each respective element is measured at a specific mass to charge ratio, m/z value expressed in amu (atomic mass units), and at a specified instrument resolution. LR: Low Resolution ~300, MR: Mid Resolution ~4000. Ref: 02-2704, National Laboratory for Environmental Testing, CCIW, Environment Canada, Burlington. Requested April 2003. | 0.05 | Y | Y | Y | 37 | 8/9/2016 | 10/16/2023 | 11/28/2023 2:21:32 PM | 2004-01-19 | 4/19/2021 1:22:00 PM | UMULLER | EMS | MassPerMass | µg/g | Y | (mass conc.) | (wet) | Iron Total (wet) (mass conc.) |  | Metals - Total; | Metals - Total; EMS code: FE-T | Chemical | Total | LAB
FE-T | mg/L | Iron Total | X555 | Metals in water, total, "in-bottle" digestion, ICP-SFMS | The whole water sample collected in a HDPE/Teflon bottle, is de-capped in a Class 100 workstation, acidified to 2% (v/v) HNO3 and 1% (v/v) HCl and recapped. The sample is "in-bottle" digested in a convection oven, set at 60 degrees C for 16 hours. The sample is cooled and introduced to an inductively coupled argon plasma-sector field mass spectrometer (ICP-SFMS) by pneumatic nebulization. Each respective element is measured at a specific mass to charge ratio, m/z value expressed in amu (atomic mass units), and at resolution. LR: Low Resolution ~300, MR: Mid Resolution ~4000 and HR: High a specified instrument Resolution ~10000. Ref: Method number 02-2003, National Laboratory for Environmental Testing, CCIW, Environment Canada, Burlington. Requested April 2003. | 0.0005 | Y | Y | Y | 17342 | 5/20/2003 | 12/6/2023 2:08:00 AM | 2/14/2024 9:05:32 AM | 2004-01-19 |  | UMULLER |  | Concentration | mg/L | Y | (fl. conc.) |  | Iron Total (fl. conc.) |  | Metals - Total; | Metals - Total; EMS code: FE-T | Chemical | Total | LAB
FE-T | mg/L | Iron Total | X285 | HNO3 Dig:ICP Analysis | HNO3 Dig:ICP Analysis | 0.002 | Y | Y | Y | 1012 | 8/10/1993 | 12/2/2010 12:05:00 PM | 12/8/2010 2:19:57 PM | 1996-04-12 | 6/3/1998 6:51:36 PM | EMS |  | Concentration | mg/L | Y | (fl. conc.) |  | Iron Total (fl. conc.) |  | Metals - Total; | Metals - Total; EMS code: FE-T | Chemical | Total | LAB
FE-T | ug/g | Iron Total | 0190 | D&G:N/P Dig:ICP Anal. | Sample is dried and grind-mixed before nitric perchloric digestion followed by ICP analysis. | 1 | Y | Y | Y | 948 |  |  | 10/24/1996 3:06:16 PM | 1996-04-12 | 6/3/1998 6:51:36 PM | EMS |  | MassPerMass | mg/kg | Y | (mass conc.) |  | Iron Total (mass conc.) |  | Metals - Total; | Metals - Total; EMS code: FE-T | Chemical | Total | LAB
FE-T | mg/L | Iron Total | X088 | Nitric/Perchlor Dig,AA | Nitric/Perchlor Dig;AA | 0.1 | Y | Y | Y | 209 |  |  | 11/15/2000 1:51:21 PM | 1996-04-12 | 6/3/1998 6:51:36 PM | EMS |  | Concentration | mg/L | Y | (fl. conc.) |  | Iron Total (fl. conc.) |  | Metals - Total; | Metals - Total; EMS code: FE-T | Chemical | Total | LAB
FE-T | ug/g wet | Iron Total | X560 | Metals in Biota (wet), total, acid digestion, ICP-MS | Five grams of wet, homogenized biota is weighed in a centrifuge tube, and 2.5 mL of Nitric Acid and 2.5 mL of Hydrochloric acid are added. The sample is digested at room temperature for one hour and then digested at 90-95 degrees C in a water bath or heat block for one more hour. The sample is cooled and the volume is made to 50 mL with DI. The digestate is then diluted at ten fold and introduced to a collision/reaction cell inductively coupled argon plasma mass spectrometer (CRC ICPMS) by pneumatic nebulization. Each respective element is measured at a specific mass to charge ratio, m/z value expressed in amu (atomic mass units). - Maxxam Analytics (BC). Added to EMS Dec 10, 2008. | 10 | Y | Y | Y | 541 | 11/15/2008 4:00:00 AM | 10/30/2023 | 11/28/2023 2:21:33 PM | 2008-12-11 | 3/11/2015 4:17:10 PM | RDGIBSON | EMS | MassPerMass | µg/g | Y | (mass conc.) | (wet) | Iron Total (wet) (mass conc.) |  | Metals - Total; | Metals - Total; EMS code: FE-T | Chemical | Total | LAB
FE-T | mg/L | Iron Total | TCLP | Toxicity Characteristic Leaching Procedure | The TCLP procedure is designed to simulate the climatic leaching action | 0.5 | Y | Y | Y | 22 | 12/8/2009 7:49:00 PM | 9/28/2011 5:53:00 PM | 11/3/2011 11:10:19 PM | 2009-03-06 | 3/6/2009 8:09:45 AM | JCHUDAK | JCHUDAK | Concentration | mg/L | Y | (fl. conc.) |  | Iron Total (fl. conc.) |  | Metals - Total; | Metals - Total; EMS code: FE-T | Chemical | Total | LAB
FE-T | ug/g wet | Iron Total | X217 | Hom,Nitric/Perchl,ICP | Sample is wet but not dried before homogenization and nitric perchloric digestion followed by ICP analysis. | 0.1 | Y | Y | Y | 74 |  |  | 10/24/1996 10:01:27 AM | 1996-04-12 | 6/3/1998 6:51:42 PM | EMS |  | MassPerMass | µg/g | Y | (mass conc.) | (wet) | Iron Total (wet) (mass conc.) |  | Metals - Total; | Metals - Total; EMS code: FE-T | Chemical | Total | LAB
FE-T | ug/g | Iron Total | FB04 | Freeze Dry,Hom,Acid Dig,ICP | Sample moisture is determined by weighing samples before and after freeze-drying. Freeze-dried samples are ground in a Braun blender before being weighed into Teflon PFA liners and digested with concentrated HNO3, concentrated HCl, and 50% H2O2 in a microwave oven; which results in a complete sample dissolution. Samples are collected and volumized with Type I water. Analysis is performed via Inductively Coupled Argon Plasma Atomic Emission Spectrometry (ICP-AES) (simultaneous multielement analysis), Graphite Furnace Atomic Absorption Spectrometry (GF-AAS) (low level Cd and Pb analysis), Cold Vapor Atomic Fluorescence Spectrometry (CV-AFS) (Mercury analysis) and/or Flame Atomic Absorption Spectrometry (F-AAS) (K analysis). | 0.4 | Y | Y | Y | 241 | 8/22/1996 6:11:00 PM | 2/1/2002 2:30:00 PM | 6/23/2003 1:15:30 PM | 1997-02-27 | 6/3/1998 6:51:46 PM | MPAK |  | MassPerMass | mg/kg | Y | (mass conc.) |  | Iron Total (mass conc.) |  | Metals - Total; | Metals - Total; EMS code: FE-T | Chemical | Total | LAB
FE-T | mg/L | Iron Total | 0042 | GI HNO3 Dig: ICP | Sample aliquot taken from unfiltered unpreserved General Ions bottle; HNO3 digestion; ICP analysis. | 0.05 | Y | Y | Y | 11507 | 9/20/1996 | 7/5/2019 3:21:00 PM | 7/12/2019 6:49:13 PM | 1996-04-12 | 6/3/1998 6:51:41 PM | EMS |  | Concentration | mg/L | Y | (fl. conc.) |  | Iron Total (fl. conc.) |  | Metals - Total; | Metals - Total; EMS code: FE-T | Chemical | Total | LAB
FE-T | ug/g | Iron Total | 0192 | Hom, Nitric/Perchlor,ICP | Sample already in a dry state so does not require drying before homogenization and nitric perchlorid digestion followed by ICP analysis. | 1 | Y | Y | Y | 708 |  |  | 10/24/1996 12:50:25 PM | 1996-04-12 | 6/3/1998 6:51:36 PM | EMS |  | MassPerMass | mg/kg | Y | (mass conc.) |  | Iron Total (mass conc.) |  | Metals - Total; | Metals - Total; EMS code: FE-T | Chemical | Total | LAB
FE-T | ug/g | Iron Total | FS01 | D&G,Nitric/HCl:ICP | Samples are dried at 60 C, then either sieved to -100mesh (150 um) and rolled to homogenize, or finely ground with a mortar and pestle or a shatter box laboratory mill and rolled (depending on the submitter's requirements). The samples are then weighed into Teflon digestion liners and digested with HNO3 and HCl in a microwave oven. Samples are cooled, volumized, and allowed to settle overnight before being decanted. Analysis is then performed via Inductively Coupled Argon Plasma Atomic Emission Spectrometry (ICP-AES) (Simultaneous multi element analysis). In ICP Spectrometry, the liquid sample is pumped to the nebulizer, where it is convertd into an aerosol. the nebulizer carrier gas flow than transports the wet aerosol past the impact bead, where the large droplets condense out and flow to the waste container. The remaining fine aerosol is carried into the plasma. There, the high temperature desolvates the aerosol and dissociates the sample material into individual atoms and ions, which are excited to emit light at wavelengths characteristic of the elements present. A diffraction grating in the spectrometer diffracts the various wavelengths of light through exit slits to the photo-multiplier tubes, which convert the light energy to electronic signals. | 0.8 | Y | Y | Y | 2642 | 8/30/1996 1:52:00 PM | 10/10/2007 | 5/24/2008 1:15:02 AM | 1996-11-20 | 6/3/1998 6:51:46 PM | MPAK |  | MassPerMass | mg/kg | Y | (mass conc.) |  | Iron Total (mass conc.) |  | Metals - Total; | Metals - Total; EMS code: FE-T | Chemical | Total | LAB
FE-T | mg/L | Iron Total | X174 | Acid Dig:Dir.Asp | Acid Dig:Dir.Asp | 0.005 | Y | Y | Y | 10 |  |  | 9/18/1997 2:16:33 PM | 1996-04-12 | 6/3/1998 6:51:36 PM | EMS |  | Concentration | mg/L | Y | (fl. conc.) |  | Iron Total (fl. conc.) |  | Metals - Total; | Metals - Total; EMS code: FE-T | Chemical | Total | LAB
FE-T | mg/L | Iron Total | X145 | Dig, ICP | Dig; ICP | 0.03 | Y | Y | Y | 867 | 5/14/2013 | 9/22/2016 | 6/13/2024 3:58:21 PM | 1996-04-12 | 6/3/1998 6:51:41 PM | EMS |  | Concentration | mg/L | Y | (fl. conc.) |  | Iron Total (fl. conc.) |  | Metals - Total; | Metals - Total; EMS code: FE-T | Chemical | Total | LAB
FE-T | mg/L | Iron Total | X074 | FA,HNO3,Plasma Spec. | FA;HNO3;Plasma Spec. | 0.1 | Y | Y | Y | 5034 |  |  | 8/28/2018 3:14:04 PM | 1996-04-12 | 6/3/1998 6:51:36 PM | EMS |  | Concentration | mg/L | Y | (fl. conc.) |  | Iron Total (fl. conc.) |  | Metals - Total; | Metals - Total; EMS code: FE-T | Chemical | Total | LAB
FE-T | mg/L | Iron Total | 0105 | 5%AR Dig:APDC Ext:AA Anl | 5%AR Dig:APDC Ext:AA Anl | 0.005 | Y | Y | Y | 273 |  |  | 10/24/1996 2:13:46 PM | 1996-04-12 | 6/3/1998 6:51:36 PM | EMS |  | Concentration | mg/L | Y | (fl. conc.) |  | Iron Total (fl. conc.) |  | Metals - Total; | Metals - Total; EMS code: FE-T | Chemical | Total | LAB
FE-T | ug/g wet | Iron Total | BPSY | MET-WET-BPSY-CCMS-VA | Metals in Animal Tissue and Vegetation (Biota) - Prescriptive. EPA 200.3/6020B This method is conducted following British Columbia Lab Manual method "Metals in Animal Tissue and Vegetation (Biota) - Prescriptive". Tissue samples are digested with nitric and hydrochloric acids; in combination with addition of hydrogen peroxide. Instrumental analysis is by collision cell inductively coupled plasma - mass spectrometry (modified from EPA Method 6020B). | 1 | Y | Y | Y | 2 | 11/7/2019 | 11/7/2019 | 6/10/2021 1:57:10 PM | 2021-06-04 |  | EMS |  | MassPerMass | µg/g | Y | (mass conc.) | (wet) | Iron Total (wet) (mass conc.) |  | Metals - Total; | Metals - Total; EMS code: FE-T | Chemical | Total | LAB
FE-T | ug/g | Iron Total | EQ02 | EQUIS (ug/g) | EQUIS (ug/g) | 1 | N | Y | Y | 800 |  |  | 10/23/1996 12:08:28 AM | 1996-04-12 | 6/3/1998 6:51:36 PM | EMS |  | MassPerMass | mg/kg | Y | (mass conc.) |  | Iron Total (mass conc.) |  | Metals - Total; | Metals - Total; EMS code: FE-T | Chemical | Total | LAB
FE-T | ug/g | Iron Total | FS82 | PBM,dry60C,10mesh,SALM-ICP-MS | PBM;dry60C;10mesh;SALM-ICP-MS | 100 | Y | Y | Y | 3559 | 5/21/2002 | 7/24/2024 | 7/29/2024 5:17:09 PM | 2002-11-12 |  | UMULLER |  | MassPerMass | mg/kg | Y | (mass conc.) |  | Iron Total (mass conc.) |  | Metals - Total; | Metals - Total; EMS code: FE-T | Chemical | Total | LAB
FE-T | % | Iron Total | MS61 | ALS Minerals ME-MS61 | ALS Minerals ME-MS61. A prepared sample is digested with HNO3; HF; and HClO4 in a graphite heating block to near dryness. After cooling; the sample is boiled in diluted HCl. Bulk to volume with deionized water; mixed and analyzed by ICP-AES and ICP-MS for all elements. | 0.01 | Y | Y | Y | 0 |  |  |  | 2024-01-29 |  | EMS |  | DimensionlessRatio | % | N | (ratio) |  | Iron Total (ratio) |  | Metals - Total; | Metals - Total; EMS code: FE-T | Chemical | Total | LAB
FE-T | ug/g | Iron Total | ICPM | Metals in Soil by CRC ICPMS | This analysis is carried out using procedures from CSR Analytical Method: "Strong Acid Leachable Metals (SALM) in Soil", BC Ministry of Environment, 26 June 2009, and procedures adapted from EPA Method 200.2. The sample is manually homogenized, dried at 60 degrees Celsius, sieved through a 63 um (230 mesh) sieve, and a representative subsample of the dry material is weighed. The sample is then digested at 95 degrees Celsius for 2 hours by block digester using concentrated nitric and hydrochloric acids. Instrumental analysis is by collision cell inductively coupled plasma - mass spectrometry (modifed from EPA Method 6020A). Method Limitation: This method is not a total digestion technique. It is a very strong acid digestion that is intended to dissolve those metals that may be environmentally available. By design, elements bound in silicate structures are not normally dissolved by this procedure as they are not usually mobile in the environment. | 50 | Y | Y | Y | 103 | 9/19/2017 | 9/24/2023 | 1/12/2024 11:16:52 AM | 2018-01-09 |  | JKROGH |  | MassPerMass | mg/kg | Y | (mass conc.) |  | Iron Total (mass conc.) |  | Metals - Total; | Metals - Total; EMS code: FE-T | Chemical | Total | LAB
FE-T | mg/L | Iron Total | X072 | FA,HNO3,Graph.Furnace,AA | Samples are field-acidified with HNO3; HNO3 digestion; analysis by GFAA. | 0.01 | Y | Y | Y | 6742 |  |  | 8/26/2024 2:41:32 PM | 1996-04-12 | 6/3/1998 6:51:36 PM | EMS |  | Concentration | mg/L | Y | (fl. conc.) |  | Iron Total (fl. conc.) |  | Metals - Total; | Metals - Total; EMS code: FE-T | Chemical | Total | LAB
FE-T | % | Iron Total | MS41 | ALS Minerals ME-MS41 | ALS Minerals ME-MS41. A prepared sample is digested with aqua regia in a graphite heating block. After cooling; the resulting solution is diluted to with deionized water; mixed and analyzed by ICP-AES. Samples are then analyzed by ICP-MS for the remaining suite of elements. | 0.01 | Y | Y | Y | 0 |  |  |  | 2024-01-29 |  | EMS |  | DimensionlessRatio | % | N | (ratio) |  | Iron Total (ratio) |  | Metals - Total; | Metals - Total; EMS code: FE-T | Chemical | Total | LAB
FE-T | ug/g | Iron Total | X492 | Homo,HCl-HNO3 dig,ICP-OES,dry calc | Homogenize samples and take sub-samples. For Total Recoverable Metals Digest with Aqua Regia. Analyses by ICPOES (Inductively Coupled Plasma Optical Emission Spectroscopy). Perform moisture analysis on separate sub-sample and calculate dry value for trace elements. | 50 | Y | Y | Y | 7 | 8/16/2022 6:31:00 PM | 8/16/2022 6:42:00 PM | 2/25/2023 1:41:51 PM | 1998-11-13 |  | NPEPPIN |  | MassPerMass | mg/kg | Y | (mass conc.) |  | Iron Total (mass conc.) |  | Metals - Total; | Metals - Total; EMS code: FE-T | Chemical | Total | LAB
FE-T | mg/L | Iron Total | F121 | Acid,ICPMS | Acidified, no digestion, ICPMS | 0.001 | Y | Y | Y | 6980 | 3/12/2010 7:57:00 PM | 8/15/2024 9:10:00 PM | 8/26/2024 8:06:03 AM | 2002-07-17 | 7/16/2012 3:28:34 PM | UMULLER | RDGIBSON | Concentration | mg/L | Y | (fl. conc.) |  | Iron Total (fl. conc.) |  | Metals - Total; | Metals - Total; EMS code: FE-T | Chemical | Total | LAB
FE-T | mg/L | Iron Total | X073 | FA,HNO3,Direct Flame,AA | Samples are field-acidified with HNO3; HNO3 digestion; analysis by AA. | 0.005 | Y | Y | Y | 304 |  |  | 1/25/2007 4:06:11 PM | 1996-04-12 | 6/3/1998 6:51:41 PM | EMS |  | Concentration | mg/L | Y | (fl. conc.) |  | Iron Total (fl. conc.) |  | Metals - Total; | Metals - Total; EMS code: FE-T | Chemical | Total | LAB
FE-T | mg/dm2/d | Iron Total | DUST | HNO3 dig.(microwave or hot block),ICPMS | HNO3 dig.(microwave or hot block);ICPMS |  | Y | Y | Y | 121 | 2/4/2020 | 7/31/2024 | 7/31/2024 11:23:12 AM | 2022-09-01 |  | EMS |  | MassFlux | mg/dm²/d | Y | (mass flux) |  | Iron Total (mass flux) |  | Metals - Total; | Metals - Total; EMS code: FE-T | Chemical | Total | LAB
FE-T | mg/L | Iron Total | X234 | Aqua Reg. Dig,AA | Aqua Reg. Dig;AA | 0.005 | Y | Y | Y | 248 |  |  | 10/24/1996 12:12:10 PM | 1996-04-12 | 6/3/1998 6:51:36 PM | EMS |  | Concentration | mg/L | Y | (fl. conc.) |  | Iron Total (fl. conc.) |  | Metals - Total; | Metals - Total; EMS code: FE-T | Chemical | Total | LAB
FE-T | ug/g | Iron Total | 0196 | D&G: LTA/HCl: ICP | D&G: LTA/HCl: ICP | 1 | Y | Y | Y | 27 |  |  | 10/24/1996 2:08:12 PM | 1996-04-12 | 6/3/1998 6:51:36 PM | EMS |  | MassPerMass | mg/kg | Y | (mass conc.) |  | Iron Total (mass conc.) |  | Metals - Total; | Metals - Total; EMS code: FE-T | Chemical | Total | LAB
FE-T | ug/g | Iron Total | E496 | CSSS (1993) 11 | CSSS (1993) 11: Plant available micro nutrients (Cu; Fe; Mn; Zn) are determined by extraction using 0.005 M DTPA followed by analysis using ICPOES. | 2 | Y | Y | Y | 0 |  |  |  | 2024-02-27 |  | RDGIBSON |  | MassPerMass | mg/kg | Y | (mass conc.) |  | Iron Total (mass conc.) |  | Metals - Total; | Metals - Total; EMS code: FE-T | Chemical | Total | LAB
FE-T | mg/L | Iron Total | X028 | FA,HNO3,AA Analysis | FA;HNO3;AA Analysis | 0.1 | Y | Y | Y | 302 |  |  | 10/29/2001 11:34:42 AM | 1996-04-12 | 6/3/1998 6:51:36 PM | EMS |  | Concentration | mg/L | Y | (fl. conc.) |  | Iron Total (fl. conc.) |  | Metals - Total; | Metals - Total; EMS code: FE-T | Chemical | Total | LAB
FE-T | mg/L | Iron Total | X561 | Total metals: FA, HNO3, HR-ICPMS (No Prep)Prep) | Ultra trace metals in water are analyzed by high resolution inductively coupled plasma mass spectrometry (HR-ICPMS), modified from US EPA 200.8, (Revision 5.5). The detection limits provided can only be met for undigested samples. This procedure is intended for colorless, non-turbid, acid-preserved water samples (i.e. pristine water samples), having turbidity < 1 NTU and no odor. Where turbidity exceeds 1 NTU, and/or the sample is colored and has an odor, results may be biased low compared to true Total Metals concentrations. ALS recommends that turbidity analysis be requested on samples submitted for this test to aid with interpretation of results. | 0.001 | Y | Y | Y | 2747 | 6/8/2015 | 10/5/2023 | 1/31/2024 2:37:41 PM | 2015-10-22 |  | RDGIBSON |  | Concentration | mg/L | Y | (fl. conc.) |  | Iron Total (fl. conc.) |  | Metals - Total; | Metals - Total; EMS code: FE-T | Chemical | Total | LAB
FE-T | mg/L | Iron Total | 0040 | HNO3 Dig: ICP Analysis | HNO3 Dig: ICP Analysis | 0.01 | Y | Y | Y | 17870 | 10/17/2000 | 10/31/2003 | 7/18/2024 8:07:20 AM | 1996-04-12 | 6/3/1998 6:51:36 PM | EMS |  | Concentration | mg/L | Y | (fl. conc.) |  | Iron Total (fl. conc.) |  | Metals - Total; | Metals - Total; EMS code: FE-T | Chemical | Total | LAB
FE-T | ug/g | Iron Total | X230 | Hom,HNO3/HCl,ICP | The sample is dried, ground to pass 100 mesh, and digested with a nitric-hydrochloric acid mixture to solubilize the solid matter and to remove the organic material by oxidation and volatilization. The completed digests are submitted to metals instrumentation personnel for metals analysis by Inductively Coupled Atomic Emission Spectrophotometry (ICP-AES), or hydride analysis by atomic absorption. | 0.5 | Y | Y | Y | 264 | 2/12/2002 | 12/30/2005 | 2/3/2006 1:13:12 AM | 1996-04-12 | 6/3/1998 6:51:36 PM | EMS |  | MassPerMass | mg/kg | Y | (mass conc.) |  | Iron Total (mass conc.) |  | Metals - Total; | Metals - Total; EMS code: FE-T | Chemical | Total | LAB
FE-T | ug/g | Iron Total | PB02 | PBM,HNO3/HCl dig,dry60C,100mesh, | Performance Based Method. Samples are dried at 60 C and sieved using a .149 mm (100 mesh) sieve. Solid samples are digested with the combination of nitric and hydrochloric acids with a 1:1 ratio of acids. For environmental purposes, strong acid leaches are used routinely since they have the advantage of speed, low cost, and superio reproducibility. All metals trapped within the refractory matrix are unlikely to have environmental consequences. Instrumental analysis of extracts produced using this digestion method can be performed by a variety of techniques. |  | Y | Y | Y | 10 | 9/13/2001 | 9/13/2001 | 11/6/2002 3:52:35 PM | 2001-07-18 |  | NPEPPIN |  | MassPerMass | mg/kg | Y | (mass conc.) |  | Iron Total (mass conc.) |  | Metals - Total; | Metals - Total; EMS code: FE-T | Chemical | Total | LAB
FE-T | mg/L | Iron Total | F038 | HNO3/HCl dig,ICP-AES | Dissolved and extractable water samples are ready for analysis once collected. Total samples are autoclaved with HNO3 and hydrochloric acid (HCI) in capped Teflon vials. Analysis is performed via Inductively Coupled Argon Plasma Atomic Emission Spectrometry (ICP-AES)(simultaneous multielement analysis), Graphite Furnace Atomic Absorption Spectrometry (GF-AAS) (low level analysis) and/or Flame Atomic Absorption Spectrometry (F-AAS). | 0.006 | Y | Y | Y | 7046 | 8/1/1996 6:04:00 PM | 5/6/2008 | 9/11/2019 8:47:42 AM | 1997-05-08 | 6/3/1998 6:51:46 PM | MPAK |  | Concentration | mg/L | Y | (fl. conc.) |  | Iron Total (fl. conc.) |  | Metals - Total; | Metals - Total; EMS code: FE-T | Chemical | Total | LAB
FE-T | mg/dm2/d | Iron Total | EQ15 | EQUIS (mg/dm2/d) | EQUIS (mg/dm2/d) | 0.01 | N | Y | Y | 4 |  |  | 10/22/1996 6:13:43 PM | 1996-04-12 | 6/3/1998 6:51:36 PM | EMS |  | MassFlux | mg/dm²/d | Y | (mass flux) |  | Iron Total (mass flux) |  | Metals - Total; | Metals - Total; EMS code: FE-T | Chemical | Total | LAB
FE-T | mg/L | Iron Total | EQ01 | EQUIS (mg/L) | EQUIS (mg/L) | 0.005 | N | Y | Y | 6621 |  |  | 10/24/1996 3:16:18 PM | 1996-04-12 | 6/3/1998 6:51:36 PM | EMS |  | Concentration | mg/L | Y | (fl. conc.) |  | Iron Total (fl. conc.) |  | Metals - Total; | Metals - Total; EMS code: FE-T | Chemical | Total | LAB
FE-T | mg/L | Iron Total | F082 | Aliquot,HNO3/HCL dig,ICPMS | Aliquot microwave digested with 3:1 HNO3/HCL. Analysis by ICPMS. For HiVol samples preliminary result are reported in mg/L and loading results are reported in ug/m3 | 0.01 | Y | Y | Y | 26055 | 1/24/2011 | 8/25/2024 | 8/26/2024 4:53:05 PM | 2015-04-15 |  | RDGIBSON |  | Concentration | mg/L | Y | (fl. conc.) |  | Iron Total (fl. conc.) |  | Metals - Total; | Metals - Total; EMS code: FE-T | Chemical | Total | LAB
FE-T | mg/L | Iron Total | ULTR | Total Metals in Water by ICPMS (Ultra) | Total Metals in Water by ICPMS (Ultra), EPA SW-846 3005A/6020A, MET-TOT-ULTRA-MS-VA This analysis is carried out using procedures adapted from "Standard Methods for the Examination of Water and Wastewater" published by the American Public Health Association, and with procedures adapted from "Test Methods for Evaluating Solid Waste" SW-846 published by the United States Environmental Protection Agency (EPA). The procedures may involve preliminary sample treatment by acid digestion, using either hotblock or microwave oven, or filtration (EPA Method 3005A). Instrumental analysis is by inductively coupled plasma - mass spectrometry (EPA Method 6020A). | 0.025 | Y | Y | Y | 956 | 5/2/2016 | 9/21/2020 | 7/21/2021 6:01:33 PM | 2016-05-25 |  | RDGIBSON |  | Concentration | mg/L | Y | (fl. conc.) |  | Iron Total (fl. conc.) |  | Metals - Total; | Metals - Total; EMS code: FE-T | Chemical | Total | LAB
FE-T | mg/L | Iron Total | F033 | Acid Dig,Centrifuge,ICPMS | For the determination of total recoverable elements in water or waste water, take 100 ml aliquot from a well mixed, acid preserved, sample containing not more than 0.25% (w/v) total solids and transfer to a 250 ml Griffin beaker. (If total solids are greater than 0.25% reduce the size of the aliquot by a proportionate amount). Add 1 ml of conc. nitric acid and 0.5 ml conc. hydrochloric acid. Heat on a hot plate until the volume has been reduced to 15-20ml, ensuring that the sample does not boil. Allow to cool and transfer to a 50 ml volumetric flask. Dilute to volume with ASTM type I water. Centrifuge the sample or allow to stand overnight to separate insoluble material. Prior analysis, pipette 20 ml into a 50 ml volumetric flask and dilute to volume with ASTM type I water. Analysis should be performed within seventy two hours of the completed preparation. Note that different laboratories may have very different field preservation, digestion procedures and detection limits for this method, so these should be documented. | 0.01 | Y | Y | Y | 26983 | 10/15/1996 9:27:00 AM | 8/22/2024 5:20:00 PM | 8/25/2024 9:47:36 PM | 1997-01-16 | 6/3/1998 6:51:46 PM | MPAK |  | Concentration | mg/L | Y | (fl. conc.) |  | Iron Total (fl. conc.) |  | Metals - Total; | Metals - Total; EMS code: FE-T | Chemical | Total | LAB
FE-T | mg/L | Iron Total | 0100 | Dig:Ext.(APDC):AA Anal. | Dig:Ext.(APDC):AA Anal. | 0.005 | Y | Y | Y | 821 |  |  | 10/24/1996 9:48:38 AM | 1996-04-12 | 6/3/1998 6:51:36 PM | EMS |  | Concentration | mg/L | Y | (fl. conc.) |  | Iron Total (fl. conc.) |  | Metals - Total; | Metals - Total; EMS code: FE-T | Chemical | Total | LAB
FE-T | ug/g wet | Iron Total | MTIW | MET-WET-CCMS-MID-VA | Metals in Tissue by CRC ICPMS (WET) Methodology: This method is conducted following British Columbia Lab Manual method "Metals in Animal Tissue and Vegetation (Biota) - Prescriptive". Tissue samples are homogenized and sub-sampled prior to hotblock digestion with nitric and hydrochloric acids, in combination with addition of hydrogen peroxide. Instrumental analysis is by collision cell inductively coupled plasma - mass spectrometry (modified from EPA Method 6020A). Method Limitation: This method employs a strong acid/peroxide digestion, and is intended to provide a conservative estimate of bio-available metals. Near complete recoveries are achieved for most toxicologically important metals, but elements associated with recalcitrant minerals may be only partially recovered. | 1 | Y | Y | Y | 24 | 11/25/2018 | 12/6/2018 | 7/21/2021 2:55:07 PM | 2018-12-24 |  | EMS |  | MassPerMass | µg/g | Y | (mass conc.) | (wet) | Iron Total (wet) (mass conc.) |  | Metals - Total; | Metals - Total; EMS code: FE-T | Chemical | Total | LAB
FE-T | mg/L | Iron Total | 0200 | 20%A.R.Dig:AA Anal. | 20%A.R.Dig:AA Anal. | 0.1 | Y | Y | Y | 34 |  |  | 10/24/1996 9:44:09 AM | 1996-04-12 | 6/3/1998 6:51:36 PM | EMS |  | Concentration | mg/L | Y | (fl. conc.) |  | Iron Total (fl. conc.) |  | Metals - Total; | Metals - Total; EMS code: FE-T | Chemical | Total | LAB
FE-T | % | Iron Total | 6041 | Aqua Regia by ICPMS | These were sediment/soils digested with Aqua-Regia (Nitric Acid & Hydrochloric Acids) and then analysis by ICPMS. Results would be reported in mg/kg and/or %. | 0.001 | Y | Y | Y | 92 | 10/13/2019 | 7/27/2024 | 7/29/2024 9:13:08 AM | 2019-09-06 |  | KATFRASE |  | DimensionlessRatio | % | N | (ratio) |  | Iron Total (ratio) |  | Metals - Total; | Metals - Total; EMS code: FE-T | Chemical | Total | LAB
FE-T | mg/L | Iron Total | X170 | ICP | This code is a catch-all for a large variety of ICP methods including water and also solids digested into solutions as an intermediate result. Different procedures will have greatly different detection limits. (a) For air particulates, preliminary results are reported in mg/L and loading results as mg/dm2/d. (b) For Hi Vol samples preliminiary results are same as reported in mg/L and loading results as ug/m3. (c) Note that for Rain samples there is no filtration nor digestion steps. | 0.03 | Y | Y | Y | 8346 | 3/28/1996 6:03:00 PM | 2/28/2019 | 10/24/2022 10:58:25 AM | 1996-04-12 | 8/15/2019 1:19:08 PM | EMS | KATFRASE | Concentration | mg/L | Y | (fl. conc.) |  | Iron Total (fl. conc.) |  | Metals - Total; | Metals - Total; EMS code: FE-T | Chemical | Total | LAB
FE-T | mg/L | Iron Total | 0020 | HNO3 Dig:AA Analysis | HNO3 Dig:AA Analysis | 0.1 | Y | Y | Y | 3393 | 6/25/2010 3:15:00 PM | 12/1/2010 | 12/9/2010 3:55:26 PM | 1996-04-12 | 6/3/1998 6:51:36 PM | EMS |  | Concentration | mg/L | Y | (fl. conc.) |  | Iron Total (fl. conc.) |  | Metals - Total; | Metals - Total; EMS code: FE-T | Chemical | Total | LAB
FE-T | % | Iron Total | EQ34 | EQUIS (%) | EQUIS (%) | 0 | N | Y | Y | 13 |  |  | 10/22/1996 7:34:23 PM | 1996-04-12 | 6/3/1998 6:51:41 PM | EMS |  | DimensionlessRatio | % | N | (ratio) |  | Iron Total (ratio) |  | Metals - Total; | Metals - Total; EMS code: FE-T | Chemical | Total | LAB
FE-T | mg/L | Iron Total | F048 | Acid dig,conc'n,nebul.,ICP-OES | This method provides an automated procedure for the Inductively Coupled Argon Plasma Optical Emission Spectrophotometric (ICP-OES) determination and quantitation of trace amounts of metals in surface, ground and precipitation waters after a manual acid digestion and concentration procedure and ultrasonic nebulisation. Organically-bound metals and metal complexes are converted to metal ions by digestion with nitric acid [method details says HCl]. To increase the signal detected by the ICP, the samples are concentrated 10 times. The liquid sample is pumped onto the piezoelectric transducer, where the ultrasonic action efficiently converts it to a fine, dense aerosol. The nebulizer carrier gas flow then transports the wet aerosol through a heated U-tube, where the solvent is vaporized. Subsequently, most of the solvent vapour is removed from the aerosol stream by passing the stream through a chilled condenser. The result is a dry, analyte-laden aerosol which is introduced to the plasma. There, the high temperature desolvates the aerosol and dissociates the sample material into individual atoms and ions, which are excited to emit light at wave lengths characteristic of the elements present. Dissolved samples are field filtered through 0.45 micron cellulose acetate filter prior to acid preservation. Samples for extractable analysis are acidified with 2 mL concentrated HNO3 per litre and allowed to stand a minimum of 24 hours before analysis. Clients providing these samples are advised to preserve samples at the collection site. | 0.0004 | Y | Y | Y | 6304 | 7/4/1997 2:36:00 PM | 6/1/2006 6:01:00 AM | 9/11/2019 8:46:38 AM | 1997-10-20 | 6/3/1998 6:51:47 PM | NPEPPIN |  | Concentration | mg/L | Y | (fl. conc.) |  | Iron Total (fl. conc.) |  | Metals - Total; | Metals - Total; EMS code: FE-T | Chemical | Total | LAB
FE-T | ug/g | Iron Total | X551 | Metals in Biota (dry), total, acid digestion, ICP-SFMS | 0.1 gram of dried, homogenized biota is weighed in a microwave TFM vessel, and 8mL of Nitric Acid and 1mL of Hydrogen Peroxide are added. The sample is digested at 200 degrees C for 15 minutes in a closed vessel high pressure microwave oven. The sample is cooled, the volume adjusted to 100mL and introduced to an iductively coupled argon plasma-sector field mass spectrometer (ICP-SFMS) by pneumatic nebulization. Each respective element is measured at a specific mass to charge ratio, m/z value expressed in amu (atomic mass units), and at a specified instrument resolution. LR: Low Resolution ~300, MR: Mid Resolution ~4000 and HR: Hich Resolution ~10000. Ref: Method number 02-2705, National Laboratory for Environmental Testing, CCIW, Environment Canada, Burlington. Requested April 2003. | 0.5 | Y | Y | Y | 57 | 8/11/2015 | 10/30/2023 | 11/28/2023 2:21:33 PM | 2004-01-19 |  | UMULLER |  | MassPerMass | mg/kg | Y | (mass conc.) |  | Iron Total (mass conc.) |  | Metals - Total; | Metals - Total; EMS code: FE-T | Chemical | Total | LAB
FE-T | mg/L | Iron Total | MWID | Metals in Water by ICPMS (digested) | Metals in Water by ICPMS (digested) Based on EPA 200.8 | 0.02 | Y | Y | Y | 36359 |  |  | 7/31/2024 1:46:41 PM | 2009-12-24 |  | RDGIBSON |  | Concentration | mg/L | Y | (fl. conc.) |  | Iron Total (fl. conc.) |  | Metals - Total; | Metals - Total; EMS code: FE-T | Chemical | Total | LAB
FE2-O3 | % | Iron (as Fe2O3) | ICP06 | ALS Minerals ME-ICP06 | ALS Minerals ME-ICP06. A prepared sample is fused with a lithium tetraborate/lithium metaborate flux and poured into a graphite crucible and fused in a furnace at 1000°C. The resulting melt is then cooled and dissolved in 100 mL of 4% HNO3/2% HCl followed by ICP-AES. | 0.01 | Y | Y | Y | 0 |  |  |  | 2024-01-29 |  | EMS |  | DimensionlessRatio | % | N | (ratio) |  | Iron (as Fe2O3) (ratio) |  | Metals; | Metals; EMS code: FE2-O3 | Chemical |  | LAB
FE2D | mg/L | Ferrous Iron; Dissolved - Iron (II) | FE2D | FE2-D-COL-VA | APHA 3500-Fe B/James Ball et al (1999) This analysis is carried out using procedures adapted from APHA 3500-Fe B and "A New Method for the Direct Determination of Dissolved Iron Concentration in Acid Mine Waters" published by James W. Ball et al (1999). The procedure involves preliminary sample filtration, and ferrous iron is determined using the "FerroZine" colourimetric method. | 0.02 | Y | Y | Y | 2 | 11/8/2019 | 11/8/2019 | 1/9/2020 3:05:21 PM | 2020-01-09 |  | RDGIBSON |  | Concentration | mg/L | Y | (fl. conc.) |  | Ferrous Iron; Dissolved - Iron (II) (fl. conc.) |  | Metals - Dissolved; | Metals - Dissolved; EMS code: FE2D | Chemical | Dissolved | LAB
FE3D | mg/L | Ferric Iron; Dissolved - Iron (III) [FE-D minus FE2D] | XCAL | Calculation (mg/L) | Calculation (mg/L) | 0.03 | Y | Y | Y | 2 | 11/14/2019 | 11/14/2019 | 1/9/2020 3:05:21 PM | 2020-01-09 |  | RDGIBSON |  | Concentration | mg/L | Y | (fl. conc.) |  | Ferric Iron; Dissolved - Iron (III) [FE-D minus FE2D] (fl. conc.) |  | Metals - Dissolved; | Metals - Dissolved; EMS code: FE3D | Chemical | Dissolved | LAB
FFLD | mg/L | Fluoride measured in the Field | FLD | Field sampling | Field sampling method | 0.01 | Y | Y | Y | 0 |  |  |  | 2024-06-03 |  | RDGIBSON |  | Concentration | mg/L | Y | (fl. conc.) |  | Fluoride measured in the Field (fl. conc.) |  |  | EMS code: FFLD | Chemical |  | FIELD_RESULT | Field sampling
FIBR | Fib/cm3 | Asbestos | X291 | Optic.Micros.Fbr Cnt-Stack | Optic.Micros.Fbr Cnt-Stack | 0.01 | Y | Y | Y | 24 |  |  | 10/23/1996 8:13:42 PM | 1996-04-12 | 6/3/1998 6:51:42 PM | EMS |  | AmountOfSubstancePerVolume | #/mL | Y | (amt./vol.) |  | Asbestos (amt./vol.) | 1332-21-4 | Physical Properties; | Physical Properties; EMS code: FIBR | Chemical |  | LAB
FIBR | Fib/cm3 | Asbestos | EQ17 | EQUIS (Fibre/cc) | EQUIS (Fibre/cc) | 0.1 | N | Y | Y | 172 |  |  | 10/22/1996 7:44:50 PM | 1996-04-12 | 6/3/1998 6:51:36 PM | EMS |  | AmountOfSubstancePerVolume | #/mL | Y | (amt./vol.) |  | Asbestos (amt./vol.) | 1332-21-4 | Physical Properties; | Physical Properties; EMS code: FIBR | Chemical |  | LAB
FIZZ | None | Fizz rating | VOL08M | ALS Minerals OA-VOL08m | ALS Minerals OA-VOL08m. Determine the acid generation power of a sample as tonne CaCO3 per kilo-tonne using modified Sobek method | 1 | Y | Y | Y | 0 |  |  |  | 2024-01-29 |  | EMS |  | None | None | N | (unitless) |  | Fizz rating (unitless) |  | Physical Properties; | Physical Properties; EMS code: FIZZ | Physical |  | LAB
FIZZ | None | Fizz rating | VOL08 | ALS Minerals OA-VOL08 | ALS Minerals OA-VOL08. Determine the acid generation power of a sample as tonne CaCO3 per kilo-tonne. | 1 | Y | Y | Y | 0 |  |  |  | 2024-01-29 |  | RDGIBSON |  | None | None | N | (unitless) |  | Fizz rating (unitless) |  | Physical Properties; | Physical Properties; EMS code: FIZZ | Physical |  | LAB
FIZZ | None | Fizz rating | VOL08N | ALS Minerals OA-VOL08mn | ALS Minerals OA-VOL08mn. Determine the acid generation power of a sample as tonne CaCO3 per kilo-tonne using MEND 1991 method. | 1 | Y | Y | Y | 0 |  |  |  | 2024-01-29 |  | EMS |  | None | None | N | (unitless) |  | Fizz rating (unitless) |  | Physical Properties; | Physical Properties; EMS code: FIZZ | Physical |  | LAB
FIZZ | None | Fizz rating | VOL08S | ALS Minerals OA-VOL08s | ALS Minerals OA-VOL08s. Determine the acid generation power of a sample as tonne CaCO3 per kilo-tonne using Siderite correction method. | 1 | Y | Y | Y | 0 |  |  |  | 2024-01-29 |  | EMS |  | None | None | N | (unitless) |  | Fizz rating (unitless) |  | Physical Properties; | Physical Properties; EMS code: FIZZ | Physical |  | LAB
FLAS | C | Flash point | D93M | ASTM D93 (mod) | ASTM D93 (mod) Standard Test Methods for Flash Point by Pensky-Martens Closed Cup Tester | 30 | Y | Y | Y | 0 |  |  |  | 2024-01-24 |  | RDGIBSON |  | Temperature | degC | Y | (temp.) |  | Flash point (temp.) |  | Physical Properties; | Physical Properties; EMS code: FLAS | Physical |  | LAB
FLRI | mg/dm2/d | Fluoridation Index | X143 | Selective Ion Electrode | The analyte is determined potentiometrically by direct measurement using a selective ion electrode with standard reference electrode and either a meter with a direct concentration scale or a pH meter with an expanded mV scale. | 0.02 | Y | Y | Y | 1915 |  |  | 11/15/2000 1:50:52 PM | 1996-04-12 | 6/3/1998 6:51:42 PM | EMS |  | MassFlux | mg/dm²/d | Y | (mass flux) |  | Fluoridation Index (mass flux) |  | Inorganics; | Inorganics; EMS code: FLRI | Chemical |  | LAB
FLRI | mg/dm2/d | Fluoridation Index | EQ15 | EQUIS (mg/dm2/d) | EQUIS (mg/dm2/d) | 0 | N | Y | Y | 1010 |  |  | 10/24/1996 2:28:58 PM | 1996-04-12 | 6/3/1998 6:51:36 PM | EMS |  | MassFlux | mg/dm²/d | Y | (mass flux) |  | Fluoridation Index (mass flux) |  | Inorganics; | Inorganics; EMS code: FLRI | Chemical |  | LAB
FLRI | mg/dm2/d | Fluoridation Index | 5004 | Selective Ion Electrode | A calcium oxide plate is exposed to ambient air for a period of 30 days. During this time fluoride in the ambient air is collected as calcium fluoride. The plates undergo an extraction and the fluoride which is isolated is determined using the selective ion electrode procedure. | 0.005 | Y | Y | Y | 359 | 10/31/2000 | 1/2/2004 | 1/20/2004 12:08:12 PM | 1996-04-12 | 6/3/1998 6:51:36 PM | EMS |  | MassFlux | mg/dm²/d | Y | (mass flux) |  | Fluoridation Index (mass flux) |  | Inorganics; | Inorganics; EMS code: FLRI | Chemical |  | LAB
FLUD | mg/L | Fludioxonil | E751 | E751D Pesticides in Water by LC-MS/MS (POS1 List) | E751D Pesticides in Water by LC-MS/MS (POS1 List) (MOECC E3501) | 0.0005 | Y | Y | Y | 39 | 12/9/2021 | 12/31/2021 | 2/15/2022 9:50:05 AM | 2021-12-14 |  | EMS |  | Concentration | mg/L | Y | (fl. conc.) |  | Fludioxonil (fl. conc.) |  | Pesticides; | Pesticides; EMS code: FLUD | Chemical |  | LAB
FLUQ | mg/L | Flumequine CAS: 42835-25-6 | XMSP | HPLC and Tri. Quadpole MS/MS | EPA 1694 An aliquot of filtered sample containing 5% organic is injected directly. Each compound is separated by reversed phase HPLC. Analyzed by Electro Spray Ionization MS/MS detection using Triple Quadrupole MS/MS detector. Quantiitation with internal or external standard. |  | Y | Y | Y | 12 | 5/8/2020 | 6/4/2024 | 8/21/2024 3:32:03 PM | 2022-05-04 |  | EMS |  | Concentration | mg/L | Y | (fl. conc.) |  | Flumequine (fl. conc.) | 42835-25-6 | Pesticides; | Pesticides; EMS code: FLUQ | Chemical |  | LAB
FLUX | mg/L | Fluoxetine CAS: 54910-89-3 | XMSP | HPLC and Tri. Quadpole MS/MS | EPA 1694 An aliquot of filtered sample containing 5% organic is injected directly. Each compound is separated by reversed phase HPLC. Analyzed by Electro Spray Ionization MS/MS detection using Triple Quadrupole MS/MS detector. Quantiitation with internal or external standard. |  | Y | Y | Y | 12 | 5/8/2020 | 6/4/2024 | 8/21/2024 3:32:03 PM | 2022-05-04 |  | EMS |  | Concentration | mg/L | Y | (fl. conc.) |  | Fluoxetine(fl. conc.) | 54910-89-3 | Pesticides; | Pesticides; EMS code: FLUX | Chemical |  | LAB
FRLQ | None | Free liquid | 9095B | EPA 9095B (mod) | EPA 9095B (mod) Free Liquids are determined by the Paint Filter Test, where a specified amount of material is placed in a paint filter. If any portion of the test material passes through the filter in the five minute test period, the material is deemed to contain free liquids. The absence of any amount of free liquid classifies the sample as a solid. |  | Y | Y | Y | 0 |  |  |  | 2024-01-24 |  | RDGIBSON |  | None | None | N | (unitless) |  | Free liquid (unitless) |  | Physical Properties; | Physical Properties; EMS code: FRLQ | Physical |  | LAB
FSS | mg/L | solids, fixed suspended | F012 | 1.2 um GF filt:Grav. 550C | A homogeneous sample is filtered through a 5.5 cm Whatman GF/C glass fibre filter (1.2 um particle retention) that has been muffled at 550 C for 20 minutes and pre-weighed. For fixed non-filterable residue (FNFR) analysis, the filter with the residue is then dried at 103 C for one hour and weighed once again to constant weight. The filter with dried and weighed residue is then ignited in a muffle at 550 C to determine FNFR. For fixed filterable residue (FFR) analysis, evaporated overnight at 90-95C, then dried for 1 hour at 103C and weighed for total filterable residue (FR), then ignited for 1 hour at 550C to determine FFR. Volatile residue is the difference in weights after drying at 103 C and asking at 550 C. | 3 | Y | Y | Y | 8 | 5/15/2024 | 7/3/2024 | 7/17/2024 10:51:03 AM | 2022-10-26 |  | EMS |  | Concentration | mg/L | Y | (fl. conc.) |  | solids, fixed suspended (fl. conc.) |  | Physical Properties; | Physical Properties; EMS code: FSS | Chemical |  | LAB
FTC102 | mg/L | Fluorotelomer carboxylic acid, 10:2 (10:2 FTCA) | E3533 | MECP E3533 | PFAs in water by direct injection LC/MS/MS | 0.00012 | Y | Y | Y | 2 | 6/3/2024 | 6/3/2024 | 8/21/2024 3:32:03 PM | 2023-11-15 | 6/27/2024 5:06:31 PM | EMS | EMS | Concentration | mg/L | Y | (fl. conc.) |  | Fluorotelomer carboxylic acid, 10:2 (10:2 FTCA) (fl. conc.) |  | Per- and polyfluoroalkyl substances (PFAS); | Per- and polyfluoroalkyl substances (PFAS); EMS code: FTC102 | Chemical |  | LAB
FTC33 | mg/L | Fluorotelomer carboxylic acid, 3:3 (3:3 FTCA) | E3533 | MECP E3533 | PFAs in water by direct injection LC/MS/MS | 2e-05 | Y | Y | Y | 2 | 6/3/2024 | 6/3/2024 | 8/21/2024 3:32:03 PM | 2023-11-15 | 6/27/2024 5:06:31 PM | EMS | EMS | Concentration | mg/L | Y | (fl. conc.) |  | Fluorotelomer carboxylic acid, 3:3 (3:3 FTCA) (fl. conc.) |  | Per- and polyfluoroalkyl substances (PFAS); | Per- and polyfluoroalkyl substances (PFAS); EMS code: FTC33 | Chemical |  | LAB
FTC53 | mg/L | Fluorotelomer carboxylic acid, 5:3 (5:3 FTCA) | E3533 | MECP E3533 | PFAs in water by direct injection LC/MS/MS | 2e-05 | Y | Y | Y | 2 | 6/3/2024 | 6/3/2024 | 8/21/2024 3:32:03 PM | 2023-11-15 | 6/27/2024 5:06:31 PM | EMS | EMS | Concentration | mg/L | Y | (fl. conc.) |  | Fluorotelomer carboxylic acid, 5:3 (5:3 FTCA) (fl. conc.) |  | Per- and polyfluoroalkyl substances (PFAS); | Per- and polyfluoroalkyl substances (PFAS); EMS code: FTC53 | Chemical |  | LAB
FTC62 | mg/L | Fluorotelomer carboxylic acid, 6:2 (6:2 FTCA) | E3533 | MECP E3533 | PFAs in water by direct injection LC/MS/MS | 0.00012 | Y | Y | Y | 2 | 6/3/2024 | 6/3/2024 | 8/21/2024 3:32:03 PM | 2023-11-15 | 6/27/2024 5:06:31 PM | EMS | EMS | Concentration | mg/L | Y | (fl. conc.) |  | Fluorotelomer carboxylic acid, 6:2 (6:2 FTCA) (fl. conc.) |  | Per- and polyfluoroalkyl substances (PFAS); | Per- and polyfluoroalkyl substances (PFAS); EMS code: FTC62 | Chemical |  | LAB
FTC73 | mg/L | Fluorotelomer carboxylic acid, 7:3 (7:3 FTCA) | E3533 | MECP E3533 | PFAs in water by direct injection LC/MS/MS | 2e-05 | Y | Y | Y | 2 | 6/3/2024 | 6/3/2024 | 8/21/2024 3:32:03 PM | 2023-11-15 | 6/27/2024 5:06:31 PM | EMS | EMS | Concentration | mg/L | Y | (fl. conc.) |  | Fluorotelomer carboxylic acid, 7:3 (7:3 FTCA) (fl. conc.) |  | Per- and polyfluoroalkyl substances (PFAS); | Per- and polyfluoroalkyl substances (PFAS); EMS code: FTC73 | Chemical |  | LAB
FTC82 | mg/L | Fluorotelomer carboxylic acid, 8:2 (8:2 FTCA) | E3533 | MECP E3533 | PFAs in water by direct injection LC/MS/MS | 0.00012 | Y | Y | Y | 2 | 6/3/2024 | 6/3/2024 | 8/21/2024 3:32:03 PM | 2023-11-15 | 6/27/2024 5:06:31 PM | EMS | EMS | Concentration | mg/L | Y | (fl. conc.) |  | Fluorotelomer carboxylic acid, 8:2 (8:2 FTCA) (fl. conc.) |  | Per- and polyfluoroalkyl substances (PFAS); | Per- and polyfluoroalkyl substances (PFAS); EMS code: FTC82 | Chemical |  | LAB
FTS102 | mg/L | Fluorotelomer sulfonic acid, 10:2 (10:2 FTS) CAS:120226-60-0 | E3533 | MECP E3533 | PFAs in water by direct injection LC/MS/MS | 2e-05 | Y | Y | Y | 2 | 6/3/2024 | 6/3/2024 | 8/21/2024 3:32:03 PM | 2023-11-15 | 6/27/2024 5:06:31 PM | EMS | EMS | Concentration | mg/L | Y | (fl. conc.) |  | Fluorotelomer sulfonic acid, 10:2 (10:2 FTS) (fl. conc.) | 120226-60-0 | Per- and polyfluoroalkyl substances (PFAS); | Per- and polyfluoroalkyl substances (PFAS); EMS code: FTS102 | Chemical |  | LAB
FTS42 | mg/L | Fluorotelomer sulfonic acid, 4:2 (4:2 FTS) CAS: 757124-72-4 | E3533 | MECP E3533 | PFAs in water by direct injection LC/MS/MS | 2e-05 | Y | Y | Y | 2 | 6/3/2024 | 6/3/2024 | 8/21/2024 3:32:03 PM | 2023-11-15 | 6/27/2024 5:06:31 PM | EMS | EMS | Concentration | mg/L | Y | (fl. conc.) |  | Fluorotelomer sulfonic acid, 4:2 (4:2 FTS) (fl. conc.) | 757124-72-4 | Per- and polyfluoroalkyl substances (PFAS); | Per- and polyfluoroalkyl substances (PFAS); EMS code: FTS42 | Chemical |  | LAB
FTS62 | mg/L | Fluorotelomer sulfonic acid, 6:2 (6:2 FTS) CAS: 27619-97-2 | E3533 | MECP E3533 | PFAs in water by direct injection LC/MS/MS | 2e-05 | Y | Y | Y | 2 | 6/3/2024 | 6/3/2024 | 8/21/2024 3:32:03 PM | 2023-11-15 | 6/27/2024 5:06:31 PM | EMS | EMS | Concentration | mg/L | Y | (fl. conc.) |  | Fluorotelomer sulfonic acid, 6:2 (6:2 FTS) (fl. conc.) | 27619-97-2 | Per- and polyfluoroalkyl substances (PFAS); | Per- and polyfluoroalkyl substances (PFAS); EMS code: FTS62 | Chemical |  | LAB
FTS82 | mg/L | Fluorotelomer sulfonic acid, 8:2 (8:2 FTS) CAS: 39108-34-4 | E3533 | MECP E3533 | PFAs in water by direct injection LC/MS/MS | 2e-05 | Y | Y | Y | 2 | 6/3/2024 | 6/3/2024 | 8/21/2024 3:32:03 PM | 2023-11-15 | 6/27/2024 5:06:31 PM | EMS | EMS | Concentration | mg/L | Y | (fl. conc.) |  | Fluorotelomer sulfonic acid, 8:2 (8:2 FTS) (fl. conc.) | 39108-34-4 | Per- and polyfluoroalkyl substances (PFAS); | Per- and polyfluoroalkyl substances (PFAS); EMS code: FTS82 | Chemical |  | LAB
FTU102 | mg/L | Fluorotelomer unsaturated carboxylic acid, 10:2 (10:2 FTUCA) | E3533 | MECP E3533 | PFAs in water by direct injection LC/MS/MS | 5e-05 | Y | Y | Y | 2 | 6/3/2024 | 6/3/2024 | 8/21/2024 3:32:03 PM | 2023-11-15 | 6/27/2024 5:06:31 PM | EMS | EMS | Concentration | mg/L | Y | (fl. conc.) |  | Fluorotelomer unsaturated carboxylic acid, 10:2 (10:2 FTUCA) (fl. conc.) |  | Per- and polyfluoroalkyl substances (PFAS); | Per- and polyfluoroalkyl substances (PFAS); EMS code: FTU102 | Chemical |  | LAB
FTU62 | mg/L | Fluorotelomer unsaturated carboxylic acid, 6:2 (6:2 FTUCA) | E3533 | MECP E3533 | PFAs in water by direct injection LC/MS/MS | 2e-05 | Y | Y | Y | 2 | 6/3/2024 | 6/3/2024 | 8/21/2024 3:32:03 PM | 2023-11-15 | 6/27/2024 5:06:31 PM | EMS | EMS | Concentration | mg/L | Y | (fl. conc.) |  | Fluorotelomer unsaturated carboxylic acid, 6:2 (6:2 FTUCA) (fl. conc.) |  | Per- and polyfluoroalkyl substances (PFAS); | Per- and polyfluoroalkyl substances (PFAS); EMS code: FTU62 | Chemical |  | LAB
FTU82 | mg/L | Fluorotelomer unsaturated carboxylic acid, 8:2 (8:2 FTUCA) | E3533 | MECP E3533 | PFAs in water by direct injection LC/MS/MS | 2e-05 | Y | Y | Y | 2 | 6/3/2024 | 6/3/2024 | 8/21/2024 3:32:03 PM | 2023-11-15 | 6/27/2024 5:06:31 PM | EMS | EMS | Concentration | mg/L | Y | (fl. conc.) |  | Fluorotelomer unsaturated carboxylic acid, 8:2 (8:2 FTUCA) (fl. conc.) |  | Per- and polyfluoroalkyl substances (PFAS); | Per- and polyfluoroalkyl substances (PFAS); EMS code: FTU82 | Chemical |  | LAB
G001 | mg/L | Glyphosate | X365 | Derivat,extr,HPLC/UV-VIS | An aliquot of the water sample is reacted with dinitrofluoro-benzene (DNFB) to form derivatives of glyphosate and AMPA. Interferences are removed from the reaction mixture by extraction with dichloromethane. A subsequent extraction with iso-butanol removes the derivatives. The derivatives are then back-extracted into a borax solution in preparation for analysis by high performance liquid chromatography. | 0.005 | Y | Y | Y | 19 | 1/3/2001 8:19:00 AM | 7/29/2005 7:01:00 AM | 6/2/2009 12:01:05 PM | 1996-04-12 | 6/3/1998 6:51:36 PM | EMS |  | Concentration | mg/L | Y | (fl. conc.) |  | Glyphosate (fl. conc.) | 1071-83-6 | Pesticides; | Pesticides; EMS code: G001 | Chemical |  | LAB
G001 | mg/L | Glyphosate | P012 | Glyphosate & AMPA in H2O | Glyphosate & AMPA in H2O | 0.005 | Y | Y | Y | 8 | 9/24/2002 10:55:00 AM | 9/24/2002 10:55:00 AM | 3/24/2004 1:04:13 AM | 1996-04-12 | 6/3/1998 6:51:36 PM | EMS |  | Concentration | mg/L | Y | (fl. conc.) |  | Glyphosate (fl. conc.) | 1071-83-6 | Pesticides; | Pesticides; EMS code: G001 | Chemical |  | LAB
G001 | ug/g | Glyphosate | P013 | Gly and AMPA | Gly and AMPA | 0.05 | Y | Y | Y | 45 | 9/10/1997 4:52:00 PM | 2/2/2022 | 1/11/2024 1:07:07 PM | 1996-04-12 | 6/3/1998 6:51:36 PM | EMS |  | MassPerMass | mg/kg | Y | (mass conc.) |  | Glyphosate (mass conc.) | 1071-83-6 | Pesticides; | Pesticides; EMS code: G001 | Chemical |  | LAB
G001 | mg/L | Glyphosate | P014 | Glyphosate | An aliquot of 4.0 ? 0.1 mL of a water sample is spiked with an Internal Standard, Glyphosate-13C2,15N, and derivatized to FMOC-Glyphosate and FMOC-AMPA, then analyzed by LC-MS/MS. | 5 | Y | Y | Y | 67 | 10/8/2019 | 1/12/2022 | 9/2/2022 6:21:53 PM | 2019-10-17 |  | KATFRASE |  | Concentration | mg/L | Y | (fl. conc.) |  | Glyphosate (fl. conc.) | 1071-83-6 | Pesticides; | Pesticides; EMS code: G001 | Chemical |  | LAB
G002 | mg/L | 41-G | P050 | MR's | MR's | 0.02 | Y | Y | Y | 5 |  |  | 10/23/1996 8:47:21 PM | 1996-04-12 | 6/3/1998 6:51:36 PM | EMS |  | Concentration | mg/L | Y | (fl. conc.) |  | 41-G (fl. conc.) |  |  | EMS code: G002 | Chemical |  | LAB
GA-D | mg/L | Gallium Dissolved | X562 | Dissolved metals: FF, FA, HNO3, HR-ICPMS | Ultra trace metals in water are analyzed by high resolution inductively coupled plasma mass spectrometry (HR-ICPMS), based on US EPA Method 200.8, (Oct. 1999). This procedure is intended for pristine field-filtered acid-preserved water samples. The detection limits (LOR) for this test are based on lab instrumental analysis only, not including filtration. ALS-supplied field filtration equipment does not support these LOR. Therefore, because of the high probability of false positives due to filtration, it is strongly recommended that a filtration blank be analyzed to aid in data interpretation. | 5e-05 | Y | Y | Y | 10 | 10/20/2015 | 10/24/2018 | 10/24/2018 12:26:13 PM | 2016-03-24 |  | RDGIBSON |  | Concentration | mg/L | Y | (fl. conc.) |  | Gallium Dissolved (fl. conc.) |  | Metals - Dissolved; | Metals - Dissolved; EMS code: GA-D | Chemical | Dissolved | LAB
GA-D | mg/L | Gallium Dissolved | F032 | Filt.,Acid,ICPMS | There is no acid digestion required for this analysis. A water sample is collected and preserved in 0.2 percent nitric acid. To analyse the extractable metals the sample is shaken, allowed to sit for at least twenty four hours, and a decanted portion is introduced to the spectrometer. When the dissolved metals are required, the samples must first be filtered through a 0.45 micron cellulose acetate filter paper and then followed by preservation in 0.2 percent nitric acid. It is important to note that different laboratories may have very different detection limits for some variables analyzed by ICPMS. | 0.001 | Y | Y | Y | 165 | 11/7/2018 4:00:00 PM | 12/19/2023 | 1/2/2024 4:32:04 PM | 1997-10-20 | 6/22/1998 2:30:34 PM | NPEPPIN | NPEPPIN | Concentration | mg/L | Y | (fl. conc.) |  | Gallium Dissolved (fl. conc.) |  | Metals - Dissolved; | Metals - Dissolved; EMS code: GA-D | Chemical | Dissolved | LAB
GA-D | mg/L | Gallium Dissolved | X556 | Metals in water, dissolved, "in-bottle" digestion, ICP-SFMS | The whole water sample is filtered through a 0.45 micron cellulose acetate filter paper and the filtrate is stored in a HDPE/Teflon bottle. The container is de-capped in a Class 100 workstation, acidified to 2% (v/v) HNO3 and 1% (v/v) HCl and recapped. The sample is "in-bottle" digested in a convection oven, set at 60 degrees C for 16 hours. The sample is cooled and introduced to an inductively coupled argon plasma-sector field mass spectrometer (ICP-SFMS) by pneumatic nebulization. Each respective element is measured at a specific mass to charge ratio, m/z value expressed in amu (atomic mass units), and at a specified intrument resolution. LR: Low Resolution ~300, MR: Mid Resolution ~4000 and HR: High Resolution ~10000. Ref: Method number 02-2003, National Laboratory for Environmental Testing, CCIW, Environment Canada, Burlington, Requested April 2003. | 1e-06 | Y | Y | Y | 2949 | 6/11/2003 | 12/6/2023 2:04:00 AM | 2/14/2024 9:05:32 AM | 2004-01-19 |  | UMULLER |  | Concentration | mg/L | Y | (fl. conc.) |  | Gallium Dissolved (fl. conc.) |  | Metals - Dissolved; | Metals - Dissolved; EMS code: GA-D | Chemical | Dissolved | LAB
GA-D | mg/L | Gallium Dissolved | METC | MET-D-NP-CCMS-VA | Ultra trace metals in water are analyzed by CRC ICPMS, based on US EPA Method 6020B (July 2014). This procedure is intended for pristine field-filtered acid-preserved water samples. The detection limits (LOR) for this test are based on lab instrumental analysis only, not including filtration. ALS-supplied field filtration equipment does not support these LOR. Therefore, because of the high probability of false positives due to filtration, it is strongly recommended that a filtration blank be analysed to aid in data interpretation. | 5e-05 | Y | Y | Y | 394 | 3/30/2020 | 4/9/2024 | 4/29/2024 4:53:06 PM | 2022-10-07 |  | EMS |  | Concentration | mg/L | Y | (fl. conc.) |  | Gallium Dissolved (fl. conc.) |  | Metals - Dissolved; | Metals - Dissolved; EMS code: GA-D | Chemical | Dissolved | LAB
GA-D | mg/L | Gallium Dissolved | MSEA | Metals in Seawater (undigested) by (CRC) ICP-MS | This method describes the multi-elemental determination of trace elements by Collision/Reaction Cell (CRC) ICP-MS. The method measures ions produced by a radio-frequency inductively coupled plasma. Analyte species originating in a liquid are pneumatically nebulized and the resulting aerosol transported by argon gas into the plasma torch. The ions produced are entrained in the plasma gas and extracted, by means of a differentially pumped vacuum interface, into a mass spectrometer. The ions flow through a Collision/Reaction Cell (CRC) where depending on the mode of analysis either the ions passes through the cell unhindered (no gas - classical ICP-MS mode), or collisions/reactions chemistry takes place between the ions and a neutral gas such as hydrogen or helium. Any polyatomic ions which give interference on the analytes are removed. The ions are then sorted according to their mass-to-charge ratios by a quadrupole mass spectrometer. The ions are then sorted according to their mass-to-charge ratios by a quadrupole mass spectrometer. The ions transmitted through the quadrupole are quantified by a channel electron multiplier and the ion information processed by a data handling system. Although most polyatomic interference is removed, any additional interference must be assessed and valid corrections applied or the data flagged to indicate problems. Interference correction must include compensation for background ions contributed by the plasma gas, reagents, and constituents of the sample matrix. Instrument drift as well as suppressions or enhancements of instrument response caused by the sample matrix must be corrected for by the use of internal standards. | 0.0005 | Y | Y | Y | 4 | 4/12/2023 | 4/29/2024 | 5/16/2024 2:00:14 PM | 2023-05-26 |  | EMS |  | Concentration | mg/L | Y | (fl. conc.) |  | Gallium Dissolved (fl. conc.) |  | Metals - Dissolved; | Metals - Dissolved; EMS code: GA-D | Chemical | Dissolved | LAB
GA-E | mg/L | Gallium Extractable | F047 | 24hr HNO3 dil'n,ICP-MS | There is no acid digestion required for this analysis. A water sample is collected and preserved in 0.2 percent nitric acid. To analyse the extractable metals the sample is shaken, allowed to sit for at least twenty four hours, and a decanted portion is introduced to the spectrometer. When the dissolved metals are required, the samples must first be filtered through a 0.45 micron cellulose acetate filter paper and then followed by preservation in 0.2 percent nitric acid. It is important to note that different laboratories may have very different detection limits for some variables analyzed by ICP/MS. | 2e-06 | Y | Y | Y | 2676 | 4/16/1998 10:21:00 AM | 6/26/2003 | 9/11/2019 8:46:18 AM | 1997-10-20 | 6/3/1998 6:51:47 PM | NPEPPIN |  | Concentration | mg/L | Y | (fl. conc.) |  | Gallium Extractable (fl. conc.) |  | Metals - Extractable; | Metals - Extractable; EMS code: GA-E | Chemical | Extractable | LAB
GA-E | mg/L | Gallium Extractable | X553 | Metals in water, extractable, ICP-SFMS | The whole water sample is collected in a HDPE/Teflon bottle. In a Class 100 workstation it is preserved with concentrated Nitric Acid at concentration of 0.2 percent (v/v). The acid preservative extracts the metals from material in the sample and a decanted portion is introduced directly to an inductively coupled argon plasma-sector field mass spectrometer (ICP-SFMS) by pneumatic nebulization. Each respective element is measured at a specific mass to charge ratio, m/z value expressed in amu (atomic mass units), and at a specified instrument resolution. LR: Low Resolution ~300, MR: Mid Resolution ~4000 and HR: High Resolution ~10000. Ref: Method number 02-2061, National Laboratory for Environmental Testing, CCIW, Environment Canada, Burlington. Requested April 2003. | 1e-06 | Y | Y | Y | 2584 | 5/22/2003 | 9/21/2007 | 4/18/2018 9:28:02 AM | 2004-01-19 |  | UMULLER |  | Concentration | mg/L | Y | (fl. conc.) |  | Gallium Extractable (fl. conc.) |  | Metals - Extractable; | Metals - Extractable; EMS code: GA-E | Chemical | Extractable | LAB
GA-T | mg/L | Gallium Total | F121 | Acid,ICPMS | Acidified, no digestion, ICPMS |  | Y | Y | Y | 62 |  |  | 4/3/2008 2:53:49 PM | 2002-07-17 |  | UMULLER |  | Concentration | mg/L | Y | (fl. conc.) |  | Gallium Total (fl. conc.) |  | Metals - Total; | Metals - Total; EMS code: GA-T | Chemical | Total | LAB
GA-T | ppm (S) | Gallium Total | 6042 | Aqua Regia by ICPMS | These were sediment/soils digested with Aqua-Regia (Nitric Acid & Hydrochloric Acids) and then analysis by ICPMS. Results would be reported in ppm s. | 0.004 | Y | Y | Y | 71 | 5/14/2022 | 7/27/2024 | 7/29/2024 9:13:08 AM | 2022-06-28 |  | EMS |  | DimensionlessRatio | ppm | N | (ratio) |  | Gallium Total (ratio) |  | Metals - Total; | Metals - Total; EMS code: GA-T | Chemical | Total | LAB
GA-T | mg/L | Gallium Total | METN | MET-T-NP-CCMS-VA | Ultra trace metals in water are analyzed by CRC ICPMS, based on US EPA Method 6020B (July 2014). The detection limits provided can only be met for undigested samples. This procedure is intended for colorless, non-turbid, acid-preserved water samples (i.e. pristine water samples), having turbidity < 1 NTU and no odor. Where turbidity exceeds 1 NTU, and/or the sample is colored and has an odor, results may be biased low compared to true Total Metals concentrations. ALS recommends that turbidity analysis be requested on samples submitted for this test to aid with interpretation of results. Where turbidity is <1NTU, undigested metals are equivalent to total metals concentrations. | 5e-05 | Y | Y | Y | 280 | 8/6/2019 | 4/4/2024 | 4/29/2024 4:53:06 PM | 2021-06-16 |  | EMS |  | Concentration | mg/L | Y | (fl. conc.) |  | Gallium Total (fl. conc.) |  | Metals - Total; | Metals - Total; EMS code: GA-T | Chemical | Total | LAB
GA-T | mg/L | Gallium Total | F082 | Aliquot,HNO3/HCL dig,ICPMS | Aliquot microwave digested with 3:1 HNO3/HCL. Analysis by ICPMS. For HiVol samples preliminary result are reported in mg/L and loading results are reported in ug/m3 | 5e-05 | Y | Y | Y | 69 | 7/30/2020 | 8/16/2024 | 8/21/2024 11:26:07 AM | 2021-11-03 |  | RDGIBSON |  | Concentration | mg/L | Y | (fl. conc.) |  | Gallium Total (fl. conc.) |  | Metals - Total; | Metals - Total; EMS code: GA-T | Chemical | Total | LAB
GA-T | mg/L | Gallium Total | X555 | Metals in water, total, "in-bottle" digestion, ICP-SFMS | The whole water sample collected in a HDPE/Teflon bottle, is de-capped in a Class 100 workstation, acidified to 2% (v/v) HNO3 and 1% (v/v) HCl and recapped. The sample is "in-bottle" digested in a convection oven, set at 60 degrees C for 16 hours. The sample is cooled and introduced to an inductively coupled argon plasma-sector field mass spectrometer (ICP-SFMS) by pneumatic nebulization. Each respective element is measured at a specific mass to charge ratio, m/z value expressed in amu (atomic mass units), and at resolution. LR: Low Resolution ~300, MR: Mid Resolution ~4000 and HR: High a specified instrument Resolution ~10000. Ref: Method number 02-2003, National Laboratory for Environmental Testing, CCIW, Environment Canada, Burlington. Requested April 2003. | 1e-06 | Y | Y | Y | 17336 | 5/20/2003 | 12/6/2023 2:08:00 AM | 2/14/2024 9:05:32 AM | 2004-01-19 |  | UMULLER |  | Concentration | mg/L | Y | (fl. conc.) |  | Gallium Total (fl. conc.) |  | Metals - Total; | Metals - Total; EMS code: GA-T | Chemical | Total | LAB
GA-T | mg/L | Gallium Total | MSEA | Metals in Seawater (undigested) by (CRC) ICP-MS | This method describes the multi-elemental determination of trace elements by Collision/Reaction Cell (CRC) ICP-MS. The method measures ions produced by a radio-frequency inductively coupled plasma. Analyte species originating in a liquid are pneumatically nebulized and the resulting aerosol transported by argon gas into the plasma torch. The ions produced are entrained in the plasma gas and extracted, by means of a differentially pumped vacuum interface, into a mass spectrometer. The ions flow through a Collision/Reaction Cell (CRC) where depending on the mode of analysis either the ions passes through the cell unhindered (no gas - classical ICP-MS mode), or collisions/reactions chemistry takes place between the ions and a neutral gas such as hydrogen or helium. Any polyatomic ions which give interference on the analytes are removed. The ions are then sorted according to their mass-to-charge ratios by a quadrupole mass spectrometer. The ions are then sorted according to their mass-to-charge ratios by a quadrupole mass spectrometer. The ions transmitted through the quadrupole are quantified by a channel electron multiplier and the ion information processed by a data handling system. Although most polyatomic interference is removed, any additional interference must be assessed and valid corrections applied or the data flagged to indicate problems. Interference correction must include compensation for background ions contributed by the plasma gas, reagents, and constituents of the sample matrix. Instrument drift as well as suppressions or enhancements of instrument response caused by the sample matrix must be corrected for by the use of internal standards. | 0.0005 | Y | Y | Y | 20 | 4/13/2023 | 4/29/2024 | 6/7/2024 10:03:16 AM | 2023-05-26 |  | EMS |  | Concentration | mg/L | Y | (fl. conc.) |  | Gallium Total (fl. conc.) |  | Metals - Total; | Metals - Total; EMS code: GA-T | Chemical | Total | LAB
GA-T | ppm (S) | Gallium Total | MS81 | ALS Minerals: ME-MS81 | ALS Minerals: ME-MS81. A prepared sample (0.100 g) is added to lithium metaborate/lithium tetraborate flux; mixed well and fused in a furnace at 1025°C. The resulting melt is then cooled and dissolved in an acid mixture containing nitric; hydrochloric and hydrofluoric acids. This solution is then analyzed by inductively coupled plasma - mass spectrometry. | 0.1 | Y | Y | Y | 0 |  |  |  | 2024-01-29 |  | EMS |  | DimensionlessRatio | ppm | N | (ratio) |  | Gallium Total (ratio) |  | Metals - Total; | Metals - Total; EMS code: GA-T | Chemical | Total | LAB
GA-T | ppm (S) | Gallium Total | MS41 | ALS Minerals ME-MS41 | ALS Minerals ME-MS41. A prepared sample is digested with aqua regia in a graphite heating block. After cooling; the resulting solution is diluted to with deionized water; mixed and analyzed by ICP-AES. Samples are then analyzed by ICP-MS for the remaining suite of elements. | 0.05 | Y | Y | Y | 0 |  |  |  | 2024-01-29 |  | EMS |  | DimensionlessRatio | ppm | N | (ratio) |  | Gallium Total (ratio) |  | Metals - Total; | Metals - Total; EMS code: GA-T | Chemical | Total | LAB
GA-T | ppm (S) | Gallium Total | MS61 | ALS Minerals ME-MS61 | ALS Minerals ME-MS61. A prepared sample is digested with HNO3; HF; and HClO4 in a graphite heating block to near dryness. After cooling; the sample is boiled in diluted HCl. Bulk to volume with deionized water; mixed and analyzed by ICP-AES and ICP-MS for all elements. | 0.05 | Y | Y | Y | 0 |  |  |  | 2024-01-29 |  | EMS |  | DimensionlessRatio | ppm | N | (ratio) |  | Gallium Total (ratio) |  | Metals - Total; | Metals - Total; EMS code: GA-T | Chemical | Total | LAB
GA-T | ppm (S) | Gallium Total | 6041 | Aqua Regia by ICPMS | These were sediment/soils digested with Aqua-Regia (Nitric Acid & Hydrochloric Acids) and then analysis by ICPMS. Results would be reported in mg/kg and/or %. | 0.004 | Y | Y | Y | 21 | 10/13/2019 | 8/5/2020 | 8/10/2020 3:52:06 PM | 2019-09-06 | 10/21/2019 11:55:03 AM | KATFRASE | KATFRASE | DimensionlessRatio | ppm | N | (ratio) |  | Gallium Total (ratio) |  | Metals - Total; | Metals - Total; EMS code: GA-T | Chemical | Total | LAB
GA-T | mg/L | Gallium Total | 1349 | Tot Metals in Seawater by CRC ICPMS (BC) | Tot Metals in Seawater by CRC ICPMS (BC) | 0.0005 | Y | Y | Y | 98 | 8/16/2019 | 12/5/2021 | 12/13/2021 12:34:16 PM | 2019-09-30 |  | KATFRASE |  | Concentration | mg/L | Y | (fl. conc.) |  | Gallium Total (fl. conc.) |  | Metals - Total; | Metals - Total; EMS code: GA-T | Chemical | Total | LAB
GA-T | mg/L | Gallium Total | X561 | Total metals: FA, HNO3, HR-ICPMS (No Prep)Prep) | Ultra trace metals in water are analyzed by high resolution inductively coupled plasma mass spectrometry (HR-ICPMS), modified from US EPA 200.8, (Revision 5.5). The detection limits provided can only be met for undigested samples. This procedure is intended for colorless, non-turbid, acid-preserved water samples (i.e. pristine water samples), having turbidity < 1 NTU and no odor. Where turbidity exceeds 1 NTU, and/or the sample is colored and has an odor, results may be biased low compared to true Total Metals concentrations. ALS recommends that turbidity analysis be requested on samples submitted for this test to aid with interpretation of results. | 5e-05 | Y | Y | Y | 19 | 10/20/2015 | 10/24/2018 | 10/24/2018 12:26:13 PM | 2016-03-24 |  | RDGIBSON |  | Concentration | mg/L | Y | (fl. conc.) |  | Gallium Total (fl. conc.) |  | Metals - Total; | Metals - Total; EMS code: GA-T | Chemical | Total | LAB
GA-T | ug/g | Gallium Total | EQ02 | EQUIS (ug/g) | EQUIS (ug/g) | 10 | N | Y | Y | 2 |  |  | 10/22/1996 9:46:36 PM | 1996-04-12 | 6/3/1998 6:51:36 PM | EMS |  | MassPerMass | mg/kg | Y | (mass conc.) |  | Gallium Total (mass conc.) |  | Metals - Total; | Metals - Total; EMS code: GA-T | Chemical | Total | LAB
GCMS | None | GCMS Scan | SCAN | SolventExtr,varioussolvents/methods,HRGC/LRMS | Sample extraction methods vary due to properties unique to the sample type. Water samples are typically liquid / liquid DCM extraction, however the analyst may decide that another solvent such as Hexane is more appropriate. For sediment and biota samples, extraction methods include Microwave Assisted Process and wrist or platform shaker. Again solvent choice is at the discretion of teh analyst and may include DCM, Methanol, Acetone, Hexane, or others. For samples such as raw products, no extraction may be necessary and simple dilution in an appropriate solvent may be all that is required. The extract is analyzed using High Resolution Gas Chromatography / Low Resolution Mass Spectrometry (HRGC/LRMS). MS data is acquired in Full Scan Mode (Scan) and normally no quantitation is performed. |  | Y | Y | Y | 16 | 3/3/2000 7:36:00 AM | 1/9/2002 8:26:00 AM | 6/24/2003 10:41:00 AM | 2000-10-25 |  | NPEPPIN |  | None | None | N | (unitless) |  | GCMS Scan (unitless) |  |  | EMS code: GCMS | Chemical |  | LAB
GD-T | ppm (S) | Gadolinium | MS81 | ALS Minerals: ME-MS81 | ALS Minerals: ME-MS81. A prepared sample (0.100 g) is added to lithium metaborate/lithium tetraborate flux; mixed well and fused in a furnace at 1025°C. The resulting melt is then cooled and dissolved in an acid mixture containing nitric; hydrochloric and hydrofluoric acids. This solution is then analyzed by inductively coupled plasma - mass spectrometry. | 0.05 | Y | Y | Y | 0 |  |  |  | 2024-01-29 |  | EMS |  | DimensionlessRatio | ppm | N | (ratio) |  | Gadolinium (ratio) |  | Metals; | Metals; EMS code: GD-T | Chemical |  | LAB
GE-T | ppm (S) | Germanium | MS81 | ALS Minerals: ME-MS81 | ALS Minerals: ME-MS81. A prepared sample (0.100 g) is added to lithium metaborate/lithium tetraborate flux; mixed well and fused in a furnace at 1025°C. The resulting melt is then cooled and dissolved in an acid mixture containing nitric; hydrochloric and hydrofluoric acids. This solution is then analyzed by inductively coupled plasma - mass spectrometry. | 5 | Y | Y | Y | 0 |  |  |  | 2024-01-29 |  | EMS |  | DimensionlessRatio | ppm | N | (ratio) |  | Germanium (ratio) |  | Metals; | Metals; EMS code: GE-T | Chemical |  | LAB
GE-T | ppm (S) | Germanium | 6041 | Aqua Regia by ICPMS | These were sediment/soils digested with Aqua-Regia (Nitric Acid & Hydrochloric Acids) and then analysis by ICPMS. Results would be reported in mg/kg and/or %. | 0.005 | Y | Y | Y | 21 | 10/13/2019 | 8/5/2020 | 8/10/2020 3:52:06 PM | 2019-09-06 | 10/21/2019 11:55:07 AM | KATFRASE | KATFRASE | DimensionlessRatio | ppm | N | (ratio) |  | Germanium (ratio) |  | Metals; | Metals; EMS code: GE-T | Chemical |  | LAB
GE-T | ppm (S) | Germanium | MS41 | ALS Minerals ME-MS41 | ALS Minerals ME-MS41. A prepared sample is digested with aqua regia in a graphite heating block. After cooling; the resulting solution is diluted to with deionized water; mixed and analyzed by ICP-AES. Samples are then analyzed by ICP-MS for the remaining suite of elements. | 0.05 | Y | Y | Y | 0 |  |  |  | 2024-01-29 |  | EMS |  | DimensionlessRatio | ppm | N | (ratio) |  | Germanium (ratio) |  | Metals; | Metals; EMS code: GE-T | Chemical |  | LAB
GE-T | ppm (S) | Germanium | MS61 | ALS Minerals ME-MS61 | ALS Minerals ME-MS61. A prepared sample is digested with HNO3; HF; and HClO4 in a graphite heating block to near dryness. After cooling; the sample is boiled in diluted HCl. Bulk to volume with deionized water; mixed and analyzed by ICP-AES and ICP-MS for all elements. | 0.05 | Y | Y | Y | 0 |  |  |  | 2024-01-29 |  | EMS |  | DimensionlessRatio | ppm | N | (ratio) |  | Germanium (ratio) |  | Metals; | Metals; EMS code: GE-T | Chemical |  | LAB
GE-T | ppm (S) | Germanium | 6042 | Aqua Regia by ICPMS | These were sediment/soils digested with Aqua-Regia (Nitric Acid & Hydrochloric Acids) and then analysis by ICPMS. Results would be reported in ppm s. | 0.005 | Y | Y | Y | 71 | 5/14/2022 | 7/27/2024 | 7/29/2024 9:13:08 AM | 2022-06-28 |  | EMS |  | DimensionlessRatio | ppm | N | (ratio) |  | Germanium (ratio) |  | Metals; | Metals; EMS code: GE-T | Chemical |  | LAB
GLBR | mg/L | Glyburide CAS: 10238-21-8 | XMSP | HPLC and Tri. Quadpole MS/MS | EPA 1694 An aliquot of filtered sample containing 5% organic is injected directly. Each compound is separated by reversed phase HPLC. Analyzed by Electro Spray Ionization MS/MS detection using Triple Quadrupole MS/MS detector. Quantiitation with internal or external standard. |  | Y | Y | Y | 12 | 5/8/2020 | 6/4/2024 | 8/21/2024 3:32:03 PM | 2022-05-04 |  | EMS |  | Concentration | mg/L | Y | (fl. conc.) |  | Glyburide (fl. conc.) | 10238-21-8 | Pesticides; | Pesticides; EMS code: GLBR | Chemical |  | LAB
GLPZ | mg/L | Glipizide CAS: 29094-61-9 | XMSP | HPLC and Tri. Quadpole MS/MS | EPA 1694 An aliquot of filtered sample containing 5% organic is injected directly. Each compound is separated by reversed phase HPLC. Analyzed by Electro Spray Ionization MS/MS detection using Triple Quadrupole MS/MS detector. Quantiitation with internal or external standard. |  | Y | Y | Y | 12 | 5/8/2020 | 6/4/2024 | 8/21/2024 3:32:03 PM | 2022-05-04 |  | EMS |  | Concentration | mg/L | Y | (fl. conc.) |  | Glipizide (fl. conc.) | 29094-61-9 | Pesticides; | Pesticides; EMS code: GLPZ | Chemical |  | LAB
H-3- | Bq/L | Tritium | X241 | Liquid Scintillation | Liquid Scintillation | 1 | Y | Y | Y | 276 |  |  | 10/24/1996 7:20:16 AM | 1996-04-12 | 6/3/1998 6:51:36 PM | EMS |  | Radioactivity | Bq/L | Y | (rad./vol.) |  | Tritium (rad./vol.) | 10028-17-8 | Radionuclide; | Radionuclide; EMS code: H-3- | Chemical |  | LAB
H-TE | ug/g | Hydrocarbons;Extract | FS07 | Freon 113 Extr,Infrared | Freon 113 soluble compounds are extracted from an acidified effluent sample and made up to a measured volume. Sediment samples are acidified then extracted using a wrist-action shaker. The "Oils and Greases" content of the Freon extracts is determined using an Infrared Spectrophotometer with commercial quantitation software. Quantitation is based on the comparison of the Absorbance difference between samples and a synthetic oil standard. As there is no sample Absorbance band at 4000 cm-1 it represents a useful reference value to compensate for minor differences between spectrophotometer quartz cells. Sample Absorbance at 2935 cm-1 represents the C-H stretching frequency of the aliphatic hydrocarbons of which typical environmental "Oils and Greases" are comprised. "Hydrocarbons" are quantitated in the same manner after treating the Freon 113 extract with Silica Gel. | 5 | Y | Y | Y | 164 | 5/25/2000 2:47:00 PM | 1/4/2002 8:07:00 AM | 5/26/2003 1:56:28 PM | 1997-04-28 | 6/3/1998 6:51:46 PM | MPAK |  | MassPerMass | mg/kg | Y | (mass conc.) |  | Hydrocarbons;Extract (mass conc.) |  | Petroleum Hydrocarbons; | Petroleum Hydrocarbons; EMS code: H-TE | Chemical |  | LAB
H-TE | mg/L | Hydrocarbons;Extract | SHEG | Shake30min,Hexane Ext, GC/FID | Water samples are spiked with 2-BBTF surrogate, and extracted by reciprocal action shaker for 30 minutes using a single micro-extraction with 2mL hexane. After extraction, hexane extracts are dispensed into GC vials for GC-FID analysis. | 0.25 | Y | Y | Y | 153 | 3/31/2016 | 1/3/2019 | 1/15/2019 12:11:08 PM | 2016-04-05 |  | RDGIBSON |  | Concentration | mg/L | Y | (fl. conc.) |  | Hydrocarbons;Extract (fl. conc.) |  | Petroleum Hydrocarbons; | Petroleum Hydrocarbons; EMS code: H-TE | Chemical |  | LAB
H-TE | mg/L | Hydrocarbons;Extract | X368 | Freon extract,IR absorb. | Freon extract,IR absorb. | 0.2 | Y | Y | Y | 87 |  |  | 1/20/2022 4:45:11 PM | 1996-04-12 | 6/3/1998 6:51:36 PM | EMS |  | Concentration | mg/L | Y | (fl. conc.) |  | Hydrocarbons;Extract (fl. conc.) |  | Petroleum Hydrocarbons; | Petroleum Hydrocarbons; EMS code: H-TE | Chemical |  | LAB
H-TE | mg/L | Hydrocarbons;Extract | F103 | Calc,LEPH+HEPH,nC10-nC32 | Calc;LEPH+HEPH;nC10-nC32 |  | Y | Y | Y | 5 | 1/3/2001 8:44:00 AM | 4/5/2001 1:33:00 PM | 9/10/2002 1:07:10 PM | 2000-11-27 |  | NPEPPIN |  | Concentration | mg/L | Y | (fl. conc.) |  | Hydrocarbons;Extract (fl. conc.) |  | Petroleum Hydrocarbons; | Petroleum Hydrocarbons; EMS code: H-TE | Chemical |  | LAB
H-TE | mg/L | Hydrocarbons;Extract | X366 | Dichlorometh extr,GC/FID | The sample is extracted with dichloromethane in a separatory funnel. The extract is dried, then reduced to a known volume. The final extract is then analyzed with capillary column gas chromatography with flame ionization detection. The results are determined by obtaining the total area under the chromato-graphic curve between n-C10 and n-C30, then quantifying by external calibration. | 1 | Y | Y | Y | 333 |  |  | 8/28/2018 3:13:59 PM | 1996-04-12 | 6/3/1998 6:51:36 PM | EMS |  | Concentration | mg/L | Y | (fl. conc.) |  | Hydrocarbons;Extract (fl. conc.) |  | Petroleum Hydrocarbons; | Petroleum Hydrocarbons; EMS code: H-TE | Chemical |  | LAB
H-TE | mg/L | Hydrocarbons;Extract | F020 | Freon 113 Extr,Infrared | Freon 113 soluble compounds are extracted from an acidified effluent sample and made up to a measured volume. Sediment samples are acidified then extracted using a wrist-action shaker. The "Oils and Greases" content of the Freon extracts is determined using an Infrared Spectrophotometer with commercial quantitation software. Quantitation is based on the comparison of the Absorbance difference between samples and a synthetic oil standard. As there is no sample Absorbance band at 4000 cm-1 it represents a useful reference value to compensate for minor differences between spectrophotometer quartz cells. Sample Absorbance at 2935 cm-1 represents the C-H stretching frequency of the aliphatic hydrcarbons of which typical environmental "Oils and Greases" are comprised. "Hydrocarbons" are quantitated in the same manner after treating the Freon 113 extract with Silica Gel. | 0.2 | Y | Y | Y | 169 | 3/24/1999 4:06:00 PM | 10/26/2001 1:55:00 PM | 5/20/2003 4:08:47 PM | 1997-04-28 | 6/3/1998 6:51:46 PM | MPAK |  | Concentration | mg/L | Y | (fl. conc.) |  | Hydrocarbons;Extract (fl. conc.) |  | Petroleum Hydrocarbons; | Petroleum Hydrocarbons; EMS code: H-TE | Chemical |  | LAB
H-TP | mg/L | Total Petroleum Hydrocarbons (TPH) | X370 | Freon extr,silicic ac,IR | The sample is extracted with 1,1,2-trichloro-1,2,2-trifluoroethane in a separatory funnel. The extract is diluted to a known volume, and analyzed with an infrared spectro-photometer. The result is calculated by comparison of the absorbance of the sample extract to that of a prepared standard mixture. When a silica gel clean-up step is incorporated into this procedure, it is referred to as Total Petroleum Hydrocarbons (TPH) rather than as Total Extractable Hydrocarbon. | 0.2 | Y | Y | Y | 4 |  |  | 12/17/2001 1:49:50 PM | 1996-04-12 | 6/3/1998 6:51:36 PM | EMS |  | Concentration | mg/L | Y | (fl. conc.) |  | Total Petroleum Hydrocarbons (TPH) (fl. conc.) |  | Petroleum Hydrocarbons; | Petroleum Hydrocarbons; EMS code: H-TP | Chemical |  | LAB
H-TP | None | Total Petroleum Hydrocarbons (TPH) | FQ01 | Hydrocarbons&Petroleum Products ID screening | Flame ionization capillary gas chromatography is the primary means used for determining and matching the type of hydrocarbons present. Initial screening and all "high-boiling" hydrocarbons (those having a boiling range greater than or equal to diesel fuel) are analysed using a bonded-phase polymethyl/phenyl siloxane GC column. All "low-boiling" hydrocarbons (those having a boiling range less than diesel fuel) are preferably analysed using a bonded-phase polymethyl siloxane GC column. Identification is based on the recognition and matching of peak patterns in the chromatogram and comparison with those produced by "source" samples or commercial petroleum hydrocarbons standards. The distribution and relative concentrations of the components are considered unique to a given oil. The carbon range is determined by comparing the sample with a commercial "n-C5" to "n-C40" hydrocarbon calibration mix. Infrared spectrophotometry (IR) is used to confirm a hydrocarbon sample as either a petroleum or vegetable-animal in origin. The spectrum is examined for the presence or absence of a strong carbonyl absorption at about 1740 cm-1. The absence of this peak is good evidence for a petroleum origin of the sample. The spectrum can also be compared to that of other hydrocarbon standards. Ultraviolet/Visible spectrophotometry (UV/Vis) is used to determine the theoretical ratio of the aromatic compounds containing 2 rings to those containing 3 or more rings. Trace metals present in the sample are examined by ICP emission spectroscopy. The Nickel-Vanadium ratio can be a good indicator when used for non-distillate petroleum hydrocarbons, however, "profile matching" of the concentrations of selected metals is generally more useful. Note: Results from this test are non-numeric and therefore are reported in the comment field. |  | Y | Y | Y | 49 | 2/3/1997 11:27:00 AM | 10/19/2001 3:46:00 PM | 5/21/2003 8:53:01 AM | 1997-01-14 | 6/3/1998 6:51:46 PM | MPAK |  | None | None | N | (unitless) |  | Total Petroleum Hydrocarbons (TPH) (unitless) |  | Petroleum Hydrocarbons; | Petroleum Hydrocarbons; EMS code: H-TP | Chemical |  | LAB
H-TP | mg/L | Total Petroleum Hydrocarbons (TPH) | E3511 | EPA 3511/BCMOE EPHw | EPA 3511/BCMOE EPHw; Hexane MicroExtraction (Base/Neutral) / Gas Chromatography (GC-FID). Sample extracts are subjected to column silica gel treatment prior to analysis by GC-FID for BC hydrocarbon fractions (EPHsg). | 0.5 | Y | Y | Y | 4 | 3/12/2024 | 5/10/2024 | 5/28/2024 1:16:02 PM | 2024-03-14 |  | RDGIBSON |  | Concentration | mg/L | Y | (fl. conc.) |  | Total Petroleum Hydrocarbons (TPH) (fl. conc.) |  | Petroleum Hydrocarbons; | Petroleum Hydrocarbons; EMS code: H-TP | Chemical |  | LAB
H-TP | mg/L | Total Petroleum Hydrocarbons (TPH) | F103 | Calc,LEPH+HEPH,nC10-nC32 | Calc;LEPH+HEPH;nC10-nC32 |  | Y | Y | Y | 1897 |  |  | 7/30/2024 6:32:12 AM | 2000-11-27 |  | NPEPPIN |  | Concentration | mg/L | Y | (fl. conc.) |  | Total Petroleum Hydrocarbons (TPH) (fl. conc.) |  | Petroleum Hydrocarbons; | Petroleum Hydrocarbons; EMS code: H-TP | Chemical |  | LAB
H001 | mg/L | Heptachlor | E660 | EPA 8270E (mod) | EPA 8270E (mod) Miscellaneous Pesticides by GC-MS | 2e-06 | Y | Y | Y | 2 | 9/14/2022 | 9/14/2022 | 12/8/2022 9:48:14 AM | 2022-12-06 |  | EMS |  | Concentration | mg/L | Y | (fl. conc.) |  | Heptachlor (fl. conc.) | 76-44-8 | Pesticides; | Pesticides; EMS code: H001 | Chemical |  | LAB
H001 | mg/L | Heptachlor | F027 | Hexane Extr,GC/ECD | Effluent samples are solvent extracted with Hexane and the extract dried over Sodium Sulfate, concentrated and cleaned up, as required, using Florisil, activated Copper metal, concentrated Sulfuric Acid and Mercury, to remove interferences. The extracts are then analysed using megabore column Gas Chromatography with Electron Capture detection. | 5e-05 | Y | Y | Y | 20 | 3/3/2000 7:36:00 AM | 1/22/2002 10:58:00 AM | 6/18/2003 2:53:44 PM | 2002-06-05 |  | UMULLER |  | Concentration | mg/L | Y | (fl. conc.) |  | Heptachlor (fl. conc.) | 76-44-8 | Pesticides; | Pesticides; EMS code: H001 | Chemical |  | LAB
H001 | ug/g | Heptachlor | ZN03 | DCM+Ac.Extr:GC/MS | DCM+Ac.Extr:GC/MS EPA 8270E (mod) | 0.05 | Y | Y | Y | 19 | 1/17/2022 | 11/14/2022 | 1/11/2024 1:07:07 PM | 1996-04-12 | 11/16/1999 2:27:40 PM | EMS | NPEPPIN | MassPerMass | mg/kg | Y | (mass conc.) |  | Heptachlor (mass conc.) | 76-44-8 | Pesticides; | Pesticides; EMS code: H001 | Chemical |  | LAB
H001 | ug/g | Heptachlor | E8081 | EPA 8081 | See methods EPA 3540, 3630, 3660, 8081 | 0.002 | Y | Y | Y | 9 | 3/18/2020 | 3/23/2020 | 2/22/2023 12:51:03 PM | 2022-07-29 |  | EMS |  | MassPerMass | mg/kg | Y | (mass conc.) |  | Heptachlor (mass conc.) | 76-44-8 | Pesticides; | Pesticides; EMS code: H001 | Chemical |  | LAB
H001 | ug/g | Heptachlor | X213 | Hex/AcExt,Part,Chrom,GC | Hex/AcExt;Part;Chrom;GC | 0.0005 | Y | Y | Y | 187 | 4/20/2002 10:10:00 AM | 10/13/2002 10:30:00 AM | 2/26/2004 1:05:28 AM | 1996-04-12 | 6/3/1998 6:51:36 PM | EMS |  | MassPerMass | mg/kg | Y | (mass conc.) |  | Heptachlor (mass conc.) | 76-44-8 | Pesticides; | Pesticides; EMS code: H001 | Chemical |  | LAB
H001 | mg/L | Heptachlor | ZN01 | DCM+Ac.Extr:GC/MS | DCM+Ac.Extr:GC/MS | 0.0005 | Y | Y | Y | 52 | 10/3/2002 | 3/29/2012 10:15:00 PM | 5/3/2012 9:27:51 AM | 1996-04-12 | 11/16/1999 4:06:27 PM | EMS | NPEPPIN | Concentration | mg/L | Y | (fl. conc.) |  | Heptachlor (fl. conc.) | 76-44-8 | Pesticides; | Pesticides; EMS code: H001 | Chemical |  | LAB
H001 | mg/L | Heptachlor | X268 | CH2Cl2 ext:GC/MS | CH2Cl2 ext:GC/MS |  | Y | Y | Y | 16 | 12/29/2021 | 9/2/2022 | 11/10/2022 9:54:07 AM | 2022-04-08 |  | EMS |  | Concentration | mg/L | Y | (fl. conc.) |  | Heptachlor (fl. conc.) | 76-44-8 | Pesticides; | Pesticides; EMS code: H001 | Chemical |  | LAB
H001 | ug/g wet | Heptachlor | X220 | Hex/AcExt,Part,Chrom,GC | Hex/AcExt;Part;Chrom;GC | 0.0002 | Y | Y | Y | 72 |  |  | 10/24/1996 10:01:25 AM | 1996-04-12 | 6/3/1998 6:51:42 PM | EMS |  | MassPerMass | µg/g | Y | (mass conc.) | (wet) | Heptachlor (wet) (mass conc.) | 76-44-8 | Pesticides; | Pesticides; EMS code: H001 | Chemical |  | LAB
H001 | mg/L | Heptachlor | EQ01 | EQUIS (mg/L) | EQUIS (mg/L) | 0 | N | Y | Y | 17 |  |  | 10/22/1996 11:26:54 PM | 1996-04-12 | 6/3/1998 6:51:36 PM | EMS |  | Concentration | mg/L | Y | (fl. conc.) |  | Heptachlor (fl. conc.) | 76-44-8 | Pesticides; | Pesticides; EMS code: H001 | Chemical |  | LAB
H001 | mg/L | Heptachlor | MW03 | OCP1-LL-SF-ECD-VA | This analysis is carried out using procedures adapted from "Test Methods for Evaluating Solid Waste" SW-846, Methods 3510, 3630, 3660 and 8081, published by the United States Environmental Protection Agency (EPA). The procedure involves extraction of the entire water sample with dichloromethane. The extract then undergoes one or more of the following clean-up procedures (if required): silica gel clean-up and/or sulphur clean-up. The final extract is analyzed by dual capillary column gas chromatography with electron capture detection (GC/ECD). | 0.0001 | Y | Y | Y | 10 | 9/20/2016 | 9/21/2016 | 8/15/2019 2:42:34 PM | 2019-08-15 |  | KATFRASE |  | Concentration | mg/L | Y | (fl. conc.) |  | Heptachlor (fl. conc.) | 76-44-8 | Pesticides; | Pesticides; EMS code: H001 | Chemical |  | LAB
H002 | mg/L | Heptachlor Epoxide | F027 | Hexane Extr,GC/ECD | Effluent samples are solvent extracted with Hexane and the extract dried over Sodium Sulfate, concentrated and cleaned up, as required, using Florisil, activated Copper metal, concentrated Sulfuric Acid and Mercury, to remove interferences. The extracts are then analysed using megabore column Gas Chromatography with Electron Capture detection. | 5e-05 | Y | Y | Y | 20 | 3/3/2000 7:36:00 AM | 1/22/2002 10:58:00 AM | 6/18/2003 2:53:44 PM | 2002-06-05 |  | UMULLER |  | Concentration | mg/L | Y | (fl. conc.) |  | Heptachlor Epoxide (fl. conc.) | 1024-57-3 | Pesticides; | Pesticides; EMS code: H002 | Chemical |  | LAB
H002 | ug/g | Heptachlor Epoxide | ZN03 | DCM+Ac.Extr:GC/MS | DCM+Ac.Extr:GC/MS EPA 8270E (mod) | 0.05 | Y | Y | Y | 19 | 1/17/2022 | 11/14/2022 | 1/11/2024 1:07:07 PM | 2003-06-16 |  | UMULLER |  | MassPerMass | mg/kg | Y | (mass conc.) |  | Heptachlor Epoxide (mass conc.) | 1024-57-3 | Pesticides; | Pesticides; EMS code: H002 | Chemical |  | LAB
H002 | ug/g | Heptachlor Epoxide | X213 | Hex/AcExt,Part,Chrom,GC | Hex/AcExt;Part;Chrom;GC | 0.01 | Y | Y | Y | 187 | 4/20/2002 10:10:00 AM | 10/13/2002 10:30:00 AM | 2/26/2004 1:05:28 AM | 1996-04-12 | 6/3/1998 6:51:36 PM | EMS |  | MassPerMass | mg/kg | Y | (mass conc.) |  | Heptachlor Epoxide (mass conc.) | 1024-57-3 | Pesticides; | Pesticides; EMS code: H002 | Chemical |  | LAB
H002 | mg/L | Heptachlor Epoxide | MW03 | OCP1-LL-SF-ECD-VA | This analysis is carried out using procedures adapted from "Test Methods for Evaluating Solid Waste" SW-846, Methods 3510, 3630, 3660 and 8081, published by the United States Environmental Protection Agency (EPA). The procedure involves extraction of the entire water sample with dichloromethane. The extract then undergoes one or more of the following clean-up procedures (if required): silica gel clean-up and/or sulphur clean-up. The final extract is analyzed by dual capillary column gas chromatography with electron capture detection (GC/ECD). | 3e-06 | Y | Y | Y | 10 | 9/20/2016 | 9/21/2016 | 8/15/2019 2:42:34 PM | 2019-08-15 |  | KATFRASE |  | Concentration | mg/L | Y | (fl. conc.) |  | Heptachlor Epoxide (fl. conc.) | 1024-57-3 | Pesticides; | Pesticides; EMS code: H002 | Chemical |  | LAB
H002 | ug/g | Heptachlor Epoxide | E8081 | EPA 8081 | See methods EPA 3540, 3630, 3660, 8081 | 0.001 | Y | Y | Y | 9 | 3/18/2020 | 3/23/2020 | 2/22/2023 12:51:03 PM | 2022-07-29 |  | EMS |  | MassPerMass | mg/kg | Y | (mass conc.) |  | Heptachlor Epoxide (mass conc.) | 1024-57-3 | Pesticides; | Pesticides; EMS code: H002 | Chemical |  | LAB
H002 | mg/L | Heptachlor Epoxide | E660 | EPA 8270E (mod) | EPA 8270E (mod) Miscellaneous Pesticides by GC-MS | 2e-06 | Y | Y | Y | 2 | 9/14/2022 | 9/14/2022 | 12/8/2022 9:48:14 AM | 2022-12-06 |  | EMS |  | Concentration | mg/L | Y | (fl. conc.) |  | Heptachlor Epoxide (fl. conc.) | 1024-57-3 | Pesticides; | Pesticides; EMS code: H002 | Chemical |  | LAB
H002 | mg/L | Heptachlor Epoxide | P017 | OC's and OP's | OC's and OP's | 4e-05 | Y | Y | Y | 1 | 7/28/2005 7:03:00 AM | 7/28/2005 7:03:00 AM | 6/2/2009 12:01:06 PM | 1996-04-12 | 6/3/1998 6:51:36 PM | EMS |  | Concentration | mg/L | Y | (fl. conc.) |  | Heptachlor Epoxide (fl. conc.) | 1024-57-3 | Pesticides; | Pesticides; EMS code: H002 | Chemical |  | LAB
H002 | mg/L | Heptachlor Epoxide | X268 | CH2Cl2 ext:GC/MS | CH2Cl2 ext:GC/MS |  | Y | Y | Y | 16 | 12/29/2021 | 9/2/2022 | 11/10/2022 9:54:07 AM | 2022-04-08 |  | EMS |  | Concentration | mg/L | Y | (fl. conc.) |  | Heptachlor Epoxide (fl. conc.) | 1024-57-3 | Pesticides; | Pesticides; EMS code: H002 | Chemical |  | LAB
H002 | ug/g wet | Heptachlor Epoxide | X220 | Hex/AcExt,Part,Chrom,GC | Hex/AcExt;Part;Chrom;GC | 0.01 | Y | Y | Y | 73 |  |  | 10/24/1996 10:01:25 AM | 1996-04-12 | 6/3/1998 6:51:45 PM | EMS |  | MassPerMass | µg/g | Y | (mass conc.) | (wet) | Heptachlor Epoxide (wet) (mass conc.) | 1024-57-3 | Pesticides; | Pesticides; EMS code: H002 | Chemical |  | LAB
H003 | ug/g | Hexachlorobenzene | ZN03 | DCM+Ac.Extr:GC/MS | DCM+Ac.Extr:GC/MS EPA 8270E (mod) | 0.02 | Y | Y | Y | 43 | 4/20/2002 10:10:00 AM | 11/14/2022 | 1/11/2024 1:07:07 PM | 1996-04-12 | 11/16/1999 2:27:40 PM | EMS | NPEPPIN | MassPerMass | mg/kg | Y | (mass conc.) |  | Hexachlorobenzene (mass conc.) | 118-74-1 | Pesticides; | Pesticides; EMS code: H003 | Chemical |  | LAB
H003 | mg/L | Hexachlorobenzene | ZN01 | DCM+Ac.Extr:GC/MS | DCM+Ac.Extr:GC/MS | 0.0002 | Y | Y | Y | 60 | 10/3/2002 | 3/29/2012 10:15:00 PM | 5/3/2012 9:27:51 AM | 1996-04-12 | 11/16/1999 4:06:27 PM | EMS | NPEPPIN | Concentration | mg/L | Y | (fl. conc.) |  | Hexachlorobenzene (fl. conc.) | 118-74-1 | Pesticides; | Pesticides; EMS code: H003 | Chemical |  | LAB
H003 | mg/L | Hexachlorobenzene | E660 | EPA 8270E (mod) | EPA 8270E (mod) Miscellaneous Pesticides by GC-MS | 2e-06 | Y | Y | Y | 2 | 9/14/2022 | 9/14/2022 | 12/8/2022 9:48:14 AM | 2022-12-06 |  | EMS |  | Concentration | mg/L | Y | (fl. conc.) |  | Hexachlorobenzene (fl. conc.) | 118-74-1 | Pesticides; | Pesticides; EMS code: H003 | Chemical |  | LAB
H003 | mg/L | Hexachlorobenzene | X268 | CH2Cl2 ext:GC/MS | CH2Cl2 ext:GC/MS |  | Y | Y | Y | 16 | 12/29/2021 | 9/2/2022 | 11/10/2022 9:54:07 AM | 2022-04-08 |  | EMS |  | Concentration | mg/L | Y | (fl. conc.) |  | Hexachlorobenzene (fl. conc.) | 118-74-1 | Pesticides; | Pesticides; EMS code: H003 | Chemical |  | LAB
H005 | mg/L | Heptenofos | ZN01 | DCM+Ac.Extr:GC/MS | DCM+Ac.Extr:GC/MS | 0.0005 | Y | Y | Y | 1 | 7/29/2005 7:10:00 AM | 7/29/2005 7:10:00 AM | 6/2/2009 12:01:06 PM | 2003-06-17 |  | UMULLER |  | Concentration | mg/L | Y | (fl. conc.) |  | Heptenofos (fl. conc.) |  | Pesticides; | Pesticides; EMS code: H005 | Chemical |  | LAB
H008 | mg/L | Hexazinone | ZN01 | DCM+Ac.Extr:GC/MS | DCM+Ac.Extr:GC/MS | 0.0001 | Y | Y | Y | 52 | 10/3/2002 | 3/29/2012 10:15:00 PM | 5/3/2012 9:27:51 AM | 1999-11-17 |  | NPEPPIN |  | Concentration | mg/L | Y | (fl. conc.) |  | Hexazinone (fl. conc.) | 51235-04-2 | Pesticides; | Pesticides; EMS code: H008 | Chemical |  | LAB
H008 | ug/g | Hexazinone | FS61 | DCM +MeOH extr,HRGC-NPD | Sediment samples are shaken with a mixture of Dichloromethane and Methanol, filtered then re-extracted with Dichloromethane. The extract is concentrated and analysed using High Resolution Gas Chromatography with a Thermionic Specific Detector (TSD), alternately referred to as a Nitrogen Phosphorus Detector (NPD). | 0.05 | Y | Y | Y | 34 | 4/20/2002 3:40:00 PM | 7/16/2002 3:20:00 PM | 2/26/2004 1:05:28 AM | 1998-10-02 |  | NPEPPIN |  | MassPerMass | mg/kg | Y | (mass conc.) |  | Hexazinone (mass conc.) | 51235-04-2 | Pesticides; | Pesticides; EMS code: H008 | Chemical |  | LAB
H008 | mg/L | Hexazinone | F056 | DCM extr,HRGC-NPD | Effluent samples are extracted with Dichloromethane. The extract is concentrated and analysed using High Resolution Gas Chromatography with a Thermionic Specific Detector (TSD), alternately referred to as a Nitrogen Phosphorus Detector (NPD). | 0.02 | Y | Y | Y | 12 | 4/3/2000 10:51:00 AM | 2/16/2001 3:36:00 PM | 9/11/2002 9:34:56 AM | 1998-10-02 |  | NPEPPIN |  | Concentration | mg/L | Y | (fl. conc.) |  | Hexazinone (fl. conc.) | 51235-04-2 | Pesticides; | Pesticides; EMS code: H008 | Chemical |  | LAB
H009 | mg/L | Petroleum Distillate | XSCN | SCAN | SCAN | 0.02 | Y | Y | Y | 7 |  |  | 10/24/1996 11:24:27 AM | 1996-04-12 | 6/3/1998 6:51:36 PM | EMS |  | Concentration | mg/L | Y | (fl. conc.) |  | Petroleum Distillate (fl. conc.) |  | Petroleum Hydrocarbons; | Petroleum Hydrocarbons; EMS code: H009 | Chemical |  | LAB
H009 | mg/L | Petroleum Distillate | P010 | HC's | HC's | 0.01 | Y | Y | Y | 61 |  |  | 2/13/1998 1:02:12 AM | 1996-04-12 | 6/3/1998 6:51:36 PM | EMS |  | Concentration | mg/L | Y | (fl. conc.) |  | Petroleum Distillate (fl. conc.) |  | Petroleum Hydrocarbons; | Petroleum Hydrocarbons; EMS code: H009 | Chemical |  | LAB
H010 | pg/g wet | 6Chldibenzodioxin | F119 | Solids,Isotope dil,HRGC/HRMS | EPA Method 1613: for Sediments and Fish tissue developed for isomer specific determination of the 2,3,7,8-substituted, tetra-through octa-chlorinated, dibenzo-p-dioxins and dibenzofurans in aqueous, solid and tissue matrices by isotope dilution, high resolution capillary column gas chromatography (HRGC)/high resolution mass spectrometry (HRMS) |  | Y | Y | Y | 5 | 6/16/2000 6:33:00 AM | 10/4/2000 2:53:00 PM | 6/18/2002 3:49:58 PM | 2001-10-01 | 10/4/2001 3:20:45 PM | NPEPPIN | NPEPPIN | MassPerMass | pg/g | Y | (mass conc.) | (wet) | 6Chldibenzodioxin (wet) (mass conc.) | 34465-46-8 | Dioxins & Furans; | Dioxins & Furans; EMS code: H010 | Chemical |  | LAB
H010 | pg/L | 6Chldibenzodioxin | F114 | Isotope dil,HRGC/HRMS | EPA Method 1613: Developed for isomer specific determination of the 2,3,7,8-substituted, tetra-through octa-chlorinated, dibenzo-p-dioxins and dibenzofurans in aqueous, solid and tissue matrices by isotope dilution, high resolution capillary column gas chromatography (HRGC)/high resolution mass spectrometry (HRMS) |  | Y | Y | Y | 18 |  |  | 2/21/2024 1:00:50 PM | 2001-05-23 | 5/24/2001 3:35:10 PM | NPEPPIN | NPEPPIN | Concentration | pg/L | Y | (fl. conc.) |  | 6Chldibenzodioxin (fl. conc.) | 34465-46-8 | Dioxins & Furans; | Dioxins & Furans; EMS code: H010 | Chemical |  | LAB
H010 | pg/L | 6Chldibenzodioxin | X253 | Tol.-Ac/Cl2CH2 ext:GC/MS | C-13 labelled PCDD and PCDF congeners are added to a 1L sample of pulp effluent, which is separated into solid and liquid phases by filtration through a glass fibre filter. The solid phase is aire-dried overnight before undergoing a Soxhlet extraction with toluene. The liquid phase is extracted with dichloromethane in a separatory funnel. The two extracts are combined, concentrated, and then passed through a series of columns which remove the bulk of the organic matrix co-extracted with PCDD/PCDF. One resulting fraction is concentrated to a known small volume for analysis by HRGC/HRMS/DS. | 1 | Y | Y | Y | 146 |  |  | 6/25/2012 2:16:02 PM | 1996-04-12 | 6/3/1998 6:51:36 PM | EMS |  | Concentration | pg/L | Y | (fl. conc.) |  | 6Chldibenzodioxin (fl. conc.) | 34465-46-8 | Dioxins & Furans; | Dioxins & Furans; EMS code: H010 | Chemical |  | SURROGATE_RESULT
H010 | pg/L | 6Chldibenzodioxin | F102 | Env.Canada(1992),DioxinsProc,Pulp Effluents | Environment Canada (1992);Dioxins Procedures for Pulp Effluents |  | Y | Y | Y | 6 | 5/1/2001 7:37:00 AM | 5/1/2001 7:37:00 AM | 11/9/2011 3:41:38 PM | 2000-10-25 |  | NPEPPIN |  | Concentration | pg/L | Y | (fl. conc.) |  | 6Chldibenzodioxin (fl. conc.) | 34465-46-8 | Dioxins & Furans; | Dioxins & Furans; EMS code: H010 | Chemical |  | LAB
H011 | pg/g wet | 6Chl2BDiox1;2;3;4;78 | F119 | Solids,Isotope dil,HRGC/HRMS | EPA Method 1613: for Sediments and Fish tissue developed for isomer specific determination of the 2,3,7,8-substituted, tetra-through octa-chlorinated, dibenzo-p-dioxins and dibenzofurans in aqueous, solid and tissue matrices by isotope dilution, high resolution capillary column gas chromatography (HRGC)/high resolution mass spectrometry (HRMS) |  | Y | Y | Y | 5 | 6/16/2000 6:33:00 AM | 10/4/2000 2:53:00 PM | 6/18/2002 3:49:58 PM | 2001-07-03 | 10/4/2001 3:20:45 PM | NPEPPIN | NPEPPIN | MassPerMass | pg/g | Y | (mass conc.) | (wet) | 6Chl2BDiox1;2;3;4;78 (wet) (mass conc.) | 39227-28-6 | Dioxins & Furans; | Dioxins & Furans; EMS code: H011 | Chemical |  | LAB
H011 | pg/L | 6Chl2BDiox1;2;3;4;78 | F114 | Isotope dil,HRGC/HRMS | EPA Method 1613: Developed for isomer specific determination of the 2,3,7,8-substituted, tetra-through octa-chlorinated, dibenzo-p-dioxins and dibenzofurans in aqueous, solid and tissue matrices by isotope dilution, high resolution capillary column gas chromatography (HRGC)/high resolution mass spectrometry (HRMS) |  | Y | Y | Y | 18 |  |  | 2/21/2024 1:00:50 PM | 2001-05-23 | 5/24/2001 3:35:10 PM | NPEPPIN | NPEPPIN | Concentration | pg/L | Y | (fl. conc.) |  | 6Chl2BDiox1;2;3;4;78 (fl. conc.) | 39227-28-6 | Dioxins & Furans; | Dioxins & Furans; EMS code: H011 | Chemical |  | LAB
H011 | pg/L | 6Chl2BDiox1;2;3;4;78 | F102 | Env.Canada(1992),DioxinsProc,Pulp Effluents | Environment Canada (1992);Dioxins Procedures for Pulp Effluents |  | Y | Y | Y | 6 | 5/1/2001 7:37:00 AM | 5/1/2001 7:37:00 AM | 11/9/2011 3:41:38 PM | 2000-10-25 |  | NPEPPIN |  | Concentration | pg/L | Y | (fl. conc.) |  | 6Chl2BDiox1;2;3;4;78 (fl. conc.) | 39227-28-6 | Dioxins & Furans; | Dioxins & Furans; EMS code: H011 | Chemical |  | LAB
H011 | pg/L | 6Chl2BDiox1;2;3;4;78 | X253 | Tol.-Ac/Cl2CH2 ext:GC/MS | C-13 labelled PCDD and PCDF congeners are added to a 1L sample of pulp effluent, which is separated into solid and liquid phases by filtration through a glass fibre filter. The solid phase is aire-dried overnight before undergoing a Soxhlet extraction with toluene. The liquid phase is extracted with dichloromethane in a separatory funnel. The two extracts are combined, concentrated, and then passed through a series of columns which remove the bulk of the organic matrix co-extracted with PCDD/PCDF. One resulting fraction is concentrated to a known small volume for analysis by HRGC/HRMS/DS. | 1 | Y | Y | Y | 155 |  |  | 6/25/2012 2:16:02 PM | 1996-04-12 | 6/3/1998 6:51:36 PM | EMS |  | Concentration | pg/L | Y | (fl. conc.) |  | 6Chl2BDiox1;2;3;4;78 (fl. conc.) | 39227-28-6 | Dioxins & Furans; | Dioxins & Furans; EMS code: H011 | Chemical |  | SURROGATE_RESULT
H012 | pg/L | 6Chl2BDiox1;2;3;6;78 | X253 | Tol.-Ac/Cl2CH2 ext:GC/MS | C-13 labelled PCDD and PCDF congeners are added to a 1L sample of pulp effluent, which is separated into solid and liquid phases by filtration through a glass fibre filter. The solid phase is aire-dried overnight before undergoing a Soxhlet extraction with toluene. The liquid phase is extracted with dichloromethane in a separatory funnel. The two extracts are combined, concentrated, and then passed through a series of columns which remove the bulk of the organic matrix co-extracted with PCDD/PCDF. One resulting fraction is concentrated to a known small volume for analysis by HRGC/HRMS/DS. | 1 | Y | Y | Y | 155 |  |  | 6/25/2012 2:16:02 PM | 1996-04-12 | 6/3/1998 6:51:36 PM | EMS |  | Concentration | pg/L | Y | (fl. conc.) |  | 6Chl2BDiox1;2;3;6;78 (fl. conc.) | 57653-85-7 | Dioxins & Furans; | Dioxins & Furans; EMS code: H012 | Chemical |  | SURROGATE_RESULT
H012 | pg/g wet | 6Chl2BDiox1;2;3;6;78 | F119 | Solids,Isotope dil,HRGC/HRMS | EPA Method 1613: for Sediments and Fish tissue developed for isomer specific determination of the 2,3,7,8-substituted, tetra-through octa-chlorinated, dibenzo-p-dioxins and dibenzofurans in aqueous, solid and tissue matrices by isotope dilution, high resolution capillary column gas chromatography (HRGC)/high resolution mass spectrometry (HRMS) |  | Y | Y | Y | 5 | 6/16/2000 6:33:00 AM | 10/4/2000 2:53:00 PM | 6/18/2002 3:49:59 PM | 2001-07-03 | 10/4/2001 3:20:45 PM | NPEPPIN | NPEPPIN | MassPerMass | pg/g | Y | (mass conc.) | (wet) | 6Chl2BDiox1;2;3;6;78 (wet) (mass conc.) | 57653-85-7 | Dioxins & Furans; | Dioxins & Furans; EMS code: H012 | Chemical |  | LAB
H012 | pg/L | 6Chl2BDiox1;2;3;6;78 | F102 | Env.Canada(1992),DioxinsProc,Pulp Effluents | Environment Canada (1992);Dioxins Procedures for Pulp Effluents |  | Y | Y | Y | 6 | 5/1/2001 7:37:00 AM | 5/1/2001 7:37:00 AM | 11/9/2011 3:41:38 PM | 2000-10-25 |  | NPEPPIN |  | Concentration | pg/L | Y | (fl. conc.) |  | 6Chl2BDiox1;2;3;6;78 (fl. conc.) | 57653-85-7 | Dioxins & Furans; | Dioxins & Furans; EMS code: H012 | Chemical |  | LAB
H012 | pg/L | 6Chl2BDiox1;2;3;6;78 | F114 | Isotope dil,HRGC/HRMS | EPA Method 1613: Developed for isomer specific determination of the 2,3,7,8-substituted, tetra-through octa-chlorinated, dibenzo-p-dioxins and dibenzofurans in aqueous, solid and tissue matrices by isotope dilution, high resolution capillary column gas chromatography (HRGC)/high resolution mass spectrometry (HRMS) |  | Y | Y | Y | 18 |  |  | 2/21/2024 1:00:50 PM | 2001-05-23 | 5/24/2001 3:35:10 PM | NPEPPIN | NPEPPIN | Concentration | pg/L | Y | (fl. conc.) |  | 6Chl2BDiox1;2;3;6;78 (fl. conc.) | 57653-85-7 | Dioxins & Furans; | Dioxins & Furans; EMS code: H012 | Chemical |  | LAB
H013 | pg/L | 6Chl2BDiox1;2;3;7;89 | X253 | Tol.-Ac/Cl2CH2 ext:GC/MS | C-13 labelled PCDD and PCDF congeners are added to a 1L sample of pulp effluent, which is separated into solid and liquid phases by filtration through a glass fibre filter. The solid phase is aire-dried overnight before undergoing a Soxhlet extraction with toluene. The liquid phase is extracted with dichloromethane in a separatory funnel. The two extracts are combined, concentrated, and then passed through a series of columns which remove the bulk of the organic matrix co-extracted with PCDD/PCDF. One resulting fraction is concentrated to a known small volume for analysis by HRGC/HRMS/DS. | 1e-05 | Y | Y | Y | 155 |  |  | 6/25/2012 2:16:02 PM | 1996-04-12 | 6/3/1998 6:51:36 PM | EMS |  | Concentration | pg/L | Y | (fl. conc.) |  | 6Chl2BDiox1;2;3;7;89 (fl. conc.) | 19408-74-3 | Dioxins & Furans; | Dioxins & Furans; EMS code: H013 | Chemical |  | SURROGATE_RESULT
H013 | pg/L | 6Chl2BDiox1;2;3;7;89 | F102 | Env.Canada(1992),DioxinsProc,Pulp Effluents | Environment Canada (1992);Dioxins Procedures for Pulp Effluents |  | Y | Y | Y | 6 | 5/1/2001 7:37:00 AM | 5/1/2001 7:37:00 AM | 11/9/2011 3:41:38 PM | 2000-10-25 |  | NPEPPIN |  | Concentration | pg/L | Y | (fl. conc.) |  | 6Chl2BDiox1;2;3;7;89 (fl. conc.) | 19408-74-3 | Dioxins & Furans; | Dioxins & Furans; EMS code: H013 | Chemical |  | LAB
H013 | pg/L | 6Chl2BDiox1;2;3;7;89 | F114 | Isotope dil,HRGC/HRMS | EPA Method 1613: Developed for isomer specific determination of the 2,3,7,8-substituted, tetra-through octa-chlorinated, dibenzo-p-dioxins and dibenzofurans in aqueous, solid and tissue matrices by isotope dilution, high resolution capillary column gas chromatography (HRGC)/high resolution mass spectrometry (HRMS) |  | Y | Y | Y | 18 |  |  | 2/21/2024 1:00:50 PM | 2001-05-23 | 5/24/2001 3:35:10 PM | NPEPPIN | NPEPPIN | Concentration | pg/L | Y | (fl. conc.) |  | 6Chl2BDiox1;2;3;7;89 (fl. conc.) | 19408-74-3 | Dioxins & Furans; | Dioxins & Furans; EMS code: H013 | Chemical |  | LAB
H013 | pg/g wet | 6Chl2BDiox1;2;3;7;89 | F119 | Solids,Isotope dil,HRGC/HRMS | EPA Method 1613: for Sediments and Fish tissue developed for isomer specific determination of the 2,3,7,8-substituted, tetra-through octa-chlorinated, dibenzo-p-dioxins and dibenzofurans in aqueous, solid and tissue matrices by isotope dilution, high resolution capillary column gas chromatography (HRGC)/high resolution mass spectrometry (HRMS) |  | Y | Y | Y | 5 | 6/16/2000 6:33:00 AM | 10/4/2000 2:53:00 PM | 6/18/2002 3:49:59 PM | 2001-07-03 | 10/4/2001 3:20:45 PM | NPEPPIN | NPEPPIN | MassPerMass | pg/g | Y | (mass conc.) | (wet) | 6Chl2BDiox1;2;3;7;89 (wet) (mass conc.) | 19408-74-3 | Dioxins & Furans; | Dioxins & Furans; EMS code: H013 | Chemical |  | LAB
H014 | pg/L | 7ChlDibenzodioxin | X253 | Tol.-Ac/Cl2CH2 ext:GC/MS | C-13 labelled PCDD and PCDF congeners are added to a 1L sample of pulp effluent, which is separated into solid and liquid phases by filtration through a glass fibre filter. The solid phase is aire-dried overnight before undergoing a Soxhlet extraction with toluene. The liquid phase is extracted with dichloromethane in a separatory funnel. The two extracts are combined, concentrated, and then passed through a series of columns which remove the bulk of the organic matrix co-extracted with PCDD/PCDF. One resulting fraction is concentrated to a known small volume for analysis by HRGC/HRMS/DS. | 1 | Y | Y | Y | 146 |  |  | 6/25/2012 2:16:02 PM | 1996-04-12 | 6/3/1998 6:51:36 PM | EMS |  | Concentration | pg/L | Y | (fl. conc.) |  | 7ChlDibenzodioxin (fl. conc.) | 37871-00-4 | Dioxins & Furans; | Dioxins & Furans; EMS code: H014 | Chemical |  | SURROGATE_RESULT
H014 | pg/L | 7ChlDibenzodioxin | F114 | Isotope dil,HRGC/HRMS | EPA Method 1613: Developed for isomer specific determination of the 2,3,7,8-substituted, tetra-through octa-chlorinated, dibenzo-p-dioxins and dibenzofurans in aqueous, solid and tissue matrices by isotope dilution, high resolution capillary column gas chromatography (HRGC)/high resolution mass spectrometry (HRMS) |  | Y | Y | Y | 18 |  |  | 2/21/2024 1:00:50 PM | 2001-05-23 | 5/24/2001 3:35:10 PM | NPEPPIN | NPEPPIN | Concentration | pg/L | Y | (fl. conc.) |  | 7ChlDibenzodioxin (fl. conc.) | 37871-00-4 | Dioxins & Furans; | Dioxins & Furans; EMS code: H014 | Chemical |  | LAB
H014 | pg/L | 7ChlDibenzodioxin | F102 | Env.Canada(1992),DioxinsProc,Pulp Effluents | Environment Canada (1992);Dioxins Procedures for Pulp Effluents |  | Y | Y | Y | 6 | 5/1/2001 7:37:00 AM | 5/1/2001 7:37:00 AM | 11/9/2011 3:41:38 PM | 2000-10-25 |  | NPEPPIN |  | Concentration | pg/L | Y | (fl. conc.) |  | 7ChlDibenzodioxin (fl. conc.) | 37871-00-4 | Dioxins & Furans; | Dioxins & Furans; EMS code: H014 | Chemical |  | LAB
H014 | pg/g wet | 7ChlDibenzodioxin | F119 | Solids,Isotope dil,HRGC/HRMS | EPA Method 1613: for Sediments and Fish tissue developed for isomer specific determination of the 2,3,7,8-substituted, tetra-through octa-chlorinated, dibenzo-p-dioxins and dibenzofurans in aqueous, solid and tissue matrices by isotope dilution, high resolution capillary column gas chromatography (HRGC)/high resolution mass spectrometry (HRMS) |  | Y | Y | Y | 5 | 6/16/2000 6:33:00 AM | 10/4/2000 2:53:00 PM | 6/18/2002 3:49:59 PM | 2001-10-01 | 10/4/2001 3:20:45 PM | NPEPPIN | NPEPPIN | MassPerMass | pg/g | Y | (mass conc.) | (wet) | 7ChlDibenzodioxin (wet) (mass conc.) | 37871-00-4 | Dioxins & Furans; | Dioxins & Furans; EMS code: H014 | Chemical |  | LAB
H015 | pg/L | 7Chl2BDiox1;2;3;4678 | X253 | Tol.-Ac/Cl2CH2 ext:GC/MS | C-13 labelled PCDD and PCDF congeners are added to a 1L sample of pulp effluent, which is separated into solid and liquid phases by filtration through a glass fibre filter. The solid phase is aire-dried overnight before undergoing a Soxhlet extraction with toluene. The liquid phase is extracted with dichloromethane in a separatory funnel. The two extracts are combined, concentrated, and then passed through a series of columns which remove the bulk of the organic matrix co-extracted with PCDD/PCDF. One resulting fraction is concentrated to a known small volume for analysis by HRGC/HRMS/DS. | 1 | Y | Y | Y | 155 |  |  | 6/25/2012 2:16:02 PM | 1996-04-12 | 6/3/1998 6:51:36 PM | EMS |  | Concentration | pg/L | Y | (fl. conc.) |  | 7Chl2BDiox1;2;3;4678 (fl. conc.) | 35822-46-9 | Dioxins & Furans; | Dioxins & Furans; EMS code: H015 | Chemical |  | SURROGATE_RESULT
H015 | pg/g wet | 7Chl2BDiox1;2;3;4678 | F119 | Solids,Isotope dil,HRGC/HRMS | EPA Method 1613: for Sediments and Fish tissue developed for isomer specific determination of the 2,3,7,8-substituted, tetra-through octa-chlorinated, dibenzo-p-dioxins and dibenzofurans in aqueous, solid and tissue matrices by isotope dilution, high resolution capillary column gas chromatography (HRGC)/high resolution mass spectrometry (HRMS) |  | Y | Y | Y | 5 | 6/16/2000 6:33:00 AM | 10/4/2000 2:53:00 PM | 6/18/2002 3:50:00 PM | 2001-07-03 | 10/4/2001 3:20:45 PM | NPEPPIN | NPEPPIN | MassPerMass | pg/g | Y | (mass conc.) | (wet) | 7Chl2BDiox1;2;3;4678 (wet) (mass conc.) | 35822-46-9 | Dioxins & Furans; | Dioxins & Furans; EMS code: H015 | Chemical |  | LAB
H015 | pg/L | 7Chl2BDiox1;2;3;4678 | F102 | Env.Canada(1992),DioxinsProc,Pulp Effluents | Environment Canada (1992);Dioxins Procedures for Pulp Effluents |  | Y | Y | Y | 6 | 5/1/2001 7:37:00 AM | 5/1/2001 7:37:00 AM | 11/9/2011 3:41:38 PM | 2000-10-25 |  | NPEPPIN |  | Concentration | pg/L | Y | (fl. conc.) |  | 7Chl2BDiox1;2;3;4678 (fl. conc.) | 35822-46-9 | Dioxins & Furans; | Dioxins & Furans; EMS code: H015 | Chemical |  | LAB
H015 | pg/L | 7Chl2BDiox1;2;3;4678 | F114 | Isotope dil,HRGC/HRMS | EPA Method 1613: Developed for isomer specific determination of the 2,3,7,8-substituted, tetra-through octa-chlorinated, dibenzo-p-dioxins and dibenzofurans in aqueous, solid and tissue matrices by isotope dilution, high resolution capillary column gas chromatography (HRGC)/high resolution mass spectrometry (HRMS) |  | Y | Y | Y | 18 |  |  | 2/21/2024 1:00:50 PM | 2001-05-23 | 5/24/2001 3:35:10 PM | NPEPPIN | NPEPPIN | Concentration | pg/L | Y | (fl. conc.) |  | 7Chl2BDiox1;2;3;4678 (fl. conc.) | 35822-46-9 | Dioxins & Furans; | Dioxins & Furans; EMS code: H015 | Chemical |  | LAB
H016 | pg/g wet | 6Chldibenzofuran | F119 | Solids,Isotope dil,HRGC/HRMS | EPA Method 1613: for Sediments and Fish tissue developed for isomer specific determination of the 2,3,7,8-substituted, tetra-through octa-chlorinated, dibenzo-p-dioxins and dibenzofurans in aqueous, solid and tissue matrices by isotope dilution, high resolution capillary column gas chromatography (HRGC)/high resolution mass spectrometry (HRMS) |  | Y | Y | Y | 5 | 6/16/2000 6:33:00 AM | 10/4/2000 2:53:00 PM | 6/18/2002 3:50:00 PM | 2001-10-01 | 10/4/2001 3:20:45 PM | NPEPPIN | NPEPPIN | MassPerMass | pg/g | Y | (mass conc.) | (wet) | 6Chldibenzofuran (wet) (mass conc.) | 55684-94-1 | Dioxins & Furans; | Dioxins & Furans; EMS code: H016 | Chemical |  | LAB
H016 | pg/L | 6Chldibenzofuran | F114 | Isotope dil,HRGC/HRMS | EPA Method 1613: Developed for isomer specific determination of the 2,3,7,8-substituted, tetra-through octa-chlorinated, dibenzo-p-dioxins and dibenzofurans in aqueous, solid and tissue matrices by isotope dilution, high resolution capillary column gas chromatography (HRGC)/high resolution mass spectrometry (HRMS) |  | Y | Y | Y | 18 |  |  | 2/21/2024 1:00:50 PM | 2001-05-23 | 5/24/2001 3:35:10 PM | NPEPPIN | NPEPPIN | Concentration | pg/L | Y | (fl. conc.) |  | 6Chldibenzofuran (fl. conc.) | 55684-94-1 | Dioxins & Furans; | Dioxins & Furans; EMS code: H016 | Chemical |  | LAB
H016 | pg/L | 6Chldibenzofuran | F102 | Env.Canada(1992),DioxinsProc,Pulp Effluents | Environment Canada (1992);Dioxins Procedures for Pulp Effluents |  | Y | Y | Y | 6 | 5/1/2001 7:37:00 AM | 5/1/2001 7:37:00 AM | 11/9/2011 3:41:38 PM | 2000-10-25 |  | NPEPPIN |  | Concentration | pg/L | Y | (fl. conc.) |  | 6Chldibenzofuran (fl. conc.) | 55684-94-1 | Dioxins & Furans; | Dioxins & Furans; EMS code: H016 | Chemical |  | LAB
H016 | pg/L | 6Chldibenzofuran | X253 | Tol.-Ac/Cl2CH2 ext:GC/MS | C-13 labelled PCDD and PCDF congeners are added to a 1L sample of pulp effluent, which is separated into solid and liquid phases by filtration through a glass fibre filter. The solid phase is aire-dried overnight before undergoing a Soxhlet extraction with toluene. The liquid phase is extracted with dichloromethane in a separatory funnel. The two extracts are combined, concentrated, and then passed through a series of columns which remove the bulk of the organic matrix co-extracted with PCDD/PCDF. One resulting fraction is concentrated to a known small volume for analysis by HRGC/HRMS/DS. | 1 | Y | Y | Y | 146 |  |  | 6/25/2012 2:16:02 PM | 1996-04-12 | 6/3/1998 6:51:36 PM | EMS |  | Concentration | pg/L | Y | (fl. conc.) |  | 6Chldibenzofuran (fl. conc.) | 55684-94-1 | Dioxins & Furans; | Dioxins & Furans; EMS code: H016 | Chemical |  | SURROGATE_RESULT
H017 | pg/g wet | 6Chl2BFuran1;2;3;478 | F119 | Solids,Isotope dil,HRGC/HRMS | EPA Method 1613: for Sediments and Fish tissue developed for isomer specific determination of the 2,3,7,8-substituted, tetra-through octa-chlorinated, dibenzo-p-dioxins and dibenzofurans in aqueous, solid and tissue matrices by isotope dilution, high resolution capillary column gas chromatography (HRGC)/high resolution mass spectrometry (HRMS) |  | Y | Y | Y | 5 | 6/16/2000 6:33:00 AM | 10/4/2000 2:53:00 PM | 6/18/2002 3:50:00 PM | 2001-07-03 | 10/4/2001 3:20:45 PM | NPEPPIN | NPEPPIN | MassPerMass | pg/g | Y | (mass conc.) | (wet) | 6Chl2BFuran1;2;3;478 (wet) (mass conc.) | 70648-26-9 | Dioxins & Furans; | Dioxins & Furans; EMS code: H017 | Chemical |  | LAB
H017 | pg/L | 6Chl2BFuran1;2;3;478 | F102 | Env.Canada(1992),DioxinsProc,Pulp Effluents | Environment Canada (1992);Dioxins Procedures for Pulp Effluents |  | Y | Y | Y | 6 | 5/1/2001 7:37:00 AM | 5/1/2001 7:37:00 AM | 11/9/2011 3:41:38 PM | 2000-10-25 |  | NPEPPIN |  | Concentration | pg/L | Y | (fl. conc.) |  | 6Chl2BFuran1;2;3;478 (fl. conc.) | 70648-26-9 | Dioxins & Furans; | Dioxins & Furans; EMS code: H017 | Chemical |  | LAB
H017 | pg/L | 6Chl2BFuran1;2;3;478 | F114 | Isotope dil,HRGC/HRMS | EPA Method 1613: Developed for isomer specific determination of the 2,3,7,8-substituted, tetra-through octa-chlorinated, dibenzo-p-dioxins and dibenzofurans in aqueous, solid and tissue matrices by isotope dilution, high resolution capillary column gas chromatography (HRGC)/high resolution mass spectrometry (HRMS) |  | Y | Y | Y | 18 |  |  | 2/21/2024 1:00:50 PM | 2001-05-23 | 5/24/2001 3:35:10 PM | NPEPPIN | NPEPPIN | Concentration | pg/L | Y | (fl. conc.) |  | 6Chl2BFuran1;2;3;478 (fl. conc.) | 70648-26-9 | Dioxins & Furans; | Dioxins & Furans; EMS code: H017 | Chemical |  | LAB
H017 | pg/L | 6Chl2BFuran1;2;3;478 | X253 | Tol.-Ac/Cl2CH2 ext:GC/MS | C-13 labelled PCDD and PCDF congeners are added to a 1L sample of pulp effluent, which is separated into solid and liquid phases by filtration through a glass fibre filter. The solid phase is aire-dried overnight before undergoing a Soxhlet extraction with toluene. The liquid phase is extracted with dichloromethane in a separatory funnel. The two extracts are combined, concentrated, and then passed through a series of columns which remove the bulk of the organic matrix co-extracted with PCDD/PCDF. One resulting fraction is concentrated to a known small volume for analysis by HRGC/HRMS/DS. | 1 | Y | Y | Y | 154 |  |  | 6/25/2012 2:16:02 PM | 1996-04-12 | 6/3/1998 6:51:45 PM | EMS |  | Concentration | pg/L | Y | (fl. conc.) |  | 6Chl2BFuran1;2;3;478 (fl. conc.) | 70648-26-9 | Dioxins & Furans; | Dioxins & Furans; EMS code: H017 | Chemical |  | SURROGATE_RESULT
H018 | pg/L | 6Chl2BFuran1;2;3;678 | F114 | Isotope dil,HRGC/HRMS | EPA Method 1613: Developed for isomer specific determination of the 2,3,7,8-substituted, tetra-through octa-chlorinated, dibenzo-p-dioxins and dibenzofurans in aqueous, solid and tissue matrices by isotope dilution, high resolution capillary column gas chromatography (HRGC)/high resolution mass spectrometry (HRMS) |  | Y | Y | Y | 18 |  |  | 2/21/2024 1:00:50 PM | 2001-05-23 | 5/24/2001 3:35:10 PM | NPEPPIN | NPEPPIN | Concentration | pg/L | Y | (fl. conc.) |  | 6Chl2BFuran1;2;3;678 (fl. conc.) | 57117-44-9 | Dioxins & Furans; | Dioxins & Furans; EMS code: H018 | Chemical |  | LAB
H018 | pg/L | 6Chl2BFuran1;2;3;678 | F102 | Env.Canada(1992),DioxinsProc,Pulp Effluents | Environment Canada (1992);Dioxins Procedures for Pulp Effluents |  | Y | Y | Y | 6 | 5/1/2001 7:37:00 AM | 5/1/2001 7:37:00 AM | 11/9/2011 3:41:38 PM | 2000-10-25 |  | NPEPPIN |  | Concentration | pg/L | Y | (fl. conc.) |  | 6Chl2BFuran1;2;3;678 (fl. conc.) | 57117-44-9 | Dioxins & Furans; | Dioxins & Furans; EMS code: H018 | Chemical |  | LAB
H018 | pg/g wet | 6Chl2BFuran1;2;3;678 | F119 | Solids,Isotope dil,HRGC/HRMS | EPA Method 1613: for Sediments and Fish tissue developed for isomer specific determination of the 2,3,7,8-substituted, tetra-through octa-chlorinated, dibenzo-p-dioxins and dibenzofurans in aqueous, solid and tissue matrices by isotope dilution, high resolution capillary column gas chromatography (HRGC)/high resolution mass spectrometry (HRMS) |  | Y | Y | Y | 5 | 6/16/2000 6:33:00 AM | 10/4/2000 2:53:00 PM | 6/18/2002 3:50:01 PM | 2001-07-03 | 10/4/2001 3:20:45 PM | NPEPPIN | NPEPPIN | MassPerMass | pg/g | Y | (mass conc.) | (wet) | 6Chl2BFuran1;2;3;678 (wet) (mass conc.) | 57117-44-9 | Dioxins & Furans; | Dioxins & Furans; EMS code: H018 | Chemical |  | LAB
H018 | pg/L | 6Chl2BFuran1;2;3;678 | X253 | Tol.-Ac/Cl2CH2 ext:GC/MS | C-13 labelled PCDD and PCDF congeners are added to a 1L sample of pulp effluent, which is separated into solid and liquid phases by filtration through a glass fibre filter. The solid phase is aire-dried overnight before undergoing a Soxhlet extraction with toluene. The liquid phase is extracted with dichloromethane in a separatory funnel. The two extracts are combined, concentrated, and then passed through a series of columns which remove the bulk of the organic matrix co-extracted with PCDD/PCDF. One resulting fraction is concentrated to a known small volume for analysis by HRGC/HRMS/DS. | 1 | Y | Y | Y | 153 |  |  | 6/25/2012 2:16:02 PM | 1996-04-12 | 6/3/1998 6:51:36 PM | EMS |  | Concentration | pg/L | Y | (fl. conc.) |  | 6Chl2BFuran1;2;3;678 (fl. conc.) | 57117-44-9 | Dioxins & Furans; | Dioxins & Furans; EMS code: H018 | Chemical |  | SURROGATE_RESULT
H019 | pg/L | 6Chl2BFuran2;3;4;678 | F114 | Isotope dil,HRGC/HRMS | EPA Method 1613: Developed for isomer specific determination of the 2,3,7,8-substituted, tetra-through octa-chlorinated, dibenzo-p-dioxins and dibenzofurans in aqueous, solid and tissue matrices by isotope dilution, high resolution capillary column gas chromatography (HRGC)/high resolution mass spectrometry (HRMS) |  | Y | Y | Y | 18 |  |  | 2/21/2024 1:00:50 PM | 2001-05-23 | 5/24/2001 3:35:10 PM | NPEPPIN | NPEPPIN | Concentration | pg/L | Y | (fl. conc.) |  | 6Chl2BFuran2;3;4;678 (fl. conc.) | 60851-34-5 | Dioxins & Furans; | Dioxins & Furans; EMS code: H019 | Chemical |  | LAB
H019 | pg/L | 6Chl2BFuran2;3;4;678 | X253 | Tol.-Ac/Cl2CH2 ext:GC/MS | C-13 labelled PCDD and PCDF congeners are added to a 1L sample of pulp effluent, which is separated into solid and liquid phases by filtration through a glass fibre filter. The solid phase is aire-dried overnight before undergoing a Soxhlet extraction with toluene. The liquid phase is extracted with dichloromethane in a separatory funnel. The two extracts are combined, concentrated, and then passed through a series of columns which remove the bulk of the organic matrix co-extracted with PCDD/PCDF. One resulting fraction is concentrated to a known small volume for analysis by HRGC/HRMS/DS. | 1 | Y | Y | Y | 154 |  |  | 6/25/2012 2:16:03 PM | 1996-04-12 | 6/3/1998 6:51:36 PM | EMS |  | Concentration | pg/L | Y | (fl. conc.) |  | 6Chl2BFuran2;3;4;678 (fl. conc.) | 60851-34-5 | Dioxins & Furans; | Dioxins & Furans; EMS code: H019 | Chemical |  | SURROGATE_RESULT
H019 | pg/g wet | 6Chl2BFuran2;3;4;678 | F119 | Solids,Isotope dil,HRGC/HRMS | EPA Method 1613: for Sediments and Fish tissue developed for isomer specific determination of the 2,3,7,8-substituted, tetra-through octa-chlorinated, dibenzo-p-dioxins and dibenzofurans in aqueous, solid and tissue matrices by isotope dilution, high resolution capillary column gas chromatography (HRGC)/high resolution mass spectrometry (HRMS) |  | Y | Y | Y | 5 | 6/16/2000 6:33:00 AM | 10/4/2000 2:53:00 PM | 6/18/2002 3:50:01 PM | 2001-07-03 | 10/4/2001 3:20:45 PM | NPEPPIN | NPEPPIN | MassPerMass | pg/g | Y | (mass conc.) | (wet) | 6Chl2BFuran2;3;4;678 (wet) (mass conc.) | 60851-34-5 | Dioxins & Furans; | Dioxins & Furans; EMS code: H019 | Chemical |  | LAB
H019 | pg/L | 6Chl2BFuran2;3;4;678 | F102 | Env.Canada(1992),DioxinsProc,Pulp Effluents | Environment Canada (1992);Dioxins Procedures for Pulp Effluents |  | Y | Y | Y | 6 | 5/1/2001 7:37:00 AM | 5/1/2001 7:37:00 AM | 11/9/2011 3:41:39 PM | 2000-10-25 |  | NPEPPIN |  | Concentration | pg/L | Y | (fl. conc.) |  | 6Chl2BFuran2;3;4;678 (fl. conc.) | 60851-34-5 | Dioxins & Furans; | Dioxins & Furans; EMS code: H019 | Chemical |  | LAB
H020 | pg/L | 6Chl2BFuran1;2;3;789 | F114 | Isotope dil,HRGC/HRMS | EPA Method 1613: Developed for isomer specific determination of the 2,3,7,8-substituted, tetra-through octa-chlorinated, dibenzo-p-dioxins and dibenzofurans in aqueous, solid and tissue matrices by isotope dilution, high resolution capillary column gas chromatography (HRGC)/high resolution mass spectrometry (HRMS) |  | Y | Y | Y | 18 |  |  | 2/21/2024 1:00:50 PM | 2001-05-23 | 5/24/2001 3:35:10 PM | NPEPPIN | NPEPPIN | Concentration | pg/L | Y | (fl. conc.) |  | 6Chl2BFuran1;2;3;789 (fl. conc.) | 72918-21-9 | Dioxins & Furans; | Dioxins & Furans; EMS code: H020 | Chemical |  | LAB
H020 | pg/g wet | 6Chl2BFuran1;2;3;789 | F119 | Solids,Isotope dil,HRGC/HRMS | EPA Method 1613: for Sediments and Fish tissue developed for isomer specific determination of the 2,3,7,8-substituted, tetra-through octa-chlorinated, dibenzo-p-dioxins and dibenzofurans in aqueous, solid and tissue matrices by isotope dilution, high resolution capillary column gas chromatography (HRGC)/high resolution mass spectrometry (HRMS) |  | Y | Y | Y | 5 | 6/16/2000 6:33:00 AM | 10/4/2000 2:53:00 PM | 6/18/2002 3:50:01 PM | 2001-07-03 | 10/4/2001 3:20:45 PM | NPEPPIN | NPEPPIN | MassPerMass | pg/g | Y | (mass conc.) | (wet) | 6Chl2BFuran1;2;3;789 (wet) (mass conc.) | 72918-21-9 | Dioxins & Furans; | Dioxins & Furans; EMS code: H020 | Chemical |  | LAB
H020 | pg/L | 6Chl2BFuran1;2;3;789 | X253 | Tol.-Ac/Cl2CH2 ext:GC/MS | C-13 labelled PCDD and PCDF congeners are added to a 1L sample of pulp effluent, which is separated into solid and liquid phases by filtration through a glass fibre filter. The solid phase is aire-dried overnight before undergoing a Soxhlet extraction with toluene. The liquid phase is extracted with dichloromethane in a separatory funnel. The two extracts are combined, concentrated, and then passed through a series of columns which remove the bulk of the organic matrix co-extracted with PCDD/PCDF. One resulting fraction is concentrated to a known small volume for analysis by HRGC/HRMS/DS. | 1 | Y | Y | Y | 155 |  |  | 6/25/2012 2:16:03 PM | 1996-04-12 | 6/3/1998 6:51:36 PM | EMS |  | Concentration | pg/L | Y | (fl. conc.) |  | 6Chl2BFuran1;2;3;789 (fl. conc.) | 72918-21-9 | Dioxins & Furans; | Dioxins & Furans; EMS code: H020 | Chemical |  | SURROGATE_RESULT
H020 | pg/L | 6Chl2BFuran1;2;3;789 | F102 | Env.Canada(1992),DioxinsProc,Pulp Effluents | Environment Canada (1992);Dioxins Procedures for Pulp Effluents |  | Y | Y | Y | 6 | 5/1/2001 7:37:00 AM | 5/1/2001 7:37:00 AM | 11/9/2011 3:41:39 PM | 2000-10-25 |  | NPEPPIN |  | Concentration | pg/L | Y | (fl. conc.) |  | 6Chl2BFuran1;2;3;789 (fl. conc.) | 72918-21-9 | Dioxins & Furans; | Dioxins & Furans; EMS code: H020 | Chemical |  | LAB
H021 | pg/L | 7Chldibenzofuran | X253 | Tol.-Ac/Cl2CH2 ext:GC/MS | C-13 labelled PCDD and PCDF congeners are added to a 1L sample of pulp effluent, which is separated into solid and liquid phases by filtration through a glass fibre filter. The solid phase is aire-dried overnight before undergoing a Soxhlet extraction with toluene. The liquid phase is extracted with dichloromethane in a separatory funnel. The two extracts are combined, concentrated, and then passed through a series of columns which remove the bulk of the organic matrix co-extracted with PCDD/PCDF. One resulting fraction is concentrated to a known small volume for analysis by HRGC/HRMS/DS. | 1 | Y | Y | Y | 146 |  |  | 6/25/2012 2:16:03 PM | 1996-04-12 | 6/3/1998 6:51:45 PM | EMS |  | Concentration | pg/L | Y | (fl. conc.) |  | 7Chldibenzofuran (fl. conc.) | 38998-75-3 | Dioxins & Furans; | Dioxins & Furans; EMS code: H021 | Chemical |  | SURROGATE_RESULT
H021 | pg/L | 7Chldibenzofuran | F114 | Isotope dil,HRGC/HRMS | EPA Method 1613: Developed for isomer specific determination of the 2,3,7,8-substituted, tetra-through octa-chlorinated, dibenzo-p-dioxins and dibenzofurans in aqueous, solid and tissue matrices by isotope dilution, high resolution capillary column gas chromatography (HRGC)/high resolution mass spectrometry (HRMS) |  | Y | Y | Y | 18 |  |  | 2/21/2024 1:00:50 PM | 2001-05-23 | 5/24/2001 3:35:10 PM | NPEPPIN | NPEPPIN | Concentration | pg/L | Y | (fl. conc.) |  | 7Chldibenzofuran (fl. conc.) | 38998-75-3 | Dioxins & Furans; | Dioxins & Furans; EMS code: H021 | Chemical |  | LAB
H021 | pg/L | 7Chldibenzofuran | F102 | Env.Canada(1992),DioxinsProc,Pulp Effluents | Environment Canada (1992);Dioxins Procedures for Pulp Effluents |  | Y | Y | Y | 6 | 5/1/2001 7:37:00 AM | 5/1/2001 7:37:00 AM | 11/9/2011 3:41:39 PM | 2000-10-25 |  | NPEPPIN |  | Concentration | pg/L | Y | (fl. conc.) |  | 7Chldibenzofuran (fl. conc.) | 38998-75-3 | Dioxins & Furans; | Dioxins & Furans; EMS code: H021 | Chemical |  | LAB
H021 | pg/g wet | 7Chldibenzofuran | F119 | Solids,Isotope dil,HRGC/HRMS | EPA Method 1613: for Sediments and Fish tissue developed for isomer specific determination of the 2,3,7,8-substituted, tetra-through octa-chlorinated, dibenzo-p-dioxins and dibenzofurans in aqueous, solid and tissue matrices by isotope dilution, high resolution capillary column gas chromatography (HRGC)/high resolution mass spectrometry (HRMS) |  | Y | Y | Y | 5 | 6/16/2000 6:33:00 AM | 10/4/2000 2:53:00 PM | 6/18/2002 3:50:02 PM | 2001-10-01 | 10/4/2001 3:20:45 PM | NPEPPIN | NPEPPIN | MassPerMass | pg/g | Y | (mass conc.) | (wet) | 7Chldibenzofuran (wet) (mass conc.) | 38998-75-3 | Dioxins & Furans; | Dioxins & Furans; EMS code: H021 | Chemical |  | LAB
H022 | pg/L | 7Chl2BFuran1;2;34678 | F102 | Env.Canada(1992),DioxinsProc,Pulp Effluents | Environment Canada (1992);Dioxins Procedures for Pulp Effluents |  | Y | Y | Y | 6 | 5/1/2001 7:37:00 AM | 5/1/2001 7:37:00 AM | 11/9/2011 3:41:39 PM | 2000-10-25 |  | NPEPPIN |  | Concentration | pg/L | Y | (fl. conc.) |  | 7Chl2BFuran1;2;34678 (fl. conc.) | 67562-39-4 | Dioxins & Furans; | Dioxins & Furans; EMS code: H022 | Chemical |  | LAB
H022 | pg/g wet | 7Chl2BFuran1;2;34678 | F119 | Solids,Isotope dil,HRGC/HRMS | EPA Method 1613: for Sediments and Fish tissue developed for isomer specific determination of the 2,3,7,8-substituted, tetra-through octa-chlorinated, dibenzo-p-dioxins and dibenzofurans in aqueous, solid and tissue matrices by isotope dilution, high resolution capillary column gas chromatography (HRGC)/high resolution mass spectrometry (HRMS) |  | Y | Y | Y | 5 | 6/16/2000 6:33:00 AM | 10/4/2000 2:53:00 PM | 6/18/2002 3:50:02 PM | 2001-07-03 | 10/4/2001 3:20:45 PM | NPEPPIN | NPEPPIN | MassPerMass | pg/g | Y | (mass conc.) | (wet) | 7Chl2BFuran1;2;34678 (wet) (mass conc.) | 67562-39-4 | Dioxins & Furans; | Dioxins & Furans; EMS code: H022 | Chemical |  | LAB
H022 | pg/L | 7Chl2BFuran1;2;34678 | X253 | Tol.-Ac/Cl2CH2 ext:GC/MS | C-13 labelled PCDD and PCDF congeners are added to a 1L sample of pulp effluent, which is separated into solid and liquid phases by filtration through a glass fibre filter. The solid phase is aire-dried overnight before undergoing a Soxhlet extraction with toluene. The liquid phase is extracted with dichloromethane in a separatory funnel. The two extracts are combined, concentrated, and then passed through a series of columns which remove the bulk of the organic matrix co-extracted with PCDD/PCDF. One resulting fraction is concentrated to a known small volume for analysis by HRGC/HRMS/DS. | 1 | Y | Y | Y | 155 |  |  | 6/25/2012 2:16:03 PM | 1996-04-12 | 6/3/1998 6:51:36 PM | EMS |  | Concentration | pg/L | Y | (fl. conc.) |  | 7Chl2BFuran1;2;34678 (fl. conc.) | 67562-39-4 | Dioxins & Furans; | Dioxins & Furans; EMS code: H022 | Chemical |  | SURROGATE_RESULT
H022 | pg/L | 7Chl2BFuran1;2;34678 | F114 | Isotope dil,HRGC/HRMS | EPA Method 1613: Developed for isomer specific determination of the 2,3,7,8-substituted, tetra-through octa-chlorinated, dibenzo-p-dioxins and dibenzofurans in aqueous, solid and tissue matrices by isotope dilution, high resolution capillary column gas chromatography (HRGC)/high resolution mass spectrometry (HRMS) |  | Y | Y | Y | 18 |  |  | 2/21/2024 1:00:50 PM | 2001-05-23 | 5/24/2001 3:35:10 PM | NPEPPIN | NPEPPIN | Concentration | pg/L | Y | (fl. conc.) |  | 7Chl2BFuran1;2;34678 (fl. conc.) | 67562-39-4 | Dioxins & Furans; | Dioxins & Furans; EMS code: H022 | Chemical |  | LAB
H023 | pg/L | 7Chl2BFuran1;2;34789 | X253 | Tol.-Ac/Cl2CH2 ext:GC/MS | C-13 labelled PCDD and PCDF congeners are added to a 1L sample of pulp effluent, which is separated into solid and liquid phases by filtration through a glass fibre filter. The solid phase is aire-dried overnight before undergoing a Soxhlet extraction with toluene. The liquid phase is extracted with dichloromethane in a separatory funnel. The two extracts are combined, concentrated, and then passed through a series of columns which remove the bulk of the organic matrix co-extracted with PCDD/PCDF. One resulting fraction is concentrated to a known small volume for analysis by HRGC/HRMS/DS. | 1 | Y | Y | Y | 155 |  |  | 6/25/2012 2:16:03 PM | 1996-04-12 | 6/3/1998 6:51:36 PM | EMS |  | Concentration | pg/L | Y | (fl. conc.) |  | 7Chl2BFuran1;2;34789 (fl. conc.) | 55673-89-7 | Dioxins & Furans; | Dioxins & Furans; EMS code: H023 | Chemical |  | SURROGATE_RESULT
H023 | pg/L | 7Chl2BFuran1;2;34789 | F114 | Isotope dil,HRGC/HRMS | EPA Method 1613: Developed for isomer specific determination of the 2,3,7,8-substituted, tetra-through octa-chlorinated, dibenzo-p-dioxins and dibenzofurans in aqueous, solid and tissue matrices by isotope dilution, high resolution capillary column gas chromatography (HRGC)/high resolution mass spectrometry (HRMS) |  | Y | Y | Y | 18 |  |  | 2/21/2024 1:00:50 PM | 2001-05-23 | 5/24/2001 3:35:10 PM | NPEPPIN | NPEPPIN | Concentration | pg/L | Y | (fl. conc.) |  | 7Chl2BFuran1;2;34789 (fl. conc.) | 55673-89-7 | Dioxins & Furans; | Dioxins & Furans; EMS code: H023 | Chemical |  | LAB
H023 | pg/g wet | 7Chl2BFuran1;2;34789 | F119 | Solids,Isotope dil,HRGC/HRMS | EPA Method 1613: for Sediments and Fish tissue developed for isomer specific determination of the 2,3,7,8-substituted, tetra-through octa-chlorinated, dibenzo-p-dioxins and dibenzofurans in aqueous, solid and tissue matrices by isotope dilution, high resolution capillary column gas chromatography (HRGC)/high resolution mass spectrometry (HRMS) |  | Y | Y | Y | 5 | 6/16/2000 6:33:00 AM | 10/4/2000 2:53:00 PM | 6/18/2002 3:50:02 PM | 2001-10-01 | 10/4/2001 3:20:45 PM | NPEPPIN | NPEPPIN | MassPerMass | pg/g | Y | (mass conc.) | (wet) | 7Chl2BFuran1;2;34789 (wet) (mass conc.) | 55673-89-7 | Dioxins & Furans; | Dioxins & Furans; EMS code: H023 | Chemical |  | LAB
H023 | pg/L | 7Chl2BFuran1;2;34789 | F102 | Env.Canada(1992),DioxinsProc,Pulp Effluents | Environment Canada (1992);Dioxins Procedures for Pulp Effluents |  | Y | Y | Y | 6 | 5/1/2001 7:37:00 AM | 5/1/2001 7:37:00 AM | 11/9/2011 3:41:39 PM | 2000-10-25 |  | NPEPPIN |  | Concentration | pg/L | Y | (fl. conc.) |  | 7Chl2BFuran1;2;34789 (fl. conc.) | 55673-89-7 | Dioxins & Furans; | Dioxins & Furans; EMS code: H023 | Chemical |  | LAB
H109 | mg/L | HydroC Extrct/C10-30 | F065 | EPH:DCM,liq-liq extr'n,GC/FID,(PAH not sub.) | Samples are extracted with dichloromethane (DCM) in a separatory funnel. Extracts are dried, concentrated and analyzed by capillary gas chromatography with flame ionization detector. Results are not corrected for specified PAH concentrations. LEPH is defined as the sum of components that elute on gc between nC10 and nC19; HEPH is defined as the sum of components that elute on gc between nC19 and nC32. Performance Base Method published circa Dec. 31, 2000, but distributed 1999. | 1 | Y | Y | Y | 467 | 4/21/2007 | 8/21/2024 | 8/22/2024 5:27:05 PM | 2007-11-05 |  | RDGIBSON |  | Concentration | mg/L | Y | (fl. conc.) |  | HydroC Extrct/C10-30 (fl. conc.) |  | Petroleum Hydrocarbons; | Petroleum Hydrocarbons; EMS code: H109 | Chemical |  | LAB
H109 | mg/L | HydroC Extrct/C10-30 | TEHW | Total Ext-Hydrocarbons in Water by GCFID | Total Ext-Hydrocarbons in Water by GCFID, BC Lab Manual | 0.25 | Y | Y | Y | 534 | 5/21/2015 | 3/19/2020 | 2/18/2021 5:29:07 PM | 2015-05-26 | 5/26/2015 11:25:31 AM | RDGIBSON | RDGIBSON | Concentration | mg/L | Y | (fl. conc.) |  | HydroC Extrct/C10-30 (fl. conc.) |  | Petroleum Hydrocarbons; | Petroleum Hydrocarbons; EMS code: H109 | Chemical |  | LAB
H109 | mg/L | HydroC Extrct/C10-30 | P109 | CH2Cl2 Extr:Ioniz'n GLC | CH2Cl2 Extr:Ioniz'n GLC | 0.01 | Y | Y | Y | 4 |  |  | 10/24/1996 3:05:29 PM | 1996-04-12 | 6/3/1998 6:51:45 PM | EMS |  | Concentration | mg/L | Y | (fl. conc.) |  | HydroC Extrct/C10-30 (fl. conc.) |  | Petroleum Hydrocarbons; | Petroleum Hydrocarbons; EMS code: H109 | Chemical |  | LAB
H109 | ug/g | HydroC Extrct/C10-30 | F086 | EPH:Hex/Actn,soxhlet extr'n,GC/FID(PAHs not sub) | Samples are dried and soxhlet extracted with 1:1 Hexane:Acetone for 4 hours. The extract is concentrated and analyzed by capillary gas chromatography with flame ionization detector. Results are not corrected for specified PAH's. LEPH is defined as the sum of components that elute on gc between nC10 and nC19; HEPH is defined as the sum of components that elute on gc between nC19 and nC32. Performance Base Method published circa Dec. 31, 2000. | 250 | Y | Y | Y | 2 | 9/3/2021 | 9/3/2021 | 10/19/2022 1:02:09 PM | 2007-11-05 |  | RDGIBSON |  | MassPerMass | mg/kg | Y | (mass conc.) |  | HydroC Extrct/C10-30 (mass conc.) |  | Petroleum Hydrocarbons; | Petroleum Hydrocarbons; EMS code: H109 | Chemical |  | LAB
H109 | mg/L | HydroC Extrct/C10-30 | GCFI | Solvent Extn, GC-FID | Hexane extraction using tumbler; Analysis by GC/FID | 0.08 | Y | Y | Y | 1263 | 2/9/2016 | 8/8/2024 6:16:00 PM | 8/19/2024 2:37:07 PM | 2010-11-16 |  | RDGIBSON |  | Concentration | mg/L | Y | (fl. conc.) |  | HydroC Extrct/C10-30 (fl. conc.) |  | Petroleum Hydrocarbons; | Petroleum Hydrocarbons; EMS code: H109 | Chemical |  | LAB
H110 | ug/g | HydroC Petrol/C10-30 | P111 | CH2Cl2+Meth Extr:Sil:GLC | CH2Cl2+Meth Extr:Sil:GLC | 5 | Y | Y | Y | 2 | 4/14/1999 11:54:00 AM | 4/14/1999 11:55:00 AM | 4/16/1999 6:12:51 PM | 1996-04-12 | 6/3/1998 6:51:36 PM | EMS |  | MassPerMass | mg/kg | Y | (mass conc.) |  | HydroC Petrol/C10-30 (mass conc.) |  | Petroleum Hydrocarbons; | Petroleum Hydrocarbons; EMS code: H110 | Chemical |  | LAB
H110 | mg/L | HydroC Petrol/C10-30 | P110 | CH2Cl2 Extr:Silica:GLC | CH2Cl2 Extr:Silica:GLC | 0.01 | Y | Y | Y | 2 | 5/8/2003 10:20:00 AM | 5/8/2003 10:20:00 AM | 8/14/2003 3:50:24 PM | 1996-04-12 | 6/3/1998 6:51:36 PM | EMS |  | Concentration | mg/L | Y | (fl. conc.) |  | HydroC Petrol/C10-30 (fl. conc.) |  | Petroleum Hydrocarbons; | Petroleum Hydrocarbons; EMS code: H110 | Chemical |  | LAB
H2O2 | mg/L | Hydrogen Peroxide | X466 | H2SO4, KMnO4, Col'm | H2SO4; KMnO4; Col'm | 0.1 | Y | Y | Y | 3826 |  |  | 7/29/2024 2:02:49 PM | 1996-10-13 | 6/3/1998 6:51:46 PM | EMS |  | Concentration | mg/L | Y | (fl. conc.) |  | Hydrogen Peroxide (fl. conc.) | 7722-84-1 | Inorganics; | Inorganics; EMS code: H2O2 | Chemical |  | LAB
H2S- | mg/L | Hydrogen Sulfide | X257 | Colorimetric | Colorimetric. General use for several parameters. E.g.Total Sulfide is determined by spectrophotometer using the methlyene blue colourimetric method. The H2S calculation converts Total Sulphide as (S2-) and reports it as Total Sulphide as (H2S). | 1 | Y | Y | Y | 259 | 4/9/2020 | 8/1/2024 | 8/12/2024 4:48:10 PM | 2020-06-09 |  | RDGIBSON |  | Concentration | mg/L | Y | (fl. conc.) |  | Hydrogen Sulfide (fl. conc.) | 7783-06-4 | Volatile Organic Compounds (VOCs); | Volatile Organic Compounds (VOCs); EMS code: H2S- | Chemical |  | LAB
H2S- | ug/m3 | Hydrogen Sulfide | 5099 | Iodine Titration-Stack | Iodine Titration-Stack | 0.001 | Y | Y | Y | 169 | 8/24/2019 | 6/30/2024 | 7/22/2024 3:17:33 PM | 1996-04-12 | 6/3/1998 6:51:36 PM | EMS |  | AirConcentration | µg/m³ | Y | (air conc.) |  | Hydrogen Sulfide (air conc.) | 7783-06-4 | Volatile Organic Compounds (VOCs); | Volatile Organic Compounds (VOCs); EMS code: H2S- | Chemical |  | LAB
H2S- | ug/g | Hydrogen Sulfide | CAL2 | Calculation (ug/g) | Calculation (ug/g) | 0 | Y | Y | Y | 62 |  |  | 10/24/1996 7:10:48 AM | 1996-04-12 | 6/3/1998 6:51:42 PM | EMS |  | MassPerMass | mg/kg | Y | (mass conc.) |  | Hydrogen Sulfide (mass conc.) | 7783-06-4 | Volatile Organic Compounds (VOCs); | Volatile Organic Compounds (VOCs); EMS code: H2S- | Chemical |  | LAB
H2S- | mg/L | Hydrogen Sulfide | 1075 | Colorimetric:Methyl.Blue | Colorimetric:Methyl.Blue | 0.011 | Y | Y | Y | 1330 | 6/15/2021 | 8/14/2024 | 8/20/2024 1:06:09 PM | 2021-07-07 |  | RDGIBSON |  | Concentration | mg/L | Y | (fl. conc.) |  | Hydrogen Sulfide (fl. conc.) | 7783-06-4 | Volatile Organic Compounds (VOCs); | Volatile Organic Compounds (VOCs); EMS code: H2S- | Chemical |  | LAB
H2S- | ug/m3 | Hydrogen Sulfide | EQ40 | EQUIS (mg/m3) | EQUIS (mg/m3) | 0 | N | Y | Y | 6036 |  |  | 10/22/1996 11:35:33 PM | 1996-04-12 | 6/3/1998 6:51:36 PM | EMS |  | AirConcentration | µg/m³ | Y | (air conc.) |  | Hydrogen Sulfide (air conc.) | 7783-06-4 | Volatile Organic Compounds (VOCs); | Volatile Organic Compounds (VOCs); EMS code: H2S- | Chemical |  | LAB
H2S- | mg/L | Hydrogen Sulfide | XCAL | Calculation (mg/L) | Calculation (mg/L) | 0.0021 | Y | Y | Y | 103 | 5/13/2022 | 6/26/2024 | 6/26/2024 4:28:10 PM | 2018-09-04 |  | RDGIBSON |  | Concentration | mg/L | Y | (fl. conc.) |  | Hydrogen Sulfide (fl. conc.) | 7783-06-4 | Volatile Organic Compounds (VOCs); | Volatile Organic Compounds (VOCs); EMS code: H2S- | Chemical |  | LAB
H2S- | ppm A | Hydrogen Sulfide | X186 | Gas Chromatography-Ave/d | Gas Chromatography-Ave/d | 0 | Y | Y | Y | 1731 |  |  | 8/22/2024 1:49:14 PM | 1996-04-12 | 8/27/2020 11:16:51 AM | EMS | RDGIBSON | DimensionlessRatio | ppm | N | (ratio) |  | Hydrogen Sulfide (ratio) | 7783-06-4 | Volatile Organic Compounds (VOCs); | Volatile Organic Compounds (VOCs); EMS code: H2S- | Chemical |  | LAB
HAA5 | mg/L | Total Haloacetic Acids | THMHAA | THM/HAA CALC | THM/HAA CALC | 0.01 | Y | Y | Y | 2 | 3/27/2023 9:58:00 AM | 5/31/2023 4:53:00 PM | 6/8/2023 12:14:57 PM | 2023-03-31 |  | EMS |  | Concentration | mg/L | Y | (fl. conc.) |  | Total Haloacetic Acids (fl. conc.) |  | Organics; | Organics; EMS code: HAA5 | Chemical | Total | LAB
HAO | mg/L | 10-Hydroxy Amitriptyline Oxalate CAS: 1246833-15-7 | XMSP | HPLC and Tri. Quadpole MS/MS | EPA 1694 An aliquot of filtered sample containing 5% organic is injected directly. Each compound is separated by reversed phase HPLC. Analyzed by Electro Spray Ionization MS/MS detection using Triple Quadrupole MS/MS detector. Quantiitation with internal or external standard. |  | Y | Y | Y | 12 | 5/8/2020 | 6/4/2024 | 8/21/2024 3:32:03 PM | 2022-05-04 |  | EMS |  | Concentration | mg/L | Y | (fl. conc.) |  | 10-Hydroxy Amitriptyline Oxalate (fl. conc.) | 1246833-15-7 | Organics; | Organics; EMS code: HAO | Chemical |  | LAB
HCBD | mg/L | Hexachlorobutadiene | E660 | EPA 8270E (mod) | EPA 8270E (mod) Miscellaneous Pesticides by GC-MS | 2e-06 | Y | Y | Y | 2 | 9/14/2022 | 9/14/2022 | 12/8/2022 9:48:14 AM | 2022-12-06 |  | EMS |  | Concentration | mg/L | Y | (fl. conc.) |  | Hexachlorobutadiene (fl. conc.) | 87-68-3 | Volatile Organic Compounds (VOCs); | Volatile Organic Compounds (VOCs); EMS code: HCBD | Chemical |  | LAB
HCBD | ug/g | Hexachlorobutadiene | ZN03 | DCM+Ac.Extr:GC/MS | DCM+Ac.Extr:GC/MS EPA 8270E (mod) |  | Y | Y | Y | 17 | 1/17/2022 | 11/14/2022 | 1/11/2024 1:07:07 PM | 2007-10-04 |  | RDGIBSON |  | MassPerMass | mg/kg | Y | (mass conc.) |  | Hexachlorobutadiene (mass conc.) | 87-68-3 | Volatile Organic Compounds (VOCs); | Volatile Organic Compounds (VOCs); EMS code: HCBD | Chemical |  | LAB
HCBD | mg/L | Hexachlorobutadiene | X268 | CH2Cl2 ext:GC/MS | CH2Cl2 ext:GC/MS |  | Y | Y | Y | 16 | 12/29/2021 | 9/2/2022 | 11/10/2022 9:54:07 AM | 2022-04-08 |  | EMS |  | Concentration | mg/L | Y | (fl. conc.) |  | Hexachlorobutadiene (fl. conc.) | 87-68-3 | Volatile Organic Compounds (VOCs); | Volatile Organic Compounds (VOCs); EMS code: HCBD | Chemical |  | LAB
HCBD | mg/L | Hexachlorobutadiene | MS04 | Static Headspace GC/MS (Water) | (Water) A static headspace procedure is used as a means of driving the analytes out of the matrix into a headspace. Once equilibrium is achieved, a fixed volume of headspeace is taken for subsequent determination by Gas Chromatography - Mass Spectroscopy (GC/MS) using Selected Ion Monitoring. | 0.0005 | Y | Y | Y | 8 | 10/5/2019 2:07:00 AM | 6/28/2024 12:29:00 AM | 7/9/2024 3:31:02 PM | 2021-04-07 |  | RDGIBSON |  | Concentration | mg/L | Y | (fl. conc.) |  | Hexachlorobutadiene (fl. conc.) | 87-68-3 | Volatile Organic Compounds (VOCs); | Volatile Organic Compounds (VOCs); EMS code: HCBD | Chemical |  | LAB
HCCH | mg/L | Total Hexachlorocyclohexane | X268 | CH2Cl2 ext:GC/MS | CH2Cl2 ext:GC/MS |  | Y | Y | Y | 16 | 12/29/2021 | 9/2/2022 | 11/10/2022 9:54:07 AM | 2022-04-08 |  | EMS |  | Concentration | mg/L | Y | (fl. conc.) |  | Total Hexachlorocyclohexane (fl. conc.) |  | Pesticides; | Pesticides; EMS code: HCCH | Chemical | Total | LAB
HCCH | ug/g | Total Hexachlorocyclohexane | ZN03 | DCM+Ac.Extr:GC/MS | DCM+Ac.Extr:GC/MS EPA 8270E (mod) | 0.0009 | Y | Y | Y | 2 | 11/14/2022 | 11/14/2022 | 12/15/2022 10:50:04 AM | 2022-12-14 |  | EMS |  | MassPerMass | mg/kg | Y | (mass conc.) |  | Total Hexachlorocyclohexane (mass conc.) |  | Pesticides; | Pesticides; EMS code: HCCH | Chemical | Total | LAB
HCCH | mg/L | Total Hexachlorocyclohexane | E660 | EPA 8270E (mod) | EPA 8270E (mod) Miscellaneous Pesticides by GC-MS | 1e-05 | Y | Y | Y | 2 | 9/14/2022 | 9/14/2022 | 12/8/2022 9:48:14 AM | 2022-12-06 |  | EMS |  | Concentration | mg/L | Y | (fl. conc.) |  | Total Hexachlorocyclohexane (fl. conc.) |  | Pesticides; | Pesticides; EMS code: HCCH | Chemical | Total | LAB
HCH | ug/g | Hexachloroethane CAS: 67-72-1 | ZN03 | DCM+Ac.Extr:GC/MS | DCM+Ac.Extr:GC/MS EPA 8270E (mod) | 0.01 | Y | Y | Y | 17 | 1/17/2022 | 11/14/2022 | 1/11/2024 1:07:07 PM | 2022-07-12 |  | EMS |  | MassPerMass | mg/kg | Y | (mass conc.) |  | Hexachloroethane (mass conc.) | 67-72-1 | Volatile Organic Compounds (VOCs); | Volatile Organic Compounds (VOCs); EMS code: HCH | Chemical |  | LAB
HCH | mg/L | Hexachloroethane CAS: 67-72-1 | X268 | CH2Cl2 ext:GC/MS | CH2Cl2 ext:GC/MS |  | Y | Y | Y | 16 | 12/29/2021 | 9/2/2022 | 11/10/2022 9:54:07 AM | 2022-04-04 |  | EMS |  | Concentration | mg/L | Y | (fl. conc.) |  | Hexachloroethane (fl. conc.) | 67-72-1 | Volatile Organic Compounds (VOCs); | Volatile Organic Compounds (VOCs); EMS code: HCH | Chemical |  | LAB
HCH | mg/L | Hexachloroethane CAS: 67-72-1 | E660 | EPA 8270E (mod) | EPA 8270E (mod) Miscellaneous Pesticides by GC-MS | 2e-06 | Y | Y | Y | 2 | 9/14/2022 | 9/14/2022 | 12/8/2022 9:48:14 AM | 2022-12-06 |  | EMS |  | Concentration | mg/L | Y | (fl. conc.) |  | Hexachloroethane (fl. conc.) | 67-72-1 | Volatile Organic Compounds (VOCs); | Volatile Organic Compounds (VOCs); EMS code: HCH | Chemical |  | LAB
HCO3 | mg/L | bicarbonate (CAS 71-52-3) | X310 | Auto Potentiometric Titration | For alkalinity, the samples are (if pH >8.3) first titrated to a pH of 8.3, which corresponds to the equivalence point for the conversion of carbonate to bicarbonate ion (phenolphthalein alkalinity). The titration is continued to a pH of 4.5, which corresponds approx. to the equivalence point for the conversion of bicarbonate ion to carbonic acid (entire titration indicates total alkalinity). An extrapolation technique based on the near proportionality of concentration of H ions to excess of titrant beyond the equivalence pt. The amount of standard acid required to reduce pH exactly 0.30pH (ie. pH4.2) unit is measured carefully. Because this change in pH corresponds to an exact coupling of the H ion concentration, a simple extrapolation can be made to the equivalence point. The actual pH of the stoichiometric end point in alkalinity determinations is the pH at which the inflection in the titration curve occurs and is directly related to the amount of carbonate alkalinity originally present in the sample. A curve can be constructed by plotting pH vs mL of acid to find this point directly. *** For acidity, the samples are titrated to a pH of 8.3, which corresponds to stoichiometric neutralization of carbonic acid to bicarbonate. Because the colour change of phenolphthalein indicator is close to pH 8.3, this value generally is accepted as a standard end point for titration of total acidity, including CO2 and most weak acids an equivalence point for the conversion of carbonate to bicarbonate ion (phenolphthalein alkalinity). The actual pH of the stoichiometric end point in acidity determinations is the pH at which the inflection in the titration curve occurs and is directly related to the amount of titratable acids originally present in the sample. A curve can be constructed by plotting pH vs. mL of NaOH to find this point directly. (To obtain the actual substances contributing to acidity/alkalinity, a pH vs mL of acid/NaOH curve should be constructed.) | 1 | Y | Y | Y | 99 | 11/24/2019 | 8/11/2024 | 8/20/2024 9:51:05 AM | 2020-01-29 |  | RDGIBSON |  | Concentration | mg/L | Y | (fl. conc.) |  | bicarbonate (fl. conc.) | 71-52-3 | Major Ions; | Major Ions; EMS code: HCO3 | Chemical |  | LAB
HD01 | mg/L | Hexadecanoic Acid | RFAS | Separation of resin and fatty acids (RFAs) in water by GC/MS | Scientific Principle The separation of resin and fatty acids (RFAs) in water samples is performed using gas chromatography/mass spectrometry (GC/MS) in the Selected Ion Monitoring (SIM) mode. The separation is based on compound size; boiling point and polarity. Analyte concentration is calculated using the internal standard method. RFA compounds are extracted from water samples using a liquid-liquid solvent extraction with DCM. Extraction occurs after the sample is brought to a pH of ~ 2 and placed in a rotary tumbler (at approx. 40 rpm) for 1 hour. The resulting extract is transferred to a vial; dried with acidified baked sodium sulfate; methylated with diazomethane and concentrated to 1mL under nitrogen prior to analysis. The resulting extracts are vialed and fortified with labeled internal standard compounds prior to analysis by gas chromatography with a (50%-Phenyl)-methylpolysiloxane (DB-17) column and a mass spectrometer detector in SIM Mode. | 0.006 | Y | Y | Y | 1 | 9/17/2021 7:49:00 PM | 9/17/2021 7:49:00 PM | 2/1/2022 2:38:14 PM | 2022-01-19 |  | EMS |  | Concentration | mg/L | Y | (fl. conc.) |  | Hexadecanoic Acid (fl. conc.) |  | Organics; | Organics; EMS code: HD01 | Chemical |  | LAB
HEPA | mg/L | Heptanoic Acid | VFAD | Volatile Fatty Acids by GC/FID | Volatile Fatty Acids by GC/FID; Standard Methods for the Examination of Water and Wastewater; 21st Edition, Method 5560 D; pages 5-59 to 5-61. 5560 D. Gas Chromatographic Method Principle: This gas chromatographic procedure may be usedto determine the individual concentrations of the following fattyacids: acetic; propionic; butyric; isobutyric; valeric; andisovaleric acids in digester sludge samples. Caproic and heptanoicacids also may be determined; but with less certaintybecause of their reduced solubility in water. Volatile fatty acidsare generally soluble in digester sludge and exist in either theprotonated or unprotonated form; depending on the pH of thesludge sample. Acetic and propionic acids usually are in higherconcentrations than the other fatty acids in digester sludge. Allacids must be converted to the protonated (volatile) form beforeinjection into the gas chromatograph (GC) so vaporization mayoccur. The fatty acids must also be separated from the solidsmaterial in the sample before introduction into the GC to minimizedegradation of the GC column. Digester sludge samples areprepared by acidification; centrifugation; and filtration. The sampleis analyzed by direct injection into a gas chromatographequipped with a flame ionization detector after both centrifugationand filtration.b. Interference: The blank amount may be high relative tosample analyte. High blanks may be caused by a buildup ofcontaminants in the injector and guard column or by samplecarry-over. The analyst must be aware of the blank levels andcorrect unacceptably high blank levels that are above the reportinglimits.c. Minimum detectable concentration: The method detectionlevel has been determined by the USEPA method1 in reagentwater to be about 3 mg/L for acetic acid and about 1 mg/L for allother target compounds. | 2 | Y | Y | Y | 5 | 3/14/2017 10:30:00 AM | 11/19/2017 5:10:00 PM | 6/20/2018 9:02:11 AM | 2018-06-19 |  | RDGIBSON |  | Concentration | mg/L | Y | (fl. conc.) |  | Heptanoic Acid (fl. conc.) |  | Organics; | Organics; EMS code: HEPA | Chemical |  | LAB
HEPH | mg/L | Heavy Extractable Petroleum Hydrocarbons | XCAL | Calculation (mg/L) | Calculation (mg/L) | 0.25 | Y | Y | Y | 464 | 8/28/2014 | 8/15/2024 | 8/16/2024 1:36:03 PM | 2009-10-06 | 8/15/2019 1:28:05 PM | RDGIBSON | KATFRASE | Concentration | mg/L | Y | (fl. conc.) |  | Heavy Extractable Petroleum Hydrocarbons (fl. conc.) |  | Petroleum Hydrocarbons; | Petroleum Hydrocarbons; EMS code: HEPH | Chemical | Extractable | LAB
HEPH | ug/g | Heavy Extractable Petroleum Hydrocarbons | F086 | EPH:Hex/Actn,soxhlet extr'n,GC/FID(PAHs not sub) | Samples are dried and soxhlet extracted with 1:1 Hexane:Acetone for 4 hours. The extract is concentrated and analyzed by capillary gas chromatography with flame ionization detector. Results are not corrected for specified PAH's. LEPH is defined as the sum of components that elute on gc between nC10 and nC19; HEPH is defined as the sum of components that elute on gc between nC19 and nC32. Performance Base Method published circa Dec. 31, 2000. | 250 | Y | Y | Y | 358 | 6/19/2000 1:42:00 PM | 9/9/2022 | 1/11/2024 1:07:07 PM | 2000-07-14 |  | NPEPPIN |  | MassPerMass | mg/kg | Y | (mass conc.) |  | Heavy Extractable Petroleum Hydrocarbons (mass conc.) |  | Petroleum Hydrocarbons; | Petroleum Hydrocarbons; EMS code: HEPH | Chemical | Extractable | LAB
HEPH | ug/g | Heavy Extractable Petroleum Hydrocarbons | F066 | DCM,soxhlet extr'n,GC/FID(PAHs sub.) | Samples are dried and soxhlet extracted with dichloromethane (DCM) for 4 hours. The extract is concentrated and analyzed by capillary gas chromatography with flame ionization detector. Results are corrected for specified PAH concentrations and percent moisture. LEPH is defined as the sum of components that elute on gc between nC1O and nC19; HEPH is defined as the sum of components that elute on gc between nC19 and nC32. Method superceded Dec. 2000. | 250 | Y | Y | Y | 10 | 7/9/2001 9:40:00 AM | 9/19/2001 11:33:00 AM | 3/17/2003 1:22:49 PM | 1999-06-28 |  | NPEPPIN |  | MassPerMass | mg/kg | Y | (mass conc.) |  | Heavy Extractable Petroleum Hydrocarbons (mass conc.) |  | Petroleum Hydrocarbons; | Petroleum Hydrocarbons; EMS code: HEPH | Chemical | Extractable | LAB
HEPH | mg/L | Heavy Extractable Petroleum Hydrocarbons | TEHW | Total Ext-Hydrocarbons in Water by GCFID | Total Ext-Hydrocarbons in Water by GCFID, BC Lab Manual | 0.25 | Y | Y | Y | 647 | 5/5/2015 | 10/12/2022 1:33:00 PM | 3/10/2023 3:28:44 PM | 2015-11-02 |  | RDGIBSON |  | Concentration | mg/L | Y | (fl. conc.) |  | Heavy Extractable Petroleum Hydrocarbons (fl. conc.) |  | Petroleum Hydrocarbons; | Petroleum Hydrocarbons; EMS code: HEPH | Chemical | Extractable | LAB
HEPH | mg/L | Heavy Extractable Petroleum Hydrocarbons | EPH2 | EPHnC19-nC32 extr. silica gel: GC/FID(PAH not subtr.) | Heavy Extractable Petroleum Hydrocarbon | 0.25 | Y | Y | Y | 54 |  |  | 7/31/2018 10:37:13 PM | 2009-10-06 |  | RDGIBSON |  | Concentration | mg/L | Y | (fl. conc.) |  | Heavy Extractable Petroleum Hydrocarbons (fl. conc.) |  | Petroleum Hydrocarbons; | Petroleum Hydrocarbons; EMS code: HEPH | Chemical | Extractable | LAB
HEPH | mg/L | Heavy Extractable Petroleum Hydrocarbons | F064 | EPH:DCM,liq-liq extr'n,GC/FID,(PAH sub.) | Samples are extracted with dichloromethane (DCM) in a separatory funnel. Extracts are dried, concentrated and analyzed by capillary gas chromatography with flame ionization detector. Results are corrected for specified PAH concentrations. LEPH is defined as the sum of components that elute on GC between nC10 and nC19; HEPH is defined as the sum of components that elute on gc between nC19 and nC32. Performance Base Method published circa Dec. 31, 2000, but distributed 1999. | 1 | Y | Y | Y | 740 | 6/12/2000 | 7/3/2024 | 7/22/2024 3:17:29 PM | 1999-06-28 |  | NPEPPIN |  | Concentration | mg/L | Y | (fl. conc.) |  | Heavy Extractable Petroleum Hydrocarbons (fl. conc.) |  | Petroleum Hydrocarbons; | Petroleum Hydrocarbons; EMS code: HEPH | Chemical | Extractable | LAB
HEPH | ug/g | Heavy Extractable Petroleum Hydrocarbons | CAL2 | Calculation (ug/g) | Calculation (ug/g) |  | Y | Y | Y | 391 | 5/21/2015 | 9/9/2022 | 7/19/2024 8:20:35 AM | 2009-10-06 |  | RDGIBSON |  | MassPerMass | mg/kg | Y | (mass conc.) |  | Heavy Extractable Petroleum Hydrocarbons (mass conc.) |  | Petroleum Hydrocarbons; | Petroleum Hydrocarbons; EMS code: HEPH | Chemical | Extractable | LAB
HEPH | ug/g | Heavy Extractable Petroleum Hydrocarbons | F067 | DCM,soxhlet extr'n,GC/FID,(PAHs not sub.) | Samples are dried and soxhlet extracted with dichloromethane (DCM) for 4 hours. The extract is concentrated and analyzed by capillary gas chromatography with flame ionization detector. Results are corrected for percent moisture, but are not corrected for specified PAH's. LEPH is defined as the sum of components that elute on gc between nC10 and nC19; HEPH is defined as the sum of components that elute on gc between nC19 and nC32. Method superceded Dec. 2000. | 250 | Y | Y | Y | 18 | 7/9/2001 9:39:00 AM | 9/19/2001 11:15:00 AM | 3/17/2003 1:22:49 PM | 1999-06-28 |  | NPEPPIN |  | MassPerMass | mg/kg | Y | (mass conc.) |  | Heavy Extractable Petroleum Hydrocarbons (mass conc.) |  | Petroleum Hydrocarbons; | Petroleum Hydrocarbons; EMS code: HEPH | Chemical | Extractable | LAB
HEPH | ug/g | Heavy Extractable Petroleum Hydrocarbons | F085 | EPH:Hex/Actn,soxhlet extr'n,GC/FID(PAHs sub) | Samples are dried and soxhlet extracted with 1:1 Hexane:Acetone for 4 hours. The extract is concentrated and analyzed by capillary gas chromatography with flame ionization detector. Results are corrected for specified PAH concentrations. LEPH is defined as the sum of components that elute on gc between nC10 and nC19; HEPH is defined as the sum of components that elute on gc between nC19 and nC32. Performance Base Method published circa Dec. 31, 2000. | 250 | Y | Y | Y | 234 | 6/19/2000 1:42:00 PM | 10/29/2022 9:40:00 AM | 1/7/2023 3:12:10 PM | 2000-07-14 |  | NPEPPIN |  | MassPerMass | mg/kg | Y | (mass conc.) |  | Heavy Extractable Petroleum Hydrocarbons (mass conc.) |  | Petroleum Hydrocarbons; | Petroleum Hydrocarbons; EMS code: HEPH | Chemical | Extractable | LAB
HEPH | mg/L | Heavy Extractable Petroleum Hydrocarbons | EPH3 | EPH, Hexane micro-extrn, GC/FID (Waters, w/o Si-Gel) | A single extraction using a small volume of hexane is conducted within the original sample container using vigorous mechanical agitation. A portion of the hexane extract is concentrated by evaporation if necessary, and is analyzed by capillary column gas chromatography with flame ionization detection. Test results for EPHw10-19 and EPHw19-32 are not corrected for PAHs. LEPHw and HEPHw are EPHw10-19 and EPHw19-32 minus the concentrations of applicable PAHs found in each fraction. Performance Based Method, revision date Nov 1, 2015. | 0.25 | Y | Y | Y | 299 | 1/21/2016 10:45:00 AM | 12/3/2022 8:48:00 AM | 12/15/2022 6:20:11 PM | 2015-10-23 |  | RDGIBSON |  | Concentration | mg/L | Y | (fl. conc.) |  | Heavy Extractable Petroleum Hydrocarbons (fl. conc.) |  | Petroleum Hydrocarbons; | Petroleum Hydrocarbons; EMS code: HEPH | Chemical | Extractable | LAB
HEPH | mg/L | Heavy Extractable Petroleum Hydrocarbons | F065 | EPH:DCM,liq-liq extr'n,GC/FID,(PAH not sub.) | Samples are extracted with dichloromethane (DCM) in a separatory funnel. Extracts are dried, concentrated and analyzed by capillary gas chromatography with flame ionization detector. Results are not corrected for specified PAH concentrations. LEPH is defined as the sum of components that elute on gc between nC10 and nC19; HEPH is defined as the sum of components that elute on gc between nC19 and nC32. Performance Base Method published circa Dec. 31, 2000, but distributed 1999. | 1 | Y | Y | Y | 1710 | 6/12/2000 | 8/21/2024 | 8/22/2024 5:27:05 PM | 1999-06-28 |  | NPEPPIN |  | Concentration | mg/L | Y | (fl. conc.) |  | Heavy Extractable Petroleum Hydrocarbons (fl. conc.) |  | Petroleum Hydrocarbons; | Petroleum Hydrocarbons; EMS code: HEPH | Chemical | Extractable | LAB
HF-T | ppm (S) | Hafnium | MS81 | ALS Minerals: ME-MS81 | ALS Minerals: ME-MS81. A prepared sample (0.100 g) is added to lithium metaborate/lithium tetraborate flux; mixed well and fused in a furnace at 1025°C. The resulting melt is then cooled and dissolved in an acid mixture containing nitric; hydrochloric and hydrofluoric acids. This solution is then analyzed by inductively coupled plasma - mass spectrometry. | 0.1 | Y | Y | Y | 0 |  |  |  | 2024-01-29 |  | EMS |  | DimensionlessRatio | ppm | N | (ratio) |  | Hafnium (ratio) |  | Metals; | Metals; EMS code: HF-T | Chemical |  | LAB
HF-T | ppm (S) | Hafnium | 6041 | Aqua Regia by ICPMS | These were sediment/soils digested with Aqua-Regia (Nitric Acid & Hydrochloric Acids) and then analysis by ICPMS. Results would be reported in mg/kg and/or %. | 0.002 | Y | Y | Y | 21 | 10/13/2019 | 8/5/2020 | 8/10/2020 3:52:06 PM | 2019-09-06 | 10/21/2019 11:55:11 AM | KATFRASE | KATFRASE | DimensionlessRatio | ppm | N | (ratio) |  | Hafnium (ratio) |  | Metals; | Metals; EMS code: HF-T | Chemical |  | LAB
HF-T | ppm (S) | Hafnium | 6042 | Aqua Regia by ICPMS | These were sediment/soils digested with Aqua-Regia (Nitric Acid & Hydrochloric Acids) and then analysis by ICPMS. Results would be reported in ppm s. | 0.002 | Y | Y | Y | 71 | 5/14/2022 | 7/27/2024 | 7/29/2024 9:13:08 AM | 2022-06-28 |  | EMS |  | DimensionlessRatio | ppm | N | (ratio) |  | Hafnium (ratio) |  | Metals; | Metals; EMS code: HF-T | Chemical |  | LAB
HF-T | ppm (S) | Hafnium | MS41 | ALS Minerals ME-MS41 | ALS Minerals ME-MS41. A prepared sample is digested with aqua regia in a graphite heating block. After cooling; the resulting solution is diluted to with deionized water; mixed and analyzed by ICP-AES. Samples are then analyzed by ICP-MS for the remaining suite of elements. | 0.02 | Y | Y | Y | 0 |  |  |  | 2024-01-29 |  | EMS |  | DimensionlessRatio | ppm | N | (ratio) |  | Hafnium (ratio) |  | Metals; | Metals; EMS code: HF-T | Chemical |  | LAB
HF-T | ppm (S) | Hafnium | MS61 | ALS Minerals ME-MS61 | ALS Minerals ME-MS61. A prepared sample is digested with HNO3; HF; and HClO4 in a graphite heating block to near dryness. After cooling; the sample is boiled in diluted HCl. Bulk to volume with deionized water; mixed and analyzed by ICP-AES and ICP-MS for all elements. | 0.1 | Y | Y | Y | 0 |  |  |  | 2024-01-29 |  | EMS |  | DimensionlessRatio | ppm | N | (ratio) |  | Hafnium (ratio) |  | Metals; | Metals; EMS code: HF-T | Chemical |  | LAB
HFPO | mg/L | Hexafluoropropylene oxide dimer acid (HFPO-DA)CAS:13252-13-6 | E3533 | MECP E3533 | PFAs in water by direct injection LC/MS/MS | 2e-05 | Y | Y | Y | 2 | 6/3/2024 | 6/3/2024 | 8/21/2024 3:32:03 PM | 2023-11-15 | 6/27/2024 5:06:31 PM | EMS | EMS | Concentration | mg/L | Y | (fl. conc.) |  | Hexafluoropropylene oxide dimer acid (HFPO-DA) (fl. conc.) | 13252-13-6 | Per- and polyfluoroalkyl substances (PFAS); | Per- and polyfluoroalkyl substances (PFAS); EMS code: HFPO | Chemical |  | LAB
HG-D | mg/L | Mercury Dissolved | EQ01 | EQUIS (mg/L) | EQUIS (mg/L) | 0 | N | Y | Y | 455 |  |  | 10/22/1996 9:50:33 PM | 1996-04-12 | 6/3/1998 6:51:36 PM | EMS |  | Concentration | mg/L | Y | (fl. conc.) |  | Mercury Dissolved (fl. conc.) |  | Metals - Dissolved; | Metals - Dissolved; EMS code: HG-D | Chemical | Dissolved | LAB
HG-D | ng/L | Mercury Dissolved | T548 | EPA1631E-T | APHA 3030B / EPA 1631E - Trace | 0.1 | Y | Y | Y | 0 |  |  |  | 2024-08-19 |  | EMS |  | Concentration | ng/L | Y | (fl. conc.) |  | Mercury Dissolved (fl. conc.) |  | Metals - Dissolved; | Metals - Dissolved; EMS code: HG-D | Chemical | Dissolved | LAB
HG-D | ng/L | Mercury Dissolved | E1631E | EPA 1631E | See EPA method 1631E | 0.5 | Y | Y | Y | 245 | 7/24/2020 | 4/4/2024 | 4/29/2024 4:53:06 PM | 2023-05-23 |  | EMS |  | Concentration | ng/L | Y | (fl. conc.) |  | Mercury Dissolved (fl. conc.) |  | Metals - Dissolved; | Metals - Dissolved; EMS code: HG-D | Chemical | Dissolved | LAB
HG-D | mg/L | Mercury Dissolved | X204 | FF:HNO3/KMnO4,Cold V,AA | FF:HNO3/KMnO4;Cold V;AA | 5e-05 | Y | Y | Y | 690 | 5/9/1998 | 12/23/2003 | 1/8/2004 3:18:34 PM | 1996-04-12 | 6/3/1998 6:51:36 PM | EMS |  | Concentration | mg/L | Y | (fl. conc.) |  | Mercury Dissolved (fl. conc.) |  | Metals - Dissolved; | Metals - Dissolved; EMS code: HG-D | Chemical | Dissolved | LAB
HG-D | mg/L | Mercury Dissolved | X347 | FF,FA:HNO3/K2Cr2O7:CVAA | The sample is field filtered (0.45 um) and field acidified (HNO3); strong oxidizing digestion; analysis by CVAA. | 5e-05 | Y | Y | Y | 331 |  |  | 1/30/2002 4:36:50 PM | 1996-04-12 | 6/3/1998 6:51:41 PM | EMS |  | Concentration | mg/L | Y | (fl. conc.) |  | Mercury Dissolved (fl. conc.) |  | Metals - Dissolved; | Metals - Dissolved; EMS code: HG-D | Chemical | Dissolved | LAB
HG-D | mg/L | Mercury Dissolved | X222 | FF,FA,HNO3,Flameless AA | FF;FA;HNO3;Flameless AA | 0 | Y | Y | Y | 34 |  |  | 10/24/1996 5:42:36 AM | 1996-04-12 | 6/3/1998 6:51:36 PM | EMS |  | Concentration | mg/L | Y | (fl. conc.) |  | Mercury Dissolved (fl. conc.) |  | Metals - Dissolved; | Metals - Dissolved; EMS code: HG-D | Chemical | Dissolved | LAB
HG-D | mg/L | Mercury Dissolved | MSCV | HG-200.2-CVAF-VA (WET) | Mercury in Soil by CVAFS (WET) | 5e-06 | Y | Y | Y | 725 | 4/11/2012 | 4/18/2020 | 12/8/2023 4:30:26 PM | 2021-05-10 | 5/10/2021 4:33:57 PM | RDGIBSON | RDGIBSON | Concentration | mg/L | Y | (fl. conc.) |  | Mercury Dissolved (fl. conc.) |  | Metals - Dissolved; | Metals - Dissolved; EMS code: HG-D | Chemical | Dissolved | LAB
HG-D | mg/L | Mercury Dissolved | X090 | Decant,Dig,Flameless AA | Decant;Dig;Flameless AA | 0 | Y | Y | Y | 192 | 3/1/2000 | 11/13/2002 | 2/23/2004 2:18:50 PM | 1996-04-12 | 6/3/1998 6:51:36 PM | EMS |  | Concentration | mg/L | Y | (fl. conc.) |  | Mercury Dissolved (fl. conc.) |  | Metals - Dissolved; | Metals - Dissolved; EMS code: HG-D | Chemical | Dissolved | LAB
HG-D | mg/L | Mercury Dissolved | CVAL | Hg CVAF low level | Mercury in Water, Leachates and Liquids by Bromination and Cold Vapour Atomic Fluorescence Spectroscopy (low level) PURPOSE/PRINCIPLE OF METHOD: Chemically generated bromine is used to oxidize all organic and inorganic forms of mercury in aqueous samples to mercury (II). Excess bromine is neutralized with hydroxylamine hydrochloride. The solution is introduced into a gas liquid separator and mixed with stannous chloride to reduce the mercury to its elemental form. The mercury is then purged from the solution with argon gas and passes through an atomic fluorescence detector where it absorbs light from a mercury vapour lamp that fluoresces at 254 nm. This is detected and amplified by a photo-multiplier tube in the detector. SCOPE: The method applies to the determination of field filtered, dissolved and total mercury in water, wastewater, effluent, as well as hot acid digested liquids and aqueous waste leachates generated by Leachate Extraction Procedure, Toxicity Characteristic Leaching Procedure (TCLP), Synthetic Precipitation Leaching Procedure (SPLP) and Simultaneously Extractable Metals (SEM) extracts. DETECTION LIMITS: Routine RDL = 0.02 ug/L Ultra-Low level RDL = 0.002 ug/L REFERENCE: Determination of Mercury in Water by Semi-Automated Cold Vapor Atomic Fluorescence Spectrometry. EPA Method 245.7, Revision 2, February 2005 | 2e-06 | Y | Y | Y | 3220 | 7/12/2011 9:03:00 AM | 8/9/2024 6:00:00 PM | 8/19/2024 2:37:08 PM | 2012-02-03 |  | RDGIBSON |  | Concentration | mg/L | Y | (fl. conc.) |  | Mercury Dissolved (fl. conc.) |  | Metals - Dissolved; | Metals - Dissolved; EMS code: HG-D | Chemical | Dissolved | LAB
HG-D | mg/L | Mercury Dissolved | F021 | Acid Dig,Auto,CV Fluo.Spec. | Mercury in organic and inorganic compounds is oxidized to mercuric ions by addition of sulfuric acid, nitric acid, potassium persulfate and heat. The mercuric ions are reduced to Hg(0) with SnCl2, are sparged from the solution with argon, and passed through a fluorescence cell in the light path of a mercury lamp. The fluorescence at the 253.7 nm mercury emission line by Hg(0) is proportional to the concentration of mercury in the sample. | 5e-05 | Y | Y | Y | 29 | 8/30/1996 11:35:00 AM | 12/3/2001 12:36:00 PM | 6/7/2024 5:30:18 PM | 1996-09-17 | 6/3/1998 6:51:45 PM | MPAK |  | Concentration | mg/L | Y | (fl. conc.) |  | Mercury Dissolved (fl. conc.) |  | Metals - Dissolved; | Metals - Dissolved; EMS code: HG-D | Chemical | Dissolved | LAB
HG-D | mg/L | Mercury Dissolved | 0300 | Autoanal/Cold Vap. | Autoanal/Cold Vap. | 5e-05 | Y | Y | Y | 26221 | 4/19/2007 | 1/15/2012 2:54:00 PM | 12/18/2023 2:41:39 PM | 1996-04-12 | 6/3/1998 6:51:36 PM | EMS |  | Concentration | mg/L | Y | (fl. conc.) |  | Mercury Dissolved (fl. conc.) |  | Metals - Dissolved; | Metals - Dissolved; EMS code: HG-D | Chemical | Dissolved | LAB
HG-D | mg/L | Mercury Dissolved | MWID | Metals in Water by ICPMS (digested) | Metals in Water by ICPMS (digested) Based on EPA 200.8 | 0.0002 | Y | Y | Y | 31 |  |  | 7/30/2024 6:32:12 AM | 2009-12-24 |  | RDGIBSON |  | Concentration | mg/L | Y | (fl. conc.) |  | Mercury Dissolved (fl. conc.) |  | Metals - Dissolved; | Metals - Dissolved; EMS code: HG-D | Chemical | Dissolved | LAB
HG-D | mg/L | Mercury Dissolved | F032 | Filt.,Acid,ICPMS | There is no acid digestion required for this analysis. A water sample is collected and preserved in 0.2 percent nitric acid. To analyse the extractable metals the sample is shaken, allowed to sit for at least twenty four hours, and a decanted portion is introduced to the spectrometer. When the dissolved metals are required, the samples must first be filtered through a 0.45 micron cellulose acetate filter paper and then followed by preservation in 0.2 percent nitric acid. It is important to note that different laboratories may have very different detection limits for some variables analyzed by ICPMS. | 1e-05 | Y | Y | Y | 2618 | 5/5/2008 2:11:00 PM | 6/24/2024 7:11:00 PM | 6/28/2024 5:01:20 PM | 2008-06-09 | 7/16/2012 3:34:31 PM | RDGIBSON | RDGIBSON | Concentration | mg/L | Y | (fl. conc.) |  | Mercury Dissolved (fl. conc.) |  | Metals - Dissolved; | Metals - Dissolved; EMS code: HG-D | Chemical | Dissolved | LAB
HG-D | mg/L | Mercury Dissolved | X273 | Cold Vapour:AA | Cold Vapour:AA | 0.005 | Y | Y | Y | 2051 | 2/2/2001 | 4/30/2008 1:31:00 PM | 5/19/2015 3:10:33 PM | 1996-04-12 | 6/3/1998 6:51:36 PM | EMS |  | Concentration | mg/L | Y | (fl. conc.) |  | Mercury Dissolved (fl. conc.) |  | Metals - Dissolved; | Metals - Dissolved; EMS code: HG-D | Chemical | Dissolved | LAB
HG-D | mg/L | Mercury Dissolved | 0310 | HNO3/H2SO4 Dig:Man.C.V. | Samples are field filtered for dissolved or not filtered for total, and preserved with K2Cr2O7/H2SO4 to pH <2. The organomercury compounds are oxidized via a H2SO4/KMnO4/K2S2O8 digestion. The mercury compounds are then reduced to elemental mercury. A stream of synthetic air introduced into the system passes through a reference cell and then, after sparging the elemental mercury from the solution, passes through the sample cell. A UV monitor possessing a mercury lamp provides the light source and the absorbance at 253.7 nm is measured against the reference. | 5e-05 | Y | Y | Y | 382 | 5/9/2003 9:40:00 AM | 9/12/2007 | 10/3/2007 1:27:54 AM | 1996-04-12 | 6/3/1998 6:51:36 PM | EMS |  | Concentration | mg/L | Y | (fl. conc.) |  | Mercury Dissolved (fl. conc.) |  | Metals - Dissolved; | Metals - Dissolved; EMS code: HG-D | Chemical | Dissolved | LAB
HG-D | mg/L | Mercury Dissolved | X548 | BrCldig, SnCl2red'n,CVAF | HNO3 preserved water samples are dispensed into pre-cleaned glass test tubes. Bromine monochloride (BrCl, a strong oxidant) is added to each sample to oxidize all mercury species (including organo-mercurais) to the Hg2+ state. The samples are left to cold digest for approx. 45 min. Upon meeting the required digestion time the BrCl is quenched with Hydroxylamine hydrochloride (NH2OH-HCl, a weak reductant). These samples are allowed to sit again for approx. 45 min. in order for the BrCl to be completely consumed by the hydroxylamine. The Hg in the samples is reduced with stannous chloride dihydrate (SnCl2-2H2O, a strong reductant) to its elemental state and instrumental analysis is performed using a cold vapor atomic fluoresence spectrophotometer. | 5e-06 | Y | Y | Y | 16754 | 10/27/2004 | 8/22/2024 | 8/26/2024 3:51:07 PM | 2002-12-16 | 8/15/2019 1:25:39 PM | EMS | KATFRASE | Concentration | mg/L | Y | (fl. conc.) |  | Mercury Dissolved (fl. conc.) |  | Metals - Dissolved; | Metals - Dissolved; EMS code: HG-D | Chemical | Dissolved | LAB
HG-D | mg/L | Mercury Dissolved | CVAF | Hg CVAF | Mercury in Water, Leachates and Liquids by Bromination and Cold Vapour Atomic Fluorescence Spectroscopy PURPOSE/PRINCIPLE OF METHOD: Chemically generated bromine is used to oxidize all organic and inorganic forms of mercury in aqueous samples to mercury (II). Excess bromine is neutralized with hydroxylamine hydrochloride. The solution is introduced into a gas liquid separator and mixed with stannous chloride to reduce the mercury to its elemental form. The mercury is then purged from the solution with argon gas and passes through an atomic fluorescence detector where it absorbs light from a mercury vapour lamp that fluoresces at 254 nm. This is detected and amplified by a photo-multiplier tube in the detector. SCOPE: The method applies to the determination of field filtered, dissolved and total mercury in water, wastewater, effluent, as well as hot acid digested liquids and aqueous waste leachates generated by Leachate Extraction Procedure, Toxicity Characteristic Leaching Procedure (TCLP), Synthetic Precipitation Leaching Procedure (SPLP) and Simultaneously Extractable Metals (SEM) extracts. DETECTION LIMITS: Routine RDL = 0.02 ug/L Ultra-Low level RDL = 0.002 ug/L REFERENCE: Determination of Mercury in Water by Semi-Automated Cold Vapor Atomic Fluorescence Spectrometry. EPA Method 245.7, Revision 2, February 2005 | 1e-05 | Y | Y | Y | 11427 | 7/8/2011 10:09:00 AM | 8/20/2024 3:34:00 PM | 8/23/2024 5:01:25 PM | 2012-02-03 | 2/3/2012 9:37:02 AM | RDGIBSON | RDGIBSON | Concentration | mg/L | Y | (fl. conc.) |  | Mercury Dissolved (fl. conc.) |  | Metals - Dissolved; | Metals - Dissolved; EMS code: HG-D | Chemical | Dissolved | LAB
HG-D | mg/L | Mercury Dissolved | X507 | BrCl cold oxidn,SnCl2redn,CVAAS | BrCl is used in a cold oxidation procedure to oxidize mercury, including organo-mercury compounds. Note that there is no type of digestion prior to the BrCl oxidation. Also, the organo-mercury compounds may not fully oxidize. The oxidized mercury is then reduced to mercury vapour using stannous chloride, and the mercury analyzed by CVAAS at 253.7 nm. | 0.0001 | Y | Y | Y | 9081 | 3/10/1998 2:07:00 PM | 10/19/2023 | 8/20/2024 7:35:41 AM | 1999-02-18 |  | NPEPPIN |  | Concentration | mg/L | Y | (fl. conc.) |  | Mercury Dissolved (fl. conc.) |  | Metals - Dissolved; | Metals - Dissolved; EMS code: HG-D | Chemical | Dissolved | LAB
HG-E | mg/L | Mercury Extractable | X177 | Extr,Auto An,Flamless AA | Extr;Auto An;Flamless AA | 0 | Y | Y | Y | 13 |  |  | 11/25/2021 1:50:15 PM | 1996-04-12 | 6/3/1998 6:51:36 PM | EMS |  | Concentration | mg/L | Y | (fl. conc.) |  | Mercury Extractable (fl. conc.) |  | Metals - Extractable; | Metals - Extractable; EMS code: HG-E | Chemical | Extractable | LAB
HG-E | ug/g | Mercury Extractable | X229 | Hom,HNO3/HCl,Cold V,UV | Hom;HNO3/HCl;Cold V;UV | 0.025 | Y | Y | Y | 70 |  |  | 10/24/1996 2:25:27 PM | 1996-04-12 | 6/3/1998 6:51:42 PM | EMS |  | MassPerMass | mg/kg | Y | (mass conc.) |  | Mercury Extractable (mass conc.) |  | Metals - Extractable; | Metals - Extractable; EMS code: HG-E | Chemical | Extractable | LAB
HG-E | umol/g | Mercury Extractable | FSEM | AVS extraction, SEM Analysis ICPMS | Determination of simultaneous extractable metals (SEM) in sediment after sample extracted with hydrochloric acid for acid volatile sulfide (AVS) | 0.02 | Y | Y | Y | 19 | 3/13/2020 | 3/25/2020 | 2/22/2023 12:51:03 PM | 2010-11-16 |  | RDGIBSON |  | ParticlesPerMass | µmol/g | Y | (part. /m) |  | Mercury Extractable (part. /m) |  | Metals - Extractable; | Metals - Extractable; EMS code: HG-E | Chemical | Extractable | LAB
HG-E | mg/L | Mercury Extractable | X510 | H2SO4/ Pot.DichromPres,H2SO4,flameless AA | H2So4 and pot. dichr preservation;overnight; H2SO4 digestion (overnight);flameless AA | 5e-05 | Y | Y | Y | 774 |  |  | 11/25/2021 1:50:15 PM | 1999-04-22 |  | NPEPPIN |  | Concentration | mg/L | Y | (fl. conc.) |  | Mercury Extractable (fl. conc.) |  | Metals - Extractable; | Metals - Extractable; EMS code: HG-E | Chemical | Extractable | LAB
HG-I | mg/dm2/d | Mercury-Insoluble | EQ15 | EQUIS (mg/dm2/d) | EQUIS (mg/dm2/d) | 0 | N | Y | Y | 28 |  |  | 10/22/1996 10:18:38 PM | 1996-04-12 | 6/3/1998 6:51:41 PM | EMS |  | MassFlux | mg/dm²/d | Y | (mass flux) |  | Mercury-Insoluble (mass flux) |  | Metals - Insoluble; | Metals - Insoluble; EMS code: HG-I | Chemical | Insoluble | LAB
HG-L | mg/L | Leachable Mercury after EPA TCLP extraction of solid sample | 1313 | EPA 1311/245.7 (CAVF) | This analysis is carried out in accordance with the extraction procedure outlined in "Test Methods for Evaluating Solid Waste - Physical/Chemical Methods Volume 1C" SW-846 EPA Method 1311, published by the United States Environmental Protection Agency (EPA). In summary, the sample is extracted at a 20:1 liquid to solids ratio for 16 to 20 hours using either extraction fluid #1 (glacial acetic acid, water and sodium hydroxide) or extraction fluid #2 (glacial acetic acid), depending on the pH of the original sample. The extract is then filtered through a 0.6 to 0.8 micron glass fibre filter and analysed using atomic fluorescence spectrophotometry (EPA 245.7). | 0.001 | Y | Y | Y | 98 |  |  | 10/24/2022 10:58:25 AM | 2011-04-27 |  | RDGIBSON |  | Concentration | mg/L | Y | (fl. conc.) |  | Leachable Mercury after EPA TCLP extraction of solid sample (fl. conc.) |  | Metals - Extractable; | Metals - Extractable; EMS code: HG-L | Chemical | Leachable | LAB
HG-L | mg/L | Leachable Mercury after EPA TCLP extraction of solid sample | E1631E | EPA 1631E | See EPA method 1631E | 5e-05 | Y | Y | Y | 0 |  |  |  | 2024-01-29 |  | RDGIBSON |  | Concentration | mg/L | Y | (fl. conc.) |  | Leachable Mercury after EPA TCLP extraction of solid sample (fl. conc.) |  | Metals - Extractable; | Metals - Extractable; EMS code: HG-L | Chemical | Leachable | LAB
HG-S | mg/dm2/d | Mercury-Soluble | EQ15 | EQUIS (mg/dm2/d) | EQUIS (mg/dm2/d) | 0 | N | Y | Y | 24 |  |  | 10/22/1996 10:18:25 PM | 1996-04-12 | 6/3/1998 6:51:36 PM | EMS |  | MassFlux | mg/dm²/d | Y | (mass flux) |  | Mercury-Soluble (mass flux) |  | Metals - Soluble; | Metals - Soluble; EMS code: HG-S | Chemical | Soluble | LAB
HG-S | umol/g | Mercury-Soluble | FSEM | AVS extraction, SEM Analysis ICPMS | Determination of simultaneous extractable metals (SEM) in sediment after sample extracted with hydrochloric acid for acid volatile sulfide (AVS) | 0.02 | Y | Y | Y | 0 |  |  |  | 2024-02-02 |  | EMS |  | ParticlesPerMass | µmol/g | Y | (part. /m) |  | Mercury-Soluble (part. /m) |  | Metals - Soluble; | Metals - Soluble; EMS code: HG-S | Chemical | Soluble | LAB
HG-T | ug/g | Mercury Total | FB08 | Freeze Dry,Hom,Acid Dig,CVAF | Sample moisture is determined by weighing samples before and after freeze-drying. freeze-dried samples are ground in a Braun blender before being weighed into Teflon PFA liners and digested with concentrated HNO3, concentrated HCI, and 50% H2O2 in a microwave oven; which results in a complete sample dissolution. Samples are collected and volumized with Type1 water. Analysis is performed via Inductively Coupled Argon Plasma Atomic Emission Spectrometry (ICP-AES) (simultaneous multielement analysis), Graphite Furnace Atomic Absorption Spectometry (GF-AAS) (low leve Cd and Pb analysis), Cold Vapor Atomic Fluorescence Spectrometry (CV-AFS)(Mercury analysis) and/or Flame Atomic Absorption Spectrometry (F-AAS) (K analysis). | 0.004 | Y | Y | Y | 117 | 8/16/1999 12:38:00 PM | 1/31/2002 2:35:00 PM | 2/24/2012 11:58:36 AM | 1997-02-27 | 6/3/1998 6:51:46 PM | MPAK |  | MassPerMass | mg/kg | Y | (mass conc.) |  | Mercury Total (mass conc.) |  | Metals - Total; | Metals - Total; EMS code: HG-T | Chemical | Total | LAB
HG-T | ug/g | Mercury Total | CVAAS | Hg in Soil by CVAAS | This analysis is carried out using procedures from CSR Analytical Method: "Strong Acid Leachable Metals (SALM) in Soil", BC Ministry of Environment, 26 June 2009, and procedures adapted from EPA Method 200.2. The sample is manually homogenized, dried at 60 degrees Celsius, sieved through a 63 um (230 mesh) sieve, and a representative subsample of the dry material is weighed. The sample is then digested at 95 degrees Celsius for 2 hours by block digester using concentrated nitric and hydrochloric acids. Instrumental analysis is by atomic absorption spectrophotometry (EPA Method 245.7). Method Limitation: This method is not a total digestion technique. It is a very strong acid digestion that is intended to dissolve those metals that may be environmentally available. By design, elements bound in silicate structures are not normally dissolved by this procedure as they are not usually mobile in the environment. Deviation from Reference Method: This procedure deviates from the BC CSR SALM method, which specifies sieving to 2 mm (10 mesh). | 0.005 | Y | Y | Y | 23 | 9/19/2017 | 10/19/2018 | 10/22/2018 3:16:04 PM | 2018-01-09 |  | JKROGH |  | MassPerMass | mg/kg | Y | (mass conc.) |  | Mercury Total (mass conc.) |  | Metals - Total; | Metals - Total; EMS code: HG-T | Chemical | Total | LAB
HG-T | ug/g wet | Mercury Total | EQ44 | EQUIS (ug/g wet) | EQUIS (ug/g wet) | 0.5 | N | Y | Y | 388 |  |  | 10/23/1996 12:08:26 AM | 1996-04-12 | 6/3/1998 6:51:36 PM | EMS |  | MassPerMass | µg/g | Y | (mass conc.) | (wet) | Mercury Total (wet) (mass conc.) |  | Metals - Total; | Metals - Total; EMS code: HG-T | Chemical | Total | LAB
HG-T | ug/g | Mercury Total | FS82 | PBM,dry60C,10mesh,SALM-ICP-MS | PBM;dry60C;10mesh;SALM-ICP-MS | 0.05 | Y | Y | Y | 1973 | 4/10/2002 9:30:00 AM | 3/16/2024 6:06:00 PM | 7/30/2024 10:00:14 AM | 2002-11-12 |  | UMULLER |  | MassPerMass | mg/kg | Y | (mass conc.) |  | Mercury Total (mass conc.) |  | Metals - Total; | Metals - Total; EMS code: HG-T | Chemical | Total | LAB
HG-T | mg/L | Mercury Total | X346 | FA:HNO3/K2Cr2O7:CVAA | The unfiltered field-acidified water sample undergoes a strongly oxidizing digestion. The sample is then analyzed by CVAAS. | 5e-05 | Y | Y | Y | 339 |  |  | 11/20/2006 3:40:13 PM | 1996-04-12 | 6/3/1998 6:51:41 PM | EMS |  | Concentration | mg/L | Y | (fl. conc.) |  | Mercury Total (fl. conc.) |  | Metals - Total; | Metals - Total; EMS code: HG-T | Chemical | Total | LAB
HG-T | mg/L | Mercury Total | X256 | HNO3/H2O2Dig:Cold V.:UV | HNO3/H2O2Dig:Cold V.:UV | 0.001 | Y | Y | Y | 4 |  |  | 6/7/2024 5:30:18 PM | 1996-04-12 | 6/3/1998 6:51:41 PM | EMS |  | Concentration | mg/L | Y | (fl. conc.) |  | Mercury Total (fl. conc.) |  | Metals - Total; | Metals - Total; EMS code: HG-T | Chemical | Total | LAB
HG-T | ug/g | Mercury Total | X246 | Hom:HNO3/HCl:CVAA | Soil or sediment samples are homogenized and sub-samples taken. Without a drying step, samples are digested with Aqua Regia either over a hot plate or in a microwave. Solubidized mercury is then analyzed by CVAAS. A separate sub-sample is analyzed for moisture and the mercury result corrected to dry value. | 0.005 | Y | Y | Y | 81 | 2/1/2002 4:15:00 PM | 10/12/2023 | 11/28/2023 2:21:33 PM | 1996-04-12 | 4/1/1999 4:04:53 PM | EMS | NPEPPIN | MassPerMass | mg/kg | Y | (mass conc.) |  | Mercury Total (mass conc.) |  | Metals - Total; | Metals - Total; EMS code: HG-T | Chemical | Total | LAB
HG-T | mg/L | Mercury Total | F033 | Acid Dig,Centrifuge,ICPMS | For the determination of total recoverable elements in water or waste water, take 100 ml aliquot from a well mixed, acid preserved, sample containing not more than 0.25% (w/v) total solids and transfer to a 250 ml Griffin beaker. (If total solids are greater than 0.25% reduce the size of the aliquot by a proportionate amount). Add 1 ml of conc. nitric acid and 0.5 ml conc. hydrochloric acid. Heat on a hot plate until the volume has been reduced to 15-20ml, ensuring that the sample does not boil. Allow to cool and transfer to a 50 ml volumetric flask. Dilute to volume with ASTM type I water. Centrifuge the sample or allow to stand overnight to separate insoluble material. Prior analysis, pipette 20 ml into a 50 ml volumetric flask and dilute to volume with ASTM type I water. Analysis should be performed within seventy two hours of the completed preparation. Note that different laboratories may have very different field preservation, digestion procedures and detection limits for this method, so these should be documented. | 1e-05 | Y | Y | Y | 1220 | 5/6/2008 4:46:00 PM | 8/20/2024 3:11:00 AM | 8/22/2024 5:01:03 PM | 2008-06-09 | 7/16/2012 3:35:33 PM | RDGIBSON | RDGIBSON | Concentration | mg/L | Y | (fl. conc.) |  | Mercury Total (fl. conc.) |  | Metals - Total; | Metals - Total; EMS code: HG-T | Chemical | Total | LAB
HG-T | mg/L | Mercury Total | F021 | Acid Dig,Auto,CV Fluo.Spec. | Mercury in organic and inorganic compounds is oxidized to mercuric ions by addition of sulfuric acid, nitric acid, potassium persulfate and heat. The mercuric ions are reduced to Hg(0) with SnCl2, are sparged from the solution with argon, and passed through a fluorescence cell in the light path of a mercury lamp. The fluorescence at the 253.7 nm mercury emission line by Hg(0) is proportional to the concentration of mercury in the sample. | 5e-05 | Y | Y | Y | 295 | 8/30/1996 11:36:00 AM | 11/26/2015 8:00:00 AM | 9/11/2019 8:46:38 AM | 1996-09-17 | 6/3/1998 6:51:45 PM | MPAK |  | Concentration | mg/L | Y | (fl. conc.) |  | Mercury Total (fl. conc.) |  | Metals - Total; | Metals - Total; EMS code: HG-T | Chemical | Total | LAB
HG-T | ug/g | Mercury Total | MSCS | HG-200.2-CVAF-VA (DRY) | Mercury in Soil by CVAFS (DRY) | 0.005 | Y | Y | Y | 127 | 5/19/2015 | 7/29/2024 | 7/29/2024 5:17:09 PM | 2015-06-03 |  | RDGIBSON |  | MassPerMass | mg/kg | Y | (mass conc.) |  | Mercury Total (mass conc.) |  | Metals - Total; | Metals - Total; EMS code: HG-T | Chemical | Total | LAB
HG-T | mg/L | Mercury Total | X111 | HCl/HNO3Dig,Flameless AA | HCl/HNO3Dig;Flameless AA | 0 | Y | Y | Y | 4542 |  |  | 7/21/1998 10:56:36 AM | 1996-04-12 | 6/3/1998 6:51:36 PM | EMS |  | Concentration | mg/L | Y | (fl. conc.) |  | Mercury Total (fl. conc.) |  | Metals - Total; | Metals - Total; EMS code: HG-T | Chemical | Total | LAB
HG-T | ug/m3 | Mercury Total | X197 | HCL/HNO3,Man C.V.-Air | HCL/HNO3;Man C.V.-Air | 0 | Y | Y | Y | 920 |  |  | 4/12/2024 1:44:40 PM | 1996-04-12 | 6/3/1998 6:51:36 PM | EMS |  | AirConcentration | µg/m³ | Y | (air conc.) |  | Mercury Total (air conc.) |  | Metals - Total; | Metals - Total; EMS code: HG-T | Chemical | Total | LAB
HG-T | ppm (S) | Mercury Total | MS41 | ALS Minerals ME-MS41 | ALS Minerals ME-MS41. A prepared sample is digested with aqua regia in a graphite heating block. After cooling; the resulting solution is diluted to with deionized water; mixed and analyzed by ICP-AES. Samples are then analyzed by ICP-MS for the remaining suite of elements. | 0.01 | Y | Y | Y | 0 |  |  |  | 2024-01-29 |  | EMS |  | DimensionlessRatio | ppm | N | (ratio) |  | Mercury Total (ratio) |  | Metals - Total; | Metals - Total; EMS code: HG-T | Chemical | Total | LAB
HG-T | mg/L | Mercury Total | EQ01 | EQUIS (mg/L) | EQUIS (mg/L) | 0 | N | Y | Y | 6518 |  |  | 10/24/1996 3:16:55 PM | 1996-04-12 | 6/3/1998 6:51:41 PM | EMS |  | Concentration | mg/L | Y | (fl. conc.) |  | Mercury Total (fl. conc.) |  | Metals - Total; | Metals - Total; EMS code: HG-T | Chemical | Total | LAB
HG-T | mg/L | Mercury Total | X169 | H2SO4:K2Cr2O7:Cold V:AA | H2SO4:K2Cr2O7:Cold V:AA | 0 | Y | Y | Y | 69 |  |  | 1/30/1998 8:32:01 AM | 1996-04-12 | 6/3/1998 6:51:36 PM | EMS |  | Concentration | mg/L | Y | (fl. conc.) |  | Mercury Total (fl. conc.) |  | Metals - Total; | Metals - Total; EMS code: HG-T | Chemical | Total | LAB
HG-T | mg | Mercury Total | 1630 | US EPA 1630 in Water | Method 1630 of the US EPA. Samples are distilled under an inert gas flow to isolate methylmercury and minimize matrix interferences. The distillate is analyzed by aqueous phase ethylation, purge and trap, desorption and GC separation. THe separated species are then pyrolized to elemental Hg and quantified by cold vapour atomic flourescence spectroscopy. | 2.5e-05 | Y | Y | Y | 0 |  |  |  | 2024-04-29 |  | RDGIBSON |  | Mass | mg | Y | (mass) |  | Mercury Total (mass) |  | Metals - Total; | Metals - Total; EMS code: HG-T | Chemical | Total | LAB
HG-T | mg/L | Mercury Total | X201 | HNO3/KMnO4 Dig,Cold V:AA | HNO3/KMnO4 Dig;Cold V:AA | 5e-05 | Y | Y | Y | 1354 | 5/3/2002 | 5/1/2008 2:33:00 PM | 5/13/2008 1:24:04 PM | 1996-04-12 | 6/3/1998 6:51:36 PM | EMS |  | Concentration | mg/L | Y | (fl. conc.) |  | Mercury Total (fl. conc.) |  | Metals - Total; | Metals - Total; EMS code: HG-T | Chemical | Total | LAB
HG-T | mg/L | Mercury Total | X451 | Hom,HNO3/HCl,Cold V,UV | Hom;HNO3/HCl;Cold V;UV | 0.005 | Y | Y | Y | 210 |  |  | 9/27/2001 9:54:54 AM | 1996-04-12 | 6/3/1998 6:51:36 PM | EMS |  | Concentration | mg/L | Y | (fl. conc.) |  | Mercury Total (fl. conc.) |  | Metals - Total; | Metals - Total; EMS code: HG-T | Chemical | Total | LAB
HG-T | ug/g | Mercury Total | E200 | EPA 200.2 | See EPA 200.2/1631 Appendix | 0.005 | Y | Y | Y | 80 | 7/14/2022 | 9/22/2023 | 1/12/2024 11:16:52 AM | 2022-07-27 |  | EMS |  | MassPerMass | mg/kg | Y | (mass conc.) |  | Mercury Total (mass conc.) |  | Metals - Total; | Metals - Total; EMS code: HG-T | Chemical | Total | LAB
HG-T | ug/m3 | Mercury Total | 5091 | KMnO4,Man Cold Vapour,UV | KMnO4;Man Cold Vapour;UV | 0.01 | Y | Y | Y | 629 |  |  | 10/23/1996 4:34:24 PM | 1996-04-12 | 6/3/1998 6:51:36 PM | EMS |  | AirConcentration | µg/m³ | Y | (air conc.) |  | Mercury Total (air conc.) |  | Metals - Total; | Metals - Total; EMS code: HG-T | Chemical | Total | LAB
HG-T | ng/L | Mercury Total | E1631E | EPA 1631E | See EPA method 1631E | 0.5 | Y | Y | Y | 501 | 7/24/2020 | 7/22/2024 | 7/24/2024 12:17:06 PM | 2022-06-28 |  | EMS |  | Concentration | ng/L | Y | (fl. conc.) |  | Mercury Total (fl. conc.) |  | Metals - Total; | Metals - Total; EMS code: HG-T | Chemical | Total | LAB
HG-T | mg/L | Mercury Total | 1348 | Total Mercury in Seawater by CVAFS | Total Mercury in Seawater by CVAFS | 5e-06 | Y | Y | Y | 118 | 8/17/2019 | 4/26/2024 | 6/7/2024 10:03:16 AM | 2019-09-30 |  | KATFRASE |  | Concentration | mg/L | Y | (fl. conc.) |  | Mercury Total (fl. conc.) |  | Metals - Total; | Metals - Total; EMS code: HG-T | Chemical | Total | LAB
HG-T | ppm (S) | Mercury Total | 6041 | Aqua Regia by ICPMS | These were sediment/soils digested with Aqua-Regia (Nitric Acid & Hydrochloric Acids) and then analysis by ICPMS. Results would be reported in mg/kg and/or %. | 0.004 | Y | Y | Y | 21 | 10/13/2019 | 8/5/2020 | 8/10/2020 3:52:06 PM | 2019-09-06 | 10/21/2019 11:55:15 AM | KATFRASE | KATFRASE | DimensionlessRatio | ppm | N | (ratio) |  | Mercury Total (ratio) |  | Metals - Total; | Metals - Total; EMS code: HG-T | Chemical | Total | LAB
HG-T | ug/g | Mercury Total | X229 | Hom,HNO3/HCl,Cold V,UV | Hom;HNO3/HCl;Cold V;UV | 0.005 | Y | Y | Y | 312 | 12/30/2005 | 9/29/2010 9:05:00 AM | 10/6/2010 10:26:58 AM | 1996-04-12 | 6/3/1998 6:51:42 PM | EMS |  | MassPerMass | mg/kg | Y | (mass conc.) |  | Mercury Total (mass conc.) |  | Metals - Total; | Metals - Total; EMS code: HG-T | Chemical | Total | LAB
HG-T | ug/g | Mercury Total | CVAL | Hg CVAF low level | Mercury in Water, Leachates and Liquids by Bromination and Cold Vapour Atomic Fluorescence Spectroscopy (low level) PURPOSE/PRINCIPLE OF METHOD: Chemically generated bromine is used to oxidize all organic and inorganic forms of mercury in aqueous samples to mercury (II). Excess bromine is neutralized with hydroxylamine hydrochloride. The solution is introduced into a gas liquid separator and mixed with stannous chloride to reduce the mercury to its elemental form. The mercury is then purged from the solution with argon gas and passes through an atomic fluorescence detector where it absorbs light from a mercury vapour lamp that fluoresces at 254 nm. This is detected and amplified by a photo-multiplier tube in the detector. SCOPE: The method applies to the determination of field filtered, dissolved and total mercury in water, wastewater, effluent, as well as hot acid digested liquids and aqueous waste leachates generated by Leachate Extraction Procedure, Toxicity Characteristic Leaching Procedure (TCLP), Synthetic Precipitation Leaching Procedure (SPLP) and Simultaneously Extractable Metals (SEM) extracts. DETECTION LIMITS: Routine RDL = 0.02 ug/L Ultra-Low level RDL = 0.002 ug/L REFERENCE: Determination of Mercury in Water by Semi-Automated Cold Vapor Atomic Fluorescence Spectrometry. EPA Method 245.7, Revision 2, February 2005 | 2e-06 | Y | Y | Y | 3255 | 7/12/2011 8:14:00 AM | 10/1/2018 8:57:00 AM | 2/16/2018 6:30:18 PM | 2012-02-03 |  | RDGIBSON |  | MassPerMass | mg/kg | Y | (mass conc.) |  | Mercury Total (mass conc.) |  | Metals - Total; | Metals - Total; EMS code: HG-T | Chemical | Total | LAB
HG-T | mg/L | Mercury Total | X295 | KMnO4/H2SO4,Cold Vap,AA | KMnO4/H2SO4;Cold Vap;AA | 0.005 | Y | Y | Y | 1030 | 5/9/1998 | 10/14/2003 | 11/18/2003 1:33:38 PM | 1996-04-12 | 6/3/1998 6:51:36 PM | EMS |  | Concentration | mg/L | Y | (fl. conc.) |  | Mercury Total (fl. conc.) |  | Metals - Total; | Metals - Total; EMS code: HG-T | Chemical | Total | LAB
HG-T | ug/g | Mercury Total | 0330 | Air D&G:N/S Dig:Man.CV-P | Vegetation samples are air dried, ground, and digested with H2SO4/HNO3. The organomercury compounds are oxidized via a H2SO4/KMnO4/K2S2O8 digestion. The mercury compounds are then reduced to elemental mercury. A stream of synthetic air introduced into the system passes through a reference cell and then, after sparging the elemental mercury from the solution, passes through the sample cell. A UV monitor possessing a mercury lamp provides the light source and the absorbance at 253.7 nm is measured against the reference. | 0.05 | Y | Y | Y | 36 |  |  | 10/24/1996 2:08:12 PM | 1996-04-12 | 6/3/1998 6:51:36 PM | EMS |  | MassPerMass | mg/kg | Y | (mass conc.) |  | Mercury Total (mass conc.) |  | Metals - Total; | Metals - Total; EMS code: HG-T | Chemical | Total | LAB
HG-T | mg/L | Mercury Total | 0310 | HNO3/H2SO4 Dig:Man.C.V. | Samples are field filtered for dissolved or not filtered for total, and preserved with K2Cr2O7/H2SO4 to pH <2. The organomercury compounds are oxidized via a H2SO4/KMnO4/K2S2O8 digestion. The mercury compounds are then reduced to elemental mercury. A stream of synthetic air introduced into the system passes through a reference cell and then, after sparging the elemental mercury from the solution, passes through the sample cell. A UV monitor possessing a mercury lamp provides the light source and the absorbance at 253.7 nm is measured against the reference. | 5e-05 | Y | Y | Y | 1769 | 12/2/1996 | 4/8/2016 12:19:00 PM | 10/31/2016 5:31:01 PM | 1996-04-12 | 6/3/1998 6:51:36 PM | EMS |  | Concentration | mg/L | Y | (fl. conc.) |  | Mercury Total (fl. conc.) |  | Metals - Total; | Metals - Total; EMS code: HG-T | Chemical | Total | LAB
HG-T | ug/g | Mercury Total | X273 | Cold Vapour:AA | Cold Vapour:AA | 0.005 | Y | Y | Y | 106 | 1/24/2001 | 4/16/2001 | 7/27/2001 11:17:47 AM | 1996-04-12 | 6/3/1998 6:51:36 PM | EMS |  | MassPerMass | mg/kg | Y | (mass conc.) |  | Mercury Total (mass conc.) |  | Metals - Total; | Metals - Total; EMS code: HG-T | Chemical | Total | LAB
HG-T | ppm (S) | Mercury Total | 6042 | Aqua Regia by ICPMS | These were sediment/soils digested with Aqua-Regia (Nitric Acid & Hydrochloric Acids) and then analysis by ICPMS. Results would be reported in ppm s. | 0.004 | Y | Y | Y | 71 | 5/14/2022 | 7/27/2024 | 7/29/2024 9:13:08 AM | 2022-06-28 |  | EMS |  | DimensionlessRatio | ppm | N | (ratio) |  | Mercury Total (ratio) |  | Metals - Total; | Metals - Total; EMS code: HG-T | Chemical | Total | LAB
HG-T | mg/L | Mercury Total | CVAF | Hg CVAF | Mercury in Water, Leachates and Liquids by Bromination and Cold Vapour Atomic Fluorescence Spectroscopy PURPOSE/PRINCIPLE OF METHOD: Chemically generated bromine is used to oxidize all organic and inorganic forms of mercury in aqueous samples to mercury (II). Excess bromine is neutralized with hydroxylamine hydrochloride. The solution is introduced into a gas liquid separator and mixed with stannous chloride to reduce the mercury to its elemental form. The mercury is then purged from the solution with argon gas and passes through an atomic fluorescence detector where it absorbs light from a mercury vapour lamp that fluoresces at 254 nm. This is detected and amplified by a photo-multiplier tube in the detector. SCOPE: The method applies to the determination of field filtered, dissolved and total mercury in water, wastewater, effluent, as well as hot acid digested liquids and aqueous waste leachates generated by Leachate Extraction Procedure, Toxicity Characteristic Leaching Procedure (TCLP), Synthetic Precipitation Leaching Procedure (SPLP) and Simultaneously Extractable Metals (SEM) extracts. DETECTION LIMITS: Routine RDL = 0.02 ug/L Ultra-Low level RDL = 0.002 ug/L REFERENCE: Determination of Mercury in Water by Semi-Automated Cold Vapor Atomic Fluorescence Spectrometry. EPA Method 245.7, Revision 2, February 2005 | 1e-05 | Y | Y | Y | 10175 | 11/11/2011 8:54:00 AM | 8/22/2024 3:13:00 PM | 8/26/2024 8:06:03 AM | 2012-02-03 |  | RDGIBSON |  | Concentration | mg/L | Y | (fl. conc.) |  | Mercury Total (fl. conc.) |  | Metals - Total; | Metals - Total; EMS code: HG-T | Chemical | Total | LAB
HG-T | mg/L | Mercury Total | X521 | KMNO4,K2S208,SnCl2,CVAA | KMNO4;K2S2O8;H2SO$@105C;HAH reduction;SnCl2;CVAA | 5e-06 | Y | Y | Y | 244 |  |  | 9/9/1999 12:07:49 PM | 1999-08-18 |  | NPEPPIN |  | Concentration | mg/L | Y | (fl. conc.) |  | Mercury Total (fl. conc.) |  | Metals - Total; | Metals - Total; EMS code: HG-T | Chemical | Total | LAB
HG-T | ug/g wet | Mercury Total | X560 | Metals in Biota (wet), total, acid digestion, ICP-MS | Five grams of wet, homogenized biota is weighed in a centrifuge tube, and 2.5 mL of Nitric Acid and 2.5 mL of Hydrochloric acid are added. The sample is digested at room temperature for one hour and then digested at 90-95 degrees C in a water bath or heat block for one more hour. The sample is cooled and the volume is made to 50 mL with DI. The digestate is then diluted at ten fold and introduced to a collision/reaction cell inductively coupled argon plasma mass spectrometer (CRC ICPMS) by pneumatic nebulization. Each respective element is measured at a specific mass to charge ratio, m/z value expressed in amu (atomic mass units). - Maxxam Analytics (BC). Added to EMS Dec 10, 2008. | 0.01 | Y | Y | Y | 196 | 2/3/2008 | 10/12/2023 | 11/28/2023 2:21:33 PM | 2008-12-11 | 3/11/2015 4:17:10 PM | RDGIBSON | EMS | MassPerMass | µg/g | Y | (mass conc.) | (wet) | Mercury Total (wet) (mass conc.) |  | Metals - Total; | Metals - Total; EMS code: HG-T | Chemical | Total | LAB
HG-T | mg/L | Mercury Total | X548 | BrCldig, SnCl2red'n,CVAF | HNO3 preserved water samples are dispensed into pre-cleaned glass test tubes. Bromine monochloride (BrCl, a strong oxidant) is added to each sample to oxidize all mercury species (including organo-mercurais) to the Hg2+ state. The samples are left to cold digest for approx. 45 min. Upon meeting the required digestion time the BrCl is quenched with Hydroxylamine hydrochloride (NH2OH-HCl, a weak reductant). These samples are allowed to sit again for approx. 45 min. in order for the BrCl to be completely consumed by the hydroxylamine. The Hg in the samples is reduced with stannous chloride dihydrate (SnCl2-2H2O, a strong reductant) to its elemental state and instrumental analysis is performed using a cold vapor atomic fluoresence spectrophotometer. | 5e-08 | Y | Y | Y | 17774 | 3/25/2004 | 8/22/2024 | 8/26/2024 3:51:07 PM | 2002-12-16 | 6/2/2022 4:45:59 PM | EMS | EMS | Concentration | mg/L | Y | (fl. conc.) |  | Mercury Total (fl. conc.) |  | Metals - Total; | Metals - Total; EMS code: HG-T | Chemical | Total | LAB
HG-T | ug/g wet | Mercury Total | MTCW | HG-WET-CVAFS-VA | Mercury in Tissue by CVAFS (WET) | 0.001 | Y | Y | Y | 171 | 8/13/2014 | 3/8/2023 | 3/23/2023 4:12:02 PM | 2015-03-09 | 3/26/2015 11:24:47 AM | RDGIBSON | RDGIBSON | MassPerMass | µg/g | Y | (mass conc.) | (wet) | Mercury Total (wet) (mass conc.) |  | Metals - Total; | Metals - Total; EMS code: HG-T | Chemical | Total | LAB
HG-T | ug/m3 | Mercury Total | EQ40 | EQUIS (mg/m3) | EQUIS (mg/m3) | 0.0001 | N | Y | Y | 2916 |  |  | 10/23/1996 12:01:28 AM | 1996-04-12 | 6/3/1998 6:51:41 PM | EMS |  | AirConcentration | µg/m³ | Y | (air conc.) |  | Mercury Total (air conc.) |  | Metals - Total; | Metals - Total; EMS code: HG-T | Chemical | Total | LAB
HG-T | mg/L | Mercury Total | 0300 | Autoanal/Cold Vap. | Autoanal/Cold Vap. | 5e-05 | Y | Y | Y | 17572 | 4/19/2007 | 12/1/2010 | 8/23/2024 8:10:35 AM | 1996-04-12 | 6/3/1998 6:51:36 PM | EMS |  | Concentration | mg/L | Y | (fl. conc.) |  | Mercury Total (fl. conc.) |  | Metals - Total; | Metals - Total; EMS code: HG-T | Chemical | Total | LAB
HG-T | mg/L | Mercury Total | X507 | BrCl cold oxidn,SnCl2redn,CVAAS | BrCl is used in a cold oxidation procedure to oxidize mercury, including organo-mercury compounds. Note that there is no type of digestion prior to the BrCl oxidation. Also, the organo-mercury compounds may not fully oxidize. The oxidized mercury is then reduced to mercury vapour using stannous chloride, and the mercury analyzed by CVAAS at 253.7 nm. | 5e-05 | Y | Y | Y | 8040 | 3/10/1998 1:58:00 PM | 10/5/2023 | 7/30/2024 12:35:06 PM | 1999-02-18 |  | NPEPPIN |  | Concentration | mg/L | Y | (fl. conc.) |  | Mercury Total (fl. conc.) |  | Metals - Total; | Metals - Total; EMS code: HG-T | Chemical | Total | LAB
HG-T | mg/L | Mercury Total | TCLP | Toxicity Characteristic Leaching Procedure | The TCLP procedure is designed to simulate the climatic leaching action | 0.002 | Y | Y | Y | 22 | 12/8/2009 7:49:00 PM | 9/28/2011 5:53:00 PM | 11/3/2011 11:10:19 PM | 2009-03-06 |  | JCHUDAK |  | Concentration | mg/L | Y | (fl. conc.) |  | Mercury Total (fl. conc.) |  | Metals - Total; | Metals - Total; EMS code: HG-T | Chemical | Total | LAB
HG-T | ug/g | Mercury Total | 0320 | Air D&G:A.R.Dig:Man.CV-S | Soil and sediment samples are air dried, ground, and digested with aqua regia. The organomercury compounds are oxidized via a H2SO4/KMnO4/K2S2O8 digestion. The mercury compounds are then reduced to elemental mercury. A stream of synthetic air introduced into the system passes through a reference cell and then, after sparging the elemental mercury from the solution, passes through the sample cell. A UV monitor possessing a mercury lamp provides the light source and the absorbance at 253.7 nm is measured against the reference. | 0.05 | Y | Y | Y | 344 |  |  | 10/24/1996 3:04:18 PM | 1996-04-12 | 6/3/1998 6:51:41 PM | EMS |  | MassPerMass | mg/kg | Y | (mass conc.) |  | Mercury Total (mass conc.) |  | Metals - Total; | Metals - Total; EMS code: HG-T | Chemical | Total | LAB
HG-T | mg/L | Mercury Total | X127 | FA,HNO3,Flameless AA | FA;HNO3;Flameless AA | 0 | Y | Y | Y | 36 |  |  | 3/6/1997 7:02:30 PM | 1996-04-12 | 6/3/1998 6:51:36 PM | EMS |  | Concentration | mg/L | Y | (fl. conc.) |  | Mercury Total (fl. conc.) |  | Metals - Total; | Metals - Total; EMS code: HG-T | Chemical | Total | LAB
HG-T | ng/L | Mercury Total | T548 | EPA1631E-T | APHA 3030B / EPA 1631E - Trace | 0.1 | Y | Y | Y | 0 |  |  |  | 2024-08-19 |  | EMS |  | Concentration | ng/L | Y | (fl. conc.) |  | Mercury Total (fl. conc.) |  | Metals - Total; | Metals - Total; EMS code: HG-T | Chemical | Total | LAB
HG-T | mg/L | Mercury Total | MTDM | HG-DRY-MICR-CVAF-VA | Mercury in Tissue by CVAFS Micro (DRY) | 5e-06 | Y | Y | Y | 4 | 8/12/2015 | 10/16/2015 | 11/24/2015 1:22:59 PM | 2015-03-09 | 3/9/2015 3:32:04 PM | RDGIBSON | RDGIBSON | Concentration | mg/L | Y | (fl. conc.) |  | Mercury Total (fl. conc.) |  | Metals - Total; | Metals - Total; EMS code: HG-T | Chemical | Total | LAB
HG-T | mg/L | Mercury Total | X272 | HNO3:ColdVapour:AA | HNO3:ColdVapour:AA | 0.005 | Y | Y | Y | 74 |  |  | 2/15/2016 10:44:41 AM | 1996-04-12 | 6/3/1998 6:51:36 PM | EMS |  | Concentration | mg/L | Y | (fl. conc.) |  | Mercury Total (fl. conc.) |  | Metals - Total; | Metals - Total; EMS code: HG-T | Chemical | Total | LAB
HG-T | ug/g | Mercury Total | EQ02 | EQUIS (ug/g) | EQUIS (ug/g) | 0.5 | N | Y | Y | 213 |  |  | 10/23/1996 12:08:28 AM | 1996-04-12 | 6/3/1998 6:51:36 PM | EMS |  | MassPerMass | mg/kg | Y | (mass conc.) |  | Mercury Total (mass conc.) |  | Metals - Total; | Metals - Total; EMS code: HG-T | Chemical | Total | LAB
HG-T | t/d | Mercury Total | CALC | Calculation (t/d) | Calculation (t/d) | 0 | Y | Y | Y | 10952 |  |  | 7/3/2024 12:59:29 PM | 1996-04-12 | 6/3/1998 6:51:36 PM | EMS |  | MassFlow | t/d | Y | (mass flow) |  | Mercury Total (mass flow) |  | Metals - Total; | Metals - Total; EMS code: HG-T | Chemical | Total | LAB
HG-T | ug/g | Mercury Total | MTCD | HG-DRY-CVAFS-VA | Mercury in Tissue by CVAFS (DRY) | 0.005 | Y | Y | Y | 107 | 8/13/2014 | 3/8/2023 | 3/23/2023 4:12:02 PM | 2015-03-09 | 3/11/2015 4:07:14 PM | RDGIBSON | RDGIBSON | MassPerMass | mg/kg | Y | (mass conc.) |  | Mercury Total (mass conc.) |  | Metals - Total; | Metals - Total; EMS code: HG-T | Chemical | Total | LAB
HG-T | mg/L | Mercury Total | MSCV | HG-200.2-CVAF-VA (WET) | Mercury in Soil by CVAFS (WET) | 5e-06 | Y | Y | Y | 9667 | 4/2/2012 | 4/18/2020 | 12/8/2023 4:30:26 PM | 2015-03-09 | 3/9/2015 3:31:40 PM | RDGIBSON | RDGIBSON | Concentration | mg/L | Y | (fl. conc.) |  | Mercury Total (fl. conc.) |  | Metals - Total; | Metals - Total; EMS code: HG-T | Chemical | Total | LAB
HG-T | mg/L | Mercury Total | MWID | Metals in Water by ICPMS (digested) | Metals in Water by ICPMS (digested) Based on EPA 200.8 | 0.0002 | Y | Y | Y | 6 |  |  | 7/30/2024 6:32:12 AM | 2009-12-24 |  | RDGIBSON |  | Concentration | mg/L | Y | (fl. conc.) |  | Mercury Total (fl. conc.) |  | Metals - Total; | Metals - Total; EMS code: HG-T | Chemical | Total | LAB
HG-T | ug/g wet | Mercury Total | 0340 | Hom, N/S Dig:Man.C.V.-AN | Animal tissues are not dried. They are homogenized and digested with H2SO4/HNO3. The organomercury compounds are oxidized via a H2SO4/KMnO4/K2S2O8 digestion. The mercury compounds are then reduced to elemental mercury. A stream of synthetic air introduced into the system passes through a reference cell and then, after sparging the elemental mercury from the solution, passes through the sample cell. A UV monitor possessing a mercury lamp provides the light source and the absorbance at 253.7 nm is measured against the reference. | 0.05 | Y | Y | Y | 690 |  |  | 10/24/1996 2:24:57 PM | 1996-04-12 | 6/22/1998 1:57:15 PM | EMS | NPEPPIN | MassPerMass | µg/g | Y | (mass conc.) | (wet) | Mercury Total (wet) (mass conc.) |  | Metals - Total; | Metals - Total; EMS code: HG-T | Chemical | Total | LAB
HG-T | ug/g wet | Mercury Total | X218 | Hom,Nitr/Perchl,ColdV,UV | Hom;Nitr/Perchl;ColdV;UV | 0.001 | Y | Y | Y | 85 |  |  | 10/24/1996 12:37:38 PM | 1996-04-12 | 6/3/1998 6:51:42 PM | EMS |  | MassPerMass | µg/g | Y | (mass conc.) | (wet) | Mercury Total (wet) (mass conc.) |  | Metals - Total; | Metals - Total; EMS code: HG-T | Chemical | Total | LAB
HG-T | mg/L | Mercury Total | F121 | Acid,ICPMS | Acidified, no digestion, ICPMS | 1e-05 | Y | Y | Y | 181 | 10/2/2010 4:03:00 AM | 10/29/2013 11:28:00 PM | 10/26/2017 9:58:33 AM | 2010-11-16 | 7/16/2012 3:29:45 PM | RDGIBSON | RDGIBSON | Concentration | mg/L | Y | (fl. conc.) |  | Mercury Total (fl. conc.) |  | Metals - Total; | Metals - Total; EMS code: HG-T | Chemical | Total | LAB
HG-T | ug/g | Mercury Total | FB15 | Homo,HotHNO3dig,H2O2dig,CVAA | Homogenized; hotplate HNO3 digestion;H2O2 digestion; CVAA | 0.005 | Y | Y | Y | 467 | 6/20/2005 | 10/10/2007 | 5/24/2008 1:15:02 AM | 1999-04-08 |  | NPEPPIN |  | MassPerMass | mg/kg | Y | (mass conc.) |  | Mercury Total (mass conc.) |  | Metals - Total; | Metals - Total; EMS code: HG-T | Chemical | Total | LAB
HG-T | ug/g wet | Mercury Total | MTMW | HG-WET-MICR-CVAF-VA | Mercury in Tissue by CVAFS Micro (WET) | 0.001 | Y | Y | Y | 38 | 11/28/2018 | 2/26/2019 | 7/21/2021 2:55:07 PM | 2015-03-09 | 4/21/2021 5:16:40 PM | RDGIBSON | RDGIBSON | MassPerMass | µg/g | Y | (mass conc.) | (wet) | Mercury Total (wet) (mass conc.) |  | Metals - Total; | Metals - Total; EMS code: HG-T | Chemical | Total | LAB
HG-T | ug/g | Mercury Total | X551 | Metals in Biota (dry), total, acid digestion, ICP-SFMS | 0.1 gram of dried, homogenized biota is weighed in a microwave TFM vessel, and 8mL of Nitric Acid and 1mL of Hydrogen Peroxide are added. The sample is digested at 200 degrees C for 15 minutes in a closed vessel high pressure microwave oven. The sample is cooled, the volume adjusted to 100mL and introduced to an iductively coupled argon plasma-sector field mass spectrometer (ICP-SFMS) by pneumatic nebulization. Each respective element is measured at a specific mass to charge ratio, m/z value expressed in amu (atomic mass units), and at a specified instrument resolution. LR: Low Resolution ~300, MR: Mid Resolution ~4000 and HR: Hich Resolution ~10000. Ref: Method number 02-2705, National Laboratory for Environmental Testing, CCIW, Environment Canada, Burlington. Requested April 2003. | 0.002 | Y | Y | Y | 3 | 11/2/2022 3:26:00 PM | 11/2/2022 3:30:00 PM | 11/10/2022 4:22:46 PM | 2022-08-18 | 11/8/2022 2:37:29 PM | EMS | EMS | MassPerMass | mg/kg | Y | (mass conc.) |  | Mercury Total (mass conc.) |  | Metals - Total; | Metals - Total; EMS code: HG-T | Chemical | Total | LAB
HG-T | ug/g | Mercury Total | FS08 | D&G,Nitric/HCl,CVAF | Samples are dried at 60 C, then either sieved to -100 mesh (150 um) and rolled to homogenize, or finely ground with a mortar and pestle or a shatter box laboratory mill and rolled (depending on the submitter's requirements). The samples are then weighed into Teflon digestion liners and digested with HNO3 and HCI in a microwave oven. Samples are cooled, volumized, and allowed to settle overnight before being decanted. Anaysis is then performed via Inductively Coupled Argon Plasma Atomic Emission Spectrometry (ICP-AES) (Simultaneous multi element analysis), Graphite Furnace Atomic Absorption spectrometry (GF-AAS)(low level analysis), Cold Vapor Atomic Fluorescence Spectrometry (CV-AFS) (Mercury analysis) and/or Flame Atomic Absorption Spectrometry (F-AAS)(K analysis). | 0.008 | Y | Y | Y | 157 | 8/30/1996 8:01:00 AM | 9/23/2021 2:28:00 PM | 10/18/2021 9:46:09 AM | 1997-04-17 | 6/3/1998 6:51:46 PM | MPAK |  | MassPerMass | mg/kg | Y | (mass conc.) |  | Mercury Total (mass conc.) |  | Metals - Total; | Metals - Total; EMS code: HG-T | Chemical | Total | LAB
HG-T | mg/dm2/d | Mercury Total | EQ15 | EQUIS (mg/dm2/d) | EQUIS (mg/dm2/d) | 0 | N | Y | Y | 21 |  |  | 10/22/1996 10:18:25 PM | 1996-04-12 | 6/3/1998 6:51:36 PM | EMS |  | MassFlux | mg/dm²/d | Y | (mass flux) |  | Mercury Total (mass flux) |  | Metals - Total; | Metals - Total; EMS code: HG-T | Chemical | Total | LAB
HG-T | mg/L | Mercury Total | 2310 | HNO3/H2SO4 dig:Auto CVAA | Samples are field filtered for dissolved or not filtered for total, and preserved with K2Cr2O7/H2SO4 to ph <2. The organomercury compounds are oxidized via a H2SO4/KMnO4/K2S2O8 digestion. The mercury compounds are then reduced to elemental mercury. A stream of synthetic air introduced into the system passes through a reference cell and then, after sparging the elemental mercury from the solution, passes through the sample cell. A UV monitor possessing a mercury lamp provides the light source and the absorbance at 253.7 nm is measured against the reference. | 5e-05 | Y | Y | Y | 5539 | 6/25/2010 3:15:00 PM | 6/25/2010 3:15:00 PM | 7/22/2010 2:30:27 PM | 1996-04-12 | 6/3/1998 6:51:36 PM | EMS |  | Concentration | mg/L | Y | (fl. conc.) |  | Mercury Total (fl. conc.) |  | Metals - Total; | Metals - Total; EMS code: HG-T | Chemical | Total | LAB
HGME | ug/g | Mercury- Methyl | MMTD | HG-MEHG-DRY-GCAFS-VA | Methyl Mercury in Tissue by GCAFS (DRY) | 0.005 | Y | Y | Y | 39 | 8/13/2014 | 9/18/2014 | 12/15/2015 10:36:27 AM | 2015-03-09 |  | RDGIBSON |  | MassPerMass | mg/kg | Y | (mass conc.) |  | Mercury- Methyl (mass conc.) | 22967-92-6 | Organometallics; | Organometallics; EMS code: HGME | Chemical |  | LAB
HGME | ug/g wet | Mercury- Methyl | MMTW | HG-MEHG-WET-GCAFS-VA | Methyl Mercury in Tissue by GCAFS (WET) | 0.001 | Y | Y | Y | 39 | 8/13/2014 | 9/18/2014 | 12/15/2015 10:36:27 AM | 2015-03-09 |  | RDGIBSON |  | MassPerMass | µg/g | Y | (mass conc.) | (wet) | Mercury- Methyl (wet) (mass conc.) | 22967-92-6 | Organometallics; | Organometallics; EMS code: HGME | Chemical |  | LAB
HGME | mg/L | Mercury- Methyl | 1630 | US EPA 1630 in Water | Method 1630 of the US EPA. Samples are distilled under an inert gas flow to isolate methylmercury and minimize matrix interferences. The distillate is analyzed by aqueous phase ethylation, purge and trap, desorption and GC separation. THe separated species are then pyrolized to elemental Hg and quantified by cold vapour atomic flourescence spectroscopy. | 2e-08 | Y | Y | Y | 1032 | 8/13/2015 11:55:00 AM | 8/9/2024 | 8/12/2024 8:54:07 AM | 2017-11-29 | 9/17/2018 2:57:51 PM | JKROGH | RDGIBSON | Concentration | mg/L | Y | (fl. conc.) |  | Mercury- Methyl (fl. conc.) | 22967-92-6 | Organometallics; | Organometallics; EMS code: HGME | Chemical |  | LAB
HGME-D | mg/L | Methylmercury (as MeHg), dissolved | 1630 | US EPA 1630 in Water | Method 1630 of the US EPA. Samples are distilled under an inert gas flow to isolate methylmercury and minimize matrix interferences. The distillate is analyzed by aqueous phase ethylation, purge and trap, desorption and GC separation. THe separated species are then pyrolized to elemental Hg and quantified by cold vapour atomic flourescence spectroscopy. | 2e-05 | Y | Y | Y | 10 | 6/22/2021 | 7/31/2023 | 8/18/2023 4:29:13 PM | 2023-08-18 |  | EMS |  | Concentration | mg/L | Y | (fl. conc.) |  | Methylmercury (as MeHg), dissolved (fl. conc.) |  | Organometallics; | Organometallics; EMS code: HGME-D | Chemical | Dissolved | LAB
HH03 | mg/L | heptachlor + heptachlor epoxide | E660 | EPA 8270E (mod) | EPA 8270E (mod) Miscellaneous Pesticides by GC-MS | 3e-06 | Y | Y | Y | 2 | 9/14/2022 | 9/14/2022 | 12/8/2022 9:48:14 AM | 2022-12-06 |  | EMS |  | Concentration | mg/L | Y | (fl. conc.) |  | heptachlor + heptachlor epoxide (fl. conc.) |  | Pesticides; | Pesticides; EMS code: HH03 | Chemical |  | LAB
HH03 | mg/L | heptachlor + heptachlor epoxide | X268 | CH2Cl2 ext:GC/MS | CH2Cl2 ext:GC/MS |  | Y | Y | Y | 16 | 12/29/2021 | 9/2/2022 | 11/10/2022 9:54:07 AM | 2022-04-04 |  | EMS |  | Concentration | mg/L | Y | (fl. conc.) |  | heptachlor + heptachlor epoxide (fl. conc.) |  | Pesticides; | Pesticides; EMS code: HH03 | Chemical |  | LAB
HO-T | ppm (S) | Holmium | MS81 | ALS Minerals: ME-MS81 | ALS Minerals: ME-MS81. A prepared sample (0.100 g) is added to lithium metaborate/lithium tetraborate flux; mixed well and fused in a furnace at 1025°C. The resulting melt is then cooled and dissolved in an acid mixture containing nitric; hydrochloric and hydrofluoric acids. This solution is then analyzed by inductively coupled plasma - mass spectrometry. | 0.01 | Y | Y | Y | 0 |  |  |  | 2024-01-29 |  | EMS |  | DimensionlessRatio | ppm | N | (ratio) |  | Holmium (ratio) |  | Metals; | Metals; EMS code: HO-T | Chemical |  | LAB
HPC- | MPN/100mL | Heterotrophic Plate Count | QMPN | Q-Tray MPN | Q-Tray MPN | 0 | Y | Y | Y | 1 | 3/9/2023 6:01:00 PM | 3/9/2023 6:01:00 PM | 4/13/2023 4:05:14 PM | 2023-04-11 |  | EMS |  | MicrobialFormation | MPN/100mL | N | (microb.) |  | Heterotrophic Plate Count (microb.) |  | Bacteriological; | Bacteriological; EMS code: HPC- | Chemical |  | LAB
HPC- | CFU/mL | Heterotrophic Plate Count | X022 | Membrane Filter Count | Non-turbid water samples are passed through a 0.45 um membrane filter which is placed on a selective medium for 18 - 24 hours incubation at 35 degrees C for growth of total coliform colonies, of fecal coliform colonies, fecal streptococci colonies, enterococci colonies, or of E. coli colonies. | 0 | Y | Y | Y | 121 | 4/29/2002 | 8/8/2024 | 8/19/2024 12:29:02 PM | 2002-07-11 | 8/27/2002 11:28:16 AM | NPEPPIN | UMULLER | MicrobialFormation | CFU/mL | N | (microb.) |  | Heterotrophic Plate Count (microb.) |  | Bacteriological; | Bacteriological; EMS code: HPC- | Chemical |  | LAB
HVEX | None | HiVol Extended Metal | 5330 | HiVol Extended Reporting | HiVol Extended Reporting | 0 | Y | Y | Y | 1302 | 11/7/1997 2:00:00 PM | 3/29/2000 3:15:00 PM | 8/3/2000 1:06:29 AM | 1996-04-12 | 6/3/1998 6:51:36 PM | EMS |  | None | None | N | (unitless) |  | HiVol Extended Metal (unitless) |  | Metals; | Metals; EMS code: HVEX | Chemical |  | LAB
HYDR | mg/L | hydroxide (CAS 14280-30-9) | X310 | Auto Potentiometric Titration | For alkalinity, the samples are (if pH >8.3) first titrated to a pH of 8.3, which corresponds to the equivalence point for the conversion of carbonate to bicarbonate ion (phenolphthalein alkalinity). The titration is continued to a pH of 4.5, which corresponds approx. to the equivalence point for the conversion of bicarbonate ion to carbonic acid (entire titration indicates total alkalinity). An extrapolation technique based on the near proportionality of concentration of H ions to excess of titrant beyond the equivalence pt. The amount of standard acid required to reduce pH exactly 0.30pH (ie. pH4.2) unit is measured carefully. Because this change in pH corresponds to an exact coupling of the H ion concentration, a simple extrapolation can be made to the equivalence point. The actual pH of the stoichiometric end point in alkalinity determinations is the pH at which the inflection in the titration curve occurs and is directly related to the amount of carbonate alkalinity originally present in the sample. A curve can be constructed by plotting pH vs mL of acid to find this point directly. *** For acidity, the samples are titrated to a pH of 8.3, which corresponds to stoichiometric neutralization of carbonic acid to bicarbonate. Because the colour change of phenolphthalein indicator is close to pH 8.3, this value generally is accepted as a standard end point for titration of total acidity, including CO2 and most weak acids an equivalence point for the conversion of carbonate to bicarbonate ion (phenolphthalein alkalinity). The actual pH of the stoichiometric end point in acidity determinations is the pH at which the inflection in the titration curve occurs and is directly related to the amount of titratable acids originally present in the sample. A curve can be constructed by plotting pH vs. mL of NaOH to find this point directly. (To obtain the actual substances contributing to acidity/alkalinity, a pH vs mL of acid/NaOH curve should be constructed.) | 1 | Y | Y | Y | 90 | 11/24/2019 | 8/11/2024 | 8/20/2024 9:51:05 AM | 2020-01-29 |  | RDGIBSON |  | Concentration | mg/L | Y | (fl. conc.) |  | hydroxide (fl. conc.) | 14280-30-9 | Inorganics; | Inorganics; EMS code: HYDR | Chemical |  | LAB
I002 | mg/L | Imazethapyr | AEWA | H2SO4 acid,CHCl2 extr.florisil. GC/MS | Effluent and water samples are acidified with Sulphuric Acid and extracted with Dichloromethane. After concentration, the extract is methylated using ethereal Diazomethane. Extracts are cleaned up as required using Florisil. The extracts are analyzed for Acid Extractable Herbicides using GC/MS. | 0.0005 | Y | Y | Y | 1 | 7/29/2005 7:10:00 AM | 7/29/2005 7:10:00 AM | 6/2/2009 12:01:06 PM | 2003-06-17 |  | UMULLER |  | Concentration | mg/L | Y | (fl. conc.) |  | Imazethapyr (fl. conc.) | 81335-77-5 | Pesticides; | Pesticides; EMS code: I002 | Chemical |  | LAB
I003 | mg/L | Iprodione | P015 | Phenoxyacid Herbicides by LC-MS/MS | Phenoxyacid Herbicides by LC-MS/MS | 0.0001 | Y | Y | Y | 1 | 12/28/2021 | 12/28/2021 | 2/15/2022 9:50:05 AM | 2021-12-14 |  | EMS |  | Concentration | mg/L | Y | (fl. conc.) |  | Iprodione (fl. conc.) | 36734-19-7 | Pesticides; | Pesticides; EMS code: I003 | Chemical |  | LAB
I003 | mg/L | Iprodione | X268 | CH2Cl2 ext:GC/MS | CH2Cl2 ext:GC/MS | 1 | Y | Y | Y | 1 | 7/29/2005 7:10:00 AM | 7/29/2005 7:10:00 AM | 6/2/2009 12:01:06 PM | 1996-04-12 | 6/3/1998 6:51:42 PM | EMS |  | Concentration | mg/L | Y | (fl. conc.) |  | Iprodione (fl. conc.) | 36734-19-7 | Pesticides; | Pesticides; EMS code: I003 | Chemical |  | LAB
I004 | mg/L | Isopimaric Acid | F009 | pH 9,Diethyl Ether:HRGC/LRMS | Effluent samples are first adjusted to pH 9 and then extracted with Diethyl Ether. The extracts are rotary evaporated and the resulting residues, containing the Resin Acids and other solvent soluble organics are then redissolved in Methanol, transferred to small glass vials, methylated with Diazomethane then evaporated to dryness under Nitrogen. [Methylation Reaction R-COOH(Resin/Fatty Acids) + CH2N2(Diazomethane) = R-COOCH3(Methyl ester) + N2) The residue is re-dissolved in Methyl t-Butyl Ether (MTBE) and injected on a High Resolution Gas Chromatograph/Low Resolution Mass Spectrometer (HRGC/LRMS) system for analysis. LRMS data is acquired in Total Ion Mode (TIM) and quantitation is performed using selected mass ions. Positive identification of the various Resin Acids is made on the basis of a comparison of absolute retention times to those of the external standards and computer based mass spectral mass/intensity library matching routines. | 0.005 | Y | Y | Y | 198 | 12/27/1996 4:15:00 PM | 5/13/2003 | 4/1/2004 1:08:46 AM | 1996-09-16 | 6/3/1998 6:51:44 PM | MPAK |  | Concentration | mg/L | Y | (fl. conc.) |  | Isopimaric Acid (fl. conc.) |  | Organics; | Organics; EMS code: I004 | Chemical |  | LAB
I004 | mg/L | Isopimaric Acid | E677 | Resins & Fatty Acids are methylated then analyzed by GCMS | EPA 8270E; AB Env. 129; NCASI RA/FA-85.02 (mod) Resins & Fatty Acids are methylated then analyzed by GCMS. Total fatty acids is the sum of arachidic acid; behenic acid; lauric acid; lignoceric acid; linoleic acid; linolenic acid; myristic acid; oleic acid; palmitic acid; and stearic acid. Total resin acids is the sum of abietic acid; 12-chlorodehydroabietic acid; 14-chlorodehydroabietic acid; dehydroabietic acid; 12;14-dichlorodehydroabietic acid; isopimaric acid; levopimaric acid; neoabietic acid; palustric acid; pimaric acid and sandaracopimaric acid. Total resin acids (BC WWQG) is the sum of abietic acid; neoabietic acid; pimaric acid; isopimaric acid; and sandaracopimaric acid. | 0.001 | Y | Y | Y | 12 | 11/28/2022 | 6/15/2024 | 7/8/2024 11:10:02 AM | 2021-12-30 |  | EMS |  | Concentration | mg/L | Y | (fl. conc.) |  | Isopimaric Acid (fl. conc.) |  | Organics; | Organics; EMS code: I004 | Chemical |  | LAB
I004 | ug/g | Isopimaric Acid | FS05 | MeOH,DCM extr:HRGC/LRMS | Sediment samples are extracted by shaking with Methanol in a screw cap centrifuge bottle. The extracts are rotary evaporated and the resulting residues, containing the Resin Acids and other solvent soluble organics, are then redissolved in Methanol, transferred to small glass vials, methylated with Diazomethane then evaporated to dryness under Nitrogen. [Methylation Reaction R-COOH(Resin/Fatty Acids) + CH2N2 (Diazomethane) = R-COOCH3 (Methyl ester) + N2] The residue is re-dissolved in Methyl t-Butyl Ether (MTBE) and injected on a High Resolution Gas Chromatograph / Low Resolution Mass Spectrometer (HRGC/LRMS) system for analysis. LRMS data is acquired in Total Ion Mode (TIM) and quantitation is performed using selected mass ions. Positive identification of the various Resin Acids is made on the basis of a comparison of absolute retention times to those of the external standards and computer based mass spectral mass/intensity library matching routines. | 0.2 | Y | Y | Y | 5 | 9/14/2001 3:29:00 PM | 12/3/2002 | 3/27/2004 1:09:21 AM | 1996-09-16 | 6/3/1998 6:51:44 PM | MPAK |  | MassPerMass | mg/kg | Y | (mass conc.) |  | Isopimaric Acid (mass conc.) |  | Organics; | Organics; EMS code: I004 | Chemical |  | LAB
I004 | mg/L | Isopimaric Acid | P030 | Resin Acids | The water sample is acidified and extracted with dichloro-methane. The raw extract is concentrated and derivatized with diazomethane to produce the corresponding methyl ester derivatives. (Suitable substitute derivatives may be produced by other available techniques.) If required, the extracts are cleaned up by Florisil column chromatography. The derivatives are analyzed by gas chromatography with flame ionization detection. | 0.001 | Y | Y | Y | 469 |  |  | 3/23/1998 11:50:50 AM | 1996-04-12 | 6/3/1998 6:51:36 PM | EMS |  | Concentration | mg/L | Y | (fl. conc.) |  | Isopimaric Acid (fl. conc.) |  | Organics; | Organics; EMS code: I004 | Chemical |  | LAB
I004 | mg/L | Isopimaric Acid | X239 | Gas Chromatography | Gas Chromatography | 0.001 | Y | Y | Y | 12 |  |  | 10/23/1996 11:56:37 PM | 1996-04-12 | 6/3/1998 6:51:36 PM | EMS |  | Concentration | mg/L | Y | (fl. conc.) |  | Isopimaric Acid (fl. conc.) |  | Organics; | Organics; EMS code: I004 | Chemical |  | LAB
I004 | mg/L | Isopimaric Acid | RFAS | Separation of resin and fatty acids (RFAs) in water by GC/MS | Scientific Principle The separation of resin and fatty acids (RFAs) in water samples is performed using gas chromatography/mass spectrometry (GC/MS) in the Selected Ion Monitoring (SIM) mode. The separation is based on compound size; boiling point and polarity. Analyte concentration is calculated using the internal standard method. RFA compounds are extracted from water samples using a liquid-liquid solvent extraction with DCM. Extraction occurs after the sample is brought to a pH of ~ 2 and placed in a rotary tumbler (at approx. 40 rpm) for 1 hour. The resulting extract is transferred to a vial; dried with acidified baked sodium sulfate; methylated with diazomethane and concentrated to 1mL under nitrogen prior to analysis. The resulting extracts are vialed and fortified with labeled internal standard compounds prior to analysis by gas chromatography with a (50%-Phenyl)-methylpolysiloxane (DB-17) column and a mass spectrometer detector in SIM Mode. | 0.006 | Y | Y | Y | 1 | 9/17/2021 7:49:00 PM | 9/17/2021 7:49:00 PM | 2/1/2022 2:38:14 PM | 2022-01-19 |  | EMS |  | Concentration | mg/L | Y | (fl. conc.) |  | Isopimaric Acid (fl. conc.) |  | Organics; | Organics; EMS code: I004 | Chemical |  | LAB
I004 | mg/L | Isopimaric Acid | P035 | DCM-DIAZO-FLOR | DCM-DIAZO-FLOR | 0.0005 | Y | Y | Y | 22 |  |  | 10/24/1996 12:44:24 PM | 1996-04-12 | 6/3/1998 6:51:36 PM | EMS |  | Concentration | mg/L | Y | (fl. conc.) |  | Isopimaric Acid (fl. conc.) |  | Organics; | Organics; EMS code: I004 | Chemical |  | LAB
I004 | mg/L | Isopimaric Acid | P033 | Resin&Fatty Acids,GC/MS | Resin&Fatty Acids;GC/MS | 0.001 | Y | Y | Y | 31 |  |  | 6/7/2024 5:30:18 PM | 1996-04-12 | 6/3/1998 6:51:36 PM | EMS |  | Concentration | mg/L | Y | (fl. conc.) |  | Isopimaric Acid (fl. conc.) |  | Organics; | Organics; EMS code: I004 | Chemical |  | LAB
I005 | mg/L | Imazapyr | AEWA | H2SO4 acid,CHCl2 extr.florisil. GC/MS | Effluent and water samples are acidified with Sulphuric Acid and extracted with Dichloromethane. After concentration, the extract is methylated using ethereal Diazomethane. Extracts are cleaned up as required using Florisil. The extracts are analyzed for Acid Extractable Herbicides using GC/MS. | 0.0005 | Y | Y | Y | 1 | 7/29/2005 7:10:00 AM | 7/29/2005 7:10:00 AM | 6/2/2009 12:01:06 PM | 2003-06-17 |  | UMULLER |  | Concentration | mg/L | Y | (fl. conc.) |  | Imazapyr (fl. conc.) |  | Pesticides; | Pesticides; EMS code: I005 | Chemical |  | LAB
I006 | mg/L | Isofenphos | ZN01 | DCM+Ac.Extr:GC/MS | DCM+Ac.Extr:GC/MS | 0.0003 | Y | Y | Y | 28 | 7/29/2005 7:10:00 AM | 3/29/2012 10:15:00 PM | 5/3/2012 9:27:51 AM | 1996-04-12 | 11/16/1999 4:06:27 PM | EMS | NPEPPIN | Concentration | mg/L | Y | (fl. conc.) |  | Isofenphos (fl. conc.) |  | Pesticides; | Pesticides; EMS code: I006 | Chemical |  | LAB
I007 | mg/L | Iodofenphos | ZN01 | DCM+Ac.Extr:GC/MS | DCM+Ac.Extr:GC/MS | 0.0001 | Y | Y | Y | 51 | 10/3/2002 | 3/29/2012 10:15:00 PM | 5/3/2012 9:27:51 AM | 1996-04-12 | 11/16/1999 4:06:27 PM | EMS | NPEPPIN | Concentration | mg/L | Y | (fl. conc.) |  | Iodofenphos (fl. conc.) |  | Pesticides; | Pesticides; EMS code: I007 | Chemical |  | LAB
I007 | ug/g | Iodofenphos | ZN03 | DCM+Ac.Extr:GC/MS | DCM+Ac.Extr:GC/MS EPA 8270E (mod) | 0.01 | Y | Y | Y | 19 | 4/20/2002 10:25:00 AM | 10/16/2002 3:20:00 PM | 2/25/2004 1:09:46 AM | 1996-04-12 | 11/16/1999 2:27:40 PM | EMS | NPEPPIN | MassPerMass | mg/kg | Y | (mass conc.) |  | Iodofenphos (mass conc.) |  | Pesticides; | Pesticides; EMS code: I007 | Chemical |  | LAB
I008 | mg/L | Isopropylbenzene (cumene) | MS04 | Static Headspace GC/MS (Water) | (Water) A static headspace procedure is used as a means of driving the analytes out of the matrix into a headspace. Once equilibrium is achieved, a fixed volume of headspeace is taken for subsequent determination by Gas Chromatography - Mass Spectroscopy (GC/MS) using Selected Ion Monitoring. | 0.02 | Y | Y | Y | 8 | 10/5/2019 2:07:00 AM | 6/28/2024 12:29:00 AM | 7/9/2024 3:31:02 PM | 2021-04-07 |  | RDGIBSON |  | Concentration | mg/L | Y | (fl. conc.) |  | Isopropylbenzene (cumene) (fl. conc.) |  | Organics; | Organics; EMS code: I008 | Chemical |  | LAB
I010 | mg/L | p-Isopropyltoluene (1-Methyl-4-(propan-2-yl)benzene) | F007 | Static Headspace:HRGC/LRMS | Static headspace analysis is a method for the determination of volatile compounds in liquids and/or solids by analysing the vapour phase of a sample in a closed system; in this case a sealed vial. Static Headspace is sometimes referred to as Equilibrium Headspace because the vapour phase above the sample in a sealed vial is brought into thermodynamic equilibrium with the sample prior to analysis. Volatile organic compounds are thermally extracted from the sample matrix by Static Headspace sampling and introduced directly to a thick-film megabore GC column. Compounds eluting from the GC column are introduced into a Mass Spectrometer and target compounds are identified by a comparison of both retention times and mass spectra with those of reference standards stored in a computerized data system. Reference retention times and mass spectra for analytes are obtained by the measurement of calibration standards under precisely the same extraction conditions used for samples. The concentration of each identified component is measured by relating the response of the sample quantitation mass ions to the response of the quantitation mass ions of a series of external standards. A deuterated surrogate recovery standard is added to provide information on matrix effects. | 0.001 | Y | Y | Y | 0 |  |  |  | 2024-07-11 |  | RDGIBSON |  | Concentration | mg/L | Y | (fl. conc.) |  | p-Isopropyltoluene (1-Methyl-4-(propan-2-yl)benzene) (fl. conc.) |  | Organics; | Organics; EMS code: I010 | Chemical |  | LAB
I100 | mg/L | Imazamox | ZN01 | DCM+Ac.Extr:GC/MS | DCM+Ac.Extr:GC/MS | 0.1 | Y | Y | Y | 1 | 7/29/2005 7:10:00 AM | 7/29/2005 7:10:00 AM | 6/2/2009 12:01:06 PM | 2009-02-20 |  | JCHUDAK |  | Concentration | mg/L | Y | (fl. conc.) |  | Imazamox (fl. conc.) |  | Pesticides; | Pesticides; EMS code: I100 | Chemical |  | LAB
I101 | mg/L | Imidachloprid | ZN01 | DCM+Ac.Extr:GC/MS | DCM+Ac.Extr:GC/MS | 0.02 | Y | Y | Y | 1 | 7/29/2005 7:10:00 AM | 7/29/2005 7:10:00 AM | 6/2/2009 12:01:06 PM | 2009-02-20 |  | JCHUDAK |  | Concentration | mg/L | Y | (fl. conc.) |  | Imidachloprid (fl. conc.) |  | Pesticides; | Pesticides; EMS code: I101 | Chemical |  | LAB
IACR | None | Index of Additive Cancer Risk | 3570 | EPA 3570/8270 | EPA 3570/8270. PAH (Low) in Soil - CCME and BCAWQG | 0.15 | Y | Y | Y | 28 | 3/16/2020 | 3/22/2020 | 2/22/2023 12:51:03 PM | 2023-01-31 |  | EMS |  | None | None | N | (unitless) |  | Index of Additive Cancer Risk (unitless) |  | Toxicity; | Toxicity; EMS code: IACR | Chemical |  | LAB
IACR | None | Index of Additive Cancer Risk | CAL2 | Calculation (ug/g) | Calculation (ug/g) | 0.15 | Y | Y | Y | 19 | 1/10/2022 | 9/9/2022 | 1/11/2024 1:07:07 PM | 2022-07-28 |  | EMS |  | None | None | N | (unitless) |  | Index of Additive Cancer Risk (unitless) |  | Toxicity; | Toxicity; EMS code: IACR | Chemical |  | LAB
IACR | ug/g | Index of Additive Cancer Risk | FS87 | DCM,Ac/Hex extr,PAM,HRGC/LRMS | DCM;Ac/Hex extr;PAM;HRGC/LRMS | 0.15 | Y | Y | Y | 31 | 9/20/2017 | 12/16/2019 | 12/20/2019 10:51:07 AM | 2018-01-08 | 1/8/2018 4:06:16 PM | JKROGH | JKROGH | MassPerMass | mg/kg | Y | (mass conc.) |  | Index of Additive Cancer Risk (mass conc.) |  | Toxicity; | Toxicity; EMS code: IACR | Chemical |  | LAB
IBUA | mg/L | Iso-Butyric Acid | VFAD | Volatile Fatty Acids by GC/FID | Volatile Fatty Acids by GC/FID; Standard Methods for the Examination of Water and Wastewater; 21st Edition, Method 5560 D; pages 5-59 to 5-61. 5560 D. Gas Chromatographic Method Principle: This gas chromatographic procedure may be usedto determine the individual concentrations of the following fattyacids: acetic; propionic; butyric; isobutyric; valeric; andisovaleric acids in digester sludge samples. Caproic and heptanoicacids also may be determined; but with less certaintybecause of their reduced solubility in water. Volatile fatty acidsare generally soluble in digester sludge and exist in either theprotonated or unprotonated form; depending on the pH of thesludge sample. Acetic and propionic acids usually are in higherconcentrations than the other fatty acids in digester sludge. Allacids must be converted to the protonated (volatile) form beforeinjection into the gas chromatograph (GC) so vaporization mayoccur. The fatty acids must also be separated from the solidsmaterial in the sample before introduction into the GC to minimizedegradation of the GC column. Digester sludge samples areprepared by acidification; centrifugation; and filtration. The sampleis analyzed by direct injection into a gas chromatographequipped with a flame ionization detector after both centrifugationand filtration.b. Interference: The blank amount may be high relative tosample analyte. High blanks may be caused by a buildup ofcontaminants in the injector and guard column or by samplecarry-over. The analyst must be aware of the blank levels andcorrect unacceptably high blank levels that are above the reportinglimits.c. Minimum detectable concentration: The method detectionlevel has been determined by the USEPA method1 in reagentwater to be about 3 mg/L for acetic acid and about 1 mg/L for allother target compounds. | 2 | Y | Y | Y | 5 | 3/14/2017 10:30:00 AM | 11/19/2017 5:10:00 PM | 6/20/2018 9:02:11 AM | 2018-06-19 |  | RDGIBSON |  | Concentration | mg/L | Y | (fl. conc.) |  | Iso-Butyric Acid (fl. conc.) |  | Organics; | Organics; EMS code: IBUA | Chemical |  | LAB
IBUOH | ug/g | Isobutanol CAS: 78-83-1 | E8015S | EPA 8015 D in soil | See EPA 8015D | 10 | Y | Y | Y | 15 | 1/12/2022 | 2/1/2022 | 1/11/2024 1:07:07 PM | 2022-06-23 |  | EMS |  | MassPerMass | mg/kg | Y | (mass conc.) |  | Isobutanol (mass conc.) | 78-83-1 | Organics; | Organics; EMS code: IBUOH | Chemical |  | LAB
IBUOH | mg/L | Isobutanol CAS: 78-83-1 | E-8015 | EPA 8015D | EPA Method 8015D (SW-846): Nonhalogenated Organics Using GC/FID |  | Y | Y | Y | 14 | 12/29/2021 | 1/12/2022 | 9/2/2022 6:21:53 PM | 2022-04-05 |  | EMS |  | Concentration | mg/L | Y | (fl. conc.) |  | Isobutanol(fl. conc.) | 78-83-1 | Organics; | Organics; EMS code: IBUOH | Chemical |  | LAB
IC50 | %(V/V) | Bio.IC50 | FB25 | S.capricomutum growth inhibition (vol) | Using the microplate technique and the Coulter Counter Model Z1, growth of the algae, S. capricornutum, exposed to the test substance is compared with algae growth in a control over a 72 hour time period. The 72 hour IC50 (the median inhibition concentration, i.e., the concentration estimated to cause a 50% reduction in growth compared to the control) or IC25 (the median inhibition concentration, i.e., the concentration estimated to cause a 25% reduction in growth compared to the control) is computed by comparing growth in the various test concentrations at the end of the test. Reference toxicant tests or positive controls are used to establish alga sensitivity. | 0 | Y | Y | Y | 10 | 1/24/2000 12:59:00 PM | 2/7/2000 9:47:00 AM | 5/29/2002 2:29:33 PM | 2000-10-25 |  | NPEPPIN |  | DimensionlessRatio | % | N | (ratio) | (by vol.) | Bio.IC50 (by vol.) (ratio) |  | Toxicity; | Toxicity; EMS code: IC50 | Chemical |  | LAB
IC63UM | % | Inorganic carbon <63um | CSSS20 | CSSS (2008) 20.2 | See Canadian Society of Soil Science (2008) section 20.2 | 0.05 | Y | Y | Y | 80 | 7/12/2022 | 9/20/2023 | 1/12/2024 11:16:52 AM | 2022-07-27 |  | EMS |  | DimensionlessRatio | % | N | (ratio) |  | Inorganic carbon <63um (ratio) |  | Carbon; | Carbon; EMS code: IC63UM | Chemical |  | LAB
ICPA | mg/L | Iso-Caproic Acid | VFAD | Volatile Fatty Acids by GC/FID | Volatile Fatty Acids by GC/FID; Standard Methods for the Examination of Water and Wastewater; 21st Edition, Method 5560 D; pages 5-59 to 5-61. 5560 D. Gas Chromatographic Method Principle: This gas chromatographic procedure may be usedto determine the individual concentrations of the following fattyacids: acetic; propionic; butyric; isobutyric; valeric; andisovaleric acids in digester sludge samples. Caproic and heptanoicacids also may be determined; but with less certaintybecause of their reduced solubility in water. Volatile fatty acidsare generally soluble in digester sludge and exist in either theprotonated or unprotonated form; depending on the pH of thesludge sample. Acetic and propionic acids usually are in higherconcentrations than the other fatty acids in digester sludge. Allacids must be converted to the protonated (volatile) form beforeinjection into the gas chromatograph (GC) so vaporization mayoccur. The fatty acids must also be separated from the solidsmaterial in the sample before introduction into the GC to minimizedegradation of the GC column. Digester sludge samples areprepared by acidification; centrifugation; and filtration. The sampleis analyzed by direct injection into a gas chromatographequipped with a flame ionization detector after both centrifugationand filtration.b. Interference: The blank amount may be high relative tosample analyte. High blanks may be caused by a buildup ofcontaminants in the injector and guard column or by samplecarry-over. The analyst must be aware of the blank levels andcorrect unacceptably high blank levels that are above the reportinglimits.c. Minimum detectable concentration: The method detectionlevel has been determined by the USEPA method1 in reagentwater to be about 3 mg/L for acetic acid and about 1 mg/L for allother target compounds. | 2 | Y | Y | Y | 5 | 3/14/2017 10:30:00 AM | 11/19/2017 5:10:00 PM | 6/20/2018 9:02:11 AM | 2018-06-19 |  | RDGIBSON |  | Concentration | mg/L | Y | (fl. conc.) |  | Iso-Caproic Acid (fl. conc.) |  | Organics; | Organics; EMS code: ICPA | Chemical |  | LAB
IMID | mg/L | Imidacloprid (Insecticide/Pesticide) | E751 | E751D Pesticides in Water by LC-MS/MS (POS1 List) | E751D Pesticides in Water by LC-MS/MS (POS1 List) (MOECC E3501) | 0.0001 | Y | Y | Y | 39 | 12/9/2021 | 12/31/2021 | 2/15/2022 9:50:05 AM | 2021-12-14 |  | EMS |  | Concentration | mg/L | Y | (fl. conc.) |  | Imidacloprid (Insecticide/Pesticide) (fl. conc.) |  | Pesticides; | Pesticides; EMS code: IMID | Chemical |  | LAB
IMID | mg/L | Imidacloprid (Insecticide/Pesticide) | 3544 | MOECC E3544 NEONIC-DI-LCMS-WT | MOECC E3544 NEONIC-DI-LCMS-WT Water sample is filtered; spiked with internal standard and analyzed by direct injection using a HPLC Triple Quadrupole MS/MS with Electro Spray Ionization. | 5e-06 | Y | Y | Y | 29 | 6/5/2020 | 7/27/2020 | 7/21/2021 6:01:34 PM | 2020-06-30 |  | RDGIBSON |  | Concentration | mg/L | Y | (fl. conc.) |  | Imidacloprid (Insecticide/Pesticide) (fl. conc.) |  | Pesticides; | Pesticides; EMS code: IMID | Chemical |  | LAB
IN-T | ppm (S) | Indium Total | MS41 | ALS Minerals ME-MS41 | ALS Minerals ME-MS41. A prepared sample is digested with aqua regia in a graphite heating block. After cooling; the resulting solution is diluted to with deionized water; mixed and analyzed by ICP-AES. Samples are then analyzed by ICP-MS for the remaining suite of elements. | 0.005 | Y | Y | Y | 0 |  |  |  | 2024-01-29 |  | EMS |  | DimensionlessRatio | ppm | N | (ratio) |  | Indium Total (ratio) |  | Metals - Total; | Metals - Total; EMS code: IN-T | Chemical | Total | LAB
IN-T | ppm (S) | Indium Total | MS61 | ALS Minerals ME-MS61 | ALS Minerals ME-MS61. A prepared sample is digested with HNO3; HF; and HClO4 in a graphite heating block to near dryness. After cooling; the sample is boiled in diluted HCl. Bulk to volume with deionized water; mixed and analyzed by ICP-AES and ICP-MS for all elements. | 0.005 | Y | Y | Y | 0 |  |  |  | 2024-01-29 |  | EMS |  | DimensionlessRatio | ppm | N | (ratio) |  | Indium Total (ratio) |  | Metals - Total; | Metals - Total; EMS code: IN-T | Chemical | Total | LAB
IN-T | ug/g | Indium Total | EQ02 | EQUIS (ug/g) | EQUIS (ug/g) | 20 | N | Y | Y | 2 |  |  | 10/22/1996 9:46:36 PM | 1996-04-12 | 6/3/1998 6:51:36 PM | EMS |  | MassPerMass | mg/kg | Y | (mass conc.) |  | Indium Total (mass conc.) |  | Metals - Total; | Metals - Total; EMS code: IN-T | Chemical | Total | LAB
IN-T | ppm (S) | Indium Total | 6042 | Aqua Regia by ICPMS | These were sediment/soils digested with Aqua-Regia (Nitric Acid & Hydrochloric Acids) and then analysis by ICPMS. Results would be reported in ppm s. | 0.005 | Y | Y | Y | 71 | 5/14/2022 | 7/27/2024 | 7/29/2024 9:13:08 AM | 2022-06-28 |  | EMS |  | DimensionlessRatio | ppm | N | (ratio) |  | Indium Total (ratio) |  | Metals - Total; | Metals - Total; EMS code: IN-T | Chemical | Total | LAB
IN-T | ppm (S) | Indium Total | 6041 | Aqua Regia by ICPMS | These were sediment/soils digested with Aqua-Regia (Nitric Acid & Hydrochloric Acids) and then analysis by ICPMS. Results would be reported in mg/kg and/or %. | 0.005 | Y | Y | Y | 21 | 10/13/2019 | 8/5/2020 | 8/10/2020 3:52:06 PM | 2019-09-06 | 10/21/2019 11:55:19 AM | KATFRASE | KATFRASE | DimensionlessRatio | ppm | N | (ratio) |  | Indium Total (ratio) |  | Metals - Total; | Metals - Total; EMS code: IN-T | Chemical | Total | LAB
INDM | mg/L | Indomethacin CAS: 53-86-1 | XMSP | HPLC and Tri. Quadpole MS/MS | EPA 1694 An aliquot of filtered sample containing 5% organic is injected directly. Each compound is separated by reversed phase HPLC. Analyzed by Electro Spray Ionization MS/MS detection using Triple Quadrupole MS/MS detector. Quantiitation with internal or external standard. |  | Y | Y | Y | 12 | 5/8/2020 | 6/4/2024 | 8/21/2024 3:32:03 PM | 2022-05-04 |  | EMS |  | Concentration | mg/L | Y | (fl. conc.) |  | Indomethacin(fl. conc.) | 53-86-1 | Organics; | Organics; EMS code: INDM | Chemical |  | LAB
IPA | mg/L | Isopropyl Alcohol, Isopropanol CAS: 67-63-0 | E-8015 | EPA 8015D | EPA Method 8015D (SW-846): Nonhalogenated Organics Using GC/FID |  | Y | Y | Y | 14 | 12/29/2021 | 1/12/2022 | 9/2/2022 6:21:53 PM | 2022-04-05 |  | EMS |  | Concentration | mg/L | Y | (fl. conc.) |  | Isopropyl Alcohol, Isopropanol (fl. conc.) | 67-63-0 | Volatile Organic Compounds (VOCs); | Volatile Organic Compounds (VOCs); EMS code: IPA | Chemical |  | LAB
IPA | ug/g | Isopropyl Alcohol, Isopropanol CAS: 67-63-0 | E8015S | EPA 8015 D in soil | See EPA 8015D | 10 | Y | Y | Y | 15 | 1/12/2022 | 2/1/2022 | 1/11/2024 1:07:07 PM | 2022-06-23 |  | EMS |  | MassPerMass | mg/kg | Y | (mass conc.) |  | Isopropyl Alcohol, Isopropanol (mass conc.) | 67-63-0 | Volatile Organic Compounds (VOCs); | Volatile Organic Compounds (VOCs); EMS code: IPA | Chemical |  | LAB
IPAC | mg/L | Isopimaric & Palustric Acids | P030 | Resin Acids | The water sample is acidified and extracted with dichloro-methane. The raw extract is concentrated and derivatized with diazomethane to produce the corresponding methyl ester derivatives. (Suitable substitute derivatives may be produced by other available techniques.) If required, the extracts are cleaned up by Florisil column chromatography. The derivatives are analyzed by gas chromatography with flame ionization detection. | 0.001 | Y | Y | Y | 12 | 8/26/2014 | 8/15/2017 | 8/15/2017 4:37:09 PM | 2015-03-10 |  | RDGIBSON |  | Concentration | mg/L | Y | (fl. conc.) |  | Isopimaric & Palustric Acids (fl. conc.) | 5835-26-7, 1945-53-5 | Organics; | Organics; EMS code: IPAC | Chemical |  | LAB
IVLA | mg/L | Iso-Valeric Acid | VFAD | Volatile Fatty Acids by GC/FID | Volatile Fatty Acids by GC/FID; Standard Methods for the Examination of Water and Wastewater; 21st Edition, Method 5560 D; pages 5-59 to 5-61. 5560 D. Gas Chromatographic Method Principle: This gas chromatographic procedure may be usedto determine the individual concentrations of the following fattyacids: acetic; propionic; butyric; isobutyric; valeric; andisovaleric acids in digester sludge samples. Caproic and heptanoicacids also may be determined; but with less certaintybecause of their reduced solubility in water. Volatile fatty acidsare generally soluble in digester sludge and exist in either theprotonated or unprotonated form; depending on the pH of thesludge sample. Acetic and propionic acids usually are in higherconcentrations than the other fatty acids in digester sludge. Allacids must be converted to the protonated (volatile) form beforeinjection into the gas chromatograph (GC) so vaporization mayoccur. The fatty acids must also be separated from the solidsmaterial in the sample before introduction into the GC to minimizedegradation of the GC column. Digester sludge samples areprepared by acidification; centrifugation; and filtration. The sampleis analyzed by direct injection into a gas chromatographequipped with a flame ionization detector after both centrifugationand filtration.b. Interference: The blank amount may be high relative tosample analyte. High blanks may be caused by a buildup ofcontaminants in the injector and guard column or by samplecarry-over. The analyst must be aware of the blank levels andcorrect unacceptably high blank levels that are above the reportinglimits.c. Minimum detectable concentration: The method detectionlevel has been determined by the USEPA method1 in reagentwater to be about 3 mg/L for acetic acid and about 1 mg/L for allother target compounds. | 2 | Y | Y | Y | 5 | 3/14/2017 10:30:00 AM | 11/19/2017 5:10:00 PM | 6/20/2018 9:02:11 AM | 2018-06-19 |  | RDGIBSON |  | Concentration | mg/L | Y | (fl. conc.) |  | Iso-Valeric Acid (fl. conc.) |  | Organics; | Organics; EMS code: IVLA | Chemical |  | LAB
K--D | mg/L | Potassium Dissolved | X170 | ICP | This code is a catch-all for a large variety of ICP methods including water and also solids digested into solutions as an intermediate result. Different procedures will have greatly different detection limits. (a) For air particulates, preliminary results are reported in mg/L and loading results as mg/dm2/d. (b) For Hi Vol samples preliminiary results are same as reported in mg/L and loading results as ug/m3. (c) Note that for Rain samples there is no filtration nor digestion steps. | 2 | Y | Y | Y | 16456 | 4/11/1996 6:01:00 PM | 11/6/2018 | 8/26/2024 2:41:32 PM | 1996-04-12 | 8/15/2019 1:19:41 PM | EMS | KATFRASE | Concentration | mg/L | Y | (fl. conc.) |  | Potassium Dissolved (fl. conc.) |  | Metals - Dissolved; | Metals - Dissolved; EMS code: K--D | Chemical | Dissolved | LAB
K--D | mg/L | Potassium Dissolved | F049 | Filt,HNO3 dig,conc'n,nebul'ICP-OES | This method provides an automated procedure for the Inductively Coupled Argon Plasma Optical Emission Spectrophotometric (ICP-OES) determination and quantitation of trace amounts of metals in surface, ground and precipitation waters after a manual acid digestion and concentration procedure and ultrasonic nebulisation. Organically-bound metals and metal complexes are converted to metal ions by digestion with nitric acid. To increase the signal detected by the ICP, the samples are concentrated 10 times. The liquid sample is pumped onto the piezoelectric transducer, where the ultrasonic action efficiently converts it to a fine, dense aerosol. The nebulizer carrier gas flow then transports the wet aerosol through a heated U-tube, where the solvent is vaporized. Subsequently, most of the solvent vapour is removed from the aerosol stream by passing the stream through a chilled condenser. The result is a dry, analyte-laden aerosol which is introduced to the plasma. There, the high temperature desolvates the aerosol and dissociates the sample material into individual atoms and ions, which are excited to emit light at wave lengths characteristic of the elements present. Dissolved samples are field filtered through 0.45 micron cellulose acetate filter prior to acid preservation. Samples for extractable analysis are acidified with 2 mL concentrated HNO3 per litre and allowed to stand a minimum of 24 hours before analysis. Clients providing these samples are advised to preserve samples at the collection site. | 2 | Y | Y | Y | 82 | 3/10/1998 2:07:00 PM | 2/25/1999 2:08:00 PM | 10/20/1999 2:21:45 AM | 1998-11-17 |  | NPEPPIN |  | Concentration | mg/L | Y | (fl. conc.) |  | Potassium Dissolved (fl. conc.) |  | Metals - Dissolved; | Metals - Dissolved; EMS code: K--D | Chemical | Dissolved | LAB
K--D | mg/L | Potassium Dissolved | F039 | Field filt + acid, ICP-AES | Dissolved and extractable water samples are ready for analysis once collected. Total samples are autoclaved with HNO3 and hydrochloric acid (HCI) in capped Teflon vials. Analysis is performed via Inductively Coupled Argon Plasma Atomic Emission Spectrometry (ICP-AES)(simultaneous multielement analysis), Graphite Furnace Atomic Absorption Spectrometry (GF-AAS)(low level analysis) and/or Flame Atomic Absorption Spectrometry (F-AAS). | 0.1 | Y | Y | Y | 8584 | 7/29/1996 6:03:00 PM | 7/5/2019 1:21:00 PM | 5/16/2022 4:21:33 PM | 1997-05-08 | 6/3/1998 6:51:46 PM | MPAK |  | Concentration | mg/L | Y | (fl. conc.) |  | Potassium Dissolved (fl. conc.) |  | Metals - Dissolved; | Metals - Dissolved; EMS code: K--D | Chemical | Dissolved | LAB
K--D | mg/L | Potassium Dissolved | EQ01 | EQUIS (mg/L) | EQUIS (mg/L) | 0.1 | N | Y | Y | 6130 |  |  | 10/24/1996 3:16:58 PM | 1996-04-12 | 6/3/1998 6:51:36 PM | EMS |  | Concentration | mg/L | Y | (fl. conc.) |  | Potassium Dissolved (fl. conc.) |  | Metals - Dissolved; | Metals - Dissolved; EMS code: K--D | Chemical | Dissolved | LAB
K--D | mg/L | Potassium Dissolved | X562 | Dissolved metals: FF, FA, HNO3, HR-ICPMS | Ultra trace metals in water are analyzed by high resolution inductively coupled plasma mass spectrometry (HR-ICPMS), based on US EPA Method 200.8, (Oct. 1999). This procedure is intended for pristine field-filtered acid-preserved water samples. The detection limits (LOR) for this test are based on lab instrumental analysis only, not including filtration. ALS-supplied field filtration equipment does not support these LOR. Therefore, because of the high probability of false positives due to filtration, it is strongly recommended that a filtration blank be analyzed to aid in data interpretation. | 0.005 | Y | Y | Y | 1260 | 6/8/2015 | 10/19/2023 | 11/9/2023 7:41:27 AM | 2015-10-22 |  | RDGIBSON |  | Concentration | mg/L | Y | (fl. conc.) |  | Potassium Dissolved (fl. conc.) |  | Metals - Dissolved; | Metals - Dissolved; EMS code: K--D | Chemical | Dissolved | LAB
K--D | mg/L | Potassium Dissolved | METC | MET-D-NP-CCMS-VA | Ultra trace metals in water are analyzed by CRC ICPMS, based on US EPA Method 6020B (July 2014). This procedure is intended for pristine field-filtered acid-preserved water samples. The detection limits (LOR) for this test are based on lab instrumental analysis only, not including filtration. ALS-supplied field filtration equipment does not support these LOR. Therefore, because of the high probability of false positives due to filtration, it is strongly recommended that a filtration blank be analysed to aid in data interpretation. | 0.02 | Y | Y | Y | 1156 | 11/13/2017 | 8/15/2024 | 8/23/2024 9:50:35 AM | 2019-08-20 |  | KATFRASE |  | Concentration | mg/L | Y | (fl. conc.) |  | Potassium Dissolved (fl. conc.) |  | Metals - Dissolved; | Metals - Dissolved; EMS code: K--D | Chemical | Dissolved | LAB
K--D | mg/L | Potassium Dissolved | MSEA | Metals in Seawater (undigested) by (CRC) ICP-MS | This method describes the multi-elemental determination of trace elements by Collision/Reaction Cell (CRC) ICP-MS. The method measures ions produced by a radio-frequency inductively coupled plasma. Analyte species originating in a liquid are pneumatically nebulized and the resulting aerosol transported by argon gas into the plasma torch. The ions produced are entrained in the plasma gas and extracted, by means of a differentially pumped vacuum interface, into a mass spectrometer. The ions flow through a Collision/Reaction Cell (CRC) where depending on the mode of analysis either the ions passes through the cell unhindered (no gas - classical ICP-MS mode), or collisions/reactions chemistry takes place between the ions and a neutral gas such as hydrogen or helium. Any polyatomic ions which give interference on the analytes are removed. The ions are then sorted according to their mass-to-charge ratios by a quadrupole mass spectrometer. The ions are then sorted according to their mass-to-charge ratios by a quadrupole mass spectrometer. The ions transmitted through the quadrupole are quantified by a channel electron multiplier and the ion information processed by a data handling system. Although most polyatomic interference is removed, any additional interference must be assessed and valid corrections applied or the data flagged to indicate problems. Interference correction must include compensation for background ions contributed by the plasma gas, reagents, and constituents of the sample matrix. Instrument drift as well as suppressions or enhancements of instrument response caused by the sample matrix must be corrected for by the use of internal standards. | 1 | Y | Y | Y | 4 | 4/12/2023 | 4/29/2024 | 5/16/2024 2:00:14 PM | 2009-09-08 | 5/26/2023 3:37:23 PM | RDGIBSON | EMS | Concentration | mg/L | Y | (fl. conc.) |  | Potassium Dissolved (fl. conc.) |  | Metals - Dissolved; | Metals - Dissolved; EMS code: K--D | Chemical | Dissolved | LAB
K--D | mg/L | Potassium Dissolved | MWID | Metals in Water by ICPMS (digested) | Metals in Water by ICPMS (digested) Based on EPA 200.8 | 1 | Y | Y | Y | 2280 |  |  | 7/31/2024 1:46:41 PM | 2009-12-24 |  | RDGIBSON |  | Concentration | mg/L | Y | (fl. conc.) |  | Potassium Dissolved (fl. conc.) |  | Metals - Dissolved; | Metals - Dissolved; EMS code: K--D | Chemical | Dissolved | LAB
K--D | mg/L | Potassium Dissolved | F032 | Filt.,Acid,ICPMS | There is no acid digestion required for this analysis. A water sample is collected and preserved in 0.2 percent nitric acid. To analyse the extractable metals the sample is shaken, allowed to sit for at least twenty four hours, and a decanted portion is introduced to the spectrometer. When the dissolved metals are required, the samples must first be filtered through a 0.45 micron cellulose acetate filter paper and then followed by preservation in 0.2 percent nitric acid. It is important to note that different laboratories may have very different detection limits for some variables analyzed by ICPMS. | 0.01 | Y | Y | Y | 80685 | 5/7/1998 | 8/26/2024 | 8/26/2024 4:53:05 PM | 1998-05-14 | 6/3/1998 6:51:49 PM | NPEPPIN |  | Concentration | mg/L | Y | (fl. conc.) |  | Potassium Dissolved (fl. conc.) |  | Metals - Dissolved; | Metals - Dissolved; EMS code: K--D | Chemical | Dissolved | LAB
K--D | mg/L | Potassium Dissolved | X285 | HNO3 Dig:ICP Analysis | HNO3 Dig:ICP Analysis | 0.1 | Y | Y | Y | 675 | 1/30/2002 | 11/13/2002 | 2/19/2004 4:22:43 PM | 1998-05-22 | 6/3/1998 6:51:49 PM | NPEPPIN |  | Concentration | mg/L | Y | (fl. conc.) |  | Potassium Dissolved (fl. conc.) |  | Metals - Dissolved; | Metals - Dissolved; EMS code: K--D | Chemical | Dissolved | LAB
K--D | mg/L | Potassium Dissolved | X085 | LF,LA,HNO3,AA | Samples are lab-filtered (0.45 um), and lab-acidified (HNO3); analysis by AA. | 0.01 | Y | Y | Y | 78 | 5/9/2003 9:40:00 AM | 9/16/2003 3:18:00 PM | 8/11/2021 12:04:11 PM | 1996-04-12 | 6/3/1998 6:51:36 PM | EMS |  | Concentration | mg/L | Y | (fl. conc.) |  | Potassium Dissolved (fl. conc.) |  | Metals - Dissolved; | Metals - Dissolved; EMS code: K--D | Chemical | Dissolved | LAB
K--D | mg/L | Potassium Dissolved | X159 | Auto Flame Emission | Auto Flame Emission | 0.1 | Y | Y | Y | 6523 | 9/29/2000 | 12/1/2010 | 12/9/2010 3:55:26 PM | 1996-04-12 | 6/3/1998 6:51:46 PM | EMS |  | Concentration | mg/L | Y | (fl. conc.) |  | Potassium Dissolved (fl. conc.) |  | Metals - Dissolved; | Metals - Dissolved; EMS code: K--D | Chemical | Dissolved | LAB
K--E | mg/L | Potassium Extractable | F040 | Field Acid, lab filt, ICP-AES | Dissolved and extractable water samples are ready for analysis once collected. Total samples are autoclaved with HNO3 and hydrochloric acid (HCI) in capped Teflon vials. Analysis is performed via Inductively Coupled Argon Plasma Atomic Emission Spectrometry (ICP-AES)(simultaneous multielement analysis), Graphite Furnace Atomic Absorption Spectrometry (GF-AAS)(low level analysis) and/or Flame Atomic Absorption Spectrometry (F-AAS). | 0.1 | Y | Y | Y | 2368 | 8/5/1997 6:06:00 PM | 3/23/2012 3:11:00 AM | 9/11/2019 8:47:41 AM | 1997-05-08 | 6/3/1998 6:51:46 PM | MPAK |  | Concentration | mg/L | Y | (fl. conc.) |  | Potassium Extractable (fl. conc.) |  | Metals - Extractable; | Metals - Extractable; EMS code: K--E | Chemical | Extractable | LAB
K--E | ug/g | Potassium Extractable | ASCE | Ammonium Acetate soil colloid exchange | Available metals - Ammonium Acetate soil colloid exchange sites displacement |  | Y | Y | Y | 7 | 10/31/2022 | 10/31/2022 | 11/22/2022 2:40:39 PM | 2007-10-04 |  | RDGIBSON |  | MassPerMass | mg/kg | Y | (mass conc.) |  | Potassium Extractable (mass conc.) |  | Metals - Extractable; | Metals - Extractable; EMS code: K--E | Chemical | Extractable | LAB
K--E | mg/L | Potassium Extractable | F055 | No Pres/No Filtr/No Dig,ICP | Shortcut method which excludes preservation, filtration and digestion steps. Results are used for estimation of water hardness. | 0.1 | Y | Y | Y | 2359 | 8/23/1996 6:04:00 PM | 6/1/2000 11:27:00 AM | 9/11/2019 8:46:38 AM | 1998-04-21 | 6/3/1998 6:51:48 PM | NPEPPIN |  | Concentration | mg/L | Y | (fl. conc.) |  | Potassium Extractable (fl. conc.) |  | Metals - Extractable; | Metals - Extractable; EMS code: K--E | Chemical | Extractable | LAB
K--E | ug/g | Potassium Extractable | CAL2 | Calculation (ug/g) | Calculation (ug/g) | 2 | Y | Y | Y | 15 | 1/11/2022 | 2/7/2022 | 1/11/2024 1:07:07 PM | 2022-07-06 |  | EMS |  | MassPerMass | mg/kg | Y | (mass conc.) |  | Potassium Extractable (mass conc.) |  | Metals - Extractable; | Metals - Extractable; EMS code: K--E | Chemical | Extractable | LAB
K--E | mg/L | Potassium Extractable | E6020B | EPA 6020B | See EPA 6020B and CSSS Ch. 15 | 4 | Y | Y | Y | 15 | 1/11/2022 | 2/4/2022 | 1/11/2024 1:07:07 PM | 2022-06-29 |  | EMS |  | Concentration | mg/L | Y | (fl. conc.) |  | Potassium Extractable (fl. conc.) |  | Metals - Extractable; | Metals - Extractable; EMS code: K--E | Chemical | Extractable | LAB
K--E | mg/L | Potassium Extractable | X515 | HNO3 add,dir Asp,AA | 5ml 11 HNO3 added;Direct Aspiration; AA | 0.01 | Y | Y | Y | 6671 |  |  | 5/8/2000 4:03:53 PM | 1999-04-22 |  | NPEPPIN |  | Concentration | mg/L | Y | (fl. conc.) |  | Potassium Extractable (fl. conc.) |  | Metals - Extractable; | Metals - Extractable; EMS code: K--E | Chemical | Extractable | LAB
K--L | mg/L | Potassium Leachable | MEND | MEND (mod) EPA 300.1 (mod) | MEND (mod) EPA 300.1 (mod). Inorganic anions are analyzed by obtaining an extract from a soil sample that has been added in a 3:1 ratio of deionized water to soil; then leached and the extract filtered prior to analysis; which is then analyzed by Ion Chromatography with conductivity and/or UV detection. | 0.05 | Y | Y | Y | 0 |  |  |  | 2024-01-29 | 1/29/2024 10:44:03 AM | RDGIBSON | RDGIBSON | Concentration | mg/L | Y | (fl. conc.) |  | Potassium Leachable (fl. conc.) |  | Nutrients; | Nutrients; EMS code: K--L | Chemical |  | LAB
K--S | mg/L | Potassium-Soluble ion content | F032 | Filt.,Acid,ICPMS | There is no acid digestion required for this analysis. A water sample is collected and preserved in 0.2 percent nitric acid. To analyse the extractable metals the sample is shaken, allowed to sit for at least twenty four hours, and a decanted portion is introduced to the spectrometer. When the dissolved metals are required, the samples must first be filtered through a 0.45 micron cellulose acetate filter paper and then followed by preservation in 0.2 percent nitric acid. It is important to note that different laboratories may have very different detection limits for some variables analyzed by ICPMS. | 0.01 | Y | Y | Y | 337 | 3/1/2002 1:15:00 PM | 2/17/2005 | 3/17/2005 1:07:53 AM | 2004-02-10 |  | UMULLER |  | Concentration | mg/L | Y | (fl. conc.) |  | Potassium-Soluble ion content (fl. conc.) |  | Nutrients; | Nutrients; EMS code: K--S | Chemical | Soluble | LAB
K--S | mg/L | Potassium-Soluble ion content | EQ01 | EQUIS (mg/L) | EQUIS (mg/L) | 0.01 | N | Y | Y | 125 |  |  | 10/22/1996 11:12:48 PM | 1996-04-12 | 6/3/1998 6:51:36 PM | EMS |  | Concentration | mg/L | Y | (fl. conc.) |  | Potassium-Soluble ion content (fl. conc.) |  | Nutrients; | Nutrients; EMS code: K--S | Chemical | Soluble | LAB
K--S | mg/L | Potassium-Soluble ion content | 5073 | Aut.Col-Flame Phot/MerTh | Aut.Col-Flame Phot/MerTh | 0.1 | Y | Y | Y | 2369 | 4/6/2005 | 4/11/2008 | 5/28/2008 2:08:44 AM | 1996-04-12 | 6/3/1998 6:51:36 PM | EMS |  | Concentration | mg/L | Y | (fl. conc.) |  | Potassium-Soluble ion content (fl. conc.) |  | Nutrients; | Nutrients; EMS code: K--S | Chemical | Soluble | LAB
K--S | mg/L | Potassium-Soluble ion content | 5071 | Ion Chromatog -Cation I | (a) For the Soluble Particulates Pkg. a portion of filtrate from the Particulate - Total procedure made up to 500 mL is analyzed by IC. Intermediate results are reported in mg/L. Loading results are calculated and reported with units of mg/dm2/d.(b) For precipitation samples, all samples are filtered through a 0.22 um fritted glass filter prior to injection. The cations of interest are separated through a cation "guard" and cation "separator" column with a methane sulphonic acid eluent. After the cations are separated they exit at various times from the column in a background of eluent. A self regenerating membrane suppressor is attached to the end of the column to neutralize the acid before detection and form the corresponding hydroxide. Note that ammonium is reported as NH4 not as N. (c) For HiVol samples, discs cut from the HiVol filter are extracted with DI water and the soluble anions are analyzed by IC. Intermediate results are reported in units of mg/L and HiVol loading results are reported as ug/m3. (*) Note that ammonium is reported as NH4 not as N | 0.01 | Y | Y | Y | 3679 | 4/30/1997 3:50:00 PM | 1/7/2002 2:33:00 PM | 6/24/2003 10:43:21 AM | 1996-04-12 | 6/3/1998 6:51:45 PM | EMS |  | Concentration | mg/L | Y | (fl. conc.) |  | Potassium-Soluble ion content (fl. conc.) |  | Nutrients; | Nutrients; EMS code: K--S | Chemical | Soluble | LAB
K--S | mg/L | Potassium-Soluble ion content | E6020B | EPA 6020B | See EPA 6020B and CSSS Ch. 15 | 4 | Y | Y | Y | 0 |  |  |  | 2024-02-02 |  | EMS |  | Concentration | mg/L | Y | (fl. conc.) |  | Potassium-Soluble ion content (fl. conc.) |  | Nutrients; | Nutrients; EMS code: K--S | Chemical | Soluble | LAB
K--S | mg/dm2/d | Potassium-Soluble ion content | EQ15 | EQUIS (mg/dm2/d) | EQUIS (mg/dm2/d) | 0.001 | N | Y | Y | 27 |  |  | 10/22/1996 6:09:53 PM | 1996-04-12 | 6/3/1998 6:51:36 PM | EMS |  | MassFlux | mg/dm²/d | Y | (mass flux) |  | Potassium-Soluble ion content (mass flux) |  | Nutrients; | Nutrients; EMS code: K--S | Chemical | Soluble | LAB
K--S | mg/L | Potassium-Soluble ion content | E6010B | EPA 6010B/CSSS 15 | See CSSS CH15/EPA 6010B/AER D50. INDUCTIVELY COUPLED PLASMA-ATOMIC EMISSION SPECTROMETRY See EPA 6010B and CSSS Ch. 15. https://www.epa.gov/sites/default/files/documents/6010b.pdf | 5 | Y | Y | Y | 0 |  |  |  | 2024-01-30 |  | EMS |  | Concentration | mg/L | Y | (fl. conc.) |  | Potassium-Soluble ion content (fl. conc.) |  | Nutrients; | Nutrients; EMS code: K--S | Chemical | Soluble | LAB
K--S | ug/g | Potassium-Soluble ion content | CAL2 | Calculation (ug/g) | Calculation (ug/g) | 2 | Y | Y | Y | 0 |  |  |  | 2024-02-02 |  | EMS |  | MassPerMass | mg/kg | Y | (mass conc.) |  | Potassium-Soluble ion content (mass conc.) |  | Nutrients; | Nutrients; EMS code: K--S | Chemical | Soluble | LAB
K--T | mg/L | Potassium Total | 1349 | Tot Metals in Seawater by CRC ICPMS (BC) | Tot Metals in Seawater by CRC ICPMS (BC) | 1 | Y | Y | Y | 98 | 8/16/2019 | 12/5/2021 | 12/13/2021 12:34:16 PM | 2019-09-30 |  | KATFRASE |  | Concentration | mg/L | Y | (fl. conc.) |  | Potassium Total (fl. conc.) |  | Metals - Total; | Metals - Total; EMS code: K--T | Chemical | Total | LAB
K--T | mg/L | Potassium Total | F121 | Acid,ICPMS | Acidified, no digestion, ICPMS | 0.01 | Y | Y | Y | 6950 | 10/18/2007 | 8/16/2024 | 8/26/2024 8:06:03 AM | 2002-07-17 | 7/16/2012 3:33:53 PM | UMULLER | RDGIBSON | Concentration | mg/L | Y | (fl. conc.) |  | Potassium Total (fl. conc.) |  | Metals - Total; | Metals - Total; EMS code: K--T | Chemical | Total | LAB
K--T | ug/g wet | Potassium Total | X552 | Metals in Biota (wet), total, acid digestion, ICP-SFMS | One gram of wet, homogenized biota is weighed in a microwave TFM vessel, and 8mL of Nitric Acid and 1mL of Hydrogen Peroxide are added. The sample is digested at 200 degrees C for 15 minutes in a closed vessel high pressure microwave oven. The sample is cooled, the volume adjusted to 100mL and introduced to an inductively coupled argon plasma-sector field mass spectrometer (ICP-SFMS) by pneumatic nebulization. Each respective element is measured at a specific mass to charge ratio, m/z value expressed in amu (atomic mass units), and at a specified instrument resolution. LR: Low Resolution ~300, MR: Mid Resolution ~4000. Ref: 02-2704, National Laboratory for Environmental Testing, CCIW, Environment Canada, Burlington. Requested April 2003. | 4 | Y | Y | Y | 37 | 8/9/2016 | 10/16/2023 | 11/28/2023 2:21:32 PM | 2015-03-11 |  | RDGIBSON |  | MassPerMass | µg/g | Y | (mass conc.) | (wet) | Potassium Total (wet) (mass conc.) |  | Metals - Total; | Metals - Total; EMS code: K--T | Chemical | Total | LAB
K--T | mg/L | Potassium Total | F048 | Acid dig,conc'n,nebul.,ICP-OES | This method provides an automated procedure for the Inductively Coupled Argon Plasma Optical Emission Spectrophotometric (ICP-OES) determination and quantitation of trace amounts of metals in surface, ground and precipitation waters after a manual acid digestion and concentration procedure and ultrasonic nebulisation. Organically-bound metals and metal complexes are converted to metal ions by digestion with nitric acid [method details says HCl]. To increase the signal detected by the ICP, the samples are concentrated 10 times. The liquid sample is pumped onto the piezoelectric transducer, where the ultrasonic action efficiently converts it to a fine, dense aerosol. The nebulizer carrier gas flow then transports the wet aerosol through a heated U-tube, where the solvent is vaporized. Subsequently, most of the solvent vapour is removed from the aerosol stream by passing the stream through a chilled condenser. The result is a dry, analyte-laden aerosol which is introduced to the plasma. There, the high temperature desolvates the aerosol and dissociates the sample material into individual atoms and ions, which are excited to emit light at wave lengths characteristic of the elements present. Dissolved samples are field filtered through 0.45 micron cellulose acetate filter prior to acid preservation. Samples for extractable analysis are acidified with 2 mL concentrated HNO3 per litre and allowed to stand a minimum of 24 hours before analysis. Clients providing these samples are advised to preserve samples at the collection site. | 2 | Y | Y | Y | 231 | 11/20/1997 9:32:00 AM | 7/17/2012 7:03:00 AM | 7/26/2012 12:28:00 PM | 1998-11-13 |  | NPEPPIN |  | Concentration | mg/L | Y | (fl. conc.) |  | Potassium Total (fl. conc.) |  | Metals - Total; | Metals - Total; EMS code: K--T | Chemical | Total | LAB
K--T | mg/L | Potassium Total | F082 | Aliquot,HNO3/HCL dig,ICPMS | Aliquot microwave digested with 3:1 HNO3/HCL. Analysis by ICPMS. For HiVol samples preliminary result are reported in mg/L and loading results are reported in ug/m3 | 0.05 | Y | Y | Y | 26422 | 1/24/2011 | 8/23/2024 | 8/26/2024 4:53:05 PM | 2015-04-15 |  | RDGIBSON |  | Concentration | mg/L | Y | (fl. conc.) |  | Potassium Total (fl. conc.) |  | Metals - Total; | Metals - Total; EMS code: K--T | Chemical | Total | LAB
K--T | mg/L | Potassium Total | 0042 | GI HNO3 Dig: ICP | Sample aliquot taken from unfiltered unpreserved General Ions bottle; HNO3 digestion; ICP analysis. | 0.4 | Y | Y | Y | 6662 | 9/20/1996 | 7/5/2019 3:21:00 PM | 7/12/2019 6:49:13 PM | 1996-04-12 | 6/3/1998 6:51:36 PM | EMS |  | Concentration | mg/L | Y | (fl. conc.) |  | Potassium Total (fl. conc.) |  | Metals - Total; | Metals - Total; EMS code: K--T | Chemical | Total | LAB
K--T | % | Potassium Total | MS61 | ALS Minerals ME-MS61 | ALS Minerals ME-MS61. A prepared sample is digested with HNO3; HF; and HClO4 in a graphite heating block to near dryness. After cooling; the sample is boiled in diluted HCl. Bulk to volume with deionized water; mixed and analyzed by ICP-AES and ICP-MS for all elements. | 0.01 | Y | Y | Y | 0 |  |  |  | 2024-01-29 |  | EMS |  | DimensionlessRatio | % | N | (ratio) |  | Potassium Total (ratio) |  | Metals - Total; | Metals - Total; EMS code: K--T | Chemical | Total | LAB
K--T | mg/L | Potassium Total | MWID | Metals in Water by ICPMS (digested) | Metals in Water by ICPMS (digested) Based on EPA 200.8 | 1 | Y | Y | Y | 36495 |  |  | 7/31/2024 1:46:41 PM | 2009-12-24 |  | RDGIBSON |  | Concentration | mg/L | Y | (fl. conc.) |  | Potassium Total (fl. conc.) |  | Metals - Total; | Metals - Total; EMS code: K--T | Chemical | Total | LAB
K--T | mg/dm2/d | Potassium Total | DUST | HNO3 dig.(microwave or hot block),ICPMS | HNO3 dig.(microwave or hot block);ICPMS |  | Y | Y | Y | 420 | 7/19/2016 1:11:00 PM | 7/31/2024 | 7/31/2024 11:23:12 AM | 2008-02-07 |  | RDGIBSON |  | MassFlux | mg/dm²/d | Y | (mass flux) |  | Potassium Total (mass flux) |  | Metals - Total; | Metals - Total; EMS code: K--T | Chemical | Total | LAB
K--T | ug/g | Potassium Total | FB04 | Freeze Dry,Hom,Acid Dig,ICP | Sample moisture is determined by weighing samples before and after freeze-drying. Freeze-dried samples are ground in a Braun blender before being weighed into Teflon PFA liners and digested with concentrated HNO3, concentrated HCl, and 50% H2O2 in a microwave oven; which results in a complete sample dissolution. Samples are collected and volumized with Type I water. Analysis is performed via Inductively Coupled Argon Plasma Atomic Emission Spectrometry (ICP-AES) (simultaneous multielement analysis), Graphite Furnace Atomic Absorption Spectrometry (GF-AAS) (low level Cd and Pb analysis), Cold Vapor Atomic Fluorescence Spectrometry (CV-AFS) (Mercury analysis) and/or Flame Atomic Absorption Spectrometry (F-AAS) (K analysis). | 8 | Y | Y | Y | 236 | 8/22/1996 6:11:00 PM | 2/1/2002 2:30:00 PM | 6/23/2003 1:15:31 PM | 1996-06-19 | 6/3/1998 6:51:44 PM | MPAK |  | MassPerMass | mg/kg | Y | (mass conc.) |  | Potassium Total (mass conc.) |  | Metals - Total; | Metals - Total; EMS code: K--T | Chemical | Total | LAB
K--T | ug/g | Potassium Total | FS82 | PBM,dry60C,10mesh,SALM-ICP-MS | PBM;dry60C;10mesh;SALM-ICP-MS | 100 | Y | Y | Y | 2861 | 5/21/2002 | 7/24/2024 | 7/29/2024 5:17:09 PM | 2002-11-12 |  | UMULLER |  | MassPerMass | mg/kg | Y | (mass conc.) |  | Potassium Total (mass conc.) |  | Metals - Total; | Metals - Total; EMS code: K--T | Chemical | Total | LAB
K--T | mg/L | Potassium Total | X285 | HNO3 Dig:ICP Analysis | HNO3 Dig:ICP Analysis | 0.1 | Y | Y | Y | 633 | 8/10/1993 | 10/5/1993 | 4/30/2004 11:32:38 AM | 1998-05-22 | 6/3/1998 6:51:49 PM | NPEPPIN |  | Concentration | mg/L | Y | (fl. conc.) |  | Potassium Total (fl. conc.) |  | Metals - Total; | Metals - Total; EMS code: K--T | Chemical | Total | LAB
K--T | ug/g | Potassium Total | X230 | Hom,HNO3/HCl,ICP | The sample is dried, ground to pass 100 mesh, and digested with a nitric-hydrochloric acid mixture to solubilize the solid matter and to remove the organic material by oxidation and volatilization. The completed digests are submitted to metals instrumentation personnel for metals analysis by Inductively Coupled Atomic Emission Spectrophotometry (ICP-AES), or hydride analysis by atomic absorption. | 10 | Y | Y | Y | 41 | 2/12/2002 | 12/30/2005 | 2/3/2006 1:13:13 AM | 1998-11-05 |  | NPEPPIN |  | MassPerMass | mg/kg | Y | (mass conc.) |  | Potassium Total (mass conc.) |  | Metals - Total; | Metals - Total; EMS code: K--T | Chemical | Total | LAB
K--T | mg/L | Potassium Total | F038 | HNO3/HCl dig,ICP-AES | Dissolved and extractable water samples are ready for analysis once collected. Total samples are autoclaved with HNO3 and hydrochloric acid (HCI) in capped Teflon vials. Analysis is performed via Inductively Coupled Argon Plasma Atomic Emission Spectrometry (ICP-AES)(simultaneous multielement analysis), Graphite Furnace Atomic Absorption Spectrometry (GF-AAS) (low level analysis) and/or Flame Atomic Absorption Spectrometry (F-AAS). | 0.1 | Y | Y | Y | 7065 | 8/1/1996 6:04:00 PM | 5/6/2008 | 9/11/2019 8:47:42 AM | 1997-05-08 | 6/3/1998 6:51:46 PM | MPAK |  | Concentration | mg/L | Y | (fl. conc.) |  | Potassium Total (fl. conc.) |  | Metals - Total; | Metals - Total; EMS code: K--T | Chemical | Total | LAB
K--T | mg/L | Potassium Total | MSEA | Metals in Seawater (undigested) by (CRC) ICP-MS | This method describes the multi-elemental determination of trace elements by Collision/Reaction Cell (CRC) ICP-MS. The method measures ions produced by a radio-frequency inductively coupled plasma. Analyte species originating in a liquid are pneumatically nebulized and the resulting aerosol transported by argon gas into the plasma torch. The ions produced are entrained in the plasma gas and extracted, by means of a differentially pumped vacuum interface, into a mass spectrometer. The ions flow through a Collision/Reaction Cell (CRC) where depending on the mode of analysis either the ions passes through the cell unhindered (no gas - classical ICP-MS mode), or collisions/reactions chemistry takes place between the ions and a neutral gas such as hydrogen or helium. Any polyatomic ions which give interference on the analytes are removed. The ions are then sorted according to their mass-to-charge ratios by a quadrupole mass spectrometer. The ions are then sorted according to their mass-to-charge ratios by a quadrupole mass spectrometer. The ions transmitted through the quadrupole are quantified by a channel electron multiplier and the ion information processed by a data handling system. Although most polyatomic interference is removed, any additional interference must be assessed and valid corrections applied or the data flagged to indicate problems. Interference correction must include compensation for background ions contributed by the plasma gas, reagents, and constituents of the sample matrix. Instrument drift as well as suppressions or enhancements of instrument response caused by the sample matrix must be corrected for by the use of internal standards. | 1 | Y | Y | Y | 441 | 8/11/2009 11:57:00 PM | 4/29/2024 | 6/7/2024 10:03:16 AM | 2009-09-08 | 5/26/2023 3:37:55 PM | RDGIBSON | EMS | Concentration | mg/L | Y | (fl. conc.) |  | Potassium Total (fl. conc.) |  | Metals - Total; | Metals - Total; EMS code: K--T | Chemical | Total | LAB
K--T | ug/g | Potassium Total | MTID | MET-DRY-CCMS-VA | Metals in Tissue by CRC ICPMS (DRY) Methodology: This method is conducted following British Columbia Lab Manual method "Metals in Animal Tissue and Vegetation (Biota) - Prescriptive". Tissue samples are homogenized and sub-sampled prior to hotblock digestion with nitric and hydrochloric acids, in combination with addition of hydrogen peroxide. Instrumental analysis is by collision cell inductively coupled plasma - mass spectrometry (modified from EPA Method 6020A). Method Limitation: This method employs a strong acid/peroxide digestion, and is intended to provide a conservative estimate of bio-available metals. Near complete recoveries are achieved for most toxicologically important metals, but elements associated with recalcitrant minerals may be only partially recovered. | 20 | Y | Y | Y | 122 | 8/13/2014 | 9/8/2022 | 10/26/2022 7:14:15 AM | 2015-03-09 | 3/11/2015 4:12:05 PM | RDGIBSON | EMS | MassPerMass | mg/kg | Y | (mass conc.) |  | Potassium Total (mass conc.) |  | Metals - Total; | Metals - Total; EMS code: K--T | Chemical | Total | LAB
K--T | mg/L | Potassium Total | F033 | Acid Dig,Centrifuge,ICPMS | For the determination of total recoverable elements in water or waste water, take 100 ml aliquot from a well mixed, acid preserved, sample containing not more than 0.25% (w/v) total solids and transfer to a 250 ml Griffin beaker. (If total solids are greater than 0.25% reduce the size of the aliquot by a proportionate amount). Add 1 ml of conc. nitric acid and 0.5 ml conc. hydrochloric acid. Heat on a hot plate until the volume has been reduced to 15-20ml, ensuring that the sample does not boil. Allow to cool and transfer to a 50 ml volumetric flask. Dilute to volume with ASTM type I water. Centrifuge the sample or allow to stand overnight to separate insoluble material. Prior analysis, pipette 20 ml into a 50 ml volumetric flask and dilute to volume with ASTM type I water. Analysis should be performed within seventy two hours of the completed preparation. Note that different laboratories may have very different field preservation, digestion procedures and detection limits for this method, so these should be documented. | 0.01 | Y | Y | Y | 26538 | 5/7/1998 | 8/22/2024 3:24:00 AM | 8/25/2024 9:47:36 PM | 1998-05-14 | 6/3/1998 6:51:49 PM | NPEPPIN |  | Concentration | mg/L | Y | (fl. conc.) |  | Potassium Total (fl. conc.) |  | Metals - Total; | Metals - Total; EMS code: K--T | Chemical | Total | LAB
K--T | mg/L | Potassium Total | X145 | Dig, ICP | Dig; ICP | 0.1 | Y | Y | Y | 6812 | 9/5/2001 | 9/22/2016 | 8/26/2024 2:41:32 PM | 1996-04-12 | 6/3/1998 6:51:45 PM | EMS |  | Concentration | mg/L | Y | (fl. conc.) |  | Potassium Total (fl. conc.) |  | Metals - Total; | Metals - Total; EMS code: K--T | Chemical | Total | LAB
K--T | % | Potassium Total | 6041 | Aqua Regia by ICPMS | These were sediment/soils digested with Aqua-Regia (Nitric Acid & Hydrochloric Acids) and then analysis by ICPMS. Results would be reported in mg/kg and/or %. | 0.005 | Y | Y | Y | 92 | 10/13/2019 | 7/27/2024 | 7/29/2024 9:13:08 AM | 2019-09-06 | 9/6/2019 2:20:27 PM | KATFRASE | KATFRASE | DimensionlessRatio | % | N | (ratio) |  | Potassium Total (ratio) |  | Metals - Total; | Metals - Total; EMS code: K--T | Chemical | Total | LAB
K--T | ug/g | Potassium Total | ICPM | Metals in Soil by CRC ICPMS | This analysis is carried out using procedures from CSR Analytical Method: "Strong Acid Leachable Metals (SALM) in Soil", BC Ministry of Environment, 26 June 2009, and procedures adapted from EPA Method 200.2. The sample is manually homogenized, dried at 60 degrees Celsius, sieved through a 63 um (230 mesh) sieve, and a representative subsample of the dry material is weighed. The sample is then digested at 95 degrees Celsius for 2 hours by block digester using concentrated nitric and hydrochloric acids. Instrumental analysis is by collision cell inductively coupled plasma - mass spectrometry (modifed from EPA Method 6020A). Method Limitation: This method is not a total digestion technique. It is a very strong acid digestion that is intended to dissolve those metals that may be environmentally available. By design, elements bound in silicate structures are not normally dissolved by this procedure as they are not usually mobile in the environment. | 100 | Y | Y | Y | 103 | 9/19/2017 | 9/24/2023 | 1/12/2024 11:16:52 AM | 2018-01-09 |  | JKROGH |  | MassPerMass | mg/kg | Y | (mass conc.) |  | Potassium Total (mass conc.) |  | Metals - Total; | Metals - Total; EMS code: K--T | Chemical | Total | LAB
K--T | mg/L | Potassium Total | X073 | FA,HNO3,Direct Flame,AA | Samples are field-acidified with HNO3; HNO3 digestion; analysis by AA. | 0.01 | Y | Y | Y | 130 |  |  | 1/25/2007 4:06:11 PM | 1996-04-12 | 6/3/1998 6:51:36 PM | EMS |  | Concentration | mg/L | Y | (fl. conc.) |  | Potassium Total (fl. conc.) |  | Metals - Total; | Metals - Total; EMS code: K--T | Chemical | Total | LAB
K--T | mg/L | Potassium Total | X170 | ICP | This code is a catch-all for a large variety of ICP methods including water and also solids digested into solutions as an intermediate result. Different procedures will have greatly different detection limits. (a) For air particulates, preliminary results are reported in mg/L and loading results as mg/dm2/d. (b) For Hi Vol samples preliminiary results are same as reported in mg/L and loading results as ug/m3. (c) Note that for Rain samples there is no filtration nor digestion steps. | 2 | Y | Y | Y | 6440 | 3/28/1996 6:03:00 PM | 11/7/2018 | 12/8/2023 4:30:26 PM | 1996-04-12 | 8/15/2019 1:20:17 PM | EMS | KATFRASE | Concentration | mg/L | Y | (fl. conc.) |  | Potassium Total (fl. conc.) |  | Metals - Total; | Metals - Total; EMS code: K--T | Chemical | Total | LAB
K--T | ug/g wet | Potassium Total | X560 | Metals in Biota (wet), total, acid digestion, ICP-MS | Five grams of wet, homogenized biota is weighed in a centrifuge tube, and 2.5 mL of Nitric Acid and 2.5 mL of Hydrochloric acid are added. The sample is digested at room temperature for one hour and then digested at 90-95 degrees C in a water bath or heat block for one more hour. The sample is cooled and the volume is made to 50 mL with DI. The digestate is then diluted at ten fold and introduced to a collision/reaction cell inductively coupled argon plasma mass spectrometer (CRC ICPMS) by pneumatic nebulization. Each respective element is measured at a specific mass to charge ratio, m/z value expressed in amu (atomic mass units). - Maxxam Analytics (BC). Added to EMS Dec 10, 2008. | 10 | Y | Y | Y | 541 | 11/15/2008 4:00:00 AM | 10/30/2023 | 11/28/2023 2:21:33 PM | 2008-12-11 | 3/11/2015 4:17:10 PM | RDGIBSON | EMS | MassPerMass | µg/g | Y | (mass conc.) | (wet) | Potassium Total (wet) (mass conc.) |  | Metals - Total; | Metals - Total; EMS code: K--T | Chemical | Total | LAB
K--T | mg/L | Potassium Total | METN | MET-T-NP-CCMS-VA | Ultra trace metals in water are analyzed by CRC ICPMS, based on US EPA Method 6020B (July 2014). The detection limits provided can only be met for undigested samples. This procedure is intended for colorless, non-turbid, acid-preserved water samples (i.e. pristine water samples), having turbidity < 1 NTU and no odor. Where turbidity exceeds 1 NTU, and/or the sample is colored and has an odor, results may be biased low compared to true Total Metals concentrations. ALS recommends that turbidity analysis be requested on samples submitted for this test to aid with interpretation of results. Where turbidity is <1NTU, undigested metals are equivalent to total metals concentrations. | 0.02 | Y | Y | Y | 3028 | 11/14/2017 | 8/19/2024 | 8/26/2024 1:03:06 PM | 2019-08-20 |  | KATFRASE |  | Concentration | mg/L | Y | (fl. conc.) |  | Potassium Total (fl. conc.) |  | Metals - Total; | Metals - Total; EMS code: K--T | Chemical | Total | LAB
K--T | ug/g wet | Potassium Total | BPSY | MET-WET-BPSY-CCMS-VA | Metals in Animal Tissue and Vegetation (Biota) - Prescriptive. EPA 200.3/6020B This method is conducted following British Columbia Lab Manual method "Metals in Animal Tissue and Vegetation (Biota) - Prescriptive". Tissue samples are digested with nitric and hydrochloric acids; in combination with addition of hydrogen peroxide. Instrumental analysis is by collision cell inductively coupled plasma - mass spectrometry (modified from EPA Method 6020B). | 4 | Y | Y | Y | 2 | 11/7/2019 | 11/7/2019 | 6/10/2021 1:57:10 PM | 2021-06-04 |  | EMS |  | MassPerMass | µg/g | Y | (mass conc.) | (wet) | Potassium Total (wet) (mass conc.) |  | Metals - Total; | Metals - Total; EMS code: K--T | Chemical | Total | LAB
K--T | mg/L | Potassium Total | EQ01 | EQUIS (mg/L) | EQUIS (mg/L) | 0 | N | Y | Y | 1202 |  |  | 10/24/1996 3:16:20 PM | 1996-04-12 | 6/3/1998 6:51:36 PM | EMS |  | Concentration | mg/L | Y | (fl. conc.) |  | Potassium Total (fl. conc.) |  | Metals - Total; | Metals - Total; EMS code: K--T | Chemical | Total | LAB
K--T | % | Potassium Total | MS41 | ALS Minerals ME-MS41 | ALS Minerals ME-MS41. A prepared sample is digested with aqua regia in a graphite heating block. After cooling; the resulting solution is diluted to with deionized water; mixed and analyzed by ICP-AES. Samples are then analyzed by ICP-MS for the remaining suite of elements. | 0.01 | Y | Y | Y | 0 |  |  |  | 2024-01-29 |  | EMS |  | DimensionlessRatio | % | N | (ratio) |  | Potassium Total (ratio) |  | Metals - Total; | Metals - Total; EMS code: K--T | Chemical | Total | LAB
K--T | ug/g wet | Potassium Total | MTIW | MET-WET-CCMS-MID-VA | Metals in Tissue by CRC ICPMS (WET) Methodology: This method is conducted following British Columbia Lab Manual method "Metals in Animal Tissue and Vegetation (Biota) - Prescriptive". Tissue samples are homogenized and sub-sampled prior to hotblock digestion with nitric and hydrochloric acids, in combination with addition of hydrogen peroxide. Instrumental analysis is by collision cell inductively coupled plasma - mass spectrometry (modified from EPA Method 6020A). Method Limitation: This method employs a strong acid/peroxide digestion, and is intended to provide a conservative estimate of bio-available metals. Near complete recoveries are achieved for most toxicologically important metals, but elements associated with recalcitrant minerals may be only partially recovered. | 4 | Y | Y | Y | 24 | 11/25/2018 | 12/6/2018 | 7/21/2021 2:55:07 PM | 2018-12-24 |  | EMS |  | MassPerMass | µg/g | Y | (mass conc.) | (wet) | Potassium Total (wet) (mass conc.) |  | Metals - Total; | Metals - Total; EMS code: K--T | Chemical | Total | LAB
K--T | ug/g | Potassium Total | X551 | Metals in Biota (dry), total, acid digestion, ICP-SFMS | 0.1 gram of dried, homogenized biota is weighed in a microwave TFM vessel, and 8mL of Nitric Acid and 1mL of Hydrogen Peroxide are added. The sample is digested at 200 degrees C for 15 minutes in a closed vessel high pressure microwave oven. The sample is cooled, the volume adjusted to 100mL and introduced to an iductively coupled argon plasma-sector field mass spectrometer (ICP-SFMS) by pneumatic nebulization. Each respective element is measured at a specific mass to charge ratio, m/z value expressed in amu (atomic mass units), and at a specified instrument resolution. LR: Low Resolution ~300, MR: Mid Resolution ~4000 and HR: Hich Resolution ~10000. Ref: Method number 02-2705, National Laboratory for Environmental Testing, CCIW, Environment Canada, Burlington. Requested April 2003. | 20 | Y | Y | Y | 49 | 8/11/2015 | 10/30/2023 | 11/28/2023 2:21:33 PM | 2015-03-11 |  | RDGIBSON |  | MassPerMass | mg/kg | Y | (mass conc.) |  | Potassium Total (mass conc.) |  | Metals - Total; | Metals - Total; EMS code: K--T | Chemical | Total | LAB
K--T | mg/L | Potassium Total | X349 | FA:HNO3:ICP | Samples are field-acidified with HNO3; HNO3 digestion; analysis by ICP/AES. | 0.1 | Y | Y | Y | 30 |  |  | 10/30/2014 8:31:17 AM | 1998-05-22 | 6/3/1998 6:51:49 PM | NPEPPIN |  | Concentration | mg/L | Y | (fl. conc.) |  | Potassium Total (fl. conc.) |  | Metals - Total; | Metals - Total; EMS code: K--T | Chemical | Total | LAB
K--T | ug/g | Potassium Total | PB02 | PBM,HNO3/HCl dig,dry60C,100mesh, | Performance Based Method. Samples are dried at 60 C and sieved using a .149 mm (100 mesh) sieve. Solid samples are digested with the combination of nitric and hydrochloric acids with a 1:1 ratio of acids. For environmental purposes, strong acid leaches are used routinely since they have the advantage of speed, low cost, and superio reproducibility. All metals trapped within the refractory matrix are unlikely to have environmental consequences. Instrumental analysis of extracts produced using this digestion method can be performed by a variety of techniques. |  | Y | Y | Y | 10 | 9/13/2001 | 9/13/2001 | 11/6/2002 3:52:36 PM | 2001-07-18 |  | NPEPPIN |  | MassPerMass | mg/kg | Y | (mass conc.) |  | Potassium Total (mass conc.) |  | Metals - Total; | Metals - Total; EMS code: K--T | Chemical | Total | LAB
K--T | mg/L | Potassium Total | 0040 | HNO3 Dig: ICP Analysis | HNO3 Dig: ICP Analysis | 0.4 | Y | Y | Y | 1620 | 10/6/2010 10:03:00 AM | 7/25/2012 7:11:00 AM | 9/26/2012 1:06:57 PM | 1996-04-12 | 6/3/1998 6:51:45 PM | EMS |  | Concentration | mg/L | Y | (fl. conc.) |  | Potassium Total (fl. conc.) |  | Metals - Total; | Metals - Total; EMS code: K--T | Chemical | Total | LAB
K--T | ug/g | Potassium Total | FS01 | D&G,Nitric/HCl:ICP | Samples are dried at 60 C, then either sieved to -100mesh (150 um) and rolled to homogenize, or finely ground with a mortar and pestle or a shatter box laboratory mill and rolled (depending on the submitter's requirements). The samples are then weighed into Teflon digestion liners and digested with HNO3 and HCl in a microwave oven. Samples are cooled, volumized, and allowed to settle overnight before being decanted. Analysis is then performed via Inductively Coupled Argon Plasma Atomic Emission Spectrometry (ICP-AES) (Simultaneous multi element analysis). In ICP Spectrometry, the liquid sample is pumped to the nebulizer, where it is convertd into an aerosol. the nebulizer carrier gas flow than transports the wet aerosol past the impact bead, where the large droplets condense out and flow to the waste container. The remaining fine aerosol is carried into the plasma. There, the high temperature desolvates the aerosol and dissociates the sample material into individual atoms and ions, which are excited to emit light at wavelengths characteristic of the elements present. A diffraction grating in the spectrometer diffracts the various wavelengths of light through exit slits to the photo-multiplier tubes, which convert the light energy to electronic signals. | 20 | Y | Y | Y | 2642 | 8/30/1996 1:52:00 PM | 10/10/2007 | 5/24/2008 1:15:02 AM | 1996-06-19 | 6/3/1998 6:51:44 PM | MPAK |  | MassPerMass | mg/kg | Y | (mass conc.) |  | Potassium Total (mass conc.) |  | Metals - Total; | Metals - Total; EMS code: K--T | Chemical | Total | LAB
K--T | mg/L | Potassium Total | X561 | Total metals: FA, HNO3, HR-ICPMS (No Prep)Prep) | Ultra trace metals in water are analyzed by high resolution inductively coupled plasma mass spectrometry (HR-ICPMS), modified from US EPA 200.8, (Revision 5.5). The detection limits provided can only be met for undigested samples. This procedure is intended for colorless, non-turbid, acid-preserved water samples (i.e. pristine water samples), having turbidity < 1 NTU and no odor. Where turbidity exceeds 1 NTU, and/or the sample is colored and has an odor, results may be biased low compared to true Total Metals concentrations. ALS recommends that turbidity analysis be requested on samples submitted for this test to aid with interpretation of results. | 0.005 | Y | Y | Y | 2746 | 6/8/2015 | 10/5/2023 | 1/31/2024 2:37:41 PM | 2015-10-22 |  | RDGIBSON |  | Concentration | mg/L | Y | (fl. conc.) |  | Potassium Total (fl. conc.) |  | Metals - Total; | Metals - Total; EMS code: K--T | Chemical | Total | LAB
K2-O | % | Potassium (as K2O) | ICP06 | ALS Minerals ME-ICP06 | ALS Minerals ME-ICP06. A prepared sample is fused with a lithium tetraborate/lithium metaborate flux and poured into a graphite crucible and fused in a furnace at 1000°C. The resulting melt is then cooled and dissolved in 100 mL of 4% HNO3/2% HCl followed by ICP-AES. | 0.01 | Y | Y | Y | 0 |  |  |  | 2024-01-29 |  | EMS |  | DimensionlessRatio | % | N | (ratio) |  | Potassium (as K2O) (ratio) |  | Nutrients; | Nutrients; EMS code: K2-O | Chemical |  | LAB
KETP | mg/L | Ketoprofen CAS: 22071-15-4 | XMSP | HPLC and Tri. Quadpole MS/MS | EPA 1694 An aliquot of filtered sample containing 5% organic is injected directly. Each compound is separated by reversed phase HPLC. Analyzed by Electro Spray Ionization MS/MS detection using Triple Quadrupole MS/MS detector. Quantiitation with internal or external standard. |  | Y | Y | Y | 12 | 5/8/2020 | 6/4/2024 | 8/21/2024 3:32:03 PM | 2022-05-04 |  | EMS |  | Concentration | mg/L | Y | (fl. conc.) |  | Ketoprofen (fl. conc.) | 22071-15-4 | Pharmaceuticals and Personal Care Products (PPCP); | Pharmaceuticals and Personal Care Products (PPCP); EMS code: KETP | Chemical |  | LAB
L002 | mg/L | Lindane | MW03 | OCP1-LL-SF-ECD-VA | This analysis is carried out using procedures adapted from "Test Methods for Evaluating Solid Waste" SW-846, Methods 3510, 3630, 3660 and 8081, published by the United States Environmental Protection Agency (EPA). The procedure involves extraction of the entire water sample with dichloromethane. The extract then undergoes one or more of the following clean-up procedures (if required): silica gel clean-up and/or sulphur clean-up. The final extract is analyzed by dual capillary column gas chromatography with electron capture detection (GC/ECD). | 1e-05 | Y | Y | Y | 10 | 9/20/2016 | 9/21/2016 | 8/15/2019 2:42:34 PM | 2019-08-15 |  | KATFRASE |  | Concentration | mg/L | Y | (fl. conc.) |  | Lindane (fl. conc.) | 58-89-9 | Pesticides; | Pesticides; EMS code: L002 | Chemical |  | LAB
L002 | mg/L | Lindane | X268 | CH2Cl2 ext:GC/MS | CH2Cl2 ext:GC/MS |  | Y | Y | Y | 16 | 12/29/2021 | 9/2/2022 | 11/10/2022 9:54:07 AM | 2022-04-08 |  | EMS |  | Concentration | mg/L | Y | (fl. conc.) |  | Lindane (fl. conc.) | 58-89-9 | Pesticides; | Pesticides; EMS code: L002 | Chemical |  | LAB
L002 | ug/g | Lindane | E8081 | EPA 8081 | See methods EPA 3540, 3630, 3660, 8081 | 0.001 | Y | Y | Y | 9 | 3/18/2020 | 3/23/2020 | 2/22/2023 12:51:03 PM | 2022-07-29 |  | EMS |  | MassPerMass | mg/kg | Y | (mass conc.) |  | Lindane (mass conc.) | 58-89-9 | Pesticides; | Pesticides; EMS code: L002 | Chemical |  | LAB
L002 | mg/L | Lindane | ZN01 | DCM+Ac.Extr:GC/MS | DCM+Ac.Extr:GC/MS | 0.0005 | Y | Y | Y | 52 | 10/3/2002 | 3/29/2012 10:15:00 PM | 5/3/2012 9:27:51 AM | 1996-04-12 | 11/16/1999 4:06:27 PM | EMS | NPEPPIN | Concentration | mg/L | Y | (fl. conc.) |  | Lindane (fl. conc.) | 58-89-9 | Pesticides; | Pesticides; EMS code: L002 | Chemical |  | LAB
L002 | ug/g wet | Lindane | X220 | Hex/AcExt,Part,Chrom,GC | Hex/AcExt;Part;Chrom;GC | 0.0005 | Y | Y | Y | 71 |  |  | 10/24/1996 10:01:25 AM | 1996-04-12 | 6/3/1998 6:51:42 PM | EMS |  | MassPerMass | µg/g | Y | (mass conc.) | (wet) | Lindane (wet) (mass conc.) | 58-89-9 | Pesticides; | Pesticides; EMS code: L002 | Chemical |  | LAB
L002 | mg/L | Lindane | E660 | EPA 8270E (mod) | EPA 8270E (mod) Miscellaneous Pesticides by GC-MS | 3e-06 | Y | Y | Y | 2 | 9/14/2022 | 9/14/2022 | 12/8/2022 9:48:14 AM | 2022-12-06 |  | EMS |  | Concentration | mg/L | Y | (fl. conc.) |  | Lindane (fl. conc.) | 58-89-9 | Pesticides; | Pesticides; EMS code: L002 | Chemical |  | LAB
L002 | ug/g | Lindane | X213 | Hex/AcExt,Part,Chrom,GC | Hex/AcExt;Part;Chrom;GC | 0.0005 | Y | Y | Y | 187 | 4/20/2002 10:10:00 AM | 10/13/2002 10:30:00 AM | 2/26/2004 1:05:28 AM | 1996-04-12 | 6/3/1998 6:51:36 PM | EMS |  | MassPerMass | mg/kg | Y | (mass conc.) |  | Lindane (mass conc.) | 58-89-9 | Pesticides; | Pesticides; EMS code: L002 | Chemical |  | LAB
L002 | ug/g | Lindane | ZN03 | DCM+Ac.Extr:GC/MS | DCM+Ac.Extr:GC/MS EPA 8270E (mod) | 0.05 | Y | Y | Y | 19 | 1/17/2022 | 11/14/2022 | 1/11/2024 1:07:07 PM | 1996-04-12 | 11/16/1999 2:27:40 PM | EMS | NPEPPIN | MassPerMass | mg/kg | Y | (mass conc.) |  | Lindane (mass conc.) | 58-89-9 | Pesticides; | Pesticides; EMS code: L002 | Chemical |  | LAB
L003 | mg/L | Levopimaric Acid | P030 | Resin Acids | The water sample is acidified and extracted with dichloro-methane. The raw extract is concentrated and derivatized with diazomethane to produce the corresponding methyl ester derivatives. (Suitable substitute derivatives may be produced by other available techniques.) If required, the extracts are cleaned up by Florisil column chromatography. The derivatives are analyzed by gas chromatography with flame ionization detection. | 0.001 | Y | Y | Y | 481 | 8/26/2014 | 8/15/2017 | 8/15/2017 4:37:09 PM | 1996-04-12 | 6/3/1998 6:51:36 PM | EMS |  | Concentration | mg/L | Y | (fl. conc.) |  | Levopimaric Acid (fl. conc.) |  | Organics; | Organics; EMS code: L003 | Chemical |  | LAB
L003 | mg/L | Levopimaric Acid | P035 | DCM-DIAZO-FLOR | DCM-DIAZO-FLOR | 0.0005 | Y | Y | Y | 22 |  |  | 10/24/1996 12:44:24 PM | 1996-04-12 | 6/3/1998 6:51:36 PM | EMS |  | Concentration | mg/L | Y | (fl. conc.) |  | Levopimaric Acid (fl. conc.) |  | Organics; | Organics; EMS code: L003 | Chemical |  | LAB
L003 | mg/L | Levopimaric Acid | E677 | Resins & Fatty Acids are methylated then analyzed by GCMS | EPA 8270E; AB Env. 129; NCASI RA/FA-85.02 (mod) Resins & Fatty Acids are methylated then analyzed by GCMS. Total fatty acids is the sum of arachidic acid; behenic acid; lauric acid; lignoceric acid; linoleic acid; linolenic acid; myristic acid; oleic acid; palmitic acid; and stearic acid. Total resin acids is the sum of abietic acid; 12-chlorodehydroabietic acid; 14-chlorodehydroabietic acid; dehydroabietic acid; 12;14-dichlorodehydroabietic acid; isopimaric acid; levopimaric acid; neoabietic acid; palustric acid; pimaric acid and sandaracopimaric acid. Total resin acids (BC WWQG) is the sum of abietic acid; neoabietic acid; pimaric acid; isopimaric acid; and sandaracopimaric acid. | 0.001 | Y | Y | Y | 12 | 11/28/2022 | 6/15/2024 | 7/8/2024 11:10:02 AM | 2021-12-30 |  | EMS |  | Concentration | mg/L | Y | (fl. conc.) |  | Levopimaric Acid (fl. conc.) |  | Organics; | Organics; EMS code: L003 | Chemical |  | LAB
L003 | mg/L | Levopimaric Acid | F009 | pH 9,Diethyl Ether:HRGC/LRMS | Effluent samples are first adjusted to pH 9 and then extracted with Diethyl Ether. The extracts are rotary evaporated and the resulting residues, containing the Resin Acids and other solvent soluble organics are then redissolved in Methanol, transferred to small glass vials, methylated with Diazomethane then evaporated to dryness under Nitrogen. [Methylation Reaction R-COOH(Resin/Fatty Acids) + CH2N2(Diazomethane) = R-COOCH3(Methyl ester) + N2) The residue is re-dissolved in Methyl t-Butyl Ether (MTBE) and injected on a High Resolution Gas Chromatograph/Low Resolution Mass Spectrometer (HRGC/LRMS) system for analysis. LRMS data is acquired in Total Ion Mode (TIM) and quantitation is performed using selected mass ions. Positive identification of the various Resin Acids is made on the basis of a comparison of absolute retention times to those of the external standards and computer based mass spectral mass/intensity library matching routines. | 0.005 | Y | Y | Y | 26 | 1/31/2002 | 8/20/2002 | 4/1/2004 1:08:46 AM | 1997-01-13 | 6/3/1998 6:51:46 PM | MPAK |  | Concentration | mg/L | Y | (fl. conc.) |  | Levopimaric Acid (fl. conc.) |  | Organics; | Organics; EMS code: L003 | Chemical |  | LAB
L003 | mg/L | Levopimaric Acid | X239 | Gas Chromatography | Gas Chromatography | 0.001 | Y | Y | Y | 12 |  |  | 10/23/1996 11:56:37 PM | 1996-04-12 | 6/3/1998 6:51:36 PM | EMS |  | Concentration | mg/L | Y | (fl. conc.) |  | Levopimaric Acid (fl. conc.) |  | Organics; | Organics; EMS code: L003 | Chemical |  | LAB
L003 | mg/L | Levopimaric Acid | P033 | Resin&Fatty Acids,GC/MS | Resin&Fatty Acids;GC/MS | 0.001 | Y | Y | Y | 30 |  |  | 6/7/2024 5:30:18 PM | 1996-04-12 | 6/3/1998 6:51:45 PM | EMS |  | Concentration | mg/L | Y | (fl. conc.) |  | Levopimaric Acid (fl. conc.) |  | Organics; | Organics; EMS code: L003 | Chemical |  | LAB
L010 | ug/g | Linuron | ZN03 | DCM+Ac.Extr:GC/MS | DCM+Ac.Extr:GC/MS EPA 8270E (mod) | 0.05 | Y | Y | Y | 17 | 1/14/2022 | 11/2/2022 4:41:00 AM | 1/11/2024 1:07:07 PM | 2003-06-16 |  | UMULLER |  | MassPerMass | mg/kg | Y | (mass conc.) |  | Linuron (mass conc.) | 330-55-2 | Pesticides; | Pesticides; EMS code: L010 | Chemical |  | LAB
L010 | mg/L | Linuron | ZN01 | DCM+Ac.Extr:GC/MS | DCM+Ac.Extr:GC/MS | 0.0005 | Y | Y | Y | 3 | 10/17/2002 1:25:00 PM | 10/17/2002 1:25:00 PM | 3/24/2004 1:04:13 AM | 2003-06-16 |  | UMULLER |  | Concentration | mg/L | Y | (fl. conc.) |  | Linuron (fl. conc.) | 330-55-2 | Pesticides; | Pesticides; EMS code: L010 | Chemical |  | LAB
L010 | mg/L | Linuron | P015 | Phenoxyacid Herbicides by LC-MS/MS | Phenoxyacid Herbicides by LC-MS/MS | 0.0001 | Y | Y | Y | 68 | 12/6/2021 | 9/9/2022 | 12/8/2022 12:57:05 PM | 2021-12-14 |  | EMS |  | Concentration | mg/L | Y | (fl. conc.) |  | Linuron (fl. conc.) | 330-55-2 | Pesticides; | Pesticides; EMS code: L010 | Chemical |  | LAB
L420 | % | Loss on ignition @ 420?C | E205C | CSSS (1978) 3.81 (mod) | CSSS (1978) 3.81 (mod): Loss On Ignition (LOI) is determined by drying a portion of an air dried and ground sample at 105?C; then igniting at 420?C for 2 hours. The weight loss after ignition is reported as % loss on ignition. LOI is reported on a dry weight basis. LOI at 420?C can be considered an estimation of Organic Matter Content according to Alberta Agriculture (1988). | 1 | Y | Y | Y | 0 |  |  |  | 2024-02-27 |  | RDGIBSON |  | DimensionlessRatio | % | N | (ratio) |  | Loss on ignition @ 420?C (ratio) |  | Physical Properties; | Physical Properties; EMS code: L420 | Physical |  | LAB
LA-D | mg/L | Lanthanum Dissolved | METC | MET-D-NP-CCMS-VA | Ultra trace metals in water are analyzed by CRC ICPMS, based on US EPA Method 6020B (July 2014). This procedure is intended for pristine field-filtered acid-preserved water samples. The detection limits (LOR) for this test are based on lab instrumental analysis only, not including filtration. ALS-supplied field filtration equipment does not support these LOR. Therefore, because of the high probability of false positives due to filtration, it is strongly recommended that a filtration blank be analysed to aid in data interpretation. | 1e-05 | Y | Y | Y | 394 | 3/30/2020 | 4/9/2024 | 4/29/2024 4:53:06 PM | 2022-10-07 |  | EMS |  | Concentration | mg/L | Y | (fl. conc.) |  | Lanthanum Dissolved (fl. conc.) |  | Metals - Dissolved; | Metals - Dissolved; EMS code: LA-D | Chemical | Dissolved | LAB
LA-D | mg/L | Lanthanum Dissolved | X562 | Dissolved metals: FF, FA, HNO3, HR-ICPMS | Ultra trace metals in water are analyzed by high resolution inductively coupled plasma mass spectrometry (HR-ICPMS), based on US EPA Method 200.8, (Oct. 1999). This procedure is intended for pristine field-filtered acid-preserved water samples. The detection limits (LOR) for this test are based on lab instrumental analysis only, not including filtration. ALS-supplied field filtration equipment does not support these LOR. Therefore, because of the high probability of false positives due to filtration, it is strongly recommended that a filtration blank be analyzed to aid in data interpretation. | 1e-05 | Y | Y | Y | 3 | 5/29/2018 | 6/6/2018 | 9/17/2018 4:21:35 PM | 2018-06-06 |  | RDGIBSON |  | Concentration | mg/L | Y | (fl. conc.) |  | Lanthanum Dissolved (fl. conc.) |  | Metals - Dissolved; | Metals - Dissolved; EMS code: LA-D | Chemical | Dissolved | LAB
LA-D | mg/L | Lanthanum Dissolved | X556 | Metals in water, dissolved, "in-bottle" digestion, ICP-SFMS | The whole water sample is filtered through a 0.45 micron cellulose acetate filter paper and the filtrate is stored in a HDPE/Teflon bottle. The container is de-capped in a Class 100 workstation, acidified to 2% (v/v) HNO3 and 1% (v/v) HCl and recapped. The sample is "in-bottle" digested in a convection oven, set at 60 degrees C for 16 hours. The sample is cooled and introduced to an inductively coupled argon plasma-sector field mass spectrometer (ICP-SFMS) by pneumatic nebulization. Each respective element is measured at a specific mass to charge ratio, m/z value expressed in amu (atomic mass units), and at a specified intrument resolution. LR: Low Resolution ~300, MR: Mid Resolution ~4000 and HR: High Resolution ~10000. Ref: Method number 02-2003, National Laboratory for Environmental Testing, CCIW, Environment Canada, Burlington, Requested April 2003. | 1e-06 | Y | Y | Y | 4239 | 6/11/2003 | 12/6/2023 2:04:00 AM | 2/14/2024 9:05:32 AM | 2004-01-19 |  | UMULLER |  | Concentration | mg/L | Y | (fl. conc.) |  | Lanthanum Dissolved (fl. conc.) |  | Metals - Dissolved; | Metals - Dissolved; EMS code: LA-D | Chemical | Dissolved | LAB
LA-D | mg/L | Lanthanum Dissolved | F032 | Filt.,Acid,ICPMS | There is no acid digestion required for this analysis. A water sample is collected and preserved in 0.2 percent nitric acid. To analyse the extractable metals the sample is shaken, allowed to sit for at least twenty four hours, and a decanted portion is introduced to the spectrometer. When the dissolved metals are required, the samples must first be filtered through a 0.45 micron cellulose acetate filter paper and then followed by preservation in 0.2 percent nitric acid. It is important to note that different laboratories may have very different detection limits for some variables analyzed by ICPMS. | 1e-06 | Y | Y | Y | 235 | 5/9/2003 9:40:00 AM | 12/19/2023 | 1/2/2024 4:32:04 PM | 1997-10-20 | 6/3/1998 6:51:47 PM | NPEPPIN |  | Concentration | mg/L | Y | (fl. conc.) |  | Lanthanum Dissolved (fl. conc.) |  | Metals - Dissolved; | Metals - Dissolved; EMS code: LA-D | Chemical | Dissolved | LAB
LA-E | mg/L | Lanthanum Extractable | F047 | 24hr HNO3 dil'n,ICP-MS | There is no acid digestion required for this analysis. A water sample is collected and preserved in 0.2 percent nitric acid. To analyse the extractable metals the sample is shaken, allowed to sit for at least twenty four hours, and a decanted portion is introduced to the spectrometer. When the dissolved metals are required, the samples must first be filtered through a 0.45 micron cellulose acetate filter paper and then followed by preservation in 0.2 percent nitric acid. It is important to note that different laboratories may have very different detection limits for some variables analyzed by ICP/MS. | 1e-06 | Y | Y | Y | 2744 | 4/16/1998 10:21:00 AM | 9/12/2003 3:16:00 PM | 9/11/2019 8:46:18 AM | 1997-10-20 | 6/3/1998 6:51:47 PM | NPEPPIN |  | Concentration | mg/L | Y | (fl. conc.) |  | Lanthanum Extractable (fl. conc.) |  | Metals - Extractable; | Metals - Extractable; EMS code: LA-E | Chemical | Extractable | LAB
LA-E | mg/L | Lanthanum Extractable | X553 | Metals in water, extractable, ICP-SFMS | The whole water sample is collected in a HDPE/Teflon bottle. In a Class 100 workstation it is preserved with concentrated Nitric Acid at concentration of 0.2 percent (v/v). The acid preservative extracts the metals from material in the sample and a decanted portion is introduced directly to an inductively coupled argon plasma-sector field mass spectrometer (ICP-SFMS) by pneumatic nebulization. Each respective element is measured at a specific mass to charge ratio, m/z value expressed in amu (atomic mass units), and at a specified instrument resolution. LR: Low Resolution ~300, MR: Mid Resolution ~4000 and HR: High Resolution ~10000. Ref: Method number 02-2061, National Laboratory for Environmental Testing, CCIW, Environment Canada, Burlington. Requested April 2003. | 1e-06 | Y | Y | Y | 2584 | 5/22/2003 | 9/21/2007 | 4/18/2018 9:28:02 AM | 2004-01-19 |  | UMULLER |  | Concentration | mg/L | Y | (fl. conc.) |  | Lanthanum Extractable (fl. conc.) |  | Metals - Extractable; | Metals - Extractable; EMS code: LA-E | Chemical | Extractable | LAB
LA-T | ppm (S) | Lathanum Total | 6042 | Aqua Regia by ICPMS | These were sediment/soils digested with Aqua-Regia (Nitric Acid & Hydrochloric Acids) and then analysis by ICPMS. Results would be reported in ppm s. | 0.002 | Y | Y | Y | 71 | 5/14/2022 | 7/27/2024 | 7/29/2024 9:13:08 AM | 2022-06-28 |  | EMS |  | DimensionlessRatio | ppm | N | (ratio) |  | Lathanum Total (ratio) |  | Metals - Total; | Metals - Total; EMS code: LA-T | Chemical | Total | LAB
LA-T | ppm (S) | Lathanum Total | MS81 | ALS Minerals: ME-MS81 | ALS Minerals: ME-MS81. A prepared sample (0.100 g) is added to lithium metaborate/lithium tetraborate flux; mixed well and fused in a furnace at 1025°C. The resulting melt is then cooled and dissolved in an acid mixture containing nitric; hydrochloric and hydrofluoric acids. This solution is then analyzed by inductively coupled plasma - mass spectrometry. | 0.1 | Y | Y | Y | 0 |  |  |  | 2024-01-29 |  | EMS |  | DimensionlessRatio | ppm | N | (ratio) |  | Lathanum Total (ratio) |  | Metals - Total; | Metals - Total; EMS code: LA-T | Chemical | Total | LAB
LA-T | ppm (S) | Lathanum Total | 6041 | Aqua Regia by ICPMS | These were sediment/soils digested with Aqua-Regia (Nitric Acid & Hydrochloric Acids) and then analysis by ICPMS. Results would be reported in mg/kg and/or %. | 0.002 | Y | Y | Y | 21 | 10/13/2019 | 8/5/2020 | 8/10/2020 3:52:06 PM | 2019-09-06 | 10/21/2019 11:55:27 AM | KATFRASE | KATFRASE | DimensionlessRatio | ppm | N | (ratio) |  | Lathanum Total (ratio) |  | Metals - Total; | Metals - Total; EMS code: LA-T | Chemical | Total | LAB
LA-T | mg/L | Lathanum Total | METN | MET-T-NP-CCMS-VA | Ultra trace metals in water are analyzed by CRC ICPMS, based on US EPA Method 6020B (July 2014). The detection limits provided can only be met for undigested samples. This procedure is intended for colorless, non-turbid, acid-preserved water samples (i.e. pristine water samples), having turbidity < 1 NTU and no odor. Where turbidity exceeds 1 NTU, and/or the sample is colored and has an odor, results may be biased low compared to true Total Metals concentrations. ALS recommends that turbidity analysis be requested on samples submitted for this test to aid with interpretation of results. Where turbidity is <1NTU, undigested metals are equivalent to total metals concentrations. | 1e-05 | Y | Y | Y | 280 | 8/6/2019 | 4/4/2024 | 4/29/2024 4:53:06 PM | 2021-06-16 |  | EMS |  | Concentration | mg/L | Y | (fl. conc.) |  | Lathanum Total (fl. conc.) |  | Metals - Total; | Metals - Total; EMS code: LA-T | Chemical | Total | LAB
LA-T | mg/L | Lathanum Total | X561 | Total metals: FA, HNO3, HR-ICPMS (No Prep)Prep) | Ultra trace metals in water are analyzed by high resolution inductively coupled plasma mass spectrometry (HR-ICPMS), modified from US EPA 200.8, (Revision 5.5). The detection limits provided can only be met for undigested samples. This procedure is intended for colorless, non-turbid, acid-preserved water samples (i.e. pristine water samples), having turbidity < 1 NTU and no odor. Where turbidity exceeds 1 NTU, and/or the sample is colored and has an odor, results may be biased low compared to true Total Metals concentrations. ALS recommends that turbidity analysis be requested on samples submitted for this test to aid with interpretation of results. | 1e-05 | Y | Y | Y | 6 | 5/4/2016 | 5/29/2018 | 6/6/2018 11:28:31 AM | 2016-05-19 |  | RDGIBSON |  | Concentration | mg/L | Y | (fl. conc.) |  | Lathanum Total (fl. conc.) |  | Metals - Total; | Metals - Total; EMS code: LA-T | Chemical | Total | LAB
LA-T | mg/L | Lathanum Total | F082 | Aliquot,HNO3/HCL dig,ICPMS | Aliquot microwave digested with 3:1 HNO3/HCL. Analysis by ICPMS. For HiVol samples preliminary result are reported in mg/L and loading results are reported in ug/m3 | 5e-05 | Y | Y | Y | 69 | 7/30/2020 | 8/16/2024 | 8/21/2024 11:26:07 AM | 2021-11-03 |  | RDGIBSON |  | Concentration | mg/L | Y | (fl. conc.) |  | Lathanum Total (fl. conc.) |  | Metals - Total; | Metals - Total; EMS code: LA-T | Chemical | Total | LAB
LA-T | ppm (S) | Lathanum Total | MS41 | ALS Minerals ME-MS41 | ALS Minerals ME-MS41. A prepared sample is digested with aqua regia in a graphite heating block. After cooling; the resulting solution is diluted to with deionized water; mixed and analyzed by ICP-AES. Samples are then analyzed by ICP-MS for the remaining suite of elements. | 0.2 | Y | Y | Y | 0 |  |  |  | 2024-01-29 |  | EMS |  | DimensionlessRatio | ppm | N | (ratio) |  | Lathanum Total (ratio) |  | Metals - Total; | Metals - Total; EMS code: LA-T | Chemical | Total | LAB
LA-T | ppm (S) | Lathanum Total | MS61 | ALS Minerals ME-MS61 | ALS Minerals ME-MS61. A prepared sample is digested with HNO3; HF; and HClO4 in a graphite heating block to near dryness. After cooling; the sample is boiled in diluted HCl. Bulk to volume with deionized water; mixed and analyzed by ICP-AES and ICP-MS for all elements. | 0.5 | Y | Y | Y | 0 |  |  |  | 2024-01-29 |  | EMS |  | DimensionlessRatio | ppm | N | (ratio) |  | Lathanum Total (ratio) |  | Metals - Total; | Metals - Total; EMS code: LA-T | Chemical | Total | LAB
LA-T | mg/L | Lathanum Total | X555 | Metals in water, total, "in-bottle" digestion, ICP-SFMS | The whole water sample collected in a HDPE/Teflon bottle, is de-capped in a Class 100 workstation, acidified to 2% (v/v) HNO3 and 1% (v/v) HCl and recapped. The sample is "in-bottle" digested in a convection oven, set at 60 degrees C for 16 hours. The sample is cooled and introduced to an inductively coupled argon plasma-sector field mass spectrometer (ICP-SFMS) by pneumatic nebulization. Each respective element is measured at a specific mass to charge ratio, m/z value expressed in amu (atomic mass units), and at resolution. LR: Low Resolution ~300, MR: Mid Resolution ~4000 and HR: High a specified instrument Resolution ~10000. Ref: Method number 02-2003, National Laboratory for Environmental Testing, CCIW, Environment Canada, Burlington. Requested April 2003. | 1e-06 | Y | Y | Y | 17405 | 5/20/2003 | 12/6/2023 2:08:00 AM | 2/14/2024 9:05:32 AM | 2004-01-19 |  | UMULLER |  | Concentration | mg/L | Y | (fl. conc.) |  | Lathanum Total (fl. conc.) |  | Metals - Total; | Metals - Total; EMS code: LA-T | Chemical | Total | LAB
LD23 | % | Lauric acid-d23 | SURR | Surrogate Recovery | Surrogate Recovery | 0.001 | Y | Y | Y | 12 | 11/28/2022 | 6/15/2024 | 7/8/2024 11:10:02 AM | 2022-01-18 | 1/3/2023 11:50:09 AM | RDGIBSON | EMS | DimensionlessRatio | % | N | (ratio) |  | Lauric acid-d23 (ratio) | 59154-43-7 | Organics; | Organics; EMS code: LD23 | Chemical |  | SURROGATE_RESULT
LEPH | ug/g | Light Extractable Petroleum Hydrocarbons | F086 | EPH:Hex/Actn,soxhlet extr'n,GC/FID(PAHs not sub) | Samples are dried and soxhlet extracted with 1:1 Hexane:Acetone for 4 hours. The extract is concentrated and analyzed by capillary gas chromatography with flame ionization detector. Results are not corrected for specified PAH's. LEPH is defined as the sum of components that elute on gc between nC10 and nC19; HEPH is defined as the sum of components that elute on gc between nC19 and nC32. Performance Base Method published circa Dec. 31, 2000. | 250 | Y | Y | Y | 327 | 6/19/2000 1:42:00 PM | 9/9/2022 | 1/11/2024 1:07:08 PM | 2000-07-14 |  | NPEPPIN |  | MassPerMass | mg/kg | Y | (mass conc.) |  | Light Extractable Petroleum Hydrocarbons (mass conc.) |  | Petroleum Hydrocarbons; | Petroleum Hydrocarbons; EMS code: LEPH | Chemical | Extractable | LAB
LEPH | ug/g | Light Extractable Petroleum Hydrocarbons | F067 | DCM,soxhlet extr'n,GC/FID,(PAHs not sub.) | Samples are dried and soxhlet extracted with dichloromethane (DCM) for 4 hours. The extract is concentrated and analyzed by capillary gas chromatography with flame ionization detector. Results are corrected for percent moisture, but are not corrected for specified PAH's. LEPH is defined as the sum of components that elute on gc between nC10 and nC19; HEPH is defined as the sum of components that elute on gc between nC19 and nC32. Method superceded Dec. 2000. | 250 | Y | Y | Y | 21 | 6/8/2001 8:52:00 AM | 9/19/2001 11:15:00 AM | 3/17/2003 1:22:50 PM | 1999-06-28 |  | NPEPPIN |  | MassPerMass | mg/kg | Y | (mass conc.) |  | Light Extractable Petroleum Hydrocarbons (mass conc.) |  | Petroleum Hydrocarbons; | Petroleum Hydrocarbons; EMS code: LEPH | Chemical | Extractable | LAB
LEPH | ug/g | Light Extractable Petroleum Hydrocarbons | F066 | DCM,soxhlet extr'n,GC/FID(PAHs sub.) | Samples are dried and soxhlet extracted with dichloromethane (DCM) for 4 hours. The extract is concentrated and analyzed by capillary gas chromatography with flame ionization detector. Results are corrected for specified PAH concentrations and percent moisture. LEPH is defined as the sum of components that elute on gc between nC1O and nC19; HEPH is defined as the sum of components that elute on gc between nC19 and nC32. Method superceded Dec. 2000. | 250 | Y | Y | Y | 10 | 7/9/2001 9:41:00 AM | 9/19/2001 11:16:00 AM | 3/17/2003 1:22:49 PM | 1999-06-28 |  | NPEPPIN |  | MassPerMass | mg/kg | Y | (mass conc.) |  | Light Extractable Petroleum Hydrocarbons (mass conc.) |  | Petroleum Hydrocarbons; | Petroleum Hydrocarbons; EMS code: LEPH | Chemical | Extractable | LAB
LEPH | mg/L | Light Extractable Petroleum Hydrocarbons | XCAL | Calculation (mg/L) | Calculation (mg/L) |  | Y | Y | Y | 485 | 8/28/2014 | 8/15/2024 | 8/16/2024 1:36:03 PM | 2009-10-06 | 1/13/2010 1:56:21 PM | RDGIBSON | RDGIBSON | Concentration | mg/L | Y | (fl. conc.) |  | Light Extractable Petroleum Hydrocarbons (fl. conc.) |  | Petroleum Hydrocarbons; | Petroleum Hydrocarbons; EMS code: LEPH | Chemical | Extractable | LAB
LEPH | ug/g | Light Extractable Petroleum Hydrocarbons | F085 | EPH:Hex/Actn,soxhlet extr'n,GC/FID(PAHs sub) | Samples are dried and soxhlet extracted with 1:1 Hexane:Acetone for 4 hours. The extract is concentrated and analyzed by capillary gas chromatography with flame ionization detector. Results are corrected for specified PAH concentrations. LEPH is defined as the sum of components that elute on gc between nC10 and nC19; HEPH is defined as the sum of components that elute on gc between nC19 and nC32. Performance Base Method published circa Dec. 31, 2000. | 250 | Y | Y | Y | 234 | 6/19/2000 1:42:00 PM | 10/29/2022 9:40:00 AM | 1/7/2023 3:12:10 PM | 2000-07-14 |  | NPEPPIN |  | MassPerMass | mg/kg | Y | (mass conc.) |  | Light Extractable Petroleum Hydrocarbons (mass conc.) |  | Petroleum Hydrocarbons; | Petroleum Hydrocarbons; EMS code: LEPH | Chemical | Extractable | LAB
LEPH | ug/g | Light Extractable Petroleum Hydrocarbons | CAL2 | Calculation (ug/g) | Calculation (ug/g) |  | Y | Y | Y | 391 | 5/21/2015 | 9/9/2022 | 7/19/2024 8:20:35 AM | 2009-10-06 |  | RDGIBSON |  | MassPerMass | mg/kg | Y | (mass conc.) |  | Light Extractable Petroleum Hydrocarbons (mass conc.) |  | Petroleum Hydrocarbons; | Petroleum Hydrocarbons; EMS code: LEPH | Chemical | Extractable | LAB
LEPH | mg/L | Light Extractable Petroleum Hydrocarbons | EPH3 | EPH, Hexane micro-extrn, GC/FID (Waters, w/o Si-Gel) | A single extraction using a small volume of hexane is conducted within the original sample container using vigorous mechanical agitation. A portion of the hexane extract is concentrated by evaporation if necessary, and is analyzed by capillary column gas chromatography with flame ionization detection. Test results for EPHw10-19 and EPHw19-32 are not corrected for PAHs. LEPHw and HEPHw are EPHw10-19 and EPHw19-32 minus the concentrations of applicable PAHs found in each fraction. Performance Based Method, revision date Nov 1, 2015. | 0.25 | Y | Y | Y | 299 | 1/21/2016 10:45:00 AM | 12/3/2022 8:48:00 AM | 12/15/2022 6:20:11 PM | 2015-10-23 |  | RDGIBSON |  | Concentration | mg/L | Y | (fl. conc.) |  | Light Extractable Petroleum Hydrocarbons (fl. conc.) |  | Petroleum Hydrocarbons; | Petroleum Hydrocarbons; EMS code: LEPH | Chemical | Extractable | LAB
LEPH | mg/L | Light Extractable Petroleum Hydrocarbons | EPH1 | EPHnC10-nC19:extr. silica gel: GC/FID (PAH not subtr.) | Light Extractable Petroleum Hydrocarbon | 0.25 | Y | Y | Y | 45 |  |  | 1/30/2017 2:57:53 PM | 2009-10-06 |  | RDGIBSON |  | Concentration | mg/L | Y | (fl. conc.) |  | Light Extractable Petroleum Hydrocarbons (fl. conc.) |  | Petroleum Hydrocarbons; | Petroleum Hydrocarbons; EMS code: LEPH | Chemical | Extractable | LAB
LEPH | mg/L | Light Extractable Petroleum Hydrocarbons | F064 | EPH:DCM,liq-liq extr'n,GC/FID,(PAH sub.) | Samples are extracted with dichloromethane (DCM) in a separatory funnel. Extracts are dried, concentrated and analyzed by capillary gas chromatography with flame ionization detector. Results are corrected for specified PAH concentrations. LEPH is defined as the sum of components that elute on GC between nC10 and nC19; HEPH is defined as the sum of components that elute on gc between nC19 and nC32. Performance Base Method published circa Dec. 31, 2000, but distributed 1999. | 1 | Y | Y | Y | 760 | 6/12/2000 | 7/3/2024 | 7/22/2024 3:17:29 PM | 1999-06-28 |  | NPEPPIN |  | Concentration | mg/L | Y | (fl. conc.) |  | Light Extractable Petroleum Hydrocarbons (fl. conc.) |  | Petroleum Hydrocarbons; | Petroleum Hydrocarbons; EMS code: LEPH | Chemical | Extractable | LAB
LEPH | mg/L | Light Extractable Petroleum Hydrocarbons | TEHW | Total Ext-Hydrocarbons in Water by GCFID | Total Ext-Hydrocarbons in Water by GCFID, BC Lab Manual | 0.25 | Y | Y | Y | 638 | 5/5/2015 | 12/12/2019 | 3/10/2023 3:28:44 PM | 2015-11-02 |  | RDGIBSON |  | Concentration | mg/L | Y | (fl. conc.) |  | Light Extractable Petroleum Hydrocarbons (fl. conc.) |  | Petroleum Hydrocarbons; | Petroleum Hydrocarbons; EMS code: LEPH | Chemical | Extractable | LAB
LEPH | mg/L | Light Extractable Petroleum Hydrocarbons | F065 | EPH:DCM,liq-liq extr'n,GC/FID,(PAH not sub.) | Samples are extracted with dichloromethane (DCM) in a separatory funnel. Extracts are dried, concentrated and analyzed by capillary gas chromatography with flame ionization detector. Results are not corrected for specified PAH concentrations. LEPH is defined as the sum of components that elute on gc between nC10 and nC19; HEPH is defined as the sum of components that elute on gc between nC19 and nC32. Performance Base Method published circa Dec. 31, 2000, but distributed 1999. | 1 | Y | Y | Y | 1710 | 6/12/2000 | 8/21/2024 | 8/22/2024 5:27:05 PM | 1999-06-28 |  | NPEPPIN |  | Concentration | mg/L | Y | (fl. conc.) |  | Light Extractable Petroleum Hydrocarbons (fl. conc.) |  | Petroleum Hydrocarbons; | Petroleum Hydrocarbons; EMS code: LEPH | Chemical | Extractable | LAB
LI-D | mg/L | Lithium Dissolved | F068 | HNO3dig,TeflonTube/Autoclave,ICP-MS | This method is applicable to determiniation of dissolved, extractable, and total recoverable metals in ground waters, surface waters, drinking waters, and wastewaters. Dissolved, extractable, and total recoverable sample fractions are acidified to pH <1.5 with analytical grade nitric acid (HNO3) to prevent the metals from adsorbing to the container material or precipitating. Extractable and total recoverable water samples are collected in 250 ml acid washed polyethylene bottles and are preserved with Nitric acid (0.5 ml acid for 100 mLs of sample). All preserved samples are stored at room temperature and can be held for up to 6 months. Unfiltered, un-preserved samples are stored at 4C for no more than 30 days. For dissolved and extractable water fractions, no sample preparation is necessary. Samples are simply poured in autosampler tubes (i.e. 15 ml centrifuge tubes), and placed on the autosampler and analyzed. Samples for total recoverable metals are digested with Nitric acid in special Teflon tubes in an autoclave prior to analyses. | 5e-05 | Y | Y | Y | 1398 | 6/1/1999 1:21:00 PM | 4/17/2015 1:06:00 PM | 8/11/2021 12:04:11 PM | 1999-06-28 |  | NPEPPIN |  | Concentration | mg/L | Y | (fl. conc.) |  | Lithium Dissolved (fl. conc.) |  | Metals - Dissolved; | Metals - Dissolved; EMS code: LI-D | Chemical | Dissolved | LAB
LI-D | mg/L | Lithium Dissolved | F039 | Field filt + acid, ICP-AES | Dissolved and extractable water samples are ready for analysis once collected. Total samples are autoclaved with HNO3 and hydrochloric acid (HCI) in capped Teflon vials. Analysis is performed via Inductively Coupled Argon Plasma Atomic Emission Spectrometry (ICP-AES)(simultaneous multielement analysis), Graphite Furnace Atomic Absorption Spectrometry (GF-AAS)(low level analysis) and/or Flame Atomic Absorption Spectrometry (F-AAS). | 0.02 | Y | Y | Y | 436 | 10/7/2013 1:59:00 PM | 7/5/2019 1:21:00 PM | 7/12/2019 6:49:16 PM | 2013-10-09 | 10/9/2013 2:54:07 PM | RDGIBSON | RDGIBSON | Concentration | mg/L | Y | (fl. conc.) |  | Lithium Dissolved (fl. conc.) |  | Metals - Dissolved; | Metals - Dissolved; EMS code: LI-D | Chemical | Dissolved | LAB
LI-D | mg/L | Lithium Dissolved | METC | MET-D-NP-CCMS-VA | Ultra trace metals in water are analyzed by CRC ICPMS, based on US EPA Method 6020B (July 2014). This procedure is intended for pristine field-filtered acid-preserved water samples. The detection limits (LOR) for this test are based on lab instrumental analysis only, not including filtration. ALS-supplied field filtration equipment does not support these LOR. Therefore, because of the high probability of false positives due to filtration, it is strongly recommended that a filtration blank be analysed to aid in data interpretation. | 0.0005 | Y | Y | Y | 1776 | 4/3/2012 | 8/15/2024 | 8/23/2024 9:50:35 AM | 2019-08-20 |  | KATFRASE |  | Concentration | mg/L | Y | (fl. conc.) |  | Lithium Dissolved (fl. conc.) |  | Metals - Dissolved; | Metals - Dissolved; EMS code: LI-D | Chemical | Dissolved | LAB
LI-D | mg/L | Lithium Dissolved | X562 | Dissolved metals: FF, FA, HNO3, HR-ICPMS | Ultra trace metals in water are analyzed by high resolution inductively coupled plasma mass spectrometry (HR-ICPMS), based on US EPA Method 200.8, (Oct. 1999). This procedure is intended for pristine field-filtered acid-preserved water samples. The detection limits (LOR) for this test are based on lab instrumental analysis only, not including filtration. ALS-supplied field filtration equipment does not support these LOR. Therefore, because of the high probability of false positives due to filtration, it is strongly recommended that a filtration blank be analyzed to aid in data interpretation. | 0.0005 | Y | Y | Y | 1264 | 6/8/2015 | 10/19/2023 | 11/9/2023 7:41:27 AM | 2015-10-22 |  | RDGIBSON |  | Concentration | mg/L | Y | (fl. conc.) |  | Lithium Dissolved (fl. conc.) |  | Metals - Dissolved; | Metals - Dissolved; EMS code: LI-D | Chemical | Dissolved | LAB
LI-D | mg/L | Lithium Dissolved | X170 | ICP | This code is a catch-all for a large variety of ICP methods including water and also solids digested into solutions as an intermediate result. Different procedures will have greatly different detection limits. (a) For air particulates, preliminary results are reported in mg/L and loading results as mg/dm2/d. (b) For Hi Vol samples preliminiary results are same as reported in mg/L and loading results as ug/m3. (c) Note that for Rain samples there is no filtration nor digestion steps. | 0.001 | Y | Y | Y | 4096 | 1/19/1998 | 11/6/2018 | 6/12/2020 1:21:11 PM | 1997-06-23 | 6/3/1998 6:51:47 PM | NPEPPIN |  | Concentration | mg/L | Y | (fl. conc.) |  | Lithium Dissolved (fl. conc.) |  | Metals - Dissolved; | Metals - Dissolved; EMS code: LI-D | Chemical | Dissolved | LAB
LI-D | mg/L | Lithium Dissolved | X556 | Metals in water, dissolved, "in-bottle" digestion, ICP-SFMS | The whole water sample is filtered through a 0.45 micron cellulose acetate filter paper and the filtrate is stored in a HDPE/Teflon bottle. The container is de-capped in a Class 100 workstation, acidified to 2% (v/v) HNO3 and 1% (v/v) HCl and recapped. The sample is "in-bottle" digested in a convection oven, set at 60 degrees C for 16 hours. The sample is cooled and introduced to an inductively coupled argon plasma-sector field mass spectrometer (ICP-SFMS) by pneumatic nebulization. Each respective element is measured at a specific mass to charge ratio, m/z value expressed in amu (atomic mass units), and at a specified intrument resolution. LR: Low Resolution ~300, MR: Mid Resolution ~4000 and HR: High Resolution ~10000. Ref: Method number 02-2003, National Laboratory for Environmental Testing, CCIW, Environment Canada, Burlington, Requested April 2003. | 0.0002 | Y | Y | Y | 4234 | 6/11/2003 | 12/6/2023 2:04:00 AM | 2/14/2024 9:05:32 AM | 2004-01-19 |  | UMULLER |  | Concentration | mg/L | Y | (fl. conc.) |  | Lithium Dissolved (fl. conc.) |  | Metals - Dissolved; | Metals - Dissolved; EMS code: LI-D | Chemical | Dissolved | LAB
LI-D | mg/L | Lithium Dissolved | MWID | Metals in Water by ICPMS (digested) | Metals in Water by ICPMS (digested) Based on EPA 200.8 | 0.01 | Y | Y | Y | 31766 | 10/26/2018 5:20:00 PM | 10/26/2018 5:20:00 PM | 7/31/2024 1:46:41 PM | 2009-12-24 |  | RDGIBSON |  | Concentration | mg/L | Y | (fl. conc.) |  | Lithium Dissolved (fl. conc.) |  | Metals - Dissolved; | Metals - Dissolved; EMS code: LI-D | Chemical | Dissolved | LAB
LI-D | mg/L | Lithium Dissolved | F032 | Filt.,Acid,ICPMS | There is no acid digestion required for this analysis. A water sample is collected and preserved in 0.2 percent nitric acid. To analyse the extractable metals the sample is shaken, allowed to sit for at least twenty four hours, and a decanted portion is introduced to the spectrometer. When the dissolved metals are required, the samples must first be filtered through a 0.45 micron cellulose acetate filter paper and then followed by preservation in 0.2 percent nitric acid. It is important to note that different laboratories may have very different detection limits for some variables analyzed by ICPMS. | 0.001 | Y | Y | Y | 74787 | 3/10/1998 2:06:00 PM | 8/26/2024 | 8/26/2024 4:53:05 PM | 1997-10-20 | 8/15/2019 10:56:19 AM | NPEPPIN | KATFRASE | Concentration | mg/L | Y | (fl. conc.) |  | Lithium Dissolved (fl. conc.) |  | Metals - Dissolved; | Metals - Dissolved; EMS code: LI-D | Chemical | Dissolved | LAB
LI-D | mg/L | Lithium Dissolved | MSEA | Metals in Seawater (undigested) by (CRC) ICP-MS | This method describes the multi-elemental determination of trace elements by Collision/Reaction Cell (CRC) ICP-MS. The method measures ions produced by a radio-frequency inductively coupled plasma. Analyte species originating in a liquid are pneumatically nebulized and the resulting aerosol transported by argon gas into the plasma torch. The ions produced are entrained in the plasma gas and extracted, by means of a differentially pumped vacuum interface, into a mass spectrometer. The ions flow through a Collision/Reaction Cell (CRC) where depending on the mode of analysis either the ions passes through the cell unhindered (no gas - classical ICP-MS mode), or collisions/reactions chemistry takes place between the ions and a neutral gas such as hydrogen or helium. Any polyatomic ions which give interference on the analytes are removed. The ions are then sorted according to their mass-to-charge ratios by a quadrupole mass spectrometer. The ions are then sorted according to their mass-to-charge ratios by a quadrupole mass spectrometer. The ions transmitted through the quadrupole are quantified by a channel electron multiplier and the ion information processed by a data handling system. Although most polyatomic interference is removed, any additional interference must be assessed and valid corrections applied or the data flagged to indicate problems. Interference correction must include compensation for background ions contributed by the plasma gas, reagents, and constituents of the sample matrix. Instrument drift as well as suppressions or enhancements of instrument response caused by the sample matrix must be corrected for by the use of internal standards. | 0.0002 | Y | Y | Y | 4 | 4/12/2023 | 4/29/2024 | 5/16/2024 2:00:14 PM | 2009-09-08 |  | RDGIBSON |  | Concentration | mg/L | Y | (fl. conc.) |  | Lithium Dissolved (fl. conc.) |  | Metals - Dissolved; | Metals - Dissolved; EMS code: LI-D | Chemical | Dissolved | LAB
LI-E | mg/L | Lithium Extractable | F047 | 24hr HNO3 dil'n,ICP-MS | There is no acid digestion required for this analysis. A water sample is collected and preserved in 0.2 percent nitric acid. To analyse the extractable metals the sample is shaken, allowed to sit for at least twenty four hours, and a decanted portion is introduced to the spectrometer. When the dissolved metals are required, the samples must first be filtered through a 0.45 micron cellulose acetate filter paper and then followed by preservation in 0.2 percent nitric acid. It is important to note that different laboratories may have very different detection limits for some variables analyzed by ICP/MS. | 2e-05 | Y | Y | Y | 2676 | 4/16/1998 10:21:00 AM | 6/26/2003 | 9/11/2019 8:46:18 AM | 1997-10-20 | 6/3/1998 6:51:47 PM | NPEPPIN |  | Concentration | mg/L | Y | (fl. conc.) |  | Lithium Extractable (fl. conc.) |  | Metals - Extractable; | Metals - Extractable; EMS code: LI-E | Chemical | Extractable | LAB
LI-E | mg/L | Lithium Extractable | X553 | Metals in water, extractable, ICP-SFMS | The whole water sample is collected in a HDPE/Teflon bottle. In a Class 100 workstation it is preserved with concentrated Nitric Acid at concentration of 0.2 percent (v/v). The acid preservative extracts the metals from material in the sample and a decanted portion is introduced directly to an inductively coupled argon plasma-sector field mass spectrometer (ICP-SFMS) by pneumatic nebulization. Each respective element is measured at a specific mass to charge ratio, m/z value expressed in amu (atomic mass units), and at a specified instrument resolution. LR: Low Resolution ~300, MR: Mid Resolution ~4000 and HR: High Resolution ~10000. Ref: Method number 02-2061, National Laboratory for Environmental Testing, CCIW, Environment Canada, Burlington. Requested April 2003. | 0.0002 | Y | Y | Y | 2583 | 5/22/2003 | 9/21/2007 | 4/18/2018 9:28:02 AM | 2004-01-19 |  | UMULLER |  | Concentration | mg/L | Y | (fl. conc.) |  | Lithium Extractable (fl. conc.) |  | Metals - Extractable; | Metals - Extractable; EMS code: LI-E | Chemical | Extractable | LAB
LI-L | mg/L | Lithium Leachable | MEND | MEND (mod) EPA 300.1 (mod) | MEND (mod) EPA 300.1 (mod). Inorganic anions are analyzed by obtaining an extract from a soil sample that has been added in a 3:1 ratio of deionized water to soil; then leached and the extract filtered prior to analysis; which is then analyzed by Ion Chromatography with conductivity and/or UV detection. | 0.005 | Y | Y | Y | 0 |  |  |  | 2024-01-29 |  | RDGIBSON |  | Concentration | mg/L | Y | (fl. conc.) |  | Lithium Leachable (fl. conc.) |  | Metals - Extractable; | Metals - Extractable; EMS code: LI-L | Chemical |  | LAB
LI-T | ppm (S) | Lithium Total | 6042 | Aqua Regia by ICPMS | These were sediment/soils digested with Aqua-Regia (Nitric Acid & Hydrochloric Acids) and then analysis by ICPMS. Results would be reported in ppm s. | 0.1 | Y | Y | Y | 71 | 5/14/2022 | 7/27/2024 | 7/29/2024 9:13:08 AM | 2022-06-28 |  | EMS |  | DimensionlessRatio | ppm | N | (ratio) |  | Lithium Total (ratio) |  | Metals - Total; | Metals - Total; EMS code: LI-T | Chemical | Total | LAB
LI-T | mg/L | Lithium Total | F081 | Aliquot,HNO3/HCl Dig,ICP | Aliquot microwave digested with 3:1 HNO3/HCL. Analysis by ICP. For HiVol samples preliminary result are reported in mg/L and loading results are reported in ug/m3 |  | Y | Y | Y | 2630 |  |  | 8/26/2024 2:41:32 PM | 2000-06-13 |  | NPEPPIN |  | Concentration | mg/L | Y | (fl. conc.) |  | Lithium Total (fl. conc.) |  | Metals - Total; | Metals - Total; EMS code: LI-T | Chemical | Total | LAB
LI-T | ppm (S) | Lithium Total | MS41 | ALS Minerals ME-MS41 | ALS Minerals ME-MS41. A prepared sample is digested with aqua regia in a graphite heating block. After cooling; the resulting solution is diluted to with deionized water; mixed and analyzed by ICP-AES. Samples are then analyzed by ICP-MS for the remaining suite of elements. | 0.1 | Y | Y | Y | 0 |  |  |  | 2024-01-29 |  | EMS |  | DimensionlessRatio | ppm | N | (ratio) |  | Lithium Total (ratio) |  | Metals - Total; | Metals - Total; EMS code: LI-T | Chemical | Total | LAB
LI-T | mg/L | Lithium Total | X285 | HNO3 Dig:ICP Analysis | HNO3 Dig:ICP Analysis | 0.01 | Y | Y | Y | 32 |  |  | 2/5/2001 4:22:09 PM | 2000-03-14 | 3/14/2000 3:52:34 PM | NPEPPIN | NPEPPIN | Concentration | mg/L | Y | (fl. conc.) |  | Lithium Total (fl. conc.) |  | Metals - Total; | Metals - Total; EMS code: LI-T | Chemical | Total | LAB
LI-T | ug/m3 | Lithium Total | F088 | Aliquot,HNO3/HCl dig, ICPMS,loading | Aliquot microwave digested with 3:1 HNO3/HCL. Analysis by ICPMS. For HiVol samples preliminary result are reported in mg/L and loading results are reported in ug/m3 |  | Y | Y | Y | 565 | 8/17/2000 10:46:00 AM | 12/14/2006 | 12/30/2006 1:01:36 AM | 2000-08-16 |  | NPEPPIN |  | AirConcentration | µg/m³ | Y | (air conc.) |  | Lithium Total (air conc.) |  | Metals - Total; | Metals - Total; EMS code: LI-T | Chemical | Total | LAB
LI-T | ppm (S) | Lithium Total | MS61 | ALS Minerals ME-MS61 | ALS Minerals ME-MS61. A prepared sample is digested with HNO3; HF; and HClO4 in a graphite heating block to near dryness. After cooling; the sample is boiled in diluted HCl. Bulk to volume with deionized water; mixed and analyzed by ICP-AES and ICP-MS for all elements. | 0.2 | Y | Y | Y | 0 |  |  |  | 2024-01-29 |  | EMS |  | DimensionlessRatio | ppm | N | (ratio) |  | Lithium Total (ratio) |  | Metals - Total; | Metals - Total; EMS code: LI-T | Chemical | Total | LAB
LI-T | mg/L | Lithium Total | F033 | Acid Dig,Centrifuge,ICPMS | For the determination of total recoverable elements in water or waste water, take 100 ml aliquot from a well mixed, acid preserved, sample containing not more than 0.25% (w/v) total solids and transfer to a 250 ml Griffin beaker. (If total solids are greater than 0.25% reduce the size of the aliquot by a proportionate amount). Add 1 ml of conc. nitric acid and 0.5 ml conc. hydrochloric acid. Heat on a hot plate until the volume has been reduced to 15-20ml, ensuring that the sample does not boil. Allow to cool and transfer to a 50 ml volumetric flask. Dilute to volume with ASTM type I water. Centrifuge the sample or allow to stand overnight to separate insoluble material. Prior analysis, pipette 20 ml into a 50 ml volumetric flask and dilute to volume with ASTM type I water. Analysis should be performed within seventy two hours of the completed preparation. Note that different laboratories may have very different field preservation, digestion procedures and detection limits for this method, so these should be documented. | 0.001 | Y | Y | Y | 34294 | 11/20/1997 9:31:00 AM | 8/22/2024 3:24:00 AM | 8/25/2024 9:47:36 PM | 1998-05-14 | 6/3/1998 6:51:49 PM | NPEPPIN |  | Concentration | mg/L | Y | (fl. conc.) |  | Lithium Total (fl. conc.) |  | Metals - Total; | Metals - Total; EMS code: LI-T | Chemical | Total | LAB
LI-T | mg/L | Lithium Total | METN | MET-T-NP-CCMS-VA | Ultra trace metals in water are analyzed by CRC ICPMS, based on US EPA Method 6020B (July 2014). The detection limits provided can only be met for undigested samples. This procedure is intended for colorless, non-turbid, acid-preserved water samples (i.e. pristine water samples), having turbidity < 1 NTU and no odor. Where turbidity exceeds 1 NTU, and/or the sample is colored and has an odor, results may be biased low compared to true Total Metals concentrations. ALS recommends that turbidity analysis be requested on samples submitted for this test to aid with interpretation of results. Where turbidity is <1NTU, undigested metals are equivalent to total metals concentrations. | 0.0005 | Y | Y | Y | 3648 | 4/3/2012 | 8/19/2024 | 8/26/2024 1:03:06 PM | 2019-08-20 |  | KATFRASE |  | Concentration | mg/L | Y | (fl. conc.) |  | Lithium Total (fl. conc.) |  | Metals - Total; | Metals - Total; EMS code: LI-T | Chemical | Total | LAB
LI-T | ug/g wet | Lithium Total | MTIW | MET-WET-CCMS-MID-VA | Metals in Tissue by CRC ICPMS (WET) Methodology: This method is conducted following British Columbia Lab Manual method "Metals in Animal Tissue and Vegetation (Biota) - Prescriptive". Tissue samples are homogenized and sub-sampled prior to hotblock digestion with nitric and hydrochloric acids, in combination with addition of hydrogen peroxide. Instrumental analysis is by collision cell inductively coupled plasma - mass spectrometry (modified from EPA Method 6020A). Method Limitation: This method employs a strong acid/peroxide digestion, and is intended to provide a conservative estimate of bio-available metals. Near complete recoveries are achieved for most toxicologically important metals, but elements associated with recalcitrant minerals may be only partially recovered. | 0.1 | Y | Y | Y | 24 | 11/25/2018 | 12/6/2018 | 7/21/2021 2:55:07 PM | 2018-12-24 |  | EMS |  | MassPerMass | µg/g | Y | (mass conc.) | (wet) | Lithium Total (wet) (mass conc.) |  | Metals - Total; | Metals - Total; EMS code: LI-T | Chemical | Total | LAB
LI-T | mg/L | Lithium Total | F048 | Acid dig,conc'n,nebul.,ICP-OES | This method provides an automated procedure for the Inductively Coupled Argon Plasma Optical Emission Spectrophotometric (ICP-OES) determination and quantitation of trace amounts of metals in surface, ground and precipitation waters after a manual acid digestion and concentration procedure and ultrasonic nebulisation. Organically-bound metals and metal complexes are converted to metal ions by digestion with nitric acid [method details says HCl]. To increase the signal detected by the ICP, the samples are concentrated 10 times. The liquid sample is pumped onto the piezoelectric transducer, where the ultrasonic action efficiently converts it to a fine, dense aerosol. The nebulizer carrier gas flow then transports the wet aerosol through a heated U-tube, where the solvent is vaporized. Subsequently, most of the solvent vapour is removed from the aerosol stream by passing the stream through a chilled condenser. The result is a dry, analyte-laden aerosol which is introduced to the plasma. There, the high temperature desolvates the aerosol and dissociates the sample material into individual atoms and ions, which are excited to emit light at wave lengths characteristic of the elements present. Dissolved samples are field filtered through 0.45 micron cellulose acetate filter prior to acid preservation. Samples for extractable analysis are acidified with 2 mL concentrated HNO3 per litre and allowed to stand a minimum of 24 hours before analysis. Clients providing these samples are advised to preserve samples at the collection site. | 0.0001 | Y | Y | Y | 6150 | 7/4/1997 2:36:00 PM | 6/13/2003 4:01:00 PM | 9/11/2019 8:46:38 AM | 1997-10-20 | 6/3/1998 6:51:47 PM | NPEPPIN |  | Concentration | mg/L | Y | (fl. conc.) |  | Lithium Total (fl. conc.) |  | Metals - Total; | Metals - Total; EMS code: LI-T | Chemical | Total | LAB
LI-T | ug/m3 | Lithium Total | 6040 | Aqua Regia Dig:ICP | Aqua regia digestion is used to bring the metals into solution prior to ICP analysis. For Hi-Vol samples, preliminary results are reported as mg/L and24 hr loadings as ug/m3 | 0.001 | Y | Y | Y | 437 | 2/13/2002 | 2/9/2005 | 2/23/2005 1:13:47 AM | 2004-02-10 |  | UMULLER |  | AirConcentration | µg/m³ | Y | (air conc.) |  | Lithium Total (air conc.) |  | Metals - Total; | Metals - Total; EMS code: LI-T | Chemical | Total | LAB
LI-T | mg/L | Lithium Total | MWID | Metals in Water by ICPMS (digested) | Metals in Water by ICPMS (digested) Based on EPA 200.8 | 0.01 | Y | Y | Y | 36810 |  |  | 7/31/2024 1:46:41 PM | 2009-12-24 |  | RDGIBSON |  | Concentration | mg/L | Y | (fl. conc.) |  | Lithium Total (fl. conc.) |  | Metals - Total; | Metals - Total; EMS code: LI-T | Chemical | Total | LAB
LI-T | ug/g | Lithium Total | FS82 | PBM,dry60C,10mesh,SALM-ICP-MS | PBM;dry60C;10mesh;SALM-ICP-MS | 0.1 | Y | Y | Y | 515 | 12/7/2009 3:52:00 PM | 7/24/2024 | 7/29/2024 5:17:09 PM | 2002-11-12 |  | UMULLER |  | MassPerMass | mg/kg | Y | (mass conc.) |  | Lithium Total (mass conc.) |  | Metals - Total; | Metals - Total; EMS code: LI-T | Chemical | Total | LAB
LI-T | ug/g | Lithium Total | PB02 | PBM,HNO3/HCl dig,dry60C,100mesh, | Performance Based Method. Samples are dried at 60 C and sieved using a .149 mm (100 mesh) sieve. Solid samples are digested with the combination of nitric and hydrochloric acids with a 1:1 ratio of acids. For environmental purposes, strong acid leaches are used routinely since they have the advantage of speed, low cost, and superio reproducibility. All metals trapped within the refractory matrix are unlikely to have environmental consequences. Instrumental analysis of extracts produced using this digestion method can be performed by a variety of techniques. |  | Y | Y | Y | 52 | 5/24/2001 3:42:00 PM | 11/26/2001 10:23:00 AM | 5/26/2003 2:27:54 PM | 2001-10-11 |  | NPEPPIN |  | MassPerMass | mg/kg | Y | (mass conc.) |  | Lithium Total (mass conc.) |  | Metals - Total; | Metals - Total; EMS code: LI-T | Chemical | Total | LAB
LI-T | mg/L | Lithium Total | X561 | Total metals: FA, HNO3, HR-ICPMS (No Prep)Prep) | Ultra trace metals in water are analyzed by high resolution inductively coupled plasma mass spectrometry (HR-ICPMS), modified from US EPA 200.8, (Revision 5.5). The detection limits provided can only be met for undigested samples. This procedure is intended for colorless, non-turbid, acid-preserved water samples (i.e. pristine water samples), having turbidity < 1 NTU and no odor. Where turbidity exceeds 1 NTU, and/or the sample is colored and has an odor, results may be biased low compared to true Total Metals concentrations. ALS recommends that turbidity analysis be requested on samples submitted for this test to aid with interpretation of results. | 0.0005 | Y | Y | Y | 2756 | 6/8/2015 | 10/5/2023 | 1/31/2024 2:37:41 PM | 2015-10-22 |  | RDGIBSON |  | Concentration | mg/L | Y | (fl. conc.) |  | Lithium Total (fl. conc.) |  | Metals - Total; | Metals - Total; EMS code: LI-T | Chemical | Total | LAB
LI-T | ug/g | Lithium Total | MTID | MET-DRY-CCMS-VA | Metals in Tissue by CRC ICPMS (DRY) Methodology: This method is conducted following British Columbia Lab Manual method "Metals in Animal Tissue and Vegetation (Biota) - Prescriptive". Tissue samples are homogenized and sub-sampled prior to hotblock digestion with nitric and hydrochloric acids, in combination with addition of hydrogen peroxide. Instrumental analysis is by collision cell inductively coupled plasma - mass spectrometry (modified from EPA Method 6020A). Method Limitation: This method employs a strong acid/peroxide digestion, and is intended to provide a conservative estimate of bio-available metals. Near complete recoveries are achieved for most toxicologically important metals, but elements associated with recalcitrant minerals may be only partially recovered. | 0.5 | Y | Y | Y | 122 | 8/13/2014 | 9/8/2022 | 10/26/2022 7:14:15 AM | 2015-03-09 | 3/11/2015 4:12:05 PM | RDGIBSON | EMS | MassPerMass | mg/kg | Y | (mass conc.) |  | Lithium Total (mass conc.) |  | Metals - Total; | Metals - Total; EMS code: LI-T | Chemical | Total | LAB
LI-T | mg/L | Lithium Total | F068 | HNO3dig,TeflonTube/Autoclave,ICP-MS | This method is applicable to determiniation of dissolved, extractable, and total recoverable metals in ground waters, surface waters, drinking waters, and wastewaters. Dissolved, extractable, and total recoverable sample fractions are acidified to pH <1.5 with analytical grade nitric acid (HNO3) to prevent the metals from adsorbing to the container material or precipitating. Extractable and total recoverable water samples are collected in 250 ml acid washed polyethylene bottles and are preserved with Nitric acid (0.5 ml acid for 100 mLs of sample). All preserved samples are stored at room temperature and can be held for up to 6 months. Unfiltered, un-preserved samples are stored at 4C for no more than 30 days. For dissolved and extractable water fractions, no sample preparation is necessary. Samples are simply poured in autosampler tubes (i.e. 15 ml centrifuge tubes), and placed on the autosampler and analyzed. Samples for total recoverable metals are digested with Nitric acid in special Teflon tubes in an autoclave prior to analyses. | 5e-05 | Y | Y | Y | 1779 | 6/1/1999 1:32:00 PM | 7/17/2008 4:45:00 PM | 7/30/2024 6:32:12 AM | 1999-06-28 |  | NPEPPIN |  | Concentration | mg/L | Y | (fl. conc.) |  | Lithium Total (fl. conc.) |  | Metals - Total; | Metals - Total; EMS code: LI-T | Chemical | Total | LAB
LI-T | ug/g wet | Lithium Total | X552 | Metals in Biota (wet), total, acid digestion, ICP-SFMS | One gram of wet, homogenized biota is weighed in a microwave TFM vessel, and 8mL of Nitric Acid and 1mL of Hydrogen Peroxide are added. The sample is digested at 200 degrees C for 15 minutes in a closed vessel high pressure microwave oven. The sample is cooled, the volume adjusted to 100mL and introduced to an inductively coupled argon plasma-sector field mass spectrometer (ICP-SFMS) by pneumatic nebulization. Each respective element is measured at a specific mass to charge ratio, m/z value expressed in amu (atomic mass units), and at a specified instrument resolution. LR: Low Resolution ~300, MR: Mid Resolution ~4000. Ref: 02-2704, National Laboratory for Environmental Testing, CCIW, Environment Canada, Burlington. Requested April 2003. | 0.01 | Y | Y | Y | 37 | 8/9/2016 | 10/16/2023 | 11/28/2023 2:21:32 PM | 2004-01-19 | 4/19/2021 1:22:00 PM | UMULLER | EMS | MassPerMass | µg/g | Y | (mass conc.) | (wet) | Lithium Total (wet) (mass conc.) |  | Metals - Total; | Metals - Total; EMS code: LI-T | Chemical | Total | LAB
LI-T | mg/L | Lithium Total | X170 | ICP | This code is a catch-all for a large variety of ICP methods including water and also solids digested into solutions as an intermediate result. Different procedures will have greatly different detection limits. (a) For air particulates, preliminary results are reported in mg/L and loading results as mg/dm2/d. (b) For Hi Vol samples preliminiary results are same as reported in mg/L and loading results as ug/m3. (c) Note that for Rain samples there is no filtration nor digestion steps. | 0.001 | Y | Y | Y | 6288 | 3/29/2001 | 11/7/2018 | 10/24/2022 10:58:25 AM | 1997-06-23 | 6/3/1998 6:51:47 PM | NPEPPIN |  | Concentration | mg/L | Y | (fl. conc.) |  | Lithium Total (fl. conc.) |  | Metals - Total; | Metals - Total; EMS code: LI-T | Chemical | Total | LAB
LI-T | mg/dm2/d | Lithium Total | DUST | HNO3 dig.(microwave or hot block),ICPMS | HNO3 dig.(microwave or hot block);ICPMS |  | Y | Y | Y | 420 | 7/19/2016 1:11:00 PM | 7/31/2024 | 7/31/2024 11:23:12 AM | 2008-02-07 |  | RDGIBSON |  | MassFlux | mg/dm²/d | Y | (mass flux) |  | Lithium Total (mass flux) |  | Metals - Total; | Metals - Total; EMS code: LI-T | Chemical | Total | LAB
LI-T | mg/L | Lithium Total | MSEA | Metals in Seawater (undigested) by (CRC) ICP-MS | This method describes the multi-elemental determination of trace elements by Collision/Reaction Cell (CRC) ICP-MS. The method measures ions produced by a radio-frequency inductively coupled plasma. Analyte species originating in a liquid are pneumatically nebulized and the resulting aerosol transported by argon gas into the plasma torch. The ions produced are entrained in the plasma gas and extracted, by means of a differentially pumped vacuum interface, into a mass spectrometer. The ions flow through a Collision/Reaction Cell (CRC) where depending on the mode of analysis either the ions passes through the cell unhindered (no gas - classical ICP-MS mode), or collisions/reactions chemistry takes place between the ions and a neutral gas such as hydrogen or helium. Any polyatomic ions which give interference on the analytes are removed. The ions are then sorted according to their mass-to-charge ratios by a quadrupole mass spectrometer. The ions are then sorted according to their mass-to-charge ratios by a quadrupole mass spectrometer. The ions transmitted through the quadrupole are quantified by a channel electron multiplier and the ion information processed by a data handling system. Although most polyatomic interference is removed, any additional interference must be assessed and valid corrections applied or the data flagged to indicate problems. Interference correction must include compensation for background ions contributed by the plasma gas, reagents, and constituents of the sample matrix. Instrument drift as well as suppressions or enhancements of instrument response caused by the sample matrix must be corrected for by the use of internal standards. | 0.0002 | Y | Y | Y | 441 | 8/12/2009 12:08:00 AM | 4/29/2024 | 6/7/2024 10:03:16 AM | 2009-09-08 |  | RDGIBSON |  | Concentration | mg/L | Y | (fl. conc.) |  | Lithium Total (fl. conc.) |  | Metals - Total; | Metals - Total; EMS code: LI-T | Chemical | Total | LAB
LI-T | ug/g | Lithium Total | FS94 | PBM,dry60C,100mesh,SALM-ICP-MS | PBM;dry60C;100mesh;SALM-ICP-MS, same as FS83 but MDL=0.5ug/g | 5 | Y | Y | Y | 507 | 8/31/2007 1:21:00 PM | 10/15/2021 5:06:00 PM | 10/18/2021 4:14:03 PM | 2009-08-27 | 8/27/2009 3:45:34 PM | RDGIBSON | RDGIBSON | MassPerMass | mg/kg | Y | (mass conc.) |  | Lithium Total (mass conc.) |  | Metals - Total; | Metals - Total; EMS code: LI-T | Chemical | Total | LAB
LI-T | mg/L | Lithium Total | F082 | Aliquot,HNO3/HCL dig,ICPMS | Aliquot microwave digested with 3:1 HNO3/HCL. Analysis by ICPMS. For HiVol samples preliminary result are reported in mg/L and loading results are reported in ug/m3 | 0.001 | Y | Y | Y | 25369 | 7/26/2005 7:04:00 AM | 8/23/2024 | 8/26/2024 4:53:05 PM | 2000-06-13 | 8/15/2019 11:07:58 AM | NPEPPIN | KATFRASE | Concentration | mg/L | Y | (fl. conc.) |  | Lithium Total (fl. conc.) |  | Metals - Total; | Metals - Total; EMS code: LI-T | Chemical | Total | LAB
LI-T | ppm (S) | Lithium Total | 6041 | Aqua Regia by ICPMS | These were sediment/soils digested with Aqua-Regia (Nitric Acid & Hydrochloric Acids) and then analysis by ICPMS. Results would be reported in mg/kg and/or %. | 0.1 | Y | Y | Y | 21 | 10/13/2019 | 8/5/2020 | 8/10/2020 3:52:06 PM | 2019-09-06 | 10/21/2019 11:55:30 AM | KATFRASE | KATFRASE | DimensionlessRatio | ppm | N | (ratio) |  | Lithium Total (ratio) |  | Metals - Total; | Metals - Total; EMS code: LI-T | Chemical | Total | LAB
LI-T | mg/L | Lithium Total | X555 | Metals in water, total, "in-bottle" digestion, ICP-SFMS | The whole water sample collected in a HDPE/Teflon bottle, is de-capped in a Class 100 workstation, acidified to 2% (v/v) HNO3 and 1% (v/v) HCl and recapped. The sample is "in-bottle" digested in a convection oven, set at 60 degrees C for 16 hours. The sample is cooled and introduced to an inductively coupled argon plasma-sector field mass spectrometer (ICP-SFMS) by pneumatic nebulization. Each respective element is measured at a specific mass to charge ratio, m/z value expressed in amu (atomic mass units), and at resolution. LR: Low Resolution ~300, MR: Mid Resolution ~4000 and HR: High a specified instrument Resolution ~10000. Ref: Method number 02-2003, National Laboratory for Environmental Testing, CCIW, Environment Canada, Burlington. Requested April 2003. | 0.0002 | Y | Y | Y | 17337 | 5/20/2003 | 12/6/2023 2:08:00 AM | 2/14/2024 9:05:32 AM | 2004-01-19 |  | UMULLER |  | Concentration | mg/L | Y | (fl. conc.) |  | Lithium Total (fl. conc.) |  | Metals - Total; | Metals - Total; EMS code: LI-T | Chemical | Total | LAB
LI-T | mg/L | Lithium Total | 0042 | GI HNO3 Dig: ICP | Sample aliquot taken from unfiltered unpreserved General Ions bottle; HNO3 digestion; ICP analysis. | 0.02 | Y | Y | Y | 53 | 10/7/2013 12:33:00 PM | 7/5/2019 3:21:00 PM | 7/12/2019 6:49:13 PM | 2013-10-09 | 10/9/2013 2:54:25 PM | RDGIBSON | RDGIBSON | Concentration | mg/L | Y | (fl. conc.) |  | Lithium Total (fl. conc.) |  | Metals - Total; | Metals - Total; EMS code: LI-T | Chemical | Total | LAB
LI-T | ug/g | Lithium Total | X551 | Metals in Biota (dry), total, acid digestion, ICP-SFMS | 0.1 gram of dried, homogenized biota is weighed in a microwave TFM vessel, and 8mL of Nitric Acid and 1mL of Hydrogen Peroxide are added. The sample is digested at 200 degrees C for 15 minutes in a closed vessel high pressure microwave oven. The sample is cooled, the volume adjusted to 100mL and introduced to an iductively coupled argon plasma-sector field mass spectrometer (ICP-SFMS) by pneumatic nebulization. Each respective element is measured at a specific mass to charge ratio, m/z value expressed in amu (atomic mass units), and at a specified instrument resolution. LR: Low Resolution ~300, MR: Mid Resolution ~4000 and HR: Hich Resolution ~10000. Ref: Method number 02-2705, National Laboratory for Environmental Testing, CCIW, Environment Canada, Burlington. Requested April 2003. | 0.1 | Y | Y | Y | 43 | 8/11/2015 | 10/30/2023 | 11/28/2023 2:21:33 PM | 2004-01-19 |  | UMULLER |  | MassPerMass | mg/kg | Y | (mass conc.) |  | Lithium Total (mass conc.) |  | Metals - Total; | Metals - Total; EMS code: LI-T | Chemical | Total | LAB
LI-T | mg/L | Lithium Total | 1349 | Tot Metals in Seawater by CRC ICPMS (BC) | Tot Metals in Seawater by CRC ICPMS (BC) | 0.02 | Y | Y | Y | 98 | 8/16/2019 | 12/5/2021 | 12/13/2021 12:34:16 PM | 2019-09-30 |  | KATFRASE |  | Concentration | mg/L | Y | (fl. conc.) |  | Lithium Total (fl. conc.) |  | Metals - Total; | Metals - Total; EMS code: LI-T | Chemical | Total | LAB
LI-T | mg/L | Lithium Total | F121 | Acid,ICPMS | Acidified, no digestion, ICPMS | 0.0005 | Y | Y | Y | 8438 | 3/12/2010 7:57:00 PM | 8/15/2024 9:10:00 PM | 8/26/2024 8:06:03 AM | 2002-07-22 | 7/16/2012 3:29:02 PM | UMULLER | RDGIBSON | Concentration | mg/L | Y | (fl. conc.) |  | Lithium Total (fl. conc.) |  | Metals - Total; | Metals - Total; EMS code: LI-T | Chemical | Total | LAB
LI-T | ug/g | Lithium Total | ICPM | Metals in Soil by CRC ICPMS | This analysis is carried out using procedures from CSR Analytical Method: "Strong Acid Leachable Metals (SALM) in Soil", BC Ministry of Environment, 26 June 2009, and procedures adapted from EPA Method 200.2. The sample is manually homogenized, dried at 60 degrees Celsius, sieved through a 63 um (230 mesh) sieve, and a representative subsample of the dry material is weighed. The sample is then digested at 95 degrees Celsius for 2 hours by block digester using concentrated nitric and hydrochloric acids. Instrumental analysis is by collision cell inductively coupled plasma - mass spectrometry (modifed from EPA Method 6020A). Method Limitation: This method is not a total digestion technique. It is a very strong acid digestion that is intended to dissolve those metals that may be environmentally available. By design, elements bound in silicate structures are not normally dissolved by this procedure as they are not usually mobile in the environment. | 5 | Y | Y | Y | 103 | 9/19/2017 | 9/24/2023 | 1/12/2024 11:16:52 AM | 2018-01-09 |  | JKROGH |  | MassPerMass | mg/kg | Y | (mass conc.) |  | Lithium Total (mass conc.) |  | Metals - Total; | Metals - Total; EMS code: LI-T | Chemical | Total | LAB
LI-T | ug/g wet | Lithium Total | X560 | Metals in Biota (wet), total, acid digestion, ICP-MS | Five grams of wet, homogenized biota is weighed in a centrifuge tube, and 2.5 mL of Nitric Acid and 2.5 mL of Hydrochloric acid are added. The sample is digested at room temperature for one hour and then digested at 90-95 degrees C in a water bath or heat block for one more hour. The sample is cooled and the volume is made to 50 mL with DI. The digestate is then diluted at ten fold and introduced to a collision/reaction cell inductively coupled argon plasma mass spectrometer (CRC ICPMS) by pneumatic nebulization. Each respective element is measured at a specific mass to charge ratio, m/z value expressed in amu (atomic mass units). - Maxxam Analytics (BC). Added to EMS Dec 10, 2008. | 0.1 | Y | Y | Y | 179 | 8/13/2014 | 10/30/2023 | 11/28/2023 2:21:33 PM | 2015-03-11 | 3/11/2015 4:17:10 PM | RDGIBSON | EMS | MassPerMass | µg/g | Y | (mass conc.) | (wet) | Lithium Total (wet) (mass conc.) |  | Metals - Total; | Metals - Total; EMS code: LI-T | Chemical | Total | LAB
LI-T | ug/g | Lithium Total | FS01 | D&G,Nitric/HCl:ICP | Samples are dried at 60 C, then either sieved to -100mesh (150 um) and rolled to homogenize, or finely ground with a mortar and pestle or a shatter box laboratory mill and rolled (depending on the submitter's requirements). The samples are then weighed into Teflon digestion liners and digested with HNO3 and HCl in a microwave oven. Samples are cooled, volumized, and allowed to settle overnight before being decanted. Analysis is then performed via Inductively Coupled Argon Plasma Atomic Emission Spectrometry (ICP-AES) (Simultaneous multi element analysis). In ICP Spectrometry, the liquid sample is pumped to the nebulizer, where it is convertd into an aerosol. the nebulizer carrier gas flow than transports the wet aerosol past the impact bead, where the large droplets condense out and flow to the waste container. The remaining fine aerosol is carried into the plasma. There, the high temperature desolvates the aerosol and dissociates the sample material into individual atoms and ions, which are excited to emit light at wavelengths characteristic of the elements present. A diffraction grating in the spectrometer diffracts the various wavelengths of light through exit slits to the photo-multiplier tubes, which convert the light energy to electronic signals. | 0 | Y | Y | Y | 6 | 6/13/2003 2:41:00 PM | 12/1/2010 | 12/9/2010 3:55:27 PM | 1997-03-19 | 6/3/1998 6:51:46 PM | MPAK |  | MassPerMass | mg/kg | Y | (mass conc.) |  | Lithium Total (mass conc.) |  | Metals - Total; | Metals - Total; EMS code: LI-T | Chemical | Total | LAB
LI-T | ug/g wet | Lithium Total | BPSY | MET-WET-BPSY-CCMS-VA | Metals in Animal Tissue and Vegetation (Biota) - Prescriptive. EPA 200.3/6020B This method is conducted following British Columbia Lab Manual method "Metals in Animal Tissue and Vegetation (Biota) - Prescriptive". Tissue samples are digested with nitric and hydrochloric acids; in combination with addition of hydrogen peroxide. Instrumental analysis is by collision cell inductively coupled plasma - mass spectrometry (modified from EPA Method 6020B). | 0.1 | Y | Y | Y | 2 | 11/7/2019 | 11/7/2019 | 6/10/2021 1:57:10 PM | 2021-06-04 |  | EMS |  | MassPerMass | µg/g | Y | (mass conc.) | (wet) | Lithium Total (wet) (mass conc.) |  | Metals - Total; | Metals - Total; EMS code: LI-T | Chemical | Total | LAB
LI04 | None | Langelier index (@ 4C) | XCAL | Calculation (mg/L) | Calculation (mg/L) | 0.01 | Y | Y | Y | 1 | 5/22/2024 | 5/22/2024 | 5/30/2024 1:43:07 PM | 2024-05-24 |  | RDGIBSON |  | None | None | N | (unitless) |  | Langelier index (@ 4C) (unitless) |  | Physical Properties; | Physical Properties; EMS code: LI04 | Physical |  | LAB
LI15 | None | Langelier index (@ 15C) | XCAL | Calculation (mg/L) | Calculation (mg/L) | 0.01 | Y | Y | Y | 1 | 5/22/2024 | 5/22/2024 | 5/30/2024 1:43:07 PM | 2024-05-24 |  | RDGIBSON |  | None | None | N | (unitless) |  | Langelier index (@ 15C) (unitless) |  | Physical Properties; | Physical Properties; EMS code: LI15 | Physical |  | LAB
LI20 | None | Langelier index (@ 20C) | XCAL | Calculation (mg/L) | Calculation (mg/L) | 0.01 | Y | Y | Y | 1 | 5/22/2024 | 5/22/2024 | 5/30/2024 1:43:07 PM | 2024-05-24 |  | RDGIBSON |  | None | None | N | (unitless) |  | Langelier index (@ 20C) (unitless) |  | Physical Properties; | Physical Properties; EMS code: LI20 | Physical |  | LAB
LI25 | None | Langelier index (@ 25C) | XCAL | Calculation (mg/L) | Calculation (mg/L) | 0.01 | Y | Y | Y | 1 | 5/22/2024 | 5/22/2024 | 5/30/2024 1:43:07 PM | 2024-05-24 |  | RDGIBSON |  | None | None | N | (unitless) |  | Langelier index (@ 25C) (unitless) |  | Physical Properties; | Physical Properties; EMS code: LI25 | Physical |  | LAB
LI375 | % | Loss on ignition @ 375?C | E205B | CSSS (2008) 28.3 (mod) | CSSS (2008) 28.3 (mod) | 1 | Y | Y | Y | 1 | 2/8/2023 | 2/8/2023 | 2/22/2023 12:19:06 PM | 2023-02-14 |  | EMS |  | DimensionlessRatio | % | N | (ratio) |  | Loss on ignition @ 375C (ratio) |  | Physical Properties; | Physical Properties; EMS code: LI375 | Physical |  | LAB
LI550 | % | Loss on ignition at 550?C | E205B | CSSS (2008) 28.3 (mod) | CSSS (2008) 28.3 (mod) | 1 | Y | Y | Y | 0 |  |  |  | 2023-12-12 |  | EMS |  | DimensionlessRatio | % | N | (ratio) |  | Loss on ignition at 550C (ratio) |  | Physical Properties; | Physical Properties; EMS code: LI550 | Physical |  | LAB
LI60 | None | Langelier index (@ 60C) | XCAL | Calculation (mg/L) | Calculation (mg/L) | 0.01 | Y | Y | Y | 1 | 5/22/2024 | 5/22/2024 | 5/30/2024 1:43:07 PM | 2024-05-24 |  | RDGIBSON |  | None | None | N | (unitless) |  | Langelier index (@ 60C) (unitless) |  | Physical Properties; | Physical Properties; EMS code: LI60 | Physical |  | LAB
LI77 | None | Langelier index (@77C) | XCAL | Calculation (mg/L) | Calculation (mg/L) | 0.01 | Y | Y | Y | 0 |  |  |  | 2024-05-24 |  | RDGIBSON |  | None | None | N | (unitless) |  | Langelier index (@77C) (unitless) |  | Physical Properties; | Physical Properties; EMS code: LI77 | Physical |  | LAB
LIDC | mg/L | Lidocaine CAS: 137-58-6 | XMSP | HPLC and Tri. Quadpole MS/MS | EPA 1694 An aliquot of filtered sample containing 5% organic is injected directly. Each compound is separated by reversed phase HPLC. Analyzed by Electro Spray Ionization MS/MS detection using Triple Quadrupole MS/MS detector. Quantiitation with internal or external standard. |  | Y | Y | Y | 12 | 5/8/2020 | 6/4/2024 | 8/21/2024 3:32:03 PM | 2022-05-04 |  | EMS |  | Concentration | mg/L | Y | (fl. conc.) |  | Lidocaine (fl. conc.) | 137-58-6 | Pharmaceuticals and Personal Care Products (PPCP); | Pharmaceuticals and Personal Care Products (PPCP); EMS code: LIDC | Chemical |  | LAB
LINC | mg/L | Lincomycin CAS: 154-21-2 | XMSP | HPLC and Tri. Quadpole MS/MS | EPA 1694 An aliquot of filtered sample containing 5% organic is injected directly. Each compound is separated by reversed phase HPLC. Analyzed by Electro Spray Ionization MS/MS detection using Triple Quadrupole MS/MS detector. Quantiitation with internal or external standard. |  | Y | Y | Y | 12 | 5/8/2020 | 6/4/2024 | 8/21/2024 3:32:03 PM | 2022-05-04 |  | EMS |  | Concentration | mg/L | Y | (fl. conc.) |  | Lincomycin (fl. conc.) | 154-21-2 | Pharmaceuticals and Personal Care Products (PPCP); | Pharmaceuticals and Personal Care Products (PPCP); EMS code: LINC | Chemical |  | LAB
LIPI | ug/g | Lipids | X232 | CHCl3/CH3OH Ext,Evap | Lipid material is extracted from the tissue with a suitable solvent, the solvent is removed by heating and the residue is determined gravimetrically. | 0.1 | Y | Y | Y | 70 |  |  | 10/24/1996 10:01:18 AM | 1996-04-12 | 6/3/1998 6:51:42 PM | EMS |  | MassPerMass | mg/kg | Y | (mass conc.) |  | Lipids (mass conc.) |  | Organics; | Organics; EMS code: LIPI | Chemical |  | LAB
LIPI | % | Lipids | X269 | CHCl3/CH3OH Ext:Evap. | Lipid material is extracted from the tissue with a suitable solvent, the solvent is removed by heating and the residue is determined gravimetrically. | 0.1 | Y | Y | Y | 133 | 12/23/2005 | 12/1/2011 3:25:00 PM | 12/6/2011 3:50:33 PM | 1996-04-12 | 6/3/1998 6:51:36 PM | EMS |  | DimensionlessRatio | % | N | (ratio) |  | Lipids (ratio) |  | Organics; | Organics; EMS code: LIPI | Chemical |  | LAB
LOSS | % | Loss on ignition @ 1000?C | ICP06 | ALS Minerals ME-ICP06 | ALS Minerals ME-ICP06. A prepared sample is fused with a lithium tetraborate/lithium metaborate flux and poured into a graphite crucible and fused in a furnace at 1000°C. The resulting melt is then cooled and dissolved in 100 mL of 4% HNO3/2% HCl followed by ICP-AES. | 0.01 | Y | Y | Y | 0 |  |  |  | 2024-01-29 |  | EMS |  | DimensionlessRatio | % | N | (ratio) |  | Loss on ignition @ 1000C (ratio) |  | Physical Properties; | Physical Properties; EMS code: LOSS | Physical |  | LAB
LOSS | None | Loss on ignition @ 1000?C | GRAV05 | ALS Minerals OA-GRA05 | ALS Minerals OA-GRA05. The determination of the mass loss from solid combustion residues upon heating in an air or oxygen atmosphere to a prescribed temperature. The mass loss can be due to the loss of moisture; carbon; sulfur; and so forth; from the decomposition or combustion of the residue. | 0.01 | Y | Y | Y | 0 |  |  |  | 2024-01-29 |  | RDGIBSON |  | None | None | N | (unitless) |  | Loss on ignition @ 1000C (unitless) |  | Physical Properties; | Physical Properties; EMS code: LOSS | Physical |  | LAB
LSI- | None | Langerlier's Saturation Index | TCAL | Calculation | Calculation | 0 | Y | Y | Y | 44 | 5/8/2002 | 8/6/2019 4:15:00 PM | 8/7/2019 5:01:07 PM | 2002-07-11 |  | NPEPPIN |  | None | None | N | (unitless) |  | Langerlier's Saturation Index (unitless) |  | Major Ions; | Major Ions; EMS code: LSI- | Chemical |  | LAB
LU-T | ppm (S) | Lutetium | MS81 | ALS Minerals: ME-MS81 | ALS Minerals: ME-MS81. A prepared sample (0.100 g) is added to lithium metaborate/lithium tetraborate flux; mixed well and fused in a furnace at 1025°C. The resulting melt is then cooled and dissolved in an acid mixture containing nitric; hydrochloric and hydrofluoric acids. This solution is then analyzed by inductively coupled plasma - mass spectrometry. | 0.01 | Y | Y | Y | 0 |  |  |  | 2024-01-29 |  | EMS |  | DimensionlessRatio | ppm | N | (ratio) |  | Lutetium (ratio) |  | Metals; | Metals; EMS code: LU-T | Chemical |  | LAB
M001 | mg/L | Malathion | ZN01 | DCM+Ac.Extr:GC/MS | DCM+Ac.Extr:GC/MS | 0.0005 | Y | Y | Y | 52 | 10/3/2002 | 3/29/2012 10:15:00 PM | 5/3/2012 9:27:51 AM | 1996-04-12 | 11/16/1999 4:06:27 PM | EMS | NPEPPIN | Concentration | mg/L | Y | (fl. conc.) |  | Malathion (fl. conc.) | 121-75-5 | Pesticides; | Pesticides; EMS code: M001 | Chemical |  | LAB
M001 | mg/L | Malathion | F053 | DCM extr'n,conc'n,HGRC-PFPD/N,P det'n | Effluent samples are extracted with Dichloromethane. The extract is then concentrated and analysed using High Resolution Gas Chromatography with Pulsed Flame Photometric Dectection. Confirmation, if required, is by High Resolution Gas Chromatography with Nitrogen Phosphorus Detection. | 5e-05 | Y | Y | Y | 31 | 3/3/2000 7:36:00 AM | 1/22/2002 10:26:00 AM | 6/18/2003 2:53:45 PM | 1997-10-20 | 6/3/1998 6:51:47 PM | NPEPPIN |  | Concentration | mg/L | Y | (fl. conc.) |  | Malathion (fl. conc.) | 121-75-5 | Pesticides; | Pesticides; EMS code: M001 | Chemical |  | LAB
M001 | ug/g | Malathion | X260 | Sox-Ac/CH2Cl2 ext:GC/NPD | Sox-Ac/CH2Cl2 ext:GC/NPD | 0.01 | Y | Y | Y | 117 | 4/20/2002 10:25:00 AM | 10/16/2002 3:20:00 PM | 2/26/2004 1:05:28 AM | 1996-04-12 | 6/3/1998 6:51:36 PM | EMS |  | MassPerMass | mg/kg | Y | (mass conc.) |  | Malathion (mass conc.) | 121-75-5 | Pesticides; | Pesticides; EMS code: M001 | Chemical |  | LAB
M001 | mg/L | Malathion | EQ01 | EQUIS (mg/L) | EQUIS (mg/L) | 0 | N | Y | Y | 2 |  |  | 10/22/1996 5:23:43 PM | 1996-04-12 | 6/3/1998 6:51:36 PM | EMS |  | Concentration | mg/L | Y | (fl. conc.) |  | Malathion (fl. conc.) | 121-75-5 | Pesticides; | Pesticides; EMS code: M001 | Chemical |  | LAB
M001 | ug/g | Malathion | ZN03 | DCM+Ac.Extr:GC/MS | DCM+Ac.Extr:GC/MS EPA 8270E (mod) | 0.05 | Y | Y | Y | 16 | 1/25/2022 | 11/2/2022 4:41:00 AM | 1/7/2023 3:12:10 PM | 1996-04-12 | 11/16/1999 2:27:40 PM | EMS | NPEPPIN | MassPerMass | mg/kg | Y | (mass conc.) |  | Malathion (mass conc.) | 121-75-5 | Pesticides; | Pesticides; EMS code: M001 | Chemical |  | LAB
M001 | mg/L | Malathion | E660 | EPA 8270E (mod) | EPA 8270E (mod) Miscellaneous Pesticides by GC-MS | 0.0001 | Y | Y | Y | 68 | 12/7/2021 | 9/12/2022 | 12/8/2022 12:57:05 PM | 2021-12-09 | 12/9/2021 3:22:12 PM | EMS | EMS | Concentration | mg/L | Y | (fl. conc.) |  | Malathion (fl. conc.) | 121-75-5 | Pesticides; | Pesticides; EMS code: M001 | Chemical |  | LAB
M002 | mg/L | Methamidaphos | ZN01 | DCM+Ac.Extr:GC/MS | DCM+Ac.Extr:GC/MS | 0.0005 | Y | Y | Y | 1 | 7/29/2005 7:10:00 AM | 7/29/2005 7:10:00 AM | 6/2/2009 12:01:07 PM | 1996-04-12 | 11/16/1999 4:06:27 PM | EMS | NPEPPIN | Concentration | mg/L | Y | (fl. conc.) |  | Methamidaphos (fl. conc.) | 10265-92-6 | Pesticides; | Pesticides; EMS code: M002 | Chemical |  | LAB
M002 | ug/g | Methamidaphos | X260 | Sox-Ac/CH2Cl2 ext:GC/NPD | Sox-Ac/CH2Cl2 ext:GC/NPD | 0.05 | Y | Y | Y | 97 | 4/20/2002 10:25:00 AM | 10/16/2002 3:20:00 PM | 2/25/2004 1:09:46 AM | 1996-04-12 | 6/3/1998 6:51:36 PM | EMS |  | MassPerMass | mg/kg | Y | (mass conc.) |  | Methamidaphos (mass conc.) | 10265-92-6 | Pesticides; | Pesticides; EMS code: M002 | Chemical |  | LAB
M003 | mg/L | Mevinphos-cis | ZN01 | DCM+Ac.Extr:GC/MS | DCM+Ac.Extr:GC/MS | 0.0001 | Y | Y | Y | 4 | 10/3/2002 | 10/10/2003 | 3/24/2004 1:04:40 AM | 1996-04-12 | 11/16/1999 4:06:27 PM | EMS | NPEPPIN | Concentration | mg/L | Y | (fl. conc.) |  | Mevinphos-cis (fl. conc.) | 298-01-1 | Pesticides; | Pesticides; EMS code: M003 | Chemical |  | LAB
M003 | ug/g | Mevinphos-cis | X260 | Sox-Ac/CH2Cl2 ext:GC/NPD | Sox-Ac/CH2Cl2 ext:GC/NPD | 0.05 | Y | Y | Y | 117 | 4/20/2002 10:25:00 AM | 10/16/2002 3:20:00 PM | 2/26/2004 1:05:28 AM | 1996-04-12 | 6/3/1998 6:51:36 PM | EMS |  | MassPerMass | mg/kg | Y | (mass conc.) |  | Mevinphos-cis (mass conc.) | 298-01-1 | Pesticides; | Pesticides; EMS code: M003 | Chemical |  | LAB
M004 | mg/L | Metalaxyl | F056 | DCM extr,HRGC-NPD | Effluent samples are extracted with Dichloromethane. The extract is concentrated and analysed using High Resolution Gas Chromatography with a Thermionic Specific Detector (TSD), alternately referred to as a Nitrogen Phosphorus Detector (NPD). | 0.005 | Y | Y | Y | 12 | 4/3/2000 10:51:00 AM | 2/16/2001 3:36:00 PM | 9/11/2002 9:34:57 AM | 1998-10-02 |  | NPEPPIN |  | Concentration | mg/L | Y | (fl. conc.) |  | Metalaxyl (fl. conc.) | 57837-19-1 | Pesticides; | Pesticides; EMS code: M004 | Chemical |  | LAB
M004 | mg/L | Metalaxyl | ZN01 | DCM+Ac.Extr:GC/MS | DCM+Ac.Extr:GC/MS | 0.0003 | Y | Y | Y | 52 | 10/3/2002 | 3/29/2012 10:15:00 PM | 5/3/2012 9:27:51 AM | 1996-04-12 | 11/16/1999 4:06:27 PM | EMS | NPEPPIN | Concentration | mg/L | Y | (fl. conc.) |  | Metalaxyl (fl. conc.) | 57837-19-1 | Pesticides; | Pesticides; EMS code: M004 | Chemical |  | LAB
M004 | ug/g | Metalaxyl | FS61 | DCM +MeOH extr,HRGC-NPD | Sediment samples are shaken with a mixture of Dichloromethane and Methanol, filtered then re-extracted with Dichloromethane. The extract is concentrated and analysed using High Resolution Gas Chromatography with a Thermionic Specific Detector (TSD), alternately referred to as a Nitrogen Phosphorus Detector (NPD). | 0.01 | Y | Y | Y | 34 | 4/20/2002 3:40:00 PM | 7/16/2002 3:20:00 PM | 2/26/2004 1:05:28 AM | 1998-10-02 |  | NPEPPIN |  | MassPerMass | mg/kg | Y | (mass conc.) |  | Metalaxyl (mass conc.) | 57837-19-1 | Pesticides; | Pesticides; EMS code: M004 | Chemical |  | LAB
M004 | mg/L | Metalaxyl | E751 | E751D Pesticides in Water by LC-MS/MS (POS1 List) | E751D Pesticides in Water by LC-MS/MS (POS1 List) (MOECC E3501) | 5e-05 | Y | Y | Y | 39 | 12/9/2021 | 12/31/2021 | 2/15/2022 9:50:05 AM | 2021-12-14 |  | EMS |  | Concentration | mg/L | Y | (fl. conc.) |  | Metalaxyl (fl. conc.) | 57837-19-1 | Pesticides; | Pesticides; EMS code: M004 | Chemical |  | LAB
M005 | mg/L | Metolachlor | ZN01 | DCM+Ac.Extr:GC/MS | DCM+Ac.Extr:GC/MS | 0.0002 | Y | Y | Y | 28 | 7/28/2005 7:03:00 AM | 3/29/2012 10:15:00 PM | 5/3/2012 9:27:51 AM | 1996-04-12 | 11/16/1999 4:06:27 PM | EMS | NPEPPIN | Concentration | mg/L | Y | (fl. conc.) |  | Metolachlor (fl. conc.) | 51218-45-2 | Pesticides; | Pesticides; EMS code: M005 | Chemical |  | LAB
M005 | mg/L | Metolachlor | E660 | EPA 8270E (mod) | EPA 8270E (mod) Miscellaneous Pesticides by GC-MS | 0.0001 | Y | Y | Y | 68 | 12/7/2021 | 9/12/2022 | 12/8/2022 12:57:05 PM | 2021-12-09 | 12/9/2021 3:22:12 PM | EMS | EMS | Concentration | mg/L | Y | (fl. conc.) |  | Metolachlor (fl. conc.) | 51218-45-2 | Pesticides; | Pesticides; EMS code: M005 | Chemical |  | LAB
M005 | ug/g | Metolachlor | ZN03 | DCM+Ac.Extr:GC/MS | DCM+Ac.Extr:GC/MS EPA 8270E (mod) | 0.02 | Y | Y | Y | 16 | 1/25/2022 | 11/2/2022 4:41:00 AM | 1/7/2023 3:12:10 PM | 1996-04-12 | 11/16/1999 2:27:40 PM | EMS | NPEPPIN | MassPerMass | mg/kg | Y | (mass conc.) |  | Metolachlor (mass conc.) | 51218-45-2 | Pesticides; | Pesticides; EMS code: M005 | Chemical |  | LAB
M006 | mg/L | Metribuzin | ZN01 | DCM+Ac.Extr:GC/MS | DCM+Ac.Extr:GC/MS | 0.0003 | Y | Y | Y | 28 | 7/28/2005 7:03:00 AM | 3/29/2012 10:15:00 PM | 5/3/2012 9:27:51 AM | 1996-04-12 | 11/16/1999 4:06:27 PM | EMS | NPEPPIN | Concentration | mg/L | Y | (fl. conc.) |  | Metribuzin (fl. conc.) | 21087-64-9 | Pesticides; | Pesticides; EMS code: M006 | Chemical |  | LAB
M006 | mg/L | Metribuzin | E660 | EPA 8270E (mod) | EPA 8270E (mod) Miscellaneous Pesticides by GC-MS | 0.0001 | Y | Y | Y | 68 | 12/7/2021 | 9/12/2022 | 12/8/2022 12:57:05 PM | 2021-12-09 | 12/9/2021 3:22:12 PM | EMS | EMS | Concentration | mg/L | Y | (fl. conc.) |  | Metribuzin (fl. conc.) | 21087-64-9 | Pesticides; | Pesticides; EMS code: M006 | Chemical |  | LAB
M006 | ug/g | Metribuzin | ZN03 | DCM+Ac.Extr:GC/MS | DCM+Ac.Extr:GC/MS EPA 8270E (mod) | 0.03 | Y | Y | Y | 16 | 1/25/2022 | 11/2/2022 4:41:00 AM | 1/7/2023 3:12:10 PM | 1996-04-12 | 11/16/1999 2:27:40 PM | EMS | NPEPPIN | MassPerMass | mg/kg | Y | (mass conc.) |  | Metribuzin (mass conc.) | 21087-64-9 | Pesticides; | Pesticides; EMS code: M006 | Chemical |  | LAB
M009 | mg/L | Malaoxon | ZN01 | DCM+Ac.Extr:GC/MS | DCM+Ac.Extr:GC/MS | 0.0003 | Y | Y | Y | 28 | 7/29/2005 7:10:00 AM | 3/29/2012 10:15:00 PM | 5/3/2012 9:27:51 AM | 1996-04-12 | 11/16/1999 4:06:27 PM | EMS | NPEPPIN | Concentration | mg/L | Y | (fl. conc.) |  | Malaoxon (fl. conc.) | 1634-78-2 | Pesticides; | Pesticides; EMS code: M009 | Chemical |  | LAB
M010 | mg/L | MCPA | AEWA | H2SO4 acid,CHCl2 extr.florisil. GC/MS | Effluent and water samples are acidified with Sulphuric Acid and extracted with Dichloromethane. After concentration, the extract is methylated using ethereal Diazomethane. Extracts are cleaned up as required using Florisil. The extracts are analyzed for Acid Extractable Herbicides using GC/MS. | 0.001 | Y | Y | Y | 6 | 6/16/2003 | 7/29/2005 7:10:00 AM | 6/2/2009 12:01:07 PM | 1998-11-17 |  | NPEPPIN |  | Concentration | mg/L | Y | (fl. conc.) |  | MCPA (fl. conc.) | 94-74-6 | Pesticides; | Pesticides; EMS code: M010 | Chemical |  | LAB
M010 | mg/L | MCPA | P015 | Phenoxyacid Herbicides by LC-MS/MS | Phenoxyacid Herbicides by LC-MS/MS | 0.5 | Y | Y | Y | 71 | 10/9/2019 | 9/9/2022 | 12/8/2022 12:57:05 PM | 2019-10-17 |  | KATFRASE |  | Concentration | mg/L | Y | (fl. conc.) |  | MCPA (fl. conc.) | 94-74-6 | Pesticides; | Pesticides; EMS code: M010 | Chemical |  | LAB
M010 | ug/g | MCPA | FS15 | DCM/MEOH acid extr'n,DCM,methn,Florisil,HRGC-EC | Sediment samples are shaken with a mixture of Dichloromethane, Methanol and Sulphuric Acid, filtered, then extracted with Dichloromethane. After concentration, the extract is methylated using ethereal Diazomethane. Extracts are cleaned up as required using Florisil. The extracts are analysed for Acid Extractable Herbicides using High Resolution Gas Chromatography with Electron Capture detection. Positive results may be verified by use of a secondary confirmation GC column. | 0.02 | Y | Y | Y | 27 | 4/4/2002 | 6/24/2003 | 2/26/2004 1:05:28 AM | 1998-11-17 |  | NPEPPIN |  | MassPerMass | mg/kg | Y | (mass conc.) |  | MCPA (mass conc.) | 94-74-6 | Pesticides; | Pesticides; EMS code: M010 | Chemical |  | LAB
M010 | ug/g | MCPA | ZN03 | DCM+Ac.Extr:GC/MS | DCM+Ac.Extr:GC/MS EPA 8270E (mod) | 0.0005 | Y | Y | Y | 15 | 1/14/2022 | 2/2/2022 | 1/11/2024 1:07:08 PM | 2022-07-27 |  | EMS |  | MassPerMass | mg/kg | Y | (mass conc.) |  | MCPA (mass conc.) | 94-74-6 | Pesticides; | Pesticides; EMS code: M010 | Chemical |  | LAB
M011 | mg/L | Methomyl | P045 | DCM Extr:HPLC | DCM Extr:HPLC | 0.0005 | Y | Y | Y | 1 | 7/29/2005 7:10:00 AM | 7/29/2005 7:10:00 AM | 6/2/2009 12:01:07 PM | 1996-04-12 | 6/3/1998 6:51:42 PM | EMS |  | Concentration | mg/L | Y | (fl. conc.) |  | Methomyl (fl. conc.) | 16752-77-5 | Pesticides; | Pesticides; EMS code: M011 | Chemical |  | LAB
M013 | mg/L | Methidathion | ZN01 | DCM+Ac.Extr:GC/MS | DCM+Ac.Extr:GC/MS | 0.0003 | Y | Y | Y | 51 | 10/3/2002 | 3/29/2012 10:15:00 PM | 5/3/2012 9:27:51 AM | 1999-11-16 |  | NPEPPIN |  | Concentration | mg/L | Y | (fl. conc.) |  | Methidathion (fl. conc.) | 950-37-8 | Pesticides; | Pesticides; EMS code: M013 | Chemical |  | LAB
M013 | ug/g | Methidathion | FS16 | DCM/MEOH extr'n,DCM,conc'n,HRGC-PFPD/N,P det'n | Sediment samples are shaken with a mixture of Dichloromethane and Methanol, filtered, then extracted with Dichloromethane. The extract is then concentrated and analysed using High Resolution Gas Chromatography with Pulsed Flame Photometric Detection. Confirmation, if required, is by High Resolution Gas Chromatography with Nitrogen Phosphorus Detection. | 0.05 | Y | Y | Y | 1 | 4/24/2002 4:00:00 PM | 4/24/2002 4:00:00 PM | 2/12/2004 1:02:36 AM | 1997-10-20 | 6/3/1998 6:51:47 PM | NPEPPIN |  | MassPerMass | mg/kg | Y | (mass conc.) |  | Methidathion (mass conc.) | 950-37-8 | Pesticides; | Pesticides; EMS code: M013 | Chemical |  | LAB
M013 | ug/g | Methidathion | ZN03 | DCM+Ac.Extr:GC/MS | DCM+Ac.Extr:GC/MS EPA 8270E (mod) | 0.03 | Y | Y | Y | 38 | 4/20/2002 10:25:00 AM | 10/16/2002 3:20:00 PM | 2/26/2004 1:05:28 AM | 1996-04-12 | 11/16/1999 2:27:40 PM | EMS | NPEPPIN | MassPerMass | mg/kg | Y | (mass conc.) |  | Methidathion (mass conc.) | 950-37-8 | Pesticides; | Pesticides; EMS code: M013 | Chemical |  | LAB
M013 | mg/L | Methidathion | F053 | DCM extr'n,conc'n,HGRC-PFPD/N,P det'n | Effluent samples are extracted with Dichloromethane. The extract is then concentrated and analysed using High Resolution Gas Chromatography with Pulsed Flame Photometric Dectection. Confirmation, if required, is by High Resolution Gas Chromatography with Nitrogen Phosphorus Detection. | 5e-05 | Y | Y | Y | 31 | 3/3/2000 7:36:00 AM | 1/22/2002 10:26:00 AM | 6/18/2003 2:53:45 PM | 1997-10-20 | 6/3/1998 6:51:47 PM | NPEPPIN |  | Concentration | mg/L | Y | (fl. conc.) |  | Methidathion (fl. conc.) | 950-37-8 | Pesticides; | Pesticides; EMS code: M013 | Chemical |  | LAB
M015 | mg/L | Methiocarb | P045 | DCM Extr:HPLC | DCM Extr:HPLC | 0.0005 | Y | Y | Y | 1 | 7/29/2005 7:10:00 AM | 7/29/2005 7:10:00 AM | 6/2/2009 12:01:07 PM | 1996-04-12 | 6/3/1998 6:51:42 PM | EMS |  | Concentration | mg/L | Y | (fl. conc.) |  | Methiocarb (fl. conc.) | 2032-65-7 | Pesticides; | Pesticides; EMS code: M015 | Chemical |  | LAB
M016 | ug/g wet | Methoxychlor | X220 | Hex/AcExt,Part,Chrom,GC | Hex/AcExt;Part;Chrom;GC | 0.005 | Y | Y | Y | 72 |  |  | 10/24/1996 10:01:25 AM | 1996-04-12 | 6/3/1998 6:51:42 PM | EMS |  | MassPerMass | µg/g | Y | (mass conc.) | (wet) | Methoxychlor (wet) (mass conc.) | 72-43-5 | Pesticides; | Pesticides; EMS code: M016 | Chemical |  | LAB
M016 | mg/L | Methoxychlor | E660 | EPA 8270E (mod) | EPA 8270E (mod) Miscellaneous Pesticides by GC-MS | 2e-06 | Y | Y | Y | 2 | 9/14/2022 | 9/14/2022 | 12/8/2022 9:48:14 AM | 2022-12-06 |  | EMS |  | Concentration | mg/L | Y | (fl. conc.) |  | Methoxychlor (fl. conc.) | 72-43-5 | Pesticides; | Pesticides; EMS code: M016 | Chemical |  | LAB
M016 | mg/L | Methoxychlor | F027 | Hexane Extr,GC/ECD | Effluent samples are solvent extracted with Hexane and the extract dried over Sodium Sulfate, concentrated and cleaned up, as required, using Florisil, activated Copper metal, concentrated Sulfuric Acid and Mercury, to remove interferences. The extracts are then analysed using megabore column Gas Chromatography with Electron Capture detection. | 5e-05 | Y | Y | Y | 20 | 3/3/2000 7:36:00 AM | 1/22/2002 10:58:00 AM | 6/18/2003 2:53:46 PM | 2002-06-05 |  | UMULLER |  | Concentration | mg/L | Y | (fl. conc.) |  | Methoxychlor (fl. conc.) | 72-43-5 | Pesticides; | Pesticides; EMS code: M016 | Chemical |  | LAB
M016 | mg/L | Methoxychlor | X268 | CH2Cl2 ext:GC/MS | CH2Cl2 ext:GC/MS |  | Y | Y | Y | 16 | 12/29/2021 | 9/2/2022 | 11/10/2022 9:54:07 AM | 2022-04-08 |  | EMS |  | Concentration | mg/L | Y | (fl. conc.) |  | Methoxychlor (fl. conc.) | 72-43-5 | Pesticides; | Pesticides; EMS code: M016 | Chemical |  | LAB
M016 | mg/L | Methoxychlor | MW03 | OCP1-LL-SF-ECD-VA | This analysis is carried out using procedures adapted from "Test Methods for Evaluating Solid Waste" SW-846, Methods 3510, 3630, 3660 and 8081, published by the United States Environmental Protection Agency (EPA). The procedure involves extraction of the entire water sample with dichloromethane. The extract then undergoes one or more of the following clean-up procedures (if required): silica gel clean-up and/or sulphur clean-up. The final extract is analyzed by dual capillary column gas chromatography with electron capture detection (GC/ECD). | 3e-05 | Y | Y | Y | 10 | 9/20/2016 | 9/21/2016 | 8/15/2019 2:42:34 PM | 2019-08-15 |  | KATFRASE |  | Concentration | mg/L | Y | (fl. conc.) |  | Methoxychlor (fl. conc.) | 72-43-5 | Pesticides; | Pesticides; EMS code: M016 | Chemical |  | LAB
M016 | ug/g | Methoxychlor | X213 | Hex/AcExt,Part,Chrom,GC | Hex/AcExt;Part;Chrom;GC | 0.005 | Y | Y | Y | 187 | 4/20/2002 10:10:00 AM | 10/13/2002 10:30:00 AM | 2/26/2004 1:05:29 AM | 1996-04-12 | 6/3/1998 6:51:36 PM | EMS |  | MassPerMass | mg/kg | Y | (mass conc.) |  | Methoxychlor (mass conc.) | 72-43-5 | Pesticides; | Pesticides; EMS code: M016 | Chemical |  | LAB
M016 | mg/L | Methoxychlor | EQ01 | EQUIS (mg/L) | EQUIS (mg/L) | 0.0001 | N | Y | Y | 17 |  |  | 10/22/1996 11:26:54 PM | 1996-04-12 | 6/3/1998 6:51:36 PM | EMS |  | Concentration | mg/L | Y | (fl. conc.) |  | Methoxychlor (fl. conc.) | 72-43-5 | Pesticides; | Pesticides; EMS code: M016 | Chemical |  | LAB
M016 | mg/L | Methoxychlor | ZN01 | DCM+Ac.Extr:GC/MS | DCM+Ac.Extr:GC/MS | 0.0001 | Y | Y | Y | 52 | 10/3/2002 | 3/29/2012 10:15:00 PM | 5/3/2012 9:27:51 AM | 1996-04-12 | 11/16/1999 4:06:27 PM | EMS | NPEPPIN | Concentration | mg/L | Y | (fl. conc.) |  | Methoxychlor (fl. conc.) | 72-43-5 | Pesticides; | Pesticides; EMS code: M016 | Chemical |  | LAB
M016 | ug/g | Methoxychlor | ZN03 | DCM+Ac.Extr:GC/MS | DCM+Ac.Extr:GC/MS EPA 8270E (mod) | 0.01 | Y | Y | Y | 19 | 1/17/2022 | 11/14/2022 | 1/11/2024 1:07:08 PM | 1996-04-12 | 11/16/1999 2:27:40 PM | EMS | NPEPPIN | MassPerMass | mg/kg | Y | (mass conc.) |  | Methoxychlor (mass conc.) | 72-43-5 | Pesticides; | Pesticides; EMS code: M016 | Chemical |  | LAB
M016 | ug/g | Methoxychlor | E8081 | EPA 8081 | See methods EPA 3540, 3630, 3660, 8081 | 0.005 | Y | Y | Y | 9 | 3/18/2020 | 3/23/2020 | 2/22/2023 12:51:03 PM | 2022-07-29 |  | EMS |  | MassPerMass | mg/kg | Y | (mass conc.) |  | Methoxychlor (mass conc.) | 72-43-5 | Pesticides; | Pesticides; EMS code: M016 | Chemical |  | LAB
M017 | mg/L | Methyl Mercaptan | X239 | Gas Chromatography | Gas Chromatography | 0.005 | Y | Y | Y | 728 |  |  | 10/24/1996 12:58:44 PM | 1996-04-12 | 6/3/1998 6:51:36 PM | EMS |  | Concentration | mg/L | Y | (fl. conc.) |  | Methyl Mercaptan (fl. conc.) | 74-93-1 | Volatile Organic Compounds (VOCs); | Volatile Organic Compounds (VOCs); EMS code: M017 | Chemical |  | LAB
M018 | ug/g | Methyl Parathion | ZN03 | DCM+Ac.Extr:GC/MS | DCM+Ac.Extr:GC/MS EPA 8270E (mod) | 0.05 | Y | Y | Y | 2 | 11/2/2022 12:53:00 AM | 11/2/2022 4:41:00 AM | 1/7/2023 3:12:11 PM | 1996-04-12 | 6/3/1998 6:51:36 PM | EMS |  | MassPerMass | mg/kg | Y | (mass conc.) |  | Methyl Parathion (mass conc.) | 298-00-0 | Pesticides; | Pesticides; EMS code: M018 | Chemical |  | LAB
M020 | mg/L | Methanol | E-8015 | EPA 8015D | EPA Method 8015D (SW-846): Nonhalogenated Organics Using GC/FID |  | Y | Y | Y | 14 | 12/29/2021 | 1/12/2022 | 9/2/2022 6:21:53 PM | 2022-04-06 |  | EMS |  | Concentration | mg/L | Y | (fl. conc.) |  | Methanol (fl. conc.) | 67-56-1 | Volatile Organic Compounds (VOCs); | Volatile Organic Compounds (VOCs); EMS code: M020 | Chemical |  | LAB
M020 | ug/g | Methanol | E8015S | EPA 8015 D in soil | See EPA 8015D | 0.5 | Y | Y | Y | 15 | 1/12/2022 | 2/1/2022 | 1/11/2024 1:07:08 PM | 2022-06-23 |  | EMS |  | MassPerMass | mg/kg | Y | (mass conc.) |  | Methanol (mass conc.) | 67-56-1 | Volatile Organic Compounds (VOCs); | Volatile Organic Compounds (VOCs); EMS code: M020 | Chemical |  | LAB
M020 | mg/L | Methanol | X239 | Gas Chromatography | Gas Chromatography | 0.005 | Y | Y | Y | 985 |  |  | 4/28/2016 4:27:22 PM | 1996-04-12 | 6/3/1998 6:51:45 PM | EMS |  | Concentration | mg/L | Y | (fl. conc.) |  | Methanol (fl. conc.) | 67-56-1 | Volatile Organic Compounds (VOCs); | Volatile Organic Compounds (VOCs); EMS code: M020 | Chemical |  | LAB
M024 | mg/L | Mirex | X268 | CH2Cl2 ext:GC/MS | CH2Cl2 ext:GC/MS |  | Y | Y | Y | 16 | 12/29/2021 | 9/2/2022 | 11/10/2022 9:54:07 AM | 2022-04-08 |  | EMS |  | Concentration | mg/L | Y | (fl. conc.) |  | Mirex (fl. conc.) | 2385-85-5 | Pesticides; | Pesticides; EMS code: M024 | Chemical |  | LAB
M024 | mg/L | Mirex | ZN01 | DCM+Ac.Extr:GC/MS | DCM+Ac.Extr:GC/MS | 0.0003 | Y | Y | Y | 52 | 10/3/2002 | 3/29/2012 10:15:00 PM | 5/3/2012 9:27:51 AM | 1996-04-12 | 11/16/1999 4:06:27 PM | EMS | NPEPPIN | Concentration | mg/L | Y | (fl. conc.) |  | Mirex (fl. conc.) | 2385-85-5 | Pesticides; | Pesticides; EMS code: M024 | Chemical |  | LAB
M024 | ug/g | Mirex | E8081 | EPA 8081 | See methods EPA 3540, 3630, 3660, 8081 | 2.5 | Y | Y | Y | 9 | 3/18/2020 | 3/23/2020 | 2/22/2023 12:51:03 PM | 2023-02-01 |  | EMS |  | MassPerMass | mg/kg | Y | (mass conc.) |  | Mirex (mass conc.) | 2385-85-5 | Pesticides; | Pesticides; EMS code: M024 | Chemical |  | LAB
M024 | ug/g | Mirex | ZN03 | DCM+Ac.Extr:GC/MS | DCM+Ac.Extr:GC/MS EPA 8270E (mod) | 0.03 | Y | Y | Y | 28 | 4/20/2002 10:10:00 AM | 11/14/2022 | 12/15/2022 10:50:04 AM | 1996-04-12 | 11/16/1999 2:27:40 PM | EMS | NPEPPIN | MassPerMass | mg/kg | Y | (mass conc.) |  | Mirex (mass conc.) | 2385-85-5 | Pesticides; | Pesticides; EMS code: M024 | Chemical |  | LAB
M024 | mg/L | Mirex | E660 | EPA 8270E (mod) | EPA 8270E (mod) Miscellaneous Pesticides by GC-MS | 2e-06 | Y | Y | Y | 2 | 9/14/2022 | 9/14/2022 | 12/8/2022 9:48:14 AM | 2022-12-06 |  | EMS |  | Concentration | mg/L | Y | (fl. conc.) |  | Mirex (fl. conc.) | 2385-85-5 | Pesticides; | Pesticides; EMS code: M024 | Chemical |  | LAB
M024 | mg/L | Mirex | MW03 | OCP1-LL-SF-ECD-VA | This analysis is carried out using procedures adapted from "Test Methods for Evaluating Solid Waste" SW-846, Methods 3510, 3630, 3660 and 8081, published by the United States Environmental Protection Agency (EPA). The procedure involves extraction of the entire water sample with dichloromethane. The extract then undergoes one or more of the following clean-up procedures (if required): silica gel clean-up and/or sulphur clean-up. The final extract is analyzed by dual capillary column gas chromatography with electron capture detection (GC/ECD). | 5e-05 | Y | Y | Y | 10 | 9/20/2016 | 9/21/2016 | 8/15/2019 2:42:34 PM | 2019-07-08 | 8/15/2019 11:46:08 AM | KATFRASE | KATFRASE | Concentration | mg/L | Y | (fl. conc.) |  | Mirex (fl. conc.) | 2385-85-5 | Pesticides; | Pesticides; EMS code: M024 | Chemical |  | LAB
M025 | mg/L | Mevinphos | F053 | DCM extr'n,conc'n,HGRC-PFPD/N,P det'n | Effluent samples are extracted with Dichloromethane. The extract is then concentrated and analysed using High Resolution Gas Chromatography with Pulsed Flame Photometric Dectection. Confirmation, if required, is by High Resolution Gas Chromatography with Nitrogen Phosphorus Detection. | 5e-05 | Y | Y | Y | 31 | 3/3/2000 7:36:00 AM | 1/22/2002 10:26:00 AM | 6/18/2003 2:53:46 PM | 1997-10-22 | 6/3/1998 6:51:47 PM | NPEPPIN |  | Concentration | mg/L | Y | (fl. conc.) |  | Mevinphos (fl. conc.) | 7786-34-7 | Pesticides; | Pesticides; EMS code: M025 | Chemical |  | LAB
M025 | mg/L | Mevinphos | ZN01 | DCM+Ac.Extr:GC/MS | DCM+Ac.Extr:GC/MS | 0.0005 | Y | Y | Y | 1 | 7/29/2005 7:10:00 AM | 7/29/2005 7:10:00 AM | 6/2/2009 12:01:07 PM | 2003-06-17 |  | UMULLER |  | Concentration | mg/L | Y | (fl. conc.) |  | Mevinphos (fl. conc.) | 7786-34-7 | Pesticides; | Pesticides; EMS code: M025 | Chemical |  | LAB
M026 | mg/L | MCPB | AEWA | H2SO4 acid,CHCl2 extr.florisil. GC/MS | Effluent and water samples are acidified with Sulphuric Acid and extracted with Dichloromethane. After concentration, the extract is methylated using ethereal Diazomethane. Extracts are cleaned up as required using Florisil. The extracts are analyzed for Acid Extractable Herbicides using GC/MS. | 0.0005 | Y | Y | Y | 1 | 7/29/2005 7:10:00 AM | 7/29/2005 7:10:00 AM | 6/2/2009 12:01:07 PM | 2003-06-17 |  | UMULLER |  | Concentration | mg/L | Y | (fl. conc.) |  | MCPB (fl. conc.) | 94-81-5 | Pesticides; | Pesticides; EMS code: M026 | Chemical |  | LAB
M027 | mg/L | Mecoprop | P015 | Phenoxyacid Herbicides by LC-MS/MS | Phenoxyacid Herbicides by LC-MS/MS | 0.5 | Y | Y | Y | 71 | 10/9/2019 | 9/9/2022 | 12/8/2022 12:57:05 PM | 2019-10-17 |  | KATFRASE |  | Concentration | mg/L | Y | (fl. conc.) |  | Mecoprop (fl. conc.) | 93-65-2 | Pesticides; | Pesticides; EMS code: M027 | Chemical |  | LAB
M027 | mg/L | Mecoprop | AEWA | H2SO4 acid,CHCl2 extr.florisil. GC/MS | Effluent and water samples are acidified with Sulphuric Acid and extracted with Dichloromethane. After concentration, the extract is methylated using ethereal Diazomethane. Extracts are cleaned up as required using Florisil. The extracts are analyzed for Acid Extractable Herbicides using GC/MS. | 0.0005 | Y | Y | Y | 1 | 7/29/2005 7:10:00 AM | 7/29/2005 7:10:00 AM | 6/2/2009 12:01:07 PM | 2003-06-17 |  | UMULLER |  | Concentration | mg/L | Y | (fl. conc.) |  | Mecoprop (fl. conc.) | 93-65-2 | Pesticides; | Pesticides; EMS code: M027 | Chemical |  | LAB
M027 | ug/g | Mecoprop | ZN03 | DCM+Ac.Extr:GC/MS | DCM+Ac.Extr:GC/MS EPA 8270E (mod) | 0.0005 | Y | Y | Y | 15 | 1/14/2022 | 2/2/2022 | 1/11/2024 1:07:08 PM | 2022-07-27 |  | EMS |  | MassPerMass | mg/kg | Y | (mass conc.) |  | Mecoprop (mass conc.) | 93-65-2 | Pesticides; | Pesticides; EMS code: M027 | Chemical |  | LAB
M029 | mg/L | MIBC - Methyl Isobutyl Carbinol or 4-Methyl-2-pentanol | P050 | MR's | MR's | 0.001 | Y | Y | Y | 5 |  |  | 10/23/1996 8:47:21 PM | 1996-04-12 | 6/3/1998 6:51:36 PM | EMS |  | Concentration | mg/L | Y | (fl. conc.) |  | MIBC - Methyl Isobutyl Carbinol or 4-Methyl-2-pentanol (fl. conc.) | 108-11-2 | Volatile Organic Compounds (VOCs); | Volatile Organic Compounds (VOCs); EMS code: M029 | Chemical |  | LAB
M029 | mg/L | MIBC - Methyl Isobutyl Carbinol or 4-Methyl-2-pentanol | MS04 | Static Headspace GC/MS (Water) | (Water) A static headspace procedure is used as a means of driving the analytes out of the matrix into a headspace. Once equilibrium is achieved, a fixed volume of headspeace is taken for subsequent determination by Gas Chromatography - Mass Spectroscopy (GC/MS) using Selected Ion Monitoring. | 0.001 | Y | Y | Y | 290 | 11/1/2011 | 7/31/2024 | 8/3/2024 9:18:08 AM | 2018-12-31 | 12/31/2018 1:20:42 PM | RDGIBSON | RDGIBSON | Concentration | mg/L | Y | (fl. conc.) |  | MIBC - Methyl Isobutyl Carbinol or 4-Methyl-2-pentanol (fl. conc.) | 108-11-2 | Volatile Organic Compounds (VOCs); | Volatile Organic Compounds (VOCs); EMS code: M029 | Chemical |  | LAB
M030 | mg/L | MIBK | P050 | MR's | MR's | 0.001 | Y | Y | Y | 5 |  |  | 10/23/1996 8:47:21 PM | 1996-04-12 | 6/3/1998 6:51:36 PM | EMS |  | Concentration | mg/L | Y | (fl. conc.) |  | MIBK (fl. conc.) | 108-10-1 | Organics; | Organics; EMS code: M030 | Chemical |  | LAB
M033 | mg/L | Myclobutanil | E751 | E751D Pesticides in Water by LC-MS/MS (POS1 List) | E751D Pesticides in Water by LC-MS/MS (POS1 List) (MOECC E3501) | 5e-05 | Y | Y | Y | 39 | 12/9/2021 | 12/31/2021 | 2/15/2022 9:50:05 AM | 2021-12-14 |  | EMS |  | Concentration | mg/L | Y | (fl. conc.) |  | Myclobutanil (fl. conc.) | 88671-89-0 | Pesticides; | Pesticides; EMS code: M033 | Chemical |  | LAB
M033 | mg/L | Myclobutanil | ZN01 | DCM+Ac.Extr:GC/MS | DCM+Ac.Extr:GC/MS | 0.0005 | Y | Y | Y | 1 | 7/29/2005 7:10:00 AM | 7/29/2005 7:10:00 AM | 6/2/2009 12:01:07 PM | 2003-06-17 |  | UMULLER |  | Concentration | mg/L | Y | (fl. conc.) |  | Myclobutanil (fl. conc.) | 88671-89-0 | Pesticides; | Pesticides; EMS code: M033 | Chemical |  | LAB
M041 | mg/L | Dichloromethane | X384 | Purge and trap GC/MS | The sample is analyzed by purge and trap gas chromatography with detection by mass spectrometry. The procedure involves purging the volatile compounds from the sample with an inert gas, and trapping them on a solid sorbent. When purging of the sample is complete, the trap is heated and the compounds are transferred to the gas chromatographic column. Analysis is then accomplished by separation of the components by gas chromatography with detection by mass spectrometry. | 0.005 | Y | Y | Y | 105 | 4/29/2005 | 3/22/2007 | 3/31/2007 1:01:06 AM | 1996-04-12 | 6/3/1998 6:51:36 PM | EMS |  | Concentration | mg/L | Y | (fl. conc.) |  | Dichloromethane (fl. conc.) | 75-09-2 | Volatile Organic Compounds (VOCs); | Volatile Organic Compounds (VOCs); EMS code: M041 | Chemical |  | LAB
M041 | ug/g | Dichloromethane | MS02 | Purge&Trap GC/MS (VOC-S) | Purge&Trap GC/MS (VOC-S) | 0.045 | Y | Y | Y | 16 | 1/14/2022 | 10/28/2022 5:04:00 PM | 1/11/2024 1:07:08 PM | 2022-07-27 |  | EMS |  | MassPerMass | mg/kg | Y | (mass conc.) |  | Dichloromethane (mass conc.) | 75-09-2 | Volatile Organic Compounds (VOCs); | Volatile Organic Compounds (VOCs); EMS code: M041 | Chemical |  | LAB
M041 | mg/L | Dichloromethane | MS01 | Purge&Trap GC/MS (VOC-W) | Purge&Trap GC/MS (VOC-W) |  | Y | Y | Y | 217 | 4/15/2002 | 6/4/2024 10:24:00 AM | 6/13/2024 5:01:27 PM | 2003-03-18 | 3/18/2003 2:49:40 PM | UMULLER | UMULLER | Concentration | mg/L | Y | (fl. conc.) |  | Dichloromethane (fl. conc.) | 75-09-2 | Volatile Organic Compounds (VOCs); | Volatile Organic Compounds (VOCs); EMS code: M041 | Chemical |  | LAB
M041 | mg/L | Dichloromethane | MS04 | Static Headspace GC/MS (Water) | (Water) A static headspace procedure is used as a means of driving the analytes out of the matrix into a headspace. Once equilibrium is achieved, a fixed volume of headspeace is taken for subsequent determination by Gas Chromatography - Mass Spectroscopy (GC/MS) using Selected Ion Monitoring. | 0.002 | Y | Y | Y | 68 | 8/13/2009 | 6/28/2024 12:29:00 AM | 7/23/2024 11:41:50 AM | 2009-12-18 |  | RDGIBSON |  | Concentration | mg/L | Y | (fl. conc.) |  | Dichloromethane (fl. conc.) | 75-09-2 | Volatile Organic Compounds (VOCs); | Volatile Organic Compounds (VOCs); EMS code: M041 | Chemical |  | LAB
M041 | mg/L | Dichloromethane | F007 | Static Headspace:HRGC/LRMS | Static headspace analysis is a method for the determination of volatile compounds in liquids and/or solids by analysing the vapour phase of a sample in a closed system; in this case a sealed vial. Static Headspace is sometimes referred to as Equilibrium Headspace because the vapour phase above the sample in a sealed vial is brought into thermodynamic equilibrium with the sample prior to analysis. Volatile organic compounds are thermally extracted from the sample matrix by Static Headspace sampling and introduced directly to a thick-film megabore GC column. Compounds eluting from the GC column are introduced into a Mass Spectrometer and target compounds are identified by a comparison of both retention times and mass spectra with those of reference standards stored in a computerized data system. Reference retention times and mass spectra for analytes are obtained by the measurement of calibration standards under precisely the same extraction conditions used for samples. The concentration of each identified component is measured by relating the response of the sample quantitation mass ions to the response of the quantitation mass ions of a series of external standards. A deuterated surrogate recovery standard is added to provide information on matrix effects. | 0.005 | Y | Y | Y | 380 | 9/28/1998 3:08:00 PM | 8/14/2024 | 8/16/2024 4:05:21 PM | 1996-09-16 | 6/3/1998 6:51:44 PM | MPAK |  | Concentration | mg/L | Y | (fl. conc.) |  | Dichloromethane (fl. conc.) | 75-09-2 | Volatile Organic Compounds (VOCs); | Volatile Organic Compounds (VOCs); EMS code: M041 | Chemical |  | LAB
M041 | ug/g | Dichloromethane | FS04 | Static Headspace:HRGC/LRMS | Static Headspace anaylsis is a method for the determination of volatile compounds in liquids and/or solids by analysing the vapour phase of a sample in a closed system; in this case a sealed vial. Static Headspace is sometimes referred to as Equilibrium Headspace because the vapour phase above the sample in a sealed vial is brought into thermodynamic equilibrium with the sample prior to analysis. Volatile organic compounds are thermally extracted from the sample matrix by Static Headspace sampling and introduced directly to a thick-film megabore GC column. Compounds eluting from the GC column are introduced into a Mass Spectrometer and target compounds are identified by a comparison of both retention times and mass spectra with those of reference standards stored in a computerized data system. Reference retention times and mass spectra for analytes are obtained by the measurement of calibration standards under precisely the same extraction conditions used for samples. The concentration of each identified component is measured by relating the response of the sample quantitation mass ions to the response of the quantitation mass ions of a series of external standards. A deuterated surrogate recovery standard is added to provide information on matrix effects. | 0.05 | Y | Y | Y | 44 | 6/30/1999 4:41:00 PM | 5/22/2015 | 7/3/2015 3:19:04 PM | 1996-09-16 | 6/3/1998 6:51:44 PM | MPAK |  | MassPerMass | mg/kg | Y | (mass conc.) |  | Dichloromethane (mass conc.) | 75-09-2 | Volatile Organic Compounds (VOCs); | Volatile Organic Compounds (VOCs); EMS code: M041 | Chemical |  | LAB
M042 | mg/L | Methylene Chloride | MS04 | Static Headspace GC/MS (Water) | (Water) A static headspace procedure is used as a means of driving the analytes out of the matrix into a headspace. Once equilibrium is achieved, a fixed volume of headspeace is taken for subsequent determination by Gas Chromatography - Mass Spectroscopy (GC/MS) using Selected Ion Monitoring. | 0.002 | Y | Y | Y | 36 |  |  | 7/23/2024 11:41:50 AM | 2010-01-13 | 1/13/2010 1:01:35 PM | RDGIBSON | RDGIBSON | Concentration | mg/L | Y | (fl. conc.) |  | Methylene Chloride (fl. conc.) | 75-09-2 | Volatile Organic Compounds (VOCs); | Volatile Organic Compounds (VOCs); EMS code: M042 | Chemical |  | LAB
M042 | mg/L | Methylene Chloride | MS01 | Purge&Trap GC/MS (VOC-W) | Purge&Trap GC/MS (VOC-W) | 0.3 | Y | Y | Y | 237 | 10/8/2004 | 5/27/2016 6:28:00 PM | 12/30/2023 10:26:22 AM | 1996-04-12 | 3/18/2003 2:49:40 PM | EMS | UMULLER | Concentration | mg/L | Y | (fl. conc.) |  | Methylene Chloride (fl. conc.) | 75-09-2 | Volatile Organic Compounds (VOCs); | Volatile Organic Compounds (VOCs); EMS code: M042 | Chemical |  | LAB
M042 | ug/g | Methylene Chloride | MS02 | Purge&Trap GC/MS (VOC-S) | Purge&Trap GC/MS (VOC-S) | 0.002 | Y | Y | Y | 11 | 11/19/2004 | 11/29/2004 1:55:00 PM | 1/6/2005 1:01:56 AM | 1996-04-12 | 6/3/1998 6:51:45 PM | EMS |  | MassPerMass | mg/kg | Y | (mass conc.) |  | Methylene Chloride (mass conc.) | 75-09-2 | Volatile Organic Compounds (VOCs); | Volatile Organic Compounds (VOCs); EMS code: M042 | Chemical |  | LAB
M050 | mg/L | Monochlorophenols | PP11 | Ac. anhydride ext:GC/MS | Ac. anhydride ext:GC/MS | 5e-05 | Y | Y | Y | 235 | 8/3/2005 8:02:00 AM | 8/3/2005 8:02:00 AM | 6/2/2009 12:01:07 PM | 1996-04-12 | 6/3/1998 6:51:36 PM | EMS |  | Concentration | mg/L | Y | (fl. conc.) |  | Monochlorophenols (fl. conc.) | 25167-80-0 | Organics; | Organics; EMS code: M050 | Chemical |  | LAB
M051 | mg/L | Monochloroguaicols | PP11 | Ac. anhydride ext:GC/MS | Ac. anhydride ext:GC/MS | 5e-05 | Y | Y | Y | 233 |  |  | 10/24/1996 5:30:20 AM | 1996-04-12 | 6/3/1998 6:51:36 PM | EMS |  | Concentration | mg/L | Y | (fl. conc.) |  | Monochloroguaicols (fl. conc.) |  | Organics; | Organics; EMS code: M051 | Chemical |  | LAB
M052 | mg/L | Monochlorocatechols | PP11 | Ac. anhydride ext:GC/MS | Ac. anhydride ext:GC/MS | 5e-05 | Y | Y | Y | 241 |  |  | 10/24/1996 5:30:20 AM | 1996-04-12 | 6/3/1998 6:51:36 PM | EMS |  | Concentration | mg/L | Y | (fl. conc.) |  | Monochlorocatechols (fl. conc.) | 4018-65-9, 2138-22-9 | Organics; | Organics; EMS code: M052 | Chemical |  | LAB
M053 | mg/L | Monochlorovanillins | PP11 | Ac. anhydride ext:GC/MS | Ac. anhydride ext:GC/MS | 5e-05 | Y | Y | Y | 235 |  |  | 10/24/1996 5:30:20 AM | 1996-04-12 | 6/3/1998 6:51:36 PM | EMS |  | Concentration | mg/L | Y | (fl. conc.) |  | Monochlorovanillins (fl. conc.) | 19463-48-0, 18268-76-3 | Organics; | Organics; EMS code: M053 | Chemical |  | LAB
M100 | mg/L | Metazachlor | ZN01 | DCM+Ac.Extr:GC/MS | DCM+Ac.Extr:GC/MS | 0.5 | Y | Y | Y | 1 | 7/29/2005 7:10:00 AM | 7/29/2005 7:10:00 AM | 6/2/2009 12:01:07 PM | 2009-02-20 |  | JCHUDAK |  | Concentration | mg/L | Y | (fl. conc.) |  | Metazachlor (fl. conc.) | 67129-08-2 | Pesticides; | Pesticides; EMS code: M100 | Chemical |  | LAB
M101 | mg/L | Methoprene | ZN01 | DCM+Ac.Extr:GC/MS | DCM+Ac.Extr:GC/MS | 0.5 | Y | Y | Y | 1 | 7/29/2005 7:10:00 AM | 7/29/2005 7:10:00 AM | 6/2/2009 12:01:08 PM | 2009-02-20 |  | JCHUDAK |  | Concentration | mg/L | Y | (fl. conc.) |  | Methoprene (fl. conc.) | 40596-69-8 | Pesticides; | Pesticides; EMS code: M101 | Chemical |  | LAB
M200 | % | 200 mesh (<.075 mm) | FSDS | Dry Sieve | Dry Sieve | 0.1 | Y | Y | Y | 7 | 9/18/2022 | 9/18/2022 | 3/9/2023 4:55:51 PM | 2023-02-28 |  | EMS |  | DimensionlessRatio | % | N | (ratio) |  | 200 mesh (<.075 mm) (ratio) |  | Sediment Size; | Sediment Size; EMS code: M200 | Physical |  | LAB
M201 | % | 200 mesh (>.075 mm) | FSDS | Dry Sieve | Dry Sieve | 0.1 | Y | Y | Y | 7 | 9/18/2022 | 9/18/2022 | 3/9/2023 4:55:51 PM | 2023-02-28 |  | EMS |  | DimensionlessRatio | % | N | (ratio) |  | 200 mesh (>.075 mm) (ratio) |  | Sediment Size; | Sediment Size; EMS code: M201 | Physical |  | LAB
M201 | % | 200 mesh (>.075 mm) | GS01 | Gravimetric, balance, Separation by sieve >.075 mm | Gravimetric, balance; Separation by sieve >.075 mm. These are particle sizing determinations performed on soil. These two parameters reported will need the > and < associated with them because the way they have been reported is not based on a number of distinct and specific sieve sizes but instead more general size of everything that is either larger or greater than a size of 0.075mm. | 0.1 | Y | Y | Y | 7 | 10/29/2022 | 10/29/2022 | 1/7/2023 3:12:11 PM | 2010-11-16 |  | RDGIBSON |  | DimensionlessRatio | % | N | (ratio) |  | 200 mesh (>.075 mm) (ratio) |  | Sediment Size; | Sediment Size; EMS code: M201 | Physical |  | LAB
MBAA | mg/L | Monobromoacetic Acid | HAAW | Haloacetic Acids,MTBE ex | Haloacetic Acids;MTBE ex | 0.8 | Y | Y | Y | 2 | 3/27/2023 9:58:00 AM | 5/31/2023 12:14:00 PM | 6/8/2023 12:14:57 PM | 1996-10-13 | 6/3/1998 6:51:46 PM | EMS |  | Concentration | mg/L | Y | (fl. conc.) |  | Monobromoacetic Acid (fl. conc.) | 79-08-3 | Organics; | Organics; EMS code: MBAA | Chemical |  | LAB
MCAA | mg/L | Monochloracetic Acid | HAAW | Haloacetic Acids,MTBE ex | Haloacetic Acids;MTBE ex | 1 | Y | Y | Y | 2 | 3/27/2023 9:58:00 AM | 5/31/2023 12:14:00 PM | 6/8/2023 12:14:57 PM | 1996-10-13 | 6/3/1998 6:51:46 PM | EMS |  | Concentration | mg/L | Y | (fl. conc.) |  | Monochloracetic Acid (fl. conc.) | 79-11-8 | Organics; | Organics; EMS code: MCAA | Chemical |  | LAB
MCPB | ug/g | butanoic acid; 4-(4-chloro-2-methylphenoxy)- [MCPB] | ZN03 | DCM+Ac.Extr:GC/MS | DCM+Ac.Extr:GC/MS EPA 8270E (mod) | 0.0005 | Y | Y | Y | 15 | 1/14/2022 | 2/2/2022 | 1/11/2024 1:07:08 PM | 2022-07-27 |  | EMS |  | MassPerMass | mg/kg | Y | (mass conc.) |  | butanoic acid; 4-(4-chloro-2-methylphenoxy)- [MCPB] (mass conc.) | 94-81-5 | Pesticides; | Pesticides; EMS code: MCPB | Chemical |  | LAB
MCPB | mg/L | butanoic acid; 4-(4-chloro-2-methylphenoxy)- [MCPB] | P015 | Phenoxyacid Herbicides by LC-MS/MS | Phenoxyacid Herbicides by LC-MS/MS | 5e-05 | Y | Y | Y | 68 | 12/6/2021 | 9/9/2022 | 12/8/2022 12:57:05 PM | 2021-12-09 | 12/9/2021 3:22:12 PM | EMS | EMS | Concentration | mg/L | Y | (fl. conc.) |  | butanoic acid; 4-(4-chloro-2-methylphenoxy)- [MCPB] (fl. conc.) | 94-81-5 | Pesticides; | Pesticides; EMS code: MCPB | Chemical |  | LAB
MCYS | mg/L | Microcystin | EP22 | Total Microcystins Enzyme-Linked ImmunoSorbent Assay ELISA | ALS Method: ENVIROLOGIX QUANTIPLATE KIT CAT. EP022 Methodology: Total Microcystins (intraceller and extraceller) in aqueous matrices is determined by the Enzyme-Linked ImmunoSorbent Assay (ELISA) method. | 0.0002 | Y | Y | Y | 21 | 9/12/2018 | 9/20/2021 | 10/14/2021 1:14:03 PM | 2018-09-24 |  | RDGIBSON |  | Concentration | mg/L | Y | (fl. conc.) |  | Microcystin (fl. conc.) | 77238-39-2 | Organics; | Organics; EMS code: MCYS | Chemical |  | LAB
MESE | mg/L | methylseleninic acid CH3SeO2H | E540 | Miekeley Spect. Acta B 60 (2005) 633-641 | E540 and E540B: Instrumental analysis is by Anion Exchange HPLC-ICPMS. Reported Total Unknown Se Species is semi-quantitative in nature and consists of a sum of all the unknown peaks observed in the chromatogram. Field filtration is recommended. | 0.0001 | Y | Y | Y | 176 | 12/13/2021 | 8/11/2024 | 8/19/2024 12:19:03 PM | 2020-04-21 |  | RDGIBSON |  | Concentration | mg/L | Y | (fl. conc.) |  | methylseleninic acid CH3SeO2H (fl. conc.) | 28274-57-9 | Organometallics; | Organometallics; EMS code: MESE | Chemical |  | LAB
MESE | mg/L | methylseleninic acid CH3SeO2H | X564 | IC-ICP-CRC-MS | Ion chromatography inductively coupled plasma collision reaction cell mass spectrometry | 5e-06 | Y | Y | Y | 6330 |  |  | 7/30/2024 12:34:59 PM | 2016-11-16 | 3/23/2017 9:37:21 AM | RRENTON | RRENTON | Concentration | mg/L | Y | (fl. conc.) |  | methylseleninic acid CH3SeO2H (fl. conc.) | 28274-57-9 | Organometallics; | Organometallics; EMS code: MESE | Chemical |  | LAB
MESE2 | mg/L | Methaneselenonic Acid CH4O3Se | X564 | IC-ICP-CRC-MS | Ion chromatography inductively coupled plasma collision reaction cell mass spectrometry | 1e-05 | Y | Y | Y | 5540 |  |  | 7/30/2024 12:34:59 PM | 2022-10-25 |  | EMS |  | Concentration | mg/L | Y | (fl. conc.) |  | Methaneselenonic Acid CH4O3Se (fl. conc.) |  |  | EMS code: MESE2 | Chemical |  | LAB
METH | mg/L | Methane; Dissolved | MEED | Methane, Ethane and Ethene, Dissolved via GC-FID | Water samples are collected in 44 mL P&T vials without headspace and preserved with Sodium Bisulfate. Headspace is then generated using Nitrogen. The sample is then shaken to allow gas equilibration. A portion of the headspace vapour is withdrawn and injected into a gas chromatograph equipped with capillary columns and a flame ionization detector (FID). The concentration of the gas in the liquid is proportional to the partial pressure of the gas above the liquid, and is calculated using Henrys Law. Ref: EPA REGION 1, NATATTEN.WPD, REV. 1 | 0.005 | Y | Y | Y | 33 | 2/20/2014 | 8/11/2016 | 9/21/2016 4:07:57 PM | 2014-05-30 |  | RDGIBSON |  | Concentration | mg/L | Y | (fl. conc.) |  | Methane; Dissolved (fl. conc.) | 74-82-8 | Volatile Organic Compounds (VOCs); | Volatile Organic Compounds (VOCs); EMS code: METH | Chemical | Dissolved | LAB
METP | mg/L | Metoprolol CAS: 51384-51-1 | XMSP | HPLC and Tri. Quadpole MS/MS | EPA 1694 An aliquot of filtered sample containing 5% organic is injected directly. Each compound is separated by reversed phase HPLC. Analyzed by Electro Spray Ionization MS/MS detection using Triple Quadrupole MS/MS detector. Quantiitation with internal or external standard. |  | Y | Y | Y | 12 | 5/8/2020 | 6/4/2024 | 8/21/2024 3:32:03 PM | 2022-05-04 |  | EMS |  | Concentration | mg/L | Y | (fl. conc.) |  | Metoprolol (fl. conc.) | 51384-51-1 | Pharmaceuticals and Personal Care Products (PPCP); | Pharmaceuticals and Personal Care Products (PPCP); EMS code: METP | Chemical |  | LAB
MFOS1 | mg/L | Methyl perfluorooctanesulfonamide,N- (NMeFOSA)CAS:31506-32-8 | E3533 | MECP E3533 | PFAs in water by direct injection LC/MS/MS | 2e-05 | Y | Y | Y | 2 | 6/3/2024 | 6/3/2024 | 8/21/2024 3:32:03 PM | 2023-11-15 | 6/27/2024 5:06:31 PM | EMS | EMS | Concentration | mg/L | Y | (fl. conc.) |  | Methyl perfluorooctanesulfonamide,N- (NMeFOSA) (fl. conc.) | 31506-32-8 | Per- and polyfluoroalkyl substances (PFAS); | Per- and polyfluoroalkyl substances (PFAS); EMS code: MFOS1 | Chemical |  | LAB
MFOS2 | mg/L | Methyl perfluorooctanesulfonamidoacetic acid, n- (NMeFOSAA) | E3533 | MECP E3533 | PFAs in water by direct injection LC/MS/MS | 2e-05 | Y | Y | Y | 2 | 6/3/2024 | 6/3/2024 | 8/21/2024 3:32:03 PM | 2023-11-15 | 6/27/2024 5:06:31 PM | EMS | EMS | Concentration | mg/L | Y | (fl. conc.) |  | Methyl perfluorooctanesulfonamidoacetic acid, n- (NMeFOSAA) (fl. conc.) |  | Per- and polyfluoroalkyl substances (PFAS); | Per- and polyfluoroalkyl substances (PFAS); EMS code: MFOS2 | Chemical |  | LAB
MFOS3 | mg/L | Methyl perfluorooctanesulfonamidoethanol, n- (NMeFOSE) | E3533 | MECP E3533 | PFAs in water by direct injection LC/MS/MS | 2e-05 | Y | Y | Y | 2 | 6/3/2024 | 6/3/2024 | 8/21/2024 3:32:03 PM | 2023-11-15 | 6/27/2024 5:06:31 PM | EMS | EMS | Concentration | mg/L | Y | (fl. conc.) |  | Methyl perfluorooctanesulfonamidoethanol, n- (NMeFOSE) (fl. conc.) |  | Per- and polyfluoroalkyl substances (PFAS); | Per- and polyfluoroalkyl substances (PFAS); EMS code: MFOS3 | Chemical |  | LAB
MG-D | mg/L | Magnesium Dissolved | X562 | Dissolved metals: FF, FA, HNO3, HR-ICPMS | Ultra trace metals in water are analyzed by high resolution inductively coupled plasma mass spectrometry (HR-ICPMS), based on US EPA Method 200.8, (Oct. 1999). This procedure is intended for pristine field-filtered acid-preserved water samples. The detection limits (LOR) for this test are based on lab instrumental analysis only, not including filtration. ALS-supplied field filtration equipment does not support these LOR. Therefore, because of the high probability of false positives due to filtration, it is strongly recommended that a filtration blank be analyzed to aid in data interpretation. | 0.01 | Y | Y | Y | 1258 | 6/8/2015 | 10/19/2023 | 11/9/2023 7:41:27 AM | 2015-10-22 |  | RDGIBSON |  | Concentration | mg/L | Y | (fl. conc.) |  | Magnesium Dissolved (fl. conc.) | 7439-95-4 | Metals - Dissolved; | Metals - Dissolved; EMS code: MG-D | Chemical | Dissolved | LAB
MG-D | mg/L | Magnesium Dissolved | F039 | Field filt + acid, ICP-AES | Dissolved and extractable water samples are ready for analysis once collected. Total samples are autoclaved with HNO3 and hydrochloric acid (HCI) in capped Teflon vials. Analysis is performed via Inductively Coupled Argon Plasma Atomic Emission Spectrometry (ICP-AES)(simultaneous multielement analysis), Graphite Furnace Atomic Absorption Spectrometry (GF-AAS)(low level analysis) and/or Flame Atomic Absorption Spectrometry (F-AAS). | 0.1 | Y | Y | Y | 14303 | 7/29/1996 6:03:00 PM | 11/22/2022 12:29:00 PM | 7/22/2024 3:17:26 PM | 1997-05-08 | 6/3/1998 6:51:46 PM | MPAK |  | Concentration | mg/L | Y | (fl. conc.) |  | Magnesium Dissolved (fl. conc.) | 7439-95-4 | Metals - Dissolved; | Metals - Dissolved; EMS code: MG-D | Chemical | Dissolved | LAB
MG-D | mg/L | Magnesium Dissolved | XCAL | Calculation (mg/L) | Calculation (mg/L) | 0.002 | Y | Y | Y | 2604 |  |  | 8/23/1999 4:10:47 PM | 1996-04-12 | 6/3/1998 6:51:42 PM | EMS |  | Concentration | mg/L | Y | (fl. conc.) |  | Magnesium Dissolved (fl. conc.) | 7439-95-4 | Metals - Dissolved; | Metals - Dissolved; EMS code: MG-D | Chemical | Dissolved | LAB
MG-D | mg/L | Magnesium Dissolved | X170 | ICP | This code is a catch-all for a large variety of ICP methods including water and also solids digested into solutions as an intermediate result. Different procedures will have greatly different detection limits. (a) For air particulates, preliminary results are reported in mg/L and loading results as mg/dm2/d. (b) For Hi Vol samples preliminiary results are same as reported in mg/L and loading results as ug/m3. (c) Note that for Rain samples there is no filtration nor digestion steps. | 0.1 | Y | Y | Y | 19481 | 4/11/1996 6:02:00 PM | 11/6/2018 | 6/13/2024 3:58:21 PM | 1996-04-12 | 8/15/2019 1:21:24 PM | EMS | KATFRASE | Concentration | mg/L | Y | (fl. conc.) |  | Magnesium Dissolved (fl. conc.) | 7439-95-4 | Metals - Dissolved; | Metals - Dissolved; EMS code: MG-D | Chemical | Dissolved | LAB
MG-D | mg/L | Magnesium Dissolved | CS01 | System Calculation | System Calculation (mg/L) |  | N | Y | N | 127 |  |  | 10/12/2023 9:37:02 AM | 1996-04-12 | 6/3/1998 6:51:43 PM | EMS |  | Concentration | mg/L | Y | (fl. conc.) |  | Magnesium Dissolved (fl. conc.) | 7439-95-4 | Metals - Dissolved; | Metals - Dissolved; EMS code: MG-D | Chemical | Dissolved | LAB
MG-D | mg/L | Magnesium Dissolved | METC | MET-D-NP-CCMS-VA | Ultra trace metals in water are analyzed by CRC ICPMS, based on US EPA Method 6020B (July 2014). This procedure is intended for pristine field-filtered acid-preserved water samples. The detection limits (LOR) for this test are based on lab instrumental analysis only, not including filtration. ALS-supplied field filtration equipment does not support these LOR. Therefore, because of the high probability of false positives due to filtration, it is strongly recommended that a filtration blank be analysed to aid in data interpretation. | 0.01 | Y | Y | Y | 1156 | 11/13/2017 | 8/15/2024 | 8/23/2024 9:50:35 AM | 2019-08-20 |  | KATFRASE |  | Concentration | mg/L | Y | (fl. conc.) |  | Magnesium Dissolved (fl. conc.) | 7439-95-4 | Metals - Dissolved; | Metals - Dissolved; EMS code: MG-D | Chemical | Dissolved | LAB
MG-D | mg/L | Magnesium Dissolved | X053 | AA,Dir Asp. | AA;Dir Asp. | 0.02 | Y | Y | Y | 2693 |  |  | 3/29/2001 2:38:48 PM | 1996-04-12 | 6/3/1998 6:51:36 PM | EMS |  | Concentration | mg/L | Y | (fl. conc.) |  | Magnesium Dissolved (fl. conc.) | 7439-95-4 | Metals - Dissolved; | Metals - Dissolved; EMS code: MG-D | Chemical | Dissolved | LAB
MG-D | mg/L | Magnesium Dissolved | X516 | LaCl3add,Dir. Asp.,AA | LaCl3 added;Direct Aspiration;AA | 0.01 | Y | Y | Y | 1456 |  |  | 8/23/1999 4:52:30 PM | 1999-04-22 |  | NPEPPIN |  | Concentration | mg/L | Y | (fl. conc.) |  | Magnesium Dissolved (fl. conc.) | 7439-95-4 | Metals - Dissolved; | Metals - Dissolved; EMS code: MG-D | Chemical | Dissolved | LAB
MG-D | mg/L | Magnesium Dissolved | F032 | Filt.,Acid,ICPMS | There is no acid digestion required for this analysis. A water sample is collected and preserved in 0.2 percent nitric acid. To analyse the extractable metals the sample is shaken, allowed to sit for at least twenty four hours, and a decanted portion is introduced to the spectrometer. When the dissolved metals are required, the samples must first be filtered through a 0.45 micron cellulose acetate filter paper and then followed by preservation in 0.2 percent nitric acid. It is important to note that different laboratories may have very different detection limits for some variables analyzed by ICPMS. | 0.0005 | Y | Y | Y | 85373 | 4/23/1981 4:00:00 PM | 8/26/2024 | 8/26/2024 4:53:05 PM | 1997-01-16 | 6/3/1998 6:51:46 PM | MPAK |  | Concentration | mg/L | Y | (fl. conc.) |  | Magnesium Dissolved (fl. conc.) | 7439-95-4 | Metals - Dissolved; | Metals - Dissolved; EMS code: MG-D | Chemical | Dissolved | LAB
MG-D | mg/L | Magnesium Dissolved | F068 | HNO3dig,TeflonTube/Autoclave,ICP-MS | This method is applicable to determiniation of dissolved, extractable, and total recoverable metals in ground waters, surface waters, drinking waters, and wastewaters. Dissolved, extractable, and total recoverable sample fractions are acidified to pH <1.5 with analytical grade nitric acid (HNO3) to prevent the metals from adsorbing to the container material or precipitating. Extractable and total recoverable water samples are collected in 250 ml acid washed polyethylene bottles and are preserved with Nitric acid (0.5 ml acid for 100 mLs of sample). All preserved samples are stored at room temperature and can be held for up to 6 months. Unfiltered, un-preserved samples are stored at 4C for no more than 30 days. For dissolved and extractable water fractions, no sample preparation is necessary. Samples are simply poured in autosampler tubes (i.e. 15 ml centrifuge tubes), and placed on the autosampler and analyzed. Samples for total recoverable metals are digested with Nitric acid in special Teflon tubes in an autoclave prior to analyses. | 5e-05 | Y | Y | Y | 772 | 6/1/1999 1:21:00 PM | 2/12/2002 2:08:00 PM | 5/28/2003 11:57:22 AM | 1999-06-28 |  | NPEPPIN |  | Concentration | mg/L | Y | (fl. conc.) |  | Magnesium Dissolved (fl. conc.) | 7439-95-4 | Metals - Dissolved; | Metals - Dissolved; EMS code: MG-D | Chemical | Dissolved | LAB
MG-D | mg/L | Magnesium Dissolved | EQ01 | EQUIS (mg/L) | EQUIS (mg/L) | 0.02 | N | Y | Y | 9861 |  |  | 10/24/1996 3:16:55 PM | 1996-04-12 | 6/3/1998 6:51:36 PM | EMS |  | Concentration | mg/L | Y | (fl. conc.) |  | Magnesium Dissolved (fl. conc.) | 7439-95-4 | Metals - Dissolved; | Metals - Dissolved; EMS code: MG-D | Chemical | Dissolved | LAB
MG-D | mg/L | Magnesium Dissolved | X203 | FF:FA,HNO3,Dir. Flame,AA | Samples are field-filtered (0.45 um), and field-acidified (HNO3); analysis by AA. | 0.001 | Y | Y | Y | 6831 |  |  | 8/26/2024 2:41:32 PM | 1996-04-12 | 6/3/1998 6:51:36 PM | EMS |  | Concentration | mg/L | Y | (fl. conc.) |  | Magnesium Dissolved (fl. conc.) | 7439-95-4 | Metals - Dissolved; | Metals - Dissolved; EMS code: MG-D | Chemical | Dissolved | LAB
MG-D | mg/L | Magnesium Dissolved | 0040 | HNO3 Dig: ICP Analysis | HNO3 Dig: ICP Analysis | 0.1 | Y | Y | Y | 78 | 5/9/2003 9:40:00 AM | 9/16/2003 3:18:00 PM | 11/28/2003 3:22:12 PM | 1996-04-12 | 8/15/2019 1:20:40 PM | EMS | KATFRASE | Concentration | mg/L | Y | (fl. conc.) |  | Magnesium Dissolved (fl. conc.) | 7439-95-4 | Metals - Dissolved; | Metals - Dissolved; EMS code: MG-D | Chemical | Dissolved | LAB
MG-D | mg/L | Magnesium Dissolved | X350 | FF,FA:HNO3:ICP | Samples are field-filtered (0.45 um); field-acidified with HNO3; analysis by ICP/AES. | 0.1 | Y | Y | Y | 168 | 9/29/2000 | 12/1/2010 | 12/9/2010 3:55:27 PM | 1996-04-12 | 6/3/1998 6:51:41 PM | EMS |  | Concentration | mg/L | Y | (fl. conc.) |  | Magnesium Dissolved (fl. conc.) | 7439-95-4 | Metals - Dissolved; | Metals - Dissolved; EMS code: MG-D | Chemical | Dissolved | LAB
MG-D | mg/L | Magnesium Dissolved | MSEA | Metals in Seawater (undigested) by (CRC) ICP-MS | This method describes the multi-elemental determination of trace elements by Collision/Reaction Cell (CRC) ICP-MS. The method measures ions produced by a radio-frequency inductively coupled plasma. Analyte species originating in a liquid are pneumatically nebulized and the resulting aerosol transported by argon gas into the plasma torch. The ions produced are entrained in the plasma gas and extracted, by means of a differentially pumped vacuum interface, into a mass spectrometer. The ions flow through a Collision/Reaction Cell (CRC) where depending on the mode of analysis either the ions passes through the cell unhindered (no gas - classical ICP-MS mode), or collisions/reactions chemistry takes place between the ions and a neutral gas such as hydrogen or helium. Any polyatomic ions which give interference on the analytes are removed. The ions are then sorted according to their mass-to-charge ratios by a quadrupole mass spectrometer. The ions are then sorted according to their mass-to-charge ratios by a quadrupole mass spectrometer. The ions transmitted through the quadrupole are quantified by a channel electron multiplier and the ion information processed by a data handling system. Although most polyatomic interference is removed, any additional interference must be assessed and valid corrections applied or the data flagged to indicate problems. Interference correction must include compensation for background ions contributed by the plasma gas, reagents, and constituents of the sample matrix. Instrument drift as well as suppressions or enhancements of instrument response caused by the sample matrix must be corrected for by the use of internal standards. | 5 | Y | Y | Y | 4 | 4/12/2023 | 4/29/2024 | 5/16/2024 2:00:14 PM | 2009-09-08 |  | RDGIBSON |  | Concentration | mg/L | Y | (fl. conc.) |  | Magnesium Dissolved (fl. conc.) | 7439-95-4 | Metals - Dissolved; | Metals - Dissolved; EMS code: MG-D | Chemical | Dissolved | LAB
MG-D | mg/L | Magnesium Dissolved | MWID | Metals in Water by ICPMS (digested) | Metals in Water by ICPMS (digested) Based on EPA 200.8 | 1 | Y | Y | Y | 2238 |  |  | 7/31/2024 1:46:41 PM | 2009-12-24 |  | RDGIBSON |  | Concentration | mg/L | Y | (fl. conc.) |  | Magnesium Dissolved (fl. conc.) | 7439-95-4 | Metals - Dissolved; | Metals - Dissolved; EMS code: MG-D | Chemical | Dissolved | LAB
MG-D | mg/L | Magnesium Dissolved | X285 | HNO3 Dig:ICP Analysis | HNO3 Dig:ICP Analysis | 0.1 | Y | Y | Y | 745 | 3/1/2000 | 11/13/2002 | 2/23/2004 2:18:50 PM | 1998-05-22 | 6/3/1998 6:51:49 PM | NPEPPIN |  | Concentration | mg/L | Y | (fl. conc.) |  | Magnesium Dissolved (fl. conc.) | 7439-95-4 | Metals - Dissolved; | Metals - Dissolved; EMS code: MG-D | Chemical | Dissolved | LAB
MG-D | mg/L | Magnesium Dissolved | 5138 | AA Analysis | AA Analysis | 0.02 | Y | Y | Y | 29 |  |  | 10/24/1996 11:35:00 AM | 1996-04-12 | 6/3/1998 6:51:36 PM | EMS |  | Concentration | mg/L | Y | (fl. conc.) |  | Magnesium Dissolved (fl. conc.) | 7439-95-4 | Metals - Dissolved; | Metals - Dissolved; EMS code: MG-D | Chemical | Dissolved | LAB
MG-D | mg/L | Magnesium Dissolved | X085 | LF,LA,HNO3,AA | Samples are lab-filtered (0.45 um), and lab-acidified (HNO3); analysis by AA. | 0.02 | Y | Y | Y | 97 | 9/12/1998 | 9/18/2015 | 10/28/2015 9:54:16 AM | 1996-04-12 | 6/3/1998 6:51:36 PM | EMS |  | Concentration | mg/L | Y | (fl. conc.) |  | Magnesium Dissolved (fl. conc.) | 7439-95-4 | Metals - Dissolved; | Metals - Dissolved; EMS code: MG-D | Chemical | Dissolved | LAB
MG-E | mg/L | Magnesium Extract. | X553 | Metals in water, extractable, ICP-SFMS | The whole water sample is collected in a HDPE/Teflon bottle. In a Class 100 workstation it is preserved with concentrated Nitric Acid at concentration of 0.2 percent (v/v). The acid preservative extracts the metals from material in the sample and a decanted portion is introduced directly to an inductively coupled argon plasma-sector field mass spectrometer (ICP-SFMS) by pneumatic nebulization. Each respective element is measured at a specific mass to charge ratio, m/z value expressed in amu (atomic mass units), and at a specified instrument resolution. LR: Low Resolution ~300, MR: Mid Resolution ~4000 and HR: High Resolution ~10000. Ref: Method number 02-2061, National Laboratory for Environmental Testing, CCIW, Environment Canada, Burlington. Requested April 2003. | 2.5 | Y | Y | Y | 1 |  |  | 11/28/2023 2:11:53 PM | 2023-11-28 |  | EMS |  | Concentration | mg/L | Y | (fl. conc.) |  | Magnesium Extract. (fl. conc.) | 7439-95-4 | Metals - Extractable; | Metals - Extractable; EMS code: MG-E | Chemical |  | LAB
MG-E | mg/L | Magnesium Extract. | E6020B | EPA 6020B | See EPA 6020B and CSSS Ch. 15 | 1 | Y | Y | Y | 15 | 1/11/2022 | 2/4/2022 | 1/11/2024 1:07:08 PM | 2022-06-29 |  | EMS |  | Concentration | mg/L | Y | (fl. conc.) |  | Magnesium Extract. (fl. conc.) | 7439-95-4 | Metals - Extractable; | Metals - Extractable; EMS code: MG-E | Chemical |  | LAB
MG-E | ug/g | Magnesium Extract. | CAL2 | Calculation (ug/g) | Calculation (ug/g) | 0.5 | Y | Y | Y | 15 | 1/11/2022 | 2/7/2022 | 1/11/2024 1:07:08 PM | 2022-07-06 |  | EMS |  | MassPerMass | mg/kg | Y | (mass conc.) |  | Magnesium Extract. (mass conc.) | 7439-95-4 | Metals - Extractable; | Metals - Extractable; EMS code: MG-E | Chemical |  | LAB
MG-E | ug/g | Magnesium Extract. | ASCE | Ammonium Acetate soil colloid exchange | Available metals - Ammonium Acetate soil colloid exchange sites displacement |  | Y | Y | Y | 7 | 10/31/2022 | 10/31/2022 | 11/22/2022 2:40:39 PM | 2007-10-04 |  | RDGIBSON |  | MassPerMass | mg/kg | Y | (mass conc.) |  | Magnesium Extract. (mass conc.) | 7439-95-4 | Metals - Extractable; | Metals - Extractable; EMS code: MG-E | Chemical |  | LAB
MG-E | ug/g | Magnesium Extract. | X230 | Hom,HNO3/HCl,ICP | The sample is dried, ground to pass 100 mesh, and digested with a nitric-hydrochloric acid mixture to solubilize the solid matter and to remove the organic material by oxidation and volatilization. The completed digests are submitted to metals instrumentation personnel for metals analysis by Inductively Coupled Atomic Emission Spectrophotometry (ICP-AES), or hydride analysis by atomic absorption. | 0.5 | Y | Y | Y | 70 |  |  | 10/24/1996 2:25:27 PM | 1996-04-12 | 6/3/1998 6:51:36 PM | EMS |  | MassPerMass | mg/kg | Y | (mass conc.) |  | Magnesium Extract. (mass conc.) | 7439-95-4 | Metals - Extractable; | Metals - Extractable; EMS code: MG-E | Chemical |  | LAB
MG-E | mg/L | Magnesium Extract. | F069 | SEM extr,6N HCl,ICP/OES | 6.0 N Hydrochloric acid (HCI) is added to an aliquot of the sediment sample. The extract produced from the addition of the acid is then analyzed for simultaneously extracted metals (SEM) using inductively coupled plasma - optical emission spectrophotometry. The detection limits depend on the metal parameters in question and the instrumentation used. | 0.05 | Y | Y | Y | 12 |  |  | 2/28/2022 4:37:31 PM | 1999-07-09 |  | NPEPPIN |  | Concentration | mg/L | Y | (fl. conc.) |  | Magnesium Extract. (fl. conc.) | 7439-95-4 | Metals - Extractable; | Metals - Extractable; EMS code: MG-E | Chemical |  | LAB
MG-E | mg/L | Magnesium Extract. | F040 | Field Acid, lab filt, ICP-AES | Dissolved and extractable water samples are ready for analysis once collected. Total samples are autoclaved with HNO3 and hydrochloric acid (HCI) in capped Teflon vials. Analysis is performed via Inductively Coupled Argon Plasma Atomic Emission Spectrometry (ICP-AES)(simultaneous multielement analysis), Graphite Furnace Atomic Absorption Spectrometry (GF-AAS)(low level analysis) and/or Flame Atomic Absorption Spectrometry (F-AAS). | 0.1 | Y | Y | Y | 9027 | 8/5/1997 6:06:00 PM | 3/23/2012 3:11:00 AM | 9/11/2019 8:47:41 AM | 1997-05-08 | 6/3/1998 6:51:46 PM | MPAK |  | Concentration | mg/L | Y | (fl. conc.) |  | Magnesium Extract. (fl. conc.) | 7439-95-4 | Metals - Extractable; | Metals - Extractable; EMS code: MG-E | Chemical |  | LAB
MG-E | mg/L | Magnesium Extract. | F055 | No Pres/No Filtr/No Dig,ICP | Shortcut method which excludes preservation, filtration and digestion steps. Results are used for estimation of water hardness. | 0.1 | Y | Y | Y | 2359 | 8/23/1996 6:04:00 PM | 6/1/2000 11:27:00 AM | 9/11/2019 8:46:38 AM | 1998-04-21 | 6/3/1998 6:51:48 PM | NPEPPIN |  | Concentration | mg/L | Y | (fl. conc.) |  | Magnesium Extract. (fl. conc.) | 7439-95-4 | Metals - Extractable; | Metals - Extractable; EMS code: MG-E | Chemical |  | LAB
MG-O | % | Magnesium (as MgO) | ICP06 | ALS Minerals ME-ICP06 | ALS Minerals ME-ICP06. A prepared sample is fused with a lithium tetraborate/lithium metaborate flux and poured into a graphite crucible and fused in a furnace at 1000°C. The resulting melt is then cooled and dissolved in 100 mL of 4% HNO3/2% HCl followed by ICP-AES. | 0.01 | Y | Y | Y | 0 |  |  |  | 2024-01-29 |  | EMS |  | DimensionlessRatio | % | N | (ratio) |  | Magnesium (as MgO) (ratio) |  | Metals; | Metals; EMS code: MG-O | Chemical |  | LAB
MG-S | mg/L | Magnesium-Soluble ion content | F032 | Filt.,Acid,ICPMS | There is no acid digestion required for this analysis. A water sample is collected and preserved in 0.2 percent nitric acid. To analyse the extractable metals the sample is shaken, allowed to sit for at least twenty four hours, and a decanted portion is introduced to the spectrometer. When the dissolved metals are required, the samples must first be filtered through a 0.45 micron cellulose acetate filter paper and then followed by preservation in 0.2 percent nitric acid. It is important to note that different laboratories may have very different detection limits for some variables analyzed by ICPMS. | 0.01 | Y | Y | Y | 337 | 3/1/2002 1:15:00 PM | 2/17/2005 | 3/17/2005 1:07:53 AM | 2004-02-10 |  | UMULLER |  | Concentration | mg/L | Y | (fl. conc.) |  | Magnesium-Soluble ion content (fl. conc.) | 7439-95-4 | Metals - Soluble; | Metals - Soluble; EMS code: MG-S | Chemical | Soluble | LAB
MG-S | mg/L | Magnesium-Soluble ion content | X170 | ICP | This code is a catch-all for a large variety of ICP methods including water and also solids digested into solutions as an intermediate result. Different procedures will have greatly different detection limits. (a) For air particulates, preliminary results are reported in mg/L and loading results as mg/dm2/d. (b) For Hi Vol samples preliminiary results are same as reported in mg/L and loading results as ug/m3. (c) Note that for Rain samples there is no filtration nor digestion steps. | 0.02 | Y | Y | Y | 3144 | 4/6/2005 | 4/11/2008 | 5/28/2008 2:08:44 AM | 1996-04-12 | 6/3/1998 6:51:36 PM | EMS |  | Concentration | mg/L | Y | (fl. conc.) |  | Magnesium-Soluble ion content (fl. conc.) | 7439-95-4 | Metals - Soluble; | Metals - Soluble; EMS code: MG-S | Chemical | Soluble | LAB
MG-S | mg/L | Magnesium-Soluble ion content | E6010B | EPA 6010B/CSSS 15 | See CSSS CH15/EPA 6010B/AER D50. INDUCTIVELY COUPLED PLASMA-ATOMIC EMISSION SPECTROMETRY See EPA 6010B and CSSS Ch. 15. https://www.epa.gov/sites/default/files/documents/6010b.pdf | 5 | Y | Y | Y | 0 |  |  |  | 2024-01-30 |  | EMS |  | Concentration | mg/L | Y | (fl. conc.) |  | Magnesium-Soluble ion content (fl. conc.) | 7439-95-4 | Metals - Soluble; | Metals - Soluble; EMS code: MG-S | Chemical | Soluble | LAB
MG-S | mg/L | Magnesium-Soluble ion content | EQ01 | EQUIS (mg/L) | EQUIS (mg/L) | 0.02 | N | Y | Y | 106 |  |  | 10/22/1996 11:12:48 PM | 1996-04-12 | 6/3/1998 6:51:41 PM | EMS |  | Concentration | mg/L | Y | (fl. conc.) |  | Magnesium-Soluble ion content (fl. conc.) | 7439-95-4 | Metals - Soluble; | Metals - Soluble; EMS code: MG-S | Chemical | Soluble | LAB
MG-S | mg/L | Magnesium-Soluble ion content | 5138 | AA Analysis | AA Analysis | 0.01 | Y | Y | Y | 337 |  |  | 10/24/1996 10:11:48 AM | 1996-04-12 | 6/3/1998 6:51:36 PM | EMS |  | Concentration | mg/L | Y | (fl. conc.) |  | Magnesium-Soluble ion content (fl. conc.) | 7439-95-4 | Metals - Soluble; | Metals - Soluble; EMS code: MG-S | Chemical | Soluble | LAB
MG-S | mg/L | Magnesium-Soluble ion content | E6020B | EPA 6020B | See EPA 6020B and CSSS Ch. 15 | 1 | Y | Y | Y | 0 |  |  |  | 2024-02-02 |  | EMS |  | Concentration | mg/L | Y | (fl. conc.) |  | Magnesium-Soluble ion content (fl. conc.) | 7439-95-4 | Metals - Soluble; | Metals - Soluble; EMS code: MG-S | Chemical | Soluble | LAB
MG-S | mg/L | Magnesium-Soluble ion content | 5070 | Ion Chromatog-Cations II | (a) For the Soluble Particulates Pkg. a portion of filtrate from the Particulate - Total procedure made up to 500 mL is analyzed by IC. Intermediate results are reported in mg/L. Loading results are calculated and reported with units of mg/dm2/d.(b) For precipitation samples, all samples are filtered through a 0.22 um fritted glass filter prior to injection. The cations of interest are separated through a cation "guard" and cation "separator" column with a methane sulphonic acid eluent. After the cations are separated they exit at various times from the column in a background of eluent. A self regeneration membrane suppressor is attached to the end of the column to neutralize the acid before detection, and form the corresponding hydroxide. Note that ammonium is reported as NH4 not as N. (c) For HiVol samples, discs cut from the HiVol filter are extracted with DI water and the soluble anions are analyzed by IC. Intermediate results are reported in units of mg/L and HiVol loading results are reported as ug/m3. | 0.01 | Y | Y | Y | 1504 | 4/30/1997 3:50:00 PM | 1/7/2002 2:33:00 PM | 6/24/2003 10:43:21 AM | 1996-04-12 | 6/3/1998 6:51:36 PM | EMS |  | Concentration | mg/L | Y | (fl. conc.) |  | Magnesium-Soluble ion content (fl. conc.) | 7439-95-4 | Metals - Soluble; | Metals - Soluble; EMS code: MG-S | Chemical | Soluble | LAB
MG-S | mg/L | Magnesium-Soluble ion content | 5060 | Ion Chrom- Cations I | Ion Chrom- Cations | 0.01 | Y | Y | Y | 345 | 5/6/1997 4:37:00 PM | 5/7/1997 12:02:00 PM | 2/17/1999 8:58:07 AM | 1996-04-12 | 6/3/1998 6:51:42 PM | EMS |  | Concentration | mg/L | Y | (fl. conc.) |  | Magnesium-Soluble ion content (fl. conc.) | 7439-95-4 | Metals - Soluble; | Metals - Soluble; EMS code: MG-S | Chemical | Soluble | LAB
MG-S | ug/g | Magnesium-Soluble ion content | CAL2 | Calculation (ug/g) | Calculation (ug/g) | 0.5 | Y | Y | Y | 0 |  |  |  | 2024-02-02 |  | EMS |  | MassPerMass | mg/kg | Y | (mass conc.) |  | Magnesium-Soluble ion content (mass conc.) | 7439-95-4 | Metals - Soluble; | Metals - Soluble; EMS code: MG-S | Chemical | Soluble | LAB
MG-T | ug/g | Magnesium Total | 0192 | Hom, Nitric/Perchlor,ICP | Sample already in a dry state so does not require drying before homogenization and nitric perchlorid digestion followed by ICP analysis. | 1 | Y | Y | Y | 708 |  |  | 10/24/1996 12:50:25 PM | 1996-04-12 | 6/3/1998 6:51:36 PM | EMS |  | MassPerMass | mg/kg | Y | (mass conc.) |  | Magnesium Total (mass conc.) | 7439-95-4 | Metals - Total; | Metals - Total; EMS code: MG-T | Chemical | Total | LAB
MG-T | mg/L | Magnesium Total | F038 | HNO3/HCl dig,ICP-AES | Dissolved and extractable water samples are ready for analysis once collected. Total samples are autoclaved with HNO3 and hydrochloric acid (HCI) in capped Teflon vials. Analysis is performed via Inductively Coupled Argon Plasma Atomic Emission Spectrometry (ICP-AES)(simultaneous multielement analysis), Graphite Furnace Atomic Absorption Spectrometry (GF-AAS) (low level analysis) and/or Flame Atomic Absorption Spectrometry (F-AAS). | 0.1 | Y | Y | Y | 5408 | 8/1/1996 6:04:00 PM | 5/6/2008 | 9/11/2019 8:47:42 AM | 1997-05-08 | 6/3/1998 6:51:46 PM | MPAK |  | Concentration | mg/L | Y | (fl. conc.) |  | Magnesium Total (fl. conc.) | 7439-95-4 | Metals - Total; | Metals - Total; EMS code: MG-T | Chemical | Total | LAB
MG-T | % | Magnesium Total | MS61 | ALS Minerals ME-MS61 | ALS Minerals ME-MS61. A prepared sample is digested with HNO3; HF; and HClO4 in a graphite heating block to near dryness. After cooling; the sample is boiled in diluted HCl. Bulk to volume with deionized water; mixed and analyzed by ICP-AES and ICP-MS for all elements. | 0.01 | Y | Y | Y | 0 |  |  |  | 2024-01-29 |  | EMS |  | DimensionlessRatio | % | N | (ratio) |  | Magnesium Total (ratio) | 7439-95-4 | Metals - Total; | Metals - Total; EMS code: MG-T | Chemical | Total | LAB
MG-T | mg/L | Magnesium Total | F121 | Acid,ICPMS | Acidified, no digestion, ICPMS | 0.01 | Y | Y | Y | 12049 | 10/18/2007 | 8/16/2024 | 8/26/2024 8:06:03 AM | 2002-07-17 | 7/16/2012 3:33:35 PM | UMULLER | RDGIBSON | Concentration | mg/L | Y | (fl. conc.) |  | Magnesium Total (fl. conc.) | 7439-95-4 | Metals - Total; | Metals - Total; EMS code: MG-T | Chemical | Total | LAB
MG-T | mg/dm2/d | Magnesium Total | DUST | HNO3 dig.(microwave or hot block),ICPMS | HNO3 dig.(microwave or hot block);ICPMS |  | Y | Y | Y | 420 | 7/19/2016 1:11:00 PM | 7/31/2024 | 7/31/2024 11:23:12 AM | 2008-02-07 |  | RDGIBSON |  | MassFlux | mg/dm²/d | Y | (mass flux) |  | Magnesium Total (mass flux) | 7439-95-4 | Metals - Total; | Metals - Total; EMS code: MG-T | Chemical | Total | LAB
MG-T | ug/g | Magnesium Total | MTID | MET-DRY-CCMS-VA | Metals in Tissue by CRC ICPMS (DRY) Methodology: This method is conducted following British Columbia Lab Manual method "Metals in Animal Tissue and Vegetation (Biota) - Prescriptive". Tissue samples are homogenized and sub-sampled prior to hotblock digestion with nitric and hydrochloric acids, in combination with addition of hydrogen peroxide. Instrumental analysis is by collision cell inductively coupled plasma - mass spectrometry (modified from EPA Method 6020A). Method Limitation: This method employs a strong acid/peroxide digestion, and is intended to provide a conservative estimate of bio-available metals. Near complete recoveries are achieved for most toxicologically important metals, but elements associated with recalcitrant minerals may be only partially recovered. | 2 | Y | Y | Y | 122 | 8/13/2014 | 9/8/2022 | 10/26/2022 7:14:15 AM | 2015-03-09 | 3/11/2015 4:12:05 PM | RDGIBSON | EMS | MassPerMass | mg/kg | Y | (mass conc.) |  | Magnesium Total (mass conc.) | 7439-95-4 | Metals - Total; | Metals - Total; EMS code: MG-T | Chemical | Total | LAB
MG-T | mg/L | Magnesium Total | X073 | FA,HNO3,Direct Flame,AA | Samples are field-acidified with HNO3; HNO3 digestion; analysis by AA. | 0.001 | Y | Y | Y | 53 |  |  | 1/25/2007 4:06:12 PM | 1996-04-12 | 6/3/1998 6:51:36 PM | EMS |  | Concentration | mg/L | Y | (fl. conc.) |  | Magnesium Total (fl. conc.) | 7439-95-4 | Metals - Total; | Metals - Total; EMS code: MG-T | Chemical | Total | LAB
MG-T | ug/g | Magnesium Total | EQ02 | EQUIS (ug/g) | EQUIS (ug/g) | 1 | N | Y | Y | 783 |  |  | 10/23/1996 12:08:28 AM | 1996-04-12 | 6/3/1998 6:51:41 PM | EMS |  | MassPerMass | mg/kg | Y | (mass conc.) |  | Magnesium Total (mass conc.) | 7439-95-4 | Metals - Total; | Metals - Total; EMS code: MG-T | Chemical | Total | LAB
MG-T | ug/g | Magnesium Total | ICPM | Metals in Soil by CRC ICPMS | This analysis is carried out using procedures from CSR Analytical Method: "Strong Acid Leachable Metals (SALM) in Soil", BC Ministry of Environment, 26 June 2009, and procedures adapted from EPA Method 200.2. The sample is manually homogenized, dried at 60 degrees Celsius, sieved through a 63 um (230 mesh) sieve, and a representative subsample of the dry material is weighed. The sample is then digested at 95 degrees Celsius for 2 hours by block digester using concentrated nitric and hydrochloric acids. Instrumental analysis is by collision cell inductively coupled plasma - mass spectrometry (modifed from EPA Method 6020A). Method Limitation: This method is not a total digestion technique. It is a very strong acid digestion that is intended to dissolve those metals that may be environmentally available. By design, elements bound in silicate structures are not normally dissolved by this procedure as they are not usually mobile in the environment. | 10 | Y | Y | Y | 103 | 9/19/2017 | 9/24/2023 | 1/12/2024 11:16:52 AM | 2018-01-09 |  | JKROGH |  | MassPerMass | mg/kg | Y | (mass conc.) |  | Magnesium Total (mass conc.) | 7439-95-4 | Metals - Total; | Metals - Total; EMS code: MG-T | Chemical | Total | LAB
MG-T | ug/m3 | Magnesium Total | 6040 | Aqua Regia Dig:ICP | Aqua regia digestion is used to bring the metals into solution prior to ICP analysis. For Hi-Vol samples, preliminary results are reported as mg/L and24 hr loadings as ug/m3 | 0.001 | Y | Y | Y | 437 | 2/13/2002 | 2/9/2005 | 2/23/2005 1:13:47 AM | 2004-02-10 |  | UMULLER |  | AirConcentration | µg/m³ | Y | (air conc.) |  | Magnesium Total (air conc.) | 7439-95-4 | Metals - Total; | Metals - Total; EMS code: MG-T | Chemical | Total | LAB
MG-T | mg/L | Magnesium Total | 1349 | Tot Metals in Seawater by CRC ICPMS (BC) | Tot Metals in Seawater by CRC ICPMS (BC) | 1 | Y | Y | Y | 98 | 8/16/2019 | 12/5/2021 | 12/13/2021 12:34:16 PM | 2019-09-30 |  | KATFRASE |  | Concentration | mg/L | Y | (fl. conc.) |  | Magnesium Total (fl. conc.) | 7439-95-4 | Metals - Total; | Metals - Total; EMS code: MG-T | Chemical | Total | LAB
MG-T | mg/L | Magnesium Total | EQ01 | EQUIS (mg/L) | EQUIS (mg/L) | 0.02 | N | Y | Y | 4254 |  |  | 10/24/1996 3:16:18 PM | 1996-04-12 | 6/3/1998 6:51:36 PM | EMS |  | Concentration | mg/L | Y | (fl. conc.) |  | Magnesium Total (fl. conc.) | 7439-95-4 | Metals - Total; | Metals - Total; EMS code: MG-T | Chemical | Total | LAB
MG-T | mg/L | Magnesium Total | F048 | Acid dig,conc'n,nebul.,ICP-OES | This method provides an automated procedure for the Inductively Coupled Argon Plasma Optical Emission Spectrophotometric (ICP-OES) determination and quantitation of trace amounts of metals in surface, ground and precipitation waters after a manual acid digestion and concentration procedure and ultrasonic nebulisation. Organically-bound metals and metal complexes are converted to metal ions by digestion with nitric acid [method details says HCl]. To increase the signal detected by the ICP, the samples are concentrated 10 times. The liquid sample is pumped onto the piezoelectric transducer, where the ultrasonic action efficiently converts it to a fine, dense aerosol. The nebulizer carrier gas flow then transports the wet aerosol through a heated U-tube, where the solvent is vaporized. Subsequently, most of the solvent vapour is removed from the aerosol stream by passing the stream through a chilled condenser. The result is a dry, analyte-laden aerosol which is introduced to the plasma. There, the high temperature desolvates the aerosol and dissociates the sample material into individual atoms and ions, which are excited to emit light at wave lengths characteristic of the elements present. Dissolved samples are field filtered through 0.45 micron cellulose acetate filter prior to acid preservation. Samples for extractable analysis are acidified with 2 mL concentrated HNO3 per litre and allowed to stand a minimum of 24 hours before analysis. Clients providing these samples are advised to preserve samples at the collection site. | 0.05 | Y | Y | Y | 89 | 7/25/2005 7:10:00 AM | 7/17/2012 7:03:00 AM | 7/26/2012 12:28:00 PM | 1998-11-13 |  | NPEPPIN |  | Concentration | mg/L | Y | (fl. conc.) |  | Magnesium Total (fl. conc.) | 7439-95-4 | Metals - Total; | Metals - Total; EMS code: MG-T | Chemical | Total | LAB
MG-T | ug/g wet | Magnesium Total | X217 | Hom,Nitric/Perchl,ICP | Sample is wet but not dried before homogenization and nitric perchloric digestion followed by ICP analysis. | 0.05 | Y | Y | Y | 74 |  |  | 10/24/1996 10:01:27 AM | 1996-04-12 | 6/3/1998 6:51:42 PM | EMS |  | MassPerMass | µg/g | Y | (mass conc.) | (wet) | Magnesium Total (wet) (mass conc.) | 7439-95-4 | Metals - Total; | Metals - Total; EMS code: MG-T | Chemical | Total | LAB
MG-T | mg/L | Magnesium Total | F068 | HNO3dig,TeflonTube/Autoclave,ICP-MS | This method is applicable to determiniation of dissolved, extractable, and total recoverable metals in ground waters, surface waters, drinking waters, and wastewaters. Dissolved, extractable, and total recoverable sample fractions are acidified to pH <1.5 with analytical grade nitric acid (HNO3) to prevent the metals from adsorbing to the container material or precipitating. Extractable and total recoverable water samples are collected in 250 ml acid washed polyethylene bottles and are preserved with Nitric acid (0.5 ml acid for 100 mLs of sample). All preserved samples are stored at room temperature and can be held for up to 6 months. Unfiltered, un-preserved samples are stored at 4C for no more than 30 days. For dissolved and extractable water fractions, no sample preparation is necessary. Samples are simply poured in autosampler tubes (i.e. 15 ml centrifuge tubes), and placed on the autosampler and analyzed. Samples for total recoverable metals are digested with Nitric acid in special Teflon tubes in an autoclave prior to analyses. | 5e-05 | Y | Y | Y | 1675 | 6/1/1999 1:32:00 PM | 4/21/2004 7:34:00 PM | 9/10/2019 3:51:21 PM | 1999-06-28 |  | NPEPPIN |  | Concentration | mg/L | Y | (fl. conc.) |  | Magnesium Total (fl. conc.) | 7439-95-4 | Metals - Total; | Metals - Total; EMS code: MG-T | Chemical | Total | LAB
MG-T | mg/L | Magnesium Total | X145 | Dig, ICP | Dig; ICP | 0.01 | Y | Y | Y | 854 | 5/8/2003 9:40:00 AM | 9/22/2016 | 6/13/2024 3:58:21 PM | 1996-04-12 | 6/3/1998 6:51:36 PM | EMS |  | Concentration | mg/L | Y | (fl. conc.) |  | Magnesium Total (fl. conc.) | 7439-95-4 | Metals - Total; | Metals - Total; EMS code: MG-T | Chemical | Total | LAB
MG-T | mg/L | Magnesium Total | X285 | HNO3 Dig:ICP Analysis | HNO3 Dig:ICP Analysis | 0.1 | Y | Y | Y | 670 | 8/10/1993 | 10/5/1993 | 4/30/2004 1:59:45 PM | 1998-05-22 | 6/3/1998 6:51:49 PM | NPEPPIN |  | Concentration | mg/L | Y | (fl. conc.) |  | Magnesium Total (fl. conc.) | 7439-95-4 | Metals - Total; | Metals - Total; EMS code: MG-T | Chemical | Total | LAB
MG-T | mg/L | Magnesium Total | F081 | Aliquot,HNO3/HCl Dig,ICP | Aliquot microwave digested with 3:1 HNO3/HCL. Analysis by ICP. For HiVol samples preliminary result are reported in mg/L and loading results are reported in ug/m3 |  | Y | Y | Y | 76 | 10/6/2010 10:03:00 AM | 7/25/2012 7:11:00 AM | 9/26/2012 1:06:57 PM | 2000-06-13 |  | NPEPPIN |  | Concentration | mg/L | Y | (fl. conc.) |  | Magnesium Total (fl. conc.) | 7439-95-4 | Metals - Total; | Metals - Total; EMS code: MG-T | Chemical | Total | LAB
MG-T | ug/g | Magnesium Total | X551 | Metals in Biota (dry), total, acid digestion, ICP-SFMS | 0.1 gram of dried, homogenized biota is weighed in a microwave TFM vessel, and 8mL of Nitric Acid and 1mL of Hydrogen Peroxide are added. The sample is digested at 200 degrees C for 15 minutes in a closed vessel high pressure microwave oven. The sample is cooled, the volume adjusted to 100mL and introduced to an iductively coupled argon plasma-sector field mass spectrometer (ICP-SFMS) by pneumatic nebulization. Each respective element is measured at a specific mass to charge ratio, m/z value expressed in amu (atomic mass units), and at a specified instrument resolution. LR: Low Resolution ~300, MR: Mid Resolution ~4000 and HR: Hich Resolution ~10000. Ref: Method number 02-2705, National Laboratory for Environmental Testing, CCIW, Environment Canada, Burlington. Requested April 2003. | 2 | Y | Y | Y | 49 | 8/11/2015 | 10/30/2023 | 11/28/2023 2:21:33 PM | 2015-03-11 |  | RDGIBSON |  | MassPerMass | mg/kg | Y | (mass conc.) |  | Magnesium Total (mass conc.) | 7439-95-4 | Metals - Total; | Metals - Total; EMS code: MG-T | Chemical | Total | LAB
MG-T | ug/g | Magnesium Total | PB02 | PBM,HNO3/HCl dig,dry60C,100mesh, | Performance Based Method. Samples are dried at 60 C and sieved using a .149 mm (100 mesh) sieve. Solid samples are digested with the combination of nitric and hydrochloric acids with a 1:1 ratio of acids. For environmental purposes, strong acid leaches are used routinely since they have the advantage of speed, low cost, and superio reproducibility. All metals trapped within the refractory matrix are unlikely to have environmental consequences. Instrumental analysis of extracts produced using this digestion method can be performed by a variety of techniques. |  | Y | Y | Y | 10 | 9/13/2001 | 9/13/2001 | 11/6/2002 3:52:36 PM | 2001-07-18 |  | NPEPPIN |  | MassPerMass | mg/kg | Y | (mass conc.) |  | Magnesium Total (mass conc.) | 7439-95-4 | Metals - Total; | Metals - Total; EMS code: MG-T | Chemical | Total | LAB
MG-T | % | Magnesium Total | 6041 | Aqua Regia by ICPMS | These were sediment/soils digested with Aqua-Regia (Nitric Acid & Hydrochloric Acids) and then analysis by ICPMS. Results would be reported in mg/kg and/or %. | 0.01 | Y | Y | Y | 92 | 10/13/2019 | 7/27/2024 | 7/29/2024 9:13:08 AM | 2019-09-06 |  | KATFRASE |  | DimensionlessRatio | % | N | (ratio) |  | Magnesium Total (ratio) | 7439-95-4 | Metals - Total; | Metals - Total; EMS code: MG-T | Chemical | Total | LAB
MG-T | ug/g | Magnesium Total | X230 | Hom,HNO3/HCl,ICP | The sample is dried, ground to pass 100 mesh, and digested with a nitric-hydrochloric acid mixture to solubilize the solid matter and to remove the organic material by oxidation and volatilization. The completed digests are submitted to metals instrumentation personnel for metals analysis by Inductively Coupled Atomic Emission Spectrophotometry (ICP-AES), or hydride analysis by atomic absorption. | 0.5 | Y | Y | Y | 182 | 2/12/2002 | 12/30/2005 | 2/3/2006 1:13:13 AM | 1996-04-12 | 6/3/1998 6:51:36 PM | EMS |  | MassPerMass | mg/kg | Y | (mass conc.) |  | Magnesium Total (mass conc.) | 7439-95-4 | Metals - Total; | Metals - Total; EMS code: MG-T | Chemical | Total | LAB
MG-T | ug/g wet | Magnesium Total | X560 | Metals in Biota (wet), total, acid digestion, ICP-MS | Five grams of wet, homogenized biota is weighed in a centrifuge tube, and 2.5 mL of Nitric Acid and 2.5 mL of Hydrochloric acid are added. The sample is digested at room temperature for one hour and then digested at 90-95 degrees C in a water bath or heat block for one more hour. The sample is cooled and the volume is made to 50 mL with DI. The digestate is then diluted at ten fold and introduced to a collision/reaction cell inductively coupled argon plasma mass spectrometer (CRC ICPMS) by pneumatic nebulization. Each respective element is measured at a specific mass to charge ratio, m/z value expressed in amu (atomic mass units). - Maxxam Analytics (BC). Added to EMS Dec 10, 2008. | 10 | Y | Y | Y | 541 | 11/15/2008 4:00:00 AM | 10/30/2023 | 11/28/2023 2:21:33 PM | 2008-12-11 | 3/11/2015 4:17:10 PM | RDGIBSON | EMS | MassPerMass | µg/g | Y | (mass conc.) | (wet) | Magnesium Total (wet) (mass conc.) | 7439-95-4 | Metals - Total; | Metals - Total; EMS code: MG-T | Chemical | Total | LAB
MG-T | mg/L | Magnesium Total | F033 | Acid Dig,Centrifuge,ICPMS | For the determination of total recoverable elements in water or waste water, take 100 ml aliquot from a well mixed, acid preserved, sample containing not more than 0.25% (w/v) total solids and transfer to a 250 ml Griffin beaker. (If total solids are greater than 0.25% reduce the size of the aliquot by a proportionate amount). Add 1 ml of conc. nitric acid and 0.5 ml conc. hydrochloric acid. Heat on a hot plate until the volume has been reduced to 15-20ml, ensuring that the sample does not boil. Allow to cool and transfer to a 50 ml volumetric flask. Dilute to volume with ASTM type I water. Centrifuge the sample or allow to stand overnight to separate insoluble material. Prior analysis, pipette 20 ml into a 50 ml volumetric flask and dilute to volume with ASTM type I water. Analysis should be performed within seventy two hours of the completed preparation. Note that different laboratories may have very different field preservation, digestion procedures and detection limits for this method, so these should be documented. | 0.0005 | Y | Y | Y | 26880 | 10/15/1996 9:26:00 AM | 8/22/2024 3:24:00 AM | 8/25/2024 9:47:36 PM | 1997-01-16 | 6/3/1998 6:51:46 PM | MPAK |  | Concentration | mg/L | Y | (fl. conc.) |  | Magnesium Total (fl. conc.) | 7439-95-4 | Metals - Total; | Metals - Total; EMS code: MG-T | Chemical | Total | LAB
MG-T | ug/g | Magnesium Total | FS01 | D&G,Nitric/HCl:ICP | Samples are dried at 60 C, then either sieved to -100mesh (150 um) and rolled to homogenize, or finely ground with a mortar and pestle or a shatter box laboratory mill and rolled (depending on the submitter's requirements). The samples are then weighed into Teflon digestion liners and digested with HNO3 and HCl in a microwave oven. Samples are cooled, volumized, and allowed to settle overnight before being decanted. Analysis is then performed via Inductively Coupled Argon Plasma Atomic Emission Spectrometry (ICP-AES) (Simultaneous multi element analysis). In ICP Spectrometry, the liquid sample is pumped to the nebulizer, where it is convertd into an aerosol. the nebulizer carrier gas flow than transports the wet aerosol past the impact bead, where the large droplets condense out and flow to the waste container. The remaining fine aerosol is carried into the plasma. There, the high temperature desolvates the aerosol and dissociates the sample material into individual atoms and ions, which are excited to emit light at wavelengths characteristic of the elements present. A diffraction grating in the spectrometer diffracts the various wavelengths of light through exit slits to the photo-multiplier tubes, which convert the light energy to electronic signals. | 20 | Y | Y | Y | 2642 | 8/30/1996 1:52:00 PM | 10/10/2007 | 5/24/2008 1:15:02 AM | 1996-11-20 | 6/3/1998 6:51:46 PM | MPAK |  | MassPerMass | mg/kg | Y | (mass conc.) |  | Magnesium Total (mass conc.) | 7439-95-4 | Metals - Total; | Metals - Total; EMS code: MG-T | Chemical | Total | LAB
MG-T | ug/g | Magnesium Total | FS82 | PBM,dry60C,10mesh,SALM-ICP-MS | PBM;dry60C;10mesh;SALM-ICP-MS | 100 | Y | Y | Y | 3307 | 5/21/2002 | 7/24/2024 | 7/29/2024 5:17:09 PM | 2002-11-12 |  | UMULLER |  | MassPerMass | mg/kg | Y | (mass conc.) |  | Magnesium Total (mass conc.) | 7439-95-4 | Metals - Total; | Metals - Total; EMS code: MG-T | Chemical | Total | LAB
MG-T | mg/L | Magnesium Total | 0042 | GI HNO3 Dig: ICP | Sample aliquot taken from unfiltered unpreserved General Ions bottle; HNO3 digestion; ICP analysis. | 0.02 | Y | Y | Y | 12776 | 9/20/1996 | 7/5/2019 3:21:00 PM | 7/12/2019 6:49:13 PM | 1996-04-12 | 6/3/1998 6:51:41 PM | EMS |  | Concentration | mg/L | Y | (fl. conc.) |  | Magnesium Total (fl. conc.) | 7439-95-4 | Metals - Total; | Metals - Total; EMS code: MG-T | Chemical | Total | LAB
MG-T | mg/L | Magnesium Total | F082 | Aliquot,HNO3/HCL dig,ICPMS | Aliquot microwave digested with 3:1 HNO3/HCL. Analysis by ICPMS. For HiVol samples preliminary result are reported in mg/L and loading results are reported in ug/m3 |  | Y | Y | Y | 24818 | 1/24/2011 | 8/23/2024 | 8/26/2024 4:53:05 PM | 2000-06-13 |  | NPEPPIN |  | Concentration | mg/L | Y | (fl. conc.) |  | Magnesium Total (fl. conc.) | 7439-95-4 | Metals - Total; | Metals - Total; EMS code: MG-T | Chemical | Total | LAB
MG-T | mg/L | Magnesium Total | CS01 | System Calculation | System Calculation (mg/L) |  | N | Y | N | 8 |  |  | 3/4/2008 6:49:51 AM | 1996-04-12 | 6/3/1998 6:51:43 PM | EMS |  | Concentration | mg/L | Y | (fl. conc.) |  | Magnesium Total (fl. conc.) | 7439-95-4 | Metals - Total; | Metals - Total; EMS code: MG-T | Chemical | Total | LAB
MG-T | ug/g | Magnesium Total | 0190 | D&G:N/P Dig:ICP Anal. | Sample is dried and grind-mixed before nitric perchloric digestion followed by ICP analysis. | 1 | Y | Y | Y | 949 |  |  | 10/24/1996 3:06:16 PM | 1996-04-12 | 6/3/1998 6:51:36 PM | EMS |  | MassPerMass | mg/kg | Y | (mass conc.) |  | Magnesium Total (mass conc.) | 7439-95-4 | Metals - Total; | Metals - Total; EMS code: MG-T | Chemical | Total | LAB
MG-T | ug/m3 | Magnesium Total | F088 | Aliquot,HNO3/HCl dig, ICPMS,loading | Aliquot microwave digested with 3:1 HNO3/HCL. Analysis by ICPMS. For HiVol samples preliminary result are reported in mg/L and loading results are reported in ug/m3 |  | Y | Y | Y | 565 | 8/17/2000 10:46:00 AM | 12/14/2006 | 12/30/2006 1:01:36 AM | 2000-08-16 |  | NPEPPIN |  | AirConcentration | µg/m³ | Y | (air conc.) |  | Magnesium Total (air conc.) | 7439-95-4 | Metals - Total; | Metals - Total; EMS code: MG-T | Chemical | Total | LAB
MG-T | mg/L | Magnesium Total | 0200 | 20%A.R.Dig:AA Anal. | 20%A.R.Dig:AA Anal. | 0.02 | Y | Y | Y | 151 |  |  | 10/24/1996 11:35:00 AM | 1996-04-12 | 6/3/1998 6:51:36 PM | EMS |  | Concentration | mg/L | Y | (fl. conc.) |  | Magnesium Total (fl. conc.) | 7439-95-4 | Metals - Total; | Metals - Total; EMS code: MG-T | Chemical | Total | LAB
MG-T | ug/g | Magnesium Total | FS89 | Homo,HNO3 dig,ICP-AES | Homo;HNO3 dig;ICP-AES | 20 | Y | Y | Y | 4 | 1/5/2001 | 8/20/2001 12:00:00 PM | 10/23/2002 1:42:45 PM | 2001-03-08 |  | NPEPPIN |  | MassPerMass | mg/kg | Y | (mass conc.) |  | Magnesium Total (mass conc.) | 7439-95-4 | Metals - Total; | Metals - Total; EMS code: MG-T | Chemical | Total | LAB
MG-T | ug/g wet | Magnesium Total | X552 | Metals in Biota (wet), total, acid digestion, ICP-SFMS | One gram of wet, homogenized biota is weighed in a microwave TFM vessel, and 8mL of Nitric Acid and 1mL of Hydrogen Peroxide are added. The sample is digested at 200 degrees C for 15 minutes in a closed vessel high pressure microwave oven. The sample is cooled, the volume adjusted to 100mL and introduced to an inductively coupled argon plasma-sector field mass spectrometer (ICP-SFMS) by pneumatic nebulization. Each respective element is measured at a specific mass to charge ratio, m/z value expressed in amu (atomic mass units), and at a specified instrument resolution. LR: Low Resolution ~300, MR: Mid Resolution ~4000. Ref: 02-2704, National Laboratory for Environmental Testing, CCIW, Environment Canada, Burlington. Requested April 2003. | 0.4 | Y | Y | Y | 37 | 8/9/2016 | 10/16/2023 | 11/28/2023 2:21:32 PM | 2015-03-11 |  | RDGIBSON |  | MassPerMass | µg/g | Y | (mass conc.) | (wet) | Magnesium Total (wet) (mass conc.) | 7439-95-4 | Metals - Total; | Metals - Total; EMS code: MG-T | Chemical | Total | LAB
MG-T | mg/L | Magnesium Total | MWID | Metals in Water by ICPMS (digested) | Metals in Water by ICPMS (digested) Based on EPA 200.8 | 1 | Y | Y | Y | 36643 |  |  | 7/31/2024 1:46:41 PM | 2009-12-24 |  | RDGIBSON |  | Concentration | mg/L | Y | (fl. conc.) |  | Magnesium Total (fl. conc.) | 7439-95-4 | Metals - Total; | Metals - Total; EMS code: MG-T | Chemical | Total | LAB
MG-T | mg/L | Magnesium Total | 0040 | HNO3 Dig: ICP Analysis | HNO3 Dig: ICP Analysis | 0.02 | Y | Y | Y | 12376 |  |  | 3/23/2010 5:08:08 PM | 1996-04-12 | 6/3/1998 6:51:36 PM | EMS |  | Concentration | mg/L | Y | (fl. conc.) |  | Magnesium Total (fl. conc.) | 7439-95-4 | Metals - Total; | Metals - Total; EMS code: MG-T | Chemical | Total | LAB
MG-T | ug/g | Magnesium Total | 0196 | D&G: LTA/HCl: ICP | D&G: LTA/HCl: ICP | 1 | Y | Y | Y | 27 |  |  | 10/24/1996 2:08:12 PM | 1996-04-12 | 6/3/1998 6:51:41 PM | EMS |  | MassPerMass | mg/kg | Y | (mass conc.) |  | Magnesium Total (mass conc.) | 7439-95-4 | Metals - Total; | Metals - Total; EMS code: MG-T | Chemical | Total | LAB
MG-T | mg/L | Magnesium Total | METN | MET-T-NP-CCMS-VA | Ultra trace metals in water are analyzed by CRC ICPMS, based on US EPA Method 6020B (July 2014). The detection limits provided can only be met for undigested samples. This procedure is intended for colorless, non-turbid, acid-preserved water samples (i.e. pristine water samples), having turbidity < 1 NTU and no odor. Where turbidity exceeds 1 NTU, and/or the sample is colored and has an odor, results may be biased low compared to true Total Metals concentrations. ALS recommends that turbidity analysis be requested on samples submitted for this test to aid with interpretation of results. Where turbidity is <1NTU, undigested metals are equivalent to total metals concentrations. | 0.01 | Y | Y | Y | 3028 | 11/14/2017 | 8/19/2024 | 8/26/2024 1:03:06 PM | 2019-08-20 |  | KATFRASE |  | Concentration | mg/L | Y | (fl. conc.) |  | Magnesium Total (fl. conc.) | 7439-95-4 | Metals - Total; | Metals - Total; EMS code: MG-T | Chemical | Total | LAB
MG-T | mg/L | Magnesium Total | X561 | Total metals: FA, HNO3, HR-ICPMS (No Prep)Prep) | Ultra trace metals in water are analyzed by high resolution inductively coupled plasma mass spectrometry (HR-ICPMS), modified from US EPA 200.8, (Revision 5.5). The detection limits provided can only be met for undigested samples. This procedure is intended for colorless, non-turbid, acid-preserved water samples (i.e. pristine water samples), having turbidity < 1 NTU and no odor. Where turbidity exceeds 1 NTU, and/or the sample is colored and has an odor, results may be biased low compared to true Total Metals concentrations. ALS recommends that turbidity analysis be requested on samples submitted for this test to aid with interpretation of results. | 0.01 | Y | Y | Y | 2755 | 6/8/2015 | 10/5/2023 | 1/31/2024 2:37:41 PM | 2015-10-22 |  | RDGIBSON |  | Concentration | mg/L | Y | (fl. conc.) |  | Magnesium Total (fl. conc.) | 7439-95-4 | Metals - Total; | Metals - Total; EMS code: MG-T | Chemical | Total | LAB
MG-T | ug/g wet | Magnesium Total | BPSY | MET-WET-BPSY-CCMS-VA | Metals in Animal Tissue and Vegetation (Biota) - Prescriptive. EPA 200.3/6020B This method is conducted following British Columbia Lab Manual method "Metals in Animal Tissue and Vegetation (Biota) - Prescriptive". Tissue samples are digested with nitric and hydrochloric acids; in combination with addition of hydrogen peroxide. Instrumental analysis is by collision cell inductively coupled plasma - mass spectrometry (modified from EPA Method 6020B). | 0.4 | Y | Y | Y | 2 | 11/7/2019 | 11/7/2019 | 6/10/2021 1:57:10 PM | 2021-06-04 |  | EMS |  | MassPerMass | µg/g | Y | (mass conc.) | (wet) | Magnesium Total (wet) (mass conc.) | 7439-95-4 | Metals - Total; | Metals - Total; EMS code: MG-T | Chemical | Total | LAB
MG-T | ug/g | Magnesium Total | FB04 | Freeze Dry,Hom,Acid Dig,ICP | Sample moisture is determined by weighing samples before and after freeze-drying. Freeze-dried samples are ground in a Braun blender before being weighed into Teflon PFA liners and digested with concentrated HNO3, concentrated HCl, and 50% H2O2 in a microwave oven; which results in a complete sample dissolution. Samples are collected and volumized with Type I water. Analysis is performed via Inductively Coupled Argon Plasma Atomic Emission Spectrometry (ICP-AES) (simultaneous multielement analysis), Graphite Furnace Atomic Absorption Spectrometry (GF-AAS) (low level Cd and Pb analysis), Cold Vapor Atomic Fluorescence Spectrometry (CV-AFS) (Mercury analysis) and/or Flame Atomic Absorption Spectrometry (F-AAS) (K analysis). | 8 | Y | Y | Y | 241 | 8/22/1996 6:11:00 PM | 2/1/2002 2:30:00 PM | 6/23/2003 1:15:31 PM | 1997-02-27 | 6/3/1998 6:51:46 PM | MPAK |  | MassPerMass | mg/kg | Y | (mass conc.) |  | Magnesium Total (mass conc.) | 7439-95-4 | Metals - Total; | Metals - Total; EMS code: MG-T | Chemical | Total | LAB
MG-T | % | Magnesium Total | MS41 | ALS Minerals ME-MS41 | ALS Minerals ME-MS41. A prepared sample is digested with aqua regia in a graphite heating block. After cooling; the resulting solution is diluted to with deionized water; mixed and analyzed by ICP-AES. Samples are then analyzed by ICP-MS for the remaining suite of elements. | 0.01 | Y | Y | Y | 0 |  |  |  | 2024-01-29 |  | EMS |  | DimensionlessRatio | % | N | (ratio) |  | Magnesium Total (ratio) | 7439-95-4 | Metals - Total; | Metals - Total; EMS code: MG-T | Chemical | Total | LAB
MG-T | mg/L | Magnesium Total | 0020 | HNO3 Dig:AA Analysis | HNO3 Dig:AA Analysis | 0.02 | Y | Y | Y | 97 | 6/25/2010 3:15:00 PM | 12/1/2010 | 12/9/2010 3:55:27 PM | 1996-04-12 | 6/3/1998 6:51:36 PM | EMS |  | Concentration | mg/L | Y | (fl. conc.) |  | Magnesium Total (fl. conc.) | 7439-95-4 | Metals - Total; | Metals - Total; EMS code: MG-T | Chemical | Total | LAB
MG-T | mg/L | Magnesium Total | X170 | ICP | This code is a catch-all for a large variety of ICP methods including water and also solids digested into solutions as an intermediate result. Different procedures will have greatly different detection limits. (a) For air particulates, preliminary results are reported in mg/L and loading results as mg/dm2/d. (b) For Hi Vol samples preliminiary results are same as reported in mg/L and loading results as ug/m3. (c) Note that for Rain samples there is no filtration nor digestion steps. | 0.1 | Y | Y | Y | 10835 | 3/28/1996 6:03:00 PM | 11/7/2018 | 12/8/2023 4:30:26 PM | 1996-04-12 | 8/15/2019 1:22:06 PM | EMS | KATFRASE | Concentration | mg/L | Y | (fl. conc.) |  | Magnesium Total (fl. conc.) | 7439-95-4 | Metals - Total; | Metals - Total; EMS code: MG-T | Chemical | Total | LAB
MG-T | mg/L | Magnesium Total | X349 | FA:HNO3:ICP | Samples are field-acidified with HNO3; HNO3 digestion; analysis by ICP/AES. | 0.1 | Y | Y | Y | 6498 |  |  | 8/26/2024 2:41:32 PM | 1996-04-12 | 6/3/1998 6:51:36 PM | EMS |  | Concentration | mg/L | Y | (fl. conc.) |  | Magnesium Total (fl. conc.) | 7439-95-4 | Metals - Total; | Metals - Total; EMS code: MG-T | Chemical | Total | LAB
MG-T | ug/g wet | Magnesium Total | MTIW | MET-WET-CCMS-MID-VA | Metals in Tissue by CRC ICPMS (WET) Methodology: This method is conducted following British Columbia Lab Manual method "Metals in Animal Tissue and Vegetation (Biota) - Prescriptive". Tissue samples are homogenized and sub-sampled prior to hotblock digestion with nitric and hydrochloric acids, in combination with addition of hydrogen peroxide. Instrumental analysis is by collision cell inductively coupled plasma - mass spectrometry (modified from EPA Method 6020A). Method Limitation: This method employs a strong acid/peroxide digestion, and is intended to provide a conservative estimate of bio-available metals. Near complete recoveries are achieved for most toxicologically important metals, but elements associated with recalcitrant minerals may be only partially recovered. | 0.4 | Y | Y | Y | 24 | 11/25/2018 | 12/6/2018 | 7/21/2021 2:55:07 PM | 2018-12-24 |  | EMS |  | MassPerMass | µg/g | Y | (mass conc.) | (wet) | Magnesium Total (wet) (mass conc.) | 7439-95-4 | Metals - Total; | Metals - Total; EMS code: MG-T | Chemical | Total | LAB
MG-T | mg/L | Magnesium Total | MSEA | Metals in Seawater (undigested) by (CRC) ICP-MS | This method describes the multi-elemental determination of trace elements by Collision/Reaction Cell (CRC) ICP-MS. The method measures ions produced by a radio-frequency inductively coupled plasma. Analyte species originating in a liquid are pneumatically nebulized and the resulting aerosol transported by argon gas into the plasma torch. The ions produced are entrained in the plasma gas and extracted, by means of a differentially pumped vacuum interface, into a mass spectrometer. The ions flow through a Collision/Reaction Cell (CRC) where depending on the mode of analysis either the ions passes through the cell unhindered (no gas - classical ICP-MS mode), or collisions/reactions chemistry takes place between the ions and a neutral gas such as hydrogen or helium. Any polyatomic ions which give interference on the analytes are removed. The ions are then sorted according to their mass-to-charge ratios by a quadrupole mass spectrometer. The ions are then sorted according to their mass-to-charge ratios by a quadrupole mass spectrometer. The ions transmitted through the quadrupole are quantified by a channel electron multiplier and the ion information processed by a data handling system. Although most polyatomic interference is removed, any additional interference must be assessed and valid corrections applied or the data flagged to indicate problems. Interference correction must include compensation for background ions contributed by the plasma gas, reagents, and constituents of the sample matrix. Instrument drift as well as suppressions or enhancements of instrument response caused by the sample matrix must be corrected for by the use of internal standards. | 5 | Y | Y | Y | 441 | 8/11/2009 11:57:00 PM | 4/29/2024 | 6/7/2024 10:03:16 AM | 2009-09-08 |  | RDGIBSON |  | Concentration | mg/L | Y | (fl. conc.) |  | Magnesium Total (fl. conc.) | 7439-95-4 | Metals - Total; | Metals - Total; EMS code: MG-T | Chemical | Total | LAB
MN-D | mg/L | Manganese Dissolved | F068 | HNO3dig,TeflonTube/Autoclave,ICP-MS | This method is applicable to determiniation of dissolved, extractable, and total recoverable metals in ground waters, surface waters, drinking waters, and wastewaters. Dissolved, extractable, and total recoverable sample fractions are acidified to pH <1.5 with analytical grade nitric acid (HNO3) to prevent the metals from adsorbing to the container material or precipitating. Extractable and total recoverable water samples are collected in 250 ml acid washed polyethylene bottles and are preserved with Nitric acid (0.5 ml acid for 100 mLs of sample). All preserved samples are stored at room temperature and can be held for up to 6 months. Unfiltered, un-preserved samples are stored at 4C for no more than 30 days. For dissolved and extractable water fractions, no sample preparation is necessary. Samples are simply poured in autosampler tubes (i.e. 15 ml centrifuge tubes), and placed on the autosampler and analyzed. Samples for total recoverable metals are digested with Nitric acid in special Teflon tubes in an autoclave prior to analyses. | 5e-06 | Y | Y | Y | 1397 | 6/1/1999 1:21:00 PM | 4/17/2015 1:38:00 PM | 8/11/2021 12:04:11 PM | 1999-06-28 |  | NPEPPIN |  | Concentration | mg/L | Y | (fl. conc.) |  | Manganese Dissolved (fl. conc.) | 7439-96-5 | Metals - Dissolved; | Metals - Dissolved; EMS code: MN-D | Chemical | Dissolved | LAB
MN-D | mg/L | Manganese Dissolved | X203 | FF:FA,HNO3,Dir. Flame,AA | Samples are field-filtered (0.45 um), and field-acidified (HNO3); analysis by AA. | 0.01 | Y | Y | Y | 69 |  |  | 10/24/1996 3:26:17 AM | 1996-04-12 | 6/3/1998 6:51:37 PM | EMS |  | Concentration | mg/L | Y | (fl. conc.) |  | Manganese Dissolved (fl. conc.) | 7439-96-5 | Metals - Dissolved; | Metals - Dissolved; EMS code: MN-D | Chemical | Dissolved | LAB
MN-D | mg/L | Manganese Dissolved | F032 | Filt.,Acid,ICPMS | There is no acid digestion required for this analysis. A water sample is collected and preserved in 0.2 percent nitric acid. To analyse the extractable metals the sample is shaken, allowed to sit for at least twenty four hours, and a decanted portion is introduced to the spectrometer. When the dissolved metals are required, the samples must first be filtered through a 0.45 micron cellulose acetate filter paper and then followed by preservation in 0.2 percent nitric acid. It is important to note that different laboratories may have very different detection limits for some variables analyzed by ICPMS. | 0.0001 | Y | Y | Y | 82799 | 3/10/1998 2:07:00 PM | 8/26/2024 | 8/26/2024 4:53:05 PM | 1997-10-20 | 8/15/2019 10:56:50 AM | NPEPPIN | KATFRASE | Concentration | mg/L | Y | (fl. conc.) |  | Manganese Dissolved (fl. conc.) | 7439-96-5 | Metals - Dissolved; | Metals - Dissolved; EMS code: MN-D | Chemical | Dissolved | LAB
MN-D | mg/L | Manganese Dissolved | X170 | ICP | This code is a catch-all for a large variety of ICP methods including water and also solids digested into solutions as an intermediate result. Different procedures will have greatly different detection limits. (a) For air particulates, preliminary results are reported in mg/L and loading results as mg/dm2/d. (b) For Hi Vol samples preliminiary results are same as reported in mg/L and loading results as ug/m3. (c) Note that for Rain samples there is no filtration nor digestion steps. | 0.01 | Y | Y | Y | 12756 | 4/11/1996 6:02:00 PM | 11/6/2018 | 6/12/2020 1:21:12 PM | 1996-04-12 | 6/3/1998 6:51:37 PM | EMS |  | Concentration | mg/L | Y | (fl. conc.) |  | Manganese Dissolved (fl. conc.) | 7439-96-5 | Metals - Dissolved; | Metals - Dissolved; EMS code: MN-D | Chemical | Dissolved | LAB
MN-D | mg/L | Manganese Dissolved | METC | MET-D-NP-CCMS-VA | Ultra trace metals in water are analyzed by CRC ICPMS, based on US EPA Method 6020B (July 2014). This procedure is intended for pristine field-filtered acid-preserved water samples. The detection limits (LOR) for this test are based on lab instrumental analysis only, not including filtration. ALS-supplied field filtration equipment does not support these LOR. Therefore, because of the high probability of false positives due to filtration, it is strongly recommended that a filtration blank be analysed to aid in data interpretation. | 5e-05 | Y | Y | Y | 1776 | 4/3/2012 | 8/15/2024 | 8/23/2024 9:50:35 AM | 2019-08-20 |  | KATFRASE |  | Concentration | mg/L | Y | (fl. conc.) |  | Manganese Dissolved (fl. conc.) | 7439-96-5 | Metals - Dissolved; | Metals - Dissolved; EMS code: MN-D | Chemical | Dissolved | LAB
MN-D | mg/L | Manganese Dissolved | 0050 | Graphite Furnace AA | Graphite Furnace AA | 0.001 | Y | Y | Y | 131 |  |  | 10/24/1996 12:33:08 PM | 1996-04-12 | 6/3/1998 6:51:36 PM | EMS |  | Concentration | mg/L | Y | (fl. conc.) |  | Manganese Dissolved (fl. conc.) | 7439-96-5 | Metals - Dissolved; | Metals - Dissolved; EMS code: MN-D | Chemical | Dissolved | LAB
MN-D | mg/L | Manganese Dissolved | MWID | Metals in Water by ICPMS (digested) | Metals in Water by ICPMS (digested) Based on EPA 200.8 | 0.001 | Y | Y | Y | 31785 | 10/26/2018 5:20:00 PM | 10/26/2018 5:20:00 PM | 7/31/2024 1:46:41 PM | 2009-12-24 |  | RDGIBSON |  | Concentration | mg/L | Y | (fl. conc.) |  | Manganese Dissolved (fl. conc.) | 7439-96-5 | Metals - Dissolved; | Metals - Dissolved; EMS code: MN-D | Chemical | Dissolved | LAB
MN-D | mg/L | Manganese Dissolved | F039 | Field filt + acid, ICP-AES | Dissolved and extractable water samples are ready for analysis once collected. Total samples are autoclaved with HNO3 and hydrochloric acid (HCI) in capped Teflon vials. Analysis is performed via Inductively Coupled Argon Plasma Atomic Emission Spectrometry (ICP-AES)(simultaneous multielement analysis), Graphite Furnace Atomic Absorption Spectrometry (GF-AAS)(low level analysis) and/or Flame Atomic Absorption Spectrometry (F-AAS). | 0.001 | Y | Y | Y | 6602 | 7/29/1996 6:03:00 PM | 7/5/2019 1:21:00 PM | 5/16/2022 4:21:33 PM | 1997-05-08 | 6/3/1998 6:51:46 PM | MPAK |  | Concentration | mg/L | Y | (fl. conc.) |  | Manganese Dissolved (fl. conc.) | 7439-96-5 | Metals - Dissolved; | Metals - Dissolved; EMS code: MN-D | Chemical | Dissolved | LAB
MN-D | mg/L | Manganese Dissolved | EQ01 | EQUIS (mg/L) | EQUIS (mg/L) | 0.001 | N | Y | Y | 4234 |  |  | 10/24/1996 3:16:55 PM | 1996-04-12 | 6/3/1998 6:51:37 PM | EMS |  | Concentration | mg/L | Y | (fl. conc.) |  | Manganese Dissolved (fl. conc.) | 7439-96-5 | Metals - Dissolved; | Metals - Dissolved; EMS code: MN-D | Chemical | Dissolved | LAB
MN-D | mg/L | Manganese Dissolved | 0040 | HNO3 Dig: ICP Analysis | HNO3 Dig: ICP Analysis | 0.01 | Y | Y | Y | 8 |  |  | 4/12/2000 3:59:24 PM | 1996-04-12 | 6/3/1998 6:51:36 PM | EMS |  | Concentration | mg/L | Y | (fl. conc.) |  | Manganese Dissolved (fl. conc.) | 7439-96-5 | Metals - Dissolved; | Metals - Dissolved; EMS code: MN-D | Chemical | Dissolved | LAB
MN-D | mg/L | Manganese Dissolved | X113 | HNO3/HCl/Perchlor,AA | HNO3/HCl/Perchlor;AA | 0.001 | Y | Y | Y | 147 |  |  | 10/23/1996 11:07:32 PM | 1996-04-12 | 6/3/1998 6:51:37 PM | EMS |  | Concentration | mg/L | Y | (fl. conc.) |  | Manganese Dissolved (fl. conc.) | 7439-96-5 | Metals - Dissolved; | Metals - Dissolved; EMS code: MN-D | Chemical | Dissolved | LAB
MN-D | mg/L | Manganese Dissolved | F049 | Filt,HNO3 dig,conc'n,nebul'ICP-OES | This method provides an automated procedure for the Inductively Coupled Argon Plasma Optical Emission Spectrophotometric (ICP-OES) determination and quantitation of trace amounts of metals in surface, ground and precipitation waters after a manual acid digestion and concentration procedure and ultrasonic nebulisation. Organically-bound metals and metal complexes are converted to metal ions by digestion with nitric acid. To increase the signal detected by the ICP, the samples are concentrated 10 times. The liquid sample is pumped onto the piezoelectric transducer, where the ultrasonic action efficiently converts it to a fine, dense aerosol. The nebulizer carrier gas flow then transports the wet aerosol through a heated U-tube, where the solvent is vaporized. Subsequently, most of the solvent vapour is removed from the aerosol stream by passing the stream through a chilled condenser. The result is a dry, analyte-laden aerosol which is introduced to the plasma. There, the high temperature desolvates the aerosol and dissociates the sample material into individual atoms and ions, which are excited to emit light at wave lengths characteristic of the elements present. Dissolved samples are field filtered through 0.45 micron cellulose acetate filter prior to acid preservation. Samples for extractable analysis are acidified with 2 mL concentrated HNO3 per litre and allowed to stand a minimum of 24 hours before analysis. Clients providing these samples are advised to preserve samples at the collection site. | 0.0001 | Y | Y | Y | 2 |  |  | 2/15/2016 10:44:41 AM | 1997-10-20 | 6/3/1998 6:51:47 PM | NPEPPIN |  | Concentration | mg/L | Y | (fl. conc.) |  | Manganese Dissolved (fl. conc.) | 7439-96-5 | Metals - Dissolved; | Metals - Dissolved; EMS code: MN-D | Chemical | Dissolved | LAB
MN-D | mg/L | Manganese Dissolved | X285 | HNO3 Dig:ICP Analysis | HNO3 Dig:ICP Analysis | 0.005 | Y | Y | Y | 7662 | 3/1/2000 | 9/16/2003 3:18:00 PM | 8/26/2024 2:41:32 PM | 1996-04-12 | 6/3/1998 6:51:41 PM | EMS |  | Concentration | mg/L | Y | (fl. conc.) |  | Manganese Dissolved (fl. conc.) | 7439-96-5 | Metals - Dissolved; | Metals - Dissolved; EMS code: MN-D | Chemical | Dissolved | LAB
MN-D | mg/L | Manganese Dissolved | X116 | FF,FA,HNO3,Graph.Furn,AA | Samples are field-filtered (0.45 um), field-acidified with HNO3; analysis by GFAA. | 0.001 | Y | Y | Y | 127 |  |  | 10/24/1996 12:57:31 PM | 1996-04-12 | 6/3/1998 6:51:41 PM | EMS |  | Concentration | mg/L | Y | (fl. conc.) |  | Manganese Dissolved (fl. conc.) | 7439-96-5 | Metals - Dissolved; | Metals - Dissolved; EMS code: MN-D | Chemical | Dissolved | LAB
MN-D | mg/L | Manganese Dissolved | X085 | LF,LA,HNO3,AA | Samples are lab-filtered (0.45 um), and lab-acidified (HNO3); analysis by AA. | 0.001 | Y | Y | Y | 216 | 12/12/1996 | 9/18/2015 | 10/28/2015 9:54:16 AM | 1996-04-12 | 6/3/1998 6:51:37 PM | EMS |  | Concentration | mg/L | Y | (fl. conc.) |  | Manganese Dissolved (fl. conc.) | 7439-96-5 | Metals - Dissolved; | Metals - Dissolved; EMS code: MN-D | Chemical | Dissolved | LAB
MN-D | mg/L | Manganese Dissolved | MSEA | Metals in Seawater (undigested) by (CRC) ICP-MS | This method describes the multi-elemental determination of trace elements by Collision/Reaction Cell (CRC) ICP-MS. The method measures ions produced by a radio-frequency inductively coupled plasma. Analyte species originating in a liquid are pneumatically nebulized and the resulting aerosol transported by argon gas into the plasma torch. The ions produced are entrained in the plasma gas and extracted, by means of a differentially pumped vacuum interface, into a mass spectrometer. The ions flow through a Collision/Reaction Cell (CRC) where depending on the mode of analysis either the ions passes through the cell unhindered (no gas - classical ICP-MS mode), or collisions/reactions chemistry takes place between the ions and a neutral gas such as hydrogen or helium. Any polyatomic ions which give interference on the analytes are removed. The ions are then sorted according to their mass-to-charge ratios by a quadrupole mass spectrometer. The ions are then sorted according to their mass-to-charge ratios by a quadrupole mass spectrometer. The ions transmitted through the quadrupole are quantified by a channel electron multiplier and the ion information processed by a data handling system. Although most polyatomic interference is removed, any additional interference must be assessed and valid corrections applied or the data flagged to indicate problems. Interference correction must include compensation for background ions contributed by the plasma gas, reagents, and constituents of the sample matrix. Instrument drift as well as suppressions or enhancements of instrument response caused by the sample matrix must be corrected for by the use of internal standards. | 0.0003 | Y | Y | Y | 4 | 4/12/2023 | 4/29/2024 | 5/16/2024 2:00:14 PM | 2009-09-08 |  | RDGIBSON |  | Concentration | mg/L | Y | (fl. conc.) |  | Manganese Dissolved (fl. conc.) | 7439-96-5 | Metals - Dissolved; | Metals - Dissolved; EMS code: MN-D | Chemical | Dissolved | LAB
MN-D | mg/L | Manganese Dissolved | X562 | Dissolved metals: FF, FA, HNO3, HR-ICPMS | Ultra trace metals in water are analyzed by high resolution inductively coupled plasma mass spectrometry (HR-ICPMS), based on US EPA Method 200.8, (Oct. 1999). This procedure is intended for pristine field-filtered acid-preserved water samples. The detection limits (LOR) for this test are based on lab instrumental analysis only, not including filtration. ALS-supplied field filtration equipment does not support these LOR. Therefore, because of the high probability of false positives due to filtration, it is strongly recommended that a filtration blank be analyzed to aid in data interpretation. | 5e-05 | Y | Y | Y | 1264 | 6/8/2015 | 10/19/2023 | 11/9/2023 7:41:27 AM | 2015-10-22 |  | RDGIBSON |  | Concentration | mg/L | Y | (fl. conc.) |  | Manganese Dissolved (fl. conc.) | 7439-96-5 | Metals - Dissolved; | Metals - Dissolved; EMS code: MN-D | Chemical | Dissolved | LAB
MN-D | mg/L | Manganese Dissolved | X350 | FF,FA:HNO3:ICP | Samples are field-filtered (0.45 um); field-acidified with HNO3; analysis by ICP/AES. | 0.005 | Y | Y | Y | 160 |  |  | 10/24/1996 2:53:40 PM | 1996-04-12 | 6/3/1998 6:51:37 PM | EMS |  | Concentration | mg/L | Y | (fl. conc.) |  | Manganese Dissolved (fl. conc.) | 7439-96-5 | Metals - Dissolved; | Metals - Dissolved; EMS code: MN-D | Chemical | Dissolved | LAB
MN-D | mg/L | Manganese Dissolved | X556 | Metals in water, dissolved, "in-bottle" digestion, ICP-SFMS | The whole water sample is filtered through a 0.45 micron cellulose acetate filter paper and the filtrate is stored in a HDPE/Teflon bottle. The container is de-capped in a Class 100 workstation, acidified to 2% (v/v) HNO3 and 1% (v/v) HCl and recapped. The sample is "in-bottle" digested in a convection oven, set at 60 degrees C for 16 hours. The sample is cooled and introduced to an inductively coupled argon plasma-sector field mass spectrometer (ICP-SFMS) by pneumatic nebulization. Each respective element is measured at a specific mass to charge ratio, m/z value expressed in amu (atomic mass units), and at a specified intrument resolution. LR: Low Resolution ~300, MR: Mid Resolution ~4000 and HR: High Resolution ~10000. Ref: Method number 02-2003, National Laboratory for Environmental Testing, CCIW, Environment Canada, Burlington, Requested April 2003. | 5e-05 | Y | Y | Y | 4234 | 6/11/2003 | 12/6/2023 2:04:00 AM | 2/14/2024 9:05:32 AM | 2004-01-19 |  | UMULLER |  | Concentration | mg/L | Y | (fl. conc.) |  | Manganese Dissolved (fl. conc.) | 7439-96-5 | Metals - Dissolved; | Metals - Dissolved; EMS code: MN-D | Chemical | Dissolved | LAB
MN-E | ug/g | Manganese Extractble | X230 | Hom,HNO3/HCl,ICP | The sample is dried, ground to pass 100 mesh, and digested with a nitric-hydrochloric acid mixture to solubilize the solid matter and to remove the organic material by oxidation and volatilization. The completed digests are submitted to metals instrumentation personnel for metals analysis by Inductively Coupled Atomic Emission Spectrophotometry (ICP-AES), or hydride analysis by atomic absorption. | 0.25 | Y | Y | Y | 70 |  |  | 10/24/1996 2:25:28 PM | 1996-04-12 | 6/3/1998 6:51:41 PM | EMS |  | MassPerMass | mg/kg | Y | (mass conc.) |  | Manganese Extractble (mass conc.) | 7439-96-5 | Metals - Extractable; | Metals - Extractable; EMS code: MN-E | Chemical |  | LAB
MN-E | mg/L | Manganese Extractble | F047 | 24hr HNO3 dil'n,ICP-MS | There is no acid digestion required for this analysis. A water sample is collected and preserved in 0.2 percent nitric acid. To analyse the extractable metals the sample is shaken, allowed to sit for at least twenty four hours, and a decanted portion is introduced to the spectrometer. When the dissolved metals are required, the samples must first be filtered through a 0.45 micron cellulose acetate filter paper and then followed by preservation in 0.2 percent nitric acid. It is important to note that different laboratories may have very different detection limits for some variables analyzed by ICP/MS. | 5e-06 | Y | Y | Y | 2676 | 4/16/1998 10:21:00 AM | 6/26/2003 | 9/11/2019 8:46:18 AM | 1997-10-20 | 6/3/1998 6:51:47 PM | NPEPPIN |  | Concentration | mg/L | Y | (fl. conc.) |  | Manganese Extractble (fl. conc.) | 7439-96-5 | Metals - Extractable; | Metals - Extractable; EMS code: MN-E | Chemical |  | LAB
MN-E | mg/L | Manganese Extractble | F040 | Field Acid, lab filt, ICP-AES | Dissolved and extractable water samples are ready for analysis once collected. Total samples are autoclaved with HNO3 and hydrochloric acid (HCI) in capped Teflon vials. Analysis is performed via Inductively Coupled Argon Plasma Atomic Emission Spectrometry (ICP-AES)(simultaneous multielement analysis), Graphite Furnace Atomic Absorption Spectrometry (GF-AAS)(low level analysis) and/or Flame Atomic Absorption Spectrometry (F-AAS). | 0.001 | Y | Y | Y | 2360 | 8/5/1997 6:06:00 PM | 3/23/2012 3:11:00 AM | 9/11/2019 8:47:41 AM | 1997-05-08 | 6/3/1998 6:51:46 PM | MPAK |  | Concentration | mg/L | Y | (fl. conc.) |  | Manganese Extractble (fl. conc.) | 7439-96-5 | Metals - Extractable; | Metals - Extractable; EMS code: MN-E | Chemical |  | LAB
MN-E | mg/L | Manganese Extractble | X513 | Acid Dig,Dir.Asp.,AA | Overnight acid digestion; Direct Aspiration; AA | 0.01 | Y | Y | Y | 876 |  |  | 8/23/1999 3:43:46 PM | 1999-04-22 |  | NPEPPIN |  | Concentration | mg/L | Y | (fl. conc.) |  | Manganese Extractble (fl. conc.) | 7439-96-5 | Metals - Extractable; | Metals - Extractable; EMS code: MN-E | Chemical |  | LAB
MN-E | mg/L | Manganese Extractble | F050 | 24hr acid,filt,acid dig,conc'n,nebul.,ICP-OES | This method provides an automated procedure for the Inductively Coupled Argon Plasma Optical Emission Spectrophotometric (ICP-OES) determination and quantitation of trace amounts of metals in surface, ground and precipitation waters after a manual acid digestion and concentration procedure and ultrasonic nebulisation. Organically-bound metals and metal complexes are converted to metal ions by digestion with nitric acid. To increase the signal detected by the ICP, the samples are concentrated 10 times. The liquid sample is pumped onto the piezoelectric transducer, where the ultrasonic action efficiently converts it to a fine, dense aerosol. The nebulizer carrier gas flow then transports the wet aerosol through a heated U-tube, where the solvent is vaporized. Subsequently, most of the solvent vapour is removed from the aerosol stream by passing the stream through a chilled condenser. The result is a dry, analyte-laden aerosol which is introduced to the plasma. There, the high temperature desolvates the aerosol and dissociates the sample material into individual atoms and ions, which are excited to emit light at wave lengths characteristic of the elements present. Dissolved samples are field filtered through 0.45 micron cellulose acetate filter prior to acid preservation. Samples for extractable analysis are acidified with 2 mL concentrated HNO3 per litre and allowed to stand a minimum of 24 hours before analysis. Clients providing these samples are advised to preserve samples at the collection site. | 0.0001 | Y | Y | Y | 4 | 12/10/2011 12:01:00 PM | 12/10/2011 12:01:00 PM | 2/9/2012 12:15:57 PM | 1997-10-20 | 6/3/1998 6:51:47 PM | NPEPPIN |  | Concentration | mg/L | Y | (fl. conc.) |  | Manganese Extractble (fl. conc.) | 7439-96-5 | Metals - Extractable; | Metals - Extractable; EMS code: MN-E | Chemical |  | LAB
MN-E | mg/L | Manganese Extractble | X553 | Metals in water, extractable, ICP-SFMS | The whole water sample is collected in a HDPE/Teflon bottle. In a Class 100 workstation it is preserved with concentrated Nitric Acid at concentration of 0.2 percent (v/v). The acid preservative extracts the metals from material in the sample and a decanted portion is introduced directly to an inductively coupled argon plasma-sector field mass spectrometer (ICP-SFMS) by pneumatic nebulization. Each respective element is measured at a specific mass to charge ratio, m/z value expressed in amu (atomic mass units), and at a specified instrument resolution. LR: Low Resolution ~300, MR: Mid Resolution ~4000 and HR: High Resolution ~10000. Ref: Method number 02-2061, National Laboratory for Environmental Testing, CCIW, Environment Canada, Burlington. Requested April 2003. | 5e-05 | Y | Y | Y | 2584 | 5/22/2003 | 9/21/2007 | 4/18/2018 9:28:02 AM | 2004-01-19 |  | UMULLER |  | Concentration | mg/L | Y | (fl. conc.) |  | Manganese Extractble (fl. conc.) | 7439-96-5 | Metals - Extractable; | Metals - Extractable; EMS code: MN-E | Chemical |  | LAB
MN-L | mg/L | Manganese Leachable | MEND | MEND (mod) EPA 300.1 (mod) | MEND (mod) EPA 300.1 (mod). Inorganic anions are analyzed by obtaining an extract from a soil sample that has been added in a 3:1 ratio of deionized water to soil; then leached and the extract filtered prior to analysis; which is then analyzed by Ion Chromatography with conductivity and/or UV detection. | 0.0005 | Y | Y | Y | 0 |  |  |  | 2024-01-29 |  | RDGIBSON |  | Concentration | mg/L | Y | (fl. conc.) |  | Manganese Leachable (fl. conc.) | 7439-96-5 | Metals - Extractable; | Metals - Extractable; EMS code: MN-L | Chemical |  | LAB
MN-O | % | Manganese (as MnO) | ICP06 | ALS Minerals ME-ICP06 | ALS Minerals ME-ICP06. A prepared sample is fused with a lithium tetraborate/lithium metaborate flux and poured into a graphite crucible and fused in a furnace at 1000°C. The resulting melt is then cooled and dissolved in 100 mL of 4% HNO3/2% HCl followed by ICP-AES. | 0.01 | Y | Y | Y | 0 |  |  |  | 2024-01-29 |  | EMS |  | DimensionlessRatio | % | N | (ratio) |  | Manganese (as MnO) (ratio) |  | Metals; | Metals; EMS code: MN-O | Chemical |  | LAB
MN-T | mg/L | Manganese Total | F033 | Acid Dig,Centrifuge,ICPMS | For the determination of total recoverable elements in water or waste water, take 100 ml aliquot from a well mixed, acid preserved, sample containing not more than 0.25% (w/v) total solids and transfer to a 250 ml Griffin beaker. (If total solids are greater than 0.25% reduce the size of the aliquot by a proportionate amount). Add 1 ml of conc. nitric acid and 0.5 ml conc. hydrochloric acid. Heat on a hot plate until the volume has been reduced to 15-20ml, ensuring that the sample does not boil. Allow to cool and transfer to a 50 ml volumetric flask. Dilute to volume with ASTM type I water. Centrifuge the sample or allow to stand overnight to separate insoluble material. Prior analysis, pipette 20 ml into a 50 ml volumetric flask and dilute to volume with ASTM type I water. Analysis should be performed within seventy two hours of the completed preparation. Note that different laboratories may have very different field preservation, digestion procedures and detection limits for this method, so these should be documented. | 0.001 | Y | Y | Y | 40414 | 11/20/1997 9:31:00 AM | 8/22/2024 3:24:00 AM | 8/25/2024 9:47:36 PM | 1998-05-14 | 6/3/1998 6:51:49 PM | NPEPPIN |  | Concentration | mg/L | Y | (fl. conc.) |  | Manganese Total (fl. conc.) | 7439-96-5 | Metals - Total; | Metals - Total; EMS code: MN-T | Chemical | Total | LAB
MN-T | mg/L | Manganese Total | F121 | Acid,ICPMS | Acidified, no digestion, ICPMS | 5e-05 | Y | Y | Y | 8436 | 3/12/2010 7:57:00 PM | 8/15/2024 9:10:00 PM | 8/26/2024 8:06:03 AM | 2002-07-22 | 7/16/2012 3:29:16 PM | UMULLER | RDGIBSON | Concentration | mg/L | Y | (fl. conc.) |  | Manganese Total (fl. conc.) | 7439-96-5 | Metals - Total; | Metals - Total; EMS code: MN-T | Chemical | Total | LAB
MN-T | ug/g | Manganese Total | 0192 | Hom, Nitric/Perchlor,ICP | Sample already in a dry state so does not require drying before homogenization and nitric perchlorid digestion followed by ICP analysis. | 1 | Y | Y | Y | 708 |  |  | 10/24/1996 12:50:25 PM | 1996-04-12 | 6/3/1998 6:51:37 PM | EMS |  | MassPerMass | mg/kg | Y | (mass conc.) |  | Manganese Total (mass conc.) | 7439-96-5 | Metals - Total; | Metals - Total; EMS code: MN-T | Chemical | Total | LAB
MN-T | mg/L | Manganese Total | X285 | HNO3 Dig:ICP Analysis | HNO3 Dig:ICP Analysis | 0.005 | Y | Y | Y | 681 | 8/10/1993 | 10/5/1993 | 4/30/2004 1:59:45 PM | 1998-06-22 |  | NPEPPIN |  | Concentration | mg/L | Y | (fl. conc.) |  | Manganese Total (fl. conc.) | 7439-96-5 | Metals - Total; | Metals - Total; EMS code: MN-T | Chemical | Total | LAB
MN-T | ug/g | Manganese Total | 0190 | D&G:N/P Dig:ICP Anal. | Sample is dried and grind-mixed before nitric perchloric digestion followed by ICP analysis. | 1 | Y | Y | Y | 949 |  |  | 10/24/1996 3:06:16 PM | 1996-04-12 | 6/3/1998 6:51:37 PM | EMS |  | MassPerMass | mg/kg | Y | (mass conc.) |  | Manganese Total (mass conc.) | 7439-96-5 | Metals - Total; | Metals - Total; EMS code: MN-T | Chemical | Total | LAB
MN-T | ug/g wet | Manganese Total | X217 | Hom,Nitric/Perchl,ICP | Sample is wet but not dried before homogenization and nitric perchloric digestion followed by ICP analysis. | 0.05 | Y | Y | Y | 74 |  |  | 10/24/1996 10:01:27 AM | 1996-04-12 | 6/3/1998 6:51:42 PM | EMS |  | MassPerMass | µg/g | Y | (mass conc.) | (wet) | Manganese Total (wet) (mass conc.) | 7439-96-5 | Metals - Total; | Metals - Total; EMS code: MN-T | Chemical | Total | LAB
MN-T | mg/L | Manganese Total | 0042 | GI HNO3 Dig: ICP | Sample aliquot taken from unfiltered unpreserved General Ions bottle; HNO3 digestion; ICP analysis. | 0.002 | Y | Y | Y | 6273 | 9/20/1996 | 7/5/2019 3:21:00 PM | 7/12/2019 6:49:13 PM | 1996-04-12 | 6/3/1998 6:51:37 PM | EMS |  | Concentration | mg/L | Y | (fl. conc.) |  | Manganese Total (fl. conc.) | 7439-96-5 | Metals - Total; | Metals - Total; EMS code: MN-T | Chemical | Total | LAB
MN-T | mg/L | Manganese Total | 0020 | HNO3 Dig:AA Analysis | HNO3 Dig:AA Analysis | 0.02 | Y | Y | Y | 3070 |  |  | 8/26/2004 4:21:03 PM | 1996-04-12 | 6/3/1998 6:51:37 PM | EMS |  | Concentration | mg/L | Y | (fl. conc.) |  | Manganese Total (fl. conc.) | 7439-96-5 | Metals - Total; | Metals - Total; EMS code: MN-T | Chemical | Total | LAB
MN-T | ug/g wet | Manganese Total | X560 | Metals in Biota (wet), total, acid digestion, ICP-MS | Five grams of wet, homogenized biota is weighed in a centrifuge tube, and 2.5 mL of Nitric Acid and 2.5 mL of Hydrochloric acid are added. The sample is digested at room temperature for one hour and then digested at 90-95 degrees C in a water bath or heat block for one more hour. The sample is cooled and the volume is made to 50 mL with DI. The digestate is then diluted at ten fold and introduced to a collision/reaction cell inductively coupled argon plasma mass spectrometer (CRC ICPMS) by pneumatic nebulization. Each respective element is measured at a specific mass to charge ratio, m/z value expressed in amu (atomic mass units). - Maxxam Analytics (BC). Added to EMS Dec 10, 2008. | 0.1 | Y | Y | Y | 541 | 11/15/2008 4:00:00 AM | 10/30/2023 | 11/28/2023 2:21:33 PM | 2008-12-11 | 3/11/2015 4:17:10 PM | RDGIBSON | EMS | MassPerMass | µg/g | Y | (mass conc.) | (wet) | Manganese Total (wet) (mass conc.) | 7439-96-5 | Metals - Total; | Metals - Total; EMS code: MN-T | Chemical | Total | LAB
MN-T | ug/g | Manganese Total | MTID | MET-DRY-CCMS-VA | Metals in Tissue by CRC ICPMS (DRY) Methodology: This method is conducted following British Columbia Lab Manual method "Metals in Animal Tissue and Vegetation (Biota) - Prescriptive". Tissue samples are homogenized and sub-sampled prior to hotblock digestion with nitric and hydrochloric acids, in combination with addition of hydrogen peroxide. Instrumental analysis is by collision cell inductively coupled plasma - mass spectrometry (modified from EPA Method 6020A). Method Limitation: This method employs a strong acid/peroxide digestion, and is intended to provide a conservative estimate of bio-available metals. Near complete recoveries are achieved for most toxicologically important metals, but elements associated with recalcitrant minerals may be only partially recovered. | 0.05 | Y | Y | Y | 122 | 8/13/2014 | 9/13/2022 | 10/26/2022 7:14:15 AM | 2015-03-09 | 3/11/2015 4:12:05 PM | RDGIBSON | EMS | MassPerMass | mg/kg | Y | (mass conc.) |  | Manganese Total (mass conc.) | 7439-96-5 | Metals - Total; | Metals - Total; EMS code: MN-T | Chemical | Total | LAB
MN-T | ug/g wet | Manganese Total | MTIW | MET-WET-CCMS-MID-VA | Metals in Tissue by CRC ICPMS (WET) Methodology: This method is conducted following British Columbia Lab Manual method "Metals in Animal Tissue and Vegetation (Biota) - Prescriptive". Tissue samples are homogenized and sub-sampled prior to hotblock digestion with nitric and hydrochloric acids, in combination with addition of hydrogen peroxide. Instrumental analysis is by collision cell inductively coupled plasma - mass spectrometry (modified from EPA Method 6020A). Method Limitation: This method employs a strong acid/peroxide digestion, and is intended to provide a conservative estimate of bio-available metals. Near complete recoveries are achieved for most toxicologically important metals, but elements associated with recalcitrant minerals may be only partially recovered. | 0.01 | Y | Y | Y | 24 | 11/25/2018 | 12/6/2018 | 7/21/2021 2:55:07 PM | 2018-12-24 |  | EMS |  | MassPerMass | µg/g | Y | (mass conc.) | (wet) | Manganese Total (wet) (mass conc.) | 7439-96-5 | Metals - Total; | Metals - Total; EMS code: MN-T | Chemical | Total | LAB
MN-T | ug/g | Manganese Total | FS89 | Homo,HNO3 dig,ICP-AES | Homo;HNO3 dig;ICP-AES | 0.2 | Y | Y | Y | 11 | 1/5/2001 | 10/31/2022 | 11/22/2022 2:40:39 PM | 2001-03-08 |  | NPEPPIN |  | MassPerMass | mg/kg | Y | (mass conc.) |  | Manganese Total (mass conc.) | 7439-96-5 | Metals - Total; | Metals - Total; EMS code: MN-T | Chemical | Total | LAB
MN-T | ug/g | Manganese Total | EQ02 | EQUIS (ug/g) | EQUIS (ug/g) | 1 | N | Y | Y | 800 |  |  | 10/23/1996 12:08:28 AM | 1996-04-12 | 6/3/1998 6:51:37 PM | EMS |  | MassPerMass | mg/kg | Y | (mass conc.) |  | Manganese Total (mass conc.) | 7439-96-5 | Metals - Total; | Metals - Total; EMS code: MN-T | Chemical | Total | LAB
MN-T | mg/L | Manganese Total | 1349 | Tot Metals in Seawater by CRC ICPMS (BC) | Tot Metals in Seawater by CRC ICPMS (BC) | 0.0002 | Y | Y | Y | 98 | 8/16/2019 | 12/5/2021 | 12/13/2021 12:34:16 PM | 2019-09-30 |  | KATFRASE |  | Concentration | mg/L | Y | (fl. conc.) |  | Manganese Total (fl. conc.) | 7439-96-5 | Metals - Total; | Metals - Total; EMS code: MN-T | Chemical | Total | LAB
MN-T | mg/L | Manganese Total | X072 | FA,HNO3,Graph.Furnace,AA | Samples are field-acidified with HNO3; HNO3 digestion; analysis by GFAA. | 0.0002 | Y | Y | Y | 1 |  |  | 5/11/1998 3:54:38 PM | 1996-04-12 | 6/3/1998 6:51:41 PM | EMS |  | Concentration | mg/L | Y | (fl. conc.) |  | Manganese Total (fl. conc.) | 7439-96-5 | Metals - Total; | Metals - Total; EMS code: MN-T | Chemical | Total | LAB
MN-T | mg/L | Manganese Total | X349 | FA:HNO3:ICP | Samples are field-acidified with HNO3; HNO3 digestion; analysis by ICP/AES. | 0.005 | Y | Y | Y | 6720 |  |  | 8/26/2024 2:41:32 PM | 1996-04-12 | 6/3/1998 6:51:37 PM | EMS |  | Concentration | mg/L | Y | (fl. conc.) |  | Manganese Total (fl. conc.) | 7439-96-5 | Metals - Total; | Metals - Total; EMS code: MN-T | Chemical | Total | LAB
MN-T | ug/g | Manganese Total | X230 | Hom,HNO3/HCl,ICP | The sample is dried, ground to pass 100 mesh, and digested with a nitric-hydrochloric acid mixture to solubilize the solid matter and to remove the organic material by oxidation and volatilization. The completed digests are submitted to metals instrumentation personnel for metals analysis by Inductively Coupled Atomic Emission Spectrophotometry (ICP-AES), or hydride analysis by atomic absorption. | 0.5 | Y | Y | Y | 259 | 2/12/2002 | 12/30/2005 | 2/3/2006 1:13:13 AM | 1996-04-12 | 6/3/1998 6:51:37 PM | EMS |  | MassPerMass | mg/kg | Y | (mass conc.) |  | Manganese Total (mass conc.) | 7439-96-5 | Metals - Total; | Metals - Total; EMS code: MN-T | Chemical | Total | LAB
MN-T | ug/g wet | Manganese Total | X552 | Metals in Biota (wet), total, acid digestion, ICP-SFMS | One gram of wet, homogenized biota is weighed in a microwave TFM vessel, and 8mL of Nitric Acid and 1mL of Hydrogen Peroxide are added. The sample is digested at 200 degrees C for 15 minutes in a closed vessel high pressure microwave oven. The sample is cooled, the volume adjusted to 100mL and introduced to an inductively coupled argon plasma-sector field mass spectrometer (ICP-SFMS) by pneumatic nebulization. Each respective element is measured at a specific mass to charge ratio, m/z value expressed in amu (atomic mass units), and at a specified instrument resolution. LR: Low Resolution ~300, MR: Mid Resolution ~4000. Ref: 02-2704, National Laboratory for Environmental Testing, CCIW, Environment Canada, Burlington. Requested April 2003. | 0.005 | Y | Y | Y | 37 | 8/9/2016 | 10/16/2023 | 11/28/2023 2:21:32 PM | 2004-01-19 | 4/19/2021 1:22:00 PM | UMULLER | EMS | MassPerMass | µg/g | Y | (mass conc.) | (wet) | Manganese Total (wet) (mass conc.) | 7439-96-5 | Metals - Total; | Metals - Total; EMS code: MN-T | Chemical | Total | LAB
MN-T | ug/g | Manganese Total | FS01 | D&G,Nitric/HCl:ICP | Samples are dried at 60 C, then either sieved to -100mesh (150 um) and rolled to homogenize, or finely ground with a mortar and pestle or a shatter box laboratory mill and rolled (depending on the submitter's requirements). The samples are then weighed into Teflon digestion liners and digested with HNO3 and HCl in a microwave oven. Samples are cooled, volumized, and allowed to settle overnight before being decanted. Analysis is then performed via Inductively Coupled Argon Plasma Atomic Emission Spectrometry (ICP-AES) (Simultaneous multi element analysis). In ICP Spectrometry, the liquid sample is pumped to the nebulizer, where it is convertd into an aerosol. the nebulizer carrier gas flow than transports the wet aerosol past the impact bead, where the large droplets condense out and flow to the waste container. The remaining fine aerosol is carried into the plasma. There, the high temperature desolvates the aerosol and dissociates the sample material into individual atoms and ions, which are excited to emit light at wavelengths characteristic of the elements present. A diffraction grating in the spectrometer diffracts the various wavelengths of light through exit slits to the photo-multiplier tubes, which convert the light energy to electronic signals. | 0.2 | Y | Y | Y | 2646 | 8/30/1996 1:52:00 PM | 12/1/2010 | 12/9/2010 3:55:27 PM | 1996-11-20 | 6/3/1998 6:51:46 PM | MPAK |  | MassPerMass | mg/kg | Y | (mass conc.) |  | Manganese Total (mass conc.) | 7439-96-5 | Metals - Total; | Metals - Total; EMS code: MN-T | Chemical | Total | LAB
MN-T | ug/g | Manganese Total | 0196 | D&G: LTA/HCl: ICP | D&G: LTA/HCl: ICP | 1 | Y | Y | Y | 27 |  |  | 10/24/1996 2:08:12 PM | 1996-04-12 | 6/3/1998 6:51:37 PM | EMS |  | MassPerMass | mg/kg | Y | (mass conc.) |  | Manganese Total (mass conc.) | 7439-96-5 | Metals - Total; | Metals - Total; EMS code: MN-T | Chemical | Total | LAB
MN-T | ug/g | Manganese Total | FB04 | Freeze Dry,Hom,Acid Dig,ICP | Sample moisture is determined by weighing samples before and after freeze-drying. Freeze-dried samples are ground in a Braun blender before being weighed into Teflon PFA liners and digested with concentrated HNO3, concentrated HCl, and 50% H2O2 in a microwave oven; which results in a complete sample dissolution. Samples are collected and volumized with Type I water. Analysis is performed via Inductively Coupled Argon Plasma Atomic Emission Spectrometry (ICP-AES) (simultaneous multielement analysis), Graphite Furnace Atomic Absorption Spectrometry (GF-AAS) (low level Cd and Pb analysis), Cold Vapor Atomic Fluorescence Spectrometry (CV-AFS) (Mercury analysis) and/or Flame Atomic Absorption Spectrometry (F-AAS) (K analysis). | 0.08 | Y | Y | Y | 241 | 8/22/1996 6:11:00 PM | 2/1/2002 2:30:00 PM | 6/23/2003 1:15:31 PM | 1997-02-27 | 6/3/1998 6:51:46 PM | MPAK |  | MassPerMass | mg/kg | Y | (mass conc.) |  | Manganese Total (mass conc.) | 7439-96-5 | Metals - Total; | Metals - Total; EMS code: MN-T | Chemical | Total | LAB
MN-T | mg/L | Manganese Total | F048 | Acid dig,conc'n,nebul.,ICP-OES | This method provides an automated procedure for the Inductively Coupled Argon Plasma Optical Emission Spectrophotometric (ICP-OES) determination and quantitation of trace amounts of metals in surface, ground and precipitation waters after a manual acid digestion and concentration procedure and ultrasonic nebulisation. Organically-bound metals and metal complexes are converted to metal ions by digestion with nitric acid [method details says HCl]. To increase the signal detected by the ICP, the samples are concentrated 10 times. The liquid sample is pumped onto the piezoelectric transducer, where the ultrasonic action efficiently converts it to a fine, dense aerosol. The nebulizer carrier gas flow then transports the wet aerosol through a heated U-tube, where the solvent is vaporized. Subsequently, most of the solvent vapour is removed from the aerosol stream by passing the stream through a chilled condenser. The result is a dry, analyte-laden aerosol which is introduced to the plasma. There, the high temperature desolvates the aerosol and dissociates the sample material into individual atoms and ions, which are excited to emit light at wave lengths characteristic of the elements present. Dissolved samples are field filtered through 0.45 micron cellulose acetate filter prior to acid preservation. Samples for extractable analysis are acidified with 2 mL concentrated HNO3 per litre and allowed to stand a minimum of 24 hours before analysis. Clients providing these samples are advised to preserve samples at the collection site. | 0.0001 | Y | Y | Y | 6161 | 7/4/1997 2:36:00 PM | 6/1/2006 6:01:00 AM | 9/11/2019 8:46:38 AM | 1997-10-20 | 6/3/1998 6:51:47 PM | NPEPPIN |  | Concentration | mg/L | Y | (fl. conc.) |  | Manganese Total (fl. conc.) | 7439-96-5 | Metals - Total; | Metals - Total; EMS code: MN-T | Chemical | Total | LAB
MN-T | ug/g | Manganese Total | ICPM | Metals in Soil by CRC ICPMS | This analysis is carried out using procedures from CSR Analytical Method: "Strong Acid Leachable Metals (SALM) in Soil", BC Ministry of Environment, 26 June 2009, and procedures adapted from EPA Method 200.2. The sample is manually homogenized, dried at 60 degrees Celsius, sieved through a 63 um (230 mesh) sieve, and a representative subsample of the dry material is weighed. The sample is then digested at 95 degrees Celsius for 2 hours by block digester using concentrated nitric and hydrochloric acids. Instrumental analysis is by collision cell inductively coupled plasma - mass spectrometry (modifed from EPA Method 6020A). Method Limitation: This method is not a total digestion technique. It is a very strong acid digestion that is intended to dissolve those metals that may be environmentally available. By design, elements bound in silicate structures are not normally dissolved by this procedure as they are not usually mobile in the environment. | 0.2 | Y | Y | Y | 103 | 9/19/2017 | 9/24/2023 | 1/12/2024 11:16:52 AM | 2018-01-09 |  | JKROGH |  | MassPerMass | mg/kg | Y | (mass conc.) |  | Manganese Total (mass conc.) | 7439-96-5 | Metals - Total; | Metals - Total; EMS code: MN-T | Chemical | Total | LAB
MN-T | ug/g | Manganese Total | E496 | CSSS (1993) 11 | CSSS (1993) 11: Plant available micro nutrients (Cu; Fe; Mn; Zn) are determined by extraction using 0.005 M DTPA followed by analysis using ICPOES. | 0.05 | Y | Y | Y | 0 |  |  |  | 2024-02-27 |  | RDGIBSON |  | MassPerMass | mg/kg | Y | (mass conc.) |  | Manganese Total (mass conc.) | 7439-96-5 | Metals - Total; | Metals - Total; EMS code: MN-T | Chemical | Total | LAB
MN-T | mg/L | Manganese Total | X073 | FA,HNO3,Direct Flame,AA | Samples are field-acidified with HNO3; HNO3 digestion; analysis by AA. | 0.01 | Y | Y | Y | 63 |  |  | 1/25/2007 4:06:12 PM | 1996-04-12 | 6/3/1998 6:51:37 PM | EMS |  | Concentration | mg/L | Y | (fl. conc.) |  | Manganese Total (fl. conc.) | 7439-96-5 | Metals - Total; | Metals - Total; EMS code: MN-T | Chemical | Total | LAB
MN-T | ug/m3 | Manganese Total | F088 | Aliquot,HNO3/HCl dig, ICPMS,loading | Aliquot microwave digested with 3:1 HNO3/HCL. Analysis by ICPMS. For HiVol samples preliminary result are reported in mg/L and loading results are reported in ug/m3 |  | Y | Y | Y | 565 | 8/17/2000 10:46:00 AM | 12/14/2006 | 12/30/2006 1:01:37 AM | 2000-08-16 |  | NPEPPIN |  | AirConcentration | µg/m³ | Y | (air conc.) |  | Manganese Total (air conc.) | 7439-96-5 | Metals - Total; | Metals - Total; EMS code: MN-T | Chemical | Total | LAB
MN-T | mg/L | Manganese Total | F068 | HNO3dig,TeflonTube/Autoclave,ICP-MS | This method is applicable to determiniation of dissolved, extractable, and total recoverable metals in ground waters, surface waters, drinking waters, and wastewaters. Dissolved, extractable, and total recoverable sample fractions are acidified to pH <1.5 with analytical grade nitric acid (HNO3) to prevent the metals from adsorbing to the container material or precipitating. Extractable and total recoverable water samples are collected in 250 ml acid washed polyethylene bottles and are preserved with Nitric acid (0.5 ml acid for 100 mLs of sample). All preserved samples are stored at room temperature and can be held for up to 6 months. Unfiltered, un-preserved samples are stored at 4C for no more than 30 days. For dissolved and extractable water fractions, no sample preparation is necessary. Samples are simply poured in autosampler tubes (i.e. 15 ml centrifuge tubes), and placed on the autosampler and analyzed. Samples for total recoverable metals are digested with Nitric acid in special Teflon tubes in an autoclave prior to analyses. | 5e-06 | Y | Y | Y | 2428 | 6/1/1999 1:33:00 PM | 7/4/2008 9:48:00 AM | 7/30/2024 6:32:12 AM | 1999-06-28 |  | NPEPPIN |  | Concentration | mg/L | Y | (fl. conc.) |  | Manganese Total (fl. conc.) | 7439-96-5 | Metals - Total; | Metals - Total; EMS code: MN-T | Chemical | Total | LAB
MN-T | mg/L | Manganese Total | EQ01 | EQUIS (mg/L) | EQUIS (mg/L) | 0.001 | N | Y | Y | 4360 |  |  | 10/24/1996 3:16:18 PM | 1996-04-12 | 6/3/1998 6:51:37 PM | EMS |  | Concentration | mg/L | Y | (fl. conc.) |  | Manganese Total (fl. conc.) | 7439-96-5 | Metals - Total; | Metals - Total; EMS code: MN-T | Chemical | Total | LAB
MN-T | mg/L | Manganese Total | 0061 | HNO3 Dig: GFAA Analysis | HNO3 Dig: GFAA Analysis | 0.002 | Y | Y | Y | 5 |  |  | 10/24/1996 11:59:39 AM | 1996-04-12 | 6/3/1998 6:51:37 PM | EMS |  | Concentration | mg/L | Y | (fl. conc.) |  | Manganese Total (fl. conc.) | 7439-96-5 | Metals - Total; | Metals - Total; EMS code: MN-T | Chemical | Total | LAB
MN-T | mg/L | Manganese Total | MWID | Metals in Water by ICPMS (digested) | Metals in Water by ICPMS (digested) Based on EPA 200.8 | 0.001 | Y | Y | Y | 36806 |  |  | 7/31/2024 1:46:41 PM | 2009-12-24 |  | RDGIBSON |  | Concentration | mg/L | Y | (fl. conc.) |  | Manganese Total (fl. conc.) | 7439-96-5 | Metals - Total; | Metals - Total; EMS code: MN-T | Chemical | Total | LAB
MN-T | mg/L | Manganese Total | 0040 | HNO3 Dig: ICP Analysis | HNO3 Dig: ICP Analysis | 0.01 | Y | Y | Y | 15108 | 5/8/2003 9:40:00 AM | 9/12/2003 3:16:00 PM | 3/23/2010 5:08:08 PM | 1996-04-12 | 6/3/1998 6:51:37 PM | EMS |  | Concentration | mg/L | Y | (fl. conc.) |  | Manganese Total (fl. conc.) | 7439-96-5 | Metals - Total; | Metals - Total; EMS code: MN-T | Chemical | Total | LAB
MN-T | mg/L | Manganese Total | MSEA | Metals in Seawater (undigested) by (CRC) ICP-MS | This method describes the multi-elemental determination of trace elements by Collision/Reaction Cell (CRC) ICP-MS. The method measures ions produced by a radio-frequency inductively coupled plasma. Analyte species originating in a liquid are pneumatically nebulized and the resulting aerosol transported by argon gas into the plasma torch. The ions produced are entrained in the plasma gas and extracted, by means of a differentially pumped vacuum interface, into a mass spectrometer. The ions flow through a Collision/Reaction Cell (CRC) where depending on the mode of analysis either the ions passes through the cell unhindered (no gas - classical ICP-MS mode), or collisions/reactions chemistry takes place between the ions and a neutral gas such as hydrogen or helium. Any polyatomic ions which give interference on the analytes are removed. The ions are then sorted according to their mass-to-charge ratios by a quadrupole mass spectrometer. The ions are then sorted according to their mass-to-charge ratios by a quadrupole mass spectrometer. The ions transmitted through the quadrupole are quantified by a channel electron multiplier and the ion information processed by a data handling system. Although most polyatomic interference is removed, any additional interference must be assessed and valid corrections applied or the data flagged to indicate problems. Interference correction must include compensation for background ions contributed by the plasma gas, reagents, and constituents of the sample matrix. Instrument drift as well as suppressions or enhancements of instrument response caused by the sample matrix must be corrected for by the use of internal standards. | 0.0003 | Y | Y | Y | 441 | 8/12/2009 12:08:00 AM | 4/29/2024 | 6/7/2024 10:03:16 AM | 2009-09-08 |  | RDGIBSON |  | Concentration | mg/L | Y | (fl. conc.) |  | Manganese Total (fl. conc.) | 7439-96-5 | Metals - Total; | Metals - Total; EMS code: MN-T | Chemical | Total | LAB
MN-T | mg/L | Manganese Total | F038 | HNO3/HCl dig,ICP-AES | Dissolved and extractable water samples are ready for analysis once collected. Total samples are autoclaved with HNO3 and hydrochloric acid (HCI) in capped Teflon vials. Analysis is performed via Inductively Coupled Argon Plasma Atomic Emission Spectrometry (ICP-AES)(simultaneous multielement analysis), Graphite Furnace Atomic Absorption Spectrometry (GF-AAS) (low level analysis) and/or Flame Atomic Absorption Spectrometry (F-AAS). | 0.001 | Y | Y | Y | 5421 | 8/1/1996 6:04:00 PM | 12/2/2010 12:05:00 PM | 9/11/2019 8:47:42 AM | 1997-05-08 | 6/3/1998 6:51:46 PM | MPAK |  | Concentration | mg/L | Y | (fl. conc.) |  | Manganese Total (fl. conc.) | 7439-96-5 | Metals - Total; | Metals - Total; EMS code: MN-T | Chemical | Total | LAB
MN-T | mg/L | Manganese Total | X555 | Metals in water, total, "in-bottle" digestion, ICP-SFMS | The whole water sample collected in a HDPE/Teflon bottle, is de-capped in a Class 100 workstation, acidified to 2% (v/v) HNO3 and 1% (v/v) HCl and recapped. The sample is "in-bottle" digested in a convection oven, set at 60 degrees C for 16 hours. The sample is cooled and introduced to an inductively coupled argon plasma-sector field mass spectrometer (ICP-SFMS) by pneumatic nebulization. Each respective element is measured at a specific mass to charge ratio, m/z value expressed in amu (atomic mass units), and at resolution. LR: Low Resolution ~300, MR: Mid Resolution ~4000 and HR: High a specified instrument Resolution ~10000. Ref: Method number 02-2003, National Laboratory for Environmental Testing, CCIW, Environment Canada, Burlington. Requested April 2003. | 5e-05 | Y | Y | Y | 17343 | 5/20/2003 | 12/6/2023 2:08:00 AM | 2/14/2024 9:05:32 AM | 2004-01-19 |  | UMULLER |  | Concentration | mg/L | Y | (fl. conc.) |  | Manganese Total (fl. conc.) | 7439-96-5 | Metals - Total; | Metals - Total; EMS code: MN-T | Chemical | Total | LAB
MN-T | mg/L | Manganese Total | X561 | Total metals: FA, HNO3, HR-ICPMS (No Prep)Prep) | Ultra trace metals in water are analyzed by high resolution inductively coupled plasma mass spectrometry (HR-ICPMS), modified from US EPA 200.8, (Revision 5.5). The detection limits provided can only be met for undigested samples. This procedure is intended for colorless, non-turbid, acid-preserved water samples (i.e. pristine water samples), having turbidity < 1 NTU and no odor. Where turbidity exceeds 1 NTU, and/or the sample is colored and has an odor, results may be biased low compared to true Total Metals concentrations. ALS recommends that turbidity analysis be requested on samples submitted for this test to aid with interpretation of results. | 5e-05 | Y | Y | Y | 2756 | 6/8/2015 | 10/5/2023 | 1/31/2024 2:37:41 PM | 2015-10-22 |  | RDGIBSON |  | Concentration | mg/L | Y | (fl. conc.) |  | Manganese Total (fl. conc.) | 7439-96-5 | Metals - Total; | Metals - Total; EMS code: MN-T | Chemical | Total | LAB
MN-T | ug/m3 | Manganese Total | 6040 | Aqua Regia Dig:ICP | Aqua regia digestion is used to bring the metals into solution prior to ICP analysis. For Hi-Vol samples, preliminary results are reported as mg/L and24 hr loadings as ug/m3 | 0.008 | Y | Y | Y | 437 | 2/13/2002 | 2/9/2005 | 2/23/2005 1:13:47 AM | 1998-05-05 | 6/3/1998 6:51:48 PM | NPEPPIN |  | AirConcentration | µg/m³ | Y | (air conc.) |  | Manganese Total (air conc.) | 7439-96-5 | Metals - Total; | Metals - Total; EMS code: MN-T | Chemical | Total | LAB
MN-T | mg/L | Manganese Total | 0200 | 20%A.R.Dig:AA Anal. | 20%A.R.Dig:AA Anal. | 0.02 | Y | Y | Y | 6 |  |  | 10/24/1996 2:35:04 AM | 1996-04-12 | 6/3/1998 6:51:37 PM | EMS |  | Concentration | mg/L | Y | (fl. conc.) |  | Manganese Total (fl. conc.) | 7439-96-5 | Metals - Total; | Metals - Total; EMS code: MN-T | Chemical | Total | LAB
MN-T | ppm (S) | Manganese Total | MS61 | ALS Minerals ME-MS61 | ALS Minerals ME-MS61. A prepared sample is digested with HNO3; HF; and HClO4 in a graphite heating block to near dryness. After cooling; the sample is boiled in diluted HCl. Bulk to volume with deionized water; mixed and analyzed by ICP-AES and ICP-MS for all elements. | 5 | Y | Y | Y | 0 |  |  |  | 2024-01-29 |  | EMS |  | DimensionlessRatio | ppm | N | (ratio) |  | Manganese Total (ratio) | 7439-96-5 | Metals - Total; | Metals - Total; EMS code: MN-T | Chemical | Total | LAB
MN-T | ppm (S) | Manganese Total | 6041 | Aqua Regia by ICPMS | These were sediment/soils digested with Aqua-Regia (Nitric Acid & Hydrochloric Acids) and then analysis by ICPMS. Results would be reported in mg/kg and/or %. | 0.1 | Y | Y | Y | 21 | 10/13/2019 | 8/5/2020 | 8/10/2020 3:52:06 PM | 2019-09-06 | 10/21/2019 11:55:35 AM | KATFRASE | KATFRASE | DimensionlessRatio | ppm | N | (ratio) |  | Manganese Total (ratio) | 7439-96-5 | Metals - Total; | Metals - Total; EMS code: MN-T | Chemical | Total | LAB
MN-T | mg/L | Manganese Total | F108 | SPR,IDA Chelation,ICPMS | SPR;IDA Chelation;ICPMS |  | Y | Y | Y | 4 | 11/30/2000 10:32:00 AM | 11/30/2000 10:32:00 AM | 9/10/2002 9:29:48 AM | 2001-03-08 |  | NPEPPIN |  | Concentration | mg/L | Y | (fl. conc.) |  | Manganese Total (fl. conc.) | 7439-96-5 | Metals - Total; | Metals - Total; EMS code: MN-T | Chemical | Total | LAB
MN-T | ppm (S) | Manganese Total | MS41 | ALS Minerals ME-MS41 | ALS Minerals ME-MS41. A prepared sample is digested with aqua regia in a graphite heating block. After cooling; the resulting solution is diluted to with deionized water; mixed and analyzed by ICP-AES. Samples are then analyzed by ICP-MS for the remaining suite of elements. | 5 | Y | Y | Y | 0 |  |  |  | 2024-01-29 |  | EMS |  | DimensionlessRatio | ppm | N | (ratio) |  | Manganese Total (ratio) | 7439-96-5 | Metals - Total; | Metals - Total; EMS code: MN-T | Chemical | Total | LAB
MN-T | mg/dm2/d | Manganese Total | DUST | HNO3 dig.(microwave or hot block),ICPMS | HNO3 dig.(microwave or hot block);ICPMS |  | Y | Y | Y | 267 | 7/19/2016 1:11:00 PM | 7/31/2024 | 7/31/2024 11:23:12 AM | 2008-02-07 |  | RDGIBSON |  | MassFlux | mg/dm²/d | Y | (mass flux) |  | Manganese Total (mass flux) | 7439-96-5 | Metals - Total; | Metals - Total; EMS code: MN-T | Chemical | Total | LAB
MN-T | ug/g | Manganese Total | X551 | Metals in Biota (dry), total, acid digestion, ICP-SFMS | 0.1 gram of dried, homogenized biota is weighed in a microwave TFM vessel, and 8mL of Nitric Acid and 1mL of Hydrogen Peroxide are added. The sample is digested at 200 degrees C for 15 minutes in a closed vessel high pressure microwave oven. The sample is cooled, the volume adjusted to 100mL and introduced to an iductively coupled argon plasma-sector field mass spectrometer (ICP-SFMS) by pneumatic nebulization. Each respective element is measured at a specific mass to charge ratio, m/z value expressed in amu (atomic mass units), and at a specified instrument resolution. LR: Low Resolution ~300, MR: Mid Resolution ~4000 and HR: Hich Resolution ~10000. Ref: Method number 02-2705, National Laboratory for Environmental Testing, CCIW, Environment Canada, Burlington. Requested April 2003. | 0.05 | Y | Y | Y | 49 | 8/11/2015 | 10/30/2023 | 11/28/2023 2:21:33 PM | 2004-01-19 |  | UMULLER |  | MassPerMass | mg/kg | Y | (mass conc.) |  | Manganese Total (mass conc.) | 7439-96-5 | Metals - Total; | Metals - Total; EMS code: MN-T | Chemical | Total | LAB
MN-T | ug/g wet | Manganese Total | BPSY | MET-WET-BPSY-CCMS-VA | Metals in Animal Tissue and Vegetation (Biota) - Prescriptive. EPA 200.3/6020B This method is conducted following British Columbia Lab Manual method "Metals in Animal Tissue and Vegetation (Biota) - Prescriptive". Tissue samples are digested with nitric and hydrochloric acids; in combination with addition of hydrogen peroxide. Instrumental analysis is by collision cell inductively coupled plasma - mass spectrometry (modified from EPA Method 6020B). | 0.01 | Y | Y | Y | 2 | 11/7/2019 | 11/7/2019 | 6/10/2021 1:57:10 PM | 2021-06-04 |  | EMS |  | MassPerMass | µg/g | Y | (mass conc.) | (wet) | Manganese Total (wet) (mass conc.) | 7439-96-5 | Metals - Total; | Metals - Total; EMS code: MN-T | Chemical | Total | LAB
MN-T | mg/L | Manganese Total | METN | MET-T-NP-CCMS-VA | Ultra trace metals in water are analyzed by CRC ICPMS, based on US EPA Method 6020B (July 2014). The detection limits provided can only be met for undigested samples. This procedure is intended for colorless, non-turbid, acid-preserved water samples (i.e. pristine water samples), having turbidity < 1 NTU and no odor. Where turbidity exceeds 1 NTU, and/or the sample is colored and has an odor, results may be biased low compared to true Total Metals concentrations. ALS recommends that turbidity analysis be requested on samples submitted for this test to aid with interpretation of results. Where turbidity is <1NTU, undigested metals are equivalent to total metals concentrations. | 5e-05 | Y | Y | Y | 3648 | 4/3/2012 | 8/19/2024 | 8/26/2024 1:03:06 PM | 2019-08-20 |  | KATFRASE |  | Concentration | mg/L | Y | (fl. conc.) |  | Manganese Total (fl. conc.) | 7439-96-5 | Metals - Total; | Metals - Total; EMS code: MN-T | Chemical | Total | LAB
MN-T | mg/L | Manganese Total | 0060 | HNO3 Dig: HGA Analysis | HNO3 Dig: HGA Analysis | 0.001 | Y | Y | Y | 286 |  |  | 10/24/1996 12:33:07 PM | 1996-04-12 | 6/3/1998 6:51:41 PM | EMS |  | Concentration | mg/L | Y | (fl. conc.) |  | Manganese Total (fl. conc.) | 7439-96-5 | Metals - Total; | Metals - Total; EMS code: MN-T | Chemical | Total | LAB
MN-T | mg/L | Manganese Total | F082 | Aliquot,HNO3/HCL dig,ICPMS | Aliquot microwave digested with 3:1 HNO3/HCL. Analysis by ICPMS. For HiVol samples preliminary result are reported in mg/L and loading results are reported in ug/m3 | 0.0003 | Y | Y | Y | 28259 | 7/25/2005 7:10:00 AM | 8/23/2024 | 8/26/2024 4:53:05 PM | 2000-06-13 | 8/15/2019 11:08:32 AM | NPEPPIN | KATFRASE | Concentration | mg/L | Y | (fl. conc.) |  | Manganese Total (fl. conc.) | 7439-96-5 | Metals - Total; | Metals - Total; EMS code: MN-T | Chemical | Total | LAB
MN-T | ug/g | Manganese Total | PB02 | PBM,HNO3/HCl dig,dry60C,100mesh, | Performance Based Method. Samples are dried at 60 C and sieved using a .149 mm (100 mesh) sieve. Solid samples are digested with the combination of nitric and hydrochloric acids with a 1:1 ratio of acids. For environmental purposes, strong acid leaches are used routinely since they have the advantage of speed, low cost, and superio reproducibility. All metals trapped within the refractory matrix are unlikely to have environmental consequences. Instrumental analysis of extracts produced using this digestion method can be performed by a variety of techniques. |  | Y | Y | Y | 52 | 5/24/2001 3:42:00 PM | 11/26/2001 10:23:00 AM | 5/26/2003 2:28:00 PM | 2001-07-18 |  | NPEPPIN |  | MassPerMass | mg/kg | Y | (mass conc.) |  | Manganese Total (mass conc.) | 7439-96-5 | Metals - Total; | Metals - Total; EMS code: MN-T | Chemical | Total | LAB
MN-T | mg/L | Manganese Total | X351 | LA:HNO3:AA | Samples are lab-acidified (HNO3); digested HNO3; analysis by AA. | 0.01 | Y | Y | Y | 3 |  |  | 6/26/2004 5:29:23 PM | 1996-04-12 | 6/3/1998 6:51:37 PM | EMS |  | Concentration | mg/L | Y | (fl. conc.) |  | Manganese Total (fl. conc.) | 7439-96-5 | Metals - Total; | Metals - Total; EMS code: MN-T | Chemical | Total | LAB
MN-T | mg/L | Manganese Total | X170 | ICP | This code is a catch-all for a large variety of ICP methods including water and also solids digested into solutions as an intermediate result. Different procedures will have greatly different detection limits. (a) For air particulates, preliminary results are reported in mg/L and loading results as mg/dm2/d. (b) For Hi Vol samples preliminiary results are same as reported in mg/L and loading results as ug/m3. (c) Note that for Rain samples there is no filtration nor digestion steps. | 0.01 | Y | Y | Y | 5889 | 3/28/1996 6:03:00 PM | 11/7/2018 | 10/24/2022 10:58:25 AM | 1996-04-12 | 6/3/1998 6:51:37 PM | EMS |  | Concentration | mg/L | Y | (fl. conc.) |  | Manganese Total (fl. conc.) | 7439-96-5 | Metals - Total; | Metals - Total; EMS code: MN-T | Chemical | Total | LAB
MN-T | ug/m3 | Manganese Total | 5312 | Tef:N/P Dig:ICP Analysis | A nitric/perchloric acid digestion if used to bring the metals into solution prior to ICP analysis. For Hi-Vol samples preliminary results are reported as mg/L and 24 hour loading results as ug/m3. | 0.008 | Y | Y | Y | 1185 | 11/7/1997 | 5/15/2000 | 11/18/2000 1:29:35 AM | 1996-10-13 | 6/3/1998 6:51:46 PM | EMS |  | AirConcentration | µg/m³ | Y | (air conc.) |  | Manganese Total (air conc.) | 7439-96-5 | Metals - Total; | Metals - Total; EMS code: MN-T | Chemical | Total | LAB
MN-T | ppm (S) | Manganese Total | 6042 | Aqua Regia by ICPMS | These were sediment/soils digested with Aqua-Regia (Nitric Acid & Hydrochloric Acids) and then analysis by ICPMS. Results would be reported in ppm s. | 0.1 | Y | Y | Y | 71 | 5/14/2022 | 7/27/2024 | 7/29/2024 9:13:08 AM | 2022-06-28 |  | EMS |  | DimensionlessRatio | ppm | N | (ratio) |  | Manganese Total (ratio) | 7439-96-5 | Metals - Total; | Metals - Total; EMS code: MN-T | Chemical | Total | LAB
MN-T | ug/g | Manganese Total | FS82 | PBM,dry60C,10mesh,SALM-ICP-MS | PBM;dry60C;10mesh;SALM-ICP-MS | 0.2 | Y | Y | Y | 3581 | 5/21/2002 | 7/24/2024 | 7/29/2024 5:17:09 PM | 2002-11-12 |  | UMULLER |  | MassPerMass | mg/kg | Y | (mass conc.) |  | Manganese Total (mass conc.) | 7439-96-5 | Metals - Total; | Metals - Total; EMS code: MN-T | Chemical | Total | LAB
MN03 | mg/L | Oil&Grease (Mineral) | MG80 | Extr:Roller mixing (h20) | Extr:Roller mixing (h20) | 1 | Y | Y | Y | 76 | 8/11/2020 | 8/19/2024 11:24:00 AM | 8/22/2024 3:35:15 PM | 1996-04-12 | 6/3/1998 6:51:45 PM | EMS |  | Concentration | mg/L | Y | (fl. conc.) |  | Oil&Grease (Mineral) (fl. conc.) |  | Petroleum Hydrocarbons; | Petroleum Hydrocarbons; EMS code: MN03 | Chemical |  | LAB
MN03 | ug/g | Oil&Grease (Mineral) | 0961 | Ultrasonic CH2Cl2 Extr:G | Ultrasonic CH2Cl2 Extr:G | 100 | Y | Y | Y | 8 | 5/4/2020 | 5/4/2020 | 5/7/2020 9:52:55 AM | 1996-04-12 | 6/3/1998 6:51:36 PM | EMS |  | MassPerMass | mg/kg | Y | (mass conc.) |  | Oil&Grease (Mineral) (mass conc.) |  | Petroleum Hydrocarbons; | Petroleum Hydrocarbons; EMS code: MN03 | Chemical |  | LAB
MNSN | ug/g | Monobutyl Tin | SN02 | Soil/Grignard/GC/MS | Soil/Grignard/GC/MS | 0.005 | N | Y | Y | 6 |  |  | 10/24/1996 1:45:43 PM | 1996-04-12 | 6/3/1998 6:51:36 PM | EMS |  | MassPerMass | mg/kg | Y | (mass conc.) |  | Monobutyl Tin (mass conc.) | 78763-54-9 | Organometallics; | Organometallics; EMS code: MNSN | Chemical |  | LAB
MNSN | mg/L | Monobutyl Tin | SN01 | Water/Grignard/GCMS | Water/Grignard/GCMS | 0.5 | Y | Y | Y | 11 |  |  | 10/24/1996 1:49:08 PM | 1996-04-12 | 6/3/1998 6:51:36 PM | EMS |  | Concentration | mg/L | Y | (fl. conc.) |  | Monobutyl Tin (fl. conc.) | 78763-54-9 | Organometallics; | Organometallics; EMS code: MNSN | Chemical |  | LAB
MO-D | mg/L | Molybdenum Dissolved | F049 | Filt,HNO3 dig,conc'n,nebul'ICP-OES | This method provides an automated procedure for the Inductively Coupled Argon Plasma Optical Emission Spectrophotometric (ICP-OES) determination and quantitation of trace amounts of metals in surface, ground and precipitation waters after a manual acid digestion and concentration procedure and ultrasonic nebulisation. Organically-bound metals and metal complexes are converted to metal ions by digestion with nitric acid. To increase the signal detected by the ICP, the samples are concentrated 10 times. The liquid sample is pumped onto the piezoelectric transducer, where the ultrasonic action efficiently converts it to a fine, dense aerosol. The nebulizer carrier gas flow then transports the wet aerosol through a heated U-tube, where the solvent is vaporized. Subsequently, most of the solvent vapour is removed from the aerosol stream by passing the stream through a chilled condenser. The result is a dry, analyte-laden aerosol which is introduced to the plasma. There, the high temperature desolvates the aerosol and dissociates the sample material into individual atoms and ions, which are excited to emit light at wave lengths characteristic of the elements present. Dissolved samples are field filtered through 0.45 micron cellulose acetate filter prior to acid preservation. Samples for extractable analysis are acidified with 2 mL concentrated HNO3 per litre and allowed to stand a minimum of 24 hours before analysis. Clients providing these samples are advised to preserve samples at the collection site. | 0.0001 | Y | Y | Y | 1485 |  |  | 5/19/2015 3:10:33 PM | 1997-10-20 | 6/3/1998 6:51:47 PM | NPEPPIN |  | Concentration | mg/L | Y | (fl. conc.) |  | Molybdenum Dissolved (fl. conc.) | 7439-98-7 | Metals - Dissolved; | Metals - Dissolved; EMS code: MO-D | Chemical | Dissolved | LAB
MO-D | mg/L | Molybdenum Dissolved | F023 | HNO3,GFAA | Dissolved and extractable water samples are ready for analysis once collected. Total samples are autoclaved with HNO3 and hydrochloric acid (HCl) in capped Teflon vials. Analysis is performed via Inductively Coupled Argon Plasma Atomic Emission Spectrometry (ICP-AES) (simultaneous multielement analysis), Graphite Furnace Atomic Absorption Spectrometry (GF-AAS) (low level analysis) and/or Flame Atomic Absorption Spectrometry (F-AAS). Note that extractable samples are acidified with 0.5 mL HNO3 per 100 mL of sample at time of sample collection. | 0.0005 | Y | Y | Y | 109 | 6/4/1997 12:50:00 PM | 8/11/1999 4:59:00 PM | 9/1/1999 1:02:44 AM | 1997-04-02 | 6/3/1998 6:51:46 PM | MPAK |  | Concentration | mg/L | Y | (fl. conc.) |  | Molybdenum Dissolved (fl. conc.) | 7439-98-7 | Metals - Dissolved; | Metals - Dissolved; EMS code: MO-D | Chemical | Dissolved | LAB
MO-D | mg/L | Molybdenum Dissolved | X205 | LF,AA,Dir Asp | LF;AA;Dir Asp | 0.1 | Y | Y | Y | 167 |  |  | 10/24/1996 9:16:37 AM | 1996-04-12 | 6/3/1998 6:51:37 PM | EMS |  | Concentration | mg/L | Y | (fl. conc.) |  | Molybdenum Dissolved (fl. conc.) | 7439-98-7 | Metals - Dissolved; | Metals - Dissolved; EMS code: MO-D | Chemical | Dissolved | LAB
MO-D | mg/L | Molybdenum Dissolved | X285 | HNO3 Dig:ICP Analysis | HNO3 Dig:ICP Analysis | 0.03 | Y | Y | Y | 671 | 1/30/2002 | 11/13/2002 | 3/18/2004 10:28:34 AM | 1998-05-22 | 6/3/1998 6:51:49 PM | NPEPPIN |  | Concentration | mg/L | Y | (fl. conc.) |  | Molybdenum Dissolved (fl. conc.) | 7439-98-7 | Metals - Dissolved; | Metals - Dissolved; EMS code: MO-D | Chemical | Dissolved | LAB
MO-D | mg/L | Molybdenum Dissolved | EQ01 | EQUIS (mg/L) | EQUIS (mg/L) | 0.0005 | N | Y | Y | 3343 |  |  | 10/23/1996 12:07:04 AM | 1996-04-12 | 6/3/1998 6:51:37 PM | EMS |  | Concentration | mg/L | Y | (fl. conc.) |  | Molybdenum Dissolved (fl. conc.) | 7439-98-7 | Metals - Dissolved; | Metals - Dissolved; EMS code: MO-D | Chemical | Dissolved | LAB
MO-D | mg/L | Molybdenum Dissolved | X357 | LF,LA:HNO3:GFAA | Samples are lab-filtered (0.45 um) and lab-acidified with HNO3; analysis by GFAA. | 0.001 | Y | Y | Y | 1 | 11/5/1996 8:30:00 AM | 11/5/1996 8:30:00 AM | 8/14/1997 7:42:16 PM | 1996-04-12 | 6/3/1998 6:51:37 PM | EMS |  | Concentration | mg/L | Y | (fl. conc.) |  | Molybdenum Dissolved (fl. conc.) | 7439-98-7 | Metals - Dissolved; | Metals - Dissolved; EMS code: MO-D | Chemical | Dissolved | LAB
MO-D | mg/L | Molybdenum Dissolved | X350 | FF,FA:HNO3:ICP | Samples are field-filtered (0.45 um); field-acidified with HNO3; analysis by ICP/AES. | 0.005 | Y | Y | Y | 1861 | 5/9/2003 9:40:00 AM | 9/16/2003 3:18:00 PM | 1/29/2010 9:45:12 AM | 1996-04-12 | 6/3/1998 6:51:41 PM | EMS |  | Concentration | mg/L | Y | (fl. conc.) |  | Molybdenum Dissolved (fl. conc.) | 7439-98-7 | Metals - Dissolved; | Metals - Dissolved; EMS code: MO-D | Chemical | Dissolved | LAB
MO-D | mg/L | Molybdenum Dissolved | F068 | HNO3dig,TeflonTube/Autoclave,ICP-MS | This method is applicable to determiniation of dissolved, extractable, and total recoverable metals in ground waters, surface waters, drinking waters, and wastewaters. Dissolved, extractable, and total recoverable sample fractions are acidified to pH <1.5 with analytical grade nitric acid (HNO3) to prevent the metals from adsorbing to the container material or precipitating. Extractable and total recoverable water samples are collected in 250 ml acid washed polyethylene bottles and are preserved with Nitric acid (0.5 ml acid for 100 mLs of sample). All preserved samples are stored at room temperature and can be held for up to 6 months. Unfiltered, un-preserved samples are stored at 4C for no more than 30 days. For dissolved and extractable water fractions, no sample preparation is necessary. Samples are simply poured in autosampler tubes (i.e. 15 ml centrifuge tubes), and placed on the autosampler and analyzed. Samples for total recoverable metals are digested with Nitric acid in special Teflon tubes in an autoclave prior to analyses. | 5e-05 | Y | Y | Y | 1398 | 6/1/1999 1:21:00 PM | 4/17/2015 1:06:00 PM | 8/11/2021 12:04:11 PM | 1999-06-28 |  | NPEPPIN |  | Concentration | mg/L | Y | (fl. conc.) |  | Molybdenum Dissolved (fl. conc.) | 7439-98-7 | Metals - Dissolved; | Metals - Dissolved; EMS code: MO-D | Chemical | Dissolved | LAB
MO-D | mg/L | Molybdenum Dissolved | X028 | FA,HNO3,AA Analysis | FA;HNO3;AA Analysis | 0.0005 | Y | Y | Y | 35 |  |  | 10/24/1996 9:16:52 AM | 1996-04-12 | 6/3/1998 6:51:37 PM | EMS |  | Concentration | mg/L | Y | (fl. conc.) |  | Molybdenum Dissolved (fl. conc.) | 7439-98-7 | Metals - Dissolved; | Metals - Dissolved; EMS code: MO-D | Chemical | Dissolved | LAB
MO-D | mg/L | Molybdenum Dissolved | X556 | Metals in water, dissolved, "in-bottle" digestion, ICP-SFMS | The whole water sample is filtered through a 0.45 micron cellulose acetate filter paper and the filtrate is stored in a HDPE/Teflon bottle. The container is de-capped in a Class 100 workstation, acidified to 2% (v/v) HNO3 and 1% (v/v) HCl and recapped. The sample is "in-bottle" digested in a convection oven, set at 60 degrees C for 16 hours. The sample is cooled and introduced to an inductively coupled argon plasma-sector field mass spectrometer (ICP-SFMS) by pneumatic nebulization. Each respective element is measured at a specific mass to charge ratio, m/z value expressed in amu (atomic mass units), and at a specified intrument resolution. LR: Low Resolution ~300, MR: Mid Resolution ~4000 and HR: High Resolution ~10000. Ref: Method number 02-2003, National Laboratory for Environmental Testing, CCIW, Environment Canada, Burlington, Requested April 2003. | 1e-05 | Y | Y | Y | 4235 | 6/11/2003 | 12/6/2023 2:04:00 AM | 2/14/2024 9:05:32 AM | 2004-01-19 |  | UMULLER |  | Concentration | mg/L | Y | (fl. conc.) |  | Molybdenum Dissolved (fl. conc.) | 7439-98-7 | Metals - Dissolved; | Metals - Dissolved; EMS code: MO-D | Chemical | Dissolved | LAB
MO-D | mg/L | Molybdenum Dissolved | METC | MET-D-NP-CCMS-VA | Ultra trace metals in water are analyzed by CRC ICPMS, based on US EPA Method 6020B (July 2014). This procedure is intended for pristine field-filtered acid-preserved water samples. The detection limits (LOR) for this test are based on lab instrumental analysis only, not including filtration. ALS-supplied field filtration equipment does not support these LOR. Therefore, because of the high probability of false positives due to filtration, it is strongly recommended that a filtration blank be analysed to aid in data interpretation. | 5e-05 | Y | Y | Y | 1776 | 4/3/2012 | 8/15/2024 | 8/23/2024 9:50:35 AM | 2019-08-20 |  | KATFRASE |  | Concentration | mg/L | Y | (fl. conc.) |  | Molybdenum Dissolved (fl. conc.) | 7439-98-7 | Metals - Dissolved; | Metals - Dissolved; EMS code: MO-D | Chemical | Dissolved | LAB
MO-D | mg/L | Molybdenum Dissolved | MSEA | Metals in Seawater (undigested) by (CRC) ICP-MS | This method describes the multi-elemental determination of trace elements by Collision/Reaction Cell (CRC) ICP-MS. The method measures ions produced by a radio-frequency inductively coupled plasma. Analyte species originating in a liquid are pneumatically nebulized and the resulting aerosol transported by argon gas into the plasma torch. The ions produced are entrained in the plasma gas and extracted, by means of a differentially pumped vacuum interface, into a mass spectrometer. The ions flow through a Collision/Reaction Cell (CRC) where depending on the mode of analysis either the ions passes through the cell unhindered (no gas - classical ICP-MS mode), or collisions/reactions chemistry takes place between the ions and a neutral gas such as hydrogen or helium. Any polyatomic ions which give interference on the analytes are removed. The ions are then sorted according to their mass-to-charge ratios by a quadrupole mass spectrometer. The ions are then sorted according to their mass-to-charge ratios by a quadrupole mass spectrometer. The ions transmitted through the quadrupole are quantified by a channel electron multiplier and the ion information processed by a data handling system. Although most polyatomic interference is removed, any additional interference must be assessed and valid corrections applied or the data flagged to indicate problems. Interference correction must include compensation for background ions contributed by the plasma gas, reagents, and constituents of the sample matrix. Instrument drift as well as suppressions or enhancements of instrument response caused by the sample matrix must be corrected for by the use of internal standards. | 2e-05 | Y | Y | Y | 4 | 4/12/2023 | 4/29/2024 | 5/16/2024 2:00:14 PM | 2009-09-08 |  | RDGIBSON |  | Concentration | mg/L | Y | (fl. conc.) |  | Molybdenum Dissolved (fl. conc.) | 7439-98-7 | Metals - Dissolved; | Metals - Dissolved; EMS code: MO-D | Chemical | Dissolved | LAB
MO-D | mg/L | Molybdenum Dissolved | X123 | FF,FA,HNO3,Nitrous Ox,AA | FF;FA;HNO3;Nitrous Ox;AA | 0.0005 | Y | Y | Y | 147 |  |  | 10/23/1996 11:07:32 PM | 1996-04-12 | 6/3/1998 6:51:37 PM | EMS |  | Concentration | mg/L | Y | (fl. conc.) |  | Molybdenum Dissolved (fl. conc.) | 7439-98-7 | Metals - Dissolved; | Metals - Dissolved; EMS code: MO-D | Chemical | Dissolved | LAB
MO-D | mg/L | Molybdenum Dissolved | X235 | HCl Dig,AA Analysis | HCl Dig;AA Analysis | 0.1 | Y | Y | Y | 318 |  |  | 10/24/1996 12:12:10 PM | 1996-04-12 | 6/3/1998 6:51:37 PM | EMS |  | Concentration | mg/L | Y | (fl. conc.) |  | Molybdenum Dissolved (fl. conc.) | 7439-98-7 | Metals - Dissolved; | Metals - Dissolved; EMS code: MO-D | Chemical | Dissolved | LAB
MO-D | mg/L | Molybdenum Dissolved | X170 | ICP | This code is a catch-all for a large variety of ICP methods including water and also solids digested into solutions as an intermediate result. Different procedures will have greatly different detection limits. (a) For air particulates, preliminary results are reported in mg/L and loading results as mg/dm2/d. (b) For Hi Vol samples preliminiary results are same as reported in mg/L and loading results as ug/m3. (c) Note that for Rain samples there is no filtration nor digestion steps. | 0.03 | Y | Y | Y | 16673 | 4/11/1996 6:02:00 PM | 11/6/2018 | 6/12/2020 1:21:12 PM | 1996-04-12 | 6/3/1998 6:51:41 PM | EMS |  | Concentration | mg/L | Y | (fl. conc.) |  | Molybdenum Dissolved (fl. conc.) | 7439-98-7 | Metals - Dissolved; | Metals - Dissolved; EMS code: MO-D | Chemical | Dissolved | LAB
MO-D | mg/L | Molybdenum Dissolved | X544 | FF,FA,HNO3,ICPMS | Analytes nebulized followed by ion entrained in plasma gas and extracted by means of a differentially pumped vacuum interface into a mass spectrometer | 0.0001 | Y | Y | Y | 90 |  |  | 10/30/2014 8:31:18 AM | 2001-09-18 | 9/18/2001 1:36:23 PM | NPEPPIN | NPEPPIN | Concentration | mg/L | Y | (fl. conc.) |  | Molybdenum Dissolved (fl. conc.) | 7439-98-7 | Metals - Dissolved; | Metals - Dissolved; EMS code: MO-D | Chemical | Dissolved | LAB
MO-D | mg/L | Molybdenum Dissolved | X562 | Dissolved metals: FF, FA, HNO3, HR-ICPMS | Ultra trace metals in water are analyzed by high resolution inductively coupled plasma mass spectrometry (HR-ICPMS), based on US EPA Method 200.8, (Oct. 1999). This procedure is intended for pristine field-filtered acid-preserved water samples. The detection limits (LOR) for this test are based on lab instrumental analysis only, not including filtration. ALS-supplied field filtration equipment does not support these LOR. Therefore, because of the high probability of false positives due to filtration, it is strongly recommended that a filtration blank be analyzed to aid in data interpretation. | 5e-05 | Y | Y | Y | 1264 | 6/8/2015 | 10/19/2023 | 11/9/2023 7:41:27 AM | 2015-10-22 |  | RDGIBSON |  | Concentration | mg/L | Y | (fl. conc.) |  | Molybdenum Dissolved (fl. conc.) | 7439-98-7 | Metals - Dissolved; | Metals - Dissolved; EMS code: MO-D | Chemical | Dissolved | LAB
MO-D | mg/L | Molybdenum Dissolved | MWID | Metals in Water by ICPMS (digested) | Metals in Water by ICPMS (digested) Based on EPA 200.8 | 0.001 | Y | Y | Y | 31768 | 10/26/2018 5:20:00 PM | 10/26/2018 5:20:00 PM | 7/31/2024 1:46:41 PM | 2009-12-24 |  | RDGIBSON |  | Concentration | mg/L | Y | (fl. conc.) |  | Molybdenum Dissolved (fl. conc.) | 7439-98-7 | Metals - Dissolved; | Metals - Dissolved; EMS code: MO-D | Chemical | Dissolved | LAB
MO-D | mg/L | Molybdenum Dissolved | F032 | Filt.,Acid,ICPMS | There is no acid digestion required for this analysis. A water sample is collected and preserved in 0.2 percent nitric acid. To analyse the extractable metals the sample is shaken, allowed to sit for at least twenty four hours, and a decanted portion is introduced to the spectrometer. When the dissolved metals are required, the samples must first be filtered through a 0.45 micron cellulose acetate filter paper and then followed by preservation in 0.2 percent nitric acid. It is important to note that different laboratories may have very different detection limits for some variables analyzed by ICPMS. | 0.001 | Y | Y | Y | 82892 | 3/10/1998 2:07:00 PM | 8/26/2024 | 8/26/2024 4:53:05 PM | 1998-05-14 | 8/15/2019 10:57:17 AM | NPEPPIN | KATFRASE | Concentration | mg/L | Y | (fl. conc.) |  | Molybdenum Dissolved (fl. conc.) | 7439-98-7 | Metals - Dissolved; | Metals - Dissolved; EMS code: MO-D | Chemical | Dissolved | LAB
MO-D | mg/L | Molybdenum Dissolved | 0070 | Ext(Oxime):AA Analysis | Ext(Oxime):AA Analysis | 0.0005 | Y | Y | Y | 163 | 9/29/2000 | 12/1/2010 | 12/9/2010 3:55:27 PM | 1996-04-12 | 6/3/1998 6:51:37 PM | EMS |  | Concentration | mg/L | Y | (fl. conc.) |  | Molybdenum Dissolved (fl. conc.) | 7439-98-7 | Metals - Dissolved; | Metals - Dissolved; EMS code: MO-D | Chemical | Dissolved | LAB
MO-D | mg/L | Molybdenum Dissolved | X085 | LF,LA,HNO3,AA | Samples are lab-filtered (0.45 um), and lab-acidified (HNO3); analysis by AA. | 0.0005 | Y | Y | Y | 315 | 12/12/1996 | 9/18/2015 | 10/28/2015 9:54:16 AM | 1996-04-12 | 6/3/1998 6:51:37 PM | EMS |  | Concentration | mg/L | Y | (fl. conc.) |  | Molybdenum Dissolved (fl. conc.) | 7439-98-7 | Metals - Dissolved; | Metals - Dissolved; EMS code: MO-D | Chemical | Dissolved | LAB
MO-D | mg/L | Molybdenum Dissolved | F039 | Field filt + acid, ICP-AES | Dissolved and extractable water samples are ready for analysis once collected. Total samples are autoclaved with HNO3 and hydrochloric acid (HCI) in capped Teflon vials. Analysis is performed via Inductively Coupled Argon Plasma Atomic Emission Spectrometry (ICP-AES)(simultaneous multielement analysis), Graphite Furnace Atomic Absorption Spectrometry (GF-AAS)(low level analysis) and/or Flame Atomic Absorption Spectrometry (F-AAS). | 0.01 | Y | Y | Y | 5969 | 7/29/1996 6:03:00 PM | 7/5/2019 1:21:00 PM | 9/11/2019 8:47:50 AM | 1997-05-08 | 6/3/1998 6:51:46 PM | MPAK |  | Concentration | mg/L | Y | (fl. conc.) |  | Molybdenum Dissolved (fl. conc.) | 7439-98-7 | Metals - Dissolved; | Metals - Dissolved; EMS code: MO-D | Chemical | Dissolved | LAB
MO-D | mg/L | Molybdenum Dissolved | X116 | FF,FA,HNO3,Graph.Furn,AA | Samples are field-filtered (0.45 um), field-acidified with HNO3; analysis by GFAA. | 0.0005 | Y | Y | Y | 7545 |  |  | 8/28/2018 3:14:04 PM | 1996-04-12 | 6/3/1998 6:51:37 PM | EMS |  | Concentration | mg/L | Y | (fl. conc.) |  | Molybdenum Dissolved (fl. conc.) | 7439-98-7 | Metals - Dissolved; | Metals - Dissolved; EMS code: MO-D | Chemical | Dissolved | LAB
MO-E | ug/g | Molybdenum Extractable | X230 | Hom,HNO3/HCl,ICP | The sample is dried, ground to pass 100 mesh, and digested with a nitric-hydrochloric acid mixture to solubilize the solid matter and to remove the organic material by oxidation and volatilization. The completed digests are submitted to metals instrumentation personnel for metals analysis by Inductively Coupled Atomic Emission Spectrophotometry (ICP-AES), or hydride analysis by atomic absorption. | 5 | Y | Y | Y | 70 |  |  | 10/24/1996 2:25:28 PM | 1996-04-12 | 6/3/1998 6:51:37 PM | EMS |  | MassPerMass | mg/kg | Y | (mass conc.) |  | Molybdenum Extractable (mass conc.) | 7439-98-7 | Metals - Extractable; | Metals - Extractable; EMS code: MO-E | Chemical |  | LAB
MO-E | mg/L | Molybdenum Extractable | X553 | Metals in water, extractable, ICP-SFMS | The whole water sample is collected in a HDPE/Teflon bottle. In a Class 100 workstation it is preserved with concentrated Nitric Acid at concentration of 0.2 percent (v/v). The acid preservative extracts the metals from material in the sample and a decanted portion is introduced directly to an inductively coupled argon plasma-sector field mass spectrometer (ICP-SFMS) by pneumatic nebulization. Each respective element is measured at a specific mass to charge ratio, m/z value expressed in amu (atomic mass units), and at a specified instrument resolution. LR: Low Resolution ~300, MR: Mid Resolution ~4000 and HR: High Resolution ~10000. Ref: Method number 02-2061, National Laboratory for Environmental Testing, CCIW, Environment Canada, Burlington. Requested April 2003. | 5e-06 | Y | Y | Y | 2584 | 5/22/2003 | 9/21/2007 | 4/18/2018 9:28:02 AM | 2004-01-19 |  | UMULLER |  | Concentration | mg/L | Y | (fl. conc.) |  | Molybdenum Extractable (fl. conc.) | 7439-98-7 | Metals - Extractable; | Metals - Extractable; EMS code: MO-E | Chemical |  | LAB
MO-E | mg/L | Molybdenum Extractable | F040 | Field Acid, lab filt, ICP-AES | Dissolved and extractable water samples are ready for analysis once collected. Total samples are autoclaved with HNO3 and hydrochloric acid (HCI) in capped Teflon vials. Analysis is performed via Inductively Coupled Argon Plasma Atomic Emission Spectrometry (ICP-AES)(simultaneous multielement analysis), Graphite Furnace Atomic Absorption Spectrometry (GF-AAS)(low level analysis) and/or Flame Atomic Absorption Spectrometry (F-AAS). | 0.01 | Y | Y | Y | 2350 | 8/5/1997 6:06:00 PM | 3/14/2008 | 9/11/2019 8:47:41 AM | 1997-05-08 | 6/3/1998 6:51:46 PM | MPAK |  | Concentration | mg/L | Y | (fl. conc.) |  | Molybdenum Extractable (fl. conc.) | 7439-98-7 | Metals - Extractable; | Metals - Extractable; EMS code: MO-E | Chemical |  | LAB
MO-L | mg/L | Molybdenum Leachable | MEND | MEND (mod) EPA 300.1 (mod) | MEND (mod) EPA 300.1 (mod). Inorganic anions are analyzed by obtaining an extract from a soil sample that has been added in a 3:1 ratio of deionized water to soil; then leached and the extract filtered prior to analysis; which is then analyzed by Ion Chromatography with conductivity and/or UV detection. | 0.0001 | Y | Y | Y | 0 |  |  |  | 2024-01-29 |  | RDGIBSON |  | Concentration | mg/L | Y | (fl. conc.) |  | Molybdenum Leachable (fl. conc.) |  | Metals - Extractable; | Metals - Extractable; EMS code: MO-L | Chemical |  | LAB
MO-T | ug/g | Molybdenum Total | FS82 | PBM,dry60C,10mesh,SALM-ICP-MS | PBM;dry60C;10mesh;SALM-ICP-MS | 0.1 | Y | Y | Y | 3687 | 5/21/2002 | 7/24/2024 | 7/30/2024 10:00:14 AM | 2002-11-12 |  | UMULLER |  | MassPerMass | mg/kg | Y | (mass conc.) |  | Molybdenum Total (mass conc.) | 7439-98-7 | Metals - Total; | Metals - Total; EMS code: MO-T | Chemical | Total | LAB
MO-T | mg/L | Molybdenum Total | 0020 | HNO3 Dig:AA Analysis | HNO3 Dig:AA Analysis | 0.1 | Y | Y | Y | 84 |  |  | 8/26/2004 4:21:03 PM | 1996-04-12 | 6/3/1998 6:51:41 PM | EMS |  | Concentration | mg/L | Y | (fl. conc.) |  | Molybdenum Total (fl. conc.) | 7439-98-7 | Metals - Total; | Metals - Total; EMS code: MO-T | Chemical | Total | LAB
MO-T | mg/L | Molybdenum Total | F024 | HNO3/HCl,GFAA | Dissolved and extractable water samples are ready for analysis once collected. Total samples are autoclaved with HNO3 and hydrochloric acid (HCl) in capped Teflon vials. Analysis is performed via Inductively Coupled Argon Plasma Atomic Emission Spectrometry (ICP-AES) (simultaneous multielement analysis), Graphite Furnace Atomic Absorption Spectrometry (GF-AAS) (low level analysis) and/or Flame Atomic Absorption Spectrometry (F-AAS). | 0.0005 | Y | Y | Y | 107 | 3/6/1997 5:15:00 PM | 10/1/1999 4:45:00 PM | 10/6/1999 1:06:23 AM | 1997-04-02 | 6/3/1998 6:51:46 PM | MPAK |  | Concentration | mg/L | Y | (fl. conc.) |  | Molybdenum Total (fl. conc.) | 7439-98-7 | Metals - Total; | Metals - Total; EMS code: MO-T | Chemical | Total | LAB
MO-T | ppm (S) | Molybdenum Total | MS41 | ALS Minerals ME-MS41 | ALS Minerals ME-MS41. A prepared sample is digested with aqua regia in a graphite heating block. After cooling; the resulting solution is diluted to with deionized water; mixed and analyzed by ICP-AES. Samples are then analyzed by ICP-MS for the remaining suite of elements. | 0.05 | Y | Y | Y | 0 |  |  |  | 2024-01-29 |  | EMS |  | DimensionlessRatio | ppm | N | (ratio) |  | Molybdenum Total (ratio) | 7439-98-7 | Metals - Total; | Metals - Total; EMS code: MO-T | Chemical | Total | LAB
MO-T | ug/g | Molybdenum Total | FB04 | Freeze Dry,Hom,Acid Dig,ICP | Sample moisture is determined by weighing samples before and after freeze-drying. Freeze-dried samples are ground in a Braun blender before being weighed into Teflon PFA liners and digested with concentrated HNO3, concentrated HCl, and 50% H2O2 in a microwave oven; which results in a complete sample dissolution. Samples are collected and volumized with Type I water. Analysis is performed via Inductively Coupled Argon Plasma Atomic Emission Spectrometry (ICP-AES) (simultaneous multielement analysis), Graphite Furnace Atomic Absorption Spectrometry (GF-AAS) (low level Cd and Pb analysis), Cold Vapor Atomic Fluorescence Spectrometry (CV-AFS) (Mercury analysis) and/or Flame Atomic Absorption Spectrometry (F-AAS) (K analysis). | 0.8 | Y | Y | Y | 241 | 8/22/1996 6:11:00 PM | 2/1/2002 2:30:00 PM | 6/23/2003 1:15:32 PM | 1997-02-27 | 6/3/1998 6:51:46 PM | MPAK |  | MassPerMass | mg/kg | Y | (mass conc.) |  | Molybdenum Total (mass conc.) | 7439-98-7 | Metals - Total; | Metals - Total; EMS code: MO-T | Chemical | Total | LAB
MO-T | mg/L | Molybdenum Total | X170 | ICP | This code is a catch-all for a large variety of ICP methods including water and also solids digested into solutions as an intermediate result. Different procedures will have greatly different detection limits. (a) For air particulates, preliminary results are reported in mg/L and loading results as mg/dm2/d. (b) For Hi Vol samples preliminiary results are same as reported in mg/L and loading results as ug/m3. (c) Note that for Rain samples there is no filtration nor digestion steps. | 0.004 | Y | Y | Y | 6288 | 3/28/1996 6:03:00 PM | 11/7/2018 | 10/24/2022 10:58:25 AM | 1996-04-12 | 6/3/1998 6:51:37 PM | EMS |  | Concentration | mg/L | Y | (fl. conc.) |  | Molybdenum Total (fl. conc.) | 7439-98-7 | Metals - Total; | Metals - Total; EMS code: MO-T | Chemical | Total | LAB
MO-T | mg/L | Molybdenum Total | 0040 | HNO3 Dig: ICP Analysis | HNO3 Dig: ICP Analysis | 0.03 | Y | Y | Y | 15198 | 6/25/2010 3:15:00 PM | 12/1/2010 | 12/9/2010 3:55:27 PM | 1996-04-12 | 6/3/1998 6:51:37 PM | EMS |  | Concentration | mg/L | Y | (fl. conc.) |  | Molybdenum Total (fl. conc.) | 7439-98-7 | Metals - Total; | Metals - Total; EMS code: MO-T | Chemical | Total | LAB
MO-T | ug/m3 | Molybdenum Total | F088 | Aliquot,HNO3/HCl dig, ICPMS,loading | Aliquot microwave digested with 3:1 HNO3/HCL. Analysis by ICPMS. For HiVol samples preliminary result are reported in mg/L and loading results are reported in ug/m3 |  | Y | Y | Y | 565 | 8/17/2000 10:46:00 AM | 12/14/2006 | 12/30/2006 1:01:37 AM | 2000-08-16 |  | NPEPPIN |  | AirConcentration | µg/m³ | Y | (air conc.) |  | Molybdenum Total (air conc.) | 7439-98-7 | Metals - Total; | Metals - Total; EMS code: MO-T | Chemical | Total | LAB
MO-T | ug/g | Molybdenum Total | PB02 | PBM,HNO3/HCl dig,dry60C,100mesh, | Performance Based Method. Samples are dried at 60 C and sieved using a .149 mm (100 mesh) sieve. Solid samples are digested with the combination of nitric and hydrochloric acids with a 1:1 ratio of acids. For environmental purposes, strong acid leaches are used routinely since they have the advantage of speed, low cost, and superio reproducibility. All metals trapped within the refractory matrix are unlikely to have environmental consequences. Instrumental analysis of extracts produced using this digestion method can be performed by a variety of techniques. |  | Y | Y | Y | 52 | 5/24/2001 3:42:00 PM | 11/26/2001 10:23:00 AM | 5/26/2003 2:28:05 PM | 2001-07-18 |  | NPEPPIN |  | MassPerMass | mg/kg | Y | (mass conc.) |  | Molybdenum Total (mass conc.) | 7439-98-7 | Metals - Total; | Metals - Total; EMS code: MO-T | Chemical | Total | LAB
MO-T | mg/L | Molybdenum Total | X561 | Total metals: FA, HNO3, HR-ICPMS (No Prep)Prep) | Ultra trace metals in water are analyzed by high resolution inductively coupled plasma mass spectrometry (HR-ICPMS), modified from US EPA 200.8, (Revision 5.5). The detection limits provided can only be met for undigested samples. This procedure is intended for colorless, non-turbid, acid-preserved water samples (i.e. pristine water samples), having turbidity < 1 NTU and no odor. Where turbidity exceeds 1 NTU, and/or the sample is colored and has an odor, results may be biased low compared to true Total Metals concentrations. ALS recommends that turbidity analysis be requested on samples submitted for this test to aid with interpretation of results. | 5e-05 | Y | Y | Y | 2756 | 6/8/2015 | 10/5/2023 | 1/31/2024 2:37:41 PM | 2015-10-22 |  | RDGIBSON |  | Concentration | mg/L | Y | (fl. conc.) |  | Molybdenum Total (fl. conc.) | 7439-98-7 | Metals - Total; | Metals - Total; EMS code: MO-T | Chemical | Total | LAB
MO-T | mg/L | Molybdenum Total | TCLP | Toxicity Characteristic Leaching Procedure | The TCLP procedure is designed to simulate the climatic leaching action | 0.1 | Y | Y | Y | 22 | 12/8/2009 7:49:00 PM | 9/28/2011 5:53:00 PM | 11/3/2011 11:10:19 PM | 2009-03-06 |  | JCHUDAK |  | Concentration | mg/L | Y | (fl. conc.) |  | Molybdenum Total (fl. conc.) | 7439-98-7 | Metals - Total; | Metals - Total; EMS code: MO-T | Chemical | Total | LAB
MO-T | mg/L | Molybdenum Total | X555 | Metals in water, total, "in-bottle" digestion, ICP-SFMS | The whole water sample collected in a HDPE/Teflon bottle, is de-capped in a Class 100 workstation, acidified to 2% (v/v) HNO3 and 1% (v/v) HCl and recapped. The sample is "in-bottle" digested in a convection oven, set at 60 degrees C for 16 hours. The sample is cooled and introduced to an inductively coupled argon plasma-sector field mass spectrometer (ICP-SFMS) by pneumatic nebulization. Each respective element is measured at a specific mass to charge ratio, m/z value expressed in amu (atomic mass units), and at resolution. LR: Low Resolution ~300, MR: Mid Resolution ~4000 and HR: High a specified instrument Resolution ~10000. Ref: Method number 02-2003, National Laboratory for Environmental Testing, CCIW, Environment Canada, Burlington. Requested April 2003. | 5e-06 | Y | Y | Y | 17337 | 5/20/2003 | 12/6/2023 2:08:00 AM | 2/14/2024 9:05:32 AM | 2004-01-19 |  | UMULLER |  | Concentration | mg/L | Y | (fl. conc.) |  | Molybdenum Total (fl. conc.) | 7439-98-7 | Metals - Total; | Metals - Total; EMS code: MO-T | Chemical | Total | LAB
MO-T | mg/L | Molybdenum Total | MSEA | Metals in Seawater (undigested) by (CRC) ICP-MS | This method describes the multi-elemental determination of trace elements by Collision/Reaction Cell (CRC) ICP-MS. The method measures ions produced by a radio-frequency inductively coupled plasma. Analyte species originating in a liquid are pneumatically nebulized and the resulting aerosol transported by argon gas into the plasma torch. The ions produced are entrained in the plasma gas and extracted, by means of a differentially pumped vacuum interface, into a mass spectrometer. The ions flow through a Collision/Reaction Cell (CRC) where depending on the mode of analysis either the ions passes through the cell unhindered (no gas - classical ICP-MS mode), or collisions/reactions chemistry takes place between the ions and a neutral gas such as hydrogen or helium. Any polyatomic ions which give interference on the analytes are removed. The ions are then sorted according to their mass-to-charge ratios by a quadrupole mass spectrometer. The ions are then sorted according to their mass-to-charge ratios by a quadrupole mass spectrometer. The ions transmitted through the quadrupole are quantified by a channel electron multiplier and the ion information processed by a data handling system. Although most polyatomic interference is removed, any additional interference must be assessed and valid corrections applied or the data flagged to indicate problems. Interference correction must include compensation for background ions contributed by the plasma gas, reagents, and constituents of the sample matrix. Instrument drift as well as suppressions or enhancements of instrument response caused by the sample matrix must be corrected for by the use of internal standards. | 2e-05 | Y | Y | Y | 441 | 8/12/2009 12:08:00 AM | 4/29/2024 | 6/7/2024 10:03:16 AM | 2009-09-08 |  | RDGIBSON |  | Concentration | mg/L | Y | (fl. conc.) |  | Molybdenum Total (fl. conc.) | 7439-98-7 | Metals - Total; | Metals - Total; EMS code: MO-T | Chemical | Total | LAB
MO-T | ug/g wet | Molybdenum Total | X215 | Hom,Nitr/Perchl,Grph Fur | Hom;Nitr/Perchl;Grph Fur | 0.05 | Y | Y | Y | 74 |  |  | 10/24/1996 10:01:25 AM | 1996-04-12 | 6/3/1998 6:51:42 PM | EMS |  | MassPerMass | µg/g | Y | (mass conc.) | (wet) | Molybdenum Total (wet) (mass conc.) | 7439-98-7 | Metals - Total; | Metals - Total; EMS code: MO-T | Chemical | Total | LAB
MO-T | mg/L | Molybdenum Total | 0042 | GI HNO3 Dig: ICP | Sample aliquot taken from unfiltered unpreserved General Ions bottle; HNO3 digestion; ICP analysis. | 0.004 | Y | Y | Y | 6239 | 9/20/1996 | 7/5/2019 3:21:00 PM | 7/12/2019 6:49:13 PM | 1996-04-12 | 6/3/1998 6:51:37 PM | EMS |  | Concentration | mg/L | Y | (fl. conc.) |  | Molybdenum Total (fl. conc.) | 7439-98-7 | Metals - Total; | Metals - Total; EMS code: MO-T | Chemical | Total | LAB
MO-T | mg/L | Molybdenum Total | X285 | HNO3 Dig:ICP Analysis | HNO3 Dig:ICP Analysis | 0.03 | Y | Y | Y | 687 | 8/10/1993 | 10/5/1993 | 4/30/2004 1:54:39 PM | 1996-04-12 | 6/3/1998 6:51:37 PM | EMS |  | Concentration | mg/L | Y | (fl. conc.) |  | Molybdenum Total (fl. conc.) | 7439-98-7 | Metals - Total; | Metals - Total; EMS code: MO-T | Chemical | Total | LAB
MO-T | ug/m3 | Molybdenum Total | 5312 | Tef:N/P Dig:ICP Analysis | A nitric/perchloric acid digestion if used to bring the metals into solution prior to ICP analysis. For Hi-Vol samples preliminary results are reported as mg/L and 24 hour loading results as ug/m3. | 0.002 | Y | Y | Y | 1185 | 11/7/1997 | 3/29/2000 | 9/28/2000 1:34:42 PM | 1996-10-13 | 6/3/1998 6:51:46 PM | EMS |  | AirConcentration | µg/m³ | Y | (air conc.) |  | Molybdenum Total (air conc.) | 7439-98-7 | Metals - Total; | Metals - Total; EMS code: MO-T | Chemical | Total | LAB
MO-T | mg/L | Molybdenum Total | X545 | FA,HNO3,ICPMS | Analytes nebulized followed by ion entrained in plasma gas and extracted by means of a differentially pumped vacuum interface into a mass spectrometer | 0.0001 | Y | Y | Y | 34 |  |  | 10/30/2014 8:31:18 AM | 2001-09-18 | 9/18/2001 1:36:23 PM | NPEPPIN | NPEPPIN | Concentration | mg/L | Y | (fl. conc.) |  | Molybdenum Total (fl. conc.) | 7439-98-7 | Metals - Total; | Metals - Total; EMS code: MO-T | Chemical | Total | LAB
MO-T | ug/g | Molybdenum Total | X551 | Metals in Biota (dry), total, acid digestion, ICP-SFMS | 0.1 gram of dried, homogenized biota is weighed in a microwave TFM vessel, and 8mL of Nitric Acid and 1mL of Hydrogen Peroxide are added. The sample is digested at 200 degrees C for 15 minutes in a closed vessel high pressure microwave oven. The sample is cooled, the volume adjusted to 100mL and introduced to an iductively coupled argon plasma-sector field mass spectrometer (ICP-SFMS) by pneumatic nebulization. Each respective element is measured at a specific mass to charge ratio, m/z value expressed in amu (atomic mass units), and at a specified instrument resolution. LR: Low Resolution ~300, MR: Mid Resolution ~4000 and HR: Hich Resolution ~10000. Ref: Method number 02-2705, National Laboratory for Environmental Testing, CCIW, Environment Canada, Burlington. Requested April 2003. | 0.01 | Y | Y | Y | 49 | 8/11/2015 | 10/30/2023 | 11/28/2023 2:21:33 PM | 2004-01-19 |  | UMULLER |  | MassPerMass | mg/kg | Y | (mass conc.) |  | Molybdenum Total (mass conc.) | 7439-98-7 | Metals - Total; | Metals - Total; EMS code: MO-T | Chemical | Total | LAB
MO-T | ug/g | Molybdenum Total | ICPM | Metals in Soil by CRC ICPMS | This analysis is carried out using procedures from CSR Analytical Method: "Strong Acid Leachable Metals (SALM) in Soil", BC Ministry of Environment, 26 June 2009, and procedures adapted from EPA Method 200.2. The sample is manually homogenized, dried at 60 degrees Celsius, sieved through a 63 um (230 mesh) sieve, and a representative subsample of the dry material is weighed. The sample is then digested at 95 degrees Celsius for 2 hours by block digester using concentrated nitric and hydrochloric acids. Instrumental analysis is by collision cell inductively coupled plasma - mass spectrometry (modifed from EPA Method 6020A). Method Limitation: This method is not a total digestion technique. It is a very strong acid digestion that is intended to dissolve those metals that may be environmentally available. By design, elements bound in silicate structures are not normally dissolved by this procedure as they are not usually mobile in the environment. | 0.1 | Y | Y | Y | 103 | 9/19/2017 | 9/24/2023 | 1/12/2024 11:16:52 AM | 2018-01-09 |  | JKROGH |  | MassPerMass | mg/kg | Y | (mass conc.) |  | Molybdenum Total (mass conc.) | 7439-98-7 | Metals - Total; | Metals - Total; EMS code: MO-T | Chemical | Total | LAB
MO-T | mg/L | Molybdenum Total | F082 | Aliquot,HNO3/HCL dig,ICPMS | Aliquot microwave digested with 3:1 HNO3/HCL. Analysis by ICPMS. For HiVol samples preliminary result are reported in mg/L and loading results are reported in ug/m3 | 0.001 | Y | Y | Y | 27689 | 1/24/2011 | 8/23/2024 | 8/26/2024 4:53:05 PM | 2000-06-13 | 8/15/2019 11:09:21 AM | NPEPPIN | KATFRASE | Concentration | mg/L | Y | (fl. conc.) |  | Molybdenum Total (fl. conc.) | 7439-98-7 | Metals - Total; | Metals - Total; EMS code: MO-T | Chemical | Total | LAB
MO-T | mg/L | Molybdenum Total | 1349 | Tot Metals in Seawater by CRC ICPMS (BC) | Tot Metals in Seawater by CRC ICPMS (BC) | 0.0001 | Y | Y | Y | 98 | 8/16/2019 | 12/5/2021 | 12/13/2021 12:34:16 PM | 2019-09-30 |  | KATFRASE |  | Concentration | mg/L | Y | (fl. conc.) |  | Molybdenum Total (fl. conc.) | 7439-98-7 | Metals - Total; | Metals - Total; EMS code: MO-T | Chemical | Total | LAB
MO-T | ug/g wet | Molybdenum Total | X560 | Metals in Biota (wet), total, acid digestion, ICP-MS | Five grams of wet, homogenized biota is weighed in a centrifuge tube, and 2.5 mL of Nitric Acid and 2.5 mL of Hydrochloric acid are added. The sample is digested at room temperature for one hour and then digested at 90-95 degrees C in a water bath or heat block for one more hour. The sample is cooled and the volume is made to 50 mL with DI. The digestate is then diluted at ten fold and introduced to a collision/reaction cell inductively coupled argon plasma mass spectrometer (CRC ICPMS) by pneumatic nebulization. Each respective element is measured at a specific mass to charge ratio, m/z value expressed in amu (atomic mass units). - Maxxam Analytics (BC). Added to EMS Dec 10, 2008. | 0.1 | Y | Y | Y | 541 | 11/15/2008 4:00:00 AM | 10/30/2023 | 11/28/2023 2:21:33 PM | 2008-12-11 | 3/11/2015 4:17:10 PM | RDGIBSON | EMS | MassPerMass | µg/g | Y | (mass conc.) | (wet) | Molybdenum Total (wet) (mass conc.) | 7439-98-7 | Metals - Total; | Metals - Total; EMS code: MO-T | Chemical | Total | LAB
MO-T | ug/g wet | Molybdenum Total | X552 | Metals in Biota (wet), total, acid digestion, ICP-SFMS | One gram of wet, homogenized biota is weighed in a microwave TFM vessel, and 8mL of Nitric Acid and 1mL of Hydrogen Peroxide are added. The sample is digested at 200 degrees C for 15 minutes in a closed vessel high pressure microwave oven. The sample is cooled, the volume adjusted to 100mL and introduced to an inductively coupled argon plasma-sector field mass spectrometer (ICP-SFMS) by pneumatic nebulization. Each respective element is measured at a specific mass to charge ratio, m/z value expressed in amu (atomic mass units), and at a specified instrument resolution. LR: Low Resolution ~300, MR: Mid Resolution ~4000. Ref: 02-2704, National Laboratory for Environmental Testing, CCIW, Environment Canada, Burlington. Requested April 2003. | 0.001 | Y | Y | Y | 37 | 8/9/2016 | 10/16/2023 | 11/28/2023 2:21:32 PM | 2004-01-19 | 4/19/2021 1:22:00 PM | UMULLER | EMS | MassPerMass | µg/g | Y | (mass conc.) | (wet) | Molybdenum Total (wet) (mass conc.) | 7439-98-7 | Metals - Total; | Metals - Total; EMS code: MO-T | Chemical | Total | LAB
MO-T | ug/g | Molybdenum Total | 0192 | Hom, Nitric/Perchlor,ICP | Sample already in a dry state so does not require drying before homogenization and nitric perchlorid digestion followed by ICP analysis. | 1 | Y | Y | Y | 708 |  |  | 10/24/1996 12:50:25 PM | 1996-04-12 | 6/3/1998 6:51:37 PM | EMS |  | MassPerMass | mg/kg | Y | (mass conc.) |  | Molybdenum Total (mass conc.) | 7439-98-7 | Metals - Total; | Metals - Total; EMS code: MO-T | Chemical | Total | LAB
MO-T | ug/g | Molybdenum Total | EQ02 | EQUIS (ug/g) | EQUIS (ug/g) | 1 | N | Y | Y | 790 |  |  | 10/23/1996 12:08:28 AM | 1996-04-12 | 6/3/1998 6:51:41 PM | EMS |  | MassPerMass | mg/kg | Y | (mass conc.) |  | Molybdenum Total (mass conc.) | 7439-98-7 | Metals - Total; | Metals - Total; EMS code: MO-T | Chemical | Total | LAB
MO-T | ug/g | Molybdenum Total | FS89 | Homo,HNO3 dig,ICP-AES | Homo;HNO3 dig;ICP-AES | 2 | Y | Y | Y | 4 | 1/5/2001 | 8/20/2001 12:00:00 PM | 10/23/2002 1:42:47 PM | 2001-03-08 |  | NPEPPIN |  | MassPerMass | mg/kg | Y | (mass conc.) |  | Molybdenum Total (mass conc.) | 7439-98-7 | Metals - Total; | Metals - Total; EMS code: MO-T | Chemical | Total | LAB
MO-T | ppm (S) | Molybdenum Total | 6042 | Aqua Regia by ICPMS | These were sediment/soils digested with Aqua-Regia (Nitric Acid & Hydrochloric Acids) and then analysis by ICPMS. Results would be reported in ppm s. | 0.01 | Y | Y | Y | 71 | 5/14/2022 | 7/27/2024 | 7/29/2024 9:13:08 AM | 2022-06-28 |  | EMS |  | DimensionlessRatio | ppm | N | (ratio) |  | Molybdenum Total (ratio) | 7439-98-7 | Metals - Total; | Metals - Total; EMS code: MO-T | Chemical | Total | LAB
MO-T | mg/L | Molybdenum Total | 0080 | Dig:Ext(Oxime):AA Anal. | Dig:Ext(Oxime):AA Anal. | 0.0005 | Y | Y | Y | 10 |  |  | 6/16/2000 4:19:26 PM | 1996-04-12 | 6/3/1998 6:51:37 PM | EMS |  | Concentration | mg/L | Y | (fl. conc.) |  | Molybdenum Total (fl. conc.) | 7439-98-7 | Metals - Total; | Metals - Total; EMS code: MO-T | Chemical | Total | LAB
MO-T | mg/L | Molybdenum Total | MWID | Metals in Water by ICPMS (digested) | Metals in Water by ICPMS (digested) Based on EPA 200.8 | 0.001 | Y | Y | Y | 36830 |  |  | 7/31/2024 1:46:41 PM | 2009-12-24 |  | RDGIBSON |  | Concentration | mg/L | Y | (fl. conc.) |  | Molybdenum Total (fl. conc.) | 7439-98-7 | Metals - Total; | Metals - Total; EMS code: MO-T | Chemical | Total | LAB
MO-T | ug/m3 | Molybdenum Total | 6040 | Aqua Regia Dig:ICP | Aqua regia digestion is used to bring the metals into solution prior to ICP analysis. For Hi-Vol samples, preliminary results are reported as mg/L and24 hr loadings as ug/m3 | 0.002 | Y | Y | Y | 437 | 2/13/2002 | 2/9/2005 | 2/23/2005 1:13:47 AM | 1998-05-05 | 6/3/1998 6:51:48 PM | NPEPPIN |  | AirConcentration | µg/m³ | Y | (air conc.) |  | Molybdenum Total (air conc.) | 7439-98-7 | Metals - Total; | Metals - Total; EMS code: MO-T | Chemical | Total | LAB
MO-T | mg/L | Molybdenum Total | 0240 | HNO3 Dig:Extr(Oxime)-AA | HNO3 Dig:Extr(Oxime)-AA | 0.0005 | Y | Y | Y | 329 |  |  | 9/18/1997 1:59:49 PM | 1996-04-12 | 6/3/1998 6:51:37 PM | EMS |  | Concentration | mg/L | Y | (fl. conc.) |  | Molybdenum Total (fl. conc.) | 7439-98-7 | Metals - Total; | Metals - Total; EMS code: MO-T | Chemical | Total | LAB
MO-T | ppm (S) | Molybdenum Total | 6041 | Aqua Regia by ICPMS | These were sediment/soils digested with Aqua-Regia (Nitric Acid & Hydrochloric Acids) and then analysis by ICPMS. Results would be reported in mg/kg and/or %. | 0.01 | Y | Y | Y | 21 | 10/13/2019 | 8/5/2020 | 8/10/2020 3:52:06 PM | 2019-09-06 | 10/21/2019 11:55:38 AM | KATFRASE | KATFRASE | DimensionlessRatio | ppm | N | (ratio) |  | Molybdenum Total (ratio) | 7439-98-7 | Metals - Total; | Metals - Total; EMS code: MO-T | Chemical | Total | LAB
MO-T | mg/L | Molybdenum Total | F048 | Acid dig,conc'n,nebul.,ICP-OES | This method provides an automated procedure for the Inductively Coupled Argon Plasma Optical Emission Spectrophotometric (ICP-OES) determination and quantitation of trace amounts of metals in surface, ground and precipitation waters after a manual acid digestion and concentration procedure and ultrasonic nebulisation. Organically-bound metals and metal complexes are converted to metal ions by digestion with nitric acid [method details says HCl]. To increase the signal detected by the ICP, the samples are concentrated 10 times. The liquid sample is pumped onto the piezoelectric transducer, where the ultrasonic action efficiently converts it to a fine, dense aerosol. The nebulizer carrier gas flow then transports the wet aerosol through a heated U-tube, where the solvent is vaporized. Subsequently, most of the solvent vapour is removed from the aerosol stream by passing the stream through a chilled condenser. The result is a dry, analyte-laden aerosol which is introduced to the plasma. There, the high temperature desolvates the aerosol and dissociates the sample material into individual atoms and ions, which are excited to emit light at wave lengths characteristic of the elements present. Dissolved samples are field filtered through 0.45 micron cellulose acetate filter prior to acid preservation. Samples for extractable analysis are acidified with 2 mL concentrated HNO3 per litre and allowed to stand a minimum of 24 hours before analysis. Clients providing these samples are advised to preserve samples at the collection site. | 0.0001 | Y | Y | Y | 6159 | 7/4/1997 2:36:00 PM | 4/23/2003 | 9/11/2019 8:46:38 AM | 1997-10-20 | 6/3/1998 6:51:47 PM | NPEPPIN |  | Concentration | mg/L | Y | (fl. conc.) |  | Molybdenum Total (fl. conc.) | 7439-98-7 | Metals - Total; | Metals - Total; EMS code: MO-T | Chemical | Total | LAB
MO-T | ppm (S) | Molybdenum Total | MS61 | ALS Minerals ME-MS61 | ALS Minerals ME-MS61. A prepared sample is digested with HNO3; HF; and HClO4 in a graphite heating block to near dryness. After cooling; the sample is boiled in diluted HCl. Bulk to volume with deionized water; mixed and analyzed by ICP-AES and ICP-MS for all elements. | 0.05 | Y | Y | Y | 0 |  |  |  | 2024-01-29 |  | EMS |  | DimensionlessRatio | ppm | N | (ratio) |  | Molybdenum Total (ratio) | 7439-98-7 | Metals - Total; | Metals - Total; EMS code: MO-T | Chemical | Total | LAB
MO-T | mg/L | Molybdenum Total | F068 | HNO3dig,TeflonTube/Autoclave,ICP-MS | This method is applicable to determiniation of dissolved, extractable, and total recoverable metals in ground waters, surface waters, drinking waters, and wastewaters. Dissolved, extractable, and total recoverable sample fractions are acidified to pH <1.5 with analytical grade nitric acid (HNO3) to prevent the metals from adsorbing to the container material or precipitating. Extractable and total recoverable water samples are collected in 250 ml acid washed polyethylene bottles and are preserved with Nitric acid (0.5 ml acid for 100 mLs of sample). All preserved samples are stored at room temperature and can be held for up to 6 months. Unfiltered, un-preserved samples are stored at 4C for no more than 30 days. For dissolved and extractable water fractions, no sample preparation is necessary. Samples are simply poured in autosampler tubes (i.e. 15 ml centrifuge tubes), and placed on the autosampler and analyzed. Samples for total recoverable metals are digested with Nitric acid in special Teflon tubes in an autoclave prior to analyses. | 5e-05 | Y | Y | Y | 1779 | 6/1/1999 1:33:00 PM | 7/4/2008 9:48:00 AM | 7/30/2024 6:32:12 AM | 1999-06-28 |  | NPEPPIN |  | Concentration | mg/L | Y | (fl. conc.) |  | Molybdenum Total (fl. conc.) | 7439-98-7 | Metals - Total; | Metals - Total; EMS code: MO-T | Chemical | Total | LAB
MO-T | ug/g | Molybdenum Total | 0190 | D&G:N/P Dig:ICP Anal. | Sample is dried and grind-mixed before nitric perchloric digestion followed by ICP analysis. | 1 | Y | Y | Y | 948 |  |  | 10/24/1996 3:06:16 PM | 1996-04-12 | 6/3/1998 6:51:41 PM | EMS |  | MassPerMass | mg/kg | Y | (mass conc.) |  | Molybdenum Total (mass conc.) | 7439-98-7 | Metals - Total; | Metals - Total; EMS code: MO-T | Chemical | Total | LAB
MO-T | mg/L | Molybdenum Total | X028 | FA,HNO3,AA Analysis | FA;HNO3;AA Analysis | 0.0005 | Y | Y | Y | 5084 |  |  | 8/28/2018 3:14:04 PM | 1996-04-12 | 6/3/1998 6:51:37 PM | EMS |  | Concentration | mg/L | Y | (fl. conc.) |  | Molybdenum Total (fl. conc.) | 7439-98-7 | Metals - Total; | Metals - Total; EMS code: MO-T | Chemical | Total | LAB
MO-T | ug/g | Molybdenum Total | FS01 | D&G,Nitric/HCl:ICP | Samples are dried at 60 C, then either sieved to -100mesh (150 um) and rolled to homogenize, or finely ground with a mortar and pestle or a shatter box laboratory mill and rolled (depending on the submitter's requirements). The samples are then weighed into Teflon digestion liners and digested with HNO3 and HCl in a microwave oven. Samples are cooled, volumized, and allowed to settle overnight before being decanted. Analysis is then performed via Inductively Coupled Argon Plasma Atomic Emission Spectrometry (ICP-AES) (Simultaneous multi element analysis). In ICP Spectrometry, the liquid sample is pumped to the nebulizer, where it is convertd into an aerosol. the nebulizer carrier gas flow than transports the wet aerosol past the impact bead, where the large droplets condense out and flow to the waste container. The remaining fine aerosol is carried into the plasma. There, the high temperature desolvates the aerosol and dissociates the sample material into individual atoms and ions, which are excited to emit light at wavelengths characteristic of the elements present. A diffraction grating in the spectrometer diffracts the various wavelengths of light through exit slits to the photo-multiplier tubes, which convert the light energy to electronic signals. | 2 | Y | Y | Y | 2641 | 8/30/1996 1:52:00 PM | 10/10/2007 | 5/24/2008 1:15:02 AM | 1996-11-20 | 6/3/1998 6:51:46 PM | MPAK |  | MassPerMass | mg/kg | Y | (mass conc.) |  | Molybdenum Total (mass conc.) | 7439-98-7 | Metals - Total; | Metals - Total; EMS code: MO-T | Chemical | Total | LAB
MO-T | ug/g wet | Molybdenum Total | BPSY | MET-WET-BPSY-CCMS-VA | Metals in Animal Tissue and Vegetation (Biota) - Prescriptive. EPA 200.3/6020B This method is conducted following British Columbia Lab Manual method "Metals in Animal Tissue and Vegetation (Biota) - Prescriptive". Tissue samples are digested with nitric and hydrochloric acids; in combination with addition of hydrogen peroxide. Instrumental analysis is by collision cell inductively coupled plasma - mass spectrometry (modified from EPA Method 6020B). | 0.008 | Y | Y | Y | 2 | 11/7/2019 | 11/7/2019 | 6/10/2021 1:57:10 PM | 2021-06-04 |  | EMS |  | MassPerMass | µg/g | Y | (mass conc.) | (wet) | Molybdenum Total (wet) (mass conc.) | 7439-98-7 | Metals - Total; | Metals - Total; EMS code: MO-T | Chemical | Total | LAB
MO-T | mg/L | Molybdenum Total | METN | MET-T-NP-CCMS-VA | Ultra trace metals in water are analyzed by CRC ICPMS, based on US EPA Method 6020B (July 2014). The detection limits provided can only be met for undigested samples. This procedure is intended for colorless, non-turbid, acid-preserved water samples (i.e. pristine water samples), having turbidity < 1 NTU and no odor. Where turbidity exceeds 1 NTU, and/or the sample is colored and has an odor, results may be biased low compared to true Total Metals concentrations. ALS recommends that turbidity analysis be requested on samples submitted for this test to aid with interpretation of results. Where turbidity is <1NTU, undigested metals are equivalent to total metals concentrations. | 5e-05 | Y | Y | Y | 3648 | 4/3/2012 | 8/19/2024 | 8/26/2024 1:03:06 PM | 2019-08-20 |  | KATFRASE |  | Concentration | mg/L | Y | (fl. conc.) |  | Molybdenum Total (fl. conc.) | 7439-98-7 | Metals - Total; | Metals - Total; EMS code: MO-T | Chemical | Total | LAB
MO-T | mg/L | Molybdenum Total | X234 | Aqua Reg. Dig,AA | Aqua Reg. Dig;AA | 0.0005 | Y | Y | Y | 248 |  |  | 10/24/1996 12:12:10 PM | 1996-04-12 | 6/3/1998 6:51:37 PM | EMS |  | Concentration | mg/L | Y | (fl. conc.) |  | Molybdenum Total (fl. conc.) | 7439-98-7 | Metals - Total; | Metals - Total; EMS code: MO-T | Chemical | Total | LAB
MO-T | mg/L | Molybdenum Total | F121 | Acid,ICPMS | Acidified, no digestion, ICPMS | 5e-05 | Y | Y | Y | 8440 | 3/12/2010 7:57:00 PM | 8/15/2024 9:10:00 PM | 8/26/2024 8:06:03 AM | 2002-07-22 | 7/16/2012 3:29:56 PM | UMULLER | RDGIBSON | Concentration | mg/L | Y | (fl. conc.) |  | Molybdenum Total (fl. conc.) | 7439-98-7 | Metals - Total; | Metals - Total; EMS code: MO-T | Chemical | Total | LAB
MO-T | ug/g wet | Molybdenum Total | MTIW | MET-WET-CCMS-MID-VA | Metals in Tissue by CRC ICPMS (WET) Methodology: This method is conducted following British Columbia Lab Manual method "Metals in Animal Tissue and Vegetation (Biota) - Prescriptive". Tissue samples are homogenized and sub-sampled prior to hotblock digestion with nitric and hydrochloric acids, in combination with addition of hydrogen peroxide. Instrumental analysis is by collision cell inductively coupled plasma - mass spectrometry (modified from EPA Method 6020A). Method Limitation: This method employs a strong acid/peroxide digestion, and is intended to provide a conservative estimate of bio-available metals. Near complete recoveries are achieved for most toxicologically important metals, but elements associated with recalcitrant minerals may be only partially recovered. | 0.008 | Y | Y | Y | 24 | 11/25/2018 | 12/6/2018 | 7/21/2021 2:55:07 PM | 2018-12-24 |  | EMS |  | MassPerMass | µg/g | Y | (mass conc.) | (wet) | Molybdenum Total (wet) (mass conc.) | 7439-98-7 | Metals - Total; | Metals - Total; EMS code: MO-T | Chemical | Total | LAB
MO-T | mg/L | Molybdenum Total | X528 | unknown method,plasma emission spec. | unknown method;plasma emission spec |  | Y | Y | Y | 7 |  |  | 8/17/1999 2:50:49 AM | 1999-07-21 |  | NPEPPIN |  | Concentration | mg/L | Y | (fl. conc.) |  | Molybdenum Total (fl. conc.) | 7439-98-7 | Metals - Total; | Metals - Total; EMS code: MO-T | Chemical | Total | LAB
MO-T | mg/L | Molybdenum Total | EQ01 | EQUIS (mg/L) | EQUIS (mg/L) | 0.0005 | N | Y | Y | 2959 |  |  | 10/24/1996 3:15:56 PM | 1996-04-12 | 6/3/1998 6:51:37 PM | EMS |  | Concentration | mg/L | Y | (fl. conc.) |  | Molybdenum Total (fl. conc.) | 7439-98-7 | Metals - Total; | Metals - Total; EMS code: MO-T | Chemical | Total | LAB
MO-T | ug/g | Molybdenum Total | 0031 | ICP | ICP | 4 | Y | Y | Y | 4 |  |  | 10/23/1996 11:03:49 PM | 1996-10-11 | 6/3/1998 6:51:46 PM | NPEPPIN |  | MassPerMass | mg/kg | Y | (mass conc.) |  | Molybdenum Total (mass conc.) | 7439-98-7 | Metals - Total; | Metals - Total; EMS code: MO-T | Chemical | Total | LAB
MO-T | mg/dm2/d | Molybdenum Total | DUST | HNO3 dig.(microwave or hot block),ICPMS | HNO3 dig.(microwave or hot block);ICPMS |  | Y | Y | Y | 705 | 7/19/2016 1:11:00 PM | 7/31/2024 | 7/31/2024 11:23:12 AM | 2008-02-07 |  | RDGIBSON |  | MassFlux | mg/dm²/d | Y | (mass flux) |  | Molybdenum Total (mass flux) | 7439-98-7 | Metals - Total; | Metals - Total; EMS code: MO-T | Chemical | Total | LAB
MO-T | mg/L | Molybdenum Total | F033 | Acid Dig,Centrifuge,ICPMS | For the determination of total recoverable elements in water or waste water, take 100 ml aliquot from a well mixed, acid preserved, sample containing not more than 0.25% (w/v) total solids and transfer to a 250 ml Griffin beaker. (If total solids are greater than 0.25% reduce the size of the aliquot by a proportionate amount). Add 1 ml of conc. nitric acid and 0.5 ml conc. hydrochloric acid. Heat on a hot plate until the volume has been reduced to 15-20ml, ensuring that the sample does not boil. Allow to cool and transfer to a 50 ml volumetric flask. Dilute to volume with ASTM type I water. Centrifuge the sample or allow to stand overnight to separate insoluble material. Prior analysis, pipette 20 ml into a 50 ml volumetric flask and dilute to volume with ASTM type I water. Analysis should be performed within seventy two hours of the completed preparation. Note that different laboratories may have very different field preservation, digestion procedures and detection limits for this method, so these should be documented. | 0.001 | Y | Y | Y | 37010 | 11/20/1997 9:32:00 AM | 8/22/2024 12:35:00 PM | 8/25/2024 9:47:36 PM | 1998-05-14 | 6/3/1998 6:51:49 PM | NPEPPIN |  | Concentration | mg/L | Y | (fl. conc.) |  | Molybdenum Total (fl. conc.) | 7439-98-7 | Metals - Total; | Metals - Total; EMS code: MO-T | Chemical | Total | LAB
MO-T | ug/g | Molybdenum Total | MTID | MET-DRY-CCMS-VA | Metals in Tissue by CRC ICPMS (DRY) Methodology: This method is conducted following British Columbia Lab Manual method "Metals in Animal Tissue and Vegetation (Biota) - Prescriptive". Tissue samples are homogenized and sub-sampled prior to hotblock digestion with nitric and hydrochloric acids, in combination with addition of hydrogen peroxide. Instrumental analysis is by collision cell inductively coupled plasma - mass spectrometry (modified from EPA Method 6020A). Method Limitation: This method employs a strong acid/peroxide digestion, and is intended to provide a conservative estimate of bio-available metals. Near complete recoveries are achieved for most toxicologically important metals, but elements associated with recalcitrant minerals may be only partially recovered. | 0.02 | Y | Y | Y | 122 | 8/13/2014 | 9/13/2022 | 10/26/2022 7:14:15 AM | 2015-03-09 | 3/11/2015 4:12:05 PM | RDGIBSON | EMS | MassPerMass | mg/kg | Y | (mass conc.) |  | Molybdenum Total (mass conc.) | 7439-98-7 | Metals - Total; | Metals - Total; EMS code: MO-T | Chemical | Total | LAB
MO-T | mg/L | Molybdenum Total | X072 | FA,HNO3,Graph.Furnace,AA | Samples are field-acidified with HNO3; HNO3 digestion; analysis by GFAA. | 0.001 | Y | Y | Y | 7 | 10/9/1996 2:13:00 PM | 12/5/1996 1:23:00 PM | 5/11/1998 3:57:23 PM | 1996-04-12 | 6/3/1998 6:51:37 PM | EMS |  | Concentration | mg/L | Y | (fl. conc.) |  | Molybdenum Total (fl. conc.) | 7439-98-7 | Metals - Total; | Metals - Total; EMS code: MO-T | Chemical | Total | LAB
MO-T | ug/g | Molybdenum Total | X230 | Hom,HNO3/HCl,ICP | The sample is dried, ground to pass 100 mesh, and digested with a nitric-hydrochloric acid mixture to solubilize the solid matter and to remove the organic material by oxidation and volatilization. The completed digests are submitted to metals instrumentation personnel for metals analysis by Inductively Coupled Atomic Emission Spectrophotometry (ICP-AES), or hydride analysis by atomic absorption. | 1 | Y | Y | Y | 264 | 2/12/2002 | 12/30/2005 | 2/3/2006 1:13:13 AM | 1996-04-12 | 6/3/1998 6:51:37 PM | EMS |  | MassPerMass | mg/kg | Y | (mass conc.) |  | Molybdenum Total (mass conc.) | 7439-98-7 | Metals - Total; | Metals - Total; EMS code: MO-T | Chemical | Total | LAB
MO-T | mg/L | Molybdenum Total | X349 | FA:HNO3:ICP | Samples are field-acidified with HNO3; HNO3 digestion; analysis by ICP/AES. | 0.005 | Y | Y | Y | 6464 |  |  | 8/26/2024 2:41:32 PM | 1996-04-12 | 6/3/1998 6:51:37 PM | EMS |  | Concentration | mg/L | Y | (fl. conc.) |  | Molybdenum Total (fl. conc.) | 7439-98-7 | Metals - Total; | Metals - Total; EMS code: MO-T | Chemical | Total | LAB
MO-T | ug/g | Molybdenum Total | 0196 | D&G: LTA/HCl: ICP | D&G: LTA/HCl: ICP | 1 | Y | Y | Y | 27 |  |  | 10/24/1996 2:08:12 PM | 1996-04-12 | 6/3/1998 6:51:37 PM | EMS |  | MassPerMass | mg/kg | Y | (mass conc.) |  | Molybdenum Total (mass conc.) | 7439-98-7 | Metals - Total; | Metals - Total; EMS code: MO-T | Chemical | Total | LAB
MO-T | mg/L | Molybdenum Total | F038 | HNO3/HCl dig,ICP-AES | Dissolved and extractable water samples are ready for analysis once collected. Total samples are autoclaved with HNO3 and hydrochloric acid (HCI) in capped Teflon vials. Analysis is performed via Inductively Coupled Argon Plasma Atomic Emission Spectrometry (ICP-AES)(simultaneous multielement analysis), Graphite Furnace Atomic Absorption Spectrometry (GF-AAS) (low level analysis) and/or Flame Atomic Absorption Spectrometry (F-AAS). | 0.01 | Y | Y | Y | 5426 | 8/1/1996 6:04:00 PM | 5/6/2008 | 9/11/2019 8:47:42 AM | 1997-05-08 | 6/3/1998 6:51:46 PM | MPAK |  | Concentration | mg/L | Y | (fl. conc.) |  | Molybdenum Total (fl. conc.) | 7439-98-7 | Metals - Total; | Metals - Total; EMS code: MO-T | Chemical | Total | LAB
MO-T | mg/L | Molybdenum Total | 0200 | 20%A.R.Dig:AA Anal. | 20%A.R.Dig:AA Anal. | 0.1 | Y | Y | Y | 2 |  |  | 10/24/1996 4:16:44 AM | 1996-04-12 | 6/3/1998 6:51:41 PM | EMS |  | Concentration | mg/L | Y | (fl. conc.) |  | Molybdenum Total (fl. conc.) | 7439-98-7 | Metals - Total; | Metals - Total; EMS code: MO-T | Chemical | Total | LAB
MTBE | ug/g | Met.Tert.Butyl Ether | MS02 | Purge&Trap GC/MS (VOC-S) | Purge&Trap GC/MS (VOC-S) |  | Y | Y | Y | 16 | 1/14/2022 | 10/28/2022 5:04:00 PM | 1/11/2024 1:07:08 PM | 2007-08-02 |  | RDGIBSON |  | MassPerMass | mg/kg | Y | (mass conc.) |  | Met.Tert.Butyl Ether (mass conc.) | 1634-04-4 | Volatile Organic Compounds (VOCs); | Volatile Organic Compounds (VOCs); EMS code: MTBE | Chemical |  | LAB
MTBE | ug/g | Met.Tert.Butyl Ether | MEOH | MeOH extrac,GC/FID,nC6-nC10 | MeOH extrac;GC/FID;nC6-nC10. EPA8260B; 5021; 5035; BC MOE | 0.2 | Y | Y | Y | 49 | 5/21/2015 | 8/11/2019 | 8/13/2019 10:56:08 AM | 2015-05-26 |  | RDGIBSON |  | MassPerMass | mg/kg | Y | (mass conc.) |  | Met.Tert.Butyl Ether (mass conc.) | 1634-04-4 | Volatile Organic Compounds (VOCs); | Volatile Organic Compounds (VOCs); EMS code: MTBE | Chemical |  | LAB
MTBE | mg/L | Met.Tert.Butyl Ether | X239 | Gas Chromatography | Gas Chromatography | 0.005 | Y | Y | Y | 628 | 3/21/2002 | 3/21/2002 | 6/14/2004 11:21:49 AM | 1996-04-12 | 6/3/1998 6:51:36 PM | EMS |  | Concentration | mg/L | Y | (fl. conc.) |  | Met.Tert.Butyl Ether (fl. conc.) | 1634-04-4 | Volatile Organic Compounds (VOCs); | Volatile Organic Compounds (VOCs); EMS code: MTBE | Chemical |  | LAB
MTBE | mg/L | Met.Tert.Butyl Ether | F105 | Static Headspace/GCMS | Static Headspace/GCMS | 0.002 | Y | Y | Y | 417 | 6/5/2000 2:49:00 PM | 7/6/2024 9:57:00 PM | 7/22/2024 3:17:33 PM | 2001-02-23 |  | NPEPPIN |  | Concentration | mg/L | Y | (fl. conc.) |  | Met.Tert.Butyl Ether (fl. conc.) | 1634-04-4 | Volatile Organic Compounds (VOCs); | Volatile Organic Compounds (VOCs); EMS code: MTBE | Chemical |  | LAB
MTBE | mg/L | Met.Tert.Butyl Ether | MS04 | Static Headspace GC/MS (Water) | (Water) A static headspace procedure is used as a means of driving the analytes out of the matrix into a headspace. Once equilibrium is achieved, a fixed volume of headspeace is taken for subsequent determination by Gas Chromatography - Mass Spectroscopy (GC/MS) using Selected Ion Monitoring. | 0.004 | Y | Y | Y | 61 | 8/13/2009 | 6/28/2024 12:29:00 AM | 7/9/2024 3:31:02 PM | 2009-12-23 |  | RDGIBSON |  | Concentration | mg/L | Y | (fl. conc.) |  | Met.Tert.Butyl Ether (fl. conc.) | 1634-04-4 | Volatile Organic Compounds (VOCs); | Volatile Organic Compounds (VOCs); EMS code: MTBE | Chemical |  | LAB
MTBE | ug/g | Met.Tert.Butyl Ether | FS04 | Static Headspace:HRGC/LRMS | Static Headspace anaylsis is a method for the determination of volatile compounds in liquids and/or solids by analysing the vapour phase of a sample in a closed system; in this case a sealed vial. Static Headspace is sometimes referred to as Equilibrium Headspace because the vapour phase above the sample in a sealed vial is brought into thermodynamic equilibrium with the sample prior to analysis. Volatile organic compounds are thermally extracted from the sample matrix by Static Headspace sampling and introduced directly to a thick-film megabore GC column. Compounds eluting from the GC column are introduced into a Mass Spectrometer and target compounds are identified by a comparison of both retention times and mass spectra with those of reference standards stored in a computerized data system. Reference retention times and mass spectra for analytes are obtained by the measurement of calibration standards under precisely the same extraction conditions used for samples. The concentration of each identified component is measured by relating the response of the sample quantitation mass ions to the response of the quantitation mass ions of a series of external standards. A deuterated surrogate recovery standard is added to provide information on matrix effects. | 0.1 | Y | Y | Y | 18 | 3/4/2013 6:06:00 PM | 3/7/2013 3:50:00 AM | 3/19/2013 1:15:55 PM | 2010-11-16 |  | RDGIBSON |  | MassPerMass | mg/kg | Y | (mass conc.) |  | Met.Tert.Butyl Ether (mass conc.) | 1634-04-4 | Volatile Organic Compounds (VOCs); | Volatile Organic Compounds (VOCs); EMS code: MTBE | Chemical |  | LAB
MTBE | mg/L | Met.Tert.Butyl Ether | F007 | Static Headspace:HRGC/LRMS | Static headspace analysis is a method for the determination of volatile compounds in liquids and/or solids by analysing the vapour phase of a sample in a closed system; in this case a sealed vial. Static Headspace is sometimes referred to as Equilibrium Headspace because the vapour phase above the sample in a sealed vial is brought into thermodynamic equilibrium with the sample prior to analysis. Volatile organic compounds are thermally extracted from the sample matrix by Static Headspace sampling and introduced directly to a thick-film megabore GC column. Compounds eluting from the GC column are introduced into a Mass Spectrometer and target compounds are identified by a comparison of both retention times and mass spectra with those of reference standards stored in a computerized data system. Reference retention times and mass spectra for analytes are obtained by the measurement of calibration standards under precisely the same extraction conditions used for samples. The concentration of each identified component is measured by relating the response of the sample quantitation mass ions to the response of the quantitation mass ions of a series of external standards. A deuterated surrogate recovery standard is added to provide information on matrix effects. | 0.0005 | Y | Y | Y | 868 | 11/23/2013 | 8/12/2024 | 8/15/2024 11:32:04 AM | 2014-01-14 |  | RDGIBSON |  | Concentration | mg/L | Y | (fl. conc.) |  | Met.Tert.Butyl Ether (fl. conc.) | 1634-04-4 | Volatile Organic Compounds (VOCs); | Volatile Organic Compounds (VOCs); EMS code: MTBE | Chemical |  | LAB
MTBE | mg/L | Met.Tert.Butyl Ether | MS01 | Purge&Trap GC/MS (VOC-W) | Purge&Trap GC/MS (VOC-W) | 0.001 | Y | Y | Y | 403 | 5/4/2007 | 6/4/2024 10:24:00 AM | 6/13/2024 5:01:27 PM | 2007-08-02 | 8/9/2007 5:29:52 PM | RDGIBSON | RDGIBSON | Concentration | mg/L | Y | (fl. conc.) |  | Met.Tert.Butyl Ether (fl. conc.) | 1634-04-4 | Volatile Organic Compounds (VOCs); | Volatile Organic Compounds (VOCs); EMS code: MTBE | Chemical |  | LAB
MU34 | mg/L | Methyl urea; 3-3-(3;4-dichlorophenyl)-1l)-1- [DCPMU] | E751 | E751D Pesticides in Water by LC-MS/MS (POS1 List) | E751D Pesticides in Water by LC-MS/MS (POS1 List) (MOECC E3501) | 5e-05 | Y | Y | Y | 39 | 12/9/2021 | 12/31/2021 | 2/15/2022 9:50:05 AM | 2021-12-14 |  | EMS |  | Concentration | mg/L | Y | (fl. conc.) |  | Methyl urea; 3-3-(3;4-dichlorophenyl)-1l)-1- [DCPMU] (fl. conc.) | 3567-62-2 | Volatile Organic Compounds (VOCs); | Volatile Organic Compounds (VOCs); EMS code: MU34 | Chemical |  | LAB
N001 | ug/g | Naled | FS16 | DCM/MEOH extr'n,DCM,conc'n,HRGC-PFPD/N,P det'n | Sediment samples are shaken with a mixture of Dichloromethane and Methanol, filtered, then extracted with Dichloromethane. The extract is then concentrated and analysed using High Resolution Gas Chromatography with Pulsed Flame Photometric Detection. Confirmation, if required, is by High Resolution Gas Chromatography with Nitrogen Phosphorus Detection. | 0.1 | Y | Y | Y | 1 | 4/24/2002 4:00:00 PM | 4/24/2002 4:00:00 PM | 2/12/2004 1:02:37 AM | 1999-06-28 |  | NPEPPIN |  | MassPerMass | mg/kg | Y | (mass conc.) |  | Naled (mass conc.) | 300-76-5 | Pesticides; | Pesticides; EMS code: N001 | Chemical |  | LAB
N001 | mg/L | Naled | ZN01 | DCM+Ac.Extr:GC/MS | DCM+Ac.Extr:GC/MS | 0.0001 | Y | Y | Y | 4 | 10/3/2002 4:45:00 PM | 10/15/2003 5:15:00 PM | 3/24/2004 1:04:40 AM | 1999-11-16 |  | NPEPPIN |  | Concentration | mg/L | Y | (fl. conc.) |  | Naled (fl. conc.) | 300-76-5 | Pesticides; | Pesticides; EMS code: N001 | Chemical |  | LAB
N001 | ug/g | Naled | ZN03 | DCM+Ac.Extr:GC/MS | DCM+Ac.Extr:GC/MS EPA 8270E (mod) | 0.01 | Y | Y | Y | 38 | 4/20/2002 10:25:00 AM | 10/16/2002 3:20:00 PM | 2/26/2004 1:05:29 AM | 1999-11-16 |  | NPEPPIN |  | MassPerMass | mg/kg | Y | (mass conc.) |  | Naled (mass conc.) | 300-76-5 | Pesticides; | Pesticides; EMS code: N001 | Chemical |  | LAB
N001 | mg/L | Naled | F053 | DCM extr'n,conc'n,HGRC-PFPD/N,P det'n | Effluent samples are extracted with Dichloromethane. The extract is then concentrated and analysed using High Resolution Gas Chromatography with Pulsed Flame Photometric Dectection. Confirmation, if required, is by High Resolution Gas Chromatography with Nitrogen Phosphorus Detection. | 0.0001 | Y | Y | Y | 31 | 3/3/2000 7:36:00 AM | 1/22/2002 10:26:00 AM | 6/18/2003 2:53:48 PM | 1999-06-28 |  | NPEPPIN |  | Concentration | mg/L | Y | (fl. conc.) |  | Naled (fl. conc.) | 300-76-5 | Pesticides; | Pesticides; EMS code: N001 | Chemical |  | LAB
N005 | mg/L | Neoabietic Acid | P030 | Resin Acids | The water sample is acidified and extracted with dichloro-methane. The raw extract is concentrated and derivatized with diazomethane to produce the corresponding methyl ester derivatives. (Suitable substitute derivatives may be produced by other available techniques.) If required, the extracts are cleaned up by Florisil column chromatography. The derivatives are analyzed by gas chromatography with flame ionization detection. | 0.001 | Y | Y | Y | 480 | 8/26/2014 | 8/15/2017 | 8/15/2017 4:37:09 PM | 1996-04-12 | 6/3/1998 6:51:37 PM | EMS |  | Concentration | mg/L | Y | (fl. conc.) |  | Neoabietic Acid (fl. conc.) | 471-77-2 | Organics; | Organics; EMS code: N005 | Chemical |  | LAB
N005 | mg/L | Neoabietic Acid | P035 | DCM-DIAZO-FLOR | DCM-DIAZO-FLOR | 0.0005 | Y | Y | Y | 22 |  |  | 10/24/1996 12:44:24 PM | 1996-04-12 | 6/3/1998 6:51:37 PM | EMS |  | Concentration | mg/L | Y | (fl. conc.) |  | Neoabietic Acid (fl. conc.) | 471-77-2 | Organics; | Organics; EMS code: N005 | Chemical |  | LAB
N005 | mg/L | Neoabietic Acid | E677 | Resins & Fatty Acids are methylated then analyzed by GCMS | EPA 8270E; AB Env. 129; NCASI RA/FA-85.02 (mod) Resins & Fatty Acids are methylated then analyzed by GCMS. Total fatty acids is the sum of arachidic acid; behenic acid; lauric acid; lignoceric acid; linoleic acid; linolenic acid; myristic acid; oleic acid; palmitic acid; and stearic acid. Total resin acids is the sum of abietic acid; 12-chlorodehydroabietic acid; 14-chlorodehydroabietic acid; dehydroabietic acid; 12;14-dichlorodehydroabietic acid; isopimaric acid; levopimaric acid; neoabietic acid; palustric acid; pimaric acid and sandaracopimaric acid. Total resin acids (BC WWQG) is the sum of abietic acid; neoabietic acid; pimaric acid; isopimaric acid; and sandaracopimaric acid. | 0.001 | Y | Y | Y | 12 | 11/28/2022 | 6/15/2024 | 7/8/2024 11:10:02 AM | 2021-12-30 |  | EMS |  | Concentration | mg/L | Y | (fl. conc.) |  | Neoabietic Acid (fl. conc.) | 471-77-2 | Organics; | Organics; EMS code: N005 | Chemical |  | LAB
N005 | mg/L | Neoabietic Acid | F009 | pH 9,Diethyl Ether:HRGC/LRMS | Effluent samples are first adjusted to pH 9 and then extracted with Diethyl Ether. The extracts are rotary evaporated and the resulting residues, containing the Resin Acids and other solvent soluble organics are then redissolved in Methanol, transferred to small glass vials, methylated with Diazomethane then evaporated to dryness under Nitrogen. [Methylation Reaction R-COOH(Resin/Fatty Acids) + CH2N2(Diazomethane) = R-COOCH3(Methyl ester) + N2) The residue is re-dissolved in Methyl t-Butyl Ether (MTBE) and injected on a High Resolution Gas Chromatograph/Low Resolution Mass Spectrometer (HRGC/LRMS) system for analysis. LRMS data is acquired in Total Ion Mode (TIM) and quantitation is performed using selected mass ions. Positive identification of the various Resin Acids is made on the basis of a comparison of absolute retention times to those of the external standards and computer based mass spectral mass/intensity library matching routines. | 0.005 | Y | Y | Y | 192 | 12/27/1996 4:16:00 PM | 5/13/2003 | 4/1/2004 1:08:46 AM | 1996-09-16 | 6/3/1998 6:51:44 PM | MPAK |  | Concentration | mg/L | Y | (fl. conc.) |  | Neoabietic Acid (fl. conc.) | 471-77-2 | Organics; | Organics; EMS code: N005 | Chemical |  | LAB
N005 | mg/L | Neoabietic Acid | X239 | Gas Chromatography | Gas Chromatography | 0.001 | Y | Y | Y | 12 |  |  | 10/23/1996 11:56:37 PM | 1996-04-12 | 6/3/1998 6:51:37 PM | EMS |  | Concentration | mg/L | Y | (fl. conc.) |  | Neoabietic Acid (fl. conc.) | 471-77-2 | Organics; | Organics; EMS code: N005 | Chemical |  | LAB
N005 | mg/L | Neoabietic Acid | RFAS | Separation of resin and fatty acids (RFAs) in water by GC/MS | Scientific Principle The separation of resin and fatty acids (RFAs) in water samples is performed using gas chromatography/mass spectrometry (GC/MS) in the Selected Ion Monitoring (SIM) mode. The separation is based on compound size; boiling point and polarity. Analyte concentration is calculated using the internal standard method. RFA compounds are extracted from water samples using a liquid-liquid solvent extraction with DCM. Extraction occurs after the sample is brought to a pH of ~ 2 and placed in a rotary tumbler (at approx. 40 rpm) for 1 hour. The resulting extract is transferred to a vial; dried with acidified baked sodium sulfate; methylated with diazomethane and concentrated to 1mL under nitrogen prior to analysis. The resulting extracts are vialed and fortified with labeled internal standard compounds prior to analysis by gas chromatography with a (50%-Phenyl)-methylpolysiloxane (DB-17) column and a mass spectrometer detector in SIM Mode. | 0.006 | Y | Y | Y | 1 | 9/17/2021 7:49:00 PM | 9/17/2021 7:49:00 PM | 2/1/2022 2:38:14 PM | 2022-01-19 |  | EMS |  | Concentration | mg/L | Y | (fl. conc.) |  | Neoabietic Acid (fl. conc.) | 471-77-2 | Organics; | Organics; EMS code: N005 | Chemical |  | LAB
N005 | ug/g | Neoabietic Acid | FS05 | MeOH,DCM extr:HRGC/LRMS | Sediment samples are extracted by shaking with Methanol in a screw cap centrifuge bottle. The extracts are rotary evaporated and the resulting residues, containing the Resin Acids and other solvent soluble organics, are then redissolved in Methanol, transferred to small glass vials, methylated with Diazomethane then evaporated to dryness under Nitrogen. [Methylation Reaction R-COOH(Resin/Fatty Acids) + CH2N2 (Diazomethane) = R-COOCH3 (Methyl ester) + N2] The residue is re-dissolved in Methyl t-Butyl Ether (MTBE) and injected on a High Resolution Gas Chromatograph / Low Resolution Mass Spectrometer (HRGC/LRMS) system for analysis. LRMS data is acquired in Total Ion Mode (TIM) and quantitation is performed using selected mass ions. Positive identification of the various Resin Acids is made on the basis of a comparison of absolute retention times to those of the external standards and computer based mass spectral mass/intensity library matching routines. | 0.2 | Y | Y | Y | 3 | 9/14/2001 3:29:00 PM | 12/13/2002 12:55:00 PM | 3/27/2004 1:09:21 AM | 1996-09-16 | 6/3/1998 6:51:44 PM | MPAK |  | MassPerMass | mg/kg | Y | (mass conc.) |  | Neoabietic Acid (mass conc.) | 471-77-2 | Organics; | Organics; EMS code: N005 | Chemical |  | LAB
N005 | mg/L | Neoabietic Acid | P033 | Resin&Fatty Acids,GC/MS | Resin&Fatty Acids;GC/MS | 0.001 | Y | Y | Y | 30 |  |  | 6/7/2024 5:30:18 PM | 1996-04-12 | 6/3/1998 6:51:37 PM | EMS |  | Concentration | mg/L | Y | (fl. conc.) |  | Neoabietic Acid (fl. conc.) | 471-77-2 | Organics; | Organics; EMS code: N005 | Chemical |  | LAB
N008 | mg/L | Norflurazon | ZN01 | DCM+Ac.Extr:GC/MS | DCM+Ac.Extr:GC/MS | 0.0005 | Y | Y | Y | 1 | 7/29/2005 7:10:00 AM | 7/29/2005 7:10:00 AM | 6/2/2009 12:01:08 PM | 2003-06-17 |  | UMULLER |  | Concentration | mg/L | Y | (fl. conc.) |  | Norflurazon (fl. conc.) | 27314-13-2 | Pesticides; | Pesticides; EMS code: N008 | Chemical |  | LAB
N010 | mg/L | Nitrofen | ZN01 | DCM+Ac.Extr:GC/MS | DCM+Ac.Extr:GC/MS | 0.0002 | Y | Y | Y | 28 | 7/28/2005 7:03:00 AM | 3/29/2012 10:15:00 PM | 5/3/2012 9:27:51 AM | 1996-04-12 | 11/16/1999 4:06:27 PM | EMS | NPEPPIN | Concentration | mg/L | Y | (fl. conc.) |  | Nitrofen (fl. conc.) | 1836-75-5 | Pesticides; | Pesticides; EMS code: N010 | Chemical |  | LAB
N030 | mg/L | Nitrophenol-2 | E-8041 | EPA 8041A | EPA 8041A Phenols by Gas Chromatography |  | Y | Y | Y | 37 | 12/29/2021 | 6/10/2024 | 6/26/2024 4:28:10 PM | 2022-04-08 |  | EMS |  | Concentration | mg/L | Y | (fl. conc.) |  | Nitrophenol-2 (fl. conc.) | 136-83-4 | Organics; | Organics; EMS code: N030 | Chemical |  | LAB
N030 | ug/g | Nitrophenol-2 | E8041S | EPA 8041A for soil | See EPA 8041A | 0.05 | Y | Y | Y | 15 | 1/14/2022 | 2/5/2022 | 1/11/2024 1:07:08 PM | 2022-06-23 |  | EMS |  | MassPerMass | mg/kg | Y | (mass conc.) |  | Nitrophenol-2 (mass conc.) | 136-83-4 | Organics; | Organics; EMS code: N030 | Chemical |  | LAB
N030 | mg/L | Nitrophenol-2 | F028 | Acid,Extr.MeCl,HRGC/LRMS:TIM,SIM | A measured volume of sample is acidified and extracted with Methylene Chloride. For samples which have a high organic content that potentially produce chromatographic interferences an initial pre-extraction under basic conditions is performed. The sample extract is then evaporated to small volume and analysed by High Resolution Gas Chromatography / Low Resolution Mass Spectrometry (HRGC/LRMS). LRMS data is acquired in Total Ion Mode (TIM) and quantitation is performed using a single mass ion. Positive Identification of the various target compounds is made by comparison of absolute retention times to those of standards combined with computer based mass spectral mass/intensity library searching and matching routines. | 0.005 | Y | Y | Y | 100 | 4/10/1997 3:18:00 PM | 4/14/2015 12:26:00 AM | 4/16/2015 6:00:11 PM | 1997-01-14 | 6/3/1998 6:51:46 PM | MPAK |  | Concentration | mg/L | Y | (fl. conc.) |  | Nitrophenol-2 (fl. conc.) | 136-83-4 | Organics; | Organics; EMS code: N030 | Chemical |  | LAB
N031 | mg/L | Nitrophenol-4 | E-8041 | EPA 8041A | EPA 8041A Phenols by Gas Chromatography |  | Y | Y | Y | 37 | 12/29/2021 | 6/10/2024 | 6/26/2024 4:28:10 PM | 2022-04-08 |  | EMS |  | Concentration | mg/L | Y | (fl. conc.) |  | Nitrophenol-4 (fl. conc.) | 104-40-5 | Organics; | Organics; EMS code: N031 | Chemical |  | LAB
N031 | mg/L | Nitrophenol-4 | F028 | Acid,Extr.MeCl,HRGC/LRMS:TIM,SIM | A measured volume of sample is acidified and extracted with Methylene Chloride. For samples which have a high organic content that potentially produce chromatographic interferences an initial pre-extraction under basic conditions is performed. The sample extract is then evaporated to small volume and analysed by High Resolution Gas Chromatography / Low Resolution Mass Spectrometry (HRGC/LRMS). LRMS data is acquired in Total Ion Mode (TIM) and quantitation is performed using a single mass ion. Positive Identification of the various target compounds is made by comparison of absolute retention times to those of standards combined with computer based mass spectral mass/intensity library searching and matching routines. | 0.02 | Y | Y | Y | 100 | 4/10/1997 3:19:00 PM | 4/14/2015 12:26:00 AM | 4/16/2015 6:00:11 PM | 1997-01-14 | 6/3/1998 6:51:46 PM | MPAK |  | Concentration | mg/L | Y | (fl. conc.) |  | Nitrophenol-4 (fl. conc.) | 104-40-5 | Organics; | Organics; EMS code: N031 | Chemical |  | LAB
N031 | ug/g | Nitrophenol-4 | E8041S | EPA 8041A for soil | See EPA 8041A | 0.05 | Y | Y | Y | 15 | 1/14/2022 | 2/5/2022 | 1/11/2024 1:07:08 PM | 2022-06-23 |  | EMS |  | MassPerMass | mg/kg | Y | (mass conc.) |  | Nitrophenol-4 (mass conc.) | 104-40-5 | Organics; | Organics; EMS code: N031 | Chemical |  | LAB
NA-D | mg/L | Sodium Dissolved | X170 | ICP | This code is a catch-all for a large variety of ICP methods including water and also solids digested into solutions as an intermediate result. Different procedures will have greatly different detection limits. (a) For air particulates, preliminary results are reported in mg/L and loading results as mg/dm2/d. (b) For Hi Vol samples preliminiary results are same as reported in mg/L and loading results as ug/m3. (c) Note that for Rain samples there is no filtration nor digestion steps. | 2 | Y | Y | Y | 9924 | 9/26/1994 | 10/29/2019 | 8/18/2020 11:41:10 AM | 1996-04-12 | 8/15/2019 1:22:30 PM | EMS | KATFRASE | Concentration | mg/L | Y | (fl. conc.) |  | Sodium Dissolved (fl. conc.) | 7440-23-5 | Metals - Dissolved; | Metals - Dissolved; EMS code: NA-D | Chemical | Dissolved | LAB
NA-D | mg/L | Sodium Dissolved | X203 | FF:FA,HNO3,Dir. Flame,AA | Samples are field-filtered (0.45 um), and field-acidified (HNO3); analysis by AA. | 0.3 | Y | Y | Y | 995 | 9/18/1996 | 11/27/2015 | 7/25/2024 10:37:57 AM | 1996-04-12 | 6/3/1998 6:51:37 PM | EMS |  | Concentration | mg/L | Y | (fl. conc.) |  | Sodium Dissolved (fl. conc.) | 7440-23-5 | Metals - Dissolved; | Metals - Dissolved; EMS code: NA-D | Chemical | Dissolved | LAB
NA-D | mg/L | Sodium Dissolved | MWID | Metals in Water by ICPMS (digested) | Metals in Water by ICPMS (digested) Based on EPA 200.8 | 1 | Y | Y | Y | 2293 |  |  | 7/31/2024 1:46:41 PM | 2009-12-24 |  | RDGIBSON |  | Concentration | mg/L | Y | (fl. conc.) |  | Sodium Dissolved (fl. conc.) | 7440-23-5 | Metals - Dissolved; | Metals - Dissolved; EMS code: NA-D | Chemical | Dissolved | LAB
NA-D | mg/L | Sodium Dissolved | X159 | Auto Flame Emission | Auto Flame Emission | 0.5 | Y | Y | Y | 10056 |  |  | 3/8/2001 9:36:29 AM | 1996-04-12 | 6/3/1998 6:51:37 PM | EMS |  | Concentration | mg/L | Y | (fl. conc.) |  | Sodium Dissolved (fl. conc.) | 7440-23-5 | Metals - Dissolved; | Metals - Dissolved; EMS code: NA-D | Chemical | Dissolved | LAB
NA-D | mg/L | Sodium Dissolved | METC | MET-D-NP-CCMS-VA | Ultra trace metals in water are analyzed by CRC ICPMS, based on US EPA Method 6020B (July 2014). This procedure is intended for pristine field-filtered acid-preserved water samples. The detection limits (LOR) for this test are based on lab instrumental analysis only, not including filtration. ALS-supplied field filtration equipment does not support these LOR. Therefore, because of the high probability of false positives due to filtration, it is strongly recommended that a filtration blank be analysed to aid in data interpretation. | 0.02 | Y | Y | Y | 1776 | 4/3/2012 | 8/15/2024 | 8/23/2024 9:50:35 AM | 2019-08-20 |  | KATFRASE |  | Concentration | mg/L | Y | (fl. conc.) |  | Sodium Dissolved (fl. conc.) | 7440-23-5 | Metals - Dissolved; | Metals - Dissolved; EMS code: NA-D | Chemical | Dissolved | LAB
NA-D | mg/L | Sodium Dissolved | X285 | HNO3 Dig:ICP Analysis | HNO3 Dig:ICP Analysis | 0.1 | Y | Y | Y | 750 | 3/1/2000 | 11/13/2002 | 2/23/2004 2:18:50 PM | 1998-05-22 | 6/3/1998 6:51:49 PM | NPEPPIN |  | Concentration | mg/L | Y | (fl. conc.) |  | Sodium Dissolved (fl. conc.) | 7440-23-5 | Metals - Dissolved; | Metals - Dissolved; EMS code: NA-D | Chemical | Dissolved | LAB
NA-D | mg/L | Sodium Dissolved | F032 | Filt.,Acid,ICPMS | There is no acid digestion required for this analysis. A water sample is collected and preserved in 0.2 percent nitric acid. To analyse the extractable metals the sample is shaken, allowed to sit for at least twenty four hours, and a decanted portion is introduced to the spectrometer. When the dissolved metals are required, the samples must first be filtered through a 0.45 micron cellulose acetate filter paper and then followed by preservation in 0.2 percent nitric acid. It is important to note that different laboratories may have very different detection limits for some variables analyzed by ICPMS. | 0.05 | Y | Y | Y | 76898 | 5/7/1998 | 8/26/2024 | 8/26/2024 4:53:05 PM | 1998-05-14 | 6/3/1998 6:51:49 PM | NPEPPIN |  | Concentration | mg/L | Y | (fl. conc.) |  | Sodium Dissolved (fl. conc.) | 7440-23-5 | Metals - Dissolved; | Metals - Dissolved; EMS code: NA-D | Chemical | Dissolved | LAB
NA-D | mg/L | Sodium Dissolved | X350 | FF,FA:HNO3:ICP | Samples are field-filtered (0.45 um); field-acidified with HNO3; analysis by ICP/AES. | 2 | Y | Y | Y | 6964 |  |  | 8/26/2024 2:41:32 PM | 1996-04-12 | 6/3/1998 6:51:41 PM | EMS |  | Concentration | mg/L | Y | (fl. conc.) |  | Sodium Dissolved (fl. conc.) | 7440-23-5 | Metals - Dissolved; | Metals - Dissolved; EMS code: NA-D | Chemical | Dissolved | LAB
NA-D | mg/L | Sodium Dissolved | F039 | Field filt + acid, ICP-AES | Dissolved and extractable water samples are ready for analysis once collected. Total samples are autoclaved with HNO3 and hydrochloric acid (HCI) in capped Teflon vials. Analysis is performed via Inductively Coupled Argon Plasma Atomic Emission Spectrometry (ICP-AES)(simultaneous multielement analysis), Graphite Furnace Atomic Absorption Spectrometry (GF-AAS)(low level analysis) and/or Flame Atomic Absorption Spectrometry (F-AAS). | 0.1 | Y | Y | Y | 8906 | 7/29/1996 6:03:00 PM | 7/5/2019 1:21:00 PM | 5/16/2022 4:21:33 PM | 1997-05-08 | 6/3/1998 6:51:46 PM | MPAK |  | Concentration | mg/L | Y | (fl. conc.) |  | Sodium Dissolved (fl. conc.) | 7440-23-5 | Metals - Dissolved; | Metals - Dissolved; EMS code: NA-D | Chemical | Dissolved | LAB
NA-D | mg/L | Sodium Dissolved | X562 | Dissolved metals: FF, FA, HNO3, HR-ICPMS | Ultra trace metals in water are analyzed by high resolution inductively coupled plasma mass spectrometry (HR-ICPMS), based on US EPA Method 200.8, (Oct. 1999). This procedure is intended for pristine field-filtered acid-preserved water samples. The detection limits (LOR) for this test are based on lab instrumental analysis only, not including filtration. ALS-supplied field filtration equipment does not support these LOR. Therefore, because of the high probability of false positives due to filtration, it is strongly recommended that a filtration blank be analyzed to aid in data interpretation. | 0.01 | Y | Y | Y | 1260 | 6/8/2015 | 10/19/2023 | 11/9/2023 7:41:27 AM | 2015-10-22 |  | RDGIBSON |  | Concentration | mg/L | Y | (fl. conc.) |  | Sodium Dissolved (fl. conc.) | 7440-23-5 | Metals - Dissolved; | Metals - Dissolved; EMS code: NA-D | Chemical | Dissolved | LAB
NA-D | mg/L | Sodium Dissolved | EQ01 | EQUIS (mg/L) | EQUIS (mg/L) | 0 | N | Y | Y | 9487 |  |  | 10/24/1996 3:16:55 PM | 1996-04-12 | 6/3/1998 6:51:37 PM | EMS |  | Concentration | mg/L | Y | (fl. conc.) |  | Sodium Dissolved (fl. conc.) | 7440-23-5 | Metals - Dissolved; | Metals - Dissolved; EMS code: NA-D | Chemical | Dissolved | LAB
NA-D | mg/L | Sodium Dissolved | F049 | Filt,HNO3 dig,conc'n,nebul'ICP-OES | This method provides an automated procedure for the Inductively Coupled Argon Plasma Optical Emission Spectrophotometric (ICP-OES) determination and quantitation of trace amounts of metals in surface, ground and precipitation waters after a manual acid digestion and concentration procedure and ultrasonic nebulisation. Organically-bound metals and metal complexes are converted to metal ions by digestion with nitric acid. To increase the signal detected by the ICP, the samples are concentrated 10 times. The liquid sample is pumped onto the piezoelectric transducer, where the ultrasonic action efficiently converts it to a fine, dense aerosol. The nebulizer carrier gas flow then transports the wet aerosol through a heated U-tube, where the solvent is vaporized. Subsequently, most of the solvent vapour is removed from the aerosol stream by passing the stream through a chilled condenser. The result is a dry, analyte-laden aerosol which is introduced to the plasma. There, the high temperature desolvates the aerosol and dissociates the sample material into individual atoms and ions, which are excited to emit light at wave lengths characteristic of the elements present. Dissolved samples are field filtered through 0.45 micron cellulose acetate filter prior to acid preservation. Samples for extractable analysis are acidified with 2 mL concentrated HNO3 per litre and allowed to stand a minimum of 24 hours before analysis. Clients providing these samples are advised to preserve samples at the collection site. | 2 | Y | Y | Y | 2245 | 3/10/1998 2:07:00 PM | 2/25/1999 2:08:00 PM | 8/28/2018 3:14:04 PM | 1998-11-17 |  | NPEPPIN |  | Concentration | mg/L | Y | (fl. conc.) |  | Sodium Dissolved (fl. conc.) | 7440-23-5 | Metals - Dissolved; | Metals - Dissolved; EMS code: NA-D | Chemical | Dissolved | LAB
NA-D | mg/L | Sodium Dissolved | X053 | AA,Dir Asp. | AA;Dir Asp. | 0.3 | Y | Y | Y | 7406 | 1/2/1997 8:00:00 AM | 2/1/2010 8:00:00 AM | 2/10/2010 12:03:37 PM | 1996-04-12 | 6/3/1998 6:51:37 PM | EMS |  | Concentration | mg/L | Y | (fl. conc.) |  | Sodium Dissolved (fl. conc.) | 7440-23-5 | Metals - Dissolved; | Metals - Dissolved; EMS code: NA-D | Chemical | Dissolved | LAB
NA-E | mg/L | Sodium Extractable | E6020B | EPA 6020B | See EPA 6020B and CSSS Ch. 15 | 2 | Y | Y | Y | 16 | 1/11/2022 | 2/7/2023 | 1/11/2024 1:07:08 PM | 2022-06-29 |  | EMS |  | Concentration | mg/L | Y | (fl. conc.) |  | Sodium Extractable (fl. conc.) | 7440-23-5 | Metals - Extractable; | Metals - Extractable; EMS code: NA-E | Chemical | Extractable | LAB
NA-E | ug/g | Sodium Extractable | CAL2 | Calculation (ug/g) | Calculation (ug/g) | 1 | Y | Y | Y | 16 | 1/11/2022 | 2/7/2023 | 1/11/2024 1:07:08 PM | 2022-07-06 |  | EMS |  | MassPerMass | mg/kg | Y | (mass conc.) |  | Sodium Extractable (mass conc.) | 7440-23-5 | Metals - Extractable; | Metals - Extractable; EMS code: NA-E | Chemical | Extractable | LAB
NA-E | ug/g | Sodium Extractable | ASCE | Ammonium Acetate soil colloid exchange | Available metals - Ammonium Acetate soil colloid exchange sites displacement |  | Y | Y | Y | 7 | 10/31/2022 | 10/31/2022 | 11/22/2022 2:40:39 PM | 2007-10-04 |  | RDGIBSON |  | MassPerMass | mg/kg | Y | (mass conc.) |  | Sodium Extractable (mass conc.) | 7440-23-5 | Metals - Extractable; | Metals - Extractable; EMS code: NA-E | Chemical | Extractable | LAB
NA-E | mg/L | Sodium Extractable | F040 | Field Acid, lab filt, ICP-AES | Dissolved and extractable water samples are ready for analysis once collected. Total samples are autoclaved with HNO3 and hydrochloric acid (HCI) in capped Teflon vials. Analysis is performed via Inductively Coupled Argon Plasma Atomic Emission Spectrometry (ICP-AES)(simultaneous multielement analysis), Graphite Furnace Atomic Absorption Spectrometry (GF-AAS)(low level analysis) and/or Flame Atomic Absorption Spectrometry (F-AAS). | 0.1 | Y | Y | Y | 8918 | 8/5/1997 6:06:00 PM | 3/23/2012 3:11:00 AM | 9/11/2019 8:47:41 AM | 1997-05-08 | 6/3/1998 6:51:46 PM | MPAK |  | Concentration | mg/L | Y | (fl. conc.) |  | Sodium Extractable (fl. conc.) | 7440-23-5 | Metals - Extractable; | Metals - Extractable; EMS code: NA-E | Chemical | Extractable | LAB
NA-E | mg/L | Sodium Extractable | F055 | No Pres/No Filtr/No Dig,ICP | Shortcut method which excludes preservation, filtration and digestion steps. Results are used for estimation of water hardness. | 0.1 | Y | Y | Y | 2359 | 8/23/1996 6:04:00 PM | 6/1/2000 11:27:00 AM | 9/11/2019 8:46:38 AM | 1998-04-21 | 6/3/1998 6:51:48 PM | NPEPPIN |  | Concentration | mg/L | Y | (fl. conc.) |  | Sodium Extractable (fl. conc.) | 7440-23-5 | Metals - Extractable; | Metals - Extractable; EMS code: NA-E | Chemical | Extractable | LAB
NA-L | mg/L | Sodium Leachable | MEND | MEND (mod) EPA 300.1 (mod) | MEND (mod) EPA 300.1 (mod). Inorganic anions are analyzed by obtaining an extract from a soil sample that has been added in a 3:1 ratio of deionized water to soil; then leached and the extract filtered prior to analysis; which is then analyzed by Ion Chromatography with conductivity and/or UV detection. | 0.05 | Y | Y | Y | 0 |  |  |  | 2024-01-29 |  | RDGIBSON |  | Concentration | mg/L | Y | (fl. conc.) |  | Sodium Leachable (fl. conc.) |  | Metals - Extractable; | Metals - Extractable; EMS code: NA-L | Chemical |  | LAB
NA-S | ug/m3 | Sodium-Soluble ion content | EQ40 | EQUIS (mg/m3) | EQUIS (mg/m3) | 0 | N | Y | Y | 2961 |  |  | 10/22/1996 10:21:01 PM | 1996-04-12 | 6/3/1998 6:51:37 PM | EMS |  | AirConcentration | µg/m³ | Y | (air conc.) |  | Sodium-Soluble ion content (air conc.) | 7440-23-5 | Metals - Soluble; | Metals - Soluble; EMS code: NA-S | Chemical | Soluble | LAB
NA-S | mg/L | Sodium-Soluble ion content | 5318 | Na - Ion Chromatog.Hivol | For HiVol samples, discs cut from the HiVol filter are extracted with DI water and the soluble ions analyzed by IC. Intermediate results are reported as mg/L and HiVol loading results as ug/m3. | 0.01 | Y | Y | Y | 82 |  |  | 10/23/1996 5:33:36 PM | 1996-04-12 | 6/3/1998 6:51:37 PM | EMS |  | Concentration | mg/L | Y | (fl. conc.) |  | Sodium-Soluble ion content (fl. conc.) | 7440-23-5 | Metals - Soluble; | Metals - Soluble; EMS code: NA-S | Chemical | Soluble | LAB
NA-S | mg/L | Sodium-Soluble ion content | E6020B | EPA 6020B | See EPA 6020B and CSSS Ch. 15 | 2 | Y | Y | Y | 0 |  |  |  | 2024-02-02 |  | EMS |  | Concentration | mg/L | Y | (fl. conc.) |  | Sodium-Soluble ion content (fl. conc.) | 7440-23-5 | Metals - Soluble; | Metals - Soluble; EMS code: NA-S | Chemical | Soluble | LAB
NA-S | mg/dm2/d | Sodium-Soluble ion content | 5061 | Ion Chromatog- Cations | For the Soluble Particulates Pkg. a portion of filtrate from the Particulate - Total procedure made up to 500 mL is analyzed by IC. Intermediate results are reported in mg/L. Loading results are calculated and reported with units of mg/dm2/d. Note that ammonium is reported as NH4 not as N. | 0.0002 | Y | Y | Y | 136 |  |  | 10/24/1996 6:07:34 AM | 1996-04-12 | 6/3/1998 6:51:41 PM | EMS |  | MassFlux | mg/dm²/d | Y | (mass flux) |  | Sodium-Soluble ion content (mass flux) | 7440-23-5 | Metals - Soluble; | Metals - Soluble; EMS code: NA-S | Chemical | Soluble | LAB
NA-S | mg/L | Sodium-Soluble ion content | E6010B | EPA 6010B/CSSS 15 | See CSSS CH15/EPA 6010B/AER D50. INDUCTIVELY COUPLED PLASMA-ATOMIC EMISSION SPECTROMETRY See EPA 6010B and CSSS Ch. 15. https://www.epa.gov/sites/default/files/documents/6010b.pdf | 5 | Y | Y | Y | 0 |  |  |  | 2024-01-30 |  | EMS |  | Concentration | mg/L | Y | (fl. conc.) |  | Sodium-Soluble ion content (fl. conc.) | 7440-23-5 | Metals - Soluble; | Metals - Soluble; EMS code: NA-S | Chemical | Soluble | LAB
NA-S | mg/dm2/d | Sodium-Soluble ion content | EQ15 | EQUIS (mg/dm2/d) | EQUIS (mg/dm2/d) | 0.002 | N | Y | Y | 563 |  |  | 10/22/1996 10:21:00 PM | 1996-04-12 | 6/3/1998 6:51:37 PM | EMS |  | MassFlux | mg/dm²/d | Y | (mass flux) |  | Sodium-Soluble ion content (mass flux) | 7440-23-5 | Metals - Soluble; | Metals - Soluble; EMS code: NA-S | Chemical | Soluble | LAB
NA-S | mg/L | Sodium-Soluble ion content | EQ01 | EQUIS (mg/L) | EQUIS (mg/L) | 0.01 | N | Y | Y | 125 |  |  | 10/22/1996 11:12:48 PM | 1996-04-12 | 6/3/1998 6:51:37 PM | EMS |  | Concentration | mg/L | Y | (fl. conc.) |  | Sodium-Soluble ion content (fl. conc.) | 7440-23-5 | Metals - Soluble; | Metals - Soluble; EMS code: NA-S | Chemical | Soluble | LAB
NA-S | ug/m3 | Sodium-Soluble ion content | 5018 | H2O Extr,Ion Chr-Cat I | For HiVol samples, discs cut from the HiVol filter are extracted with DI water and the soluble ions analyzed by IC. Intermediate results are reported as mg/L and HiVol loading results as ug/m3. | 0.01 | Y | Y | Y | 1306 |  |  | 10/24/1996 6:37:04 AM | 1996-04-12 | 6/3/1998 6:51:37 PM | EMS |  | AirConcentration | µg/m³ | Y | (air conc.) |  | Sodium-Soluble ion content (air conc.) | 7440-23-5 | Metals - Soluble; | Metals - Soluble; EMS code: NA-S | Chemical | Soluble | LAB
NA-S | mg/L | Sodium-Soluble ion content | F032 | Filt.,Acid,ICPMS | There is no acid digestion required for this analysis. A water sample is collected and preserved in 0.2 percent nitric acid. To analyse the extractable metals the sample is shaken, allowed to sit for at least twenty four hours, and a decanted portion is introduced to the spectrometer. When the dissolved metals are required, the samples must first be filtered through a 0.45 micron cellulose acetate filter paper and then followed by preservation in 0.2 percent nitric acid. It is important to note that different laboratories may have very different detection limits for some variables analyzed by ICPMS. | 0.01 | Y | Y | Y | 337 | 3/1/2002 1:15:00 PM | 2/17/2005 | 3/17/2005 1:07:53 AM | 2004-02-10 |  | UMULLER |  | Concentration | mg/L | Y | (fl. conc.) |  | Sodium-Soluble ion content (fl. conc.) | 7440-23-5 | Metals - Soluble; | Metals - Soluble; EMS code: NA-S | Chemical | Soluble | LAB
NA-S | mg/L | Sodium-Soluble ion content | 5071 | Ion Chromatog -Cation I | (a) For the Soluble Particulates Pkg. a portion of filtrate from the Particulate - Total procedure made up to 500 mL is analyzed by IC. Intermediate results are reported in mg/L. Loading results are calculated and reported with units of mg/dm2/d.(b) For precipitation samples, all samples are filtered through a 0.22 um fritted glass filter prior to injection. The cations of interest are separated through a cation "guard" and cation "separator" column with a methane sulphonic acid eluent. After the cations are separated they exit at various times from the column in a background of eluent. A self regenerating membrane suppressor is attached to the end of the column to neutralize the acid before detection and form the corresponding hydroxide. Note that ammonium is reported as NH4 not as N. (c) For HiVol samples, discs cut from the HiVol filter are extracted with DI water and the soluble anions are analyzed by IC. Intermediate results are reported in units of mg/L and HiVol loading results are reported as ug/m3. (*) Note that ammonium is reported as NH4 not as N | 0.01 | Y | Y | Y | 3669 | 4/30/1997 3:50:00 PM | 1/7/2002 2:33:00 PM | 6/24/2003 10:43:22 AM | 1996-04-12 | 6/3/1998 6:51:37 PM | EMS |  | Concentration | mg/L | Y | (fl. conc.) |  | Sodium-Soluble ion content (fl. conc.) | 7440-23-5 | Metals - Soluble; | Metals - Soluble; EMS code: NA-S | Chemical | Soluble | LAB
NA-S | mg/L | Sodium-Soluble ion content | 5073 | Aut.Col-Flame Phot/MerTh | Aut.Col-Flame Phot/MerTh | 0.1 | Y | Y | Y | 2364 | 4/6/2005 | 4/11/2008 | 5/28/2008 2:08:44 AM | 1996-04-12 | 6/3/1998 6:51:37 PM | EMS |  | Concentration | mg/L | Y | (fl. conc.) |  | Sodium-Soluble ion content (fl. conc.) | 7440-23-5 | Metals - Soluble; | Metals - Soluble; EMS code: NA-S | Chemical | Soluble | LAB
NA-S | ug/g | Sodium-Soluble ion content | CAL2 | Calculation (ug/g) | Calculation (ug/g) | 1 | Y | Y | Y | 0 |  |  |  | 2024-02-02 |  | EMS |  | MassPerMass | mg/kg | Y | (mass conc.) |  | Sodium-Soluble ion content (mass conc.) | 7440-23-5 | Metals - Soluble; | Metals - Soluble; EMS code: NA-S | Chemical | Soluble | LAB
NA-T | mg/L | Sodium Total | F033 | Acid Dig,Centrifuge,ICPMS | For the determination of total recoverable elements in water or waste water, take 100 ml aliquot from a well mixed, acid preserved, sample containing not more than 0.25% (w/v) total solids and transfer to a 250 ml Griffin beaker. (If total solids are greater than 0.25% reduce the size of the aliquot by a proportionate amount). Add 1 ml of conc. nitric acid and 0.5 ml conc. hydrochloric acid. Heat on a hot plate until the volume has been reduced to 15-20ml, ensuring that the sample does not boil. Allow to cool and transfer to a 50 ml volumetric flask. Dilute to volume with ASTM type I water. Centrifuge the sample or allow to stand overnight to separate insoluble material. Prior analysis, pipette 20 ml into a 50 ml volumetric flask and dilute to volume with ASTM type I water. Analysis should be performed within seventy two hours of the completed preparation. Note that different laboratories may have very different field preservation, digestion procedures and detection limits for this method, so these should be documented. | 0.05 | Y | Y | Y | 24776 | 5/7/1998 | 8/22/2024 3:24:00 AM | 8/25/2024 9:47:36 PM | 1998-05-14 | 6/3/1998 6:51:49 PM | NPEPPIN |  | Concentration | mg/L | Y | (fl. conc.) |  | Sodium Total (fl. conc.) | 7440-23-5 | Metals - Total; | Metals - Total; EMS code: NA-T | Chemical | Total | LAB
NA-T | mg/L | Sodium Total | 0040 | HNO3 Dig: ICP Analysis | HNO3 Dig: ICP Analysis | 0.01 | Y | Y | Y | 1715 | 5/8/2003 9:40:00 AM | 7/25/2012 7:11:00 AM | 9/26/2012 1:06:57 PM | 1996-04-12 | 6/3/1998 6:51:41 PM | EMS |  | Concentration | mg/L | Y | (fl. conc.) |  | Sodium Total (fl. conc.) | 7440-23-5 | Metals - Total; | Metals - Total; EMS code: NA-T | Chemical | Total | LAB
NA-T | % | Sodium Total | MS41 | ALS Minerals ME-MS41 | ALS Minerals ME-MS41. A prepared sample is digested with aqua regia in a graphite heating block. After cooling; the resulting solution is diluted to with deionized water; mixed and analyzed by ICP-AES. Samples are then analyzed by ICP-MS for the remaining suite of elements. | 0.01 | Y | Y | Y | 0 |  |  |  | 2024-01-29 |  | EMS |  | DimensionlessRatio | % | N | (ratio) |  | Sodium Total (ratio) | 7440-23-5 | Metals - Total; | Metals - Total; EMS code: NA-T | Chemical | Total | LAB
NA-T | mg/L | Sodium Total | X170 | ICP | This code is a catch-all for a large variety of ICP methods including water and also solids digested into solutions as an intermediate result. Different procedures will have greatly different detection limits. (a) For air particulates, preliminary results are reported in mg/L and loading results as mg/dm2/d. (b) For Hi Vol samples preliminiary results are same as reported in mg/L and loading results as ug/m3. (c) Note that for Rain samples there is no filtration nor digestion steps. | 2 | Y | Y | Y | 5881 | 3/28/1996 6:03:00 PM | 11/7/2018 | 10/24/2022 10:58:25 AM | 1996-04-12 | 8/15/2019 1:22:52 PM | EMS | KATFRASE | Concentration | mg/L | Y | (fl. conc.) |  | Sodium Total (fl. conc.) | 7440-23-5 | Metals - Total; | Metals - Total; EMS code: NA-T | Chemical | Total | LAB
NA-T | mg/L | Sodium Total | X561 | Total metals: FA, HNO3, HR-ICPMS (No Prep)Prep) | Ultra trace metals in water are analyzed by high resolution inductively coupled plasma mass spectrometry (HR-ICPMS), modified from US EPA 200.8, (Revision 5.5). The detection limits provided can only be met for undigested samples. This procedure is intended for colorless, non-turbid, acid-preserved water samples (i.e. pristine water samples), having turbidity < 1 NTU and no odor. Where turbidity exceeds 1 NTU, and/or the sample is colored and has an odor, results may be biased low compared to true Total Metals concentrations. ALS recommends that turbidity analysis be requested on samples submitted for this test to aid with interpretation of results. | 0.01 | Y | Y | Y | 2746 | 6/8/2015 | 10/5/2023 | 1/31/2024 2:37:41 PM | 2015-10-22 |  | RDGIBSON |  | Concentration | mg/L | Y | (fl. conc.) |  | Sodium Total (fl. conc.) | 7440-23-5 | Metals - Total; | Metals - Total; EMS code: NA-T | Chemical | Total | LAB
NA-T | ug/g | Sodium Total | PB02 | PBM,HNO3/HCl dig,dry60C,100mesh, | Performance Based Method. Samples are dried at 60 C and sieved using a .149 mm (100 mesh) sieve. Solid samples are digested with the combination of nitric and hydrochloric acids with a 1:1 ratio of acids. For environmental purposes, strong acid leaches are used routinely since they have the advantage of speed, low cost, and superio reproducibility. All metals trapped within the refractory matrix are unlikely to have environmental consequences. Instrumental analysis of extracts produced using this digestion method can be performed by a variety of techniques. |  | Y | Y | Y | 10 | 9/13/2001 | 9/13/2001 | 11/6/2002 3:52:37 PM | 2001-07-18 |  | NPEPPIN |  | MassPerMass | mg/kg | Y | (mass conc.) |  | Sodium Total (mass conc.) | 7440-23-5 | Metals - Total; | Metals - Total; EMS code: NA-T | Chemical | Total | LAB
NA-T | ug/g | Sodium Total | X551 | Metals in Biota (dry), total, acid digestion, ICP-SFMS | 0.1 gram of dried, homogenized biota is weighed in a microwave TFM vessel, and 8mL of Nitric Acid and 1mL of Hydrogen Peroxide are added. The sample is digested at 200 degrees C for 15 minutes in a closed vessel high pressure microwave oven. The sample is cooled, the volume adjusted to 100mL and introduced to an iductively coupled argon plasma-sector field mass spectrometer (ICP-SFMS) by pneumatic nebulization. Each respective element is measured at a specific mass to charge ratio, m/z value expressed in amu (atomic mass units), and at a specified instrument resolution. LR: Low Resolution ~300, MR: Mid Resolution ~4000 and HR: Hich Resolution ~10000. Ref: Method number 02-2705, National Laboratory for Environmental Testing, CCIW, Environment Canada, Burlington. Requested April 2003. | 20 | Y | Y | Y | 49 | 8/11/2015 | 10/30/2023 | 11/28/2023 2:21:33 PM | 2015-03-11 |  | RDGIBSON |  | MassPerMass | mg/kg | Y | (mass conc.) |  | Sodium Total (mass conc.) | 7440-23-5 | Metals - Total; | Metals - Total; EMS code: NA-T | Chemical | Total | LAB
NA-T | mg/L | Sodium Total | MSEA | Metals in Seawater (undigested) by (CRC) ICP-MS | This method describes the multi-elemental determination of trace elements by Collision/Reaction Cell (CRC) ICP-MS. The method measures ions produced by a radio-frequency inductively coupled plasma. Analyte species originating in a liquid are pneumatically nebulized and the resulting aerosol transported by argon gas into the plasma torch. The ions produced are entrained in the plasma gas and extracted, by means of a differentially pumped vacuum interface, into a mass spectrometer. The ions flow through a Collision/Reaction Cell (CRC) where depending on the mode of analysis either the ions passes through the cell unhindered (no gas - classical ICP-MS mode), or collisions/reactions chemistry takes place between the ions and a neutral gas such as hydrogen or helium. Any polyatomic ions which give interference on the analytes are removed. The ions are then sorted according to their mass-to-charge ratios by a quadrupole mass spectrometer. The ions are then sorted according to their mass-to-charge ratios by a quadrupole mass spectrometer. The ions transmitted through the quadrupole are quantified by a channel electron multiplier and the ion information processed by a data handling system. Although most polyatomic interference is removed, any additional interference must be assessed and valid corrections applied or the data flagged to indicate problems. Interference correction must include compensation for background ions contributed by the plasma gas, reagents, and constituents of the sample matrix. Instrument drift as well as suppressions or enhancements of instrument response caused by the sample matrix must be corrected for by the use of internal standards. | 10 | Y | Y | Y | 113 | 8/11/2009 11:57:00 PM | 9/1/2011 | 9/10/2011 2:30:04 PM | 2009-09-08 |  | RDGIBSON |  | Concentration | mg/L | Y | (fl. conc.) |  | Sodium Total (fl. conc.) | 7440-23-5 | Metals - Total; | Metals - Total; EMS code: NA-T | Chemical | Total | LAB
NA-T | mg/L | Sodium Total | X285 | HNO3 Dig:ICP Analysis | HNO3 Dig:ICP Analysis | 0.1 | Y | Y | Y | 674 | 8/10/1993 | 4/24/2013 4:03:00 AM | 6/3/2014 3:32:28 PM | 1998-05-22 | 6/3/1998 6:51:49 PM | NPEPPIN |  | Concentration | mg/L | Y | (fl. conc.) |  | Sodium Total (fl. conc.) | 7440-23-5 | Metals - Total; | Metals - Total; EMS code: NA-T | Chemical | Total | LAB
NA-T | mg/L | Sodium Total | 0042 | GI HNO3 Dig: ICP | Sample aliquot taken from unfiltered unpreserved General Ions bottle; HNO3 digestion; ICP analysis. | 0.4 | Y | Y | Y | 7407 | 9/20/1996 | 7/5/2019 3:21:00 PM | 7/12/2019 6:49:13 PM | 1996-04-12 | 6/3/1998 6:51:37 PM | EMS |  | Concentration | mg/L | Y | (fl. conc.) |  | Sodium Total (fl. conc.) | 7440-23-5 | Metals - Total; | Metals - Total; EMS code: NA-T | Chemical | Total | LAB
NA-T | ug/g | Sodium Total | FS89 | Homo,HNO3 dig,ICP-AES | Homo;HNO3 dig;ICP-AES | 20 | Y | Y | Y | 4 | 1/5/2001 | 8/20/2001 12:00:00 PM | 10/23/2002 1:42:48 PM | 2001-03-08 |  | NPEPPIN |  | MassPerMass | mg/kg | Y | (mass conc.) |  | Sodium Total (mass conc.) | 7440-23-5 | Metals - Total; | Metals - Total; EMS code: NA-T | Chemical | Total | LAB
NA-T | ug/g | Sodium Total | ICPM | Metals in Soil by CRC ICPMS | This analysis is carried out using procedures from CSR Analytical Method: "Strong Acid Leachable Metals (SALM) in Soil", BC Ministry of Environment, 26 June 2009, and procedures adapted from EPA Method 200.2. The sample is manually homogenized, dried at 60 degrees Celsius, sieved through a 63 um (230 mesh) sieve, and a representative subsample of the dry material is weighed. The sample is then digested at 95 degrees Celsius for 2 hours by block digester using concentrated nitric and hydrochloric acids. Instrumental analysis is by collision cell inductively coupled plasma - mass spectrometry (modifed from EPA Method 6020A). Method Limitation: This method is not a total digestion technique. It is a very strong acid digestion that is intended to dissolve those metals that may be environmentally available. By design, elements bound in silicate structures are not normally dissolved by this procedure as they are not usually mobile in the environment. | 100 | Y | Y | Y | 103 | 9/19/2017 | 9/24/2023 | 1/12/2024 11:16:52 AM | 2018-01-09 |  | JKROGH |  | MassPerMass | mg/kg | Y | (mass conc.) |  | Sodium Total (mass conc.) | 7440-23-5 | Metals - Total; | Metals - Total; EMS code: NA-T | Chemical | Total | LAB
NA-T | ug/g | Sodium Total | FB04 | Freeze Dry,Hom,Acid Dig,ICP | Sample moisture is determined by weighing samples before and after freeze-drying. Freeze-dried samples are ground in a Braun blender before being weighed into Teflon PFA liners and digested with concentrated HNO3, concentrated HCl, and 50% H2O2 in a microwave oven; which results in a complete sample dissolution. Samples are collected and volumized with Type I water. Analysis is performed via Inductively Coupled Argon Plasma Atomic Emission Spectrometry (ICP-AES) (simultaneous multielement analysis), Graphite Furnace Atomic Absorption Spectrometry (GF-AAS) (low level Cd and Pb analysis), Cold Vapor Atomic Fluorescence Spectrometry (CV-AFS) (Mercury analysis) and/or Flame Atomic Absorption Spectrometry (F-AAS) (K analysis). | 8 | Y | Y | Y | 241 | 8/22/1996 6:11:00 PM | 2/1/2002 2:30:00 PM | 6/23/2003 1:15:32 PM | 1996-06-19 | 6/3/1998 6:51:44 PM | MPAK |  | MassPerMass | mg/kg | Y | (mass conc.) |  | Sodium Total (mass conc.) | 7440-23-5 | Metals - Total; | Metals - Total; EMS code: NA-T | Chemical | Total | LAB
NA-T | % | Sodium Total | 6041 | Aqua Regia by ICPMS | These were sediment/soils digested with Aqua-Regia (Nitric Acid & Hydrochloric Acids) and then analysis by ICPMS. Results would be reported in mg/kg and/or %. | 0.001 | Y | Y | Y | 92 | 10/13/2019 | 7/27/2024 | 7/29/2024 9:13:08 AM | 2019-09-06 |  | KATFRASE |  | DimensionlessRatio | % | N | (ratio) |  | Sodium Total (ratio) | 7440-23-5 | Metals - Total; | Metals - Total; EMS code: NA-T | Chemical | Total | LAB
NA-T | % | Sodium Total | MS61 | ALS Minerals ME-MS61 | ALS Minerals ME-MS61. A prepared sample is digested with HNO3; HF; and HClO4 in a graphite heating block to near dryness. After cooling; the sample is boiled in diluted HCl. Bulk to volume with deionized water; mixed and analyzed by ICP-AES and ICP-MS for all elements. | 0.01 | Y | Y | Y | 0 |  |  |  | 2024-01-29 |  | EMS |  | DimensionlessRatio | % | N | (ratio) |  | Sodium Total (ratio) | 7440-23-5 | Metals - Total; | Metals - Total; EMS code: NA-T | Chemical | Total | LAB
NA-T | ug/g wet | Sodium Total | X560 | Metals in Biota (wet), total, acid digestion, ICP-MS | Five grams of wet, homogenized biota is weighed in a centrifuge tube, and 2.5 mL of Nitric Acid and 2.5 mL of Hydrochloric acid are added. The sample is digested at room temperature for one hour and then digested at 90-95 degrees C in a water bath or heat block for one more hour. The sample is cooled and the volume is made to 50 mL with DI. The digestate is then diluted at ten fold and introduced to a collision/reaction cell inductively coupled argon plasma mass spectrometer (CRC ICPMS) by pneumatic nebulization. Each respective element is measured at a specific mass to charge ratio, m/z value expressed in amu (atomic mass units). - Maxxam Analytics (BC). Added to EMS Dec 10, 2008. | 10 | Y | Y | Y | 541 | 11/15/2008 4:00:00 AM | 10/30/2023 | 11/28/2023 2:21:33 PM | 2008-12-11 | 3/11/2015 4:17:10 PM | RDGIBSON | EMS | MassPerMass | µg/g | Y | (mass conc.) | (wet) | Sodium Total (wet) (mass conc.) | 7440-23-5 | Metals - Total; | Metals - Total; EMS code: NA-T | Chemical | Total | LAB
NA-T | mg/L | Sodium Total | X351 | LA:HNO3:AA | Samples are lab-acidified (HNO3); digested HNO3; analysis by AA. | 0.002 | Y | Y | Y | 326 |  |  | 2/27/2007 1:16:57 PM | 1996-04-12 | 6/3/1998 6:51:37 PM | EMS |  | Concentration | mg/L | Y | (fl. conc.) |  | Sodium Total (fl. conc.) | 7440-23-5 | Metals - Total; | Metals - Total; EMS code: NA-T | Chemical | Total | LAB
NA-T | mg/L | Sodium Total | MWID | Metals in Water by ICPMS (digested) | Metals in Water by ICPMS (digested) Based on EPA 200.8 | 1 | Y | Y | Y | 36502 |  |  | 7/31/2024 1:46:41 PM | 2009-12-24 |  | RDGIBSON |  | Concentration | mg/L | Y | (fl. conc.) |  | Sodium Total (fl. conc.) | 7440-23-5 | Metals - Total; | Metals - Total; EMS code: NA-T | Chemical | Total | LAB
NA-T | ug/g | Sodium Total | MTID | MET-DRY-CCMS-VA | Metals in Tissue by CRC ICPMS (DRY) Methodology: This method is conducted following British Columbia Lab Manual method "Metals in Animal Tissue and Vegetation (Biota) - Prescriptive". Tissue samples are homogenized and sub-sampled prior to hotblock digestion with nitric and hydrochloric acids, in combination with addition of hydrogen peroxide. Instrumental analysis is by collision cell inductively coupled plasma - mass spectrometry (modified from EPA Method 6020A). Method Limitation: This method employs a strong acid/peroxide digestion, and is intended to provide a conservative estimate of bio-available metals. Near complete recoveries are achieved for most toxicologically important metals, but elements associated with recalcitrant minerals may be only partially recovered. | 20 | Y | Y | Y | 122 | 8/13/2014 | 9/8/2022 | 10/26/2022 7:14:15 AM | 2015-03-09 | 3/11/2015 4:12:05 PM | RDGIBSON | EMS | MassPerMass | mg/kg | Y | (mass conc.) |  | Sodium Total (mass conc.) | 7440-23-5 | Metals - Total; | Metals - Total; EMS code: NA-T | Chemical | Total | LAB
NA-T | mg/L | Sodium Total | F121 | Acid,ICPMS | Acidified, no digestion, ICPMS | 0.01 | Y | Y | Y | 6504 | 10/18/2007 | 8/16/2024 | 8/26/2024 8:06:03 AM | 2002-07-22 | 7/16/2012 3:33:45 PM | UMULLER | RDGIBSON | Concentration | mg/L | Y | (fl. conc.) |  | Sodium Total (fl. conc.) | 7440-23-5 | Metals - Total; | Metals - Total; EMS code: NA-T | Chemical | Total | LAB
NA-T | ug/g wet | Sodium Total | X552 | Metals in Biota (wet), total, acid digestion, ICP-SFMS | One gram of wet, homogenized biota is weighed in a microwave TFM vessel, and 8mL of Nitric Acid and 1mL of Hydrogen Peroxide are added. The sample is digested at 200 degrees C for 15 minutes in a closed vessel high pressure microwave oven. The sample is cooled, the volume adjusted to 100mL and introduced to an inductively coupled argon plasma-sector field mass spectrometer (ICP-SFMS) by pneumatic nebulization. Each respective element is measured at a specific mass to charge ratio, m/z value expressed in amu (atomic mass units), and at a specified instrument resolution. LR: Low Resolution ~300, MR: Mid Resolution ~4000. Ref: 02-2704, National Laboratory for Environmental Testing, CCIW, Environment Canada, Burlington. Requested April 2003. | 4 | Y | Y | Y | 37 | 8/9/2016 | 10/16/2023 | 11/28/2023 2:21:32 PM | 2015-03-11 |  | RDGIBSON |  | MassPerMass | µg/g | Y | (mass conc.) | (wet) | Sodium Total (wet) (mass conc.) | 7440-23-5 | Metals - Total; | Metals - Total; EMS code: NA-T | Chemical | Total | LAB
NA-T | mg/L | Sodium Total | 1349 | Tot Metals in Seawater by CRC ICPMS (BC) | Tot Metals in Seawater by CRC ICPMS (BC) | 2.5 | Y | Y | Y | 18 | 4/13/2023 | 4/29/2024 | 6/7/2024 10:03:16 AM | 2023-05-26 |  | EMS |  | Concentration | mg/L | Y | (fl. conc.) |  | Sodium Total (fl. conc.) | 7440-23-5 | Metals - Total; | Metals - Total; EMS code: NA-T | Chemical | Total | LAB
NA-T | mg/L | Sodium Total | X073 | FA,HNO3,Direct Flame,AA | Samples are field-acidified with HNO3; HNO3 digestion; analysis by AA. | 0.002 | Y | Y | Y | 268 |  |  | 1/25/2007 4:06:12 PM | 1996-04-12 | 6/3/1998 6:51:37 PM | EMS |  | Concentration | mg/L | Y | (fl. conc.) |  | Sodium Total (fl. conc.) | 7440-23-5 | Metals - Total; | Metals - Total; EMS code: NA-T | Chemical | Total | LAB
NA-T | mg/dm2/d | Sodium Total | DUST | HNO3 dig.(microwave or hot block),ICPMS | HNO3 dig.(microwave or hot block);ICPMS |  | Y | Y | Y | 267 | 7/19/2016 1:11:00 PM | 7/31/2024 | 7/31/2024 11:23:12 AM | 2008-02-07 |  | RDGIBSON |  | MassFlux | mg/dm²/d | Y | (mass flux) |  | Sodium Total (mass flux) | 7440-23-5 | Metals - Total; | Metals - Total; EMS code: NA-T | Chemical | Total | LAB
NA-T | mg/L | Sodium Total | X349 | FA:HNO3:ICP | Samples are field-acidified with HNO3; HNO3 digestion; analysis by ICP/AES. | 2 | Y | Y | Y | 6473 |  |  | 8/26/2024 2:41:32 PM | 1996-04-12 | 6/3/1998 6:51:41 PM | EMS |  | Concentration | mg/L | Y | (fl. conc.) |  | Sodium Total (fl. conc.) | 7440-23-5 | Metals - Total; | Metals - Total; EMS code: NA-T | Chemical | Total | LAB
NA-T | mg/L | Sodium Total | F038 | HNO3/HCl dig,ICP-AES | Dissolved and extractable water samples are ready for analysis once collected. Total samples are autoclaved with HNO3 and hydrochloric acid (HCI) in capped Teflon vials. Analysis is performed via Inductively Coupled Argon Plasma Atomic Emission Spectrometry (ICP-AES)(simultaneous multielement analysis), Graphite Furnace Atomic Absorption Spectrometry (GF-AAS) (low level analysis) and/or Flame Atomic Absorption Spectrometry (F-AAS). | 0.1 | Y | Y | Y | 7279 | 8/1/1996 6:04:00 PM | 5/6/2008 | 9/11/2019 8:47:42 AM | 1997-05-08 | 6/3/1998 6:51:46 PM | MPAK |  | Concentration | mg/L | Y | (fl. conc.) |  | Sodium Total (fl. conc.) | 7440-23-5 | Metals - Total; | Metals - Total; EMS code: NA-T | Chemical | Total | LAB
NA-T | mg/L | Sodium Total | ULTR | Total Metals in Water by ICPMS (Ultra) | Total Metals in Water by ICPMS (Ultra), EPA SW-846 3005A/6020A, MET-TOT-ULTRA-MS-VA This analysis is carried out using procedures adapted from "Standard Methods for the Examination of Water and Wastewater" published by the American Public Health Association, and with procedures adapted from "Test Methods for Evaluating Solid Waste" SW-846 published by the United States Environmental Protection Agency (EPA). The procedures may involve preliminary sample treatment by acid digestion, using either hotblock or microwave oven, or filtration (EPA Method 3005A). Instrumental analysis is by inductively coupled plasma - mass spectrometry (EPA Method 6020A). | 0.01 | Y | Y | Y | 177 | 2/1/2017 | 10/26/2019 | 11/13/2019 10:17:12 AM | 2012-07-05 |  | RDGIBSON |  | Concentration | mg/L | Y | (fl. conc.) |  | Sodium Total (fl. conc.) | 7440-23-5 | Metals - Total; | Metals - Total; EMS code: NA-T | Chemical | Total | LAB
NA-T | mg/L | Sodium Total | F082 | Aliquot,HNO3/HCL dig,ICPMS | Aliquot microwave digested with 3:1 HNO3/HCL. Analysis by ICPMS. For HiVol samples preliminary result are reported in mg/L and loading results are reported in ug/m3 | 0.05 | Y | Y | Y | 26620 | 1/24/2011 | 8/23/2024 | 8/26/2024 4:53:05 PM | 2015-03-11 |  | RDGIBSON |  | Concentration | mg/L | Y | (fl. conc.) |  | Sodium Total (fl. conc.) | 7440-23-5 | Metals - Total; | Metals - Total; EMS code: NA-T | Chemical | Total | LAB
NA-T | ug/g wet | Sodium Total | MTIW | MET-WET-CCMS-MID-VA | Metals in Tissue by CRC ICPMS (WET) Methodology: This method is conducted following British Columbia Lab Manual method "Metals in Animal Tissue and Vegetation (Biota) - Prescriptive". Tissue samples are homogenized and sub-sampled prior to hotblock digestion with nitric and hydrochloric acids, in combination with addition of hydrogen peroxide. Instrumental analysis is by collision cell inductively coupled plasma - mass spectrometry (modified from EPA Method 6020A). Method Limitation: This method employs a strong acid/peroxide digestion, and is intended to provide a conservative estimate of bio-available metals. Near complete recoveries are achieved for most toxicologically important metals, but elements associated with recalcitrant minerals may be only partially recovered. | 4 | Y | Y | Y | 24 | 11/25/2018 | 12/6/2018 | 7/21/2021 2:55:07 PM | 2018-12-24 |  | EMS |  | MassPerMass | µg/g | Y | (mass conc.) | (wet) | Sodium Total (wet) (mass conc.) | 7440-23-5 | Metals - Total; | Metals - Total; EMS code: NA-T | Chemical | Total | LAB
NA-T | ug/g wet | Sodium Total | BPSY | MET-WET-BPSY-CCMS-VA | Metals in Animal Tissue and Vegetation (Biota) - Prescriptive. EPA 200.3/6020B This method is conducted following British Columbia Lab Manual method "Metals in Animal Tissue and Vegetation (Biota) - Prescriptive". Tissue samples are digested with nitric and hydrochloric acids; in combination with addition of hydrogen peroxide. Instrumental analysis is by collision cell inductively coupled plasma - mass spectrometry (modified from EPA Method 6020B). | 4 | Y | Y | Y | 2 | 11/7/2019 | 11/7/2019 | 6/10/2021 1:57:10 PM | 2021-06-04 |  | EMS |  | MassPerMass | µg/g | Y | (mass conc.) | (wet) | Sodium Total (wet) (mass conc.) | 7440-23-5 | Metals - Total; | Metals - Total; EMS code: NA-T | Chemical | Total | LAB
NA-T | ug/g | Sodium Total | FS01 | D&G,Nitric/HCl:ICP | Samples are dried at 60 C, then either sieved to -100mesh (150 um) and rolled to homogenize, or finely ground with a mortar and pestle or a shatter box laboratory mill and rolled (depending on the submitter's requirements). The samples are then weighed into Teflon digestion liners and digested with HNO3 and HCl in a microwave oven. Samples are cooled, volumized, and allowed to settle overnight before being decanted. Analysis is then performed via Inductively Coupled Argon Plasma Atomic Emission Spectrometry (ICP-AES) (Simultaneous multi element analysis). In ICP Spectrometry, the liquid sample is pumped to the nebulizer, where it is convertd into an aerosol. the nebulizer carrier gas flow than transports the wet aerosol past the impact bead, where the large droplets condense out and flow to the waste container. The remaining fine aerosol is carried into the plasma. There, the high temperature desolvates the aerosol and dissociates the sample material into individual atoms and ions, which are excited to emit light at wavelengths characteristic of the elements present. A diffraction grating in the spectrometer diffracts the various wavelengths of light through exit slits to the photo-multiplier tubes, which convert the light energy to electronic signals. | 20 | Y | Y | Y | 2642 | 8/30/1996 1:52:00 PM | 10/10/2007 | 5/24/2008 1:15:03 AM | 1996-06-19 | 6/3/1998 6:51:44 PM | MPAK |  | MassPerMass | mg/kg | Y | (mass conc.) |  | Sodium Total (mass conc.) | 7440-23-5 | Metals - Total; | Metals - Total; EMS code: NA-T | Chemical | Total | LAB
NA-T | mg/L | Sodium Total | F048 | Acid dig,conc'n,nebul.,ICP-OES | This method provides an automated procedure for the Inductively Coupled Argon Plasma Optical Emission Spectrophotometric (ICP-OES) determination and quantitation of trace amounts of metals in surface, ground and precipitation waters after a manual acid digestion and concentration procedure and ultrasonic nebulisation. Organically-bound metals and metal complexes are converted to metal ions by digestion with nitric acid [method details says HCl]. To increase the signal detected by the ICP, the samples are concentrated 10 times. The liquid sample is pumped onto the piezoelectric transducer, where the ultrasonic action efficiently converts it to a fine, dense aerosol. The nebulizer carrier gas flow then transports the wet aerosol through a heated U-tube, where the solvent is vaporized. Subsequently, most of the solvent vapour is removed from the aerosol stream by passing the stream through a chilled condenser. The result is a dry, analyte-laden aerosol which is introduced to the plasma. There, the high temperature desolvates the aerosol and dissociates the sample material into individual atoms and ions, which are excited to emit light at wave lengths characteristic of the elements present. Dissolved samples are field filtered through 0.45 micron cellulose acetate filter prior to acid preservation. Samples for extractable analysis are acidified with 2 mL concentrated HNO3 per litre and allowed to stand a minimum of 24 hours before analysis. Clients providing these samples are advised to preserve samples at the collection site. | 2 | Y | Y | Y | 1600 | 11/20/1997 9:32:00 AM | 7/17/2012 7:03:00 AM | 8/28/2018 3:14:04 PM | 1998-11-13 |  | NPEPPIN |  | Concentration | mg/L | Y | (fl. conc.) |  | Sodium Total (fl. conc.) | 7440-23-5 | Metals - Total; | Metals - Total; EMS code: NA-T | Chemical | Total | LAB
NA-T | ug/g | Sodium Total | FS82 | PBM,dry60C,10mesh,SALM-ICP-MS | PBM;dry60C;10mesh;SALM-ICP-MS | 100 | Y | Y | Y | 3307 | 5/21/2002 | 7/24/2024 | 7/29/2024 5:17:09 PM | 2002-11-12 |  | UMULLER |  | MassPerMass | mg/kg | Y | (mass conc.) |  | Sodium Total (mass conc.) | 7440-23-5 | Metals - Total; | Metals - Total; EMS code: NA-T | Chemical | Total | LAB
NA-T | mg/L | Sodium Total | METN | MET-T-NP-CCMS-VA | Ultra trace metals in water are analyzed by CRC ICPMS, based on US EPA Method 6020B (July 2014). The detection limits provided can only be met for undigested samples. This procedure is intended for colorless, non-turbid, acid-preserved water samples (i.e. pristine water samples), having turbidity < 1 NTU and no odor. Where turbidity exceeds 1 NTU, and/or the sample is colored and has an odor, results may be biased low compared to true Total Metals concentrations. ALS recommends that turbidity analysis be requested on samples submitted for this test to aid with interpretation of results. Where turbidity is <1NTU, undigested metals are equivalent to total metals concentrations. | 0.02 | Y | Y | Y | 3648 | 4/3/2012 | 8/19/2024 | 8/26/2024 1:03:06 PM | 2019-08-20 |  | KATFRASE |  | Concentration | mg/L | Y | (fl. conc.) |  | Sodium Total (fl. conc.) | 7440-23-5 | Metals - Total; | Metals - Total; EMS code: NA-T | Chemical | Total | LAB
NA-T | ug/g | Sodium Total | X230 | Hom,HNO3/HCl,ICP | The sample is dried, ground to pass 100 mesh, and digested with a nitric-hydrochloric acid mixture to solubilize the solid matter and to remove the organic material by oxidation and volatilization. The completed digests are submitted to metals instrumentation personnel for metals analysis by Inductively Coupled Atomic Emission Spectrophotometry (ICP-AES), or hydride analysis by atomic absorption. | 10 | Y | Y | Y | 45 | 2/12/2002 | 12/30/2005 | 2/3/2006 1:13:13 AM | 1998-11-05 |  | NPEPPIN |  | MassPerMass | mg/kg | Y | (mass conc.) |  | Sodium Total (mass conc.) | 7440-23-5 | Metals - Total; | Metals - Total; EMS code: NA-T | Chemical | Total | LAB
NA2-O | % | Sodium (as Na2O) | ICP06 | ALS Minerals ME-ICP06 | ALS Minerals ME-ICP06. A prepared sample is fused with a lithium tetraborate/lithium metaborate flux and poured into a graphite crucible and fused in a furnace at 1000°C. The resulting melt is then cooled and dissolved in 100 mL of 4% HNO3/2% HCl followed by ICP-AES. | 0.01 | Y | Y | Y | 0 |  |  |  | 2024-01-29 |  | EMS |  | DimensionlessRatio | % | N | (ratio) |  | Sodium (as Na2O) (ratio) |  | Metals; | Metals; EMS code: NA2-O | Chemical |  | LAB
NAPH | % | Naphthalene d8 | X262 | CH2Cl2 Ext:Silica:GC/MS | The sample is placed in a separatory funnel, the pH of the sample is adjusted to alkaline and then acidic conditions, and extracted each time with dichloromethane to selectively remove compounds. The final extracts (basic and acidic) can be combined or run separately. The extract is cleaned by silica gel column chromatography, then analyzed by GC/MS. | 1 | Y | Y | Y | 1 | 8/28/2014 | 8/28/2014 | 3/11/2015 3:57:18 PM | 2015-03-11 |  | RDGIBSON |  | DimensionlessRatio | % | N | (ratio) |  | Naphthalene d8 (ratio) | 1146-65-2 | Polycyclic Aromatic Hydrocarbons (PAHs); | Polycyclic Aromatic Hydrocarbons (PAHs); EMS code: NAPH | Chemical |  | SURROGATE_RESULT
NAPH | % | Naphthalene d8 | SURR | Surrogate Recovery | Surrogate Recovery | 30 | Y | Y | Y | 964 | 1/13/2015 | 8/15/2024 | 8/16/2024 1:36:03 PM | 1996-04-12 | 6/3/1998 6:51:41 PM | EMS |  | DimensionlessRatio | % | N | (ratio) |  | Naphthalene d8 (ratio) | 1146-65-2 | Polycyclic Aromatic Hydrocarbons (PAHs); | Polycyclic Aromatic Hydrocarbons (PAHs); EMS code: NAPH | Chemical |  | SURROGATE_RESULT
NAPH | mg/L | Naphthalene d8 | MW02 | DCM extrn,GC/MS-SIM (PAH) | Water samples are liquid - liquid extracted with dichloromethane. The extract is then cleaned by silica gel column chromatography and analyzed by gas chromatography - mass spectroscopy. | 1e-05 | Y | Y | Y | 116 | 3/28/2011 | 12/3/2022 5:46:00 PM | 12/15/2022 6:20:11 PM | 1998-09-28 |  | NPEPPIN |  | Concentration | mg/L | Y | (fl. conc.) |  | Naphthalene d8 (fl. conc.) | 1146-65-2 | Polycyclic Aromatic Hydrocarbons (PAHs); | Polycyclic Aromatic Hydrocarbons (PAHs); EMS code: NAPH | Chemical |  | SURROGATE_RESULT
NAPH | % | Naphthalene d8 | FS87 | DCM,Ac/Hex extr,PAM,HRGC/LRMS | DCM;Ac/Hex extr;PAM;HRGC/LRMS | 1 | Y | Y | Y | 43 | 5/20/2015 | 5/23/2015 | 7/3/2015 3:19:04 PM | 2015-05-27 |  | RDGIBSON |  | DimensionlessRatio | % | N | (ratio) |  | Naphthalene d8 (ratio) | 1146-65-2 | Polycyclic Aromatic Hydrocarbons (PAHs); | Polycyclic Aromatic Hydrocarbons (PAHs); EMS code: NAPH | Chemical |  | SURROGATE_RESULT
NB-D | mg/L | Niobium Dissolved | F032 | Filt.,Acid,ICPMS | There is no acid digestion required for this analysis. A water sample is collected and preserved in 0.2 percent nitric acid. To analyse the extractable metals the sample is shaken, allowed to sit for at least twenty four hours, and a decanted portion is introduced to the spectrometer. When the dissolved metals are required, the samples must first be filtered through a 0.45 micron cellulose acetate filter paper and then followed by preservation in 0.2 percent nitric acid. It is important to note that different laboratories may have very different detection limits for some variables analyzed by ICPMS. | 0.0001 | Y | Y | Y | 48 | 7/30/2020 | 12/19/2023 | 1/2/2024 4:32:04 PM | 2021-11-05 |  | RDGIBSON |  | Concentration | mg/L | Y | (fl. conc.) |  | Niobium Dissolved (fl. conc.) | 7440-03-1 | Metals - Dissolved; | Metals - Dissolved; EMS code: NB-D | Chemical | Dissolved | LAB
NB-D | mg/L | Niobium Dissolved | X556 | Metals in water, dissolved, "in-bottle" digestion, ICP-SFMS | The whole water sample is filtered through a 0.45 micron cellulose acetate filter paper and the filtrate is stored in a HDPE/Teflon bottle. The container is de-capped in a Class 100 workstation, acidified to 2% (v/v) HNO3 and 1% (v/v) HCl and recapped. The sample is "in-bottle" digested in a convection oven, set at 60 degrees C for 16 hours. The sample is cooled and introduced to an inductively coupled argon plasma-sector field mass spectrometer (ICP-SFMS) by pneumatic nebulization. Each respective element is measured at a specific mass to charge ratio, m/z value expressed in amu (atomic mass units), and at a specified intrument resolution. LR: Low Resolution ~300, MR: Mid Resolution ~4000 and HR: High Resolution ~10000. Ref: Method number 02-2003, National Laboratory for Environmental Testing, CCIW, Environment Canada, Burlington, Requested April 2003. | 5e-06 | Y | Y | Y | 4158 | 6/27/2004 | 12/6/2023 2:04:00 AM | 2/14/2024 9:05:32 AM | 2008-06-04 | 6/4/2008 4:51:55 PM | RDGIBSON | RDGIBSON | Concentration | mg/L | Y | (fl. conc.) |  | Niobium Dissolved (fl. conc.) | 7440-03-1 | Metals - Dissolved; | Metals - Dissolved; EMS code: NB-D | Chemical | Dissolved | LAB
NB-D | mg/L | Niobium Dissolved | METC | MET-D-NP-CCMS-VA | Ultra trace metals in water are analyzed by CRC ICPMS, based on US EPA Method 6020B (July 2014). This procedure is intended for pristine field-filtered acid-preserved water samples. The detection limits (LOR) for this test are based on lab instrumental analysis only, not including filtration. ALS-supplied field filtration equipment does not support these LOR. Therefore, because of the high probability of false positives due to filtration, it is strongly recommended that a filtration blank be analysed to aid in data interpretation. | 0.0001 | Y | Y | Y | 394 | 3/30/2020 | 4/9/2024 | 4/29/2024 4:53:06 PM | 2022-10-07 |  | EMS |  | Concentration | mg/L | Y | (fl. conc.) |  | Niobium Dissolved (fl. conc.) | 7440-03-1 | Metals - Dissolved; | Metals - Dissolved; EMS code: NB-D | Chemical | Dissolved | LAB
NB-E | mg/L | Niobium Extractable | X553 | Metals in water, extractable, ICP-SFMS | The whole water sample is collected in a HDPE/Teflon bottle. In a Class 100 workstation it is preserved with concentrated Nitric Acid at concentration of 0.2 percent (v/v). The acid preservative extracts the metals from material in the sample and a decanted portion is introduced directly to an inductively coupled argon plasma-sector field mass spectrometer (ICP-SFMS) by pneumatic nebulization. Each respective element is measured at a specific mass to charge ratio, m/z value expressed in amu (atomic mass units), and at a specified instrument resolution. LR: Low Resolution ~300, MR: Mid Resolution ~4000 and HR: High Resolution ~10000. Ref: Method number 02-2061, National Laboratory for Environmental Testing, CCIW, Environment Canada, Burlington. Requested April 2003. |  | Y | Y | Y | 1715 | 6/27/2004 | 9/21/2007 | 4/18/2018 9:28:02 AM | 2009-05-15 |  | RDGIBSON |  | Concentration | mg/L | Y | (fl. conc.) |  | Niobium Extractable (fl. conc.) | 7440-03-1 | Metals - Extractable; | Metals - Extractable; EMS code: NB-E | Chemical | Extractable | LAB
NB-T | ppm (S) | Niobium Total | MS41 | ALS Minerals ME-MS41 | ALS Minerals ME-MS41. A prepared sample is digested with aqua regia in a graphite heating block. After cooling; the resulting solution is diluted to with deionized water; mixed and analyzed by ICP-AES. Samples are then analyzed by ICP-MS for the remaining suite of elements. | 0.05 | Y | Y | Y | 0 |  |  |  | 2024-01-29 |  | EMS |  | DimensionlessRatio | ppm | N | (ratio) |  | Niobium Total (ratio) | 7440-03-1 | Metals - Total; | Metals - Total; EMS code: NB-T | Chemical | Total | LAB
NB-T | ppm (S) | Niobium Total | MS61 | ALS Minerals ME-MS61 | ALS Minerals ME-MS61. A prepared sample is digested with HNO3; HF; and HClO4 in a graphite heating block to near dryness. After cooling; the sample is boiled in diluted HCl. Bulk to volume with deionized water; mixed and analyzed by ICP-AES and ICP-MS for all elements. | 0.1 | Y | Y | Y | 0 |  |  |  | 2024-01-29 |  | EMS |  | DimensionlessRatio | ppm | N | (ratio) |  | Niobium Total (ratio) | 7440-03-1 | Metals - Total; | Metals - Total; EMS code: NB-T | Chemical | Total | LAB
NB-T | ppm (S) | Niobium Total | 6042 | Aqua Regia by ICPMS | These were sediment/soils digested with Aqua-Regia (Nitric Acid & Hydrochloric Acids) and then analysis by ICPMS. Results would be reported in ppm s. | 0.002 | Y | Y | Y | 71 | 5/14/2022 | 7/27/2024 | 7/29/2024 9:13:08 AM | 2022-06-28 |  | EMS |  | DimensionlessRatio | ppm | N | (ratio) |  | Niobium Total (ratio) | 7440-03-1 | Metals - Total; | Metals - Total; EMS code: NB-T | Chemical | Total | LAB
NB-T | ppm (S) | Niobium Total | 6041 | Aqua Regia by ICPMS | These were sediment/soils digested with Aqua-Regia (Nitric Acid & Hydrochloric Acids) and then analysis by ICPMS. Results would be reported in mg/kg and/or %. | 0.002 | Y | Y | Y | 21 | 10/13/2019 | 8/5/2020 | 8/10/2020 3:52:06 PM | 2019-09-06 | 10/21/2019 11:55:45 AM | KATFRASE | KATFRASE | DimensionlessRatio | ppm | N | (ratio) |  | Niobium Total (ratio) | 7440-03-1 | Metals - Total; | Metals - Total; EMS code: NB-T | Chemical | Total | LAB
NB-T | ppm (S) | Niobium Total | MS81 | ALS Minerals: ME-MS81 | ALS Minerals: ME-MS81. A prepared sample (0.100 g) is added to lithium metaborate/lithium tetraborate flux; mixed well and fused in a furnace at 1025°C. The resulting melt is then cooled and dissolved in an acid mixture containing nitric; hydrochloric and hydrofluoric acids. This solution is then analyzed by inductively coupled plasma - mass spectrometry. | 0.1 | Y | Y | Y | 0 |  |  |  | 2024-01-29 |  | EMS |  | DimensionlessRatio | ppm | N | (ratio) |  | Niobium Total (ratio) | 7440-03-1 | Metals - Total; | Metals - Total; EMS code: NB-T | Chemical | Total | LAB
NB-T | mg/L | Niobium Total | F082 | Aliquot,HNO3/HCL dig,ICPMS | Aliquot microwave digested with 3:1 HNO3/HCL. Analysis by ICPMS. For HiVol samples preliminary result are reported in mg/L and loading results are reported in ug/m3 | 0.0001 | Y | Y | Y | 69 | 7/30/2020 | 8/16/2024 | 8/21/2024 11:26:07 AM | 2021-11-03 |  | RDGIBSON |  | Concentration | mg/L | Y | (fl. conc.) |  | Niobium Total (fl. conc.) | 7440-03-1 | Metals - Total; | Metals - Total; EMS code: NB-T | Chemical | Total | LAB
NB-T | mg/L | Niobium Total | METN | MET-T-NP-CCMS-VA | Ultra trace metals in water are analyzed by CRC ICPMS, based on US EPA Method 6020B (July 2014). The detection limits provided can only be met for undigested samples. This procedure is intended for colorless, non-turbid, acid-preserved water samples (i.e. pristine water samples), having turbidity < 1 NTU and no odor. Where turbidity exceeds 1 NTU, and/or the sample is colored and has an odor, results may be biased low compared to true Total Metals concentrations. ALS recommends that turbidity analysis be requested on samples submitted for this test to aid with interpretation of results. Where turbidity is <1NTU, undigested metals are equivalent to total metals concentrations. | 0.0001 | Y | Y | Y | 280 | 8/6/2019 | 4/4/2024 | 4/29/2024 4:53:06 PM | 2021-06-10 |  | RDGIBSON |  | Concentration | mg/L | Y | (fl. conc.) |  | Niobium Total (fl. conc.) | 7440-03-1 | Metals - Total; | Metals - Total; EMS code: NB-T | Chemical | Total | LAB
NB-T | mg/L | Niobium Total | X555 | Metals in water, total, "in-bottle" digestion, ICP-SFMS | The whole water sample collected in a HDPE/Teflon bottle, is de-capped in a Class 100 workstation, acidified to 2% (v/v) HNO3 and 1% (v/v) HCl and recapped. The sample is "in-bottle" digested in a convection oven, set at 60 degrees C for 16 hours. The sample is cooled and introduced to an inductively coupled argon plasma-sector field mass spectrometer (ICP-SFMS) by pneumatic nebulization. Each respective element is measured at a specific mass to charge ratio, m/z value expressed in amu (atomic mass units), and at resolution. LR: Low Resolution ~300, MR: Mid Resolution ~4000 and HR: High a specified instrument Resolution ~10000. Ref: Method number 02-2003, National Laboratory for Environmental Testing, CCIW, Environment Canada, Burlington. Requested April 2003. | 1e-06 | Y | Y | Y | 15775 | 6/27/2004 | 12/6/2023 2:08:00 AM | 2/14/2024 9:05:32 AM | 2009-03-31 |  | JCHUDAK |  | Concentration | mg/L | Y | (fl. conc.) |  | Niobium Total (fl. conc.) | 7440-03-1 | Metals - Total; | Metals - Total; EMS code: NB-T | Chemical | Total | LAB
ND-T | ppm (S) | Neodymium | MS81 | ALS Minerals: ME-MS81 | ALS Minerals: ME-MS81. A prepared sample (0.100 g) is added to lithium metaborate/lithium tetraborate flux; mixed well and fused in a furnace at 1025°C. The resulting melt is then cooled and dissolved in an acid mixture containing nitric; hydrochloric and hydrofluoric acids. This solution is then analyzed by inductively coupled plasma - mass spectrometry. | 0.1 | Y | Y | Y | 0 |  |  |  | 2024-01-29 |  | EMS |  | DimensionlessRatio | ppm | N | (ratio) |  | Neodymium (ratio) |  | Metals; | Metals; EMS code: ND-T | Chemical |  | LAB
NDE | ug/g | Nonylphenol Diethoxylates | A849 | J. of Chrom. A849 (1999) 467-482 | Sample is derivatized and injected directly. Each compound is separated by reversed phase HPLC and analyzed by Electro Spray Ionization MS/MS detection using Triple Quadrupole MS/MS detector. Internal standards are used for quantitation. | 0.0625 | Y | Y | Y | 8 | 3/14/2020 | 3/16/2020 | 2/21/2023 2:10:31 PM | 2022-08-08 |  | EMS |  | MassPerMass | mg/kg | Y | (mass conc.) |  | Nonylphenol Diethoxylates (mass conc.) | 27176-93-8 | Organics; | Organics; EMS code: NDE | Chemical |  | LAB
NDIN | mg/L | Nitrogen Dissolved Inorganic | XCAL | Calculation (mg/L) | Calculation (mg/L) | 0.025 | Y | Y | Y | 41 | 1/26/2021 | 4/1/2024 | 4/11/2024 11:55:20 AM | 2021-11-03 | 11/3/2021 11:23:26 AM | RDGIBSON | RDGIBSON | Concentration | mg/L | Y | (fl. conc.) |  | Nitrogen Dissolved Inorganic (fl. conc.) |  | Nutrients; | Nutrients; EMS code: NDIN | Chemical | Dissolved | LAB
NFDH | mg/L | Nonafluoro-3,6-dioxaheptanoic acid (NFDHA) | E3533 | MECP E3533 | PFAs in water by direct injection LC/MS/MS | 2e-05 | Y | Y | Y | 2 | 6/3/2024 | 6/3/2024 | 8/21/2024 3:32:03 PM | 2023-11-15 | 6/27/2024 5:06:31 PM | EMS | EMS | Concentration | mg/L | Y | (fl. conc.) |  | Nonafluoro-3,6-dioxaheptanoic acid (NFDHA) (fl. conc.) | 151772-58-6 |  | EMS code: NFDH | Chemical |  | LAB
NGST | m3 | Natural Gas (Sweet) | X469 | Blowdown Calculation | Blowdown Calculation | 0 | Y | Y | Y | 1636 | 1/1/1997 | 12/29/1999 | 1/31/2000 5:17:03 PM | 1996-12-09 | 12/14/2001 10:31:53 AM | MPAK | NPEPPIN | Volume | m³ | Y | (vol.) |  | Natural Gas (Sweet) (vol.) | 8006-14-2 | Volatile Organic Compounds (VOCs); | Volatile Organic Compounds (VOCs); EMS code: NGST | Chemical |  | LAB
NGST | m3 | Natural Gas (Sweet) | X468 | Starting Gas Calculation | Starting Gas Calculation | 0 | Y | Y | Y | 394 | 1/1/1997 | 12/1/1999 | 1/31/2000 5:17:00 PM | 1996-12-09 | 12/14/2001 10:31:53 AM | MPAK | NPEPPIN | Volume | m³ | Y | (vol.) |  | Natural Gas (Sweet) (vol.) | 8006-14-2 | Volatile Organic Compounds (VOCs); | Volatile Organic Compounds (VOCs); EMS code: NGST | Chemical |  | LAB
NH04 | mg/L | Ammonia; Un-ionized (as N); 4?C; lab pH | CALN | Un-ionized Ammonia Calculation | The concentration of unionized ammonia in the water referred to in subsection (1) must be determined in accordance with the following formula: total ammonia × 1 ÷ (1 +10^(pKa pH)) where total ammonia is the concentration of total ammonia in that water namely, unionized ammonia (NH3) plus ionized ammonia (NH4+) determined in accordance with subsection (4), expressed in mg/L as nitrogen (N) pKa is 0.09018 + 2729.92/T, where T is the ambient water temperature in kelvin and pH is the pH of that water. REF: from the new National Wastewater Systems Effluent Regulations | 0.001 | Y | Y | Y | 17 | 8/20/2021 | 1/24/2023 | 2/9/2023 1:13:05 PM | 2021-12-30 |  | RDGIBSON |  | Concentration | mg/L | Y | (fl. conc.) |  | Ammonia; Un-ionized (as N); 4C; lab pH (fl. conc.) | 7664-41-7 | Nutrients; | Nutrients; EMS code: NH04 | Chemical |  | LAB
NH15 | mg/L | Ammonia; Un-ionized (as N); 15?C; lab pH | CALN | Un-ionized Ammonia Calculation | The concentration of unionized ammonia in the water referred to in subsection (1) must be determined in accordance with the following formula: total ammonia × 1 ÷ (1 +10^(pKa pH)) where total ammonia is the concentration of total ammonia in that water namely, unionized ammonia (NH3) plus ionized ammonia (NH4+) determined in accordance with subsection (4), expressed in mg/L as nitrogen (N) pKa is 0.09018 + 2729.92/T, where T is the ambient water temperature in kelvin and pH is the pH of that water. REF: from the new National Wastewater Systems Effluent Regulations | 0.001 | Y | Y | Y | 17 | 8/20/2021 | 1/24/2023 | 2/9/2023 1:13:05 PM | 2021-12-30 |  | RDGIBSON |  | Concentration | mg/L | Y | (fl. conc.) |  | Ammonia; Un-ionized (as N); 15C; lab pH (fl. conc.) | 7664-41-7 | Nutrients; | Nutrients; EMS code: NH15 | Chemical |  | LAB
NH20 | mg/L | Ammonia; Un-ionized (as N); 20?C; lab pH | CALN | Un-ionized Ammonia Calculation | The concentration of unionized ammonia in the water referred to in subsection (1) must be determined in accordance with the following formula: total ammonia × 1 ÷ (1 +10^(pKa pH)) where total ammonia is the concentration of total ammonia in that water namely, unionized ammonia (NH3) plus ionized ammonia (NH4+) determined in accordance with subsection (4), expressed in mg/L as nitrogen (N) pKa is 0.09018 + 2729.92/T, where T is the ambient water temperature in kelvin and pH is the pH of that water. REF: from the new National Wastewater Systems Effluent Regulations | 0.001 | Y | Y | Y | 17 | 8/20/2021 | 1/24/2023 | 2/9/2023 1:13:05 PM | 2021-12-30 |  | RDGIBSON |  | Concentration | mg/L | Y | (fl. conc.) |  | Ammonia; Un-ionized (as N); 20C; lab pH (fl. conc.) | 7664-41-7 | Nutrients; | Nutrients; EMS code: NH20 | Chemical |  | LAB
NH25 | mg/L | Ammonia; Un-ionized (as N); 25?C; lab pH | CALN | Un-ionized Ammonia Calculation | The concentration of unionized ammonia in the water referred to in subsection (1) must be determined in accordance with the following formula: total ammonia × 1 ÷ (1 +10^(pKa pH)) where total ammonia is the concentration of total ammonia in that water namely, unionized ammonia (NH3) plus ionized ammonia (NH4+) determined in accordance with subsection (4), expressed in mg/L as nitrogen (N) pKa is 0.09018 + 2729.92/T, where T is the ambient water temperature in kelvin and pH is the pH of that water. REF: from the new National Wastewater Systems Effluent Regulations | 0.001 | Y | Y | Y | 17 | 8/20/2021 | 1/24/2023 | 2/9/2023 1:13:05 PM | 2021-12-30 |  | RDGIBSON |  | Concentration | mg/L | Y | (fl. conc.) |  | Ammonia; Un-ionized (as N); 25C; lab pH (fl. conc.) | 7664-41-7 | Nutrients; | Nutrients; EMS code: NH25 | Chemical |  | LAB
NH4- | mg/L | Ammonium | 5072 | Auto. Color. Diazo/Nitr. | Auto. Color. Diazo/Nitr. | 0.009 | Y | Y | Y | 110 |  |  | 6/18/2018 9:20:13 AM | 1996-04-12 | 6/3/1998 6:51:37 PM | EMS |  | Concentration | mg/L | Y | (fl. conc.) |  | Ammonium (fl. conc.) | 14798-03-9 | Nutrients; | Nutrients; EMS code: NH4- | Chemical |  | LAB
NH4- | mg/L | Ammonium | 1351 | Automated Bertholot meth | Automated Bertholot meth | 0.01 | Y | Y | Y | 2281 |  |  | 10/24/1996 2:10:08 PM | 1996-04-12 | 6/3/1998 6:51:37 PM | EMS |  | Concentration | mg/L | Y | (fl. conc.) |  | Ammonium (fl. conc.) | 14798-03-9 | Nutrients; | Nutrients; EMS code: NH4- | Chemical |  | LAB
NH4- | mg/dm2/d | Ammonium | EQ15 | EQUIS (mg/dm2/d) | EQUIS (mg/dm2/d) | 0.002 | N | Y | Y | 5 |  |  | 10/22/1996 8:20:30 PM | 1996-04-12 | 6/3/1998 6:51:37 PM | EMS |  | MassFlux | mg/dm²/d | Y | (mass flux) |  | Ammonium (mass flux) | 14798-03-9 | Nutrients; | Nutrients; EMS code: NH4- | Chemical |  | LAB
NH4- | mg/L | Ammonium | 5071 | Ion Chromatog -Cation I | (a) For the Soluble Particulates Pkg. a portion of filtrate from the Particulate - Total procedure made up to 500 mL is analyzed by IC. Intermediate results are reported in mg/L. Loading results are calculated and reported with units of mg/dm2/d.(b) For precipitation samples, all samples are filtered through a 0.22 um fritted glass filter prior to injection. The cations of interest are separated through a cation "guard" and cation "separator" column with a methane sulphonic acid eluent. After the cations are separated they exit at various times from the column in a background of eluent. A self regenerating membrane suppressor is attached to the end of the column to neutralize the acid before detection and form the corresponding hydroxide. Note that ammonium is reported as NH4 not as N. (c) For HiVol samples, discs cut from the HiVol filter are extracted with DI water and the soluble anions are analyzed by IC. Intermediate results are reported in units of mg/L and HiVol loading results are reported as ug/m3. (*) Note that ammonium is reported as NH4 not as N | 0.01 | Y | Y | Y | 4024 | 4/30/1997 3:50:00 PM | 10/4/2006 | 11/8/2006 1:00:13 AM | 1996-04-12 | 6/3/1998 6:51:37 PM | EMS |  | Concentration | mg/L | Y | (fl. conc.) |  | Ammonium (fl. conc.) | 14798-03-9 | Nutrients; | Nutrients; EMS code: NH4- | Chemical |  | LAB
NH4- | ug/m3 | Ammonium | 5018 | H2O Extr,Ion Chr-Cat I | For HiVol samples, discs cut from the HiVol filter are extracted with DI water and the soluble ions analyzed by IC. Intermediate results are reported as mg/L and HiVol loading results as ug/m3. | 0.01 | Y | Y | Y | 14 |  |  | 10/24/1996 6:37:10 AM | 1996-04-12 | 6/3/1998 6:51:37 PM | EMS |  | AirConcentration | µg/m³ | Y | (air conc.) |  | Ammonium (air conc.) | 14798-03-9 | Nutrients; | Nutrients; EMS code: NH4- | Chemical |  | LAB
NH4- | mg/L | Ammonium | EQ01 | EQUIS (mg/L) | EQUIS (mg/L) | 0.009 | N | Y | Y | 125 |  |  | 10/22/1996 11:12:48 PM | 1996-04-12 | 6/3/1998 6:51:45 PM | EMS |  | Concentration | mg/L | Y | (fl. conc.) |  | Ammonium (fl. conc.) | 14798-03-9 | Nutrients; | Nutrients; EMS code: NH4- | Chemical |  | LAB
NI-D | mg/L | Nickel Dissolved | X072 | FA,HNO3,Graph.Furnace,AA | Samples are field-acidified with HNO3; HNO3 digestion; analysis by GFAA. | 0.001 | Y | Y | Y | 29 |  |  | 10/24/1996 1:00:47 PM | 1996-04-12 | 6/3/1998 6:51:37 PM | EMS |  | Concentration | mg/L | Y | (fl. conc.) |  | Nickel Dissolved (fl. conc.) | 7440-02-0 | Metals - Dissolved; | Metals - Dissolved; EMS code: NI-D | Chemical | Dissolved | LAB
NI-D | mg/L | Nickel Dissolved | X170 | ICP | This code is a catch-all for a large variety of ICP methods including water and also solids digested into solutions as an intermediate result. Different procedures will have greatly different detection limits. (a) For air particulates, preliminary results are reported in mg/L and loading results as mg/dm2/d. (b) For Hi Vol samples preliminiary results are same as reported in mg/L and loading results as ug/m3. (c) Note that for Rain samples there is no filtration nor digestion steps. | 0.05 | Y | Y | Y | 14176 | 4/11/1996 6:02:00 PM | 10/29/2019 | 6/12/2020 1:21:12 PM | 1996-04-12 | 6/3/1998 6:51:37 PM | EMS |  | Concentration | mg/L | Y | (fl. conc.) |  | Nickel Dissolved (fl. conc.) | 7440-02-0 | Metals - Dissolved; | Metals - Dissolved; EMS code: NI-D | Chemical | Dissolved | LAB
NI-D | mg/L | Nickel Dissolved | MWID | Metals in Water by ICPMS (digested) | Metals in Water by ICPMS (digested) Based on EPA 200.8 | 0.001 | Y | Y | Y | 31791 | 10/26/2018 5:20:00 PM | 10/26/2018 5:20:00 PM | 7/31/2024 1:46:41 PM | 2009-12-24 |  | RDGIBSON |  | Concentration | mg/L | Y | (fl. conc.) |  | Nickel Dissolved (fl. conc.) | 7440-02-0 | Metals - Dissolved; | Metals - Dissolved; EMS code: NI-D | Chemical | Dissolved | LAB
NI-D | mg/L | Nickel Dissolved | 0020 | HNO3 Dig:AA Analysis | HNO3 Dig:AA Analysis | 0.1 | Y | Y | Y | 10 | 9/29/2000 | 12/1/2010 | 12/9/2010 3:55:27 PM | 1996-04-12 | 6/3/1998 6:51:37 PM | EMS |  | Concentration | mg/L | Y | (fl. conc.) |  | Nickel Dissolved (fl. conc.) | 7440-02-0 | Metals - Dissolved; | Metals - Dissolved; EMS code: NI-D | Chemical | Dissolved | LAB
NI-D | mg/L | Nickel Dissolved | X116 | FF,FA,HNO3,Graph.Furn,AA | Samples are field-filtered (0.45 um), field-acidified with HNO3; analysis by GFAA. | 0.001 | Y | Y | Y | 6992 |  |  | 8/26/2024 2:41:32 PM | 1996-04-12 | 6/3/1998 6:51:37 PM | EMS |  | Concentration | mg/L | Y | (fl. conc.) |  | Nickel Dissolved (fl. conc.) | 7440-02-0 | Metals - Dissolved; | Metals - Dissolved; EMS code: NI-D | Chemical | Dissolved | LAB
NI-D | mg/L | Nickel Dissolved | X556 | Metals in water, dissolved, "in-bottle" digestion, ICP-SFMS | The whole water sample is filtered through a 0.45 micron cellulose acetate filter paper and the filtrate is stored in a HDPE/Teflon bottle. The container is de-capped in a Class 100 workstation, acidified to 2% (v/v) HNO3 and 1% (v/v) HCl and recapped. The sample is "in-bottle" digested in a convection oven, set at 60 degrees C for 16 hours. The sample is cooled and introduced to an inductively coupled argon plasma-sector field mass spectrometer (ICP-SFMS) by pneumatic nebulization. Each respective element is measured at a specific mass to charge ratio, m/z value expressed in amu (atomic mass units), and at a specified intrument resolution. LR: Low Resolution ~300, MR: Mid Resolution ~4000 and HR: High Resolution ~10000. Ref: Method number 02-2003, National Laboratory for Environmental Testing, CCIW, Environment Canada, Burlington, Requested April 2003. | 2e-05 | Y | Y | Y | 4234 | 6/11/2003 | 12/6/2023 2:04:00 AM | 2/14/2024 9:05:32 AM | 2004-01-19 |  | UMULLER |  | Concentration | mg/L | Y | (fl. conc.) |  | Nickel Dissolved (fl. conc.) | 7440-02-0 | Metals - Dissolved; | Metals - Dissolved; EMS code: NI-D | Chemical | Dissolved | LAB
NI-D | mg/L | Nickel Dissolved | F039 | Field filt + acid, ICP-AES | Dissolved and extractable water samples are ready for analysis once collected. Total samples are autoclaved with HNO3 and hydrochloric acid (HCI) in capped Teflon vials. Analysis is performed via Inductively Coupled Argon Plasma Atomic Emission Spectrometry (ICP-AES)(simultaneous multielement analysis), Graphite Furnace Atomic Absorption Spectrometry (GF-AAS)(low level analysis) and/or Flame Atomic Absorption Spectrometry (F-AAS). | 0.02 | Y | Y | Y | 5995 | 7/29/1996 6:03:00 PM | 7/5/2019 1:21:00 PM | 9/11/2019 8:47:50 AM | 1997-05-15 | 6/3/1998 6:51:47 PM | MPAK |  | Concentration | mg/L | Y | (fl. conc.) |  | Nickel Dissolved (fl. conc.) | 7440-02-0 | Metals - Dissolved; | Metals - Dissolved; EMS code: NI-D | Chemical | Dissolved | LAB
NI-D | mg/L | Nickel Dissolved | 0050 | Graphite Furnace AA | Graphite Furnace AA | 0.001 | Y | Y | Y | 533 |  |  | 2/13/1998 12:49:06 AM | 1996-04-12 | 6/3/1998 6:51:37 PM | EMS |  | Concentration | mg/L | Y | (fl. conc.) |  | Nickel Dissolved (fl. conc.) | 7440-02-0 | Metals - Dissolved; | Metals - Dissolved; EMS code: NI-D | Chemical | Dissolved | LAB
NI-D | mg/L | Nickel Dissolved | F068 | HNO3dig,TeflonTube/Autoclave,ICP-MS | This method is applicable to determiniation of dissolved, extractable, and total recoverable metals in ground waters, surface waters, drinking waters, and wastewaters. Dissolved, extractable, and total recoverable sample fractions are acidified to pH <1.5 with analytical grade nitric acid (HNO3) to prevent the metals from adsorbing to the container material or precipitating. Extractable and total recoverable water samples are collected in 250 ml acid washed polyethylene bottles and are preserved with Nitric acid (0.5 ml acid for 100 mLs of sample). All preserved samples are stored at room temperature and can be held for up to 6 months. Unfiltered, un-preserved samples are stored at 4C for no more than 30 days. For dissolved and extractable water fractions, no sample preparation is necessary. Samples are simply poured in autosampler tubes (i.e. 15 ml centrifuge tubes), and placed on the autosampler and analyzed. Samples for total recoverable metals are digested with Nitric acid in special Teflon tubes in an autoclave prior to analyses. | 5e-05 | Y | Y | Y | 1398 | 6/1/1999 1:21:00 PM | 4/17/2015 1:06:00 PM | 8/11/2021 12:04:11 PM | 1999-06-28 |  | NPEPPIN |  | Concentration | mg/L | Y | (fl. conc.) |  | Nickel Dissolved (fl. conc.) | 7440-02-0 | Metals - Dissolved; | Metals - Dissolved; EMS code: NI-D | Chemical | Dissolved | LAB
NI-D | mg/L | Nickel Dissolved | MSEA | Metals in Seawater (undigested) by (CRC) ICP-MS | This method describes the multi-elemental determination of trace elements by Collision/Reaction Cell (CRC) ICP-MS. The method measures ions produced by a radio-frequency inductively coupled plasma. Analyte species originating in a liquid are pneumatically nebulized and the resulting aerosol transported by argon gas into the plasma torch. The ions produced are entrained in the plasma gas and extracted, by means of a differentially pumped vacuum interface, into a mass spectrometer. The ions flow through a Collision/Reaction Cell (CRC) where depending on the mode of analysis either the ions passes through the cell unhindered (no gas - classical ICP-MS mode), or collisions/reactions chemistry takes place between the ions and a neutral gas such as hydrogen or helium. Any polyatomic ions which give interference on the analytes are removed. The ions are then sorted according to their mass-to-charge ratios by a quadrupole mass spectrometer. The ions are then sorted according to their mass-to-charge ratios by a quadrupole mass spectrometer. The ions transmitted through the quadrupole are quantified by a channel electron multiplier and the ion information processed by a data handling system. Although most polyatomic interference is removed, any additional interference must be assessed and valid corrections applied or the data flagged to indicate problems. Interference correction must include compensation for background ions contributed by the plasma gas, reagents, and constituents of the sample matrix. Instrument drift as well as suppressions or enhancements of instrument response caused by the sample matrix must be corrected for by the use of internal standards. | 0.0001 | Y | Y | Y | 4 | 4/12/2023 | 4/29/2024 | 5/16/2024 2:00:14 PM | 2009-09-08 |  | RDGIBSON |  | Concentration | mg/L | Y | (fl. conc.) |  | Nickel Dissolved (fl. conc.) | 7440-02-0 | Metals - Dissolved; | Metals - Dissolved; EMS code: NI-D | Chemical | Dissolved | LAB
NI-D | mg/L | Nickel Dissolved | X285 | HNO3 Dig:ICP Analysis | HNO3 Dig:ICP Analysis | 0.02 | Y | Y | Y | 742 | 1/30/2002 | 9/16/2003 3:18:00 PM | 3/18/2004 10:28:34 AM | 1996-04-12 | 6/3/1998 6:51:41 PM | EMS |  | Concentration | mg/L | Y | (fl. conc.) |  | Nickel Dissolved (fl. conc.) | 7440-02-0 | Metals - Dissolved; | Metals - Dissolved; EMS code: NI-D | Chemical | Dissolved | LAB
NI-D | mg/L | Nickel Dissolved | METC | MET-D-NP-CCMS-VA | Ultra trace metals in water are analyzed by CRC ICPMS, based on US EPA Method 6020B (July 2014). This procedure is intended for pristine field-filtered acid-preserved water samples. The detection limits (LOR) for this test are based on lab instrumental analysis only, not including filtration. ALS-supplied field filtration equipment does not support these LOR. Therefore, because of the high probability of false positives due to filtration, it is strongly recommended that a filtration blank be analysed to aid in data interpretation. | 5e-05 | Y | Y | Y | 1776 | 4/3/2012 | 8/15/2024 | 8/23/2024 9:50:35 AM | 2019-08-20 |  | KATFRASE |  | Concentration | mg/L | Y | (fl. conc.) |  | Nickel Dissolved (fl. conc.) | 7440-02-0 | Metals - Dissolved; | Metals - Dissolved; EMS code: NI-D | Chemical | Dissolved | LAB
NI-D | mg/L | Nickel Dissolved | X562 | Dissolved metals: FF, FA, HNO3, HR-ICPMS | Ultra trace metals in water are analyzed by high resolution inductively coupled plasma mass spectrometry (HR-ICPMS), based on US EPA Method 200.8, (Oct. 1999). This procedure is intended for pristine field-filtered acid-preserved water samples. The detection limits (LOR) for this test are based on lab instrumental analysis only, not including filtration. ALS-supplied field filtration equipment does not support these LOR. Therefore, because of the high probability of false positives due to filtration, it is strongly recommended that a filtration blank be analyzed to aid in data interpretation. | 5e-05 | Y | Y | Y | 1264 | 6/8/2015 | 10/19/2023 | 11/9/2023 7:41:27 AM | 2015-10-22 |  | RDGIBSON |  | Concentration | mg/L | Y | (fl. conc.) |  | Nickel Dissolved (fl. conc.) | 7440-02-0 | Metals - Dissolved; | Metals - Dissolved; EMS code: NI-D | Chemical | Dissolved | LAB
NI-D | mg/L | Nickel Dissolved | X028 | FA,HNO3,AA Analysis | FA;HNO3;AA Analysis | 0.01 | Y | Y | Y | 30 |  |  | 10/24/1996 1:00:35 PM | 1996-04-12 | 6/3/1998 6:51:37 PM | EMS |  | Concentration | mg/L | Y | (fl. conc.) |  | Nickel Dissolved (fl. conc.) | 7440-02-0 | Metals - Dissolved; | Metals - Dissolved; EMS code: NI-D | Chemical | Dissolved | LAB
NI-D | mg/L | Nickel Dissolved | EQ01 | EQUIS (mg/L) | EQUIS (mg/L) | 0.01 | N | Y | Y | 1907 |  |  | 10/24/1996 3:16:55 PM | 1996-04-12 | 6/3/1998 6:51:41 PM | EMS |  | Concentration | mg/L | Y | (fl. conc.) |  | Nickel Dissolved (fl. conc.) | 7440-02-0 | Metals - Dissolved; | Metals - Dissolved; EMS code: NI-D | Chemical | Dissolved | LAB
NI-D | mg/L | Nickel Dissolved | X085 | LF,LA,HNO3,AA | Samples are lab-filtered (0.45 um), and lab-acidified (HNO3); analysis by AA. | 0.01 | Y | Y | Y | 207 | 12/12/1996 | 9/18/2015 | 10/28/2015 9:54:16 AM | 1996-04-12 | 6/3/1998 6:51:37 PM | EMS |  | Concentration | mg/L | Y | (fl. conc.) |  | Nickel Dissolved (fl. conc.) | 7440-02-0 | Metals - Dissolved; | Metals - Dissolved; EMS code: NI-D | Chemical | Dissolved | LAB
NI-D | mg/L | Nickel Dissolved | F032 | Filt.,Acid,ICPMS | There is no acid digestion required for this analysis. A water sample is collected and preserved in 0.2 percent nitric acid. To analyse the extractable metals the sample is shaken, allowed to sit for at least twenty four hours, and a decanted portion is introduced to the spectrometer. When the dissolved metals are required, the samples must first be filtered through a 0.45 micron cellulose acetate filter paper and then followed by preservation in 0.2 percent nitric acid. It is important to note that different laboratories may have very different detection limits for some variables analyzed by ICPMS. | 0.001 | Y | Y | Y | 80982 | 8/16/1996 2:35:00 PM | 8/26/2024 | 8/26/2024 4:53:05 PM | 1997-01-16 | 8/15/2019 10:58:04 AM | MPAK | KATFRASE | Concentration | mg/L | Y | (fl. conc.) |  | Nickel Dissolved (fl. conc.) | 7440-02-0 | Metals - Dissolved; | Metals - Dissolved; EMS code: NI-D | Chemical | Dissolved | LAB
NI-E | mg/L | Nickel Extractable | F047 | 24hr HNO3 dil'n,ICP-MS | There is no acid digestion required for this analysis. A water sample is collected and preserved in 0.2 percent nitric acid. To analyse the extractable metals the sample is shaken, allowed to sit for at least twenty four hours, and a decanted portion is introduced to the spectrometer. When the dissolved metals are required, the samples must first be filtered through a 0.45 micron cellulose acetate filter paper and then followed by preservation in 0.2 percent nitric acid. It is important to note that different laboratories may have very different detection limits for some variables analyzed by ICP/MS. | 2e-05 | Y | Y | Y | 2676 | 4/16/1998 10:21:00 AM | 6/26/2003 | 9/11/2019 8:46:18 AM | 1997-10-20 | 6/3/1998 6:51:47 PM | NPEPPIN |  | Concentration | mg/L | Y | (fl. conc.) |  | Nickel Extractable (fl. conc.) | 7440-02-0 | Metals - Extractable; | Metals - Extractable; EMS code: NI-E | Chemical | Extractable | LAB
NI-E | mg/L | Nickel Extractable | X553 | Metals in water, extractable, ICP-SFMS | The whole water sample is collected in a HDPE/Teflon bottle. In a Class 100 workstation it is preserved with concentrated Nitric Acid at concentration of 0.2 percent (v/v). The acid preservative extracts the metals from material in the sample and a decanted portion is introduced directly to an inductively coupled argon plasma-sector field mass spectrometer (ICP-SFMS) by pneumatic nebulization. Each respective element is measured at a specific mass to charge ratio, m/z value expressed in amu (atomic mass units), and at a specified instrument resolution. LR: Low Resolution ~300, MR: Mid Resolution ~4000 and HR: High Resolution ~10000. Ref: Method number 02-2061, National Laboratory for Environmental Testing, CCIW, Environment Canada, Burlington. Requested April 2003. | 2e-05 | Y | Y | Y | 2597 | 5/22/2003 | 9/21/2007 | 11/28/2023 2:11:53 PM | 2004-01-19 |  | UMULLER |  | Concentration | mg/L | Y | (fl. conc.) |  | Nickel Extractable (fl. conc.) | 7440-02-0 | Metals - Extractable; | Metals - Extractable; EMS code: NI-E | Chemical | Extractable | LAB
NI-E | ug/g | Nickel Extractable | X230 | Hom,HNO3/HCl,ICP | The sample is dried, ground to pass 100 mesh, and digested with a nitric-hydrochloric acid mixture to solubilize the solid matter and to remove the organic material by oxidation and volatilization. The completed digests are submitted to metals instrumentation personnel for metals analysis by Inductively Coupled Atomic Emission Spectrophotometry (ICP-AES), or hydride analysis by atomic absorption. | 1 | Y | Y | Y | 70 |  |  | 10/24/1996 2:25:28 PM | 1996-04-12 | 6/3/1998 6:51:37 PM | EMS |  | MassPerMass | mg/kg | Y | (mass conc.) |  | Nickel Extractable (mass conc.) | 7440-02-0 | Metals - Extractable; | Metals - Extractable; EMS code: NI-E | Chemical | Extractable | LAB
NI-E | mg/L | Nickel Extractable | F040 | Field Acid, lab filt, ICP-AES | Dissolved and extractable water samples are ready for analysis once collected. Total samples are autoclaved with HNO3 and hydrochloric acid (HCI) in capped Teflon vials. Analysis is performed via Inductively Coupled Argon Plasma Atomic Emission Spectrometry (ICP-AES)(simultaneous multielement analysis), Graphite Furnace Atomic Absorption Spectrometry (GF-AAS)(low level analysis) and/or Flame Atomic Absorption Spectrometry (F-AAS). | 0.02 | Y | Y | Y | 2350 | 8/5/1997 6:06:00 PM | 3/14/2008 | 9/11/2019 8:47:41 AM | 1997-05-15 | 6/3/1998 6:51:47 PM | MPAK |  | Concentration | mg/L | Y | (fl. conc.) |  | Nickel Extractable (fl. conc.) | 7440-02-0 | Metals - Extractable; | Metals - Extractable; EMS code: NI-E | Chemical | Extractable | LAB
NI-E | mg/L | Nickel Extractable | F069 | SEM extr,6N HCl,ICP/OES | 6.0 N Hydrochloric acid (HCI) is added to an aliquot of the sediment sample. The extract produced from the addition of the acid is then analyzed for simultaneously extracted metals (SEM) using inductively coupled plasma - optical emission spectrophotometry. The detection limits depend on the metal parameters in question and the instrumentation used. | 0.02 | Y | Y | Y | 1 |  |  | 11/25/2021 1:50:15 PM | 1999-07-09 |  | NPEPPIN |  | Concentration | mg/L | Y | (fl. conc.) |  | Nickel Extractable (fl. conc.) | 7440-02-0 | Metals - Extractable; | Metals - Extractable; EMS code: NI-E | Chemical | Extractable | LAB
NI-E | umol/g | Nickel Extractable | FSEM | AVS extraction, SEM Analysis ICPMS | Determination of simultaneous extractable metals (SEM) in sediment after sample extracted with hydrochloric acid for acid volatile sulfide (AVS) | 1 | Y | Y | Y | 19 | 3/12/2020 | 3/24/2020 | 2/22/2023 12:51:03 PM | 2010-11-16 |  | RDGIBSON |  | ParticlesPerMass | µmol/g | Y | (part. /m) |  | Nickel Extractable (part. /m) | 7440-02-0 | Metals - Extractable; | Metals - Extractable; EMS code: NI-E | Chemical | Extractable | LAB
NI-S | umol/g | Nickel Soluble ion content | FSEM | AVS extraction, SEM Analysis ICPMS | Determination of simultaneous extractable metals (SEM) in sediment after sample extracted with hydrochloric acid for acid volatile sulfide (AVS) | 1 | Y | Y | Y | 0 |  |  |  | 2024-02-02 |  | EMS |  | ParticlesPerMass | µmol/g | Y | (part. /m) |  | Nickel Soluble ion content (part. /m) |  | Metals - Soluble; | Metals - Soluble; EMS code: NI-S | Chemical |  | LAB
NI-T | mg/L | Nickel Total | 1349 | Tot Metals in Seawater by CRC ICPMS (BC) | Tot Metals in Seawater by CRC ICPMS (BC) | 0.0005 | Y | Y | Y | 98 | 8/16/2019 | 12/5/2021 | 12/13/2021 12:34:16 PM | 2019-09-30 |  | KATFRASE |  | Concentration | mg/L | Y | (fl. conc.) |  | Nickel Total (fl. conc.) | 7440-02-0 | Metals - Total; | Metals - Total; EMS code: NI-T | Chemical | Total | LAB
NI-T | mg/L | Nickel Total | METN | MET-T-NP-CCMS-VA | Ultra trace metals in water are analyzed by CRC ICPMS, based on US EPA Method 6020B (July 2014). The detection limits provided can only be met for undigested samples. This procedure is intended for colorless, non-turbid, acid-preserved water samples (i.e. pristine water samples), having turbidity < 1 NTU and no odor. Where turbidity exceeds 1 NTU, and/or the sample is colored and has an odor, results may be biased low compared to true Total Metals concentrations. ALS recommends that turbidity analysis be requested on samples submitted for this test to aid with interpretation of results. Where turbidity is <1NTU, undigested metals are equivalent to total metals concentrations. | 5e-05 | Y | Y | Y | 3648 | 4/3/2012 | 8/19/2024 | 8/26/2024 1:03:06 PM | 2019-08-20 |  | KATFRASE |  | Concentration | mg/L | Y | (fl. conc.) |  | Nickel Total (fl. conc.) | 7440-02-0 | Metals - Total; | Metals - Total; EMS code: NI-T | Chemical | Total | LAB
NI-T | ug/g | Nickel Total | EQ02 | EQUIS (ug/g) | EQUIS (ug/g) | 5 | N | Y | Y | 791 |  |  | 10/23/1996 12:08:28 AM | 1996-04-12 | 6/3/1998 6:51:37 PM | EMS |  | MassPerMass | mg/kg | Y | (mass conc.) |  | Nickel Total (mass conc.) | 7440-02-0 | Metals - Total; | Metals - Total; EMS code: NI-T | Chemical | Total | LAB
NI-T | mg/L | Nickel Total | F068 | HNO3dig,TeflonTube/Autoclave,ICP-MS | This method is applicable to determiniation of dissolved, extractable, and total recoverable metals in ground waters, surface waters, drinking waters, and wastewaters. Dissolved, extractable, and total recoverable sample fractions are acidified to pH <1.5 with analytical grade nitric acid (HNO3) to prevent the metals from adsorbing to the container material or precipitating. Extractable and total recoverable water samples are collected in 250 ml acid washed polyethylene bottles and are preserved with Nitric acid (0.5 ml acid for 100 mLs of sample). All preserved samples are stored at room temperature and can be held for up to 6 months. Unfiltered, un-preserved samples are stored at 4C for no more than 30 days. For dissolved and extractable water fractions, no sample preparation is necessary. Samples are simply poured in autosampler tubes (i.e. 15 ml centrifuge tubes), and placed on the autosampler and analyzed. Samples for total recoverable metals are digested with Nitric acid in special Teflon tubes in an autoclave prior to analyses. | 5e-05 | Y | Y | Y | 1779 | 6/1/1999 1:33:00 PM | 7/17/2008 4:44:00 PM | 7/30/2024 6:32:12 AM | 1999-06-28 |  | NPEPPIN |  | Concentration | mg/L | Y | (fl. conc.) |  | Nickel Total (fl. conc.) | 7440-02-0 | Metals - Total; | Metals - Total; EMS code: NI-T | Chemical | Total | LAB
NI-T | mg/L | Nickel Total | 0200 | 20%A.R.Dig:AA Anal. | 20%A.R.Dig:AA Anal. | 0.1 | Y | Y | Y | 6 |  |  | 10/23/1996 11:23:40 PM | 1996-04-12 | 6/3/1998 6:51:37 PM | EMS |  | Concentration | mg/L | Y | (fl. conc.) |  | Nickel Total (fl. conc.) | 7440-02-0 | Metals - Total; | Metals - Total; EMS code: NI-T | Chemical | Total | LAB
NI-T | ug/g | Nickel Total | MTID | MET-DRY-CCMS-VA | Metals in Tissue by CRC ICPMS (DRY) Methodology: This method is conducted following British Columbia Lab Manual method "Metals in Animal Tissue and Vegetation (Biota) - Prescriptive". Tissue samples are homogenized and sub-sampled prior to hotblock digestion with nitric and hydrochloric acids, in combination with addition of hydrogen peroxide. Instrumental analysis is by collision cell inductively coupled plasma - mass spectrometry (modified from EPA Method 6020A). Method Limitation: This method employs a strong acid/peroxide digestion, and is intended to provide a conservative estimate of bio-available metals. Near complete recoveries are achieved for most toxicologically important metals, but elements associated with recalcitrant minerals may be only partially recovered. | 0.2 | Y | Y | Y | 122 | 8/13/2014 | 9/13/2022 | 10/26/2022 7:14:15 AM | 2015-03-09 | 3/11/2015 4:12:05 PM | RDGIBSON | EMS | MassPerMass | mg/kg | Y | (mass conc.) |  | Nickel Total (mass conc.) | 7440-02-0 | Metals - Total; | Metals - Total; EMS code: NI-T | Chemical | Total | LAB
NI-T | ug/g | Nickel Total | X551 | Metals in Biota (dry), total, acid digestion, ICP-SFMS | 0.1 gram of dried, homogenized biota is weighed in a microwave TFM vessel, and 8mL of Nitric Acid and 1mL of Hydrogen Peroxide are added. The sample is digested at 200 degrees C for 15 minutes in a closed vessel high pressure microwave oven. The sample is cooled, the volume adjusted to 100mL and introduced to an iductively coupled argon plasma-sector field mass spectrometer (ICP-SFMS) by pneumatic nebulization. Each respective element is measured at a specific mass to charge ratio, m/z value expressed in amu (atomic mass units), and at a specified instrument resolution. LR: Low Resolution ~300, MR: Mid Resolution ~4000 and HR: Hich Resolution ~10000. Ref: Method number 02-2705, National Laboratory for Environmental Testing, CCIW, Environment Canada, Burlington. Requested April 2003. | 0.05 | Y | Y | Y | 70 | 8/11/2015 | 10/30/2023 | 11/28/2023 2:21:33 PM | 2004-01-19 |  | UMULLER |  | MassPerMass | mg/kg | Y | (mass conc.) |  | Nickel Total (mass conc.) | 7440-02-0 | Metals - Total; | Metals - Total; EMS code: NI-T | Chemical | Total | LAB
NI-T | mg/L | Nickel Total | X170 | ICP | This code is a catch-all for a large variety of ICP methods including water and also solids digested into solutions as an intermediate result. Different procedures will have greatly different detection limits. (a) For air particulates, preliminary results are reported in mg/L and loading results as mg/dm2/d. (b) For Hi Vol samples preliminiary results are same as reported in mg/L and loading results as ug/m3. (c) Note that for Rain samples there is no filtration nor digestion steps. | 0.008 | Y | Y | Y | 6978 | 3/28/1996 6:03:00 PM | 11/7/2018 | 10/24/2022 10:58:25 AM | 1996-04-12 | 6/3/1998 6:51:37 PM | EMS |  | Concentration | mg/L | Y | (fl. conc.) |  | Nickel Total (fl. conc.) | 7440-02-0 | Metals - Total; | Metals - Total; EMS code: NI-T | Chemical | Total | LAB
NI-T | ug/g wet | Nickel Total | MTIW | MET-WET-CCMS-MID-VA | Metals in Tissue by CRC ICPMS (WET) Methodology: This method is conducted following British Columbia Lab Manual method "Metals in Animal Tissue and Vegetation (Biota) - Prescriptive". Tissue samples are homogenized and sub-sampled prior to hotblock digestion with nitric and hydrochloric acids, in combination with addition of hydrogen peroxide. Instrumental analysis is by collision cell inductively coupled plasma - mass spectrometry (modified from EPA Method 6020A). Method Limitation: This method employs a strong acid/peroxide digestion, and is intended to provide a conservative estimate of bio-available metals. Near complete recoveries are achieved for most toxicologically important metals, but elements associated with recalcitrant minerals may be only partially recovered. | 0.04 | Y | Y | Y | 24 | 11/25/2018 | 12/6/2018 | 7/21/2021 2:55:07 PM | 2018-12-24 |  | EMS |  | MassPerMass | µg/g | Y | (mass conc.) | (wet) | Nickel Total (wet) (mass conc.) | 7440-02-0 | Metals - Total; | Metals - Total; EMS code: NI-T | Chemical | Total | LAB
NI-T | mg/L | Nickel Total | EQ01 | EQUIS (mg/L) | EQUIS (mg/L) | 0.01 | N | Y | Y | 3249 |  |  | 10/24/1996 3:16:00 PM | 1996-04-12 | 6/3/1998 6:51:41 PM | EMS |  | Concentration | mg/L | Y | (fl. conc.) |  | Nickel Total (fl. conc.) | 7440-02-0 | Metals - Total; | Metals - Total; EMS code: NI-T | Chemical | Total | LAB
NI-T | mg/L | Nickel Total | TCLP | Toxicity Characteristic Leaching Procedure | The TCLP procedure is designed to simulate the climatic leaching action | 0.1 | Y | Y | Y | 22 | 12/8/2009 7:49:00 PM | 9/28/2011 5:53:00 PM | 11/3/2011 11:10:19 PM | 2009-03-06 |  | JCHUDAK |  | Concentration | mg/L | Y | (fl. conc.) |  | Nickel Total (fl. conc.) | 7440-02-0 | Metals - Total; | Metals - Total; EMS code: NI-T | Chemical | Total | LAB
NI-T | ppm (S) | Nickel Total | 6041 | Aqua Regia by ICPMS | These were sediment/soils digested with Aqua-Regia (Nitric Acid & Hydrochloric Acids) and then analysis by ICPMS. Results would be reported in mg/kg and/or %. | 0.04 | Y | Y | Y | 21 | 10/13/2019 | 8/5/2020 | 8/10/2020 3:52:06 PM | 2019-09-06 | 10/21/2019 11:55:49 AM | KATFRASE | KATFRASE | DimensionlessRatio | ppm | N | (ratio) |  | Nickel Total (ratio) | 7440-02-0 | Metals - Total; | Metals - Total; EMS code: NI-T | Chemical | Total | LAB
NI-T | mg/L | Nickel Total | F121 | Acid,ICPMS | Acidified, no digestion, ICPMS | 2e-05 | Y | Y | Y | 8440 | 3/12/2010 7:57:00 PM | 8/15/2024 9:10:00 PM | 8/26/2024 8:06:03 AM | 2002-07-22 | 7/16/2012 3:30:03 PM | UMULLER | RDGIBSON | Concentration | mg/L | Y | (fl. conc.) |  | Nickel Total (fl. conc.) | 7440-02-0 | Metals - Total; | Metals - Total; EMS code: NI-T | Chemical | Total | LAB
NI-T | mg/L | Nickel Total | 0061 | HNO3 Dig: GFAA Analysis | HNO3 Dig: GFAA Analysis | 0.005 | Y | Y | Y | 17 |  |  | 10/24/1996 3:06:36 PM | 1996-04-12 | 6/3/1998 6:51:37 PM | EMS |  | Concentration | mg/L | Y | (fl. conc.) |  | Nickel Total (fl. conc.) | 7440-02-0 | Metals - Total; | Metals - Total; EMS code: NI-T | Chemical | Total | LAB
NI-T | mg/L | Nickel Total | MSEA | Metals in Seawater (undigested) by (CRC) ICP-MS | This method describes the multi-elemental determination of trace elements by Collision/Reaction Cell (CRC) ICP-MS. The method measures ions produced by a radio-frequency inductively coupled plasma. Analyte species originating in a liquid are pneumatically nebulized and the resulting aerosol transported by argon gas into the plasma torch. The ions produced are entrained in the plasma gas and extracted, by means of a differentially pumped vacuum interface, into a mass spectrometer. The ions flow through a Collision/Reaction Cell (CRC) where depending on the mode of analysis either the ions passes through the cell unhindered (no gas - classical ICP-MS mode), or collisions/reactions chemistry takes place between the ions and a neutral gas such as hydrogen or helium. Any polyatomic ions which give interference on the analytes are removed. The ions are then sorted according to their mass-to-charge ratios by a quadrupole mass spectrometer. The ions are then sorted according to their mass-to-charge ratios by a quadrupole mass spectrometer. The ions transmitted through the quadrupole are quantified by a channel electron multiplier and the ion information processed by a data handling system. Although most polyatomic interference is removed, any additional interference must be assessed and valid corrections applied or the data flagged to indicate problems. Interference correction must include compensation for background ions contributed by the plasma gas, reagents, and constituents of the sample matrix. Instrument drift as well as suppressions or enhancements of instrument response caused by the sample matrix must be corrected for by the use of internal standards. | 0.0001 | Y | Y | Y | 432 | 8/12/2009 12:08:00 AM | 4/29/2024 | 6/7/2024 10:03:16 AM | 2009-09-08 |  | RDGIBSON |  | Concentration | mg/L | Y | (fl. conc.) |  | Nickel Total (fl. conc.) | 7440-02-0 | Metals - Total; | Metals - Total; EMS code: NI-T | Chemical | Total | LAB
NI-T | ug/g | Nickel Total | X230 | Hom,HNO3/HCl,ICP | The sample is dried, ground to pass 100 mesh, and digested with a nitric-hydrochloric acid mixture to solubilize the solid matter and to remove the organic material by oxidation and volatilization. The completed digests are submitted to metals instrumentation personnel for metals analysis by Inductively Coupled Atomic Emission Spectrophotometry (ICP-AES), or hydride analysis by atomic absorption. | 1.5 | Y | Y | Y | 247 | 2/12/2002 | 12/30/2005 | 2/3/2006 1:13:13 AM | 1996-04-12 | 6/3/1998 6:51:37 PM | EMS |  | MassPerMass | mg/kg | Y | (mass conc.) |  | Nickel Total (mass conc.) | 7440-02-0 | Metals - Total; | Metals - Total; EMS code: NI-T | Chemical | Total | LAB
NI-T | mg/dm2/d | Nickel Total | DUST | HNO3 dig.(microwave or hot block),ICPMS | HNO3 dig.(microwave or hot block);ICPMS |  | Y | Y | Y | 267 | 7/19/2016 1:11:00 PM | 7/31/2024 | 7/31/2024 11:23:12 AM | 2008-02-07 |  | RDGIBSON |  | MassFlux | mg/dm²/d | Y | (mass flux) |  | Nickel Total (mass flux) | 7440-02-0 | Metals - Total; | Metals - Total; EMS code: NI-T | Chemical | Total | LAB
NI-T | mg/L | Nickel Total | F033 | Acid Dig,Centrifuge,ICPMS | For the determination of total recoverable elements in water or waste water, take 100 ml aliquot from a well mixed, acid preserved, sample containing not more than 0.25% (w/v) total solids and transfer to a 250 ml Griffin beaker. (If total solids are greater than 0.25% reduce the size of the aliquot by a proportionate amount). Add 1 ml of conc. nitric acid and 0.5 ml conc. hydrochloric acid. Heat on a hot plate until the volume has been reduced to 15-20ml, ensuring that the sample does not boil. Allow to cool and transfer to a 50 ml volumetric flask. Dilute to volume with ASTM type I water. Centrifuge the sample or allow to stand overnight to separate insoluble material. Prior analysis, pipette 20 ml into a 50 ml volumetric flask and dilute to volume with ASTM type I water. Analysis should be performed within seventy two hours of the completed preparation. Note that different laboratories may have very different field preservation, digestion procedures and detection limits for this method, so these should be documented. | 0.0005 | Y | Y | Y | 40377 | 10/15/1996 9:27:00 AM | 8/22/2024 3:24:00 AM | 8/25/2024 9:47:36 PM | 1997-01-16 | 6/3/1998 6:51:46 PM | MPAK |  | Concentration | mg/L | Y | (fl. conc.) |  | Nickel Total (fl. conc.) | 7440-02-0 | Metals - Total; | Metals - Total; EMS code: NI-T | Chemical | Total | LAB
NI-T | ug/g | Nickel Total | 0196 | D&G: LTA/HCl: ICP | D&G: LTA/HCl: ICP | 5 | Y | Y | Y | 27 |  |  | 10/24/1996 2:08:12 PM | 1996-04-12 | 6/3/1998 6:51:37 PM | EMS |  | MassPerMass | mg/kg | Y | (mass conc.) |  | Nickel Total (mass conc.) | 7440-02-0 | Metals - Total; | Metals - Total; EMS code: NI-T | Chemical | Total | LAB
NI-T | mg/L | Nickel Total | X179 | LA,HNO3,Graphite Furnace | Samples are lab-acidified with HNO3; HNO3 digestion; analysis by GFAA. | 0.001 | Y | Y | Y | 1 | 6/3/2002 | 6/3/2002 | 6/10/2002 5:07:09 PM | 1996-04-12 | 6/3/1998 6:51:37 PM | EMS |  | Concentration | mg/L | Y | (fl. conc.) |  | Nickel Total (fl. conc.) | 7440-02-0 | Metals - Total; | Metals - Total; EMS code: NI-T | Chemical | Total | LAB
NI-T | mg/L | Nickel Total | ULTR | Total Metals in Water by ICPMS (Ultra) | Total Metals in Water by ICPMS (Ultra), EPA SW-846 3005A/6020A, MET-TOT-ULTRA-MS-VA This analysis is carried out using procedures adapted from "Standard Methods for the Examination of Water and Wastewater" published by the American Public Health Association, and with procedures adapted from "Test Methods for Evaluating Solid Waste" SW-846 published by the United States Environmental Protection Agency (EPA). The procedures may involve preliminary sample treatment by acid digestion, using either hotblock or microwave oven, or filtration (EPA Method 3005A). Instrumental analysis is by inductively coupled plasma - mass spectrometry (EPA Method 6020A). | 0.0001 | Y | Y | Y | 176 | 2/1/2017 | 10/26/2019 | 11/13/2019 10:17:12 AM | 2012-07-05 |  | RDGIBSON |  | Concentration | mg/L | Y | (fl. conc.) |  | Nickel Total (fl. conc.) | 7440-02-0 | Metals - Total; | Metals - Total; EMS code: NI-T | Chemical | Total | LAB
NI-T | mg/L | Nickel Total | 0060 | HNO3 Dig: HGA Analysis | HNO3 Dig: HGA Analysis | 0.001 | Y | Y | Y | 987 |  |  | 9/17/1997 1:35:59 PM | 1996-04-12 | 6/3/1998 6:51:37 PM | EMS |  | Concentration | mg/L | Y | (fl. conc.) |  | Nickel Total (fl. conc.) | 7440-02-0 | Metals - Total; | Metals - Total; EMS code: NI-T | Chemical | Total | LAB
NI-T | mg/L | Nickel Total | X072 | FA,HNO3,Graph.Furnace,AA | Samples are field-acidified with HNO3; HNO3 digestion; analysis by GFAA. | 0.001 | Y | Y | Y | 6529 |  |  | 8/26/2024 2:41:32 PM | 1996-04-12 | 6/3/1998 6:51:37 PM | EMS |  | Concentration | mg/L | Y | (fl. conc.) |  | Nickel Total (fl. conc.) | 7440-02-0 | Metals - Total; | Metals - Total; EMS code: NI-T | Chemical | Total | LAB
NI-T | ug/g wet | Nickel Total | X552 | Metals in Biota (wet), total, acid digestion, ICP-SFMS | One gram of wet, homogenized biota is weighed in a microwave TFM vessel, and 8mL of Nitric Acid and 1mL of Hydrogen Peroxide are added. The sample is digested at 200 degrees C for 15 minutes in a closed vessel high pressure microwave oven. The sample is cooled, the volume adjusted to 100mL and introduced to an inductively coupled argon plasma-sector field mass spectrometer (ICP-SFMS) by pneumatic nebulization. Each respective element is measured at a specific mass to charge ratio, m/z value expressed in amu (atomic mass units), and at a specified instrument resolution. LR: Low Resolution ~300, MR: Mid Resolution ~4000. Ref: 02-2704, National Laboratory for Environmental Testing, CCIW, Environment Canada, Burlington. Requested April 2003. | 0.005 | Y | Y | Y | 37 | 8/9/2016 | 10/16/2023 | 11/28/2023 2:21:32 PM | 2004-01-19 | 4/19/2021 1:22:00 PM | UMULLER | EMS | MassPerMass | µg/g | Y | (mass conc.) | (wet) | Nickel Total (wet) (mass conc.) | 7440-02-0 | Metals - Total; | Metals - Total; EMS code: NI-T | Chemical | Total | LAB
NI-T | ug/g | Nickel Total | 0192 | Hom, Nitric/Perchlor,ICP | Sample already in a dry state so does not require drying before homogenization and nitric perchlorid digestion followed by ICP analysis. | 5 | Y | Y | Y | 708 |  |  | 10/24/1996 12:50:25 PM | 1996-04-12 | 6/3/1998 6:51:41 PM | EMS |  | MassPerMass | mg/kg | Y | (mass conc.) |  | Nickel Total (mass conc.) | 7440-02-0 | Metals - Total; | Metals - Total; EMS code: NI-T | Chemical | Total | LAB
NI-T | mg/L | Nickel Total | F108 | SPR,IDA Chelation,ICPMS | SPR;IDA Chelation;ICPMS |  | Y | Y | Y | 4 | 11/30/2000 10:32:00 AM | 11/30/2000 10:32:00 AM | 9/10/2002 9:29:49 AM | 2001-03-08 |  | NPEPPIN |  | Concentration | mg/L | Y | (fl. conc.) |  | Nickel Total (fl. conc.) | 7440-02-0 | Metals - Total; | Metals - Total; EMS code: NI-T | Chemical | Total | LAB
NI-T | ppm (S) | Nickel Total | MS41 | ALS Minerals ME-MS41 | ALS Minerals ME-MS41. A prepared sample is digested with aqua regia in a graphite heating block. After cooling; the resulting solution is diluted to with deionized water; mixed and analyzed by ICP-AES. Samples are then analyzed by ICP-MS for the remaining suite of elements. | 0.2 | Y | Y | Y | 0 |  |  |  | 2024-01-29 |  | EMS |  | DimensionlessRatio | ppm | N | (ratio) |  | Nickel Total (ratio) | 7440-02-0 | Metals - Total; | Metals - Total; EMS code: NI-T | Chemical | Total | LAB
NI-T | mg/L | Nickel Total | F038 | HNO3/HCl dig,ICP-AES | Dissolved and extractable water samples are ready for analysis once collected. Total samples are autoclaved with HNO3 and hydrochloric acid (HCI) in capped Teflon vials. Analysis is performed via Inductively Coupled Argon Plasma Atomic Emission Spectrometry (ICP-AES)(simultaneous multielement analysis), Graphite Furnace Atomic Absorption Spectrometry (GF-AAS) (low level analysis) and/or Flame Atomic Absorption Spectrometry (F-AAS). | 0.02 | Y | Y | Y | 5426 | 8/1/1996 6:04:00 PM | 5/6/2008 | 9/11/2019 8:47:42 AM | 1997-05-15 | 6/3/1998 6:51:47 PM | MPAK |  | Concentration | mg/L | Y | (fl. conc.) |  | Nickel Total (fl. conc.) | 7440-02-0 | Metals - Total; | Metals - Total; EMS code: NI-T | Chemical | Total | LAB
NI-T | mg/L | Nickel Total | 0100 | Dig:Ext.(APDC):AA Anal. | Dig:Ext.(APDC):AA Anal. | 0.01 | Y | Y | Y | 590 |  |  | 6/16/2000 4:19:26 PM | 1996-04-12 | 6/3/1998 6:51:41 PM | EMS |  | Concentration | mg/L | Y | (fl. conc.) |  | Nickel Total (fl. conc.) | 7440-02-0 | Metals - Total; | Metals - Total; EMS code: NI-T | Chemical | Total | LAB
NI-T | ug/g | Nickel Total | ICPM | Metals in Soil by CRC ICPMS | This analysis is carried out using procedures from CSR Analytical Method: "Strong Acid Leachable Metals (SALM) in Soil", BC Ministry of Environment, 26 June 2009, and procedures adapted from EPA Method 200.2. The sample is manually homogenized, dried at 60 degrees Celsius, sieved through a 63 um (230 mesh) sieve, and a representative subsample of the dry material is weighed. The sample is then digested at 95 degrees Celsius for 2 hours by block digester using concentrated nitric and hydrochloric acids. Instrumental analysis is by collision cell inductively coupled plasma - mass spectrometry (modifed from EPA Method 6020A). Method Limitation: This method is not a total digestion technique. It is a very strong acid digestion that is intended to dissolve those metals that may be environmentally available. By design, elements bound in silicate structures are not normally dissolved by this procedure as they are not usually mobile in the environment. | 0.5 | Y | Y | Y | 103 | 9/19/2017 | 9/24/2023 | 1/12/2024 11:16:52 AM | 2018-01-09 |  | JKROGH |  | MassPerMass | mg/kg | Y | (mass conc.) |  | Nickel Total (mass conc.) | 7440-02-0 | Metals - Total; | Metals - Total; EMS code: NI-T | Chemical | Total | LAB
NI-T | ug/g | Nickel Total | FB04 | Freeze Dry,Hom,Acid Dig,ICP | Sample moisture is determined by weighing samples before and after freeze-drying. Freeze-dried samples are ground in a Braun blender before being weighed into Teflon PFA liners and digested with concentrated HNO3, concentrated HCl, and 50% H2O2 in a microwave oven; which results in a complete sample dissolution. Samples are collected and volumized with Type I water. Analysis is performed via Inductively Coupled Argon Plasma Atomic Emission Spectrometry (ICP-AES) (simultaneous multielement analysis), Graphite Furnace Atomic Absorption Spectrometry (GF-AAS) (low level Cd and Pb analysis), Cold Vapor Atomic Fluorescence Spectrometry (CV-AFS) (Mercury analysis) and/or Flame Atomic Absorption Spectrometry (F-AAS) (K analysis). | 2 | Y | Y | Y | 241 | 8/22/1996 6:11:00 PM | 2/1/2002 2:30:00 PM | 6/23/2003 1:15:32 PM | 1997-02-27 | 6/3/1998 6:51:46 PM | MPAK |  | MassPerMass | mg/kg | Y | (mass conc.) |  | Nickel Total (mass conc.) | 7440-02-0 | Metals - Total; | Metals - Total; EMS code: NI-T | Chemical | Total | LAB
NI-T | ug/g | Nickel Total | FS01 | D&G,Nitric/HCl:ICP | Samples are dried at 60 C, then either sieved to -100mesh (150 um) and rolled to homogenize, or finely ground with a mortar and pestle or a shatter box laboratory mill and rolled (depending on the submitter's requirements). The samples are then weighed into Teflon digestion liners and digested with HNO3 and HCl in a microwave oven. Samples are cooled, volumized, and allowed to settle overnight before being decanted. Analysis is then performed via Inductively Coupled Argon Plasma Atomic Emission Spectrometry (ICP-AES) (Simultaneous multi element analysis). In ICP Spectrometry, the liquid sample is pumped to the nebulizer, where it is convertd into an aerosol. the nebulizer carrier gas flow than transports the wet aerosol past the impact bead, where the large droplets condense out and flow to the waste container. The remaining fine aerosol is carried into the plasma. There, the high temperature desolvates the aerosol and dissociates the sample material into individual atoms and ions, which are excited to emit light at wavelengths characteristic of the elements present. A diffraction grating in the spectrometer diffracts the various wavelengths of light through exit slits to the photo-multiplier tubes, which convert the light energy to electronic signals. | 3 | Y | Y | Y | 2642 | 8/30/1996 1:52:00 PM | 10/10/2007 | 5/24/2008 1:15:03 AM | 1996-11-20 | 6/3/1998 6:51:46 PM | MPAK |  | MassPerMass | mg/kg | Y | (mass conc.) |  | Nickel Total (mass conc.) | 7440-02-0 | Metals - Total; | Metals - Total; EMS code: NI-T | Chemical | Total | LAB
NI-T | ug/g wet | Nickel Total | X215 | Hom,Nitr/Perchl,Grph Fur | Hom;Nitr/Perchl;Grph Fur | 0.01 | Y | Y | Y | 74 |  |  | 10/24/1996 10:01:25 AM | 1996-04-12 | 6/3/1998 6:51:42 PM | EMS |  | MassPerMass | µg/g | Y | (mass conc.) | (wet) | Nickel Total (wet) (mass conc.) | 7440-02-0 | Metals - Total; | Metals - Total; EMS code: NI-T | Chemical | Total | LAB
NI-T | mg/L | Nickel Total | MWID | Metals in Water by ICPMS (digested) | Metals in Water by ICPMS (digested) Based on EPA 200.8 | 0.001 | Y | Y | Y | 36957 |  |  | 7/31/2024 1:46:41 PM | 2009-12-24 |  | RDGIBSON |  | Concentration | mg/L | Y | (fl. conc.) |  | Nickel Total (fl. conc.) | 7440-02-0 | Metals - Total; | Metals - Total; EMS code: NI-T | Chemical | Total | LAB
NI-T | ug/g | Nickel Total | PB02 | PBM,HNO3/HCl dig,dry60C,100mesh, | Performance Based Method. Samples are dried at 60 C and sieved using a .149 mm (100 mesh) sieve. Solid samples are digested with the combination of nitric and hydrochloric acids with a 1:1 ratio of acids. For environmental purposes, strong acid leaches are used routinely since they have the advantage of speed, low cost, and superio reproducibility. All metals trapped within the refractory matrix are unlikely to have environmental consequences. Instrumental analysis of extracts produced using this digestion method can be performed by a variety of techniques. |  | Y | Y | Y | 52 | 5/24/2001 3:42:00 PM | 11/26/2001 10:23:00 AM | 5/26/2003 2:28:11 PM | 2001-07-18 |  | NPEPPIN |  | MassPerMass | mg/kg | Y | (mass conc.) |  | Nickel Total (mass conc.) | 7440-02-0 | Metals - Total; | Metals - Total; EMS code: NI-T | Chemical | Total | LAB
NI-T | mg/L | Nickel Total | F048 | Acid dig,conc'n,nebul.,ICP-OES | This method provides an automated procedure for the Inductively Coupled Argon Plasma Optical Emission Spectrophotometric (ICP-OES) determination and quantitation of trace amounts of metals in surface, ground and precipitation waters after a manual acid digestion and concentration procedure and ultrasonic nebulisation. Organically-bound metals and metal complexes are converted to metal ions by digestion with nitric acid [method details says HCl]. To increase the signal detected by the ICP, the samples are concentrated 10 times. The liquid sample is pumped onto the piezoelectric transducer, where the ultrasonic action efficiently converts it to a fine, dense aerosol. The nebulizer carrier gas flow then transports the wet aerosol through a heated U-tube, where the solvent is vaporized. Subsequently, most of the solvent vapour is removed from the aerosol stream by passing the stream through a chilled condenser. The result is a dry, analyte-laden aerosol which is introduced to the plasma. There, the high temperature desolvates the aerosol and dissociates the sample material into individual atoms and ions, which are excited to emit light at wave lengths characteristic of the elements present. Dissolved samples are field filtered through 0.45 micron cellulose acetate filter prior to acid preservation. Samples for extractable analysis are acidified with 2 mL concentrated HNO3 per litre and allowed to stand a minimum of 24 hours before analysis. Clients providing these samples are advised to preserve samples at the collection site. | 0.0002 | Y | Y | Y | 6159 | 7/4/1997 2:36:00 PM | 4/23/2003 | 9/11/2019 8:46:38 AM | 1997-10-20 | 6/3/1998 6:51:47 PM | NPEPPIN |  | Concentration | mg/L | Y | (fl. conc.) |  | Nickel Total (fl. conc.) | 7440-02-0 | Metals - Total; | Metals - Total; EMS code: NI-T | Chemical | Total | LAB
NI-T | ug/m3 | Nickel Total | F088 | Aliquot,HNO3/HCl dig, ICPMS,loading | Aliquot microwave digested with 3:1 HNO3/HCL. Analysis by ICPMS. For HiVol samples preliminary result are reported in mg/L and loading results are reported in ug/m3 |  | Y | Y | Y | 565 | 8/17/2000 10:46:00 AM | 12/14/2006 | 12/30/2006 1:01:37 AM | 2000-08-16 |  | NPEPPIN |  | AirConcentration | µg/m³ | Y | (air conc.) |  | Nickel Total (air conc.) | 7440-02-0 | Metals - Total; | Metals - Total; EMS code: NI-T | Chemical | Total | LAB
NI-T | ug/g | Nickel Total | FS82 | PBM,dry60C,10mesh,SALM-ICP-MS | PBM;dry60C;10mesh;SALM-ICP-MS | 0.8 | Y | Y | Y | 3696 | 5/21/2002 | 7/24/2024 | 7/30/2024 10:00:14 AM | 2002-11-12 |  | UMULLER |  | MassPerMass | mg/kg | Y | (mass conc.) |  | Nickel Total (mass conc.) | 7440-02-0 | Metals - Total; | Metals - Total; EMS code: NI-T | Chemical | Total | LAB
NI-T | ug/g wet | Nickel Total | BPSY | MET-WET-BPSY-CCMS-VA | Metals in Animal Tissue and Vegetation (Biota) - Prescriptive. EPA 200.3/6020B This method is conducted following British Columbia Lab Manual method "Metals in Animal Tissue and Vegetation (Biota) - Prescriptive". Tissue samples are digested with nitric and hydrochloric acids; in combination with addition of hydrogen peroxide. Instrumental analysis is by collision cell inductively coupled plasma - mass spectrometry (modified from EPA Method 6020B). | 0.04 | Y | Y | Y | 2 | 11/7/2019 | 11/7/2019 | 6/10/2021 1:57:10 PM | 2021-06-04 |  | EMS |  | MassPerMass | µg/g | Y | (mass conc.) | (wet) | Nickel Total (wet) (mass conc.) | 7440-02-0 | Metals - Total; | Metals - Total; EMS code: NI-T | Chemical | Total | LAB
NI-T | mg/L | Nickel Total | 0042 | GI HNO3 Dig: ICP | Sample aliquot taken from unfiltered unpreserved General Ions bottle; HNO3 digestion; ICP analysis. | 0.01 | Y | Y | Y | 6812 | 9/20/1996 | 7/5/2019 3:21:00 PM | 7/12/2019 6:49:13 PM | 1996-04-12 | 6/3/1998 6:51:37 PM | EMS |  | Concentration | mg/L | Y | (fl. conc.) |  | Nickel Total (fl. conc.) | 7440-02-0 | Metals - Total; | Metals - Total; EMS code: NI-T | Chemical | Total | LAB
NI-T | ppm (S) | Nickel Total | MS61 | ALS Minerals ME-MS61 | ALS Minerals ME-MS61. A prepared sample is digested with HNO3; HF; and HClO4 in a graphite heating block to near dryness. After cooling; the sample is boiled in diluted HCl. Bulk to volume with deionized water; mixed and analyzed by ICP-AES and ICP-MS for all elements. | 0.2 | Y | Y | Y | 0 |  |  |  | 2024-01-29 |  | EMS |  | DimensionlessRatio | ppm | N | (ratio) |  | Nickel Total (ratio) | 7440-02-0 | Metals - Total; | Metals - Total; EMS code: NI-T | Chemical | Total | LAB
NI-T | mg/L | Nickel Total | 0040 | HNO3 Dig: ICP Analysis | HNO3 Dig: ICP Analysis | 0.05 | Y | Y | Y | 14342 |  |  | 3/23/2010 5:08:09 PM | 1996-04-12 | 6/3/1998 6:51:41 PM | EMS |  | Concentration | mg/L | Y | (fl. conc.) |  | Nickel Total (fl. conc.) | 7440-02-0 | Metals - Total; | Metals - Total; EMS code: NI-T | Chemical | Total | LAB
NI-T | mg/L | Nickel Total | 0050 | Graphite Furnace AA | Graphite Furnace AA | 0.0005 | Y | Y | Y | 21 | 1/6/2003 | 6/6/2003 | 6/9/2003 10:18:24 AM | 2003-01-13 | 1/13/2003 1:27:47 PM | UMULLER | UMULLER | Concentration | mg/L | Y | (fl. conc.) |  | Nickel Total (fl. conc.) | 7440-02-0 | Metals - Total; | Metals - Total; EMS code: NI-T | Chemical | Total | LAB
NI-T | mg/L | Nickel Total | X561 | Total metals: FA, HNO3, HR-ICPMS (No Prep)Prep) | Ultra trace metals in water are analyzed by high resolution inductively coupled plasma mass spectrometry (HR-ICPMS), modified from US EPA 200.8, (Revision 5.5). The detection limits provided can only be met for undigested samples. This procedure is intended for colorless, non-turbid, acid-preserved water samples (i.e. pristine water samples), having turbidity < 1 NTU and no odor. Where turbidity exceeds 1 NTU, and/or the sample is colored and has an odor, results may be biased low compared to true Total Metals concentrations. ALS recommends that turbidity analysis be requested on samples submitted for this test to aid with interpretation of results. | 5e-05 | Y | Y | Y | 2756 | 6/8/2015 | 10/5/2023 | 1/31/2024 2:37:41 PM | 2015-10-22 |  | RDGIBSON |  | Concentration | mg/L | Y | (fl. conc.) |  | Nickel Total (fl. conc.) | 7440-02-0 | Metals - Total; | Metals - Total; EMS code: NI-T | Chemical | Total | LAB
NI-T | ppm (S) | Nickel Total | 6042 | Aqua Regia by ICPMS | These were sediment/soils digested with Aqua-Regia (Nitric Acid & Hydrochloric Acids) and then analysis by ICPMS. Results would be reported in ppm s. | 0.04 | Y | Y | Y | 71 | 5/14/2022 | 7/27/2024 | 7/29/2024 9:13:08 AM | 2022-06-28 |  | EMS |  | DimensionlessRatio | ppm | N | (ratio) |  | Nickel Total (ratio) | 7440-02-0 | Metals - Total; | Metals - Total; EMS code: NI-T | Chemical | Total | LAB
NI-T | mg/L | Nickel Total | X349 | FA:HNO3:ICP | Samples are field-acidified with HNO3; HNO3 digestion; analysis by ICP/AES. | 0.02 | Y | Y | Y | 30 |  |  | 10/30/2014 8:31:18 AM | 1998-05-22 | 6/3/1998 6:51:49 PM | NPEPPIN |  | Concentration | mg/L | Y | (fl. conc.) |  | Nickel Total (fl. conc.) | 7440-02-0 | Metals - Total; | Metals - Total; EMS code: NI-T | Chemical | Total | LAB
NI-T | mg/L | Nickel Total | X285 | HNO3 Dig:ICP Analysis | HNO3 Dig:ICP Analysis | 0.02 | Y | Y | Y | 690 | 8/10/1993 | 10/5/1993 | 4/30/2004 1:59:45 PM | 1998-05-22 | 6/3/1998 6:51:49 PM | NPEPPIN |  | Concentration | mg/L | Y | (fl. conc.) |  | Nickel Total (fl. conc.) | 7440-02-0 | Metals - Total; | Metals - Total; EMS code: NI-T | Chemical | Total | LAB
NI-T | mg/L | Nickel Total | X555 | Metals in water, total, "in-bottle" digestion, ICP-SFMS | The whole water sample collected in a HDPE/Teflon bottle, is de-capped in a Class 100 workstation, acidified to 2% (v/v) HNO3 and 1% (v/v) HCl and recapped. The sample is "in-bottle" digested in a convection oven, set at 60 degrees C for 16 hours. The sample is cooled and introduced to an inductively coupled argon plasma-sector field mass spectrometer (ICP-SFMS) by pneumatic nebulization. Each respective element is measured at a specific mass to charge ratio, m/z value expressed in amu (atomic mass units), and at resolution. LR: Low Resolution ~300, MR: Mid Resolution ~4000 and HR: High a specified instrument Resolution ~10000. Ref: Method number 02-2003, National Laboratory for Environmental Testing, CCIW, Environment Canada, Burlington. Requested April 2003. | 2e-05 | Y | Y | Y | 17341 | 5/20/2003 | 12/6/2023 2:08:00 AM | 2/14/2024 9:05:32 AM | 2004-01-19 |  | UMULLER |  | Concentration | mg/L | Y | (fl. conc.) |  | Nickel Total (fl. conc.) | 7440-02-0 | Metals - Total; | Metals - Total; EMS code: NI-T | Chemical | Total | LAB
NI-T | ug/g | Nickel Total | FS89 | Homo,HNO3 dig,ICP-AES | Homo;HNO3 dig;ICP-AES | 3 | Y | Y | Y | 4 | 1/5/2001 | 8/20/2001 12:00:00 PM | 10/23/2002 1:42:49 PM | 2001-03-08 |  | NPEPPIN |  | MassPerMass | mg/kg | Y | (mass conc.) |  | Nickel Total (mass conc.) | 7440-02-0 | Metals - Total; | Metals - Total; EMS code: NI-T | Chemical | Total | LAB
NI-T | mg/L | Nickel Total | 0105 | 5%AR Dig:APDC Ext:AA Anl | 5%AR Dig:APDC Ext:AA Anl | 0.01 | Y | Y | Y | 235 |  |  | 10/24/1996 12:38:17 PM | 1996-04-12 | 6/3/1998 6:51:37 PM | EMS |  | Concentration | mg/L | Y | (fl. conc.) |  | Nickel Total (fl. conc.) | 7440-02-0 | Metals - Total; | Metals - Total; EMS code: NI-T | Chemical | Total | LAB
NI-T | ug/g | Nickel Total | 0190 | D&G:N/P Dig:ICP Anal. | Sample is dried and grind-mixed before nitric perchloric digestion followed by ICP analysis. | 5 | Y | Y | Y | 949 |  |  | 10/24/1996 3:06:16 PM | 1996-04-12 | 6/3/1998 6:51:37 PM | EMS |  | MassPerMass | mg/kg | Y | (mass conc.) |  | Nickel Total (mass conc.) | 7440-02-0 | Metals - Total; | Metals - Total; EMS code: NI-T | Chemical | Total | LAB
NI-T | ug/g wet | Nickel Total | X560 | Metals in Biota (wet), total, acid digestion, ICP-MS | Five grams of wet, homogenized biota is weighed in a centrifuge tube, and 2.5 mL of Nitric Acid and 2.5 mL of Hydrochloric acid are added. The sample is digested at room temperature for one hour and then digested at 90-95 degrees C in a water bath or heat block for one more hour. The sample is cooled and the volume is made to 50 mL with DI. The digestate is then diluted at ten fold and introduced to a collision/reaction cell inductively coupled argon plasma mass spectrometer (CRC ICPMS) by pneumatic nebulization. Each respective element is measured at a specific mass to charge ratio, m/z value expressed in amu (atomic mass units). - Maxxam Analytics (BC). Added to EMS Dec 10, 2008. | 0.1 | Y | Y | Y | 541 | 11/15/2008 4:00:00 AM | 10/30/2023 | 11/28/2023 2:21:33 PM | 2008-12-11 | 3/11/2015 4:17:10 PM | RDGIBSON | EMS | MassPerMass | µg/g | Y | (mass conc.) | (wet) | Nickel Total (wet) (mass conc.) | 7440-02-0 | Metals - Total; | Metals - Total; EMS code: NI-T | Chemical | Total | LAB
NI-T | mg/L | Nickel Total | 0020 | HNO3 Dig:AA Analysis | HNO3 Dig:AA Analysis | 0.1 | Y | Y | Y | 78 | 6/25/2010 3:15:00 PM | 12/1/2010 | 12/9/2010 3:55:27 PM | 1996-04-12 | 6/3/1998 6:51:37 PM | EMS |  | Concentration | mg/L | Y | (fl. conc.) |  | Nickel Total (fl. conc.) | 7440-02-0 | Metals - Total; | Metals - Total; EMS code: NI-T | Chemical | Total | LAB
NI-T | mg/L | Nickel Total | X512 | HNO3,pHto7.5,MIK extr'n,AA | HNO3 Digestion;pH to 7.5; MIK solvent extraction;AA | 0.001 | Y | Y | Y | 535 |  |  | 8/18/1999 4:40:42 PM | 1999-04-22 |  | NPEPPIN |  | Concentration | mg/L | Y | (fl. conc.) |  | Nickel Total (fl. conc.) | 7440-02-0 | Metals - Total; | Metals - Total; EMS code: NI-T | Chemical | Total | LAB
NI-T | mg/L | Nickel Total | F082 | Aliquot,HNO3/HCL dig,ICPMS | Aliquot microwave digested with 3:1 HNO3/HCL. Analysis by ICPMS. For HiVol samples preliminary result are reported in mg/L and loading results are reported in ug/m3 | 0.001 | Y | Y | Y | 27484 | 1/24/2011 | 8/23/2024 | 8/26/2024 4:53:05 PM | 2000-06-13 | 8/15/2019 11:09:55 AM | NPEPPIN | KATFRASE | Concentration | mg/L | Y | (fl. conc.) |  | Nickel Total (fl. conc.) | 7440-02-0 | Metals - Total; | Metals - Total; EMS code: NI-T | Chemical | Total | LAB
NI-T | ug/m3 | Nickel Total | 5312 | Tef:N/P Dig:ICP Analysis | A nitric/perchloric acid digestion if used to bring the metals into solution prior to ICP analysis. For Hi-Vol samples preliminary results are reported as mg/L and 24 hour loading results as ug/m3. | 0.01 | Y | Y | Y | 1185 | 11/7/1997 | 3/29/2000 | 9/28/2000 1:34:43 PM | 1996-10-13 | 6/3/1998 6:51:46 PM | EMS |  | AirConcentration | µg/m³ | Y | (air conc.) |  | Nickel Total (air conc.) | 7440-02-0 | Metals - Total; | Metals - Total; EMS code: NI-T | Chemical | Total | LAB
NI-T | ug/m3 | Nickel Total | 6040 | Aqua Regia Dig:ICP | Aqua regia digestion is used to bring the metals into solution prior to ICP analysis. For Hi-Vol samples, preliminary results are reported as mg/L and24 hr loadings as ug/m3 | 0.01 | Y | Y | Y | 437 | 2/13/2002 | 2/9/2005 | 2/23/2005 1:13:47 AM | 1998-05-05 | 6/3/1998 6:51:48 PM | NPEPPIN |  | AirConcentration | µg/m³ | Y | (air conc.) |  | Nickel Total (air conc.) | 7440-02-0 | Metals - Total; | Metals - Total; EMS code: NI-T | Chemical | Total | LAB
NI-T | ug/g | Nickel Total | 0031 | ICP | ICP | 2 | Y | Y | Y | 103 | 7/26/2005 7:04:00 AM | 5/7/2024 | 7/11/2024 10:43:43 AM | 1996-10-11 | 6/3/1998 6:51:46 PM | NPEPPIN |  | MassPerMass | mg/kg | Y | (mass conc.) |  | Nickel Total (mass conc.) | 7440-02-0 | Metals - Total; | Metals - Total; EMS code: NI-T | Chemical | Total | LAB
NI-T | ug/g | Nickel Total | X492 | Homo,HCl-HNO3 dig,ICP-OES,dry calc | Homogenize samples and take sub-samples. For Total Recoverable Metals Digest with Aqua Regia. Analyses by ICPOES (Inductively Coupled Plasma Optical Emission Spectroscopy). Perform moisture analysis on separate sub-sample and calculate dry value for trace elements. | 2 | Y | Y | Y | 7 | 8/16/2022 6:35:00 PM | 8/17/2022 12:36:00 PM | 2/25/2023 1:41:51 PM | 1998-11-13 |  | NPEPPIN |  | MassPerMass | mg/kg | Y | (mass conc.) |  | Nickel Total (mass conc.) | 7440-02-0 | Metals - Total; | Metals - Total; EMS code: NI-T | Chemical | Total | LAB
NICR | mg/L | Nicarbazin | P015 | Phenoxyacid Herbicides by LC-MS/MS | Phenoxyacid Herbicides by LC-MS/MS | 1e-05 | Y | Y | Y | 68 | 12/6/2021 | 9/9/2022 | 12/8/2022 12:57:05 PM | 2021-12-09 | 12/9/2021 3:22:12 PM | EMS | EMS | Concentration | mg/L | Y | (fl. conc.) |  | Nicarbazin (fl. conc.) | 330-95-0 | Pesticides; | Pesticides; EMS code: NICR | Chemical |  | LAB
NITA | ug/g | Nitrate ¿ Available | CAL2 | Calculation (ug/g) | Calculation (ug/g) | 2 | Y | Y | Y | 2 | 9/3/2021 | 9/3/2021 | 10/19/2022 1:02:09 PM | 2022-10-19 | 10/19/2022 10:39:33 AM | EMS | EMS | MassPerMass | mg/kg | Y | (mass conc.) |  | Nitrate Available (mass conc.) | 14797-55-8 | Nutrients; | Nutrients; EMS code: NITA | Chemical |  | LAB
NO2- | ppb A | Nitrogen Dioxide | DIOXR | Dioxides by Radiello | Dioxides in air by Radiello | 1 | Y | Y | Y | 64 | 8/25/2023 | 8/21/2024 | 8/22/2024 2:48:06 PM | 2023-08-29 |  | EMS |  | DimensionlessRatio | ppb | N | (ratio) |  | Nitrogen Dioxide (ratio) | 10102-44-0 | Inorganics; | Inorganics; EMS code: NO2- | Chemical |  | LAB
NO2- | ppb A | Nitrogen Dioxide | PNO2 | All-Season Passive atmospheric NO2 Sampler | The NO2 passive sample system uses a specially designed rain shelter to allow the passive sampler to be installed face downwards and to allow the air movement to cross the surface of the diffusion barrier. After exposure, the sampler media is extracted in the laboratory with de-ionized (DI) water. The extract is analyzed by Ion Chromatography. The sampling rate of the passive sampler is calculated using an equation which takes into account the dependence on associated meteorological factors. | 0.1 | Y | Y | Y | 49 | 7/15/2024 12:27:00 PM | 8/19/2024 9:20:00 AM | 8/26/2024 11:29:38 AM | 2024-07-31 |  | RDGIBSON |  | DimensionlessRatio | ppb | N | (ratio) |  | Nitrogen Dioxide (ratio) | 10102-44-0 | Inorganics; | Inorganics; EMS code: NO2- | Chemical |  | LAB
NO2- | ppm A | Nitrogen Dioxide | X301 | Alkaline/MnO4,Col'metric | Alkaline/MnO4;Col'metric | 0 | Y | Y | Y | 389 |  |  | 8/22/2024 1:49:14 PM | 1996-04-12 | 8/27/2020 11:18:18 AM | EMS | RDGIBSON | DimensionlessRatio | ppm | N | (ratio) |  | Nitrogen Dioxide (ratio) | 10102-44-0 | Inorganics; | Inorganics; EMS code: NO2- | Chemical |  | LAB
NO3- | ug/m3 | Nitrate-Soluble | 5022 | Water Extr,Ion Chr-Anion | Discs cut from the HiVol filter are extracted with deionized water and the resulting anions are separated an measured using an ion chromatograph. (See Anions-Ion Chromatography-Precipitation, methods NO3-5068, S04-5068, CL-S5068). Values are then converted to units of ug/m3. | 5 | Y | Y | Y | 71 |  |  | 10/24/1996 6:45:17 AM | 1996-04-12 | 6/3/1998 6:51:37 PM | EMS |  | AirConcentration | µg/m³ | Y | (air conc.) |  | Nitrate-Soluble (air conc.) | 14797-55-8 | Nutrients; | Nutrients; EMS code: NO3- | Chemical | Soluble | LAB
NO3- | ug/m3 | Nitrate-Soluble | EQ40 | EQUIS (mg/m3) | EQUIS (mg/m3) | 0 | N | Y | Y | 27 |  |  | 10/22/1996 8:19:52 PM | 1996-04-12 | 6/3/1998 6:51:37 PM | EMS |  | AirConcentration | µg/m³ | Y | (air conc.) |  | Nitrate-Soluble (air conc.) | 14797-55-8 | Nutrients; | Nutrients; EMS code: NO3- | Chemical | Soluble | LAB
NO3- | mg/L | Nitrate-Soluble | 5068 | Ion Chromatography-Anion | (a) For the Soluble Particulates Pkg., a portion of filtrate from the Particulate-Total procedure made up to 500 mL is analyzed by Ion Chromatography. Intermediate results are in mg/L. Loading results are then calculated with units of mg/dm2/d. (b) For precipitation samples, samples and standards are spiked with concentrated carbonate/bicarbonate solution to give the sample the same background as the eluent used (avoiding a "water dip" effect). The anions of interest are separated through an anion "guard" and anion "separator" column. An anion micro membrane suppressor following the separator columns is used to reduce the background eluent conductivity by converting the carbonate and bicarbonate species to carbonic acid, while enhancing the conductivity of the ions of interest by converting these ions to their corresponding acids. A conductivity detector senses the same species in direct proportion to their initial concentration. (c) For HiVol samples, discs are cut from the HiVol filter, extracted with DI water, and anions analyzed by ion chromatograph, Preliminary results are reported in mg/L and intermediate HiVol loading results are reported as ug/m3. (*) Note that the air-convention is to report on basis of ion weight not element, e.g. as NO3 for nitrate, and as PO4 for phosphate. | 0.01 | Y | Y | Y | 6600 | 4/30/1997 3:50:00 PM | 7/26/2012 7:03:00 AM | 9/26/2012 1:07:08 PM | 1996-04-12 | 6/3/1998 6:51:37 PM | EMS |  | Concentration | mg/L | Y | (fl. conc.) |  | Nitrate-Soluble (fl. conc.) | 14797-55-8 | Nutrients; | Nutrients; EMS code: NO3- | Chemical | Soluble | LAB
NO3- | mg/dm2/d | Nitrate-Soluble | EQ15 | EQUIS (mg/dm2/d) | EQUIS (mg/dm2/d) | 0.002 | N | Y | Y | 5 |  |  | 10/22/1996 8:20:30 PM | 1996-04-12 | 6/3/1998 6:51:37 PM | EMS |  | MassFlux | mg/dm²/d | Y | (mass flux) |  | Nitrate-Soluble (mass flux) | 14797-55-8 | Nutrients; | Nutrients; EMS code: NO3- | Chemical | Soluble | LAB
NO3- | mg/L | Nitrate-Soluble | EQ01 | EQUIS (mg/L) | EQUIS (mg/L) | 0.01 | N | Y | Y | 130 |  |  | 10/22/1996 11:12:48 PM | 1996-04-12 | 6/3/1998 6:51:45 PM | EMS |  | Concentration | mg/L | Y | (fl. conc.) |  | Nitrate-Soluble (fl. conc.) | 14797-55-8 | Nutrients; | Nutrients; EMS code: NO3- | Chemical | Soluble | LAB
NONA | % | Nonadecanoic | SURR | Surrogate Recovery | Surrogate Recovery | 0.001 | Y | Y | Y | 86 | 1/31/2002 | 5/13/2003 | 4/1/2004 1:08:46 AM | 1996-04-12 | 1/3/2023 11:50:38 AM | EMS | EMS | DimensionlessRatio | % | N | (ratio) |  | Nonadecanoic (ratio) | 646-30-0 | Organics; | Organics; EMS code: NONA | Chemical |  | SURROGATE_RESULT
NONA | ug/g | Nonadecanoic | MP02 | GC/MS | GC/MS | 0.05 | Y | Y | Y | 16 | 6/21/2002 12:00:00 PM | 6/21/2002 12:25:00 PM | 3/27/2004 1:07:20 AM | 1996-04-12 | 6/3/1998 6:51:42 PM | EMS |  | MassPerMass | mg/kg | Y | (mass conc.) |  | Nonadecanoic (mass conc.) | 646-30-0 | Organics; | Organics; EMS code: NONA | Chemical |  | LAB
NORG | mg/L | Norgestimate CAS: 20799-27-3 | XMSP | HPLC and Tri. Quadpole MS/MS | EPA 1694 An aliquot of filtered sample containing 5% organic is injected directly. Each compound is separated by reversed phase HPLC. Analyzed by Electro Spray Ionization MS/MS detection using Triple Quadrupole MS/MS detector. Quantiitation with internal or external standard. |  | Y | Y | Y | 12 | 5/8/2020 | 6/4/2024 | 8/21/2024 3:32:03 PM | 2022-05-04 |  | EMS |  | Concentration | mg/L | Y | (fl. conc.) |  | Norgestimate (fl. conc.) | 20799-27-3 | Organics; | Organics; EMS code: NORG | Chemical |  | LAB
NOVB | mg/L | Novobiocin CAS: 303-81-1 | XMSP | HPLC and Tri. Quadpole MS/MS | EPA 1694 An aliquot of filtered sample containing 5% organic is injected directly. Each compound is separated by reversed phase HPLC. Analyzed by Electro Spray Ionization MS/MS detection using Triple Quadrupole MS/MS detector. Quantiitation with internal or external standard. |  | Y | Y | Y | 12 | 5/8/2020 | 6/4/2024 | 8/21/2024 3:32:03 PM | 2022-05-04 |  | EMS |  | Concentration | mg/L | Y | (fl. conc.) |  | Novobiocin (fl. conc.) | 303-81-1 | Organics; | Organics; EMS code: NOVB | Chemical |  | LAB
NOX- | ug/m3 | Nitrogen Oxides | 5102 | Auto Col. Diazo'n- Stack | Auto Col. Diazo'n- Stack | 0.001 | Y | Y | Y | 76 | 3/24/2014 8:00:00 AM | 6/28/2024 8:00:00 AM | 7/30/2024 10:08:09 AM | 1996-04-12 | 6/3/1998 6:51:37 PM | EMS |  | AirConcentration | µg/m³ | Y | (air conc.) |  | Nitrogen Oxides (air conc.) | 11104-93-1 | Inorganics; | Inorganics; EMS code: NOX- | Chemical |  | LAB
NOX- | ug/m3 | Nitrogen Oxides | 5098 | Ion Chrom-Anions Stack | Ion Chrom-Anions Stack | 0.001 | Y | Y | Y | 2 |  |  | 5/28/1998 8:52:12 AM | 1996-04-12 | 6/3/1998 6:51:37 PM | EMS |  | AirConcentration | µg/m³ | Y | (air conc.) |  | Nitrogen Oxides (air conc.) | 11104-93-1 | Inorganics; | Inorganics; EMS code: NOX- | Chemical |  | LAB
NOX- | ug/m3 | Nitrogen Oxides | EQ40 | EQUIS (mg/m3) | EQUIS (mg/m3) | 0.001 | N | Y | Y | 5 |  |  | 10/22/1996 9:05:29 PM | 1996-04-12 | 6/3/1998 6:51:37 PM | EMS |  | AirConcentration | µg/m³ | Y | (air conc.) |  | Nitrogen Oxides (air conc.) | 11104-93-1 | Inorganics; | Inorganics; EMS code: NOX- | Chemical |  | LAB
NPES | ug/g | Nonylphenol Monoethoxylates | A849 | J. of Chrom. A849 (1999) 467-482 | Sample is derivatized and injected directly. Each compound is separated by reversed phase HPLC and analyzed by Electro Spray Ionization MS/MS detection using Triple Quadrupole MS/MS detector. Internal standards are used for quantitation. | 0.25 | Y | Y | Y | 8 | 3/14/2020 | 3/16/2020 | 2/21/2023 2:10:31 PM | 2022-08-08 |  | EMS |  | MassPerMass | mg/kg | Y | (mass conc.) |  | Nonylphenol Monoethoxylates (mass conc.) | 27986-36-3 | Organics; | Organics; EMS code: NPES | Chemical |  | LAB
NPHEN | ug/g | n-Nonylphenol CAS: 25154-52-3 | A849 | J. of Chrom. A849 (1999) 467-482 | Sample is derivatized and injected directly. Each compound is separated by reversed phase HPLC and analyzed by Electro Spray Ionization MS/MS detection using Triple Quadrupole MS/MS detector. Internal standards are used for quantitation. | 0.25 | Y | Y | Y | 8 | 3/14/2020 | 3/16/2020 | 2/21/2023 2:10:31 PM | 2022-08-11 |  | EMS |  | MassPerMass | mg/kg | Y | (mass conc.) |  | n-Nonylphenol (mass conc.) | 25154-52-3 | Organics; | Organics; EMS code: NPHEN | Chemical |  | LAB
NPR | None | Neutralization potential Ratio [NPR]; (NP/MPA) | VOL08N | ALS Minerals OA-VOL08mn | ALS Minerals OA-VOL08mn. Determine the acid generation power of a sample as tonne CaCO3 per kilo-tonne using MEND 1991 method. | 0.01 | Y | Y | Y | 0 |  |  |  | 2024-01-29 |  | EMS |  | None | None | N | (unitless) |  | Neutralization potential Ratio [NPR]; (NP/MPA) (unitless) |  | Physical Properties; | Physical Properties; EMS code: NPR | Physical |  | LAB
NPR | None | Neutralization potential Ratio [NPR]; (NP/MPA) | VOL08M | ALS Minerals OA-VOL08m | ALS Minerals OA-VOL08m. Determine the acid generation power of a sample as tonne CaCO3 per kilo-tonne using modified Sobek method | 0.01 | Y | Y | Y | 0 |  |  |  | 2024-01-29 |  | EMS |  | None | None | N | (unitless) |  | Neutralization potential Ratio [NPR]; (NP/MPA) (unitless) |  | Physical Properties; | Physical Properties; EMS code: NPR | Physical |  | LAB
NPR | None | Neutralization potential Ratio [NPR]; (NP/MPA) | VOL08 | ALS Minerals OA-VOL08 | ALS Minerals OA-VOL08. Determine the acid generation power of a sample as tonne CaCO3 per kilo-tonne. | 0.01 | Y | Y | Y | 0 |  |  |  | 2024-01-29 |  | RDGIBSON |  | None | None | N | (unitless) |  | Neutralization potential Ratio [NPR]; (NP/MPA) (unitless) |  | Physical Properties; | Physical Properties; EMS code: NPR | Physical |  | LAB
NPR | None | Neutralization potential Ratio [NPR]; (NP/MPA) | VOL08S | ALS Minerals OA-VOL08s | ALS Minerals OA-VOL08s. Determine the acid generation power of a sample as tonne CaCO3 per kilo-tonne using Siderite correction method. | 0.01 | Y | Y | Y | 0 |  |  |  | 2024-01-29 |  | EMS |  | None | None | N | (unitless) |  | Neutralization potential Ratio [NPR]; (NP/MPA) (unitless) |  | Physical Properties; | Physical Properties; EMS code: NPR | Physical |  | LAB
NTIN | mg/L | Nitrogen Total Inorganic | XCAL | Calculation (mg/L) | Calculation (mg/L) | 0.025 | Y | Y | Y | 106 | 1/27/2021 | 8/14/2024 | 8/20/2024 12:44:02 PM | 2021-08-31 |  | RDGIBSON |  | Concentration | mg/L | Y | (fl. conc.) |  | Nitrogen Total Inorganic (fl. conc.) |  | Nutrients; | Nutrients; EMS code: NTIN | Chemical | Total | LAB
O101 | pg/g wet | 8Chldibenzodioxin | F119 | Solids,Isotope dil,HRGC/HRMS | EPA Method 1613: for Sediments and Fish tissue developed for isomer specific determination of the 2,3,7,8-substituted, tetra-through octa-chlorinated, dibenzo-p-dioxins and dibenzofurans in aqueous, solid and tissue matrices by isotope dilution, high resolution capillary column gas chromatography (HRGC)/high resolution mass spectrometry (HRMS) |  | Y | Y | Y | 5 | 6/16/2000 6:33:00 AM | 10/4/2000 2:53:00 PM | 6/18/2002 3:50:03 PM | 2001-10-01 | 10/4/2001 3:20:45 PM | NPEPPIN | NPEPPIN | MassPerMass | pg/g | Y | (mass conc.) | (wet) | 8Chldibenzodioxin (wet) (mass conc.) | 3268-87-9 | Dioxins & Furans; | Dioxins & Furans; EMS code: O101 | Chemical |  | LAB
O101 | pg/L | 8Chldibenzodioxin | X253 | Tol.-Ac/Cl2CH2 ext:GC/MS | C-13 labelled PCDD and PCDF congeners are added to a 1L sample of pulp effluent, which is separated into solid and liquid phases by filtration through a glass fibre filter. The solid phase is aire-dried overnight before undergoing a Soxhlet extraction with toluene. The liquid phase is extracted with dichloromethane in a separatory funnel. The two extracts are combined, concentrated, and then passed through a series of columns which remove the bulk of the organic matrix co-extracted with PCDD/PCDF. One resulting fraction is concentrated to a known small volume for analysis by HRGC/HRMS/DS. | 1 | Y | Y | Y | 155 |  |  | 6/25/2012 2:16:03 PM | 1996-04-12 | 6/3/1998 6:51:37 PM | EMS |  | Concentration | pg/L | Y | (fl. conc.) |  | 8Chldibenzodioxin (fl. conc.) | 3268-87-9 | Dioxins & Furans; | Dioxins & Furans; EMS code: O101 | Chemical |  | SURROGATE_RESULT
O101 | pg/L | 8Chldibenzodioxin | F102 | Env.Canada(1992),DioxinsProc,Pulp Effluents | Environment Canada (1992);Dioxins Procedures for Pulp Effluents |  | Y | Y | Y | 6 | 5/1/2001 7:37:00 AM | 5/1/2001 7:37:00 AM | 11/9/2011 3:41:39 PM | 2000-10-25 |  | NPEPPIN |  | Concentration | pg/L | Y | (fl. conc.) |  | 8Chldibenzodioxin (fl. conc.) | 3268-87-9 | Dioxins & Furans; | Dioxins & Furans; EMS code: O101 | Chemical |  | LAB
O101 | pg/L | 8Chldibenzodioxin | F114 | Isotope dil,HRGC/HRMS | EPA Method 1613: Developed for isomer specific determination of the 2,3,7,8-substituted, tetra-through octa-chlorinated, dibenzo-p-dioxins and dibenzofurans in aqueous, solid and tissue matrices by isotope dilution, high resolution capillary column gas chromatography (HRGC)/high resolution mass spectrometry (HRMS) |  | Y | Y | Y | 18 |  |  | 2/21/2024 1:00:50 PM | 2001-05-24 |  | NPEPPIN |  | Concentration | pg/L | Y | (fl. conc.) |  | 8Chldibenzodioxin (fl. conc.) | 3268-87-9 | Dioxins & Furans; | Dioxins & Furans; EMS code: O101 | Chemical |  | LAB
O102 | pg/L | 8Chldibenzofuran | F114 | Isotope dil,HRGC/HRMS | EPA Method 1613: Developed for isomer specific determination of the 2,3,7,8-substituted, tetra-through octa-chlorinated, dibenzo-p-dioxins and dibenzofurans in aqueous, solid and tissue matrices by isotope dilution, high resolution capillary column gas chromatography (HRGC)/high resolution mass spectrometry (HRMS) |  | Y | Y | Y | 18 |  |  | 2/21/2024 1:00:50 PM | 2001-05-24 |  | NPEPPIN |  | Concentration | pg/L | Y | (fl. conc.) |  | 8Chldibenzofuran (fl. conc.) | 39001-02-0 | Dioxins & Furans; | Dioxins & Furans; EMS code: O102 | Chemical |  | LAB
O102 | pg/g wet | 8Chldibenzofuran | F119 | Solids,Isotope dil,HRGC/HRMS | EPA Method 1613: for Sediments and Fish tissue developed for isomer specific determination of the 2,3,7,8-substituted, tetra-through octa-chlorinated, dibenzo-p-dioxins and dibenzofurans in aqueous, solid and tissue matrices by isotope dilution, high resolution capillary column gas chromatography (HRGC)/high resolution mass spectrometry (HRMS) |  | Y | Y | Y | 5 | 6/16/2000 6:33:00 AM | 10/4/2000 2:53:00 PM | 6/18/2002 3:50:03 PM | 2001-10-01 | 10/4/2001 3:20:45 PM | NPEPPIN | NPEPPIN | MassPerMass | pg/g | Y | (mass conc.) | (wet) | 8Chldibenzofuran (wet) (mass conc.) | 39001-02-0 | Dioxins & Furans; | Dioxins & Furans; EMS code: O102 | Chemical |  | LAB
O102 | pg/L | 8Chldibenzofuran | X253 | Tol.-Ac/Cl2CH2 ext:GC/MS | C-13 labelled PCDD and PCDF congeners are added to a 1L sample of pulp effluent, which is separated into solid and liquid phases by filtration through a glass fibre filter. The solid phase is aire-dried overnight before undergoing a Soxhlet extraction with toluene. The liquid phase is extracted with dichloromethane in a separatory funnel. The two extracts are combined, concentrated, and then passed through a series of columns which remove the bulk of the organic matrix co-extracted with PCDD/PCDF. One resulting fraction is concentrated to a known small volume for analysis by HRGC/HRMS/DS. | 1e-05 | Y | Y | Y | 155 |  |  | 6/25/2012 2:16:03 PM | 1996-04-12 | 6/3/1998 6:51:37 PM | EMS |  | Concentration | pg/L | Y | (fl. conc.) |  | 8Chldibenzofuran (fl. conc.) | 39001-02-0 | Dioxins & Furans; | Dioxins & Furans; EMS code: O102 | Chemical |  | SURROGATE_RESULT
O102 | pg/L | 8Chldibenzofuran | F102 | Env.Canada(1992),DioxinsProc,Pulp Effluents | Environment Canada (1992);Dioxins Procedures for Pulp Effluents |  | Y | Y | Y | 6 | 5/1/2001 7:37:00 AM | 5/1/2001 7:37:00 AM | 11/9/2011 3:41:39 PM | 2000-10-25 |  | NPEPPIN |  | Concentration | pg/L | Y | (fl. conc.) |  | 8Chldibenzofuran (fl. conc.) | 39001-02-0 | Dioxins & Furans; | Dioxins & Furans; EMS code: O102 | Chemical |  | LAB
O3 | mg/L | Ozone for water | 8311 | INDIGO | Sample adjusted to pH to 2.5; indigo reagent reacts with ozone. Measured at 600 nm with spectrophotometers or 610 nm with colorimeter. | 0.01 | Y | Y | Y | 2969 |  |  | 7/29/2024 2:00:43 PM | 2019-04-11 |  | RPWILLIA |  | Concentration | mg/L | Y | (fl. conc.) |  | Ozone for water (fl. conc.) | 10028-15-6 | Inorganics; | Inorganics; EMS code: O3 | Chemical |  | LAB
OCPS | None | Organo-chlorines | SCA1 | Organochlorine Pest Scan | Organochlorine Pest Scan | 0 | N | Y | Y | 2 |  |  | 10/24/1996 9:39:03 AM | 1996-04-12 | 6/3/1998 6:51:37 PM | EMS |  | None | None | N | (unitless) |  | Organo-chlorines (unitless) |  |  | EMS code: OCPS | Chemical |  | LAB
OCTA | mg/L | Octadecanoic Acid | RFAS | Separation of resin and fatty acids (RFAs) in water by GC/MS | Scientific Principle The separation of resin and fatty acids (RFAs) in water samples is performed using gas chromatography/mass spectrometry (GC/MS) in the Selected Ion Monitoring (SIM) mode. The separation is based on compound size; boiling point and polarity. Analyte concentration is calculated using the internal standard method. RFA compounds are extracted from water samples using a liquid-liquid solvent extraction with DCM. Extraction occurs after the sample is brought to a pH of ~ 2 and placed in a rotary tumbler (at approx. 40 rpm) for 1 hour. The resulting extract is transferred to a vial; dried with acidified baked sodium sulfate; methylated with diazomethane and concentrated to 1mL under nitrogen prior to analysis. The resulting extracts are vialed and fortified with labeled internal standard compounds prior to analysis by gas chromatography with a (50%-Phenyl)-methylpolysiloxane (DB-17) column and a mass spectrometer detector in SIM Mode. | 0.006 | Y | Y | Y | 1 | 9/17/2021 7:49:00 PM | 9/17/2021 7:49:00 PM | 2/1/2022 2:38:14 PM | 2022-01-19 |  | EMS |  | Concentration | mg/L | Y | (fl. conc.) |  | Octadecanoic Acid (fl. conc.) | 57-11-4 | Organics; | Organics; EMS code: OCTA | Chemical |  | LAB
ODE | ug/g | Octylphenol Diethoxylates | A849 | J. of Chrom. A849 (1999) 467-482 | Sample is derivatized and injected directly. Each compound is separated by reversed phase HPLC and analyzed by Electro Spray Ionization MS/MS detection using Triple Quadrupole MS/MS detector. Internal standards are used for quantitation. | 0.025 | Y | Y | Y | 8 | 3/14/2020 | 3/16/2020 | 2/21/2023 2:10:31 PM | 2022-08-08 |  | EMS |  | MassPerMass | mg/kg | Y | (mass conc.) |  | Octylphenol Diethoxylates (mass conc.) | 51437-90-2 | Organics; | Organics; EMS code: ODE | Chemical |  | LAB
OME | ug/g | Octylphenol Monoethoxylates | A849 | J. of Chrom. A849 (1999) 467-482 | Sample is derivatized and injected directly. Each compound is separated by reversed phase HPLC and analyzed by Electro Spray Ionization MS/MS detection using Triple Quadrupole MS/MS detector. Internal standards are used for quantitation. | 0.5 | Y | Y | Y | 8 | 3/14/2020 | 3/16/2020 | 2/21/2023 2:10:31 PM | 2022-08-08 |  | EMS |  | MassPerMass | mg/kg | Y | (mass conc.) |  | Octylphenol Monoethoxylates (mass conc.) | 51437-89-9 | Organics; | Organics; EMS code: OME | Chemical |  | LAB
OMET | ug/g | Omethoate | ZN03 | DCM+Ac.Extr:GC/MS | DCM+Ac.Extr:GC/MS EPA 8270E (mod) | 0.1 | Y | Y | Y | 19 | 4/20/2002 10:25:00 AM | 10/16/2002 3:20:00 PM | 2/25/2004 1:09:47 AM | 1996-04-12 | 11/16/1999 2:27:40 PM | EMS | NPEPPIN | MassPerMass | mg/kg | Y | (mass conc.) |  | Omethoate (mass conc.) | 1113-02-6 | Pesticides; | Pesticides; EMS code: OMET | Chemical |  | LAB
OMET | mg/L | Omethoate | ZN01 | DCM+Ac.Extr:GC/MS | DCM+Ac.Extr:GC/MS | 0.001 | Y | Y | Y | 52 | 10/3/2002 | 3/29/2012 10:15:00 PM | 5/3/2012 9:27:51 AM | 1996-04-12 | 11/16/1999 4:06:27 PM | EMS | NPEPPIN | Concentration | mg/L | Y | (fl. conc.) |  | Omethoate (fl. conc.) | 1113-02-6 | Pesticides; | Pesticides; EMS code: OMET | Chemical |  | LAB
OPPS | None | Organo-phosphorus | SCA1 | Organochlorine Pest Scan | Organochlorine Pest Scan | 0 | N | Y | Y | 2 |  |  | 10/24/1996 9:39:03 AM | 1996-04-12 | 6/3/1998 6:51:37 PM | EMS |  | None | None | N | (unitless) |  | Organo-phosphorus (unitless) |  | Organics; | Organics; EMS code: OPPS | Chemical |  | LAB
ORSP | None | Organics Special | ORSP | Org. Special Analysis | Org. Special Analysis | 0 | N | Y | N | 11 |  |  | 10/24/1996 9:41:24 AM | 1996-04-12 | 6/3/1998 6:51:45 PM | EMS |  | None | None | N | (unitless) |  | Organics Special (unitless) |  |  | EMS code: ORSP | Chemical |  | LAB
ORSP | SWU | Organics Special | X543 | pH adj,Filt,Spect,243.7nm | UV absorption is a useful surrogate measure of selected organic constituents in fresh waters,salt waters and waste water. Sample pH may be adjusted with HCl or NaOH. In poorly buffered samples, an appropriate non-UV-absorbing buffer system such as a phosphate buffer may be used. Once the sample pH has been measured/adjusted, filter sample. Equilibrate spectrophotometer and measure UV absorbance at 253.7 nm of at least two filtered portions at room temperature. | 0.005 | Y | Y | Y | 35 | 5/18/2001 | 3/26/2002 | 4/3/2002 1:05:16 PM | 2001-06-18 |  | NPEPPIN |  | Apperance | SWU | N | (app.) |  | Organics Special (app.) |  |  | EMS code: ORSP | Chemical |  | LAB
ORYZ | mg/L | Oryzalin | P015 | Phenoxyacid Herbicides by LC-MS/MS | Phenoxyacid Herbicides by LC-MS/MS | 5e-05 | Y | Y | Y | 68 | 12/6/2021 | 9/9/2022 | 12/8/2022 12:57:05 PM | 2021-12-09 | 12/9/2021 3:22:12 PM | EMS | EMS | Concentration | mg/L | Y | (fl. conc.) |  | Oryzalin (fl. conc.) | 19044-88-3 | Pesticides; | Pesticides; EMS code: ORYZ | Chemical |  | LAB
OX02 | mg/L | Oxamyl | P045 | DCM Extr:HPLC | DCM Extr:HPLC | 0.0005 | Y | Y | Y | 1 | 7/29/2005 7:10:00 AM | 7/29/2005 7:10:00 AM | 6/2/2009 12:01:08 PM | 1996-04-12 | 6/3/1998 6:51:42 PM | EMS |  | Concentration | mg/L | Y | (fl. conc.) |  | Oxamyl (fl. conc.) | 23135-22-0 | Pesticides; | Pesticides; EMS code: OX02 | Chemical |  | LAB
P--D | mg/L | Phosphorus Total Dissolved | X158 | FF:Dig. Auto Ascorbic Ac | The filtered sample is acidified, potassium persulfate is added and the mixture is digested at elevated temperature and pressure in a steam autoclave. After digestion, the sample is reacted with a mixture of ammonium molybdate and potassium antimonyl tartrate in acid solution. Ascorbic acid is then added to produce a blue coloured product with an absorbance maximum at 880 nm. The absorbance of the solution is measured and the phosphorus concentration is determined by comparison with standards treated in the same manner. | 0.002 | Y | Y | Y | 44 |  |  | 3/22/2001 8:16:27 AM | 1996-04-12 | 6/3/1998 6:51:45 PM | EMS |  | Concentration | mg/L | Y | (fl. conc.) |  | Phosphorus Total Dissolved (fl. conc.) | 7723-14-0 | Metals - Dissolved; | Metals - Dissolved; EMS code: P--D | Chemical | Dissolved | LAB
P--D | mg/L | Phosphorus Total Dissolved | F122 | Persulfate Dig,Col'm | Aliquot of the sample is filtered. The filtrate is digestested by pwersulphate followed by a colourimetric reading |  | Y | Y | Y | 41138 | 3/25/2011 | 8/24/2024 | 8/26/2024 3:51:07 PM | 2002-07-22 |  | UMULLER |  | Concentration | mg/L | Y | (fl. conc.) |  | Phosphorus Total Dissolved (fl. conc.) | 7723-14-0 | Metals - Dissolved; | Metals - Dissolved; EMS code: P--D | Chemical | Dissolved | LAB
P--D | mg/L | Phosphorus Total Dissolved | 2710 | Froz.Dig.Aut.Ascorbic Ac | Froz.Dig.Aut.Ascorbic Ac | 0.003 | Y | Y | Y | 2134 |  |  | 10/24/1996 2:20:17 PM | 1996-04-12 | 6/3/1998 6:51:37 PM | EMS |  | Concentration | mg/L | Y | (fl. conc.) |  | Phosphorus Total Dissolved (fl. conc.) | 7723-14-0 | Metals - Dissolved; | Metals - Dissolved; EMS code: P--D | Chemical | Dissolved | LAB
P--D | mg/L | Phosphorus Total Dissolved | MSEA | Metals in Seawater (undigested) by (CRC) ICP-MS | This method describes the multi-elemental determination of trace elements by Collision/Reaction Cell (CRC) ICP-MS. The method measures ions produced by a radio-frequency inductively coupled plasma. Analyte species originating in a liquid are pneumatically nebulized and the resulting aerosol transported by argon gas into the plasma torch. The ions produced are entrained in the plasma gas and extracted, by means of a differentially pumped vacuum interface, into a mass spectrometer. The ions flow through a Collision/Reaction Cell (CRC) where depending on the mode of analysis either the ions passes through the cell unhindered (no gas - classical ICP-MS mode), or collisions/reactions chemistry takes place between the ions and a neutral gas such as hydrogen or helium. Any polyatomic ions which give interference on the analytes are removed. The ions are then sorted according to their mass-to-charge ratios by a quadrupole mass spectrometer. The ions are then sorted according to their mass-to-charge ratios by a quadrupole mass spectrometer. The ions transmitted through the quadrupole are quantified by a channel electron multiplier and the ion information processed by a data handling system. Although most polyatomic interference is removed, any additional interference must be assessed and valid corrections applied or the data flagged to indicate problems. Interference correction must include compensation for background ions contributed by the plasma gas, reagents, and constituents of the sample matrix. Instrument drift as well as suppressions or enhancements of instrument response caused by the sample matrix must be corrected for by the use of internal standards. | 0.01 | Y | Y | Y | 320 | 7/29/2011 4:05:00 AM | 4/29/2024 | 5/16/2024 2:00:14 PM | 2009-09-08 |  | RDGIBSON |  | Concentration | mg/L | Y | (fl. conc.) |  | Phosphorus Total Dissolved (fl. conc.) | 7723-14-0 | Metals - Dissolved; | Metals - Dissolved; EMS code: P--D | Chemical | Dissolved | LAB
P--D | mg/L | Phosphorus Total Dissolved | X285 | HNO3 Dig:ICP Analysis | HNO3 Dig:ICP Analysis | 0.03 | Y | Y | Y | 651 | 12/9/1992 4:16:00 PM | 11/13/2002 | 2/19/2004 4:22:43 PM | 1998-05-22 | 6/3/1998 6:51:49 PM | NPEPPIN |  | Concentration | mg/L | Y | (fl. conc.) |  | Phosphorus Total Dissolved (fl. conc.) | 7723-14-0 | Metals - Dissolved; | Metals - Dissolved; EMS code: P--D | Chemical | Dissolved | LAB
P--D | mg/L | Phosphorus Total Dissolved | MWID | Metals in Water by ICPMS (digested) | Metals in Water by ICPMS (digested) Based on EPA 200.8 | 0.04 | Y | Y | Y | 840 |  |  | 7/31/2024 1:46:41 PM | 2009-12-24 |  | RDGIBSON |  | Concentration | mg/L | Y | (fl. conc.) |  | Phosphorus Total Dissolved (fl. conc.) | 7723-14-0 | Metals - Dissolved; | Metals - Dissolved; EMS code: P--D | Chemical | Dissolved | LAB
P--D | mg/L | Phosphorus Total Dissolved | X030 | Dig,Ascorbic,Col'metric | Dig;Ascorbic;Col'metric | 0.003 | Y | Y | Y | 5369 | 6/3/1997 | 4/4/2023 | 8/20/2024 8:39:43 AM | 1996-04-12 | 6/3/1998 6:51:45 PM | EMS |  | Concentration | mg/L | Y | (fl. conc.) |  | Phosphorus Total Dissolved (fl. conc.) | 7723-14-0 | Metals - Dissolved; | Metals - Dissolved; EMS code: P--D | Chemical | Dissolved | LAB
P--D | mg/L | Phosphorus Total Dissolved | X185 | Dig. Auto.Ascorbic Acid | The unfiltered sample is acidified, potassium persulfate is added and the mixture is digested at elevated temperature and pressure in a steam autoclave. After digestion, the sample is reacted with a mixture of ammonium molybdate and potassium antimonyl tartrate in acid solution. Ascorbic acid is then added to produce a blue coloured product with an absorbance maximum at 880 nm. The absorbance of the solution is measured and the phosphorus concentration is determined by comparison with standards treated in the same manner. Automated procedure. This unfiltered method is sometimes used for a clear turbid-free solution where filtering is not considered necessary. | 0.003 | Y | Y | Y | 27246 | 8/10/2006 | 3/13/2015 | 1/12/2016 5:01:27 PM | 1996-04-12 | 6/3/1998 6:51:37 PM | EMS |  | Concentration | mg/L | Y | (fl. conc.) |  | Phosphorus Total Dissolved (fl. conc.) | 7723-14-0 | Metals - Dissolved; | Metals - Dissolved; EMS code: P--D | Chemical | Dissolved | LAB
P--D | mg/L | Phosphorus Total Dissolved | F039 | Field filt + acid, ICP-AES | Dissolved and extractable water samples are ready for analysis once collected. Total samples are autoclaved with HNO3 and hydrochloric acid (HCI) in capped Teflon vials. Analysis is performed via Inductively Coupled Argon Plasma Atomic Emission Spectrometry (ICP-AES)(simultaneous multielement analysis), Graphite Furnace Atomic Absorption Spectrometry (GF-AAS)(low level analysis) and/or Flame Atomic Absorption Spectrometry (F-AAS). | 0.1 | Y | Y | Y | 7954 | 7/29/1996 6:03:00 PM | 7/5/2019 1:21:00 PM | 9/11/2019 8:47:50 AM | 1997-05-08 | 6/3/1998 6:51:46 PM | MPAK |  | Concentration | mg/L | Y | (fl. conc.) |  | Phosphorus Total Dissolved (fl. conc.) | 7723-14-0 | Metals - Dissolved; | Metals - Dissolved; EMS code: P--D | Chemical | Dissolved | LAB
P--D | mg/L | Phosphorus Total Dissolved | F032 | Filt.,Acid,ICPMS | There is no acid digestion required for this analysis. A water sample is collected and preserved in 0.2 percent nitric acid. To analyse the extractable metals the sample is shaken, allowed to sit for at least twenty four hours, and a decanted portion is introduced to the spectrometer. When the dissolved metals are required, the samples must first be filtered through a 0.45 micron cellulose acetate filter paper and then followed by preservation in 0.2 percent nitric acid. It is important to note that different laboratories may have very different detection limits for some variables analyzed by ICPMS. | 0.01 | Y | Y | Y | 41955 | 5/7/1998 | 8/26/2024 | 8/26/2024 4:53:05 PM | 1998-05-14 | 4/29/2011 8:50:50 AM | NPEPPIN | RDGIBSON | Concentration | mg/L | Y | (fl. conc.) |  | Phosphorus Total Dissolved (fl. conc.) | 7723-14-0 | Metals - Dissolved; | Metals - Dissolved; EMS code: P--D | Chemical | Dissolved | LAB
P--D | mg/L | Phosphorus Total Dissolved | F049 | Filt,HNO3 dig,conc'n,nebul'ICP-OES | This method provides an automated procedure for the Inductively Coupled Argon Plasma Optical Emission Spectrophotometric (ICP-OES) determination and quantitation of trace amounts of metals in surface, ground and precipitation waters after a manual acid digestion and concentration procedure and ultrasonic nebulisation. Organically-bound metals and metal complexes are converted to metal ions by digestion with nitric acid. To increase the signal detected by the ICP, the samples are concentrated 10 times. The liquid sample is pumped onto the piezoelectric transducer, where the ultrasonic action efficiently converts it to a fine, dense aerosol. The nebulizer carrier gas flow then transports the wet aerosol through a heated U-tube, where the solvent is vaporized. Subsequently, most of the solvent vapour is removed from the aerosol stream by passing the stream through a chilled condenser. The result is a dry, analyte-laden aerosol which is introduced to the plasma. There, the high temperature desolvates the aerosol and dissociates the sample material into individual atoms and ions, which are excited to emit light at wave lengths characteristic of the elements present. Dissolved samples are field filtered through 0.45 micron cellulose acetate filter prior to acid preservation. Samples for extractable analysis are acidified with 2 mL concentrated HNO3 per litre and allowed to stand a minimum of 24 hours before analysis. Clients providing these samples are advised to preserve samples at the collection site. | 0.3 | Y | Y | Y | 82 | 3/10/1998 2:07:00 PM | 2/25/1999 2:08:00 PM | 10/20/1999 2:21:47 AM | 1998-11-17 |  | NPEPPIN |  | Concentration | mg/L | Y | (fl. conc.) |  | Phosphorus Total Dissolved (fl. conc.) | 7723-14-0 | Metals - Dissolved; | Metals - Dissolved; EMS code: P--D | Chemical | Dissolved | LAB
P--D | mg/L | Phosphorus Total Dissolved | X562 | Dissolved metals: FF, FA, HNO3, HR-ICPMS | Ultra trace metals in water are analyzed by high resolution inductively coupled plasma mass spectrometry (HR-ICPMS), based on US EPA Method 200.8, (Oct. 1999). This procedure is intended for pristine field-filtered acid-preserved water samples. The detection limits (LOR) for this test are based on lab instrumental analysis only, not including filtration. ALS-supplied field filtration equipment does not support these LOR. Therefore, because of the high probability of false positives due to filtration, it is strongly recommended that a filtration blank be analyzed to aid in data interpretation. | 0.3 | Y | Y | Y | 16 | 5/30/2018 | 10/19/2023 | 11/9/2023 7:41:27 AM | 2018-06-06 | 6/6/2018 10:07:50 AM | RDGIBSON | RDGIBSON | Concentration | mg/L | Y | (fl. conc.) |  | Phosphorus Total Dissolved (fl. conc.) | 7723-14-0 | Metals - Dissolved; | Metals - Dissolved; EMS code: P--D | Chemical | Dissolved | LAB
P--D | mg/L | Phosphorus Total Dissolved | EQ01 | EQUIS (mg/L) | EQUIS (mg/L) | 0.003 | N | Y | Y | 11373 |  |  | 10/24/1996 3:16:58 PM | 1996-04-12 | 6/3/1998 6:51:37 PM | EMS |  | Concentration | mg/L | Y | (fl. conc.) |  | Phosphorus Total Dissolved (fl. conc.) | 7723-14-0 | Metals - Dissolved; | Metals - Dissolved; EMS code: P--D | Chemical | Dissolved | LAB
P--D | mg/L | Phosphorus Total Dissolved | METC | MET-D-NP-CCMS-VA | Ultra trace metals in water are analyzed by CRC ICPMS, based on US EPA Method 6020B (July 2014). This procedure is intended for pristine field-filtered acid-preserved water samples. The detection limits (LOR) for this test are based on lab instrumental analysis only, not including filtration. ALS-supplied field filtration equipment does not support these LOR. Therefore, because of the high probability of false positives due to filtration, it is strongly recommended that a filtration blank be analysed to aid in data interpretation. | 0.05 | Y | Y | Y | 559 | 11/13/2017 | 4/9/2024 | 4/29/2024 4:53:06 PM | 2021-06-09 |  | RDGIBSON |  | Concentration | mg/L | Y | (fl. conc.) |  | Phosphorus Total Dissolved (fl. conc.) | 7723-14-0 | Metals - Dissolved; | Metals - Dissolved; EMS code: P--D | Chemical | Dissolved | LAB
P--D | mg/L | Phosphorus Total Dissolved | 139A | Dig.Auto AscorbicA | Phosphorus present in the sample is converted to soluble ortho-phosphate by digesting the sample with persulphate in an acidic medium. The resulting ortho-phosphate then reacts with ammonium molybdate and antimony tartrate in an acidic medium to form an antimony-phophomolybdate complex, which is reduced to an intensely blue coloured complex by ascorbic acid. The absorbance of this complex is measured at 880 nm. | 0.002 | Y | Y | Y | 61521 | 1/9/1997 6:17:00 PM | 8/7/2024 5:01:00 PM | 8/26/2024 2:41:32 PM | 1996-06-20 | 6/3/1998 6:51:44 PM | MPAK |  | Concentration | mg/L | Y | (fl. conc.) |  | Phosphorus Total Dissolved (fl. conc.) | 7723-14-0 | Metals - Dissolved; | Metals - Dissolved; EMS code: P--D | Chemical | Dissolved | LAB
P--D | mg/L | Phosphorus Total Dissolved | X096 | Ascorbic Acid | Ascorbic Acid | 0.003 | Y | Y | Y | 2 | 6/19/2001 | 6/19/2001 | 7/13/2001 12:44:27 PM | 1996-04-12 | 6/3/1998 6:51:37 PM | EMS |  | Concentration | mg/L | Y | (fl. conc.) |  | Phosphorus Total Dissolved (fl. conc.) | 7723-14-0 | Metals - Dissolved; | Metals - Dissolved; EMS code: P--D | Chemical | Dissolved | LAB
P--D | mg/L | Phosphorus Total Dissolved | X125 | Dig,Stannous Cl2,Col'met | Dig;Stannous Cl2;Col'met | 0.003 | Y | Y | Y | 1983 | 3/17/1997 | 1/19/2006 10:15:00 AM | 2/6/2006 4:19:50 PM | 1996-04-12 | 6/3/1998 6:51:37 PM | EMS |  | Concentration | mg/L | Y | (fl. conc.) |  | Phosphorus Total Dissolved (fl. conc.) | 7723-14-0 | Metals - Dissolved; | Metals - Dissolved; EMS code: P--D | Chemical | Dissolved | LAB
P--D | mg/L | Phosphorus Total Dissolved | X170 | ICP | This code is a catch-all for a large variety of ICP methods including water and also solids digested into solutions as an intermediate result. Different procedures will have greatly different detection limits. (a) For air particulates, preliminary results are reported in mg/L and loading results as mg/dm2/d. (b) For Hi Vol samples preliminiary results are same as reported in mg/L and loading results as ug/m3. (c) Note that for Rain samples there is no filtration nor digestion steps. | 0.04 | Y | Y | Y | 8583 | 4/11/1996 6:02:00 PM | 9/23/2020 | 6/13/2024 3:58:21 PM | 1996-04-12 | 6/3/1998 6:51:37 PM | EMS |  | Concentration | mg/L | Y | (fl. conc.) |  | Phosphorus Total Dissolved (fl. conc.) | 7723-14-0 | Metals - Dissolved; | Metals - Dissolved; EMS code: P--D | Chemical | Dissolved | LAB
P--D | mg/L | Phosphorus Total Dissolved | F002 | Dig,Auto Asc.Acid,Col'm (whole bottle) | Phosphorus present in the sample is converted to soluble ortho-phosphate by digesting the sample with persulphate in an acidic medium. The resulting ortho-phosphate then reacts with ammonium molybdate and antimony tartrate in an acidic medium to form an antimony-phosphomolybdate complex, which is reduced to an intensely blue coloured complex by ascorbic acid. The absorbance of this complex is measured at 880 nm. | 0.002 | Y | Y | Y | 2738 | 8/21/1996 6:17:00 PM | 9/29/2000 | 10/5/2010 10:15:38 AM | 1996-06-20 | 6/3/1998 6:51:44 PM | MPAK |  | Concentration | mg/L | Y | (fl. conc.) |  | Phosphorus Total Dissolved (fl. conc.) | 7723-14-0 | Metals - Dissolved; | Metals - Dissolved; EMS code: P--D | Chemical | Dissolved | LAB
P--D | mg/L | Phosphorus Total Dissolved | X467 | H2SO4:ICP | H2SO4:ICP | 0.1 | Y | Y | Y | 59 | 9/26/1994 | 10/29/1996 | 3/20/2001 1:39:17 PM | 1996-10-13 | 6/3/1998 6:51:46 PM | EMS |  | Concentration | mg/L | Y | (fl. conc.) |  | Phosphorus Total Dissolved (fl. conc.) | 7723-14-0 | Metals - Dissolved; | Metals - Dissolved; EMS code: P--D | Chemical | Dissolved | LAB
P--E | mg/L | Phosphorus Extractable | F040 | Field Acid, lab filt, ICP-AES | Dissolved and extractable water samples are ready for analysis once collected. Total samples are autoclaved with HNO3 and hydrochloric acid (HCI) in capped Teflon vials. Analysis is performed via Inductively Coupled Argon Plasma Atomic Emission Spectrometry (ICP-AES)(simultaneous multielement analysis), Graphite Furnace Atomic Absorption Spectrometry (GF-AAS)(low level analysis) and/or Flame Atomic Absorption Spectrometry (F-AAS). | 0.1 | Y | Y | Y | 2371 | 8/5/1997 6:06:00 PM | 3/23/2012 3:11:00 AM | 9/11/2019 8:47:41 AM | 1997-05-08 | 6/3/1998 6:51:46 PM | MPAK |  | Concentration | mg/L | Y | (fl. conc.) |  | Phosphorus Extractable (fl. conc.) | 7723-14-0 | Nutrients; | Nutrients; EMS code: P--E | Chemical | Extractable | LAB
P--E | ug/g | Phosphorus Extractable | ASCE | Ammonium Acetate soil colloid exchange | Available metals - Ammonium Acetate soil colloid exchange sites displacement |  | Y | Y | Y | 14 | 10/31/2022 | 10/31/2023 | 11/23/2023 7:26:23 AM | 2007-10-04 |  | RDGIBSON |  | MassPerMass | mg/kg | Y | (mass conc.) |  | Phosphorus Extractable (mass conc.) | 7723-14-0 | Nutrients; | Nutrients; EMS code: P--E | Chemical | Extractable | LAB
P--L | mg/L | Phosphorus Leachable | MEND | MEND (mod) EPA 300.1 (mod) | MEND (mod) EPA 300.1 (mod). Inorganic anions are analyzed by obtaining an extract from a soil sample that has been added in a 3:1 ratio of deionized water to soil; then leached and the extract filtered prior to analysis; which is then analyzed by Ion Chromatography with conductivity and/or UV detection. | 0.3 | Y | Y | Y | 0 |  |  |  | 2024-01-29 |  | RDGIBSON |  | Concentration | mg/L | Y | (fl. conc.) |  | Phosphorus Leachable (fl. conc.) | 7723-14-0 | Nutrients; | Nutrients; EMS code: P--L | Chemical |  | LAB
P--T | mg/L | Phosphorus Total | EQ01 | EQUIS (mg/L) | EQUIS (mg/L) | 0.003 | N | Y | Y | 18381 |  |  | 10/24/1996 3:17:07 PM | 1996-04-12 | 6/3/1998 6:51:37 PM | EMS |  | Concentration | mg/L | Y | (fl. conc.) |  | Phosphorus Total (fl. conc.) | 7723-14-0 | Nutrients; | Nutrients; EMS code: P--T | Chemical | Total | LAB
P--T | ug/g | Phosphorus Total | 0192 | Hom, Nitric/Perchlor,ICP | Sample already in a dry state so does not require drying before homogenization and nitric perchlorid digestion followed by ICP analysis. | 50 | Y | Y | Y | 708 |  |  | 10/24/1996 12:50:25 PM | 1996-04-12 | 6/3/1998 6:51:37 PM | EMS |  | MassPerMass | mg/kg | Y | (mass conc.) |  | Phosphorus Total (mass conc.) | 7723-14-0 | Nutrients; | Nutrients; EMS code: P--T | Chemical | Total | LAB
P--T | mg/L | Phosphorus Total | 0042 | GI HNO3 Dig: ICP | Sample aliquot taken from unfiltered unpreserved General Ions bottle; HNO3 digestion; ICP analysis. | 0.04 | Y | Y | Y | 6268 | 9/20/1996 | 7/5/2019 3:21:00 PM | 7/12/2019 6:49:13 PM | 1996-04-12 | 6/3/1998 6:51:37 PM | EMS |  | Concentration | mg/L | Y | (fl. conc.) |  | Phosphorus Total (fl. conc.) | 7723-14-0 | Nutrients; | Nutrients; EMS code: P--T | Chemical | Total | LAB
P--T | ug/g | Phosphorus Total | MTID | MET-DRY-CCMS-VA | Metals in Tissue by CRC ICPMS (DRY) Methodology: This method is conducted following British Columbia Lab Manual method "Metals in Animal Tissue and Vegetation (Biota) - Prescriptive". Tissue samples are homogenized and sub-sampled prior to hotblock digestion with nitric and hydrochloric acids, in combination with addition of hydrogen peroxide. Instrumental analysis is by collision cell inductively coupled plasma - mass spectrometry (modified from EPA Method 6020A). Method Limitation: This method employs a strong acid/peroxide digestion, and is intended to provide a conservative estimate of bio-available metals. Near complete recoveries are achieved for most toxicologically important metals, but elements associated with recalcitrant minerals may be only partially recovered. | 10 | Y | Y | Y | 122 | 8/13/2014 | 9/13/2022 | 10/26/2022 7:14:15 AM | 2015-03-09 | 3/11/2015 4:12:05 PM | RDGIBSON | EMS | MassPerMass | mg/kg | Y | (mass conc.) |  | Phosphorus Total (mass conc.) | 7723-14-0 | Nutrients; | Nutrients; EMS code: P--T | Chemical | Total | LAB
P--T | % | Phosphorus Total | ICP06 | ALS Minerals ME-ICP06 | ALS Minerals ME-ICP06. A prepared sample is fused with a lithium tetraborate/lithium metaborate flux and poured into a graphite crucible and fused in a furnace at 1000°C. The resulting melt is then cooled and dissolved in 100 mL of 4% HNO3/2% HCl followed by ICP-AES. | 0.004 | Y | Y | Y | 0 |  |  |  | 2024-01-29 |  | EMS |  | DimensionlessRatio | % | N | (ratio) |  | Phosphorus Total (ratio) | 7723-14-0 | Nutrients; | Nutrients; EMS code: P--T | Chemical | Total | LAB
P--T | mg/L | Phosphorus Total | X170 | ICP | This code is a catch-all for a large variety of ICP methods including water and also solids digested into solutions as an intermediate result. Different procedures will have greatly different detection limits. (a) For air particulates, preliminary results are reported in mg/L and loading results as mg/dm2/d. (b) For Hi Vol samples preliminiary results are same as reported in mg/L and loading results as ug/m3. (c) Note that for Rain samples there is no filtration nor digestion steps. | 0.04 | Y | Y | Y | 7406 | 3/28/1996 6:03:00 PM | 11/7/2018 | 12/8/2023 4:30:26 PM | 1996-04-12 | 6/3/1998 6:51:37 PM | EMS |  | Concentration | mg/L | Y | (fl. conc.) |  | Phosphorus Total (fl. conc.) | 7723-14-0 | Nutrients; | Nutrients; EMS code: P--T | Chemical | Total | LAB
P--T | ug/g | Phosphorus Total | X551 | Metals in Biota (dry), total, acid digestion, ICP-SFMS | 0.1 gram of dried, homogenized biota is weighed in a microwave TFM vessel, and 8mL of Nitric Acid and 1mL of Hydrogen Peroxide are added. The sample is digested at 200 degrees C for 15 minutes in a closed vessel high pressure microwave oven. The sample is cooled, the volume adjusted to 100mL and introduced to an iductively coupled argon plasma-sector field mass spectrometer (ICP-SFMS) by pneumatic nebulization. Each respective element is measured at a specific mass to charge ratio, m/z value expressed in amu (atomic mass units), and at a specified instrument resolution. LR: Low Resolution ~300, MR: Mid Resolution ~4000 and HR: Hich Resolution ~10000. Ref: Method number 02-2705, National Laboratory for Environmental Testing, CCIW, Environment Canada, Burlington. Requested April 2003. | 10 | Y | Y | Y | 49 | 8/11/2015 | 10/30/2023 | 11/28/2023 2:21:33 PM | 2015-03-11 |  | RDGIBSON |  | MassPerMass | mg/kg | Y | (mass conc.) |  | Phosphorus Total (mass conc.) | 7723-14-0 | Nutrients; | Nutrients; EMS code: P--T | Chemical | Total | LAB
P--T | mg/L | Phosphorus Total | MSEA | Metals in Seawater (undigested) by (CRC) ICP-MS | This method describes the multi-elemental determination of trace elements by Collision/Reaction Cell (CRC) ICP-MS. The method measures ions produced by a radio-frequency inductively coupled plasma. Analyte species originating in a liquid are pneumatically nebulized and the resulting aerosol transported by argon gas into the plasma torch. The ions produced are entrained in the plasma gas and extracted, by means of a differentially pumped vacuum interface, into a mass spectrometer. The ions flow through a Collision/Reaction Cell (CRC) where depending on the mode of analysis either the ions passes through the cell unhindered (no gas - classical ICP-MS mode), or collisions/reactions chemistry takes place between the ions and a neutral gas such as hydrogen or helium. Any polyatomic ions which give interference on the analytes are removed. The ions are then sorted according to their mass-to-charge ratios by a quadrupole mass spectrometer. The ions are then sorted according to their mass-to-charge ratios by a quadrupole mass spectrometer. The ions transmitted through the quadrupole are quantified by a channel electron multiplier and the ion information processed by a data handling system. Although most polyatomic interference is removed, any additional interference must be assessed and valid corrections applied or the data flagged to indicate problems. Interference correction must include compensation for background ions contributed by the plasma gas, reagents, and constituents of the sample matrix. Instrument drift as well as suppressions or enhancements of instrument response caused by the sample matrix must be corrected for by the use of internal standards. | 0.01 | Y | Y | Y | 110 | 8/12/2009 12:08:00 AM | 4/29/2024 | 6/7/2024 10:03:16 AM | 2009-09-08 |  | RDGIBSON |  | Concentration | mg/L | Y | (fl. conc.) |  | Phosphorus Total (fl. conc.) | 7723-14-0 | Nutrients; | Nutrients; EMS code: P--T | Chemical | Total | LAB
P--T | mg/L | Phosphorus Total | X561 | Total metals: FA, HNO3, HR-ICPMS (No Prep)Prep) | Ultra trace metals in water are analyzed by high resolution inductively coupled plasma mass spectrometry (HR-ICPMS), modified from US EPA 200.8, (Revision 5.5). The detection limits provided can only be met for undigested samples. This procedure is intended for colorless, non-turbid, acid-preserved water samples (i.e. pristine water samples), having turbidity < 1 NTU and no odor. Where turbidity exceeds 1 NTU, and/or the sample is colored and has an odor, results may be biased low compared to true Total Metals concentrations. ALS recommends that turbidity analysis be requested on samples submitted for this test to aid with interpretation of results. | 0.002 | Y | Y | Y | 5 | 6/1/2016 | 10/5/2023 | 1/31/2024 2:37:41 PM | 2016-05-20 |  | RDGIBSON |  | Concentration | mg/L | Y | (fl. conc.) |  | Phosphorus Total (fl. conc.) | 7723-14-0 | Nutrients; | Nutrients; EMS code: P--T | Chemical | Total | LAB
P--T | mg/L | Phosphorus Total | X145 | Dig, ICP | Dig; ICP | 0.05 | Y | Y | Y | 414 | 5/14/2013 | 9/22/2016 | 6/13/2024 3:58:21 PM | 2015-03-11 |  | RDGIBSON |  | Concentration | mg/L | Y | (fl. conc.) |  | Phosphorus Total (fl. conc.) | 7723-14-0 | Nutrients; | Nutrients; EMS code: P--T | Chemical | Total | LAB
P--T | ppm (S) | Phosphorus Total | MS61 | ALS Minerals ME-MS61 | ALS Minerals ME-MS61. A prepared sample is digested with HNO3; HF; and HClO4 in a graphite heating block to near dryness. After cooling; the sample is boiled in diluted HCl. Bulk to volume with deionized water; mixed and analyzed by ICP-AES and ICP-MS for all elements. | 10 | Y | Y | Y | 0 |  |  |  | 2024-01-29 |  | EMS |  | DimensionlessRatio | ppm | N | (ratio) |  | Phosphorus Total (ratio) | 7723-14-0 | Nutrients; | Nutrients; EMS code: P--T | Chemical | Total | LAB
P--T | ug/g | Phosphorus Total | 0190 | D&G:N/P Dig:ICP Anal. | Sample is dried and grind-mixed before nitric perchloric digestion followed by ICP analysis. | 50 | Y | Y | Y | 854 |  |  | 10/24/1996 3:06:16 PM | 1996-04-12 | 6/3/1998 6:51:37 PM | EMS |  | MassPerMass | mg/kg | Y | (mass conc.) |  | Phosphorus Total (mass conc.) | 7723-14-0 | Nutrients; | Nutrients; EMS code: P--T | Chemical | Total | LAB
P--T | mg/L | Phosphorus Total | X307 | Persulf/H2SO4 Dig:Col'm | Persulf/H2SO4 Dig:Col'm | 0.003 | Y | Y | Y | 27479 | 9/3/1996 | 8/20/2024 3:00:00 PM | 8/25/2024 9:47:36 PM | 1996-04-12 | 6/3/1998 6:51:37 PM | EMS |  | Concentration | mg/L | Y | (fl. conc.) |  | Phosphorus Total (fl. conc.) | 7723-14-0 | Nutrients; | Nutrients; EMS code: P--T | Chemical | Total | LAB
P--T | mg/L | Phosphorus Total | F033 | Acid Dig,Centrifuge,ICPMS | For the determination of total recoverable elements in water or waste water, take 100 ml aliquot from a well mixed, acid preserved, sample containing not more than 0.25% (w/v) total solids and transfer to a 250 ml Griffin beaker. (If total solids are greater than 0.25% reduce the size of the aliquot by a proportionate amount). Add 1 ml of conc. nitric acid and 0.5 ml conc. hydrochloric acid. Heat on a hot plate until the volume has been reduced to 15-20ml, ensuring that the sample does not boil. Allow to cool and transfer to a 50 ml volumetric flask. Dilute to volume with ASTM type I water. Centrifuge the sample or allow to stand overnight to separate insoluble material. Prior analysis, pipette 20 ml into a 50 ml volumetric flask and dilute to volume with ASTM type I water. Analysis should be performed within seventy two hours of the completed preparation. Note that different laboratories may have very different field preservation, digestion procedures and detection limits for this method, so these should be documented. | 0.01 | Y | Y | Y | 16329 | 5/7/1998 | 8/22/2024 3:24:00 AM | 8/25/2024 9:47:36 PM | 1998-05-14 | 6/3/1998 6:51:49 PM | NPEPPIN |  | Concentration | mg/L | Y | (fl. conc.) |  | Phosphorus Total (fl. conc.) | 7723-14-0 | Nutrients; | Nutrients; EMS code: P--T | Chemical | Total | LAB
P--T | mg/L | Phosphorus Total | F121 | Acid,ICPMS | Acidified, no digestion, ICPMS |  | Y | Y | Y | 481 | 3/6/2013 1:42:00 PM | 8/15/2024 9:10:00 PM | 8/26/2024 8:06:03 AM | 2002-07-22 |  | UMULLER |  | Concentration | mg/L | Y | (fl. conc.) |  | Phosphorus Total (fl. conc.) | 7723-14-0 | Nutrients; | Nutrients; EMS code: P--T | Chemical | Total | LAB
P--T | mg/L | Phosphorus Total | X349 | FA:HNO3:ICP | Samples are field-acidified with HNO3; HNO3 digestion; analysis by ICP/AES. | 0.1 | Y | Y | Y | 30 |  |  | 10/30/2014 8:31:18 AM | 1998-05-22 | 6/3/1998 6:51:49 PM | NPEPPIN |  | Concentration | mg/L | Y | (fl. conc.) |  | Phosphorus Total (fl. conc.) | 7723-14-0 | Nutrients; | Nutrients; EMS code: P--T | Chemical | Total | LAB
P--T | mg/L | Phosphorus Total | X030 | Dig,Ascorbic,Col'metric | Dig;Ascorbic;Col'metric | 0.003 | Y | Y | Y | 4734 | 9/18/1996 | 7/27/2024 8:00:00 AM | 8/20/2024 8:39:43 AM | 1996-04-12 | 6/3/1998 6:51:37 PM | EMS |  | Concentration | mg/L | Y | (fl. conc.) |  | Phosphorus Total (fl. conc.) | 7723-14-0 | Nutrients; | Nutrients; EMS code: P--T | Chemical | Total | LAB
P--T | ug/g | Phosphorus Total | FS82 | PBM,dry60C,10mesh,SALM-ICP-MS | PBM;dry60C;10mesh;SALM-ICP-MS | 10 | Y | Y | Y | 3310 | 5/21/2002 | 7/24/2024 | 7/29/2024 5:17:09 PM | 2002-11-12 |  | UMULLER |  | MassPerMass | mg/kg | Y | (mass conc.) |  | Phosphorus Total (mass conc.) | 7723-14-0 | Nutrients; | Nutrients; EMS code: P--T | Chemical | Total | LAB
P--T | mg/L | Phosphorus Total | ULTR | Total Metals in Water by ICPMS (Ultra) | Total Metals in Water by ICPMS (Ultra), EPA SW-846 3005A/6020A, MET-TOT-ULTRA-MS-VA This analysis is carried out using procedures adapted from "Standard Methods for the Examination of Water and Wastewater" published by the American Public Health Association, and with procedures adapted from "Test Methods for Evaluating Solid Waste" SW-846 published by the United States Environmental Protection Agency (EPA). The procedures may involve preliminary sample treatment by acid digestion, using either hotblock or microwave oven, or filtration (EPA Method 3005A). Instrumental analysis is by inductively coupled plasma - mass spectrometry (EPA Method 6020A). | 0.05 | Y | Y | Y | 929 | 5/2/2016 | 9/21/2020 | 7/21/2021 6:01:33 PM | 2016-05-25 |  | RDGIBSON |  | Concentration | mg/L | Y | (fl. conc.) |  | Phosphorus Total (fl. conc.) | 7723-14-0 | Nutrients; | Nutrients; EMS code: P--T | Chemical | Total | LAB
P--T | ppm (S) | Phosphorus Total | MS41 | ALS Minerals ME-MS41 | ALS Minerals ME-MS41. A prepared sample is digested with aqua regia in a graphite heating block. After cooling; the resulting solution is diluted to with deionized water; mixed and analyzed by ICP-AES. Samples are then analyzed by ICP-MS for the remaining suite of elements. | 10 | Y | Y | Y | 0 |  |  |  | 2024-01-29 |  | EMS |  | DimensionlessRatio | ppm | N | (ratio) |  | Phosphorus Total (ratio) | 7723-14-0 | Nutrients; | Nutrients; EMS code: P--T | Chemical | Total | LAB
P--T | % | Phosphorus Total | 6041 | Aqua Regia by ICPMS | These were sediment/soils digested with Aqua-Regia (Nitric Acid & Hydrochloric Acids) and then analysis by ICPMS. Results would be reported in mg/kg and/or %. | 0.001 | Y | Y | Y | 18 | 10/13/2019 | 3/12/2020 | 3/24/2020 4:13:17 PM | 2019-09-06 |  | KATFRASE |  | DimensionlessRatio | % | N | (ratio) |  | Phosphorus Total (ratio) | 7723-14-0 | Nutrients; | Nutrients; EMS code: P--T | Chemical | Total | LAB
P--T | mg/L | Phosphorus Total | 1349 | Tot Metals in Seawater by CRC ICPMS (BC) | Tot Metals in Seawater by CRC ICPMS (BC) | 0.05 | Y | Y | Y | 98 | 8/16/2019 | 12/5/2021 | 12/13/2021 12:34:16 PM | 2019-09-30 |  | KATFRASE |  | Concentration | mg/L | Y | (fl. conc.) |  | Phosphorus Total (fl. conc.) | 7723-14-0 | Nutrients; | Nutrients; EMS code: P--T | Chemical | Total | LAB
P--T | mg/L | Phosphorus Total | X285 | HNO3 Dig:ICP Analysis | HNO3 Dig:ICP Analysis | 0.03 | Y | Y | Y | 634 | 8/10/1993 | 9/12/2003 3:16:00 PM | 4/30/2004 11:32:38 AM | 1996-10-13 | 6/3/1998 6:51:46 PM | EMS |  | Concentration | mg/L | Y | (fl. conc.) |  | Phosphorus Total (fl. conc.) | 7723-14-0 | Nutrients; | Nutrients; EMS code: P--T | Chemical | Total | LAB
P--T | ug/g wet | Phosphorus Total | MTIW | MET-WET-CCMS-MID-VA | Metals in Tissue by CRC ICPMS (WET) Methodology: This method is conducted following British Columbia Lab Manual method "Metals in Animal Tissue and Vegetation (Biota) - Prescriptive". Tissue samples are homogenized and sub-sampled prior to hotblock digestion with nitric and hydrochloric acids, in combination with addition of hydrogen peroxide. Instrumental analysis is by collision cell inductively coupled plasma - mass spectrometry (modified from EPA Method 6020A). Method Limitation: This method employs a strong acid/peroxide digestion, and is intended to provide a conservative estimate of bio-available metals. Near complete recoveries are achieved for most toxicologically important metals, but elements associated with recalcitrant minerals may be only partially recovered. | 2 | Y | Y | Y | 24 | 11/25/2018 | 12/6/2018 | 7/21/2021 2:55:08 PM | 2018-12-24 |  | EMS |  | MassPerMass | µg/g | Y | (mass conc.) | (wet) | Phosphorus Total (wet) (mass conc.) | 7723-14-0 | Nutrients; | Nutrients; EMS code: P--T | Chemical | Total | LAB
P--T | ug/g | Phosphorus Total | X230 | Hom,HNO3/HCl,ICP | The sample is dried, ground to pass 100 mesh, and digested with a nitric-hydrochloric acid mixture to solubilize the solid matter and to remove the organic material by oxidation and volatilization. The completed digests are submitted to metals instrumentation personnel for metals analysis by Inductively Coupled Atomic Emission Spectrophotometry (ICP-AES), or hydride analysis by atomic absorption. | 10 | Y | Y | Y | 41 | 2/12/2002 | 12/30/2005 | 2/3/2006 1:13:13 AM | 1998-11-05 |  | NPEPPIN |  | MassPerMass | mg/kg | Y | (mass conc.) |  | Phosphorus Total (mass conc.) | 7723-14-0 | Nutrients; | Nutrients; EMS code: P--T | Chemical | Total | LAB
P--T | ppm (S) | Phosphorus Total | 6042 | Aqua Regia by ICPMS | These were sediment/soils digested with Aqua-Regia (Nitric Acid & Hydrochloric Acids) and then analysis by ICPMS. Results would be reported in ppm s. | 0.001 | Y | Y | Y | 71 | 5/14/2022 | 7/27/2024 | 7/29/2024 9:13:08 AM | 2022-06-01 | 12/14/2022 11:46:30 AM | EMS | EMS | DimensionlessRatio | ppm | N | (ratio) |  | Phosphorus Total (ratio) | 7723-14-0 | Nutrients; | Nutrients; EMS code: P--T | Chemical | Total | LAB
P--T | mg/L | Phosphorus Total | METN | MET-T-NP-CCMS-VA | Ultra trace metals in water are analyzed by CRC ICPMS, based on US EPA Method 6020B (July 2014). The detection limits provided can only be met for undigested samples. This procedure is intended for colorless, non-turbid, acid-preserved water samples (i.e. pristine water samples), having turbidity < 1 NTU and no odor. Where turbidity exceeds 1 NTU, and/or the sample is colored and has an odor, results may be biased low compared to true Total Metals concentrations. ALS recommends that turbidity analysis be requested on samples submitted for this test to aid with interpretation of results. Where turbidity is <1NTU, undigested metals are equivalent to total metals concentrations. | 0.05 | Y | Y | Y | 439 | 11/14/2017 | 4/4/2024 | 4/29/2024 4:53:06 PM | 2021-06-09 |  | RDGIBSON |  | Concentration | mg/L | Y | (fl. conc.) |  | Phosphorus Total (fl. conc.) | 7723-14-0 | Nutrients; | Nutrients; EMS code: P--T | Chemical | Total | LAB
P--T | mg/L | Phosphorus Total | X025 | Acid Dig,Colorimetric | Acid Dig;Colorimetric | 0.003 | Y | Y | Y | 471 |  |  | 10/2/2007 10:05:03 AM | 1996-04-12 | 6/3/1998 6:51:37 PM | EMS |  | Concentration | mg/L | Y | (fl. conc.) |  | Phosphorus Total (fl. conc.) | 7723-14-0 | Nutrients; | Nutrients; EMS code: P--T | Chemical | Total | LAB
P--T | ug/g | Phosphorus Total | FS89 | Homo,HNO3 dig,ICP-AES | Homo;HNO3 dig;ICP-AES | 20 | Y | Y | Y | 4 | 1/5/2001 | 8/20/2001 12:00:00 PM | 10/23/2002 1:42:50 PM | 2001-03-08 |  | NPEPPIN |  | MassPerMass | mg/kg | Y | (mass conc.) |  | Phosphorus Total (mass conc.) | 7723-14-0 | Nutrients; | Nutrients; EMS code: P--T | Chemical | Total | LAB
P--T | ug/m3 | Phosphorus Total | 5312 | Tef:N/P Dig:ICP Analysis | A nitric/perchloric acid digestion if used to bring the metals into solution prior to ICP analysis. For Hi-Vol samples preliminary results are reported as mg/L and 24 hour loading results as ug/m3. | 0.02 | Y | Y | Y | 1221 | 11/7/1997 | 5/15/2000 | 11/18/2000 1:29:36 AM | 1996-10-13 | 6/3/1998 6:51:46 PM | EMS |  | AirConcentration | µg/m³ | Y | (air conc.) |  | Phosphorus Total (air conc.) | 7723-14-0 | Nutrients; | Nutrients; EMS code: P--T | Chemical | Total | LAB
P--T | mg/L | Phosphorus Total | 2710 | Froz.Dig.Aut.Ascorbic Ac | Froz.Dig.Aut.Ascorbic Ac | 0.003 | Y | Y | Y | 3247 |  |  | 10/24/1996 3:06:02 PM | 1996-04-12 | 6/3/1998 6:51:37 PM | EMS |  | Concentration | mg/L | Y | (fl. conc.) |  | Phosphorus Total (fl. conc.) | 7723-14-0 | Nutrients; | Nutrients; EMS code: P--T | Chemical | Total | LAB
P--T | mg/L | Phosphorus Total | F082 | Aliquot,HNO3/HCL dig,ICPMS | Aliquot microwave digested with 3:1 HNO3/HCL. Analysis by ICPMS. For HiVol samples preliminary result are reported in mg/L and loading results are reported in ug/m3 | 0.3 | Y | Y | Y | 22016 | 1/24/2011 | 8/23/2024 | 8/26/2024 4:53:05 PM | 2015-04-15 |  | RDGIBSON |  | Concentration | mg/L | Y | (fl. conc.) |  | Phosphorus Total (fl. conc.) | 7723-14-0 | Nutrients; | Nutrients; EMS code: P--T | Chemical | Total | LAB
P--T | ug/g | Phosphorus Total | EQ02 | EQUIS (ug/g) | EQUIS (ug/g) | 50 | N | Y | Y | 797 |  |  | 10/23/1996 12:08:28 AM | 1996-04-12 | 6/3/1998 6:51:37 PM | EMS |  | MassPerMass | mg/kg | Y | (mass conc.) |  | Phosphorus Total (mass conc.) | 7723-14-0 | Nutrients; | Nutrients; EMS code: P--T | Chemical | Total | LAB
P--T | mg/L | Phosphorus Total | MWID | Metals in Water by ICPMS (digested) | Metals in Water by ICPMS (digested) Based on EPA 200.8 | 0.04 | Y | Y | Y | 2308 |  |  | 7/31/2024 1:46:41 PM | 2009-12-24 |  | RDGIBSON |  | Concentration | mg/L | Y | (fl. conc.) |  | Phosphorus Total (fl. conc.) | 7723-14-0 | Nutrients; | Nutrients; EMS code: P--T | Chemical | Total | LAB
P--T | ug/g | Phosphorus Total | FB04 | Freeze Dry,Hom,Acid Dig,ICP | Sample moisture is determined by weighing samples before and after freeze-drying. Freeze-dried samples are ground in a Braun blender before being weighed into Teflon PFA liners and digested with concentrated HNO3, concentrated HCl, and 50% H2O2 in a microwave oven; which results in a complete sample dissolution. Samples are collected and volumized with Type I water. Analysis is performed via Inductively Coupled Argon Plasma Atomic Emission Spectrometry (ICP-AES) (simultaneous multielement analysis), Graphite Furnace Atomic Absorption Spectrometry (GF-AAS) (low level Cd and Pb analysis), Cold Vapor Atomic Fluorescence Spectrometry (CV-AFS) (Mercury analysis) and/or Flame Atomic Absorption Spectrometry (F-AAS) (K analysis). | 8 | Y | Y | Y | 241 | 8/22/1996 6:11:00 PM | 2/1/2002 2:30:00 PM | 6/23/2003 1:15:32 PM | 1997-02-27 | 6/3/1998 6:51:46 PM | MPAK |  | MassPerMass | mg/kg | Y | (mass conc.) |  | Phosphorus Total (mass conc.) | 7723-14-0 | Nutrients; | Nutrients; EMS code: P--T | Chemical | Total | LAB
P--T | mg/L | Phosphorus Total | X185 | Dig. Auto.Ascorbic Acid | The unfiltered sample is acidified, potassium persulfate is added and the mixture is digested at elevated temperature and pressure in a steam autoclave. After digestion, the sample is reacted with a mixture of ammonium molybdate and potassium antimonyl tartrate in acid solution. Ascorbic acid is then added to produce a blue coloured product with an absorbance maximum at 880 nm. The absorbance of the solution is measured and the phosphorus concentration is determined by comparison with standards treated in the same manner. Automated procedure. This unfiltered method is sometimes used for a clear turbid-free solution where filtering is not considered necessary. | 0.003 | Y | Y | Y | 28175 | 1/30/2002 1:00:00 PM | 3/13/2015 | 7/29/2024 12:59:01 PM | 1996-04-12 | 6/3/1998 6:51:37 PM | EMS |  | Concentration | mg/L | Y | (fl. conc.) |  | Phosphorus Total (fl. conc.) | 7723-14-0 | Nutrients; | Nutrients; EMS code: P--T | Chemical | Total | LAB
P--T | mg/L | Phosphorus Total | X023 | Dig Manual Ascorbic Acid | Dig Manual Ascorbic Acid | 0.003 | Y | Y | Y | 37988 |  |  | 8/20/2024 7:35:41 AM | 1996-04-12 | 6/3/1998 6:51:45 PM | EMS |  | Concentration | mg/L | Y | (fl. conc.) |  | Phosphorus Total (fl. conc.) | 7723-14-0 | Nutrients; | Nutrients; EMS code: P--T | Chemical | Total | LAB
P--T | mg/L | Phosphorus Total | X257 | Colorimetric | Colorimetric. General use for several parameters. E.g.Total Sulfide is determined by spectrophotometer using the methlyene blue colourimetric method. The H2S calculation converts Total Sulphide as (S2-) and reports it as Total Sulphide as (H2S). | 0.003 | Y | Y | Y | 20395 | 5/5/2015 | 8/24/2024 | 8/26/2024 3:51:07 PM | 1996-04-12 | 6/3/1998 6:51:46 PM | EMS |  | Concentration | mg/L | Y | (fl. conc.) |  | Phosphorus Total (fl. conc.) | 7723-14-0 | Nutrients; | Nutrients; EMS code: P--T | Chemical | Total | LAB
P--T | ug/g | Phosphorus Total | X247 | Dig:Ascorbic:Col'metric | Dig:Ascorbic:Col'metric | 50 | Y | Y | Y | 16 |  |  | 10/24/1996 12:04:29 PM | 1996-04-12 | 6/3/1998 6:51:37 PM | EMS |  | MassPerMass | mg/kg | Y | (mass conc.) |  | Phosphorus Total (mass conc.) | 7723-14-0 | Nutrients; | Nutrients; EMS code: P--T | Chemical | Total | LAB
P--T | ug/g wet | Phosphorus Total | BPSY | MET-WET-BPSY-CCMS-VA | Metals in Animal Tissue and Vegetation (Biota) - Prescriptive. EPA 200.3/6020B This method is conducted following British Columbia Lab Manual method "Metals in Animal Tissue and Vegetation (Biota) - Prescriptive". Tissue samples are digested with nitric and hydrochloric acids; in combination with addition of hydrogen peroxide. Instrumental analysis is by collision cell inductively coupled plasma - mass spectrometry (modified from EPA Method 6020B). | 4 | Y | Y | Y | 2 | 11/7/2019 | 11/7/2019 | 6/10/2021 1:57:10 PM | 2021-06-04 |  | EMS |  | MassPerMass | µg/g | Y | (mass conc.) | (wet) | Phosphorus Total (wet) (mass conc.) | 7723-14-0 | Nutrients; | Nutrients; EMS code: P--T | Chemical | Total | LAB
P--T | mg/L | Phosphorus Total | X238 | K2S2O8/H2SO4Dig,Col'met | K2S2O8/H2SO4Dig;Col'met | 0.003 | Y | Y | Y | 316 | 1/26/2001 | 1/19/2006 10:15:00 AM | 2/6/2006 4:19:50 PM | 1996-04-12 | 6/3/1998 6:51:45 PM | EMS |  | Concentration | mg/L | Y | (fl. conc.) |  | Phosphorus Total (fl. conc.) | 7723-14-0 | Nutrients; | Nutrients; EMS code: P--T | Chemical | Total | LAB
P--T | mg/dm2/d | Phosphorus Total | DUST | HNO3 dig.(microwave or hot block),ICPMS | HNO3 dig.(microwave or hot block);ICPMS |  | Y | Y | Y | 121 | 2/4/2020 | 7/31/2024 | 7/31/2024 11:23:12 AM | 2022-09-01 |  | EMS |  | MassFlux | mg/dm²/d | Y | (mass flux) |  | Phosphorus Total (mass flux) | 7723-14-0 | Nutrients; | Nutrients; EMS code: P--T | Chemical | Total | LAB
P--T | ug/g | Phosphorus Total | 0031 | ICP | ICP | 50 | Y | Y | Y | 61 | 10/31/2022 | 10/31/2023 | 11/23/2023 7:26:23 AM | 1996-04-12 | 6/3/1998 6:51:42 PM | EMS |  | MassPerMass | mg/kg | Y | (mass conc.) |  | Phosphorus Total (mass conc.) | 7723-14-0 | Nutrients; | Nutrients; EMS code: P--T | Chemical | Total | LAB
P--T | mg/L | Phosphorus Total | 0040 | HNO3 Dig: ICP Analysis | HNO3 Dig: ICP Analysis | 0.04 | Y | Y | Y | 974 | 10/6/2010 10:03:00 AM | 7/25/2012 7:11:00 AM | 3/13/2023 12:19:57 PM | 1996-04-12 | 6/3/1998 6:51:37 PM | EMS |  | Concentration | mg/L | Y | (fl. conc.) |  | Phosphorus Total (fl. conc.) | 7723-14-0 | Nutrients; | Nutrients; EMS code: P--T | Chemical | Total | LAB
P--T | ug/g | Phosphorus Total | ICPM | Metals in Soil by CRC ICPMS | This analysis is carried out using procedures from CSR Analytical Method: "Strong Acid Leachable Metals (SALM) in Soil", BC Ministry of Environment, 26 June 2009, and procedures adapted from EPA Method 200.2. The sample is manually homogenized, dried at 60 degrees Celsius, sieved through a 63 um (230 mesh) sieve, and a representative subsample of the dry material is weighed. The sample is then digested at 95 degrees Celsius for 2 hours by block digester using concentrated nitric and hydrochloric acids. Instrumental analysis is by collision cell inductively coupled plasma - mass spectrometry (modifed from EPA Method 6020A). Method Limitation: This method is not a total digestion technique. It is a very strong acid digestion that is intended to dissolve those metals that may be environmentally available. By design, elements bound in silicate structures are not normally dissolved by this procedure as they are not usually mobile in the environment. | 50 | Y | Y | Y | 103 | 9/19/2017 | 9/24/2023 | 1/12/2024 11:16:52 AM | 2018-01-09 |  | JKROGH |  | MassPerMass | mg/kg | Y | (mass conc.) |  | Phosphorus Total (mass conc.) | 7723-14-0 | Nutrients; | Nutrients; EMS code: P--T | Chemical | Total | LAB
P--T | mg/L | Phosphorus Total | F002 | Dig,Auto Asc.Acid,Col'm (whole bottle) | Phosphorus present in the sample is converted to soluble ortho-phosphate by digesting the sample with persulphate in an acidic medium. The resulting ortho-phosphate then reacts with ammonium molybdate and antimony tartrate in an acidic medium to form an antimony-phosphomolybdate complex, which is reduced to an intensely blue coloured complex by ascorbic acid. The absorbance of this complex is measured at 880 nm. | 0.002 | Y | Y | Y | 5249 | 8/14/1996 6:18:00 PM | 3/30/1998 6:16:00 PM | 11/24/2005 1:31:18 PM | 1996-06-20 | 6/3/1998 6:51:44 PM | MPAK |  | Concentration | mg/L | Y | (fl. conc.) |  | Phosphorus Total (fl. conc.) | 7723-14-0 | Nutrients; | Nutrients; EMS code: P--T | Chemical | Total | LAB
P--T | mg/L | Phosphorus Total | F128 | Dig,Auto Asc.Acid,Col'm | Phosphorus present in the sample is converted to soluble ortho-phosphate by digesting the sample with persulphate in an acidic medium. The resulting ortho-phosphate then reacts with ammonium molybdate and antimony tartrate in an acidic medium to form an antimony-phosphomolybdate complex, which is reduced to an intensely blue coloured complex by ascorbic acid. The absorbance of this complex is measured at 880 nm. | 0.002 | Y | Y | Y | 16903 | 10/16/2006 | 4/12/2019 5:16:00 PM | 10/16/2019 9:08:55 AM | 2008-12-10 | 12/10/2008 9:40:44 AM | RDGIBSON | RDGIBSON | Concentration | mg/L | Y | (fl. conc.) |  | Phosphorus Total (fl. conc.) | 7723-14-0 | Nutrients; | Nutrients; EMS code: P--T | Chemical | Total | LAB
P--T | mg/L | Phosphorus Total | F038 | HNO3/HCl dig,ICP-AES | Dissolved and extractable water samples are ready for analysis once collected. Total samples are autoclaved with HNO3 and hydrochloric acid (HCI) in capped Teflon vials. Analysis is performed via Inductively Coupled Argon Plasma Atomic Emission Spectrometry (ICP-AES)(simultaneous multielement analysis), Graphite Furnace Atomic Absorption Spectrometry (GF-AAS) (low level analysis) and/or Flame Atomic Absorption Spectrometry (F-AAS). | 0.1 | Y | Y | Y | 7041 | 8/1/1996 6:04:00 PM | 12/1/2010 | 9/11/2019 8:47:42 AM | 1997-05-08 | 6/3/1998 6:51:46 PM | MPAK |  | Concentration | mg/L | Y | (fl. conc.) |  | Phosphorus Total (fl. conc.) | 7723-14-0 | Nutrients; | Nutrients; EMS code: P--T | Chemical | Total | LAB
P--T | ug/g | Phosphorus Total | PB02 | PBM,HNO3/HCl dig,dry60C,100mesh, | Performance Based Method. Samples are dried at 60 C and sieved using a .149 mm (100 mesh) sieve. Solid samples are digested with the combination of nitric and hydrochloric acids with a 1:1 ratio of acids. For environmental purposes, strong acid leaches are used routinely since they have the advantage of speed, low cost, and superio reproducibility. All metals trapped within the refractory matrix are unlikely to have environmental consequences. Instrumental analysis of extracts produced using this digestion method can be performed by a variety of techniques. |  | Y | Y | Y | 10 | 9/13/2001 | 9/13/2001 | 11/6/2002 3:52:38 PM | 2001-07-18 |  | NPEPPIN |  | MassPerMass | mg/kg | Y | (mass conc.) |  | Phosphorus Total (mass conc.) | 7723-14-0 | Nutrients; | Nutrients; EMS code: P--T | Chemical | Total | LAB
P--T | ug/g wet | Phosphorus Total | X560 | Metals in Biota (wet), total, acid digestion, ICP-MS | Five grams of wet, homogenized biota is weighed in a centrifuge tube, and 2.5 mL of Nitric Acid and 2.5 mL of Hydrochloric acid are added. The sample is digested at room temperature for one hour and then digested at 90-95 degrees C in a water bath or heat block for one more hour. The sample is cooled and the volume is made to 50 mL with DI. The digestate is then diluted at ten fold and introduced to a collision/reaction cell inductively coupled argon plasma mass spectrometer (CRC ICPMS) by pneumatic nebulization. Each respective element is measured at a specific mass to charge ratio, m/z value expressed in amu (atomic mass units). - Maxxam Analytics (BC). Added to EMS Dec 10, 2008. | 10 | Y | Y | Y | 541 | 11/15/2008 4:00:00 AM | 10/30/2023 | 11/28/2023 2:21:33 PM | 2008-12-11 | 3/11/2015 4:17:10 PM | RDGIBSON | EMS | MassPerMass | µg/g | Y | (mass conc.) | (wet) | Phosphorus Total (wet) (mass conc.) | 7723-14-0 | Nutrients; | Nutrients; EMS code: P--T | Chemical | Total | LAB
P--T | ug/g | Phosphorus Total | FS01 | D&G,Nitric/HCl:ICP | Samples are dried at 60 C, then either sieved to -100mesh (150 um) and rolled to homogenize, or finely ground with a mortar and pestle or a shatter box laboratory mill and rolled (depending on the submitter's requirements). The samples are then weighed into Teflon digestion liners and digested with HNO3 and HCl in a microwave oven. Samples are cooled, volumized, and allowed to settle overnight before being decanted. Analysis is then performed via Inductively Coupled Argon Plasma Atomic Emission Spectrometry (ICP-AES) (Simultaneous multi element analysis). In ICP Spectrometry, the liquid sample is pumped to the nebulizer, where it is convertd into an aerosol. the nebulizer carrier gas flow than transports the wet aerosol past the impact bead, where the large droplets condense out and flow to the waste container. The remaining fine aerosol is carried into the plasma. There, the high temperature desolvates the aerosol and dissociates the sample material into individual atoms and ions, which are excited to emit light at wavelengths characteristic of the elements present. A diffraction grating in the spectrometer diffracts the various wavelengths of light through exit slits to the photo-multiplier tubes, which convert the light energy to electronic signals. | 20 | Y | Y | Y | 2650 | 8/30/1996 1:52:00 PM | 10/10/2007 | 5/24/2008 1:15:03 AM | 1996-11-20 | 6/3/1998 6:51:46 PM | MPAK |  | MassPerMass | mg/kg | Y | (mass conc.) |  | Phosphorus Total (mass conc.) | 7723-14-0 | Nutrients; | Nutrients; EMS code: P--T | Chemical | Total | LAB
P--T | mg/L | Phosphorus Total | X125 | Dig,Stannous Cl2,Col'met | Dig;Stannous Cl2;Col'met | 0.003 | Y | Y | Y | 2025 | 5/5/1998 | 10/5/2001 | 9/19/2008 10:48:04 AM | 1996-04-12 | 6/3/1998 6:51:37 PM | EMS |  | Concentration | mg/L | Y | (fl. conc.) |  | Phosphorus Total (fl. conc.) | 7723-14-0 | Nutrients; | Nutrients; EMS code: P--T | Chemical | Total | LAB
P--T | mg/L | Phosphorus Total | AMMO | Dig.Auto AmmMol/SNCl2/H2SO4/K2S208 | COLOURIMETRY ON AN AUTOANALYZER WITH AMMONIUM MOLYBDATE AND SNCL2. H2SO4 AND K2S2O8 SOLNS. ARE ADDED TO AN ALIQUOT OF THE SHAKEN SAMPLE, WHICH IS THEN BOILED 90 MIN., THE VOL. OF THE ALIQUOT IS MAINTAINED. THEN, IF TURBID, THE ALIQUOT IS PASSED THROUGH A 0.45 U MEMBRANE FILTER. THE FILTRATE IS MIXED WITH A PREMIXED SOLN. OF (NH4)6MO7O24 AND SNCL2. THE RESULTING MOLYBDENUM BLUE COLOUR IS MEASURED SPECTROPHOTO-METRICALLY AT 660 MU AND COMPARED WITH THOSE OF IDENTICALLY PREPD. STD. PO4 ION SOLNS. AND REAGENT BLANKS. INTERFEREN-CES: HG CONCN. OF 1 MG/L, AS. REF: ANALYTICAL METHODS MANUAL, WATER QUALITY BRANCH, ENVIRONMENT CANADA, OTTAWA, 1974 | 0.0005 | Y | Y | Y | 11195 | 3/27/2003 | 12/9/2021 | 2/14/2023 10:17:03 AM | 2008-06-05 | 6/11/2008 10:16:20 AM | RDGIBSON | RDGIBSON | Concentration | mg/L | Y | (fl. conc.) |  | Phosphorus Total (fl. conc.) | 7723-14-0 | Nutrients; | Nutrients; EMS code: P--T | Chemical | Total | LAB
P--T | mg/L | Phosphorus Total | 139A | Dig.Auto AscorbicA | Phosphorus present in the sample is converted to soluble ortho-phosphate by digesting the sample with persulphate in an acidic medium. The resulting ortho-phosphate then reacts with ammonium molybdate and antimony tartrate in an acidic medium to form an antimony-phophomolybdate complex, which is reduced to an intensely blue coloured complex by ascorbic acid. The absorbance of this complex is measured at 880 nm. | 0.002 | Y | Y | Y | 92492 | 1/9/1997 6:17:00 PM | 8/16/2024 2:37:00 PM | 8/26/2024 2:41:32 PM | 1996-04-12 | 6/3/1998 6:51:37 PM | EMS |  | Concentration | mg/L | Y | (fl. conc.) |  | Phosphorus Total (fl. conc.) | 7723-14-0 | Nutrients; | Nutrients; EMS code: P--T | Chemical | Total | LAB
P--T | t/d | Phosphorus Total | CALC | Calculation (t/d) | Calculation (t/d) | 0 | Y | Y | Y | 930 |  |  | 2/24/2012 12:45:46 PM | 1996-04-12 | 6/3/1998 6:51:37 PM | EMS |  | MassFlow | t/d | Y | (mass flow) |  | Phosphorus Total (mass flow) | 7723-14-0 | Nutrients; | Nutrients; EMS code: P--T | Chemical | Total | LAB
P--T | mg/L | Phosphorus Total | F048 | Acid dig,conc'n,nebul.,ICP-OES | This method provides an automated procedure for the Inductively Coupled Argon Plasma Optical Emission Spectrophotometric (ICP-OES) determination and quantitation of trace amounts of metals in surface, ground and precipitation waters after a manual acid digestion and concentration procedure and ultrasonic nebulisation. Organically-bound metals and metal complexes are converted to metal ions by digestion with nitric acid [method details says HCl]. To increase the signal detected by the ICP, the samples are concentrated 10 times. The liquid sample is pumped onto the piezoelectric transducer, where the ultrasonic action efficiently converts it to a fine, dense aerosol. The nebulizer carrier gas flow then transports the wet aerosol through a heated U-tube, where the solvent is vaporized. Subsequently, most of the solvent vapour is removed from the aerosol stream by passing the stream through a chilled condenser. The result is a dry, analyte-laden aerosol which is introduced to the plasma. There, the high temperature desolvates the aerosol and dissociates the sample material into individual atoms and ions, which are excited to emit light at wave lengths characteristic of the elements present. Dissolved samples are field filtered through 0.45 micron cellulose acetate filter prior to acid preservation. Samples for extractable analysis are acidified with 2 mL concentrated HNO3 per litre and allowed to stand a minimum of 24 hours before analysis. Clients providing these samples are advised to preserve samples at the collection site. | 0.3 | Y | Y | Y | 172 | 11/20/1997 9:32:00 AM | 2/25/1999 1:48:00 PM | 7/6/2001 11:10:13 AM | 1998-11-13 |  | NPEPPIN |  | Concentration | mg/L | Y | (fl. conc.) |  | Phosphorus Total (fl. conc.) | 7723-14-0 | Nutrients; | Nutrients; EMS code: P--T | Chemical | Total | LAB
P--T | ug/g | Phosphorus Total | 0196 | D&G: LTA/HCl: ICP | D&G: LTA/HCl: ICP | 50 | Y | Y | Y | 32 |  |  | 10/24/1996 2:08:13 PM | 1996-04-12 | 6/3/1998 6:51:37 PM | EMS |  | MassPerMass | mg/kg | Y | (mass conc.) |  | Phosphorus Total (mass conc.) | 7723-14-0 | Nutrients; | Nutrients; EMS code: P--T | Chemical | Total | LAB
P--T | ug/g wet | Phosphorus Total | X552 | Metals in Biota (wet), total, acid digestion, ICP-SFMS | One gram of wet, homogenized biota is weighed in a microwave TFM vessel, and 8mL of Nitric Acid and 1mL of Hydrogen Peroxide are added. The sample is digested at 200 degrees C for 15 minutes in a closed vessel high pressure microwave oven. The sample is cooled, the volume adjusted to 100mL and introduced to an inductively coupled argon plasma-sector field mass spectrometer (ICP-SFMS) by pneumatic nebulization. Each respective element is measured at a specific mass to charge ratio, m/z value expressed in amu (atomic mass units), and at a specified instrument resolution. LR: Low Resolution ~300, MR: Mid Resolution ~4000. Ref: 02-2704, National Laboratory for Environmental Testing, CCIW, Environment Canada, Burlington. Requested April 2003. | 2 | Y | Y | Y | 37 | 8/9/2016 | 10/16/2023 | 11/28/2023 2:21:33 PM | 2015-03-11 |  | RDGIBSON |  | MassPerMass | µg/g | Y | (mass conc.) | (wet) | Phosphorus Total (wet) (mass conc.) | 7723-14-0 | Nutrients; | Nutrients; EMS code: P--T | Chemical | Total | LAB
P--T | mg/L | Phosphorus Total | X467 | H2SO4:ICP | H2SO4:ICP | 0.1 | Y | Y | Y | 114 | 9/26/1994 | 10/29/1996 | 10/3/2005 9:51:09 AM | 1996-10-13 | 6/3/1998 6:51:46 PM | EMS |  | Concentration | mg/L | Y | (fl. conc.) |  | Phosphorus Total (fl. conc.) | 7723-14-0 | Nutrients; | Nutrients; EMS code: P--T | Chemical | Total | LAB
P001 | mg/L | Parathion | ZN01 | DCM+Ac.Extr:GC/MS | DCM+Ac.Extr:GC/MS | 0.0005 | Y | Y | Y | 52 | 10/3/2002 | 3/29/2012 10:15:00 PM | 5/3/2012 9:27:51 AM | 1996-04-12 | 11/16/1999 4:06:27 PM | EMS | NPEPPIN | Concentration | mg/L | Y | (fl. conc.) |  | Parathion (fl. conc.) | 56-38-2 | Pesticides; | Pesticides; EMS code: P001 | Chemical |  | LAB
P001 | mg/L | Parathion | EQ01 | EQUIS (mg/L) | EQUIS (mg/L) | 0.001 | N | Y | Y | 2 |  |  | 10/22/1996 5:23:43 PM | 1996-04-12 | 6/3/1998 6:51:37 PM | EMS |  | Concentration | mg/L | Y | (fl. conc.) |  | Parathion (fl. conc.) | 56-38-2 | Pesticides; | Pesticides; EMS code: P001 | Chemical |  | LAB
P001 | mg/L | Parathion | E660 | EPA 8270E (mod) | EPA 8270E (mod) Miscellaneous Pesticides by GC-MS | 0.0001 | Y | Y | Y | 68 | 12/7/2021 | 9/12/2022 | 12/8/2022 12:57:05 PM | 2021-12-09 | 12/9/2021 3:22:12 PM | EMS | EMS | Concentration | mg/L | Y | (fl. conc.) |  | Parathion (fl. conc.) | 56-38-2 | Pesticides; | Pesticides; EMS code: P001 | Chemical |  | LAB
P001 | ug/g | Parathion | X260 | Sox-Ac/CH2Cl2 ext:GC/NPD | Sox-Ac/CH2Cl2 ext:GC/NPD | 0.01 | Y | Y | Y | 117 | 4/20/2002 10:25:00 AM | 10/16/2002 3:20:00 PM | 2/26/2004 1:05:29 AM | 1996-04-12 | 6/3/1998 6:51:37 PM | EMS |  | MassPerMass | mg/kg | Y | (mass conc.) |  | Parathion (mass conc.) | 56-38-2 | Pesticides; | Pesticides; EMS code: P001 | Chemical |  | LAB
P001 | ug/g | Parathion | ZN03 | DCM+Ac.Extr:GC/MS | DCM+Ac.Extr:GC/MS EPA 8270E (mod) | 0.05 | Y | Y | Y | 16 | 1/25/2022 | 11/2/2022 4:41:00 AM | 1/7/2023 3:12:11 PM | 1996-04-12 | 11/16/1999 2:27:40 PM | EMS | NPEPPIN | MassPerMass | mg/kg | Y | (mass conc.) |  | Parathion (mass conc.) | 56-38-2 | Pesticides; | Pesticides; EMS code: P001 | Chemical |  | LAB
P001 | mg/L | Parathion | F053 | DCM extr'n,conc'n,HGRC-PFPD/N,P det'n | Effluent samples are extracted with Dichloromethane. The extract is then concentrated and analysed using High Resolution Gas Chromatography with Pulsed Flame Photometric Dectection. Confirmation, if required, is by High Resolution Gas Chromatography with Nitrogen Phosphorus Detection. | 5e-05 | Y | Y | Y | 31 | 3/3/2000 7:36:00 AM | 1/22/2002 10:26:00 AM | 6/18/2003 2:53:51 PM | 1997-10-20 | 6/3/1998 6:51:47 PM | NPEPPIN |  | Concentration | mg/L | Y | (fl. conc.) |  | Parathion (fl. conc.) | 56-38-2 | Pesticides; | Pesticides; EMS code: P001 | Chemical |  | LAB
P002 | mg/L | Parathion Methyl | ZN01 | DCM+Ac.Extr:GC/MS | DCM+Ac.Extr:GC/MS | 0.0005 | Y | Y | Y | 52 | 10/3/2002 | 3/29/2012 10:15:00 PM | 5/3/2012 9:27:51 AM | 1999-11-16 |  | NPEPPIN |  | Concentration | mg/L | Y | (fl. conc.) |  | Parathion Methyl (fl. conc.) | 298-00-0 | Pesticides; | Pesticides; EMS code: P002 | Chemical |  | LAB
P002 | ug/g | Parathion Methyl | ZN03 | DCM+Ac.Extr:GC/MS | DCM+Ac.Extr:GC/MS EPA 8270E (mod) | 0.05 | Y | Y | Y | 14 | 1/25/2022 | 9/8/2022 | 12/15/2022 10:50:07 AM | 2003-06-16 |  | UMULLER |  | MassPerMass | mg/kg | Y | (mass conc.) |  | Parathion Methyl (mass conc.) | 298-00-0 | Pesticides; | Pesticides; EMS code: P002 | Chemical |  | LAB
P002 | ug/g | Parathion Methyl | X260 | Sox-Ac/CH2Cl2 ext:GC/NPD | Sox-Ac/CH2Cl2 ext:GC/NPD | 0.02 | Y | Y | Y | 97 | 4/20/2002 10:25:00 AM | 10/16/2002 3:20:00 PM | 2/25/2004 1:09:47 AM | 1996-04-12 | 6/3/1998 6:51:37 PM | EMS |  | MassPerMass | mg/kg | Y | (mass conc.) |  | Parathion Methyl (mass conc.) | 298-00-0 | Pesticides; | Pesticides; EMS code: P002 | Chemical |  | LAB
P002 | mg/L | Parathion Methyl | E660 | EPA 8270E (mod) | EPA 8270E (mod) Miscellaneous Pesticides by GC-MS | 0.0001 | Y | Y | Y | 68 | 12/7/2021 | 9/12/2022 | 12/8/2022 12:57:05 PM | 2021-12-09 | 12/9/2021 3:22:12 PM | EMS | EMS | Concentration | mg/L | Y | (fl. conc.) |  | Parathion Methyl (fl. conc.) | 298-00-0 | Pesticides; | Pesticides; EMS code: P002 | Chemical |  | LAB
P003 | mg/L | Phosmet | ZN01 | DCM+Ac.Extr:GC/MS | DCM+Ac.Extr:GC/MS | 0.0002 | Y | Y | Y | 52 | 10/3/2002 | 3/29/2012 10:15:00 PM | 5/3/2012 9:27:51 AM | 1996-04-12 | 11/16/1999 4:06:27 PM | EMS | NPEPPIN | Concentration | mg/L | Y | (fl. conc.) |  | Phosmet (fl. conc.) | 732-11-6 | Pesticides; | Pesticides; EMS code: P003 | Chemical |  | LAB
P003 | ug/g | Phosmet | X260 | Sox-Ac/CH2Cl2 ext:GC/NPD | Sox-Ac/CH2Cl2 ext:GC/NPD | 0.03 | Y | Y | Y | 97 | 4/20/2002 10:25:00 AM | 10/16/2002 3:20:00 PM | 2/25/2004 1:09:47 AM | 1996-04-12 | 6/3/1998 6:51:37 PM | EMS |  | MassPerMass | mg/kg | Y | (mass conc.) |  | Phosmet (mass conc.) | 732-11-6 | Pesticides; | Pesticides; EMS code: P003 | Chemical |  | LAB
P004 | mg/L | Phosalone | ZN01 | DCM+Ac.Extr:GC/MS | DCM+Ac.Extr:GC/MS | 0.0002 | Y | Y | Y | 52 | 10/3/2002 | 3/29/2012 10:15:00 PM | 5/3/2012 9:27:51 AM | 1996-04-12 | 11/16/1999 4:06:27 PM | EMS | NPEPPIN | Concentration | mg/L | Y | (fl. conc.) |  | Phosalone (fl. conc.) | 2310-17-0 | Pesticides; | Pesticides; EMS code: P004 | Chemical |  | LAB
P004 | ug/g | Phosalone | ZN03 | DCM+Ac.Extr:GC/MS | DCM+Ac.Extr:GC/MS EPA 8270E (mod) | 0.02 | Y | Y | Y | 19 | 4/20/2002 10:25:00 AM | 10/16/2002 3:20:00 PM | 2/25/2004 1:09:47 AM | 1996-04-12 | 11/16/1999 2:27:40 PM | EMS | NPEPPIN | MassPerMass | mg/kg | Y | (mass conc.) |  | Phosalone (mass conc.) | 2310-17-0 | Pesticides; | Pesticides; EMS code: P004 | Chemical |  | LAB
P005 | mg/L | Phorate | ZN01 | DCM+Ac.Extr:GC/MS | DCM+Ac.Extr:GC/MS | 0.0005 | Y | Y | Y | 52 | 10/3/2002 | 3/29/2012 10:15:00 PM | 5/3/2012 9:27:51 AM | 1996-04-12 | 11/16/1999 4:06:27 PM | EMS | NPEPPIN | Concentration | mg/L | Y | (fl. conc.) |  | Phorate (fl. conc.) | 298-02-2 | Pesticides; | Pesticides; EMS code: P005 | Chemical |  | LAB
P005 | mg/L | Phorate | E660 | EPA 8270E (mod) | EPA 8270E (mod) Miscellaneous Pesticides by GC-MS | 0.0001 | Y | Y | Y | 68 | 12/7/2021 | 9/12/2022 | 12/8/2022 12:57:05 PM | 2021-12-09 | 12/9/2021 3:22:12 PM | EMS | EMS | Concentration | mg/L | Y | (fl. conc.) |  | Phorate (fl. conc.) | 298-02-2 | Pesticides; | Pesticides; EMS code: P005 | Chemical |  | LAB
P005 | ug/g | Phorate | ZN03 | DCM+Ac.Extr:GC/MS | DCM+Ac.Extr:GC/MS EPA 8270E (mod) | 0.05 | Y | Y | Y | 35 | 4/20/2002 10:25:00 AM | 11/2/2022 4:41:00 AM | 1/7/2023 3:12:11 PM | 1996-04-12 | 11/16/1999 2:27:40 PM | EMS | NPEPPIN | MassPerMass | mg/kg | Y | (mass conc.) |  | Phorate (mass conc.) | 298-02-2 | Pesticides; | Pesticides; EMS code: P005 | Chemical |  | LAB
P006 | mg/L | Pirimicarb | ZN01 | DCM+Ac.Extr:GC/MS | DCM+Ac.Extr:GC/MS | 0.0001 | Y | Y | Y | 52 | 10/3/2002 | 3/29/2012 10:15:00 PM | 5/3/2012 9:27:51 AM | 1996-04-12 | 11/16/1999 4:06:27 PM | EMS | NPEPPIN | Concentration | mg/L | Y | (fl. conc.) |  | Pirimicarb (fl. conc.) | 23103-98-2 | Pesticides; | Pesticides; EMS code: P006 | Chemical |  | LAB
P008 | mg/L | Phosphamidon | ZN01 | DCM+Ac.Extr:GC/MS | DCM+Ac.Extr:GC/MS | 0.0002 | Y | Y | Y | 52 | 10/3/2002 | 3/29/2012 10:15:00 PM | 5/3/2012 9:27:52 AM | 1996-04-12 | 11/16/1999 4:06:27 PM | EMS | NPEPPIN | Concentration | mg/L | Y | (fl. conc.) |  | Phosphamidon (fl. conc.) | 13171-21-6 | Pesticides; | Pesticides; EMS code: P008 | Chemical |  | LAB
P008 | ug/g | Phosphamidon | ZN03 | DCM+Ac.Extr:GC/MS | DCM+Ac.Extr:GC/MS EPA 8270E (mod) | 0.02 | Y | Y | Y | 19 | 4/20/2002 10:25:00 AM | 10/16/2002 3:20:00 PM | 2/25/2004 1:09:47 AM | 1996-04-12 | 11/16/1999 2:27:40 PM | EMS | NPEPPIN | MassPerMass | mg/kg | Y | (mass conc.) |  | Phosphamidon (mass conc.) | 13171-21-6 | Pesticides; | Pesticides; EMS code: P008 | Chemical |  | LAB
P009 | mg/L | Bis(2ethylhex)Phthal | X457 | Liq-liq extract'n GC-ECD | Liq-liq extract'n GC-ECD | 0.001 | Y | Y | Y | 3 |  |  | 2/15/2016 10:44:41 AM | 1996-04-12 | 6/3/1998 6:51:37 PM | EMS |  | Concentration | mg/L | Y | (fl. conc.) |  | Bis(2ethylhex)Phthal (fl. conc.) | 117-81-7 | Organics; | Organics; EMS code: P009 | Chemical |  | LAB
P009 | ug/g | Bis(2ethylhex)Phthal | FS91 | DCM,MAP,AC/Hex extr,HRGC/LRMS | DCM;MAP;AC/Hex extr;HRGC/LRMS | 0.02 | Y | Y | Y | 4 | 9/1/2000 2:05:00 PM | 9/1/2000 3:18:00 PM | 9/9/2002 1:59:43 PM | 2001-03-08 |  | NPEPPIN |  | MassPerMass | mg/kg | Y | (mass conc.) |  | Bis(2ethylhex)Phthal (mass conc.) | 117-81-7 | Organics; | Organics; EMS code: P009 | Chemical |  | LAB
P009 | ug/g | Bis(2ethylhex)Phthal | X213 | Hex/AcExt,Part,Chrom,GC | Hex/AcExt;Part;Chrom;GC | 0.05 | Y | Y | Y | 83 |  |  | 10/24/1996 10:24:19 AM | 1996-04-12 | 6/3/1998 6:51:37 PM | EMS |  | MassPerMass | mg/kg | Y | (mass conc.) |  | Bis(2ethylhex)Phthal (mass conc.) | 117-81-7 | Organics; | Organics; EMS code: P009 | Chemical |  | LAB
P009 | ug/g wet | Bis(2ethylhex)Phthal | X220 | Hex/AcExt,Part,Chrom,GC | Hex/AcExt;Part;Chrom;GC | 0.01 | Y | Y | Y | 72 |  |  | 10/24/1996 10:01:25 AM | 1996-04-12 | 6/3/1998 6:51:42 PM | EMS |  | MassPerMass | µg/g | Y | (mass conc.) | (wet) | Bis(2ethylhex)Phthal (wet) (mass conc.) | 117-81-7 | Organics; | Organics; EMS code: P009 | Chemical |  | LAB
P009 | mg/L | Bis(2ethylhex)Phthal | F110 | DCM,liq/liq extr,HRGC/LRMS | DCM;liq/liq extr;HRGC/LRMS | 0.0002 | Y | Y | Y | 107 | 12/3/2002 | 6/4/2024 | 7/23/2024 11:41:50 AM | 2001-03-08 |  | NPEPPIN |  | Concentration | mg/L | Y | (fl. conc.) |  | Bis(2ethylhex)Phthal (fl. conc.) | 117-81-7 | Organics; | Organics; EMS code: P009 | Chemical |  | LAB
P010 | ug/g | Dimethyl Phthalate | FS91 | DCM,MAP,AC/Hex extr,HRGC/LRMS | DCM;MAP;AC/Hex extr;HRGC/LRMS | 0.02 | Y | Y | Y | 4 | 9/1/2000 2:05:00 PM | 9/1/2000 3:18:00 PM | 9/9/2002 1:59:43 PM | 2001-03-08 |  | NPEPPIN |  | MassPerMass | mg/kg | Y | (mass conc.) |  | Dimethyl Phthalate (mass conc.) | 131-11-3 | Organics; | Organics; EMS code: P010 | Chemical |  | LAB
P010 | ug/g wet | Dimethyl Phthalate | X220 | Hex/AcExt,Part,Chrom,GC | Hex/AcExt;Part;Chrom;GC | 0.01 | Y | Y | Y | 73 |  |  | 10/24/1996 10:01:25 AM | 1996-04-12 | 6/3/1998 6:51:42 PM | EMS |  | MassPerMass | µg/g | Y | (mass conc.) | (wet) | Dimethyl Phthalate (wet) (mass conc.) | 131-11-3 | Organics; | Organics; EMS code: P010 | Chemical |  | LAB
P010 | mg/L | Dimethyl Phthalate | F110 | DCM,liq/liq extr,HRGC/LRMS | DCM;liq/liq extr;HRGC/LRMS | 0.0001 | Y | Y | Y | 90 | 12/3/2002 | 6/4/2024 | 7/4/2024 11:06:09 AM | 2001-03-08 |  | NPEPPIN |  | Concentration | mg/L | Y | (fl. conc.) |  | Dimethyl Phthalate (fl. conc.) | 131-11-3 | Organics; | Organics; EMS code: P010 | Chemical |  | LAB
P010 | ug/g | Dimethyl Phthalate | X213 | Hex/AcExt,Part,Chrom,GC | Hex/AcExt;Part;Chrom;GC | 0.01 | Y | Y | Y | 83 |  |  | 10/24/1996 10:24:14 AM | 1996-04-12 | 6/3/1998 6:51:37 PM | EMS |  | MassPerMass | mg/kg | Y | (mass conc.) |  | Dimethyl Phthalate (mass conc.) | 131-11-3 | Organics; | Organics; EMS code: P010 | Chemical |  | LAB
P011 | mg/L | Diethyl Phthalate | F110 | DCM,liq/liq extr,HRGC/LRMS | DCM;liq/liq extr;HRGC/LRMS | 0.0002 | Y | Y | Y | 90 | 12/3/2002 | 6/4/2024 | 7/4/2024 11:06:09 AM | 2001-03-08 |  | NPEPPIN |  | Concentration | mg/L | Y | (fl. conc.) |  | Diethyl Phthalate (fl. conc.) | 84-66-2 | Organics; | Organics; EMS code: P011 | Chemical |  | LAB
P011 | ug/g | Diethyl Phthalate | FS91 | DCM,MAP,AC/Hex extr,HRGC/LRMS | DCM;MAP;AC/Hex extr;HRGC/LRMS | 0.02 | Y | Y | Y | 4 | 9/1/2000 2:05:00 PM | 9/1/2000 3:18:00 PM | 9/9/2002 1:59:43 PM | 2001-03-08 |  | NPEPPIN |  | MassPerMass | mg/kg | Y | (mass conc.) |  | Diethyl Phthalate (mass conc.) | 84-66-2 | Organics; | Organics; EMS code: P011 | Chemical |  | LAB
P011 | ug/g | Diethyl Phthalate | X213 | Hex/AcExt,Part,Chrom,GC | Hex/AcExt;Part;Chrom;GC | 0.01 | Y | Y | Y | 83 |  |  | 10/24/1996 10:24:14 AM | 1996-04-12 | 6/3/1998 6:51:37 PM | EMS |  | MassPerMass | mg/kg | Y | (mass conc.) |  | Diethyl Phthalate (mass conc.) | 84-66-2 | Organics; | Organics; EMS code: P011 | Chemical |  | LAB
P011 | ug/g wet | Diethyl Phthalate | X220 | Hex/AcExt,Part,Chrom,GC | Hex/AcExt;Part;Chrom;GC | 0.004 | Y | Y | Y | 71 |  |  | 10/24/1996 10:01:25 AM | 1996-04-12 | 6/3/1998 6:51:42 PM | EMS |  | MassPerMass | µg/g | Y | (mass conc.) | (wet) | Diethyl Phthalate (wet) (mass conc.) | 84-66-2 | Organics; | Organics; EMS code: P011 | Chemical |  | LAB
P012 | mg/L | Di-n-Butyl Phthalate | X457 | Liq-liq extract'n GC-ECD | Liq-liq extract'n GC-ECD | 0.01 | Y | Y | Y | 3 |  |  | 2/15/2016 10:44:41 AM | 1996-04-12 | 6/3/1998 6:51:37 PM | EMS |  | Concentration | mg/L | Y | (fl. conc.) |  | Di-n-Butyl Phthalate (fl. conc.) | 84-74-2 | Organics; | Organics; EMS code: P012 | Chemical |  | LAB
P012 | mg/L | Di-n-Butyl Phthalate | F110 | DCM,liq/liq extr,HRGC/LRMS | DCM;liq/liq extr;HRGC/LRMS | 0.0002 | Y | Y | Y | 98 | 2/17/2022 | 6/4/2024 | 7/23/2024 11:41:50 AM | 2001-03-08 |  | NPEPPIN |  | Concentration | mg/L | Y | (fl. conc.) |  | Di-n-Butyl Phthalate (fl. conc.) | 84-74-2 | Organics; | Organics; EMS code: P012 | Chemical |  | LAB
P012 | ug/g | Di-n-Butyl Phthalate | FS91 | DCM,MAP,AC/Hex extr,HRGC/LRMS | DCM;MAP;AC/Hex extr;HRGC/LRMS | 0.02 | Y | Y | Y | 4 | 9/1/2000 2:05:00 PM | 9/1/2000 3:18:00 PM | 9/9/2002 1:59:44 PM | 2001-03-08 |  | NPEPPIN |  | MassPerMass | mg/kg | Y | (mass conc.) |  | Di-n-Butyl Phthalate (mass conc.) | 84-74-2 | Organics; | Organics; EMS code: P012 | Chemical |  | LAB
P012 | ug/g | Di-n-Butyl Phthalate | X213 | Hex/AcExt,Part,Chrom,GC | Hex/AcExt;Part;Chrom;GC | 0.01 | Y | Y | Y | 83 |  |  | 10/24/1996 10:24:14 AM | 1996-04-12 | 6/3/1998 6:51:37 PM | EMS |  | MassPerMass | mg/kg | Y | (mass conc.) |  | Di-n-Butyl Phthalate (mass conc.) | 84-74-2 | Organics; | Organics; EMS code: P012 | Chemical |  | LAB
P012 | ug/g wet | Di-n-Butyl Phthalate | X220 | Hex/AcExt,Part,Chrom,GC | Hex/AcExt;Part;Chrom;GC | 0.004 | Y | Y | Y | 71 |  |  | 10/24/1996 10:01:25 AM | 1996-04-12 | 6/3/1998 6:51:42 PM | EMS |  | MassPerMass | µg/g | Y | (mass conc.) | (wet) | Di-n-Butyl Phthalate (wet) (mass conc.) | 84-74-2 | Organics; | Organics; EMS code: P012 | Chemical |  | LAB
P013 | ug/g wet | ButylBenzyl Phthalat | X220 | Hex/AcExt,Part,Chrom,GC | Hex/AcExt;Part;Chrom;GC | 0.004 | Y | Y | Y | 72 |  |  | 10/24/1996 10:01:25 AM | 1996-04-12 | 6/3/1998 6:51:42 PM | EMS |  | MassPerMass | µg/g | Y | (mass conc.) | (wet) | ButylBenzyl Phthalat (wet) (mass conc.) | 85-68-7 | Organics; | Organics; EMS code: P013 | Chemical |  | LAB
P013 | ug/g | ButylBenzyl Phthalat | FS91 | DCM,MAP,AC/Hex extr,HRGC/LRMS | DCM;MAP;AC/Hex extr;HRGC/LRMS | 0.02 | Y | Y | Y | 4 | 9/1/2000 2:05:00 PM | 9/1/2000 3:18:00 PM | 9/9/2002 1:59:44 PM | 2001-03-08 |  | NPEPPIN |  | MassPerMass | mg/kg | Y | (mass conc.) |  | ButylBenzyl Phthalat (mass conc.) | 85-68-7 | Organics; | Organics; EMS code: P013 | Chemical |  | LAB
P013 | mg/L | ButylBenzyl Phthalat | F110 | DCM,liq/liq extr,HRGC/LRMS | DCM;liq/liq extr;HRGC/LRMS | 0.0001 | Y | Y | Y | 90 | 12/3/2002 | 6/4/2024 | 7/4/2024 11:06:09 AM | 2001-03-08 |  | NPEPPIN |  | Concentration | mg/L | Y | (fl. conc.) |  | ButylBenzyl Phthalat (fl. conc.) | 85-68-7 | Organics; | Organics; EMS code: P013 | Chemical |  | LAB
P013 | ug/g | ButylBenzyl Phthalat | X213 | Hex/AcExt,Part,Chrom,GC | Hex/AcExt;Part;Chrom;GC | 0.01 | Y | Y | Y | 83 |  |  | 10/24/1996 10:24:14 AM | 1996-04-12 | 6/3/1998 6:51:37 PM | EMS |  | MassPerMass | mg/kg | Y | (mass conc.) |  | ButylBenzyl Phthalat (mass conc.) | 85-68-7 | Organics; | Organics; EMS code: P013 | Chemical |  | LAB
P014 | ug/g | Di-n-Octyl Phthalate | FS91 | DCM,MAP,AC/Hex extr,HRGC/LRMS | DCM;MAP;AC/Hex extr;HRGC/LRMS | 0.02 | Y | Y | Y | 4 | 9/1/2000 2:05:00 PM | 9/1/2000 3:18:00 PM | 9/9/2002 1:59:44 PM | 2001-03-08 |  | NPEPPIN |  | MassPerMass | mg/kg | Y | (mass conc.) |  | Di-n-Octyl Phthalate (mass conc.) | 117-84-0 | Organics; | Organics; EMS code: P014 | Chemical |  | LAB
P014 | ug/g | Di-n-Octyl Phthalate | X213 | Hex/AcExt,Part,Chrom,GC | Hex/AcExt;Part;Chrom;GC | 0.01 | Y | Y | Y | 83 |  |  | 10/24/1996 10:24:14 AM | 1996-04-12 | 6/3/1998 6:51:37 PM | EMS |  | MassPerMass | mg/kg | Y | (mass conc.) |  | Di-n-Octyl Phthalate (mass conc.) | 117-84-0 | Organics; | Organics; EMS code: P014 | Chemical |  | LAB
P014 | mg/L | Di-n-Octyl Phthalate | F110 | DCM,liq/liq extr,HRGC/LRMS | DCM;liq/liq extr;HRGC/LRMS | 0.0001 | Y | Y | Y | 91 | 12/3/2002 | 6/4/2024 | 7/4/2024 11:06:09 AM | 2001-03-08 |  | NPEPPIN |  | Concentration | mg/L | Y | (fl. conc.) |  | Di-n-Octyl Phthalate (fl. conc.) | 117-84-0 | Organics; | Organics; EMS code: P014 | Chemical |  | LAB
P014 | ug/g wet | Di-n-Octyl Phthalate | X220 | Hex/AcExt,Part,Chrom,GC | Hex/AcExt;Part;Chrom;GC | 0.01 | Y | Y | Y | 73 |  |  | 10/24/1996 10:01:25 AM | 1996-04-12 | 6/3/1998 6:51:42 PM | EMS |  | MassPerMass | µg/g | Y | (mass conc.) | (wet) | Di-n-Octyl Phthalate (wet) (mass conc.) | 117-84-0 | Organics; | Organics; EMS code: P014 | Chemical |  | LAB
P015 | ug/g | Propazine | ZN03 | DCM+Ac.Extr:GC/MS | DCM+Ac.Extr:GC/MS EPA 8270E (mod) | 0.01 | Y | Y | Y | 14 | 1/25/2022 | 9/8/2022 | 12/15/2022 10:50:07 AM | 1996-04-12 | 11/16/1999 2:27:40 PM | EMS | NPEPPIN | MassPerMass | mg/kg | Y | (mass conc.) |  | Propazine (mass conc.) | 139-40-2 | Pesticides; | Pesticides; EMS code: P015 | Chemical |  | LAB
P015 | mg/L | Propazine | E660 | EPA 8270E (mod) | EPA 8270E (mod) Miscellaneous Pesticides by GC-MS | 0.0001 | Y | Y | Y | 68 | 12/7/2021 | 9/12/2022 | 12/8/2022 12:57:05 PM | 2021-12-09 | 12/9/2021 3:22:12 PM | EMS | EMS | Concentration | mg/L | Y | (fl. conc.) |  | Propazine (fl. conc.) | 139-40-2 | Pesticides; | Pesticides; EMS code: P015 | Chemical |  | LAB
P015 | mg/L | Propazine | ZN01 | DCM+Ac.Extr:GC/MS | DCM+Ac.Extr:GC/MS | 0.0001 | Y | Y | Y | 52 | 10/3/2002 | 3/29/2012 10:15:00 PM | 5/3/2012 9:27:52 AM | 1996-04-12 | 11/16/1999 4:06:27 PM | EMS | NPEPPIN | Concentration | mg/L | Y | (fl. conc.) |  | Propazine (fl. conc.) | 139-40-2 | Pesticides; | Pesticides; EMS code: P015 | Chemical |  | LAB
P015 | ug/g | Propazine | FS61 | DCM +MeOH extr,HRGC-NPD | Sediment samples are shaken with a mixture of Dichloromethane and Methanol, filtered then re-extracted with Dichloromethane. The extract is concentrated and analysed using High Resolution Gas Chromatography with a Thermionic Specific Detector (TSD), alternately referred to as a Nitrogen Phosphorus Detector (NPD). | 0.01 | Y | Y | Y | 34 | 4/20/2002 3:40:00 PM | 7/16/2002 3:20:00 PM | 2/26/2004 1:05:29 AM | 1998-10-02 |  | NPEPPIN |  | MassPerMass | mg/kg | Y | (mass conc.) |  | Propazine (mass conc.) | 139-40-2 | Pesticides; | Pesticides; EMS code: P015 | Chemical |  | LAB
P015 | mg/L | Propazine | F056 | DCM extr,HRGC-NPD | Effluent samples are extracted with Dichloromethane. The extract is concentrated and analysed using High Resolution Gas Chromatography with a Thermionic Specific Detector (TSD), alternately referred to as a Nitrogen Phosphorus Detector (NPD). | 0.0005 | Y | Y | Y | 12 | 4/3/2000 10:51:00 AM | 2/16/2001 3:36:00 PM | 9/11/2002 9:34:57 AM | 1998-10-02 |  | NPEPPIN |  | Concentration | mg/L | Y | (fl. conc.) |  | Propazine (fl. conc.) | 139-40-2 | Pesticides; | Pesticides; EMS code: P015 | Chemical |  | LAB
P016 | mg/L | Profluralin | ZN01 | DCM+Ac.Extr:GC/MS | DCM+Ac.Extr:GC/MS | 0.0005 | Y | Y | Y | 28 | 7/28/2005 7:03:00 AM | 3/29/2012 10:15:00 PM | 5/3/2012 9:27:52 AM | 1996-04-12 | 11/16/1999 4:06:27 PM | EMS | NPEPPIN | Concentration | mg/L | Y | (fl. conc.) |  | Profluralin (fl. conc.) | 26399-36-0 | Pesticides; | Pesticides; EMS code: P016 | Chemical |  | LAB
P017 | ug/g | Prometryne | ZN03 | DCM+Ac.Extr:GC/MS | DCM+Ac.Extr:GC/MS EPA 8270E (mod) | 0.02 | Y | Y | Y | 16 | 1/25/2022 | 11/2/2022 4:41:00 AM | 1/7/2023 3:12:11 PM | 1996-04-12 | 11/16/1999 2:27:40 PM | EMS | NPEPPIN | MassPerMass | mg/kg | Y | (mass conc.) |  | Prometryne (mass conc.) | 7287-19-6 | Pesticides; | Pesticides; EMS code: P017 | Chemical |  | LAB
P017 | mg/L | Prometryne | E660 | EPA 8270E (mod) | EPA 8270E (mod) Miscellaneous Pesticides by GC-MS | 0.0001 | Y | Y | Y | 68 | 12/7/2021 | 9/12/2022 | 12/8/2022 12:57:05 PM | 2021-12-09 | 12/9/2021 3:22:12 PM | EMS | EMS | Concentration | mg/L | Y | (fl. conc.) |  | Prometryne (fl. conc.) | 7287-19-6 | Pesticides; | Pesticides; EMS code: P017 | Chemical |  | LAB
P017 | mg/L | Prometryne | ZN01 | DCM+Ac.Extr:GC/MS | DCM+Ac.Extr:GC/MS | 0.0002 | Y | Y | Y | 28 | 7/28/2005 7:03:00 AM | 3/29/2012 10:15:00 PM | 5/3/2012 9:27:52 AM | 1996-04-12 | 11/16/1999 4:06:27 PM | EMS | NPEPPIN | Concentration | mg/L | Y | (fl. conc.) |  | Prometryne (fl. conc.) | 7287-19-6 | Pesticides; | Pesticides; EMS code: P017 | Chemical |  | LAB
P018 | mg/L | Propoxur | P045 | DCM Extr:HPLC | DCM Extr:HPLC | 0.0005 | Y | Y | Y | 1 | 7/29/2005 7:10:00 AM | 7/29/2005 7:10:00 AM | 6/2/2009 12:01:09 PM | 1996-04-12 | 6/3/1998 6:51:42 PM | EMS |  | Concentration | mg/L | Y | (fl. conc.) |  | Propoxur (fl. conc.) | 114-26-1 | Pesticides; | Pesticides; EMS code: P018 | Chemical |  | LAB
P019 | ug/g | PCBs - Total | X383 | Extr, GC/MS or GC/ECD | Determination of Polychloriated Biphenyls (PCBs) In Soil by GC/ECD. PURPOSE/PRINCIPLE OF METHOD: This document details the requirements and procedures for the analysis of polychlorinated biphenyls (PCBs) in soil by gas chromatography electron capture detector (GC/ECD). A homogenized soil sample is extracted using dichloromethane (DCM). The extract may be cleaned up using column chromatography with 2% Deactivated Florisil. When contamination with sulphur is suspected, further cleanup using copper is required. Analysis is performed using a GC/ECD. Similar methods have been developed for other matrices including waste and transformer oils, water, swabs, florisil tubes, biota. Generally, the sample extraction procedures differ while the GC analysis methods are similar or identical for all matrices. SCOPE: This procedure describes the determination of PCBs in soil and sediments. Other materials such as wood waste may also be analyzed by this method. DETECTION LIMITS: Target Compound Reported Detection Limit (mg/g) Arochlor 1242 0.03, Arochlor 1248 0.03, Arochlor 1254 0.03, Arochlor 1260 0.03, INTERFERENCES: Interferences (e.g. co-eluting peaks on the chromatogram) may be due to sulphur, phthalate esters, or from chlorinated compounds. Suspect samples must be reanalyzed using appropriate cleanup steps. SAMPLE HANDLING & PRESERVATION: Sample Requirements: Whenever possible, samples must completely fill the bottle such that little or no headspace is present to eliminate the loss of volatile hydrocarbon components. (~125g of sample) Sample Containers: Glass jars that have been solvent rinsed with acetone then hexane, then dried. Shipment and Storage: Samples are to be refrigerated (at 4±2°C) during shipping and prior to analysis. Samples suspected to contain high levels of analyte are to be segregated from other samples. Holding Time: Holding time prior to extraction is 14 days. REFERENCE: U.S. EPA Method 8080:Organochlorine Pesticides and PCBs Analysis. | 0.01 | Y | Y | Y | 18 | 1/11/2022 | 9/7/2022 | 1/11/2024 1:07:08 PM | 2022-07-07 |  | EMS |  | MassPerMass | mg/kg | Y | (mass conc.) |  | PCBs - Total (mass conc.) |  | Polychlorinated biphenylds (PCBs); | Polychlorinated biphenylds (PCBs); EMS code: P019 | Chemical | Total | LAB
P019 | mg/L | PCBs - Total | E-8082 | EPA-8082A | See reference for EPA method 8082A POLYCHLORINATED BIPHENYLS (PCBs) BY GAS CHROMATOGRAPHY |  | Y | Y | Y | 14 | 12/28/2021 | 1/11/2022 | 9/2/2022 6:21:54 PM | 2022-04-06 |  | EMS |  | Concentration | mg/L | Y | (fl. conc.) |  | PCBs - Total (fl. conc.) |  | Polychlorinated biphenylds (PCBs); | Polychlorinated biphenylds (PCBs); EMS code: P019 | Chemical | Total | LAB
P019 | mg/L | PCBs - Total | F018 | Hexane extr,GC/ECD | Effluent and swab samples are extracted with hexane, the extract concentrated, and then treated using Florisil, concentrated Sulfuric Acid, activated Copper and Mercury (as required) to remove interferences. The samples are then analysed using Megabore Column Gas Chromatography with Electron Capture Detection. Oil samples are first diluted with Hexane then treated in the manner outlined for effluent and swab samples. The samples are then analysed using Megabore Column Gas Chromatography with Electron Capture Detection. | 0.0001 | Y | Y | Y | 59 | 10/21/1998 6:49:00 PM | 9/7/2023 | 7/23/2024 11:41:50 AM | 1996-09-17 | 6/3/1998 6:51:45 PM | MPAK |  | Concentration | mg/L | Y | (fl. conc.) |  | PCBs - Total (fl. conc.) |  | Polychlorinated biphenylds (PCBs); | Polychlorinated biphenylds (PCBs); EMS code: P019 | Chemical | Total | LAB
P019 | mg/L | PCBs - Total | E660 | EPA 8270E (mod) | EPA 8270E (mod) Miscellaneous Pesticides by GC-MS | 6e-05 | Y | Y | Y | 8 | 5/24/2023 | 11/1/2023 | 11/16/2023 3:38:10 PM | 2023-08-16 |  | EMS |  | Concentration | mg/L | Y | (fl. conc.) |  | PCBs - Total (fl. conc.) |  | Polychlorinated biphenylds (PCBs); | Polychlorinated biphenylds (PCBs); EMS code: P019 | Chemical | Total | LAB
P019 | ug/g | PCBs - Total | PCBT | Calculation of Total PCB | Calculation of Total PCB. Total PCB is the sum of the concentrations of PCB aroclors 1016, 1221, 1232, 1242, 1248, 1254, 1260, 1262, and 1268. Results below detection limit (DL) are treated as zero. The Total PCB detection limit is equal to the highest of the aroclor detection limits used in the sum. | 0.02 | Y | Y | Y | 28 | 3/16/2020 | 3/25/2020 | 2/22/2023 12:51:03 PM | 2020-08-28 |  | RDGIBSON |  | MassPerMass | mg/kg | Y | (mass conc.) |  | PCBs - Total (mass conc.) |  | Polychlorinated biphenylds (PCBs); | Polychlorinated biphenylds (PCBs); EMS code: P019 | Chemical | Total | LAB
P019 | ug/g | PCBs - Total | F115 | Hex dil,Hex extr'n,GC/ECD | Effluent and swab samples are extracted with hexane, the extract concentrated, and then treated using Florisil, concentrated Sulfuric Acid, activated Copper and Mercury (as required) to remove interferences. The samples are then analysed using Megabore Column Gas Chromatography with Electron Capture Detection. Oil samples are first diluted with Hexane then treated in the manner outlined for effluent and swab samples. The samples are then analysed using Megabore Column Gas Chromatography with Electron Capture Detection. | 0.1 | Y | Y | Y | 1 | 4/24/2002 9:20:00 AM | 4/24/2002 9:20:00 AM | 2/13/2004 1:01:32 AM | 2001-05-24 |  | NPEPPIN |  | MassPerMass | mg/kg | Y | (mass conc.) |  | PCBs - Total (mass conc.) |  | Polychlorinated biphenylds (PCBs); | Polychlorinated biphenylds (PCBs); EMS code: P019 | Chemical | Total | LAB
P019 | ug/g | PCBs - Total | P023 | PCB's | PCB's | 0.5 | Y | Y | Y | 387 | 4/20/2002 10:10:00 AM | 4/3/2003 10:35:00 AM | 2/26/2004 1:05:29 AM | 1996-04-12 | 6/3/1998 6:51:37 PM | EMS |  | MassPerMass | mg/kg | Y | (mass conc.) |  | PCBs - Total (mass conc.) |  | Polychlorinated biphenylds (PCBs); | Polychlorinated biphenylds (PCBs); EMS code: P019 | Chemical | Total | LAB
P019 | mg/L | PCBs - Total | P020 | PCB's | PCB's | 0.0004 | Y | Y | Y | 82 |  |  | 10/24/1996 2:54:25 PM | 1996-04-12 | 6/3/1998 6:51:42 PM | EMS |  | Concentration | mg/L | Y | (fl. conc.) |  | PCBs - Total (fl. conc.) |  | Polychlorinated biphenylds (PCBs); | Polychlorinated biphenylds (PCBs); EMS code: P019 | Chemical | Total | LAB
P020 | mg/L | Propylene Glycol | F026 | Acetone Dilution, HRGC-FID | Effluent samples are diluted with Acetone and analysed using direct injection into a High Resolutioin Gas Chromatograph using Flame Ionization Detection (HRGC-FID). | 0.05 | Y | Y | Y | 4 | 12/8/1998 12:28:00 PM | 3/27/2000 10:38:00 AM | 8/19/2000 2:23:13 AM | 1997-01-14 | 6/3/1998 6:51:46 PM | MPAK |  | Concentration | mg/L | Y | (fl. conc.) |  | Propylene Glycol (fl. conc.) | 57-55-6 | Organics; | Organics; EMS code: P020 | Chemical |  | LAB
P020 | mg/L | Propylene Glycol | GLYC | Glycols in Soil and Water by GC/FID | GLYCOLS IN SOIL AND WATER PURPOSE/PRINCIPLE OF METHOD: Soils are extracted with water and liquids are injected directly or diluted with water. The final extract is analyzed by direct injection gas chromatography with a flame ionization detector (GC/FID) due to the high solubility of glycols. SCOPE: Glycols are commonly used for many applications such as a de-icing agent in aircraft and runway, as an anti-freeze in heating and cooling systems, as an industrial solvent, and in paint and plastics. The end data is used for various purposes including contamination assessment, environmental regulatory compliance and land reclamation. REFERENCES: Nonhalogenated Organics Using GC/FID. US EPA. Test Methods for Evaluating Solid Waste, Physical/Chemical Methods -- SW-846, Method 8015B Revision 2, December 1996 | 10 | Y | Y | Y | 17 | 4/14/2015 10:15:00 PM | 9/8/2018 | 12/10/2019 3:45:08 PM | 2012-05-30 | 6/4/2012 12:03:55 PM | RDGIBSON | RDGIBSON | Concentration | mg/L | Y | (fl. conc.) |  | Propylene Glycol (fl. conc.) | 57-55-6 | Organics; | Organics; EMS code: P020 | Chemical |  | LAB
P021 | ug/g | Pentachloroanisole | FS11 | Acetone extr,Hexane,Diazometh.Methyl'n,HRGC/Electron Capture | Sediment and homogenized biota samples to be analysed for Chlorinated Phenols are Acetone extracted using a wrist-action shaker, partitioned with 2% NaCl and Hexane, methylated, then cleaned up and analysed in the manner outlined for effluent samples. Sediment and homogenized biota samples to be analysed for Chlorinated Anisoles are Acetone extracted using a wrist-action shaker, partitioned with 2% NaCl and Hexane, then cleaned up and analysed in the manner outlined for effluent samples. The extracts are analysed for Chlorinated Anisoles and/or Chlorinated Phenols using High Resolution Gas Chromatography with Electron Capture detection. | 0.0001 | Y | Y | Y | 4 | 8/29/2000 4:27:00 PM | 8/29/2000 4:29:00 PM | 6/18/2002 3:50:39 PM | 1997-01-14 | 6/3/1998 6:51:46 PM | MPAK |  | MassPerMass | mg/kg | Y | (mass conc.) |  | Pentachloroanisole (mass conc.) | 1825-21-4 | Pesticides; | Pesticides; EMS code: P021 | Chemical |  | LAB
P021 | mg/L | Pentachloroanisole | F030 | Acid,Hexane extr,Diazomethane Methylation,HRGC/Electron Capt | Effluent samples to be analysed for Chlorinated Anisoles are extracted with Hexane. After concentration the extract is cleaned up using Florisil, concentrated Sulfuric Acid and Mercury. The extracts are analysed for Chlorinated Anisoles and/or Chlorinated Phenols using High Resolution Gas Chromatography with Electron Capture detection. | 1e-06 | Y | Y | Y | 3 | 8/25/2000 3:03:00 PM | 9/1/2000 9:20:00 AM | 6/18/2002 1:43:30 PM | 1997-01-14 | 6/3/1998 6:51:46 PM | MPAK |  | Concentration | mg/L | Y | (fl. conc.) |  | Pentachloroanisole (fl. conc.) | 1825-21-4 | Pesticides; | Pesticides; EMS code: P021 | Chemical |  | LAB
P022 | mg/L | Pentachlorophenol | P008 | Chlor. Phenols | An acidified water sample is extracted with dichloromethane. The raw extract is concentrated and the phenols and acids are reacted with diazomethane, or other suitable derivatizing agent, to produce the corresponding derivatives. If required, the extracts are subjected to Florisil column chromatography and divided into different fractions. The derivatives are analyzed by electron capture gas liquid chromatography. | 0.0001 | Y | Y | Y | 336 | 8/11/2004 9:03:00 AM | 10/14/2022 2:44:00 AM | 10/20/2022 3:55:15 PM | 1996-04-12 | 6/3/1998 6:51:37 PM | EMS |  | Concentration | mg/L | Y | (fl. conc.) |  | Pentachlorophenol (fl. conc.) | 87-86-5 | Pesticides; | Pesticides; EMS code: P022 | Chemical |  | LAB
P022 | ug/g | Pentachlorophenol | P028 | Chlor. Phenols | Chlor. Phenols | 0.005 | Y | Y | Y | 167 |  |  | 10/24/1996 12:44:22 PM | 1996-04-12 | 6/3/1998 6:51:45 PM | EMS |  | MassPerMass | mg/kg | Y | (mass conc.) |  | Pentachlorophenol (mass conc.) | 87-86-5 | Pesticides; | Pesticides; EMS code: P022 | Chemical |  | LAB
P022 | ug/g | Pentachlorophenol | X212 | Hex/AcExtr,Part,IonEx,GC | Hex/AcExtr;Part;IonEx;GC | 0.001 | Y | Y | Y | 161 |  |  | 10/24/1996 2:26:15 PM | 1996-04-12 | 6/3/1998 6:51:37 PM | EMS |  | MassPerMass | mg/kg | Y | (mass conc.) |  | Pentachlorophenol (mass conc.) | 87-86-5 | Pesticides; | Pesticides; EMS code: P022 | Chemical |  | LAB
P022 | mg/L | Pentachlorophenol | EQ01 | EQUIS (mg/L) | EQUIS (mg/L) | 0.0001 | N | Y | Y | 9 |  |  | 10/22/1996 9:17:28 PM | 1996-04-12 | 6/3/1998 6:51:37 PM | EMS |  | Concentration | mg/L | Y | (fl. conc.) |  | Pentachlorophenol (fl. conc.) | 87-86-5 | Pesticides; | Pesticides; EMS code: P022 | Chemical |  | LAB
P022 | ug/g wet | Pentachlorophenol | X219 | Hex/AcExtr,Part,IonEx,GC | Hex/AcExtr;Part;IonEx;GC | 0.0002 | Y | Y | Y | 72 |  |  | 10/24/1996 10:01:21 AM | 1996-04-12 | 6/3/1998 6:51:42 PM | EMS |  | MassPerMass | µg/g | Y | (mass conc.) | (wet) | Pentachlorophenol (wet) (mass conc.) | 87-86-5 | Pesticides; | Pesticides; EMS code: P022 | Chemical |  | LAB
P022 | mg/L | Pentachlorophenol | AEWA | H2SO4 acid,CHCl2 extr.florisil. GC/MS | Effluent and water samples are acidified with Sulphuric Acid and extracted with Dichloromethane. After concentration, the extract is methylated using ethereal Diazomethane. Extracts are cleaned up as required using Florisil. The extracts are analyzed for Acid Extractable Herbicides using GC/MS. | 0.0001 | Y | Y | Y | 69 | 6/16/2003 | 6/16/2003 | 2/10/2004 1:02:12 AM | 1996-04-12 | 6/3/1998 6:51:42 PM | EMS |  | Concentration | mg/L | Y | (fl. conc.) |  | Pentachlorophenol (fl. conc.) | 87-86-5 | Pesticides; | Pesticides; EMS code: P022 | Chemical |  | LAB
P022 | mg/L | Pentachlorophenol | F028 | Acid,Extr.MeCl,HRGC/LRMS:TIM,SIM | A measured volume of sample is acidified and extracted with Methylene Chloride. For samples which have a high organic content that potentially produce chromatographic interferences an initial pre-extraction under basic conditions is performed. The sample extract is then evaporated to small volume and analysed by High Resolution Gas Chromatography / Low Resolution Mass Spectrometry (HRGC/LRMS). LRMS data is acquired in Total Ion Mode (TIM) and quantitation is performed using a single mass ion. Positive Identification of the various target compounds is made by comparison of absolute retention times to those of standards combined with computer based mass spectral mass/intensity library searching and matching routines. | 0.005 | Y | Y | Y | 149 | 4/10/1997 3:19:00 PM | 11/1/2022 4:39:00 AM | 4/21/2023 4:26:45 PM | 1997-01-14 | 6/3/1998 6:51:46 PM | MPAK |  | Concentration | mg/L | Y | (fl. conc.) |  | Pentachlorophenol (fl. conc.) | 87-86-5 | Pesticides; | Pesticides; EMS code: P022 | Chemical |  | LAB
P022 | mg/L | Pentachlorophenol | X268 | CH2Cl2 ext:GC/MS | CH2Cl2 ext:GC/MS |  | Y | Y | Y | 34 | 12/28/2021 | 6/7/2024 | 6/26/2024 4:28:10 PM | 2022-04-21 |  | EMS |  | Concentration | mg/L | Y | (fl. conc.) |  | Pentachlorophenol (fl. conc.) | 87-86-5 | Pesticides; | Pesticides; EMS code: P022 | Chemical |  | LAB
P022 | mg/L | Pentachlorophenol | F029 | Acid,DiethylEther extr,Diazometh.Methyl'n,HRGC/Electron Capt | Effluent samples to be analysed for Chlorinated Phenols are acidified with dilute Sulfuric Acid and extracted with Diethyl Ether. After concentration the extract is methylated using ethereal Diazomethane and cleaned up with concentrated Sulfuric Acid and Mercury. The extracts are analysed for Chlorinated Anisoles and/or Chlorinated Phenols using High Resolution Gas Chromatography with Electron Capture detection. | 1e-06 | Y | Y | Y | 38 | 2/21/2000 10:00:00 AM | 8/20/2001 12:43:00 PM | 9/10/2002 9:55:58 AM | 1997-01-14 | 6/3/1998 6:51:46 PM | MPAK |  | Concentration | mg/L | Y | (fl. conc.) |  | Pentachlorophenol (fl. conc.) | 87-86-5 | Pesticides; | Pesticides; EMS code: P022 | Chemical |  | LAB
P022 | mg/L | Pentachlorophenol | PP11 | Ac. anhydride ext:GC/MS | Ac. anhydride ext:GC/MS | 5e-05 | Y | Y | Y | 225 |  |  | 10/24/1996 5:30:20 AM | 1996-04-12 | 6/3/1998 6:51:37 PM | EMS |  | Concentration | mg/L | Y | (fl. conc.) |  | Pentachlorophenol (fl. conc.) | 87-86-5 | Pesticides; | Pesticides; EMS code: P022 | Chemical |  | LAB
P022 | ug/g | Pentachlorophenol | FS15 | DCM/MEOH acid extr'n,DCM,methn,Florisil,HRGC-EC | Sediment samples are shaken with a mixture of Dichloromethane, Methanol and Sulphuric Acid, filtered, then extracted with Dichloromethane. After concentration, the extract is methylated using ethereal Diazomethane. Extracts are cleaned up as required using Florisil. The extracts are analysed for Acid Extractable Herbicides using High Resolution Gas Chromatography with Electron Capture detection. Positive results may be verified by use of a secondary confirmation GC column. | 0.005 | Y | Y | Y | 27 | 4/4/2002 | 6/24/2003 | 2/26/2004 1:05:29 AM | 1998-12-30 |  | NPEPPIN |  | MassPerMass | mg/kg | Y | (mass conc.) |  | Pentachlorophenol (mass conc.) | 87-86-5 | Pesticides; | Pesticides; EMS code: P022 | Chemical |  | LAB
P022 | ug/g | Pentachlorophenol | P029 | Chlor Phen | Chlor Phen | 0.01 | Y | Y | Y | 101 |  |  | 10/24/1996 12:38:17 PM | 1996-04-12 | 6/3/1998 6:51:37 PM | EMS |  | MassPerMass | mg/kg | Y | (mass conc.) |  | Pentachlorophenol (mass conc.) | 87-86-5 | Pesticides; | Pesticides; EMS code: P022 | Chemical |  | LAB
P022 | ug/g | Pentachlorophenol | AEAN | H2SO4,CHCL2/MeOH extr.florisil. GC/MS | Sediment samples are shaken with a mixture of Dichloromethane, Methanol and Sulphuric Acid, filtered, then extracted with Dichloromethane. After concentration, the extract is methylated using ethereal Diazomethane. Extracts are cleaned up as required using Florisil. The extracts are analyzed for Acid Extractable Herbicides using GC/MS. | 0.005 | Y | Y | Y | 39 |  |  | 10/24/1996 12:38:38 PM | 1996-04-12 | 6/3/1998 6:51:37 PM | EMS |  | MassPerMass | mg/kg | Y | (mass conc.) |  | Pentachlorophenol (mass conc.) | 87-86-5 | Pesticides; | Pesticides; EMS code: P022 | Chemical |  | LAB
P022 | ug/g | Pentachlorophenol | DCMA | DCM/ Acetone/MTBE extr, AA/TMA deriv, GCMS | DCM/ Acetone/MTBE extr; AA/TMA deriv; GCMS. EPA 3570: 8270: Knapp(1979) | 0.02 | Y | Y | Y | 51 | 5/22/2015 | 3/20/2020 | 2/21/2023 2:10:29 PM | 2015-05-26 |  | RDGIBSON |  | MassPerMass | mg/kg | Y | (mass conc.) |  | Pentachlorophenol (mass conc.) | 87-86-5 | Pesticides; | Pesticides; EMS code: P022 | Chemical |  | LAB
P022 | mg/L | Pentachlorophenol | X265 | CH2Cl2Ext:Ac/Base:GC/ECD | CH2Cl2Ext:Ac/Base:GC/ECD | 0.001 | Y | Y | Y | 51 |  |  | 8/23/2000 11:41:33 AM | 1996-04-12 | 6/3/1998 6:51:45 PM | EMS |  | Concentration | mg/L | Y | (fl. conc.) |  | Pentachlorophenol (fl. conc.) | 87-86-5 | Pesticides; | Pesticides; EMS code: P022 | Chemical |  | LAB
P022 | mg/L | Pentachlorophenol | F034 | High Res GC-ECD/Chloroacetylation | Potassium Carbonate is dissolved in a measured volume of neutralized sample. Any Chlorophenols in the solution are then chloroacetylated in situ with Chloroacetic Anhydride in the presence of Petroleum Ether (30-60 C). The Petroleum Ether extract is dried over anhydrous Sodium Sulfate, evaporated to small volume and analysed using High Resolution Gas Chromatography with Electron Capture Detection (HRGC-ECD). | 5e-05 | Y | Y | Y | 30 | 7/16/1997 4:10:00 PM | 12/22/1999 3:42:00 PM | 12/25/1999 1:12:13 AM | 1997-01-16 | 6/3/1998 6:51:46 PM | MPAK |  | Concentration | mg/L | Y | (fl. conc.) |  | Pentachlorophenol (fl. conc.) | 87-86-5 | Pesticides; | Pesticides; EMS code: P022 | Chemical |  | LAB
P022 | ug/g | Pentachlorophenol | FS11 | Acetone extr,Hexane,Diazometh.Methyl'n,HRGC/Electron Capture | Sediment and homogenized biota samples to be analysed for Chlorinated Phenols are Acetone extracted using a wrist-action shaker, partitioned with 2% NaCl and Hexane, methylated, then cleaned up and analysed in the manner outlined for effluent samples. Sediment and homogenized biota samples to be analysed for Chlorinated Anisoles are Acetone extracted using a wrist-action shaker, partitioned with 2% NaCl and Hexane, then cleaned up and analysed in the manner outlined for effluent samples. The extracts are analysed for Chlorinated Anisoles and/or Chlorinated Phenols using High Resolution Gas Chromatography with Electron Capture detection. | 0.0001 | Y | Y | Y | 44 | 2/22/2000 8:31:00 AM | 1/9/2001 4:56:00 PM | 8/21/2002 2:10:03 PM | 1997-01-14 | 6/3/1998 6:51:46 PM | MPAK |  | MassPerMass | mg/kg | Y | (mass conc.) |  | Pentachlorophenol (mass conc.) | 87-86-5 | Pesticides; | Pesticides; EMS code: P022 | Chemical |  | LAB
P022 | ug/g | Pentachlorophenol | 1332 | Phenolics by Tumbler/GC-MS | Phenolics by Tumbler/GC-MS | 0.02 | Y | Y | Y | 22 | 1/17/2022 | 9/22/2022 8:16:00 PM | 1/11/2024 1:07:08 PM | 2022-07-07 |  | EMS |  | MassPerMass | mg/kg | Y | (mass conc.) |  | Pentachlorophenol (mass conc.) | 87-86-5 | Pesticides; | Pesticides; EMS code: P022 | Chemical |  | LAB
P023 | mg/L | Polychlor. Biphenyls | X303 | CH2Cl2 Extr:GC/ECD | Polychlorinated Biphenyls (PCBs) In Water by Gas Chromatography-Electron Capture Detector (GC/ECD). A water sample, usually 1000 mL, is extracted using dichloromethane. Sulphuric acid clean-up step is performed in every sample. Dirty samples with dark colored extracts are put through a clean-up step using 2% Deactivated Florisil. If sulphur is present in the water samples, especially for the river or stream water samples, a copper metal clean-up step may be used. A final analysis is performed using a GC equipped with an ECD. SCOPE: This procedure describes the determination of PCBs in water. If Organochlorine Pesticides (OC's) and PCB's are to be analyzed on the same water sample, a separate test method (Method for the Determination of Pesticides in Water by GC/MS and LC/MS or HPLC Fluorescence Detection DCN 62-C-043) is used. DETECTION LIMITS: Target Compound Reported Detection Limit (mg/L) Arochlor 1221 0.1 Arochlor 1242 0.1 Arochlor 1248 0.1 Arochlor 1254 0.1 Arochlor 1260 0.1 INTERFERENCES: Interferences (e.g. co-eluting peaks on the chromatogram) may be due to sulfur,phthalate esters, or from chlorinated compounds. Suspect samples must be reanalyzed using appropriate cleanup steps. SAMPLE HANDLING & PRESERVATION: Sample Containers: For sample collection, use 1 litre amber glass bottles that have been pre-rinsed with acetone and hexane then dried. Bottle lids must be teflon-lined, or alternatively lined with solvent-rinsed aluminium foil. Shipment and Storage: Samples must be refrigerated (at 2-8°C) and stored in the dark. Holding Time: Samples must be extracted within 7 days of sampling; REFERENCE: U.S. EPA Method 8080: Organochlorine Pesticides and PCB's Analysis. SW-846 3rd Ed.,September, 1986. | 0.001 | N | N | Y | 1 |  |  | 10/23/1996 11:03:46 PM | 1996-04-12 | 1/9/2013 4:36:43 PM | EMS | RDGIBSON | Concentration | mg/L | Y | (fl. conc.) |  | Polychlor. Biphenyls (fl. conc.) |  | Polychlorinated biphenylds (PCBs); | Polychlorinated biphenylds (PCBs); EMS code: P023 | Chemical |  | LAB
P023 | ug/g | Polychlor. Biphenyls | X383 | Extr, GC/MS or GC/ECD | Determination of Polychloriated Biphenyls (PCBs) In Soil by GC/ECD. PURPOSE/PRINCIPLE OF METHOD: This document details the requirements and procedures for the analysis of polychlorinated biphenyls (PCBs) in soil by gas chromatography electron capture detector (GC/ECD). A homogenized soil sample is extracted using dichloromethane (DCM). The extract may be cleaned up using column chromatography with 2% Deactivated Florisil. When contamination with sulphur is suspected, further cleanup using copper is required. Analysis is performed using a GC/ECD. Similar methods have been developed for other matrices including waste and transformer oils, water, swabs, florisil tubes, biota. Generally, the sample extraction procedures differ while the GC analysis methods are similar or identical for all matrices. SCOPE: This procedure describes the determination of PCBs in soil and sediments. Other materials such as wood waste may also be analyzed by this method. DETECTION LIMITS: Target Compound Reported Detection Limit (mg/g) Arochlor 1242 0.03, Arochlor 1248 0.03, Arochlor 1254 0.03, Arochlor 1260 0.03, INTERFERENCES: Interferences (e.g. co-eluting peaks on the chromatogram) may be due to sulphur, phthalate esters, or from chlorinated compounds. Suspect samples must be reanalyzed using appropriate cleanup steps. SAMPLE HANDLING & PRESERVATION: Sample Requirements: Whenever possible, samples must completely fill the bottle such that little or no headspace is present to eliminate the loss of volatile hydrocarbon components. (~125g of sample) Sample Containers: Glass jars that have been solvent rinsed with acetone then hexane, then dried. Shipment and Storage: Samples are to be refrigerated (at 4±2°C) during shipping and prior to analysis. Samples suspected to contain high levels of analyte are to be segregated from other samples. Holding Time: Holding time prior to extraction is 14 days. REFERENCE: U.S. EPA Method 8080:Organochlorine Pesticides and PCBs Analysis. | 0.05 | N | N | Y | 4 |  |  | 10/23/1996 11:03:49 PM | 1996-04-12 | 1/9/2013 4:36:43 PM | EMS | RDGIBSON | MassPerMass | mg/kg | Y | (mass conc.) |  | Polychlor. Biphenyls (mass conc.) |  | Polychlorinated biphenylds (PCBs); | Polychlorinated biphenylds (PCBs); EMS code: P023 | Chemical |  | LAB
P023 | ug/g | Polychlor. Biphenyls | X213 | Hex/AcExt,Part,Chrom,GC | Hex/AcExt;Part;Chrom;GC | 0.005 | N | N | Y | 83 |  |  | 10/24/1996 10:24:14 AM | 1996-04-12 | 1/9/2013 4:36:43 PM | EMS | RDGIBSON | MassPerMass | mg/kg | Y | (mass conc.) |  | Polychlor. Biphenyls (mass conc.) |  | Polychlorinated biphenylds (PCBs); | Polychlorinated biphenylds (PCBs); EMS code: P023 | Chemical |  | LAB
P023 | ug/g wet | Polychlor. Biphenyls | X220 | Hex/AcExt,Part,Chrom,GC | Hex/AcExt;Part;Chrom;GC | 0.001 | N | N | Y | 74 |  |  | 10/24/1996 10:01:25 AM | 1996-04-12 | 1/9/2013 4:36:43 PM | EMS | RDGIBSON | MassPerMass | µg/g | Y | (mass conc.) | (wet) | Polychlor. Biphenyls (wet) (mass conc.) |  | Polychlorinated biphenylds (PCBs); | Polychlorinated biphenylds (PCBs); EMS code: P023 | Chemical |  | LAB
P025 | mg/L | Pimaric Acid | F009 | pH 9,Diethyl Ether:HRGC/LRMS | Effluent samples are first adjusted to pH 9 and then extracted with Diethyl Ether. The extracts are rotary evaporated and the resulting residues, containing the Resin Acids and other solvent soluble organics are then redissolved in Methanol, transferred to small glass vials, methylated with Diazomethane then evaporated to dryness under Nitrogen. [Methylation Reaction R-COOH(Resin/Fatty Acids) + CH2N2(Diazomethane) = R-COOCH3(Methyl ester) + N2) The residue is re-dissolved in Methyl t-Butyl Ether (MTBE) and injected on a High Resolution Gas Chromatograph/Low Resolution Mass Spectrometer (HRGC/LRMS) system for analysis. LRMS data is acquired in Total Ion Mode (TIM) and quantitation is performed using selected mass ions. Positive identification of the various Resin Acids is made on the basis of a comparison of absolute retention times to those of the external standards and computer based mass spectral mass/intensity library matching routines. | 0.005 | Y | Y | Y | 198 | 12/27/1996 4:16:00 PM | 5/13/2003 | 4/1/2004 1:08:47 AM | 1996-09-16 | 6/3/1998 6:51:44 PM | MPAK |  | Concentration | mg/L | Y | (fl. conc.) |  | Pimaric Acid (fl. conc.) | 127-27-5 | Organics; | Organics; EMS code: P025 | Chemical |  | LAB
P025 | mg/L | Pimaric Acid | P035 | DCM-DIAZO-FLOR | DCM-DIAZO-FLOR | 0.0005 | Y | Y | Y | 22 |  |  | 10/24/1996 12:44:24 PM | 1996-04-12 | 6/3/1998 6:51:37 PM | EMS |  | Concentration | mg/L | Y | (fl. conc.) |  | Pimaric Acid (fl. conc.) | 127-27-5 | Organics; | Organics; EMS code: P025 | Chemical |  | LAB
P025 | mg/L | Pimaric Acid | X266 | CH2Cl2Ext:PEBrT:GC/ECD | CH2Cl2Ext:PEBrT:GC/ECD | 0.01 | Y | Y | Y | 2 |  |  | 6/7/2024 5:30:18 PM | 1996-04-12 | 6/3/1998 6:51:37 PM | EMS |  | Concentration | mg/L | Y | (fl. conc.) |  | Pimaric Acid (fl. conc.) | 127-27-5 | Organics; | Organics; EMS code: P025 | Chemical |  | LAB
P025 | mg/L | Pimaric Acid | X239 | Gas Chromatography | Gas Chromatography | 0.001 | Y | Y | Y | 12 |  |  | 10/23/1996 11:56:37 PM | 1996-04-12 | 6/3/1998 6:51:37 PM | EMS |  | Concentration | mg/L | Y | (fl. conc.) |  | Pimaric Acid (fl. conc.) | 127-27-5 | Organics; | Organics; EMS code: P025 | Chemical |  | LAB
P025 | ug/g | Pimaric Acid | FS05 | MeOH,DCM extr:HRGC/LRMS | Sediment samples are extracted by shaking with Methanol in a screw cap centrifuge bottle. The extracts are rotary evaporated and the resulting residues, containing the Resin Acids and other solvent soluble organics, are then redissolved in Methanol, transferred to small glass vials, methylated with Diazomethane then evaporated to dryness under Nitrogen. [Methylation Reaction R-COOH(Resin/Fatty Acids) + CH2N2 (Diazomethane) = R-COOCH3 (Methyl ester) + N2] The residue is re-dissolved in Methyl t-Butyl Ether (MTBE) and injected on a High Resolution Gas Chromatograph / Low Resolution Mass Spectrometer (HRGC/LRMS) system for analysis. LRMS data is acquired in Total Ion Mode (TIM) and quantitation is performed using selected mass ions. Positive identification of the various Resin Acids is made on the basis of a comparison of absolute retention times to those of the external standards and computer based mass spectral mass/intensity library matching routines. | 0.2 | Y | Y | Y | 5 | 9/14/2001 3:29:00 PM | 12/3/2002 | 3/27/2004 1:09:21 AM | 1996-09-16 | 6/3/1998 6:51:44 PM | MPAK |  | MassPerMass | mg/kg | Y | (mass conc.) |  | Pimaric Acid (mass conc.) | 127-27-5 | Organics; | Organics; EMS code: P025 | Chemical |  | LAB
P025 | mg/L | Pimaric Acid | E677 | Resins & Fatty Acids are methylated then analyzed by GCMS | EPA 8270E; AB Env. 129; NCASI RA/FA-85.02 (mod) Resins & Fatty Acids are methylated then analyzed by GCMS. Total fatty acids is the sum of arachidic acid; behenic acid; lauric acid; lignoceric acid; linoleic acid; linolenic acid; myristic acid; oleic acid; palmitic acid; and stearic acid. Total resin acids is the sum of abietic acid; 12-chlorodehydroabietic acid; 14-chlorodehydroabietic acid; dehydroabietic acid; 12;14-dichlorodehydroabietic acid; isopimaric acid; levopimaric acid; neoabietic acid; palustric acid; pimaric acid and sandaracopimaric acid. Total resin acids (BC WWQG) is the sum of abietic acid; neoabietic acid; pimaric acid; isopimaric acid; and sandaracopimaric acid. | 0.001 | Y | Y | Y | 12 | 11/28/2022 | 6/15/2024 | 7/8/2024 11:10:02 AM | 2021-12-30 |  | EMS |  | Concentration | mg/L | Y | (fl. conc.) |  | Pimaric Acid (fl. conc.) | 127-27-5 | Organics; | Organics; EMS code: P025 | Chemical |  | LAB
P025 | mg/L | Pimaric Acid | RFAS | Separation of resin and fatty acids (RFAs) in water by GC/MS | Scientific Principle The separation of resin and fatty acids (RFAs) in water samples is performed using gas chromatography/mass spectrometry (GC/MS) in the Selected Ion Monitoring (SIM) mode. The separation is based on compound size; boiling point and polarity. Analyte concentration is calculated using the internal standard method. RFA compounds are extracted from water samples using a liquid-liquid solvent extraction with DCM. Extraction occurs after the sample is brought to a pH of ~ 2 and placed in a rotary tumbler (at approx. 40 rpm) for 1 hour. The resulting extract is transferred to a vial; dried with acidified baked sodium sulfate; methylated with diazomethane and concentrated to 1mL under nitrogen prior to analysis. The resulting extracts are vialed and fortified with labeled internal standard compounds prior to analysis by gas chromatography with a (50%-Phenyl)-methylpolysiloxane (DB-17) column and a mass spectrometer detector in SIM Mode. | 0.006 | Y | Y | Y | 1 | 9/17/2021 7:49:00 PM | 9/17/2021 7:49:00 PM | 2/1/2022 2:38:14 PM | 2022-01-19 |  | EMS |  | Concentration | mg/L | Y | (fl. conc.) |  | Pimaric Acid (fl. conc.) | 127-27-5 | Organics; | Organics; EMS code: P025 | Chemical |  | LAB
P025 | mg/L | Pimaric Acid | P030 | Resin Acids | The water sample is acidified and extracted with dichloro-methane. The raw extract is concentrated and derivatized with diazomethane to produce the corresponding methyl ester derivatives. (Suitable substitute derivatives may be produced by other available techniques.) If required, the extracts are cleaned up by Florisil column chromatography. The derivatives are analyzed by gas chromatography with flame ionization detection. | 0.001 | Y | Y | Y | 481 | 8/26/2014 | 8/15/2017 | 8/15/2017 4:37:09 PM | 1996-04-12 | 6/3/1998 6:51:37 PM | EMS |  | Concentration | mg/L | Y | (fl. conc.) |  | Pimaric Acid (fl. conc.) | 127-27-5 | Organics; | Organics; EMS code: P025 | Chemical |  | LAB
P025 | mg/L | Pimaric Acid | P033 | Resin&Fatty Acids,GC/MS | Resin&Fatty Acids;GC/MS | 0.001 | Y | Y | Y | 28 |  |  | 6/7/2024 5:30:17 PM | 1996-04-12 | 6/3/1998 6:51:37 PM | EMS |  | Concentration | mg/L | Y | (fl. conc.) |  | Pimaric Acid (fl. conc.) | 127-27-5 | Organics; | Organics; EMS code: P025 | Chemical |  | LAB
P028 | mg/L | Xanthates (as ethyl xanthate) | P034 | Xanthates (as etX) by spectrophotometry | Xanthates (as etX) by spectrophotometry. | 0.5 | Y | Y | Y | 66 | 8/20/2019 | 7/25/2024 | 8/3/2024 9:18:08 AM | 2019-09-16 |  | KATFRASE |  | Concentration | mg/L | Y | (fl. conc.) |  | Xanthates (as ethyl xanthate) (fl. conc.) | 151-01-9 | Organics; | Organics; EMS code: P028 | Chemical |  | LAB
P040 | pg/L | 5Chl2BenDiox1;2;3;78 | X253 | Tol.-Ac/Cl2CH2 ext:GC/MS | C-13 labelled PCDD and PCDF congeners are added to a 1L sample of pulp effluent, which is separated into solid and liquid phases by filtration through a glass fibre filter. The solid phase is aire-dried overnight before undergoing a Soxhlet extraction with toluene. The liquid phase is extracted with dichloromethane in a separatory funnel. The two extracts are combined, concentrated, and then passed through a series of columns which remove the bulk of the organic matrix co-extracted with PCDD/PCDF. One resulting fraction is concentrated to a known small volume for analysis by HRGC/HRMS/DS. | 1e-05 | Y | Y | Y | 154 |  |  | 6/25/2012 2:16:03 PM | 1996-04-12 | 6/3/1998 6:51:37 PM | EMS |  | Concentration | pg/L | Y | (fl. conc.) |  | 5Chl2BenDiox1;2;3;78 (fl. conc.) | 40321-76-4 | Dioxins & Furans; | Dioxins & Furans; EMS code: P040 | Chemical |  | SURROGATE_RESULT
P040 | pg/L | 5Chl2BenDiox1;2;3;78 | F114 | Isotope dil,HRGC/HRMS | EPA Method 1613: Developed for isomer specific determination of the 2,3,7,8-substituted, tetra-through octa-chlorinated, dibenzo-p-dioxins and dibenzofurans in aqueous, solid and tissue matrices by isotope dilution, high resolution capillary column gas chromatography (HRGC)/high resolution mass spectrometry (HRMS) |  | Y | Y | Y | 18 |  |  | 2/21/2024 1:00:50 PM | 2001-05-23 | 5/24/2001 3:35:10 PM | NPEPPIN | NPEPPIN | Concentration | pg/L | Y | (fl. conc.) |  | 5Chl2BenDiox1;2;3;78 (fl. conc.) | 40321-76-4 | Dioxins & Furans; | Dioxins & Furans; EMS code: P040 | Chemical |  | LAB
P040 | pg/L | 5Chl2BenDiox1;2;3;78 | F102 | Env.Canada(1992),DioxinsProc,Pulp Effluents | Environment Canada (1992);Dioxins Procedures for Pulp Effluents |  | Y | Y | Y | 6 | 5/1/2001 7:37:00 AM | 5/1/2001 7:37:00 AM | 11/9/2011 3:41:39 PM | 2000-10-25 |  | NPEPPIN |  | Concentration | pg/L | Y | (fl. conc.) |  | 5Chl2BenDiox1;2;3;78 (fl. conc.) | 40321-76-4 | Dioxins & Furans; | Dioxins & Furans; EMS code: P040 | Chemical |  | LAB
P040 | pg/g wet | 5Chl2BenDiox1;2;3;78 | F119 | Solids,Isotope dil,HRGC/HRMS | EPA Method 1613: for Sediments and Fish tissue developed for isomer specific determination of the 2,3,7,8-substituted, tetra-through octa-chlorinated, dibenzo-p-dioxins and dibenzofurans in aqueous, solid and tissue matrices by isotope dilution, high resolution capillary column gas chromatography (HRGC)/high resolution mass spectrometry (HRMS) |  | Y | Y | Y | 5 | 6/16/2000 6:33:00 AM | 10/4/2000 2:53:00 PM | 6/18/2002 3:50:03 PM | 2001-07-03 | 10/4/2001 3:20:45 PM | NPEPPIN | NPEPPIN | MassPerMass | pg/g | Y | (mass conc.) | (wet) | 5Chl2BenDiox1;2;3;78 (wet) (mass conc.) | 40321-76-4 | Dioxins & Furans; | Dioxins & Furans; EMS code: P040 | Chemical |  | LAB
P041 | pg/L | 5Chldibenzodioxin | F102 | Env.Canada(1992),DioxinsProc,Pulp Effluents | Environment Canada (1992);Dioxins Procedures for Pulp Effluents |  | Y | Y | Y | 6 | 5/1/2001 7:37:00 AM | 5/1/2001 7:37:00 AM | 11/9/2011 3:41:39 PM | 2000-10-25 |  | NPEPPIN |  | Concentration | pg/L | Y | (fl. conc.) |  | 5Chldibenzodioxin (fl. conc.) | 36088-22-9 | Dioxins & Furans; | Dioxins & Furans; EMS code: P041 | Chemical |  | LAB
P041 | pg/L | 5Chldibenzodioxin | X253 | Tol.-Ac/Cl2CH2 ext:GC/MS | C-13 labelled PCDD and PCDF congeners are added to a 1L sample of pulp effluent, which is separated into solid and liquid phases by filtration through a glass fibre filter. The solid phase is aire-dried overnight before undergoing a Soxhlet extraction with toluene. The liquid phase is extracted with dichloromethane in a separatory funnel. The two extracts are combined, concentrated, and then passed through a series of columns which remove the bulk of the organic matrix co-extracted with PCDD/PCDF. One resulting fraction is concentrated to a known small volume for analysis by HRGC/HRMS/DS. | 1 | Y | Y | Y | 146 |  |  | 6/25/2012 2:16:03 PM | 1996-04-12 | 6/3/1998 6:51:45 PM | EMS |  | Concentration | pg/L | Y | (fl. conc.) |  | 5Chldibenzodioxin (fl. conc.) | 36088-22-9 | Dioxins & Furans; | Dioxins & Furans; EMS code: P041 | Chemical |  | SURROGATE_RESULT
P041 | pg/g wet | 5Chldibenzodioxin | F119 | Solids,Isotope dil,HRGC/HRMS | EPA Method 1613: for Sediments and Fish tissue developed for isomer specific determination of the 2,3,7,8-substituted, tetra-through octa-chlorinated, dibenzo-p-dioxins and dibenzofurans in aqueous, solid and tissue matrices by isotope dilution, high resolution capillary column gas chromatography (HRGC)/high resolution mass spectrometry (HRMS) |  | Y | Y | Y | 5 | 6/16/2000 6:33:00 AM | 10/4/2000 2:53:00 PM | 6/18/2002 3:50:04 PM | 2001-10-01 | 10/4/2001 3:20:45 PM | NPEPPIN | NPEPPIN | MassPerMass | pg/g | Y | (mass conc.) | (wet) | 5Chldibenzodioxin (wet) (mass conc.) | 36088-22-9 | Dioxins & Furans; | Dioxins & Furans; EMS code: P041 | Chemical |  | LAB
P041 | pg/L | 5Chldibenzodioxin | F114 | Isotope dil,HRGC/HRMS | EPA Method 1613: Developed for isomer specific determination of the 2,3,7,8-substituted, tetra-through octa-chlorinated, dibenzo-p-dioxins and dibenzofurans in aqueous, solid and tissue matrices by isotope dilution, high resolution capillary column gas chromatography (HRGC)/high resolution mass spectrometry (HRMS) |  | Y | Y | Y | 18 |  |  | 2/21/2024 1:00:50 PM | 2001-05-23 | 5/24/2001 3:35:10 PM | NPEPPIN | NPEPPIN | Concentration | pg/L | Y | (fl. conc.) |  | 5Chldibenzodioxin (fl. conc.) | 36088-22-9 | Dioxins & Furans; | Dioxins & Furans; EMS code: P041 | Chemical |  | LAB
P042 | pg/L | 5Chldibenzofuran | X253 | Tol.-Ac/Cl2CH2 ext:GC/MS | C-13 labelled PCDD and PCDF congeners are added to a 1L sample of pulp effluent, which is separated into solid and liquid phases by filtration through a glass fibre filter. The solid phase is aire-dried overnight before undergoing a Soxhlet extraction with toluene. The liquid phase is extracted with dichloromethane in a separatory funnel. The two extracts are combined, concentrated, and then passed through a series of columns which remove the bulk of the organic matrix co-extracted with PCDD/PCDF. One resulting fraction is concentrated to a known small volume for analysis by HRGC/HRMS/DS. | 1 | Y | Y | Y | 146 |  |  | 6/25/2012 2:16:03 PM | 1996-04-12 | 6/3/1998 6:51:37 PM | EMS |  | Concentration | pg/L | Y | (fl. conc.) |  | 5Chldibenzofuran (fl. conc.) | 30402-15-4 | Dioxins & Furans; | Dioxins & Furans; EMS code: P042 | Chemical |  | SURROGATE_RESULT
P042 | pg/L | 5Chldibenzofuran | F102 | Env.Canada(1992),DioxinsProc,Pulp Effluents | Environment Canada (1992);Dioxins Procedures for Pulp Effluents |  | Y | Y | Y | 6 | 5/1/2001 7:37:00 AM | 5/1/2001 7:37:00 AM | 11/9/2011 3:41:39 PM | 2000-10-25 |  | NPEPPIN |  | Concentration | pg/L | Y | (fl. conc.) |  | 5Chldibenzofuran (fl. conc.) | 30402-15-4 | Dioxins & Furans; | Dioxins & Furans; EMS code: P042 | Chemical |  | LAB
P042 | pg/g wet | 5Chldibenzofuran | F119 | Solids,Isotope dil,HRGC/HRMS | EPA Method 1613: for Sediments and Fish tissue developed for isomer specific determination of the 2,3,7,8-substituted, tetra-through octa-chlorinated, dibenzo-p-dioxins and dibenzofurans in aqueous, solid and tissue matrices by isotope dilution, high resolution capillary column gas chromatography (HRGC)/high resolution mass spectrometry (HRMS) |  | Y | Y | Y | 5 | 6/16/2000 6:33:00 AM | 10/4/2000 2:53:00 PM | 6/18/2002 3:50:04 PM | 2001-10-01 | 10/4/2001 3:20:45 PM | NPEPPIN | NPEPPIN | MassPerMass | pg/g | Y | (mass conc.) | (wet) | 5Chldibenzofuran (wet) (mass conc.) | 30402-15-4 | Dioxins & Furans; | Dioxins & Furans; EMS code: P042 | Chemical |  | LAB
P042 | pg/L | 5Chldibenzofuran | F114 | Isotope dil,HRGC/HRMS | EPA Method 1613: Developed for isomer specific determination of the 2,3,7,8-substituted, tetra-through octa-chlorinated, dibenzo-p-dioxins and dibenzofurans in aqueous, solid and tissue matrices by isotope dilution, high resolution capillary column gas chromatography (HRGC)/high resolution mass spectrometry (HRMS) |  | Y | Y | Y | 18 |  |  | 2/21/2024 1:00:50 PM | 2001-05-23 | 5/24/2001 3:35:10 PM | NPEPPIN | NPEPPIN | Concentration | pg/L | Y | (fl. conc.) |  | 5Chldibenzofuran (fl. conc.) | 30402-15-4 | Dioxins & Furans; | Dioxins & Furans; EMS code: P042 | Chemical |  | LAB
P043 | pg/g wet | 5Chl2BFuran1;2;3;7;8 | F119 | Solids,Isotope dil,HRGC/HRMS | EPA Method 1613: for Sediments and Fish tissue developed for isomer specific determination of the 2,3,7,8-substituted, tetra-through octa-chlorinated, dibenzo-p-dioxins and dibenzofurans in aqueous, solid and tissue matrices by isotope dilution, high resolution capillary column gas chromatography (HRGC)/high resolution mass spectrometry (HRMS) |  | Y | Y | Y | 5 | 6/16/2000 6:33:00 AM | 10/4/2000 2:53:00 PM | 6/18/2002 3:50:04 PM | 2001-07-03 | 10/4/2001 3:20:45 PM | NPEPPIN | NPEPPIN | MassPerMass | pg/g | Y | (mass conc.) | (wet) | 5Chl2BFuran1;2;3;7;8 (wet) (mass conc.) | 57117-41-6 | Dioxins & Furans; | Dioxins & Furans; EMS code: P043 | Chemical |  | LAB
P043 | pg/L | 5Chl2BFuran1;2;3;7;8 | X253 | Tol.-Ac/Cl2CH2 ext:GC/MS | C-13 labelled PCDD and PCDF congeners are added to a 1L sample of pulp effluent, which is separated into solid and liquid phases by filtration through a glass fibre filter. The solid phase is aire-dried overnight before undergoing a Soxhlet extraction with toluene. The liquid phase is extracted with dichloromethane in a separatory funnel. The two extracts are combined, concentrated, and then passed through a series of columns which remove the bulk of the organic matrix co-extracted with PCDD/PCDF. One resulting fraction is concentrated to a known small volume for analysis by HRGC/HRMS/DS. | 1 | Y | Y | Y | 154 |  |  | 6/25/2012 2:16:03 PM | 1996-04-12 | 6/3/1998 6:51:37 PM | EMS |  | Concentration | pg/L | Y | (fl. conc.) |  | 5Chl2BFuran1;2;3;7;8 (fl. conc.) | 57117-41-6 | Dioxins & Furans; | Dioxins & Furans; EMS code: P043 | Chemical |  | SURROGATE_RESULT
P043 | pg/L | 5Chl2BFuran1;2;3;7;8 | F114 | Isotope dil,HRGC/HRMS | EPA Method 1613: Developed for isomer specific determination of the 2,3,7,8-substituted, tetra-through octa-chlorinated, dibenzo-p-dioxins and dibenzofurans in aqueous, solid and tissue matrices by isotope dilution, high resolution capillary column gas chromatography (HRGC)/high resolution mass spectrometry (HRMS) |  | Y | Y | Y | 18 |  |  | 2/21/2024 1:00:50 PM | 2001-05-23 | 5/24/2001 3:35:10 PM | NPEPPIN | NPEPPIN | Concentration | pg/L | Y | (fl. conc.) |  | 5Chl2BFuran1;2;3;7;8 (fl. conc.) | 57117-41-6 | Dioxins & Furans; | Dioxins & Furans; EMS code: P043 | Chemical |  | LAB
P043 | pg/L | 5Chl2BFuran1;2;3;7;8 | F102 | Env.Canada(1992),DioxinsProc,Pulp Effluents | Environment Canada (1992);Dioxins Procedures for Pulp Effluents |  | Y | Y | Y | 6 | 5/1/2001 7:37:00 AM | 5/1/2001 7:37:00 AM | 11/9/2011 3:41:39 PM | 2000-10-25 |  | NPEPPIN |  | Concentration | pg/L | Y | (fl. conc.) |  | 5Chl2BFuran1;2;3;7;8 (fl. conc.) | 57117-41-6 | Dioxins & Furans; | Dioxins & Furans; EMS code: P043 | Chemical |  | LAB
P044 | pg/L | 5Chl2BFuran2;3;4;7;8 | F102 | Env.Canada(1992),DioxinsProc,Pulp Effluents | Environment Canada (1992);Dioxins Procedures for Pulp Effluents |  | Y | Y | Y | 6 | 5/1/2001 7:37:00 AM | 5/1/2001 7:37:00 AM | 11/9/2011 3:41:39 PM | 2000-10-25 |  | NPEPPIN |  | Concentration | pg/L | Y | (fl. conc.) |  | 5Chl2BFuran2;3;4;7;8 (fl. conc.) | 57117-31-4 | Dioxins & Furans; | Dioxins & Furans; EMS code: P044 | Chemical |  | LAB
P044 | pg/g wet | 5Chl2BFuran2;3;4;7;8 | F119 | Solids,Isotope dil,HRGC/HRMS | EPA Method 1613: for Sediments and Fish tissue developed for isomer specific determination of the 2,3,7,8-substituted, tetra-through octa-chlorinated, dibenzo-p-dioxins and dibenzofurans in aqueous, solid and tissue matrices by isotope dilution, high resolution capillary column gas chromatography (HRGC)/high resolution mass spectrometry (HRMS) |  | Y | Y | Y | 5 | 6/16/2000 6:33:00 AM | 10/4/2000 2:53:00 PM | 6/18/2002 3:50:05 PM | 2001-07-03 | 10/4/2001 3:20:45 PM | NPEPPIN | NPEPPIN | MassPerMass | pg/g | Y | (mass conc.) | (wet) | 5Chl2BFuran2;3;4;7;8 (wet) (mass conc.) | 57117-31-4 | Dioxins & Furans; | Dioxins & Furans; EMS code: P044 | Chemical |  | LAB
P044 | pg/L | 5Chl2BFuran2;3;4;7;8 | F114 | Isotope dil,HRGC/HRMS | EPA Method 1613: Developed for isomer specific determination of the 2,3,7,8-substituted, tetra-through octa-chlorinated, dibenzo-p-dioxins and dibenzofurans in aqueous, solid and tissue matrices by isotope dilution, high resolution capillary column gas chromatography (HRGC)/high resolution mass spectrometry (HRMS) |  | Y | Y | Y | 18 |  |  | 2/21/2024 1:00:50 PM | 2001-05-23 | 5/24/2001 3:35:10 PM | NPEPPIN | NPEPPIN | Concentration | pg/L | Y | (fl. conc.) |  | 5Chl2BFuran2;3;4;7;8 (fl. conc.) | 57117-31-4 | Dioxins & Furans; | Dioxins & Furans; EMS code: P044 | Chemical |  | LAB
P044 | pg/L | 5Chl2BFuran2;3;4;7;8 | X253 | Tol.-Ac/Cl2CH2 ext:GC/MS | C-13 labelled PCDD and PCDF congeners are added to a 1L sample of pulp effluent, which is separated into solid and liquid phases by filtration through a glass fibre filter. The solid phase is aire-dried overnight before undergoing a Soxhlet extraction with toluene. The liquid phase is extracted with dichloromethane in a separatory funnel. The two extracts are combined, concentrated, and then passed through a series of columns which remove the bulk of the organic matrix co-extracted with PCDD/PCDF. One resulting fraction is concentrated to a known small volume for analysis by HRGC/HRMS/DS. | 1 | Y | Y | Y | 154 |  |  | 6/25/2012 2:16:03 PM | 1996-04-12 | 6/3/1998 6:51:45 PM | EMS |  | Concentration | pg/L | Y | (fl. conc.) |  | 5Chl2BFuran2;3;4;7;8 (fl. conc.) | 57117-31-4 | Dioxins & Furans; | Dioxins & Furans; EMS code: P044 | Chemical |  | SURROGATE_RESULT
P045 | ug/g | Prometon | ZN03 | DCM+Ac.Extr:GC/MS | DCM+Ac.Extr:GC/MS EPA 8270E (mod) | 0.05 | Y | Y | Y | 16 | 1/25/2022 | 11/2/2022 4:41:00 AM | 1/7/2023 3:12:11 PM | 2003-06-17 |  | UMULLER |  | MassPerMass | mg/kg | Y | (mass conc.) |  | Prometon (mass conc.) | 1610-18-0 | Pesticides; | Pesticides; EMS code: P045 | Chemical |  | LAB
P045 | mg/L | Prometon | E660 | EPA 8270E (mod) | EPA 8270E (mod) Miscellaneous Pesticides by GC-MS | 0.0001 | Y | Y | Y | 68 | 12/7/2021 | 9/12/2022 | 12/8/2022 12:57:05 PM | 2021-12-09 | 12/9/2021 3:22:12 PM | EMS | EMS | Concentration | mg/L | Y | (fl. conc.) |  | Prometon (fl. conc.) | 1610-18-0 | Pesticides; | Pesticides; EMS code: P045 | Chemical |  | LAB
P046 | mg/L | Propanil | P015 | Phenoxyacid Herbicides by LC-MS/MS | Phenoxyacid Herbicides by LC-MS/MS | 1e-05 | Y | Y | Y | 68 | 12/6/2021 | 9/9/2022 | 12/8/2022 12:57:05 PM | 2021-12-09 | 12/9/2021 3:22:12 PM | EMS | EMS | Concentration | mg/L | Y | (fl. conc.) |  | Propanil (fl. conc.) | 709-98-8 | Pesticides; | Pesticides; EMS code: P046 | Chemical |  | LAB
P046 | mg/L | Propanil | ZN01 | DCM+Ac.Extr:GC/MS | DCM+Ac.Extr:GC/MS | 0.0005 | Y | Y | Y | 1 | 7/29/2005 7:10:00 AM | 7/29/2005 7:10:00 AM | 6/2/2009 12:01:09 PM | 2003-06-17 |  | UMULLER |  | Concentration | mg/L | Y | (fl. conc.) |  | Propanil (fl. conc.) | 709-98-8 | Pesticides; | Pesticides; EMS code: P046 | Chemical |  | LAB
P048 | mg/L | Propiconazole | E751 | E751D Pesticides in Water by LC-MS/MS (POS1 List) | E751D Pesticides in Water by LC-MS/MS (POS1 List) (MOECC E3501) | 5e-05 | Y | Y | Y | 39 | 12/9/2021 | 12/31/2021 | 2/15/2022 9:50:05 AM | 2021-12-14 |  | EMS |  | Concentration | mg/L | Y | (fl. conc.) |  | Propiconazole (fl. conc.) | 60207-90-1 | Pesticides; | Pesticides; EMS code: P048 | Chemical |  | LAB
P048 | mg/L | Propiconazole | ZN01 | DCM+Ac.Extr:GC/MS | DCM+Ac.Extr:GC/MS | 0.0005 | Y | Y | Y | 1 | 7/29/2005 7:10:00 AM | 7/29/2005 7:10:00 AM | 6/2/2009 12:01:09 PM | 2003-06-17 |  | UMULLER |  | Concentration | mg/L | Y | (fl. conc.) |  | Propiconazole (fl. conc.) | 60207-90-1 | Pesticides; | Pesticides; EMS code: P048 | Chemical |  | LAB
P051 | mg/L | Permethrin; cis | ZN01 | DCM+Ac.Extr:GC/MS | DCM+Ac.Extr:GC/MS | 0.0002 | Y | Y | Y | 1 | 7/28/2005 7:03:00 AM | 7/28/2005 7:03:00 AM | 6/2/2009 12:01:09 PM | 1996-04-12 | 11/16/1999 4:06:27 PM | EMS | NPEPPIN | Concentration | mg/L | Y | (fl. conc.) |  | Permethrin; cis (fl. conc.) |  | Pesticides; | Pesticides; EMS code: P051 | Chemical |  | LAB
P051 | ug/g | Permethrin; cis | ZN03 | DCM+Ac.Extr:GC/MS | DCM+Ac.Extr:GC/MS EPA 8270E (mod) | 0.02 | Y | Y | Y | 2 | 11/2/2022 12:53:00 AM | 11/2/2022 4:41:00 AM | 1/7/2023 3:12:11 PM | 1996-04-12 | 11/16/1999 2:27:40 PM | EMS | NPEPPIN | MassPerMass | mg/kg | Y | (mass conc.) |  | Permethrin; cis (mass conc.) |  | Pesticides; | Pesticides; EMS code: P051 | Chemical |  | LAB
P052 | mg/L | Permethrin; trans | ZN01 | DCM+Ac.Extr:GC/MS | DCM+Ac.Extr:GC/MS | 0.0001 | Y | Y | Y | 1 | 7/28/2005 7:03:00 AM | 7/28/2005 7:03:00 AM | 6/2/2009 12:01:09 PM | 1996-04-12 | 11/16/1999 4:06:27 PM | EMS | NPEPPIN | Concentration | mg/L | Y | (fl. conc.) |  | Permethrin; trans (fl. conc.) |  | Pesticides; | Pesticides; EMS code: P052 | Chemical |  | LAB
P053 | mg/L | Pirimiphos-methyl | ZN01 | DCM+Ac.Extr:GC/MS | DCM+Ac.Extr:GC/MS | 0.0002 | Y | Y | Y | 28 | 7/29/2005 7:10:00 AM | 3/29/2012 10:15:00 PM | 5/3/2012 9:27:52 AM | 1996-04-12 | 11/16/1999 4:06:27 PM | EMS | NPEPPIN | Concentration | mg/L | Y | (fl. conc.) |  | Pirimiphos-methyl (fl. conc.) | 29232-93-7 | Pesticides; | Pesticides; EMS code: P053 | Chemical |  | LAB
P054 | mg/L | Procymidone | ZN01 | DCM+Ac.Extr:GC/MS | DCM+Ac.Extr:GC/MS | 0.0002 | Y | Y | Y | 28 | 7/28/2005 7:03:00 AM | 3/29/2012 10:15:00 PM | 5/3/2012 9:27:52 AM | 1996-04-12 | 11/16/1999 4:06:27 PM | EMS | NPEPPIN | Concentration | mg/L | Y | (fl. conc.) |  | Procymidone (fl. conc.) |  | Pesticides; | Pesticides; EMS code: P054 | Chemical |  | LAB
P055 | mg/L | Pronamide | ZN01 | DCM+Ac.Extr:GC/MS | DCM+Ac.Extr:GC/MS | 0.0002 | Y | Y | Y | 27 | 11/1/2008 | 3/29/2012 10:15:00 PM | 5/3/2012 9:27:52 AM | 1996-04-12 | 11/16/1999 4:06:28 PM | EMS | NPEPPIN | Concentration | mg/L | Y | (fl. conc.) |  | Pronamide (fl. conc.) |  | Pesticides; | Pesticides; EMS code: P055 | Chemical |  | LAB
P056 | mg/L | Paraoxon | ZN01 | DCM+Ac.Extr:GC/MS | DCM+Ac.Extr:GC/MS | 0.0005 | Y | Y | Y | 1 | 7/29/2005 7:10:00 AM | 7/29/2005 7:10:00 AM | 6/2/2009 12:01:10 PM | 2003-06-16 |  | UMULLER |  | Concentration | mg/L | Y | (fl. conc.) |  | Paraoxon (fl. conc.) |  | Pesticides; | Pesticides; EMS code: P056 | Chemical |  | LAB
P057 | mg/L | Pebulate | ZN01 | DCM+Ac.Extr:GC/MS | DCM+Ac.Extr:GC/MS | 0.0005 | Y | Y | Y | 1 | 7/29/2005 7:10:00 AM | 7/29/2005 7:10:00 AM | 6/2/2009 12:01:10 PM | 2003-06-16 |  | UMULLER |  | Concentration | mg/L | Y | (fl. conc.) |  | Pebulate (fl. conc.) | 1114-71-2 | Pesticides; | Pesticides; EMS code: P057 | Chemical |  | LAB
P060 | mg/L | Pirimiphos Ethyl | ZN01 | DCM+Ac.Extr:GC/MS | DCM+Ac.Extr:GC/MS | 0.0005 | Y | Y | Y | 28 | 7/29/2005 7:10:00 AM | 3/29/2012 10:15:00 PM | 5/3/2012 9:27:52 AM | 1996-04-12 | 11/16/1999 4:06:28 PM | EMS | NPEPPIN | Concentration | mg/L | Y | (fl. conc.) |  | Pirimiphos Ethyl (fl. conc.) |  | Pesticides; | Pesticides; EMS code: P060 | Chemical |  | LAB
P061 | mg/L | Profenofos | ZN01 | DCM+Ac.Extr:GC/MS | DCM+Ac.Extr:GC/MS | 0.0005 | Y | Y | Y | 1 | 7/29/2005 7:10:00 AM | 7/29/2005 7:10:00 AM | 6/2/2009 12:01:10 PM | 1996-04-12 | 11/16/1999 4:06:28 PM | EMS | NPEPPIN | Concentration | mg/L | Y | (fl. conc.) |  | Profenofos (fl. conc.) |  | Pesticides; | Pesticides; EMS code: P061 | Chemical |  | LAB
P062 | mg/L | Pyrazophos | ZN01 | DCM+Ac.Extr:GC/MS | DCM+Ac.Extr:GC/MS | 0.0001 | Y | Y | Y | 1 | 7/29/2005 7:10:00 AM | 7/29/2005 7:10:00 AM | 6/2/2009 12:01:10 PM | 1996-04-12 | 11/16/1999 4:06:28 PM | EMS | NPEPPIN | Concentration | mg/L | Y | (fl. conc.) |  | Pyrazophos (fl. conc.) |  | Pesticides; | Pesticides; EMS code: P062 | Chemical |  | LAB
P063 | mg/L | Paraquat | F123 | ph10,Solid Phase Extr. HPLC-UV | Primary Reference - USEPA Method 549.1 "Determination of Diquat and Paraquat in Drinking Water by Liquid-Solid Extraction and HPLC-UV Detection". Approximately 200mL of water is adjusted to pH 10.5. Target organics are extracted using a prepared C18 solid sorbant cartridge by reverse phase/Ion Pair SPE. Compounds are then extracted from the cartridge by eluting with an acidic aqueous solvent. The eluate is injected into a HPLC for separation and quantification is achieved by measuring absorbance at 308nm for Diquat and 258nm for Paraquat. | 0.01 | Y | Y | Y | 3 | 10/2/2002 1:55:00 PM | 10/2/2002 1:55:00 PM | 3/24/2004 1:04:15 AM | 2002-11-05 |  | UMULLER |  | Concentration | mg/L | Y | (fl. conc.) |  | Paraquat (fl. conc.) | 1910-42-5 | Pesticides; | Pesticides; EMS code: P063 | Chemical |  | LAB
P100 | mg/L | Propargite | ZN01 | DCM+Ac.Extr:GC/MS | DCM+Ac.Extr:GC/MS | 0.5 | Y | Y | Y | 1 | 7/29/2005 7:10:00 AM | 7/29/2005 7:10:00 AM | 6/2/2009 12:01:10 PM | 2009-02-20 |  | JCHUDAK |  | Concentration | mg/L | Y | (fl. conc.) |  | Propargite (fl. conc.) | 2312-35-8 | Pesticides; | Pesticides; EMS code: P100 | Chemical |  | LAB
P102 | mg/L | Pyridaben | ZN01 | DCM+Ac.Extr:GC/MS | DCM+Ac.Extr:GC/MS | 0.5 | Y | Y | Y | 1 | 7/29/2005 7:10:00 AM | 7/29/2005 7:10:00 AM | 6/2/2009 12:01:10 PM | 2009-02-20 |  | JCHUDAK |  | Concentration | mg/L | Y | (fl. conc.) |  | Pyridaben (fl. conc.) |  | Pesticides; | Pesticides; EMS code: P102 | Chemical |  | LAB
P103 | mg/L | Propachlor | ZN01 | DCM+Ac.Extr:GC/MS | DCM+Ac.Extr:GC/MS | 0.1 | Y | Y | Y | 1 | 7/28/2005 7:03:00 AM | 7/28/2005 7:03:00 AM | 6/2/2009 12:01:10 PM | 2009-02-20 |  | JCHUDAK |  | Concentration | mg/L | Y | (fl. conc.) |  | Propachlor (fl. conc.) | 1918-16-7 | Pesticides; | Pesticides; EMS code: P103 | Chemical |  | LAB
P104 | mg/L | Propyzamide | ZN01 | DCM+Ac.Extr:GC/MS | DCM+Ac.Extr:GC/MS | 0.5 | Y | Y | Y | 1 | 7/28/2005 7:03:00 AM | 7/28/2005 7:03:00 AM | 6/2/2009 12:01:11 PM | 2009-02-20 |  | JCHUDAK |  | Concentration | mg/L | Y | (fl. conc.) |  | Propyzamide (fl. conc.) | 23950-58-5 | Pesticides; | Pesticides; EMS code: P104 | Chemical |  | LAB
P105 | mg/L | Promecarb | ZN01 | DCM+Ac.Extr:GC/MS | DCM+Ac.Extr:GC/MS | 0.02 | Y | Y | Y | 1 | 7/29/2005 7:10:00 AM | 7/29/2005 7:10:00 AM | 6/2/2009 12:01:11 PM | 2009-02-20 |  | JCHUDAK |  | Concentration | mg/L | Y | (fl. conc.) |  | Promecarb (fl. conc.) |  | Pesticides; | Pesticides; EMS code: P105 | Chemical |  | LAB
P2-O5 | % | Phosphorus (as P2O5) | ICP06 | ALS Minerals ME-ICP06 | ALS Minerals ME-ICP06. A prepared sample is fused with a lithium tetraborate/lithium metaborate flux and poured into a graphite crucible and fused in a furnace at 1000°C. The resulting melt is then cooled and dissolved in 100 mL of 4% HNO3/2% HCl followed by ICP-AES. | 0.01 | Y | Y | Y | 0 |  |  |  | 2024-01-29 |  | EMS |  | DimensionlessRatio | % | N | (ratio) |  | Phosphorus (as P2O5) (ratio) |  | Nutrients; | Nutrients; EMS code: P2-O5 | Chemical |  | LAB
P2MM | mg/L | PAHs; total (P2MMP) P2 Marine Monitoring Permit | HEGC | Hexane extraction: GC/MS | Hexane extraction: GC/MS | 3e-05 | Y | Y | Y | 9 | 5/17/2024 | 7/28/2024 | 8/6/2024 3:27:04 PM | 2021-03-15 |  | RDGIBSON |  | Concentration | mg/L | Y | (fl. conc.) |  | PAHs; total (P2MMP) P2 Marine Monitoring Permit (fl. conc.) |  | Polycyclic Aromatic Hydrocarbons (PAHs); | Polycyclic Aromatic Hydrocarbons (PAHs); EMS code: P2MM | Chemical | Total | LAB
PA01 | mg/L | Acenaphthene | X262 | CH2Cl2 Ext:Silica:GC/MS | The sample is placed in a separatory funnel, the pH of the sample is adjusted to alkaline and then acidic conditions, and extracted each time with dichloromethane to selectively remove compounds. The final extracts (basic and acidic) can be combined or run separately. The extract is cleaned by silica gel column chromatography, then analyzed by GC/MS. | 0.001 | Y | Y | Y | 124 | 6/12/2000 | 1/13/2015 | 6/7/2024 5:30:18 PM | 1996-04-12 | 6/3/1998 6:51:37 PM | EMS |  | Concentration | mg/L | Y | (fl. conc.) |  | Acenaphthene (fl. conc.) | 83-32-9 | Polycyclic Aromatic Hydrocarbons (PAHs); | Polycyclic Aromatic Hydrocarbons (PAHs); EMS code: PA01 | Chemical |  | LAB
PA01 | mg/L | Acenaphthene | HEGC | Hexane extraction: GC/MS | Hexane extraction: GC/MS | 1e-05 | Y | Y | Y | 905 | 6/8/2015 | 8/15/2024 | 8/16/2024 1:36:03 PM | 2015-07-07 | 8/15/2019 11:20:10 AM | RDGIBSON | KATFRASE | Concentration | mg/L | Y | (fl. conc.) |  | Acenaphthene (fl. conc.) | 83-32-9 | Polycyclic Aromatic Hydrocarbons (PAHs); | Polycyclic Aromatic Hydrocarbons (PAHs); EMS code: PA01 | Chemical |  | LAB
PA01 | mg/L | Acenaphthene | F010 | Extract disks,DCM extr:HRGC/LRMS | Water samples are spiked with a deuterated PAH Recovery Standards mix just prior to extraction with Empore extraction disks. Disks are eluted with Dichloromethane (DCM) and the extract concentrated. Deuterated PAH Internal Standards are added and the extract analyzed using High Resolution Gas Chromatography / Low Resolution Mass Spectrometry (HRGC/LRMS). MS data is acquired in Total Ion Mode (TIM) and quantitation is by multilevel Internal Standards. Recovery efficiency is calculated and reported but results are not recovery corrected. | 0.0002 | Y | Y | Y | 86 | 3/27/1997 4:25:00 PM | 3/11/2003 | 4/1/2004 1:13:26 AM | 1996-09-16 | 6/3/1998 6:51:44 PM | MPAK |  | Concentration | mg/L | Y | (fl. conc.) |  | Acenaphthene (fl. conc.) | 83-32-9 | Polycyclic Aromatic Hydrocarbons (PAHs); | Polycyclic Aromatic Hydrocarbons (PAHs); EMS code: PA01 | Chemical |  | LAB
PA01 | ug/g | Acenaphthene | MP02 | GC/MS | GC/MS | 0.001 | Y | Y | Y | 523 | 3/28/2002 | 9/23/2021 3:24:00 AM | 10/18/2021 9:46:09 AM | 1996-04-12 | 6/3/1998 6:51:37 PM | EMS |  | MassPerMass | mg/kg | Y | (mass conc.) |  | Acenaphthene (mass conc.) | 83-32-9 | Polycyclic Aromatic Hydrocarbons (PAHs); | Polycyclic Aromatic Hydrocarbons (PAHs); EMS code: PA01 | Chemical |  | LAB
PA01 | ug/g wet | Acenaphthene | X221 | Mod.Proc21.001(AOAC,'84) | Mod.Proc21.001(AOAC,'84) | 0.004 | Y | Y | Y | 73 |  |  | 10/24/1996 10:01:25 AM | 1996-04-12 | 6/3/1998 6:51:42 PM | EMS |  | MassPerMass | µg/g | Y | (mass conc.) | (wet) | Acenaphthene (wet) (mass conc.) | 83-32-9 | Polycyclic Aromatic Hydrocarbons (PAHs); | Polycyclic Aromatic Hydrocarbons (PAHs); EMS code: PA01 | Chemical |  | LAB
PA01 | ug/m3 | Acenaphthene | PAH1 | PUF:GC/MS | To understand the extent of human exposure to B[a]P and other PAHs, a reliable sampling and analytical method has been established which involves a combination quartz fibre filter/absorbent cartridge with subsequent extraction and analysis by gas chromatography (GC) with mass spectrometry (MS) detection (GC/MS). The analytical methods are a modification of EPA Test Method 625, Methods for Organic Chemical Analysis of Municipal and Industrial Wastewater, and Methods 8270, Test Methods for Evaluation of Solid Waste. PAH1 utilizes PUF as an adsorbant; PAH2 utilizes XAD-2 as an adsorbant. Note that isotopic surrogates may be added and are reported as % recovery. | 5e-05 | Y | Y | Y | 372 | 10/14/1997 2:50:00 PM | 4/22/1998 5:40:00 PM | 1/19/1999 6:45:09 PM | 1996-04-12 | 6/3/1998 6:51:37 PM | EMS |  | AirConcentration | µg/m³ | Y | (air conc.) |  | Acenaphthene (air conc.) | 83-32-9 | Polycyclic Aromatic Hydrocarbons (PAHs); | Polycyclic Aromatic Hydrocarbons (PAHs); EMS code: PA01 | Chemical |  | LAB
PA01 | ug/g | Acenaphthene | MI01 | Microwave Extractn /GC-MS(PAH) | Soil samples are ground with anhydrous sodium sulphate and microwave extracted with dichloromethane. The extract is then cleaned up by silica gel column chromatography and analyzed by gc-ms in selective ion mode. | 0.001 | Y | Y | Y | 51 | 6/1/2000 11:42:00 AM | 9/19/2001 11:39:00 AM | 3/17/2003 1:22:50 PM | 1996-10-13 | 6/3/1998 6:51:46 PM | EMS |  | MassPerMass | mg/kg | Y | (mass conc.) |  | Acenaphthene (mass conc.) | 83-32-9 | Polycyclic Aromatic Hydrocarbons (PAHs); | Polycyclic Aromatic Hydrocarbons (PAHs); EMS code: PA01 | Chemical |  | LAB
PA01 | ug/g | Acenaphthene | X214 | Mod.Proc21.001(AOAC,'84) | Mod.Proc21.001(AOAC,'84) | 0.005 | Y | Y | Y | 160 |  |  | 10/24/1996 2:26:15 PM | 1996-04-12 | 6/3/1998 6:51:37 PM | EMS |  | MassPerMass | mg/kg | Y | (mass conc.) |  | Acenaphthene (mass conc.) | 83-32-9 | Polycyclic Aromatic Hydrocarbons (PAHs); | Polycyclic Aromatic Hydrocarbons (PAHs); EMS code: PA01 | Chemical |  | LAB
PA01 | ug/m3 | Acenaphthene | PAH2 | XAD2: GC/MS | To understand the extent of human exposure to B[a]P, and other PAHs, a reliable sampling and analytical method has been established which involves a combination quartz fibre filter/adsorbent cartridge with subsequent extraction and analysis by gas chromatography (GC) with mass spectrometry (MS) detection (GC/MS). The analytical methods are a modification of EPA Test Method 625, Methods for Organic Chemical Analysis of Municipal and Industrial Wastewater, and Methods 8270, Test Methods for Evaluation of Solid Waste. PAH1 utilizes PUF as an adsorbant; PAH2 utilizes XAD-2 as an adsorbant. Note that isotopic surrogates may be added and are reported as % recovery. | 5e-05 | Y | Y | Y | 223 |  |  | 10/24/1996 2:07:37 PM | 1996-04-12 | 6/3/1998 6:51:37 PM | EMS |  | AirConcentration | µg/m³ | Y | (air conc.) |  | Acenaphthene (air conc.) | 83-32-9 | Polycyclic Aromatic Hydrocarbons (PAHs); | Polycyclic Aromatic Hydrocarbons (PAHs); EMS code: PA01 | Chemical |  | LAB
PA01 | mg/L | Acenaphthene | MW02 | DCM extrn,GC/MS-SIM (PAH) | Water samples are liquid - liquid extracted with dichloromethane. The extract is then cleaned by silica gel column chromatography and analyzed by gas chromatography - mass spectroscopy. | 1e-05 | Y | Y | Y | 1108 | 2/5/2002 | 7/10/2024 5:51:00 PM | 8/21/2024 11:42:55 AM | 1996-04-12 | 6/3/1998 6:51:42 PM | EMS |  | Concentration | mg/L | Y | (fl. conc.) |  | Acenaphthene (fl. conc.) | 83-32-9 | Polycyclic Aromatic Hydrocarbons (PAHs); | Polycyclic Aromatic Hydrocarbons (PAHs); EMS code: PA01 | Chemical |  | LAB
PA01 | ug/g | Acenaphthene | FS87 | DCM,Ac/Hex extr,PAM,HRGC/LRMS | DCM;Ac/Hex extr;PAM;HRGC/LRMS | 0.02 | Y | Y | Y | 250 | 4/12/2000 2:52:00 PM | 10/29/2022 8:18:00 AM | 1/11/2024 1:07:08 PM | 2001-02-23 |  | NPEPPIN |  | MassPerMass | mg/kg | Y | (mass conc.) |  | Acenaphthene (mass conc.) | 83-32-9 | Polycyclic Aromatic Hydrocarbons (PAHs); | Polycyclic Aromatic Hydrocarbons (PAHs); EMS code: PA01 | Chemical |  | LAB
PA02 | mg/L | Acenaphthylene | X262 | CH2Cl2 Ext:Silica:GC/MS | The sample is placed in a separatory funnel, the pH of the sample is adjusted to alkaline and then acidic conditions, and extracted each time with dichloromethane to selectively remove compounds. The final extracts (basic and acidic) can be combined or run separately. The extract is cleaned by silica gel column chromatography, then analyzed by GC/MS. | 0.001 | Y | Y | Y | 126 | 6/12/2000 | 1/13/2015 | 6/7/2024 5:30:18 PM | 1996-04-12 | 6/3/1998 6:51:37 PM | EMS |  | Concentration | mg/L | Y | (fl. conc.) |  | Acenaphthylene (fl. conc.) | 208-96-8 | Polycyclic Aromatic Hydrocarbons (PAHs); | Polycyclic Aromatic Hydrocarbons (PAHs); EMS code: PA02 | Chemical |  | LAB
PA02 | ug/g | Acenaphthylene | X214 | Mod.Proc21.001(AOAC,'84) | Mod.Proc21.001(AOAC,'84) | 0.005 | Y | Y | Y | 160 |  |  | 10/24/1996 2:26:15 PM | 1996-04-12 | 6/3/1998 6:51:45 PM | EMS |  | MassPerMass | mg/kg | Y | (mass conc.) |  | Acenaphthylene (mass conc.) | 208-96-8 | Polycyclic Aromatic Hydrocarbons (PAHs); | Polycyclic Aromatic Hydrocarbons (PAHs); EMS code: PA02 | Chemical |  | LAB
PA02 | mg/L | Acenaphthylene | HEGC | Hexane extraction: GC/MS | Hexane extraction: GC/MS | 1e-05 | Y | Y | Y | 898 | 6/8/2015 | 8/15/2024 | 8/16/2024 1:36:03 PM | 2015-07-07 | 8/15/2019 11:22:35 AM | RDGIBSON | KATFRASE | Concentration | mg/L | Y | (fl. conc.) |  | Acenaphthylene (fl. conc.) | 208-96-8 | Polycyclic Aromatic Hydrocarbons (PAHs); | Polycyclic Aromatic Hydrocarbons (PAHs); EMS code: PA02 | Chemical |  | LAB
PA02 | ug/g | Acenaphthylene | MP02 | GC/MS | GC/MS | 0.001 | Y | Y | Y | 523 | 3/28/2002 | 9/23/2021 3:24:00 AM | 10/18/2021 9:46:09 AM | 1996-04-12 | 6/3/1998 6:51:37 PM | EMS |  | MassPerMass | mg/kg | Y | (mass conc.) |  | Acenaphthylene (mass conc.) | 208-96-8 | Polycyclic Aromatic Hydrocarbons (PAHs); | Polycyclic Aromatic Hydrocarbons (PAHs); EMS code: PA02 | Chemical |  | LAB
PA02 | ug/g | Acenaphthylene | MI01 | Microwave Extractn /GC-MS(PAH) | Soil samples are ground with anhydrous sodium sulphate and microwave extracted with dichloromethane. The extract is then cleaned up by silica gel column chromatography and analyzed by gc-ms in selective ion mode. | 0.001 | Y | Y | Y | 51 | 6/1/2000 11:42:00 AM | 9/19/2001 11:39:00 AM | 3/17/2003 1:22:50 PM | 1996-10-13 | 6/3/1998 6:51:46 PM | EMS |  | MassPerMass | mg/kg | Y | (mass conc.) |  | Acenaphthylene (mass conc.) | 208-96-8 | Polycyclic Aromatic Hydrocarbons (PAHs); | Polycyclic Aromatic Hydrocarbons (PAHs); EMS code: PA02 | Chemical |  | LAB
PA02 | ug/m3 | Acenaphthylene | PAH2 | XAD2: GC/MS | To understand the extent of human exposure to B[a]P, and other PAHs, a reliable sampling and analytical method has been established which involves a combination quartz fibre filter/adsorbent cartridge with subsequent extraction and analysis by gas chromatography (GC) with mass spectrometry (MS) detection (GC/MS). The analytical methods are a modification of EPA Test Method 625, Methods for Organic Chemical Analysis of Municipal and Industrial Wastewater, and Methods 8270, Test Methods for Evaluation of Solid Waste. PAH1 utilizes PUF as an adsorbant; PAH2 utilizes XAD-2 as an adsorbant. Note that isotopic surrogates may be added and are reported as % recovery. | 5e-05 | Y | Y | Y | 223 |  |  | 10/24/1996 2:07:37 PM | 1996-04-12 | 6/3/1998 6:51:37 PM | EMS |  | AirConcentration | µg/m³ | Y | (air conc.) |  | Acenaphthylene (air conc.) | 208-96-8 | Polycyclic Aromatic Hydrocarbons (PAHs); | Polycyclic Aromatic Hydrocarbons (PAHs); EMS code: PA02 | Chemical |  | LAB
PA02 | mg/L | Acenaphthylene | MW02 | DCM extrn,GC/MS-SIM (PAH) | Water samples are liquid - liquid extracted with dichloromethane. The extract is then cleaned by silica gel column chromatography and analyzed by gas chromatography - mass spectroscopy. | 1e-05 | Y | Y | Y | 994 | 2/8/2002 4:20:00 PM | 7/10/2024 5:51:00 PM | 8/21/2024 11:42:55 AM | 1996-04-12 | 6/3/1998 6:51:42 PM | EMS |  | Concentration | mg/L | Y | (fl. conc.) |  | Acenaphthylene (fl. conc.) | 208-96-8 | Polycyclic Aromatic Hydrocarbons (PAHs); | Polycyclic Aromatic Hydrocarbons (PAHs); EMS code: PA02 | Chemical |  | LAB
PA02 | mg/L | Acenaphthylene | F010 | Extract disks,DCM extr:HRGC/LRMS | Water samples are spiked with a deuterated PAH Recovery Standards mix just prior to extraction with Empore extraction disks. Disks are eluted with Dichloromethane (DCM) and the extract concentrated. Deuterated PAH Internal Standards are added and the extract analyzed using High Resolution Gas Chromatography / Low Resolution Mass Spectrometry (HRGC/LRMS). MS data is acquired in Total Ion Mode (TIM) and quantitation is by multilevel Internal Standards. Recovery efficiency is calculated and reported but results are not recovery corrected. | 0.0002 | Y | Y | Y | 86 | 3/27/1997 4:22:00 PM | 3/11/2003 | 4/1/2004 1:13:26 AM | 1996-09-16 | 6/3/1998 6:51:44 PM | MPAK |  | Concentration | mg/L | Y | (fl. conc.) |  | Acenaphthylene (fl. conc.) | 208-96-8 | Polycyclic Aromatic Hydrocarbons (PAHs); | Polycyclic Aromatic Hydrocarbons (PAHs); EMS code: PA02 | Chemical |  | LAB
PA02 | ug/g wet | Acenaphthylene | X221 | Mod.Proc21.001(AOAC,'84) | Mod.Proc21.001(AOAC,'84) | 0.004 | Y | Y | Y | 73 |  |  | 10/24/1996 10:01:25 AM | 1996-04-12 | 6/3/1998 6:51:42 PM | EMS |  | MassPerMass | µg/g | Y | (mass conc.) | (wet) | Acenaphthylene (wet) (mass conc.) | 208-96-8 | Polycyclic Aromatic Hydrocarbons (PAHs); | Polycyclic Aromatic Hydrocarbons (PAHs); EMS code: PA02 | Chemical |  | LAB
PA02 | ug/m3 | Acenaphthylene | PAH1 | PUF:GC/MS | To understand the extent of human exposure to B[a]P and other PAHs, a reliable sampling and analytical method has been established which involves a combination quartz fibre filter/absorbent cartridge with subsequent extraction and analysis by gas chromatography (GC) with mass spectrometry (MS) detection (GC/MS). The analytical methods are a modification of EPA Test Method 625, Methods for Organic Chemical Analysis of Municipal and Industrial Wastewater, and Methods 8270, Test Methods for Evaluation of Solid Waste. PAH1 utilizes PUF as an adsorbant; PAH2 utilizes XAD-2 as an adsorbant. Note that isotopic surrogates may be added and are reported as % recovery. | 5e-05 | Y | Y | Y | 372 | 10/14/1997 2:50:00 PM | 4/22/1998 5:40:00 PM | 1/19/1999 6:45:09 PM | 1996-04-12 | 6/3/1998 6:51:37 PM | EMS |  | AirConcentration | µg/m³ | Y | (air conc.) |  | Acenaphthylene (air conc.) | 208-96-8 | Polycyclic Aromatic Hydrocarbons (PAHs); | Polycyclic Aromatic Hydrocarbons (PAHs); EMS code: PA02 | Chemical |  | LAB
PA02 | ug/g | Acenaphthylene | FS87 | DCM,Ac/Hex extr,PAM,HRGC/LRMS | DCM;Ac/Hex extr;PAM;HRGC/LRMS | 0.02 | Y | Y | Y | 250 | 4/12/2000 2:52:00 PM | 10/29/2022 8:18:00 AM | 1/11/2024 1:07:08 PM | 2001-02-23 |  | NPEPPIN |  | MassPerMass | mg/kg | Y | (mass conc.) |  | Acenaphthylene (mass conc.) | 208-96-8 | Polycyclic Aromatic Hydrocarbons (PAHs); | Polycyclic Aromatic Hydrocarbons (PAHs); EMS code: PA02 | Chemical |  | LAB
PA03 | ug/g wet | Anthracene | X221 | Mod.Proc21.001(AOAC,'84) | Mod.Proc21.001(AOAC,'84) | 0.004 | Y | Y | Y | 73 |  |  | 10/24/1996 10:01:25 AM | 1996-04-12 | 6/3/1998 6:51:42 PM | EMS |  | MassPerMass | µg/g | Y | (mass conc.) | (wet) | Anthracene (wet) (mass conc.) | 120-12-7 | Polycyclic Aromatic Hydrocarbons (PAHs); | Polycyclic Aromatic Hydrocarbons (PAHs); EMS code: PA03 | Chemical |  | LAB
PA03 | mg/L | Anthracene | X262 | CH2Cl2 Ext:Silica:GC/MS | The sample is placed in a separatory funnel, the pH of the sample is adjusted to alkaline and then acidic conditions, and extracted each time with dichloromethane to selectively remove compounds. The final extracts (basic and acidic) can be combined or run separately. The extract is cleaned by silica gel column chromatography, then analyzed by GC/MS. | 0.001 | Y | Y | Y | 126 | 6/12/2000 | 1/13/2015 | 6/7/2024 5:30:18 PM | 1996-04-12 | 6/3/1998 6:51:37 PM | EMS |  | Concentration | mg/L | Y | (fl. conc.) |  | Anthracene (fl. conc.) | 120-12-7 | Polycyclic Aromatic Hydrocarbons (PAHs); | Polycyclic Aromatic Hydrocarbons (PAHs); EMS code: PA03 | Chemical |  | LAB
PA03 | ug/m3 | Anthracene | PAH2 | XAD2: GC/MS | To understand the extent of human exposure to B[a]P, and other PAHs, a reliable sampling and analytical method has been established which involves a combination quartz fibre filter/adsorbent cartridge with subsequent extraction and analysis by gas chromatography (GC) with mass spectrometry (MS) detection (GC/MS). The analytical methods are a modification of EPA Test Method 625, Methods for Organic Chemical Analysis of Municipal and Industrial Wastewater, and Methods 8270, Test Methods for Evaluation of Solid Waste. PAH1 utilizes PUF as an adsorbant; PAH2 utilizes XAD-2 as an adsorbant. Note that isotopic surrogates may be added and are reported as % recovery. | 5e-05 | Y | Y | Y | 223 |  |  | 10/24/1996 2:07:37 PM | 1996-04-12 | 6/3/1998 6:51:37 PM | EMS |  | AirConcentration | µg/m³ | Y | (air conc.) |  | Anthracene (air conc.) | 120-12-7 | Polycyclic Aromatic Hydrocarbons (PAHs); | Polycyclic Aromatic Hydrocarbons (PAHs); EMS code: PA03 | Chemical |  | LAB
PA03 | ug/g | Anthracene | MI01 | Microwave Extractn /GC-MS(PAH) | Soil samples are ground with anhydrous sodium sulphate and microwave extracted with dichloromethane. The extract is then cleaned up by silica gel column chromatography and analyzed by gc-ms in selective ion mode. | 0.001 | Y | Y | Y | 51 | 6/1/2000 11:42:00 AM | 9/19/2001 11:39:00 AM | 3/17/2003 1:22:50 PM | 1996-10-13 | 6/3/1998 6:51:46 PM | EMS |  | MassPerMass | mg/kg | Y | (mass conc.) |  | Anthracene (mass conc.) | 120-12-7 | Polycyclic Aromatic Hydrocarbons (PAHs); | Polycyclic Aromatic Hydrocarbons (PAHs); EMS code: PA03 | Chemical |  | LAB
PA03 | mg/L | Anthracene | F010 | Extract disks,DCM extr:HRGC/LRMS | Water samples are spiked with a deuterated PAH Recovery Standards mix just prior to extraction with Empore extraction disks. Disks are eluted with Dichloromethane (DCM) and the extract concentrated. Deuterated PAH Internal Standards are added and the extract analyzed using High Resolution Gas Chromatography / Low Resolution Mass Spectrometry (HRGC/LRMS). MS data is acquired in Total Ion Mode (TIM) and quantitation is by multilevel Internal Standards. Recovery efficiency is calculated and reported but results are not recovery corrected. | 0.0002 | Y | Y | Y | 86 | 3/27/1997 4:22:00 PM | 3/11/2003 | 4/1/2004 1:13:26 AM | 1996-09-16 | 6/3/1998 6:51:44 PM | MPAK |  | Concentration | mg/L | Y | (fl. conc.) |  | Anthracene (fl. conc.) | 120-12-7 | Polycyclic Aromatic Hydrocarbons (PAHs); | Polycyclic Aromatic Hydrocarbons (PAHs); EMS code: PA03 | Chemical |  | LAB
PA03 | mg/L | Anthracene | HEGC | Hexane extraction: GC/MS | Hexane extraction: GC/MS | 1e-05 | Y | Y | Y | 905 | 6/8/2015 | 8/15/2024 | 8/16/2024 1:36:03 PM | 2015-07-07 | 8/15/2019 11:23:24 AM | RDGIBSON | KATFRASE | Concentration | mg/L | Y | (fl. conc.) |  | Anthracene (fl. conc.) | 120-12-7 | Polycyclic Aromatic Hydrocarbons (PAHs); | Polycyclic Aromatic Hydrocarbons (PAHs); EMS code: PA03 | Chemical |  | LAB
PA03 | mg/L | Anthracene | MW02 | DCM extrn,GC/MS-SIM (PAH) | Water samples are liquid - liquid extracted with dichloromethane. The extract is then cleaned by silica gel column chromatography and analyzed by gas chromatography - mass spectroscopy. | 1e-05 | Y | Y | Y | 1130 | 2/8/2002 4:20:00 PM | 7/10/2024 5:51:00 PM | 8/21/2024 11:42:55 AM | 1996-04-12 | 6/3/1998 6:51:42 PM | EMS |  | Concentration | mg/L | Y | (fl. conc.) |  | Anthracene (fl. conc.) | 120-12-7 | Polycyclic Aromatic Hydrocarbons (PAHs); | Polycyclic Aromatic Hydrocarbons (PAHs); EMS code: PA03 | Chemical |  | LAB
PA03 | ug/m3 | Anthracene | PAH1 | PUF:GC/MS | To understand the extent of human exposure to B[a]P and other PAHs, a reliable sampling and analytical method has been established which involves a combination quartz fibre filter/absorbent cartridge with subsequent extraction and analysis by gas chromatography (GC) with mass spectrometry (MS) detection (GC/MS). The analytical methods are a modification of EPA Test Method 625, Methods for Organic Chemical Analysis of Municipal and Industrial Wastewater, and Methods 8270, Test Methods for Evaluation of Solid Waste. PAH1 utilizes PUF as an adsorbant; PAH2 utilizes XAD-2 as an adsorbant. Note that isotopic surrogates may be added and are reported as % recovery. | 5e-05 | Y | Y | Y | 372 | 10/14/1997 2:50:00 PM | 4/22/1998 5:40:00 PM | 1/19/1999 6:45:10 PM | 1996-04-12 | 6/3/1998 6:51:42 PM | EMS |  | AirConcentration | µg/m³ | Y | (air conc.) |  | Anthracene (air conc.) | 120-12-7 | Polycyclic Aromatic Hydrocarbons (PAHs); | Polycyclic Aromatic Hydrocarbons (PAHs); EMS code: PA03 | Chemical |  | LAB
PA03 | ug/g | Anthracene | FS87 | DCM,Ac/Hex extr,PAM,HRGC/LRMS | DCM;Ac/Hex extr;PAM;HRGC/LRMS | 0.02 | Y | Y | Y | 250 | 4/12/2000 2:52:00 PM | 10/29/2022 8:18:00 AM | 1/11/2024 1:07:08 PM | 2001-02-23 |  | NPEPPIN |  | MassPerMass | mg/kg | Y | (mass conc.) |  | Anthracene (mass conc.) | 120-12-7 | Polycyclic Aromatic Hydrocarbons (PAHs); | Polycyclic Aromatic Hydrocarbons (PAHs); EMS code: PA03 | Chemical |  | LAB
PA03 | ug/g | Anthracene | X214 | Mod.Proc21.001(AOAC,'84) | Mod.Proc21.001(AOAC,'84) | 0.005 | Y | Y | Y | 160 |  |  | 10/24/1996 2:26:15 PM | 1996-04-12 | 6/3/1998 6:51:37 PM | EMS |  | MassPerMass | mg/kg | Y | (mass conc.) |  | Anthracene (mass conc.) | 120-12-7 | Polycyclic Aromatic Hydrocarbons (PAHs); | Polycyclic Aromatic Hydrocarbons (PAHs); EMS code: PA03 | Chemical |  | LAB
PA03 | ug/g | Anthracene | MP02 | GC/MS | GC/MS | 0.001 | Y | Y | Y | 523 | 3/28/2002 | 9/23/2021 3:24:00 AM | 10/18/2021 9:46:09 AM | 1996-04-12 | 6/3/1998 6:51:37 PM | EMS |  | MassPerMass | mg/kg | Y | (mass conc.) |  | Anthracene (mass conc.) | 120-12-7 | Polycyclic Aromatic Hydrocarbons (PAHs); | Polycyclic Aromatic Hydrocarbons (PAHs); EMS code: PA03 | Chemical |  | LAB
PA04 | ug/g | Benzo(a)anthracene | X214 | Mod.Proc21.001(AOAC,'84) | Mod.Proc21.001(AOAC,'84) | 0.01 | Y | Y | Y | 160 |  |  | 10/24/1996 2:26:15 PM | 1996-04-12 | 6/3/1998 6:51:37 PM | EMS |  | MassPerMass | mg/kg | Y | (mass conc.) |  | Benzo(a)anthracene (mass conc.) | 56-55-3 | Polycyclic Aromatic Hydrocarbons (PAHs); | Polycyclic Aromatic Hydrocarbons (PAHs); EMS code: PA04 | Chemical |  | LAB
PA04 | ug/g | Benzo(a)anthracene | FS87 | DCM,Ac/Hex extr,PAM,HRGC/LRMS | DCM;Ac/Hex extr;PAM;HRGC/LRMS | 0.02 | Y | Y | Y | 250 | 4/12/2000 2:52:00 PM | 10/29/2022 8:18:00 AM | 1/11/2024 1:07:08 PM | 2001-02-23 |  | NPEPPIN |  | MassPerMass | mg/kg | Y | (mass conc.) |  | Benzo(a)anthracene (mass conc.) | 56-55-3 | Polycyclic Aromatic Hydrocarbons (PAHs); | Polycyclic Aromatic Hydrocarbons (PAHs); EMS code: PA04 | Chemical |  | LAB
PA04 | mg/L | Benzo(a)anthracene | MW02 | DCM extrn,GC/MS-SIM (PAH) | Water samples are liquid - liquid extracted with dichloromethane. The extract is then cleaned by silica gel column chromatography and analyzed by gas chromatography - mass spectroscopy. | 1e-05 | Y | Y | Y | 1130 | 2/5/2002 | 7/10/2024 5:51:00 PM | 8/21/2024 11:42:55 AM | 1996-04-12 | 6/3/1998 6:51:42 PM | EMS |  | Concentration | mg/L | Y | (fl. conc.) |  | Benzo(a)anthracene (fl. conc.) | 56-55-3 | Polycyclic Aromatic Hydrocarbons (PAHs); | Polycyclic Aromatic Hydrocarbons (PAHs); EMS code: PA04 | Chemical |  | LAB
PA04 | mg/L | Benzo(a)anthracene | HEGC | Hexane extraction: GC/MS | Hexane extraction: GC/MS | 1e-05 | Y | Y | Y | 905 | 6/8/2015 | 8/15/2024 | 8/16/2024 1:36:03 PM | 2015-07-07 | 8/15/2019 11:23:53 AM | RDGIBSON | KATFRASE | Concentration | mg/L | Y | (fl. conc.) |  | Benzo(a)anthracene (fl. conc.) | 56-55-3 | Polycyclic Aromatic Hydrocarbons (PAHs); | Polycyclic Aromatic Hydrocarbons (PAHs); EMS code: PA04 | Chemical |  | LAB
PA04 | mg/L | Benzo(a)anthracene | X262 | CH2Cl2 Ext:Silica:GC/MS | The sample is placed in a separatory funnel, the pH of the sample is adjusted to alkaline and then acidic conditions, and extracted each time with dichloromethane to selectively remove compounds. The final extracts (basic and acidic) can be combined or run separately. The extract is cleaned by silica gel column chromatography, then analyzed by GC/MS. | 0.001 | Y | Y | Y | 126 | 6/12/2000 | 1/13/2015 | 6/7/2024 5:30:18 PM | 1996-04-12 | 6/3/1998 6:51:37 PM | EMS |  | Concentration | mg/L | Y | (fl. conc.) |  | Benzo(a)anthracene (fl. conc.) | 56-55-3 | Polycyclic Aromatic Hydrocarbons (PAHs); | Polycyclic Aromatic Hydrocarbons (PAHs); EMS code: PA04 | Chemical |  | LAB
PA04 | ug/g | Benzo(a)anthracene | MI01 | Microwave Extractn /GC-MS(PAH) | Soil samples are ground with anhydrous sodium sulphate and microwave extracted with dichloromethane. The extract is then cleaned up by silica gel column chromatography and analyzed by gc-ms in selective ion mode. | 0.001 | Y | Y | Y | 51 | 6/1/2000 11:42:00 AM | 9/19/2001 11:39:00 AM | 3/17/2003 1:22:51 PM | 1996-10-13 | 6/3/1998 6:51:46 PM | EMS |  | MassPerMass | mg/kg | Y | (mass conc.) |  | Benzo(a)anthracene (mass conc.) | 56-55-3 | Polycyclic Aromatic Hydrocarbons (PAHs); | Polycyclic Aromatic Hydrocarbons (PAHs); EMS code: PA04 | Chemical |  | LAB
PA04 | ug/g wet | Benzo(a)anthracene | X221 | Mod.Proc21.001(AOAC,'84) | Mod.Proc21.001(AOAC,'84) | 0.01 | Y | Y | Y | 72 |  |  | 10/24/1996 10:01:25 AM | 1996-04-12 | 6/3/1998 6:51:42 PM | EMS |  | MassPerMass | µg/g | Y | (mass conc.) | (wet) | Benzo(a)anthracene (wet) (mass conc.) | 56-55-3 | Polycyclic Aromatic Hydrocarbons (PAHs); | Polycyclic Aromatic Hydrocarbons (PAHs); EMS code: PA04 | Chemical |  | LAB
PA04 | ug/m3 | Benzo(a)anthracene | PAH2 | XAD2: GC/MS | To understand the extent of human exposure to B[a]P, and other PAHs, a reliable sampling and analytical method has been established which involves a combination quartz fibre filter/adsorbent cartridge with subsequent extraction and analysis by gas chromatography (GC) with mass spectrometry (MS) detection (GC/MS). The analytical methods are a modification of EPA Test Method 625, Methods for Organic Chemical Analysis of Municipal and Industrial Wastewater, and Methods 8270, Test Methods for Evaluation of Solid Waste. PAH1 utilizes PUF as an adsorbant; PAH2 utilizes XAD-2 as an adsorbant. Note that isotopic surrogates may be added and are reported as % recovery. | 5e-05 | Y | Y | Y | 223 |  |  | 10/24/1996 2:07:37 PM | 1996-04-12 | 6/3/1998 6:51:37 PM | EMS |  | AirConcentration | µg/m³ | Y | (air conc.) |  | Benzo(a)anthracene (air conc.) | 56-55-3 | Polycyclic Aromatic Hydrocarbons (PAHs); | Polycyclic Aromatic Hydrocarbons (PAHs); EMS code: PA04 | Chemical |  | LAB
PA04 | ug/m3 | Benzo(a)anthracene | PAH1 | PUF:GC/MS | To understand the extent of human exposure to B[a]P and other PAHs, a reliable sampling and analytical method has been established which involves a combination quartz fibre filter/absorbent cartridge with subsequent extraction and analysis by gas chromatography (GC) with mass spectrometry (MS) detection (GC/MS). The analytical methods are a modification of EPA Test Method 625, Methods for Organic Chemical Analysis of Municipal and Industrial Wastewater, and Methods 8270, Test Methods for Evaluation of Solid Waste. PAH1 utilizes PUF as an adsorbant; PAH2 utilizes XAD-2 as an adsorbant. Note that isotopic surrogates may be added and are reported as % recovery. | 5e-05 | Y | Y | Y | 372 | 10/14/1997 2:50:00 PM | 4/22/1998 5:40:00 PM | 1/19/1999 6:45:10 PM | 1996-04-12 | 6/3/1998 6:51:37 PM | EMS |  | AirConcentration | µg/m³ | Y | (air conc.) |  | Benzo(a)anthracene (air conc.) | 56-55-3 | Polycyclic Aromatic Hydrocarbons (PAHs); | Polycyclic Aromatic Hydrocarbons (PAHs); EMS code: PA04 | Chemical |  | LAB
PA04 | ug/g | Benzo(a)anthracene | MP02 | GC/MS | GC/MS | 0.001 | Y | Y | Y | 523 | 3/28/2002 | 9/23/2021 3:24:00 AM | 10/18/2021 9:46:09 AM | 1996-04-12 | 6/3/1998 6:51:37 PM | EMS |  | MassPerMass | mg/kg | Y | (mass conc.) |  | Benzo(a)anthracene (mass conc.) | 56-55-3 | Polycyclic Aromatic Hydrocarbons (PAHs); | Polycyclic Aromatic Hydrocarbons (PAHs); EMS code: PA04 | Chemical |  | LAB
PA04 | mg/L | Benzo(a)anthracene | F010 | Extract disks,DCM extr:HRGC/LRMS | Water samples are spiked with a deuterated PAH Recovery Standards mix just prior to extraction with Empore extraction disks. Disks are eluted with Dichloromethane (DCM) and the extract concentrated. Deuterated PAH Internal Standards are added and the extract analyzed using High Resolution Gas Chromatography / Low Resolution Mass Spectrometry (HRGC/LRMS). MS data is acquired in Total Ion Mode (TIM) and quantitation is by multilevel Internal Standards. Recovery efficiency is calculated and reported but results are not recovery corrected. | 0.0005 | Y | Y | Y | 86 | 3/27/1997 4:23:00 PM | 3/11/2003 | 4/1/2004 1:13:26 AM | 1996-09-16 | 6/3/1998 6:51:44 PM | MPAK |  | Concentration | mg/L | Y | (fl. conc.) |  | Benzo(a)anthracene (fl. conc.) | 56-55-3 | Polycyclic Aromatic Hydrocarbons (PAHs); | Polycyclic Aromatic Hydrocarbons (PAHs); EMS code: PA04 | Chemical |  | LAB
PA05 | mg/L | Benzo(a)pyrene | F010 | Extract disks,DCM extr:HRGC/LRMS | Water samples are spiked with a deuterated PAH Recovery Standards mix just prior to extraction with Empore extraction disks. Disks are eluted with Dichloromethane (DCM) and the extract concentrated. Deuterated PAH Internal Standards are added and the extract analyzed using High Resolution Gas Chromatography / Low Resolution Mass Spectrometry (HRGC/LRMS). MS data is acquired in Total Ion Mode (TIM) and quantitation is by multilevel Internal Standards. Recovery efficiency is calculated and reported but results are not recovery corrected. | 0.005 | Y | Y | Y | 86 | 3/27/1997 4:23:00 PM | 3/11/2003 | 4/1/2004 1:13:26 AM | 1996-09-16 | 6/3/1998 6:51:44 PM | MPAK |  | Concentration | mg/L | Y | (fl. conc.) |  | Benzo(a)pyrene (fl. conc.) | 50-32-8 | Polycyclic Aromatic Hydrocarbons (PAHs); | Polycyclic Aromatic Hydrocarbons (PAHs); EMS code: PA05 | Chemical |  | LAB
PA05 | ug/g wet | Benzo(a)pyrene | X221 | Mod.Proc21.001(AOAC,'84) | Mod.Proc21.001(AOAC,'84) | 0.02 | Y | Y | Y | 73 |  |  | 10/24/1996 10:01:25 AM | 1996-04-12 | 6/3/1998 6:51:42 PM | EMS |  | MassPerMass | µg/g | Y | (mass conc.) | (wet) | Benzo(a)pyrene (wet) (mass conc.) | 50-32-8 | Polycyclic Aromatic Hydrocarbons (PAHs); | Polycyclic Aromatic Hydrocarbons (PAHs); EMS code: PA05 | Chemical |  | LAB
PA05 | ug/m3 | Benzo(a)pyrene | PAH2 | XAD2: GC/MS | To understand the extent of human exposure to B[a]P, and other PAHs, a reliable sampling and analytical method has been established which involves a combination quartz fibre filter/adsorbent cartridge with subsequent extraction and analysis by gas chromatography (GC) with mass spectrometry (MS) detection (GC/MS). The analytical methods are a modification of EPA Test Method 625, Methods for Organic Chemical Analysis of Municipal and Industrial Wastewater, and Methods 8270, Test Methods for Evaluation of Solid Waste. PAH1 utilizes PUF as an adsorbant; PAH2 utilizes XAD-2 as an adsorbant. Note that isotopic surrogates may be added and are reported as % recovery. | 5e-05 | Y | Y | Y | 223 |  |  | 10/24/1996 2:07:37 PM | 1996-04-12 | 6/3/1998 6:51:37 PM | EMS |  | AirConcentration | µg/m³ | Y | (air conc.) |  | Benzo(a)pyrene (air conc.) | 50-32-8 | Polycyclic Aromatic Hydrocarbons (PAHs); | Polycyclic Aromatic Hydrocarbons (PAHs); EMS code: PA05 | Chemical |  | LAB
PA05 | ug/m3 | Benzo(a)pyrene | PAH1 | PUF:GC/MS | To understand the extent of human exposure to B[a]P and other PAHs, a reliable sampling and analytical method has been established which involves a combination quartz fibre filter/absorbent cartridge with subsequent extraction and analysis by gas chromatography (GC) with mass spectrometry (MS) detection (GC/MS). The analytical methods are a modification of EPA Test Method 625, Methods for Organic Chemical Analysis of Municipal and Industrial Wastewater, and Methods 8270, Test Methods for Evaluation of Solid Waste. PAH1 utilizes PUF as an adsorbant; PAH2 utilizes XAD-2 as an adsorbant. Note that isotopic surrogates may be added and are reported as % recovery. | 5e-05 | Y | Y | Y | 372 | 10/14/1997 2:50:00 PM | 4/22/1998 5:40:00 PM | 1/19/1999 6:45:10 PM | 1996-04-12 | 6/3/1998 6:51:37 PM | EMS |  | AirConcentration | µg/m³ | Y | (air conc.) |  | Benzo(a)pyrene (air conc.) | 50-32-8 | Polycyclic Aromatic Hydrocarbons (PAHs); | Polycyclic Aromatic Hydrocarbons (PAHs); EMS code: PA05 | Chemical |  | LAB
PA05 | ug/g | Benzo(a)pyrene | ZN03 | DCM+Ac.Extr:GC/MS | DCM+Ac.Extr:GC/MS EPA 8270E (mod) | 0.05 | Y | Y | Y | 14 | 1/25/2022 | 9/8/2022 | 12/15/2022 10:50:07 AM | 2022-07-06 |  | EMS |  | MassPerMass | mg/kg | Y | (mass conc.) |  | Benzo(a)pyrene (mass conc.) | 50-32-8 | Polycyclic Aromatic Hydrocarbons (PAHs); | Polycyclic Aromatic Hydrocarbons (PAHs); EMS code: PA05 | Chemical |  | LAB
PA05 | mg/L | Benzo(a)pyrene | MW02 | DCM extrn,GC/MS-SIM (PAH) | Water samples are liquid - liquid extracted with dichloromethane. The extract is then cleaned by silica gel column chromatography and analyzed by gas chromatography - mass spectroscopy. | 1e-05 | Y | Y | Y | 1132 | 2/5/2002 | 7/10/2024 5:51:00 PM | 8/21/2024 11:42:55 AM | 1996-04-12 | 6/3/1998 6:51:42 PM | EMS |  | Concentration | mg/L | Y | (fl. conc.) |  | Benzo(a)pyrene (fl. conc.) | 50-32-8 | Polycyclic Aromatic Hydrocarbons (PAHs); | Polycyclic Aromatic Hydrocarbons (PAHs); EMS code: PA05 | Chemical |  | LAB
PA05 | mg/L | Benzo(a)pyrene | X262 | CH2Cl2 Ext:Silica:GC/MS | The sample is placed in a separatory funnel, the pH of the sample is adjusted to alkaline and then acidic conditions, and extracted each time with dichloromethane to selectively remove compounds. The final extracts (basic and acidic) can be combined or run separately. The extract is cleaned by silica gel column chromatography, then analyzed by GC/MS. | 0.001 | Y | Y | Y | 126 | 6/12/2000 | 1/13/2015 | 6/7/2024 5:30:18 PM | 1996-04-12 | 6/3/1998 6:51:37 PM | EMS |  | Concentration | mg/L | Y | (fl. conc.) |  | Benzo(a)pyrene (fl. conc.) | 50-32-8 | Polycyclic Aromatic Hydrocarbons (PAHs); | Polycyclic Aromatic Hydrocarbons (PAHs); EMS code: PA05 | Chemical |  | LAB
PA05 | mg/L | Benzo(a)pyrene | HEGC | Hexane extraction: GC/MS | Hexane extraction: GC/MS | 5e-06 | Y | Y | Y | 905 | 6/8/2015 | 8/15/2024 | 8/16/2024 1:36:03 PM | 2015-07-07 | 8/15/2019 11:24:26 AM | RDGIBSON | KATFRASE | Concentration | mg/L | Y | (fl. conc.) |  | Benzo(a)pyrene (fl. conc.) | 50-32-8 | Polycyclic Aromatic Hydrocarbons (PAHs); | Polycyclic Aromatic Hydrocarbons (PAHs); EMS code: PA05 | Chemical |  | LAB
PA05 | ug/g | Benzo(a)pyrene | MI01 | Microwave Extractn /GC-MS(PAH) | Soil samples are ground with anhydrous sodium sulphate and microwave extracted with dichloromethane. The extract is then cleaned up by silica gel column chromatography and analyzed by gc-ms in selective ion mode. | 0.001 | Y | Y | Y | 51 | 6/1/2000 11:42:00 AM | 9/19/2001 11:39:00 AM | 3/17/2003 1:22:51 PM | 1996-10-13 | 6/3/1998 6:51:46 PM | EMS |  | MassPerMass | mg/kg | Y | (mass conc.) |  | Benzo(a)pyrene (mass conc.) | 50-32-8 | Polycyclic Aromatic Hydrocarbons (PAHs); | Polycyclic Aromatic Hydrocarbons (PAHs); EMS code: PA05 | Chemical |  | LAB
PA05 | ug/g | Benzo(a)pyrene | FS87 | DCM,Ac/Hex extr,PAM,HRGC/LRMS | DCM;Ac/Hex extr;PAM;HRGC/LRMS | 0.02 | Y | Y | Y | 250 | 4/12/2000 2:52:00 PM | 10/29/2022 8:18:00 AM | 1/11/2024 1:07:08 PM | 2001-02-23 |  | NPEPPIN |  | MassPerMass | mg/kg | Y | (mass conc.) |  | Benzo(a)pyrene (mass conc.) | 50-32-8 | Polycyclic Aromatic Hydrocarbons (PAHs); | Polycyclic Aromatic Hydrocarbons (PAHs); EMS code: PA05 | Chemical |  | LAB
PA05 | ug/g | Benzo(a)pyrene | MP02 | GC/MS | GC/MS | 0.001 | Y | Y | Y | 523 | 3/28/2002 | 9/23/2021 3:24:00 AM | 10/18/2021 9:46:09 AM | 1996-04-12 | 6/3/1998 6:51:37 PM | EMS |  | MassPerMass | mg/kg | Y | (mass conc.) |  | Benzo(a)pyrene (mass conc.) | 50-32-8 | Polycyclic Aromatic Hydrocarbons (PAHs); | Polycyclic Aromatic Hydrocarbons (PAHs); EMS code: PA05 | Chemical |  | LAB
PA05 | ug/g | Benzo(a)pyrene | X214 | Mod.Proc21.001(AOAC,'84) | Mod.Proc21.001(AOAC,'84) | 0.02 | Y | Y | Y | 160 |  |  | 10/24/1996 2:26:15 PM | 1996-04-12 | 6/3/1998 6:51:37 PM | EMS |  | MassPerMass | mg/kg | Y | (mass conc.) |  | Benzo(a)pyrene (mass conc.) | 50-32-8 | Polycyclic Aromatic Hydrocarbons (PAHs); | Polycyclic Aromatic Hydrocarbons (PAHs); EMS code: PA05 | Chemical |  | LAB
PA06 | ug/g | Benzo(b)fluoranthene | FS87 | DCM,Ac/Hex extr,PAM,HRGC/LRMS | DCM;Ac/Hex extr;PAM;HRGC/LRMS | 0.02 | Y | Y | Y | 172 | 4/12/2000 2:52:00 PM | 10/29/2022 8:18:00 AM | 1/7/2023 3:12:11 PM | 2001-02-23 |  | NPEPPIN |  | MassPerMass | mg/kg | Y | (mass conc.) |  | Benzo(b)fluoranthene (mass conc.) | 205-99-2 | Polycyclic Aromatic Hydrocarbons (PAHs); | Polycyclic Aromatic Hydrocarbons (PAHs); EMS code: PA06 | Chemical |  | LAB
PA06 | mg/L | Benzo(b)fluoranthene | MW02 | DCM extrn,GC/MS-SIM (PAH) | Water samples are liquid - liquid extracted with dichloromethane. The extract is then cleaned by silica gel column chromatography and analyzed by gas chromatography - mass spectroscopy. | 1e-05 | Y | Y | Y | 76 | 2/5/2002 | 7/10/2024 5:51:00 PM | 8/21/2024 11:42:55 AM | 1996-04-12 | 6/3/1998 6:51:42 PM | EMS |  | Concentration | mg/L | Y | (fl. conc.) |  | Benzo(b)fluoranthene (fl. conc.) | 205-99-2 | Polycyclic Aromatic Hydrocarbons (PAHs); | Polycyclic Aromatic Hydrocarbons (PAHs); EMS code: PA06 | Chemical |  | LAB
PA06 | ug/m3 | Benzo(b)fluoranthene | PAH1 | PUF:GC/MS | To understand the extent of human exposure to B[a]P and other PAHs, a reliable sampling and analytical method has been established which involves a combination quartz fibre filter/absorbent cartridge with subsequent extraction and analysis by gas chromatography (GC) with mass spectrometry (MS) detection (GC/MS). The analytical methods are a modification of EPA Test Method 625, Methods for Organic Chemical Analysis of Municipal and Industrial Wastewater, and Methods 8270, Test Methods for Evaluation of Solid Waste. PAH1 utilizes PUF as an adsorbant; PAH2 utilizes XAD-2 as an adsorbant. Note that isotopic surrogates may be added and are reported as % recovery. | 5e-05 | Y | Y | Y | 27 | 10/14/1997 2:50:00 PM | 4/22/1998 5:40:00 PM | 1/19/1999 6:45:10 PM | 1996-04-12 | 6/3/1998 6:51:37 PM | EMS |  | AirConcentration | µg/m³ | Y | (air conc.) |  | Benzo(b)fluoranthene (air conc.) | 205-99-2 | Polycyclic Aromatic Hydrocarbons (PAHs); | Polycyclic Aromatic Hydrocarbons (PAHs); EMS code: PA06 | Chemical |  | LAB
PA06 | mg/L | Benzo(b)fluoranthene | X262 | CH2Cl2 Ext:Silica:GC/MS | The sample is placed in a separatory funnel, the pH of the sample is adjusted to alkaline and then acidic conditions, and extracted each time with dichloromethane to selectively remove compounds. The final extracts (basic and acidic) can be combined or run separately. The extract is cleaned by silica gel column chromatography, then analyzed by GC/MS. | 0.001 | Y | Y | Y | 126 | 6/12/2000 | 1/13/2015 | 6/7/2024 5:30:18 PM | 1996-04-12 | 6/3/1998 6:51:37 PM | EMS |  | Concentration | mg/L | Y | (fl. conc.) |  | Benzo(b)fluoranthene (fl. conc.) | 205-99-2 | Polycyclic Aromatic Hydrocarbons (PAHs); | Polycyclic Aromatic Hydrocarbons (PAHs); EMS code: PA06 | Chemical |  | LAB
PA06 | ug/g | Benzo(b)fluoranthene | MI01 | Microwave Extractn /GC-MS(PAH) | Soil samples are ground with anhydrous sodium sulphate and microwave extracted with dichloromethane. The extract is then cleaned up by silica gel column chromatography and analyzed by gc-ms in selective ion mode. | 0.001 | Y | Y | Y | 51 | 6/1/2000 11:42:00 AM | 9/19/2001 11:39:00 AM | 3/17/2003 1:22:51 PM | 1996-10-13 | 6/3/1998 6:51:46 PM | EMS |  | MassPerMass | mg/kg | Y | (mass conc.) |  | Benzo(b)fluoranthene (mass conc.) | 205-99-2 | Polycyclic Aromatic Hydrocarbons (PAHs); | Polycyclic Aromatic Hydrocarbons (PAHs); EMS code: PA06 | Chemical |  | LAB
PA06 | mg/L | Benzo(b)fluoranthene | HEGC | Hexane extraction: GC/MS | Hexane extraction: GC/MS | 1e-05 | Y | Y | Y | 272 | 6/8/2015 | 7/20/2017 | 8/15/2019 2:42:34 PM | 2015-07-07 | 8/15/2019 11:24:53 AM | RDGIBSON | KATFRASE | Concentration | mg/L | Y | (fl. conc.) |  | Benzo(b)fluoranthene (fl. conc.) | 205-99-2 | Polycyclic Aromatic Hydrocarbons (PAHs); | Polycyclic Aromatic Hydrocarbons (PAHs); EMS code: PA06 | Chemical |  | LAB
PA06 | mg/L | Benzo(b)fluoranthene | F010 | Extract disks,DCM extr:HRGC/LRMS | Water samples are spiked with a deuterated PAH Recovery Standards mix just prior to extraction with Empore extraction disks. Disks are eluted with Dichloromethane (DCM) and the extract concentrated. Deuterated PAH Internal Standards are added and the extract analyzed using High Resolution Gas Chromatography / Low Resolution Mass Spectrometry (HRGC/LRMS). MS data is acquired in Total Ion Mode (TIM) and quantitation is by multilevel Internal Standards. Recovery efficiency is calculated and reported but results are not recovery corrected. | 0.0005 | Y | Y | Y | 86 | 3/27/1997 4:23:00 PM | 3/11/2003 | 4/1/2004 1:13:26 AM | 1996-09-16 | 6/3/1998 6:51:44 PM | MPAK |  | Concentration | mg/L | Y | (fl. conc.) |  | Benzo(b)fluoranthene (fl. conc.) | 205-99-2 | Polycyclic Aromatic Hydrocarbons (PAHs); | Polycyclic Aromatic Hydrocarbons (PAHs); EMS code: PA06 | Chemical |  | LAB
PA06 | ug/g | Benzo(b)fluoranthene | MP02 | GC/MS | GC/MS | 0.001 | Y | Y | Y | 116 | 3/28/2002 | 9/23/2021 3:24:00 AM | 10/18/2021 9:46:09 AM | 1996-04-12 | 6/3/1998 6:51:37 PM | EMS |  | MassPerMass | mg/kg | Y | (mass conc.) |  | Benzo(b)fluoranthene (mass conc.) | 205-99-2 | Polycyclic Aromatic Hydrocarbons (PAHs); | Polycyclic Aromatic Hydrocarbons (PAHs); EMS code: PA06 | Chemical |  | LAB
PA06 | ug/g | Benzo(b)fluoranthene | X214 | Mod.Proc21.001(AOAC,'84) | Mod.Proc21.001(AOAC,'84) | 0.02 | Y | Y | Y | 160 |  |  | 10/24/1996 2:26:15 PM | 1996-04-12 | 6/3/1998 6:51:37 PM | EMS |  | MassPerMass | mg/kg | Y | (mass conc.) |  | Benzo(b)fluoranthene (mass conc.) | 205-99-2 | Polycyclic Aromatic Hydrocarbons (PAHs); | Polycyclic Aromatic Hydrocarbons (PAHs); EMS code: PA06 | Chemical |  | LAB
PA06 | ug/g wet | Benzo(b)fluoranthene | X221 | Mod.Proc21.001(AOAC,'84) | Mod.Proc21.001(AOAC,'84) | 0.02 | Y | Y | Y | 73 |  |  | 10/24/1996 10:01:25 AM | 1996-04-12 | 6/3/1998 6:51:42 PM | EMS |  | MassPerMass | µg/g | Y | (mass conc.) | (wet) | Benzo(b)fluoranthene (wet) (mass conc.) | 205-99-2 | Polycyclic Aromatic Hydrocarbons (PAHs); | Polycyclic Aromatic Hydrocarbons (PAHs); EMS code: PA06 | Chemical |  | LAB
PA07 | ug/m3 | Benzo(g;h;i)perylene | PAH1 | PUF:GC/MS | To understand the extent of human exposure to B[a]P and other PAHs, a reliable sampling and analytical method has been established which involves a combination quartz fibre filter/absorbent cartridge with subsequent extraction and analysis by gas chromatography (GC) with mass spectrometry (MS) detection (GC/MS). The analytical methods are a modification of EPA Test Method 625, Methods for Organic Chemical Analysis of Municipal and Industrial Wastewater, and Methods 8270, Test Methods for Evaluation of Solid Waste. PAH1 utilizes PUF as an adsorbant; PAH2 utilizes XAD-2 as an adsorbant. Note that isotopic surrogates may be added and are reported as % recovery. | 0.0001 | Y | Y | Y | 371 | 10/14/1997 2:50:00 PM | 4/22/1998 5:40:00 PM | 1/19/1999 6:45:10 PM | 1996-04-12 | 6/3/1998 6:51:37 PM | EMS |  | AirConcentration | µg/m³ | Y | (air conc.) |  | Benzo(g;h;i)perylene (air conc.) | 191-24-2 | Polycyclic Aromatic Hydrocarbons (PAHs); | Polycyclic Aromatic Hydrocarbons (PAHs); EMS code: PA07 | Chemical |  | LAB
PA07 | ug/g | Benzo(g;h;i)perylene | MP02 | GC/MS | GC/MS | 0.001 | Y | Y | Y | 523 | 3/28/2002 | 9/23/2021 3:24:00 AM | 10/18/2021 9:46:09 AM | 1996-04-12 | 6/3/1998 6:51:37 PM | EMS |  | MassPerMass | mg/kg | Y | (mass conc.) |  | Benzo(g;h;i)perylene (mass conc.) | 191-24-2 | Polycyclic Aromatic Hydrocarbons (PAHs); | Polycyclic Aromatic Hydrocarbons (PAHs); EMS code: PA07 | Chemical |  | LAB
PA07 | mg/L | Benzo(g;h;i)perylene | HEGC | Hexane extraction: GC/MS | Hexane extraction: GC/MS | 1e-05 | Y | Y | Y | 898 | 6/8/2015 | 8/15/2024 | 8/16/2024 1:36:03 PM | 2015-07-07 | 8/15/2019 11:25:23 AM | RDGIBSON | KATFRASE | Concentration | mg/L | Y | (fl. conc.) |  | Benzo(g;h;i)perylene (fl. conc.) | 191-24-2 | Polycyclic Aromatic Hydrocarbons (PAHs); | Polycyclic Aromatic Hydrocarbons (PAHs); EMS code: PA07 | Chemical |  | LAB
PA07 | ug/g wet | Benzo(g;h;i)perylene | X221 | Mod.Proc21.001(AOAC,'84) | Mod.Proc21.001(AOAC,'84) | 0.02 | Y | Y | Y | 71 |  |  | 10/24/1996 10:01:25 AM | 1996-04-12 | 6/3/1998 6:51:42 PM | EMS |  | MassPerMass | µg/g | Y | (mass conc.) | (wet) | Benzo(g;h;i)perylene (wet) (mass conc.) | 191-24-2 | Polycyclic Aromatic Hydrocarbons (PAHs); | Polycyclic Aromatic Hydrocarbons (PAHs); EMS code: PA07 | Chemical |  | LAB
PA07 | mg/L | Benzo(g;h;i)perylene | F010 | Extract disks,DCM extr:HRGC/LRMS | Water samples are spiked with a deuterated PAH Recovery Standards mix just prior to extraction with Empore extraction disks. Disks are eluted with Dichloromethane (DCM) and the extract concentrated. Deuterated PAH Internal Standards are added and the extract analyzed using High Resolution Gas Chromatography / Low Resolution Mass Spectrometry (HRGC/LRMS). MS data is acquired in Total Ion Mode (TIM) and quantitation is by multilevel Internal Standards. Recovery efficiency is calculated and reported but results are not recovery corrected. | 0.005 | Y | Y | Y | 86 | 3/27/1997 4:23:00 PM | 3/11/2003 | 4/1/2004 1:13:26 AM | 1996-09-16 | 6/3/1998 6:51:44 PM | MPAK |  | Concentration | mg/L | Y | (fl. conc.) |  | Benzo(g;h;i)perylene (fl. conc.) | 191-24-2 | Polycyclic Aromatic Hydrocarbons (PAHs); | Polycyclic Aromatic Hydrocarbons (PAHs); EMS code: PA07 | Chemical |  | LAB
PA07 | ug/g | Benzo(g;h;i)perylene | MI01 | Microwave Extractn /GC-MS(PAH) | Soil samples are ground with anhydrous sodium sulphate and microwave extracted with dichloromethane. The extract is then cleaned up by silica gel column chromatography and analyzed by gc-ms in selective ion mode. | 0.001 | Y | Y | Y | 51 | 6/1/2000 11:42:00 AM | 9/19/2001 11:39:00 AM | 3/17/2003 1:22:52 PM | 1996-10-13 | 6/3/1998 6:51:46 PM | EMS |  | MassPerMass | mg/kg | Y | (mass conc.) |  | Benzo(g;h;i)perylene (mass conc.) | 191-24-2 | Polycyclic Aromatic Hydrocarbons (PAHs); | Polycyclic Aromatic Hydrocarbons (PAHs); EMS code: PA07 | Chemical |  | LAB
PA07 | ug/g | Benzo(g;h;i)perylene | X214 | Mod.Proc21.001(AOAC,'84) | Mod.Proc21.001(AOAC,'84) | 0.02 | Y | Y | Y | 160 |  |  | 10/24/1996 2:26:15 PM | 1996-04-12 | 6/3/1998 6:51:37 PM | EMS |  | MassPerMass | mg/kg | Y | (mass conc.) |  | Benzo(g;h;i)perylene (mass conc.) | 191-24-2 | Polycyclic Aromatic Hydrocarbons (PAHs); | Polycyclic Aromatic Hydrocarbons (PAHs); EMS code: PA07 | Chemical |  | LAB
PA07 | ug/m3 | Benzo(g;h;i)perylene | PAH2 | XAD2: GC/MS | To understand the extent of human exposure to B[a]P, and other PAHs, a reliable sampling and analytical method has been established which involves a combination quartz fibre filter/adsorbent cartridge with subsequent extraction and analysis by gas chromatography (GC) with mass spectrometry (MS) detection (GC/MS). The analytical methods are a modification of EPA Test Method 625, Methods for Organic Chemical Analysis of Municipal and Industrial Wastewater, and Methods 8270, Test Methods for Evaluation of Solid Waste. PAH1 utilizes PUF as an adsorbant; PAH2 utilizes XAD-2 as an adsorbant. Note that isotopic surrogates may be added and are reported as % recovery. | 0.0001 | Y | Y | Y | 223 |  |  | 10/24/1996 2:07:37 PM | 1996-04-12 | 6/3/1998 6:51:37 PM | EMS |  | AirConcentration | µg/m³ | Y | (air conc.) |  | Benzo(g;h;i)perylene (air conc.) | 191-24-2 | Polycyclic Aromatic Hydrocarbons (PAHs); | Polycyclic Aromatic Hydrocarbons (PAHs); EMS code: PA07 | Chemical |  | LAB
PA07 | mg/L | Benzo(g;h;i)perylene | X262 | CH2Cl2 Ext:Silica:GC/MS | The sample is placed in a separatory funnel, the pH of the sample is adjusted to alkaline and then acidic conditions, and extracted each time with dichloromethane to selectively remove compounds. The final extracts (basic and acidic) can be combined or run separately. The extract is cleaned by silica gel column chromatography, then analyzed by GC/MS. | 0.001 | Y | Y | Y | 126 | 6/12/2000 | 1/13/2015 | 6/7/2024 5:30:18 PM | 1996-04-12 | 6/3/1998 6:51:37 PM | EMS |  | Concentration | mg/L | Y | (fl. conc.) |  | Benzo(g;h;i)perylene (fl. conc.) | 191-24-2 | Polycyclic Aromatic Hydrocarbons (PAHs); | Polycyclic Aromatic Hydrocarbons (PAHs); EMS code: PA07 | Chemical |  | LAB
PA07 | mg/L | Benzo(g;h;i)perylene | MW02 | DCM extrn,GC/MS-SIM (PAH) | Water samples are liquid - liquid extracted with dichloromethane. The extract is then cleaned by silica gel column chromatography and analyzed by gas chromatography - mass spectroscopy. | 2e-05 | Y | Y | Y | 1018 | 2/5/2002 | 7/10/2024 5:51:00 PM | 8/21/2024 11:42:55 AM | 1996-04-12 | 6/3/1998 6:51:42 PM | EMS |  | Concentration | mg/L | Y | (fl. conc.) |  | Benzo(g;h;i)perylene (fl. conc.) | 191-24-2 | Polycyclic Aromatic Hydrocarbons (PAHs); | Polycyclic Aromatic Hydrocarbons (PAHs); EMS code: PA07 | Chemical |  | LAB
PA07 | ug/g | Benzo(g;h;i)perylene | FS87 | DCM,Ac/Hex extr,PAM,HRGC/LRMS | DCM;Ac/Hex extr;PAM;HRGC/LRMS | 0.02 | Y | Y | Y | 250 | 4/12/2000 2:52:00 PM | 10/29/2022 8:18:00 AM | 1/11/2024 1:07:08 PM | 2001-02-23 |  | NPEPPIN |  | MassPerMass | mg/kg | Y | (mass conc.) |  | Benzo(g;h;i)perylene (mass conc.) | 191-24-2 | Polycyclic Aromatic Hydrocarbons (PAHs); | Polycyclic Aromatic Hydrocarbons (PAHs); EMS code: PA07 | Chemical |  | LAB
PA08 | ug/g | Benzo(k)fluoranthene | MP02 | GC/MS | GC/MS | 0.001 | Y | Y | Y | 494 | 3/28/2002 | 9/23/2021 3:24:00 AM | 10/18/2021 9:46:09 AM | 1996-04-12 | 6/3/1998 6:51:37 PM | EMS |  | MassPerMass | mg/kg | Y | (mass conc.) |  | Benzo(k)fluoranthene (mass conc.) | 207-08-9 | Polycyclic Aromatic Hydrocarbons (PAHs); | Polycyclic Aromatic Hydrocarbons (PAHs); EMS code: PA08 | Chemical |  | LAB
PA08 | ug/g | Benzo(k)fluoranthene | MI01 | Microwave Extractn /GC-MS(PAH) | Soil samples are ground with anhydrous sodium sulphate and microwave extracted with dichloromethane. The extract is then cleaned up by silica gel column chromatography and analyzed by gc-ms in selective ion mode. | 0.001 | Y | Y | Y | 51 | 6/1/2000 11:42:00 AM | 9/19/2001 11:39:00 AM | 3/17/2003 1:22:52 PM | 1996-10-13 | 6/3/1998 6:51:46 PM | EMS |  | MassPerMass | mg/kg | Y | (mass conc.) |  | Benzo(k)fluoranthene (mass conc.) | 207-08-9 | Polycyclic Aromatic Hydrocarbons (PAHs); | Polycyclic Aromatic Hydrocarbons (PAHs); EMS code: PA08 | Chemical |  | LAB
PA08 | mg/L | Benzo(k)fluoranthene | HEGC | Hexane extraction: GC/MS | Hexane extraction: GC/MS | 1e-05 | Y | Y | Y | 898 | 6/8/2015 | 8/15/2024 | 8/16/2024 1:36:03 PM | 2015-07-07 | 8/15/2019 11:25:46 AM | RDGIBSON | KATFRASE | Concentration | mg/L | Y | (fl. conc.) |  | Benzo(k)fluoranthene (fl. conc.) | 207-08-9 | Polycyclic Aromatic Hydrocarbons (PAHs); | Polycyclic Aromatic Hydrocarbons (PAHs); EMS code: PA08 | Chemical |  | LAB
PA08 | ug/g | Benzo(k)fluoranthene | FS87 | DCM,Ac/Hex extr,PAM,HRGC/LRMS | DCM;Ac/Hex extr;PAM;HRGC/LRMS | 0.02 | Y | Y | Y | 250 | 4/12/2000 2:52:00 PM | 10/29/2022 8:18:00 AM | 1/11/2024 1:07:08 PM | 2001-02-23 |  | NPEPPIN |  | MassPerMass | mg/kg | Y | (mass conc.) |  | Benzo(k)fluoranthene (mass conc.) | 207-08-9 | Polycyclic Aromatic Hydrocarbons (PAHs); | Polycyclic Aromatic Hydrocarbons (PAHs); EMS code: PA08 | Chemical |  | LAB
PA08 | ug/g wet | Benzo(k)fluoranthene | X221 | Mod.Proc21.001(AOAC,'84) | Mod.Proc21.001(AOAC,'84) | 0.02 | Y | Y | Y | 72 |  |  | 10/24/1996 10:01:25 AM | 1996-04-12 | 6/3/1998 6:51:42 PM | EMS |  | MassPerMass | µg/g | Y | (mass conc.) | (wet) | Benzo(k)fluoranthene (wet) (mass conc.) | 207-08-9 | Polycyclic Aromatic Hydrocarbons (PAHs); | Polycyclic Aromatic Hydrocarbons (PAHs); EMS code: PA08 | Chemical |  | LAB
PA08 | ug/m3 | Benzo(k)fluoranthene | PAH1 | PUF:GC/MS | To understand the extent of human exposure to B[a]P and other PAHs, a reliable sampling and analytical method has been established which involves a combination quartz fibre filter/absorbent cartridge with subsequent extraction and analysis by gas chromatography (GC) with mass spectrometry (MS) detection (GC/MS). The analytical methods are a modification of EPA Test Method 625, Methods for Organic Chemical Analysis of Municipal and Industrial Wastewater, and Methods 8270, Test Methods for Evaluation of Solid Waste. PAH1 utilizes PUF as an adsorbant; PAH2 utilizes XAD-2 as an adsorbant. Note that isotopic surrogates may be added and are reported as % recovery. | 5e-05 | Y | Y | Y | 28 | 10/14/1997 2:50:00 PM | 4/22/1998 5:40:00 PM | 1/19/1999 6:45:10 PM | 1996-04-12 | 6/3/1998 6:51:37 PM | EMS |  | AirConcentration | µg/m³ | Y | (air conc.) |  | Benzo(k)fluoranthene (air conc.) | 207-08-9 | Polycyclic Aromatic Hydrocarbons (PAHs); | Polycyclic Aromatic Hydrocarbons (PAHs); EMS code: PA08 | Chemical |  | LAB
PA08 | ug/g | Benzo(k)fluoranthene | X214 | Mod.Proc21.001(AOAC,'84) | Mod.Proc21.001(AOAC,'84) | 0.02 | Y | Y | Y | 160 |  |  | 10/24/1996 2:26:15 PM | 1996-04-12 | 6/3/1998 6:51:37 PM | EMS |  | MassPerMass | mg/kg | Y | (mass conc.) |  | Benzo(k)fluoranthene (mass conc.) | 207-08-9 | Polycyclic Aromatic Hydrocarbons (PAHs); | Polycyclic Aromatic Hydrocarbons (PAHs); EMS code: PA08 | Chemical |  | LAB
PA08 | mg/L | Benzo(k)fluoranthene | MW02 | DCM extrn,GC/MS-SIM (PAH) | Water samples are liquid - liquid extracted with dichloromethane. The extract is then cleaned by silica gel column chromatography and analyzed by gas chromatography - mass spectroscopy. | 1e-05 | Y | Y | Y | 994 | 2/5/2002 | 7/10/2024 5:51:00 PM | 8/21/2024 11:42:55 AM | 1996-04-12 | 6/3/1998 6:51:42 PM | EMS |  | Concentration | mg/L | Y | (fl. conc.) |  | Benzo(k)fluoranthene (fl. conc.) | 207-08-9 | Polycyclic Aromatic Hydrocarbons (PAHs); | Polycyclic Aromatic Hydrocarbons (PAHs); EMS code: PA08 | Chemical |  | LAB
PA08 | mg/L | Benzo(k)fluoranthene | X262 | CH2Cl2 Ext:Silica:GC/MS | The sample is placed in a separatory funnel, the pH of the sample is adjusted to alkaline and then acidic conditions, and extracted each time with dichloromethane to selectively remove compounds. The final extracts (basic and acidic) can be combined or run separately. The extract is cleaned by silica gel column chromatography, then analyzed by GC/MS. | 0.001 | Y | Y | Y | 117 | 10/6/2000 | 1/13/2015 | 6/7/2024 5:30:18 PM | 1996-04-12 | 6/3/1998 6:51:37 PM | EMS |  | Concentration | mg/L | Y | (fl. conc.) |  | Benzo(k)fluoranthene (fl. conc.) | 207-08-9 | Polycyclic Aromatic Hydrocarbons (PAHs); | Polycyclic Aromatic Hydrocarbons (PAHs); EMS code: PA08 | Chemical |  | LAB
PA08 | mg/L | Benzo(k)fluoranthene | F010 | Extract disks,DCM extr:HRGC/LRMS | Water samples are spiked with a deuterated PAH Recovery Standards mix just prior to extraction with Empore extraction disks. Disks are eluted with Dichloromethane (DCM) and the extract concentrated. Deuterated PAH Internal Standards are added and the extract analyzed using High Resolution Gas Chromatography / Low Resolution Mass Spectrometry (HRGC/LRMS). MS data is acquired in Total Ion Mode (TIM) and quantitation is by multilevel Internal Standards. Recovery efficiency is calculated and reported but results are not recovery corrected. | 0.0005 | Y | Y | Y | 86 | 3/27/1997 4:23:00 PM | 3/11/2003 | 4/1/2004 1:13:26 AM | 1996-09-16 | 6/3/1998 6:51:44 PM | MPAK |  | Concentration | mg/L | Y | (fl. conc.) |  | Benzo(k)fluoranthene (fl. conc.) | 207-08-9 | Polycyclic Aromatic Hydrocarbons (PAHs); | Polycyclic Aromatic Hydrocarbons (PAHs); EMS code: PA08 | Chemical |  | LAB
PA09 | mg/L | Chrysene | HEGC | Hexane extraction: GC/MS | Hexane extraction: GC/MS | 1e-05 | Y | Y | Y | 905 | 6/8/2015 | 8/15/2024 | 8/16/2024 1:36:03 PM | 2015-07-07 | 8/15/2019 11:26:09 AM | RDGIBSON | KATFRASE | Concentration | mg/L | Y | (fl. conc.) |  | Chrysene (fl. conc.) | 218-01-9 | Polycyclic Aromatic Hydrocarbons (PAHs); | Polycyclic Aromatic Hydrocarbons (PAHs); EMS code: PA09 | Chemical |  | LAB
PA09 | ug/m3 | Chrysene | PAH1 | PUF:GC/MS | To understand the extent of human exposure to B[a]P and other PAHs, a reliable sampling and analytical method has been established which involves a combination quartz fibre filter/absorbent cartridge with subsequent extraction and analysis by gas chromatography (GC) with mass spectrometry (MS) detection (GC/MS). The analytical methods are a modification of EPA Test Method 625, Methods for Organic Chemical Analysis of Municipal and Industrial Wastewater, and Methods 8270, Test Methods for Evaluation of Solid Waste. PAH1 utilizes PUF as an adsorbant; PAH2 utilizes XAD-2 as an adsorbant. Note that isotopic surrogates may be added and are reported as % recovery. | 5e-05 | Y | Y | Y | 371 | 10/14/1997 2:50:00 PM | 4/22/1998 5:40:00 PM | 1/19/1999 6:45:10 PM | 1996-04-12 | 6/3/1998 6:51:37 PM | EMS |  | AirConcentration | µg/m³ | Y | (air conc.) |  | Chrysene (air conc.) | 218-01-9 | Polycyclic Aromatic Hydrocarbons (PAHs); | Polycyclic Aromatic Hydrocarbons (PAHs); EMS code: PA09 | Chemical |  | LAB
PA09 | ug/g | Chrysene | MP02 | GC/MS | GC/MS | 0.001 | Y | Y | Y | 523 | 3/28/2002 | 9/23/2021 3:24:00 AM | 10/18/2021 9:46:09 AM | 1996-04-12 | 6/3/1998 6:51:37 PM | EMS |  | MassPerMass | mg/kg | Y | (mass conc.) |  | Chrysene (mass conc.) | 218-01-9 | Polycyclic Aromatic Hydrocarbons (PAHs); | Polycyclic Aromatic Hydrocarbons (PAHs); EMS code: PA09 | Chemical |  | LAB
PA09 | ug/g wet | Chrysene | X221 | Mod.Proc21.001(AOAC,'84) | Mod.Proc21.001(AOAC,'84) | 0.01 | Y | Y | Y | 72 |  |  | 10/24/1996 10:01:25 AM | 1996-04-12 | 6/3/1998 6:51:42 PM | EMS |  | MassPerMass | µg/g | Y | (mass conc.) | (wet) | Chrysene (wet) (mass conc.) | 218-01-9 | Polycyclic Aromatic Hydrocarbons (PAHs); | Polycyclic Aromatic Hydrocarbons (PAHs); EMS code: PA09 | Chemical |  | LAB
PA09 | mg/L | Chrysene | F010 | Extract disks,DCM extr:HRGC/LRMS | Water samples are spiked with a deuterated PAH Recovery Standards mix just prior to extraction with Empore extraction disks. Disks are eluted with Dichloromethane (DCM) and the extract concentrated. Deuterated PAH Internal Standards are added and the extract analyzed using High Resolution Gas Chromatography / Low Resolution Mass Spectrometry (HRGC/LRMS). MS data is acquired in Total Ion Mode (TIM) and quantitation is by multilevel Internal Standards. Recovery efficiency is calculated and reported but results are not recovery corrected. | 0.0005 | Y | Y | Y | 86 | 3/27/1997 4:23:00 PM | 3/11/2003 | 4/1/2004 1:13:26 AM | 1996-09-16 | 6/3/1998 6:51:44 PM | MPAK |  | Concentration | mg/L | Y | (fl. conc.) |  | Chrysene (fl. conc.) | 218-01-9 | Polycyclic Aromatic Hydrocarbons (PAHs); | Polycyclic Aromatic Hydrocarbons (PAHs); EMS code: PA09 | Chemical |  | LAB
PA09 | ug/m3 | Chrysene | PAH2 | XAD2: GC/MS | To understand the extent of human exposure to B[a]P, and other PAHs, a reliable sampling and analytical method has been established which involves a combination quartz fibre filter/adsorbent cartridge with subsequent extraction and analysis by gas chromatography (GC) with mass spectrometry (MS) detection (GC/MS). The analytical methods are a modification of EPA Test Method 625, Methods for Organic Chemical Analysis of Municipal and Industrial Wastewater, and Methods 8270, Test Methods for Evaluation of Solid Waste. PAH1 utilizes PUF as an adsorbant; PAH2 utilizes XAD-2 as an adsorbant. Note that isotopic surrogates may be added and are reported as % recovery. | 5e-05 | Y | Y | Y | 223 |  |  | 10/24/1996 2:07:37 PM | 1996-04-12 | 6/3/1998 6:51:37 PM | EMS |  | AirConcentration | µg/m³ | Y | (air conc.) |  | Chrysene (air conc.) | 218-01-9 | Polycyclic Aromatic Hydrocarbons (PAHs); | Polycyclic Aromatic Hydrocarbons (PAHs); EMS code: PA09 | Chemical |  | LAB
PA09 | mg/L | Chrysene | MW02 | DCM extrn,GC/MS-SIM (PAH) | Water samples are liquid - liquid extracted with dichloromethane. The extract is then cleaned by silica gel column chromatography and analyzed by gas chromatography - mass spectroscopy. | 1e-05 | Y | Y | Y | 1065 | 2/5/2002 | 7/10/2024 5:51:00 PM | 8/21/2024 11:42:55 AM | 1996-04-12 | 6/3/1998 6:51:42 PM | EMS |  | Concentration | mg/L | Y | (fl. conc.) |  | Chrysene (fl. conc.) | 218-01-9 | Polycyclic Aromatic Hydrocarbons (PAHs); | Polycyclic Aromatic Hydrocarbons (PAHs); EMS code: PA09 | Chemical |  | LAB
PA09 | ug/g | Chrysene | X214 | Mod.Proc21.001(AOAC,'84) | Mod.Proc21.001(AOAC,'84) | 0.01 | Y | Y | Y | 160 |  |  | 10/24/1996 2:26:15 PM | 1996-04-12 | 6/3/1998 6:51:45 PM | EMS |  | MassPerMass | mg/kg | Y | (mass conc.) |  | Chrysene (mass conc.) | 218-01-9 | Polycyclic Aromatic Hydrocarbons (PAHs); | Polycyclic Aromatic Hydrocarbons (PAHs); EMS code: PA09 | Chemical |  | LAB
PA09 | ug/g | Chrysene | MI01 | Microwave Extractn /GC-MS(PAH) | Soil samples are ground with anhydrous sodium sulphate and microwave extracted with dichloromethane. The extract is then cleaned up by silica gel column chromatography and analyzed by gc-ms in selective ion mode. | 0.001 | Y | Y | Y | 51 | 6/1/2000 11:42:00 AM | 9/19/2001 11:39:00 AM | 3/17/2003 1:22:52 PM | 1996-10-13 | 6/3/1998 6:51:46 PM | EMS |  | MassPerMass | mg/kg | Y | (mass conc.) |  | Chrysene (mass conc.) | 218-01-9 | Polycyclic Aromatic Hydrocarbons (PAHs); | Polycyclic Aromatic Hydrocarbons (PAHs); EMS code: PA09 | Chemical |  | LAB
PA09 | mg/L | Chrysene | X262 | CH2Cl2 Ext:Silica:GC/MS | The sample is placed in a separatory funnel, the pH of the sample is adjusted to alkaline and then acidic conditions, and extracted each time with dichloromethane to selectively remove compounds. The final extracts (basic and acidic) can be combined or run separately. The extract is cleaned by silica gel column chromatography, then analyzed by GC/MS. | 0.001 | Y | Y | Y | 126 | 6/12/2000 | 1/13/2015 | 6/7/2024 5:30:18 PM | 1996-04-12 | 6/3/1998 6:51:37 PM | EMS |  | Concentration | mg/L | Y | (fl. conc.) |  | Chrysene (fl. conc.) | 218-01-9 | Polycyclic Aromatic Hydrocarbons (PAHs); | Polycyclic Aromatic Hydrocarbons (PAHs); EMS code: PA09 | Chemical |  | LAB
PA09 | ug/g | Chrysene | FS87 | DCM,Ac/Hex extr,PAM,HRGC/LRMS | DCM;Ac/Hex extr;PAM;HRGC/LRMS | 0.02 | Y | Y | Y | 250 | 4/12/2000 2:52:00 PM | 10/29/2022 8:18:00 AM | 1/11/2024 1:07:08 PM | 2001-02-23 |  | NPEPPIN |  | MassPerMass | mg/kg | Y | (mass conc.) |  | Chrysene (mass conc.) | 218-01-9 | Polycyclic Aromatic Hydrocarbons (PAHs); | Polycyclic Aromatic Hydrocarbons (PAHs); EMS code: PA09 | Chemical |  | LAB
PA10 | ug/g | Dibenzo(a;h)anthracene | FS87 | DCM,Ac/Hex extr,PAM,HRGC/LRMS | DCM;Ac/Hex extr;PAM;HRGC/LRMS | 0.02 | Y | Y | Y | 250 | 4/12/2000 2:52:00 PM | 10/29/2022 8:18:00 AM | 1/11/2024 1:07:08 PM | 2001-02-23 |  | NPEPPIN |  | MassPerMass | mg/kg | Y | (mass conc.) |  | Dibenzo(a;h)anthracene (mass conc.) | 53-70-3 | Polycyclic Aromatic Hydrocarbons (PAHs); | Polycyclic Aromatic Hydrocarbons (PAHs); EMS code: PA10 | Chemical |  | LAB
PA10 | ug/m3 | Dibenzo(a;h)anthracene | PAH2 | XAD2: GC/MS | To understand the extent of human exposure to B[a]P, and other PAHs, a reliable sampling and analytical method has been established which involves a combination quartz fibre filter/adsorbent cartridge with subsequent extraction and analysis by gas chromatography (GC) with mass spectrometry (MS) detection (GC/MS). The analytical methods are a modification of EPA Test Method 625, Methods for Organic Chemical Analysis of Municipal and Industrial Wastewater, and Methods 8270, Test Methods for Evaluation of Solid Waste. PAH1 utilizes PUF as an adsorbant; PAH2 utilizes XAD-2 as an adsorbant. Note that isotopic surrogates may be added and are reported as % recovery. | 0.0001 | Y | Y | Y | 223 |  |  | 10/24/1996 2:07:37 PM | 1996-04-12 | 6/3/1998 6:51:37 PM | EMS |  | AirConcentration | µg/m³ | Y | (air conc.) |  | Dibenzo(a;h)anthracene (air conc.) | 53-70-3 | Polycyclic Aromatic Hydrocarbons (PAHs); | Polycyclic Aromatic Hydrocarbons (PAHs); EMS code: PA10 | Chemical |  | LAB
PA10 | ug/g | Dibenzo(a;h)anthracene | MP02 | GC/MS | GC/MS | 0.001 | Y | Y | Y | 523 | 3/28/2002 | 9/23/2021 3:24:00 AM | 10/18/2021 9:46:09 AM | 1996-04-12 | 6/3/1998 6:51:37 PM | EMS |  | MassPerMass | mg/kg | Y | (mass conc.) |  | Dibenzo(a;h)anthracene (mass conc.) | 53-70-3 | Polycyclic Aromatic Hydrocarbons (PAHs); | Polycyclic Aromatic Hydrocarbons (PAHs); EMS code: PA10 | Chemical |  | LAB
PA10 | mg/L | Dibenzo(a;h)anthracene | X262 | CH2Cl2 Ext:Silica:GC/MS | The sample is placed in a separatory funnel, the pH of the sample is adjusted to alkaline and then acidic conditions, and extracted each time with dichloromethane to selectively remove compounds. The final extracts (basic and acidic) can be combined or run separately. The extract is cleaned by silica gel column chromatography, then analyzed by GC/MS. | 0.001 | Y | Y | Y | 126 | 6/12/2000 | 1/13/2015 | 6/7/2024 5:30:18 PM | 1996-04-12 | 6/3/1998 6:51:37 PM | EMS |  | Concentration | mg/L | Y | (fl. conc.) |  | Dibenzo(a;h)anthracene (fl. conc.) | 53-70-3 | Polycyclic Aromatic Hydrocarbons (PAHs); | Polycyclic Aromatic Hydrocarbons (PAHs); EMS code: PA10 | Chemical |  | LAB
PA10 | ug/g | Dibenzo(a;h)anthracene | X214 | Mod.Proc21.001(AOAC,'84) | Mod.Proc21.001(AOAC,'84) | 0.02 | Y | Y | Y | 160 |  |  | 10/24/1996 2:26:15 PM | 1996-04-12 | 6/3/1998 6:51:37 PM | EMS |  | MassPerMass | mg/kg | Y | (mass conc.) |  | Dibenzo(a;h)anthracene (mass conc.) | 53-70-3 | Polycyclic Aromatic Hydrocarbons (PAHs); | Polycyclic Aromatic Hydrocarbons (PAHs); EMS code: PA10 | Chemical |  | LAB
PA10 | ug/g | Dibenzo(a;h)anthracene | MI01 | Microwave Extractn /GC-MS(PAH) | Soil samples are ground with anhydrous sodium sulphate and microwave extracted with dichloromethane. The extract is then cleaned up by silica gel column chromatography and analyzed by gc-ms in selective ion mode. | 0.001 | Y | Y | Y | 51 | 6/1/2000 11:42:00 AM | 9/19/2001 11:39:00 AM | 3/17/2003 1:22:53 PM | 1996-10-13 | 6/3/1998 6:51:46 PM | EMS |  | MassPerMass | mg/kg | Y | (mass conc.) |  | Dibenzo(a;h)anthracene (mass conc.) | 53-70-3 | Polycyclic Aromatic Hydrocarbons (PAHs); | Polycyclic Aromatic Hydrocarbons (PAHs); EMS code: PA10 | Chemical |  | LAB
PA10 | mg/L | Dibenzo(a;h)anthracene | HEGC | Hexane extraction: GC/MS | Hexane extraction: GC/MS | 5e-06 | Y | Y | Y | 898 | 6/8/2015 | 8/15/2024 | 8/16/2024 1:36:03 PM | 2015-07-07 | 8/15/2019 11:26:41 AM | RDGIBSON | KATFRASE | Concentration | mg/L | Y | (fl. conc.) |  | Dibenzo(a;h)anthracene (fl. conc.) | 53-70-3 | Polycyclic Aromatic Hydrocarbons (PAHs); | Polycyclic Aromatic Hydrocarbons (PAHs); EMS code: PA10 | Chemical |  | LAB
PA10 | ug/m3 | Dibenzo(a;h)anthracene | PAH1 | PUF:GC/MS | To understand the extent of human exposure to B[a]P and other PAHs, a reliable sampling and analytical method has been established which involves a combination quartz fibre filter/absorbent cartridge with subsequent extraction and analysis by gas chromatography (GC) with mass spectrometry (MS) detection (GC/MS). The analytical methods are a modification of EPA Test Method 625, Methods for Organic Chemical Analysis of Municipal and Industrial Wastewater, and Methods 8270, Test Methods for Evaluation of Solid Waste. PAH1 utilizes PUF as an adsorbant; PAH2 utilizes XAD-2 as an adsorbant. Note that isotopic surrogates may be added and are reported as % recovery. | 0.0001 | Y | Y | Y | 372 | 10/14/1997 2:50:00 PM | 4/22/1998 5:40:00 PM | 1/19/1999 6:45:10 PM | 1996-04-12 | 6/3/1998 6:51:42 PM | EMS |  | AirConcentration | µg/m³ | Y | (air conc.) |  | Dibenzo(a;h)anthracene (air conc.) | 53-70-3 | Polycyclic Aromatic Hydrocarbons (PAHs); | Polycyclic Aromatic Hydrocarbons (PAHs); EMS code: PA10 | Chemical |  | LAB
PA10 | ug/g wet | Dibenzo(a;h)anthracene | X221 | Mod.Proc21.001(AOAC,'84) | Mod.Proc21.001(AOAC,'84) | 0.02 | Y | Y | Y | 73 |  |  | 10/24/1996 10:01:25 AM | 1996-04-12 | 6/3/1998 6:51:42 PM | EMS |  | MassPerMass | µg/g | Y | (mass conc.) | (wet) | Dibenzo(a;h)anthracene (wet) (mass conc.) | 53-70-3 | Polycyclic Aromatic Hydrocarbons (PAHs); | Polycyclic Aromatic Hydrocarbons (PAHs); EMS code: PA10 | Chemical |  | LAB
PA10 | mg/L | Dibenzo(a;h)anthracene | F010 | Extract disks,DCM extr:HRGC/LRMS | Water samples are spiked with a deuterated PAH Recovery Standards mix just prior to extraction with Empore extraction disks. Disks are eluted with Dichloromethane (DCM) and the extract concentrated. Deuterated PAH Internal Standards are added and the extract analyzed using High Resolution Gas Chromatography / Low Resolution Mass Spectrometry (HRGC/LRMS). MS data is acquired in Total Ion Mode (TIM) and quantitation is by multilevel Internal Standards. Recovery efficiency is calculated and reported but results are not recovery corrected. | 0.005 | Y | Y | Y | 86 | 3/27/1997 4:23:00 PM | 3/11/2003 | 4/1/2004 1:13:26 AM | 1996-09-16 | 6/3/1998 6:51:46 PM | MPAK |  | Concentration | mg/L | Y | (fl. conc.) |  | Dibenzo(a;h)anthracene (fl. conc.) | 53-70-3 | Polycyclic Aromatic Hydrocarbons (PAHs); | Polycyclic Aromatic Hydrocarbons (PAHs); EMS code: PA10 | Chemical |  | LAB
PA10 | mg/L | Dibenzo(a;h)anthracene | MW02 | DCM extrn,GC/MS-SIM (PAH) | Water samples are liquid - liquid extracted with dichloromethane. The extract is then cleaned by silica gel column chromatography and analyzed by gas chromatography - mass spectroscopy. | 2e-05 | Y | Y | Y | 1018 | 2/5/2002 | 7/10/2024 5:51:00 PM | 8/21/2024 11:42:55 AM | 1996-04-12 | 6/3/1998 6:51:42 PM | EMS |  | Concentration | mg/L | Y | (fl. conc.) |  | Dibenzo(a;h)anthracene (fl. conc.) | 53-70-3 | Polycyclic Aromatic Hydrocarbons (PAHs); | Polycyclic Aromatic Hydrocarbons (PAHs); EMS code: PA10 | Chemical |  | LAB
PA11 | mg/L | Fluoranthene | MW02 | DCM extrn,GC/MS-SIM (PAH) | Water samples are liquid - liquid extracted with dichloromethane. The extract is then cleaned by silica gel column chromatography and analyzed by gas chromatography - mass spectroscopy. | 1e-05 | Y | Y | Y | 1130 | 2/5/2002 | 7/10/2024 5:51:00 PM | 8/21/2024 11:42:55 AM | 1996-04-12 | 6/3/1998 6:51:45 PM | EMS |  | Concentration | mg/L | Y | (fl. conc.) |  | Fluoranthene (fl. conc.) | 206-44-0 | Polycyclic Aromatic Hydrocarbons (PAHs); | Polycyclic Aromatic Hydrocarbons (PAHs); EMS code: PA11 | Chemical |  | LAB
PA11 | ug/g | Fluoranthene | X214 | Mod.Proc21.001(AOAC,'84) | Mod.Proc21.001(AOAC,'84) | 0.01 | Y | Y | Y | 160 |  |  | 10/24/1996 2:26:15 PM | 1996-04-12 | 6/3/1998 6:51:37 PM | EMS |  | MassPerMass | mg/kg | Y | (mass conc.) |  | Fluoranthene (mass conc.) | 206-44-0 | Polycyclic Aromatic Hydrocarbons (PAHs); | Polycyclic Aromatic Hydrocarbons (PAHs); EMS code: PA11 | Chemical |  | LAB
PA11 | mg/L | Fluoranthene | X262 | CH2Cl2 Ext:Silica:GC/MS | The sample is placed in a separatory funnel, the pH of the sample is adjusted to alkaline and then acidic conditions, and extracted each time with dichloromethane to selectively remove compounds. The final extracts (basic and acidic) can be combined or run separately. The extract is cleaned by silica gel column chromatography, then analyzed by GC/MS. | 0.001 | Y | Y | Y | 126 | 6/12/2000 | 1/13/2015 | 6/7/2024 5:30:18 PM | 1996-04-12 | 6/3/1998 6:51:37 PM | EMS |  | Concentration | mg/L | Y | (fl. conc.) |  | Fluoranthene (fl. conc.) | 206-44-0 | Polycyclic Aromatic Hydrocarbons (PAHs); | Polycyclic Aromatic Hydrocarbons (PAHs); EMS code: PA11 | Chemical |  | LAB
PA11 | ug/m3 | Fluoranthene | PAH1 | PUF:GC/MS | To understand the extent of human exposure to B[a]P and other PAHs, a reliable sampling and analytical method has been established which involves a combination quartz fibre filter/absorbent cartridge with subsequent extraction and analysis by gas chromatography (GC) with mass spectrometry (MS) detection (GC/MS). The analytical methods are a modification of EPA Test Method 625, Methods for Organic Chemical Analysis of Municipal and Industrial Wastewater, and Methods 8270, Test Methods for Evaluation of Solid Waste. PAH1 utilizes PUF as an adsorbant; PAH2 utilizes XAD-2 as an adsorbant. Note that isotopic surrogates may be added and are reported as % recovery. | 5e-05 | Y | Y | Y | 372 | 10/14/1997 2:50:00 PM | 4/22/1998 5:40:00 PM | 1/19/1999 6:45:10 PM | 1996-04-12 | 6/3/1998 6:51:37 PM | EMS |  | AirConcentration | µg/m³ | Y | (air conc.) |  | Fluoranthene (air conc.) | 206-44-0 | Polycyclic Aromatic Hydrocarbons (PAHs); | Polycyclic Aromatic Hydrocarbons (PAHs); EMS code: PA11 | Chemical |  | LAB
PA11 | mg/L | Fluoranthene | HEGC | Hexane extraction: GC/MS | Hexane extraction: GC/MS | 1e-05 | Y | Y | Y | 905 | 6/8/2015 | 8/15/2024 | 8/16/2024 1:36:03 PM | 2015-07-07 | 8/15/2019 11:27:02 AM | RDGIBSON | KATFRASE | Concentration | mg/L | Y | (fl. conc.) |  | Fluoranthene (fl. conc.) | 206-44-0 | Polycyclic Aromatic Hydrocarbons (PAHs); | Polycyclic Aromatic Hydrocarbons (PAHs); EMS code: PA11 | Chemical |  | LAB
PA11 | ug/g wet | Fluoranthene | X221 | Mod.Proc21.001(AOAC,'84) | Mod.Proc21.001(AOAC,'84) | 0.01 | Y | Y | Y | 73 |  |  | 10/24/1996 10:01:25 AM | 1996-04-12 | 6/3/1998 6:51:42 PM | EMS |  | MassPerMass | µg/g | Y | (mass conc.) | (wet) | Fluoranthene (wet) (mass conc.) | 206-44-0 | Polycyclic Aromatic Hydrocarbons (PAHs); | Polycyclic Aromatic Hydrocarbons (PAHs); EMS code: PA11 | Chemical |  | LAB
PA11 | ug/g | Fluoranthene | FS87 | DCM,Ac/Hex extr,PAM,HRGC/LRMS | DCM;Ac/Hex extr;PAM;HRGC/LRMS | 0.02 | Y | Y | Y | 250 | 4/12/2000 2:52:00 PM | 10/29/2022 8:18:00 AM | 1/11/2024 1:07:08 PM | 2001-02-23 |  | NPEPPIN |  | MassPerMass | mg/kg | Y | (mass conc.) |  | Fluoranthene (mass conc.) | 206-44-0 | Polycyclic Aromatic Hydrocarbons (PAHs); | Polycyclic Aromatic Hydrocarbons (PAHs); EMS code: PA11 | Chemical |  | LAB
PA11 | ug/m3 | Fluoranthene | PAH2 | XAD2: GC/MS | To understand the extent of human exposure to B[a]P, and other PAHs, a reliable sampling and analytical method has been established which involves a combination quartz fibre filter/adsorbent cartridge with subsequent extraction and analysis by gas chromatography (GC) with mass spectrometry (MS) detection (GC/MS). The analytical methods are a modification of EPA Test Method 625, Methods for Organic Chemical Analysis of Municipal and Industrial Wastewater, and Methods 8270, Test Methods for Evaluation of Solid Waste. PAH1 utilizes PUF as an adsorbant; PAH2 utilizes XAD-2 as an adsorbant. Note that isotopic surrogates may be added and are reported as % recovery. | 5e-05 | Y | Y | Y | 223 |  |  | 10/24/1996 2:07:37 PM | 1996-04-12 | 6/3/1998 6:51:37 PM | EMS |  | AirConcentration | µg/m³ | Y | (air conc.) |  | Fluoranthene (air conc.) | 206-44-0 | Polycyclic Aromatic Hydrocarbons (PAHs); | Polycyclic Aromatic Hydrocarbons (PAHs); EMS code: PA11 | Chemical |  | LAB
PA11 | mg/L | Fluoranthene | F010 | Extract disks,DCM extr:HRGC/LRMS | Water samples are spiked with a deuterated PAH Recovery Standards mix just prior to extraction with Empore extraction disks. Disks are eluted with Dichloromethane (DCM) and the extract concentrated. Deuterated PAH Internal Standards are added and the extract analyzed using High Resolution Gas Chromatography / Low Resolution Mass Spectrometry (HRGC/LRMS). MS data is acquired in Total Ion Mode (TIM) and quantitation is by multilevel Internal Standards. Recovery efficiency is calculated and reported but results are not recovery corrected. | 0.0002 | Y | Y | Y | 86 | 3/27/1997 4:23:00 PM | 3/11/2003 | 4/1/2004 1:13:26 AM | 1996-09-16 | 6/3/1998 6:51:44 PM | MPAK |  | Concentration | mg/L | Y | (fl. conc.) |  | Fluoranthene (fl. conc.) | 206-44-0 | Polycyclic Aromatic Hydrocarbons (PAHs); | Polycyclic Aromatic Hydrocarbons (PAHs); EMS code: PA11 | Chemical |  | LAB
PA11 | ug/g | Fluoranthene | MP02 | GC/MS | GC/MS | 0.001 | Y | Y | Y | 523 | 3/28/2002 | 9/23/2021 3:24:00 AM | 10/18/2021 9:46:09 AM | 1996-04-12 | 6/3/1998 6:51:37 PM | EMS |  | MassPerMass | mg/kg | Y | (mass conc.) |  | Fluoranthene (mass conc.) | 206-44-0 | Polycyclic Aromatic Hydrocarbons (PAHs); | Polycyclic Aromatic Hydrocarbons (PAHs); EMS code: PA11 | Chemical |  | LAB
PA11 | ug/g | Fluoranthene | MI01 | Microwave Extractn /GC-MS(PAH) | Soil samples are ground with anhydrous sodium sulphate and microwave extracted with dichloromethane. The extract is then cleaned up by silica gel column chromatography and analyzed by gc-ms in selective ion mode. | 0.001 | Y | Y | Y | 51 | 6/1/2000 11:42:00 AM | 9/19/2001 11:39:00 AM | 3/17/2003 1:22:53 PM | 1996-10-13 | 6/3/1998 6:51:46 PM | EMS |  | MassPerMass | mg/kg | Y | (mass conc.) |  | Fluoranthene (mass conc.) | 206-44-0 | Polycyclic Aromatic Hydrocarbons (PAHs); | Polycyclic Aromatic Hydrocarbons (PAHs); EMS code: PA11 | Chemical |  | LAB
PA12 | ug/g | Fluorene | MI01 | Microwave Extractn /GC-MS(PAH) | Soil samples are ground with anhydrous sodium sulphate and microwave extracted with dichloromethane. The extract is then cleaned up by silica gel column chromatography and analyzed by gc-ms in selective ion mode. | 0.001 | Y | Y | Y | 51 | 6/1/2000 11:42:00 AM | 9/19/2001 11:39:00 AM | 3/17/2003 1:22:53 PM | 1996-10-13 | 6/3/1998 6:51:46 PM | EMS |  | MassPerMass | mg/kg | Y | (mass conc.) |  | Fluorene (mass conc.) | 86-73-7 | Polycyclic Aromatic Hydrocarbons (PAHs); | Polycyclic Aromatic Hydrocarbons (PAHs); EMS code: PA12 | Chemical |  | LAB
PA12 | ug/g | Fluorene | FS87 | DCM,Ac/Hex extr,PAM,HRGC/LRMS | DCM;Ac/Hex extr;PAM;HRGC/LRMS | 0.02 | Y | Y | Y | 250 | 4/12/2000 2:52:00 PM | 10/29/2022 8:18:00 AM | 1/11/2024 1:07:08 PM | 2001-02-23 |  | NPEPPIN |  | MassPerMass | mg/kg | Y | (mass conc.) |  | Fluorene (mass conc.) | 86-73-7 | Polycyclic Aromatic Hydrocarbons (PAHs); | Polycyclic Aromatic Hydrocarbons (PAHs); EMS code: PA12 | Chemical |  | LAB
PA12 | ug/m3 | Fluorene | PAH1 | PUF:GC/MS | To understand the extent of human exposure to B[a]P and other PAHs, a reliable sampling and analytical method has been established which involves a combination quartz fibre filter/absorbent cartridge with subsequent extraction and analysis by gas chromatography (GC) with mass spectrometry (MS) detection (GC/MS). The analytical methods are a modification of EPA Test Method 625, Methods for Organic Chemical Analysis of Municipal and Industrial Wastewater, and Methods 8270, Test Methods for Evaluation of Solid Waste. PAH1 utilizes PUF as an adsorbant; PAH2 utilizes XAD-2 as an adsorbant. Note that isotopic surrogates may be added and are reported as % recovery. | 5e-05 | Y | Y | Y | 372 | 10/14/1997 2:50:00 PM | 4/22/1998 5:40:00 PM | 1/19/1999 6:45:10 PM | 1996-04-12 | 6/3/1998 6:51:37 PM | EMS |  | AirConcentration | µg/m³ | Y | (air conc.) |  | Fluorene (air conc.) | 86-73-7 | Polycyclic Aromatic Hydrocarbons (PAHs); | Polycyclic Aromatic Hydrocarbons (PAHs); EMS code: PA12 | Chemical |  | LAB
PA12 | ug/m3 | Fluorene | PAH2 | XAD2: GC/MS | To understand the extent of human exposure to B[a]P, and other PAHs, a reliable sampling and analytical method has been established which involves a combination quartz fibre filter/adsorbent cartridge with subsequent extraction and analysis by gas chromatography (GC) with mass spectrometry (MS) detection (GC/MS). The analytical methods are a modification of EPA Test Method 625, Methods for Organic Chemical Analysis of Municipal and Industrial Wastewater, and Methods 8270, Test Methods for Evaluation of Solid Waste. PAH1 utilizes PUF as an adsorbant; PAH2 utilizes XAD-2 as an adsorbant. Note that isotopic surrogates may be added and are reported as % recovery. | 5e-05 | Y | Y | Y | 223 |  |  | 10/24/1996 2:07:37 PM | 1996-04-12 | 6/3/1998 6:51:37 PM | EMS |  | AirConcentration | µg/m³ | Y | (air conc.) |  | Fluorene (air conc.) | 86-73-7 | Polycyclic Aromatic Hydrocarbons (PAHs); | Polycyclic Aromatic Hydrocarbons (PAHs); EMS code: PA12 | Chemical |  | LAB
PA12 | ug/g wet | Fluorene | X221 | Mod.Proc21.001(AOAC,'84) | Mod.Proc21.001(AOAC,'84) | 0.004 | Y | Y | Y | 73 |  |  | 10/24/1996 10:01:25 AM | 1996-04-12 | 6/3/1998 6:51:42 PM | EMS |  | MassPerMass | µg/g | Y | (mass conc.) | (wet) | Fluorene (wet) (mass conc.) | 86-73-7 | Polycyclic Aromatic Hydrocarbons (PAHs); | Polycyclic Aromatic Hydrocarbons (PAHs); EMS code: PA12 | Chemical |  | LAB
PA12 | mg/L | Fluorene | HEGC | Hexane extraction: GC/MS | Hexane extraction: GC/MS | 1e-05 | Y | Y | Y | 905 | 6/8/2015 | 8/15/2024 | 8/16/2024 1:36:03 PM | 2015-07-07 | 8/15/2019 11:27:22 AM | RDGIBSON | KATFRASE | Concentration | mg/L | Y | (fl. conc.) |  | Fluorene (fl. conc.) | 86-73-7 | Polycyclic Aromatic Hydrocarbons (PAHs); | Polycyclic Aromatic Hydrocarbons (PAHs); EMS code: PA12 | Chemical |  | LAB
PA12 | ug/g | Fluorene | X214 | Mod.Proc21.001(AOAC,'84) | Mod.Proc21.001(AOAC,'84) | 0.005 | Y | Y | Y | 160 |  |  | 10/24/1996 2:26:15 PM | 1996-04-12 | 6/3/1998 6:51:45 PM | EMS |  | MassPerMass | mg/kg | Y | (mass conc.) |  | Fluorene (mass conc.) | 86-73-7 | Polycyclic Aromatic Hydrocarbons (PAHs); | Polycyclic Aromatic Hydrocarbons (PAHs); EMS code: PA12 | Chemical |  | LAB
PA12 | mg/L | Fluorene | X262 | CH2Cl2 Ext:Silica:GC/MS | The sample is placed in a separatory funnel, the pH of the sample is adjusted to alkaline and then acidic conditions, and extracted each time with dichloromethane to selectively remove compounds. The final extracts (basic and acidic) can be combined or run separately. The extract is cleaned by silica gel column chromatography, then analyzed by GC/MS. | 0.001 | Y | Y | Y | 126 | 6/12/2000 | 1/13/2015 | 6/7/2024 5:30:18 PM | 1996-04-12 | 6/3/1998 6:51:37 PM | EMS |  | Concentration | mg/L | Y | (fl. conc.) |  | Fluorene (fl. conc.) | 86-73-7 | Polycyclic Aromatic Hydrocarbons (PAHs); | Polycyclic Aromatic Hydrocarbons (PAHs); EMS code: PA12 | Chemical |  | LAB
PA12 | mg/L | Fluorene | MW02 | DCM extrn,GC/MS-SIM (PAH) | Water samples are liquid - liquid extracted with dichloromethane. The extract is then cleaned by silica gel column chromatography and analyzed by gas chromatography - mass spectroscopy. | 1e-05 | Y | Y | Y | 1130 | 2/5/2002 | 7/10/2024 5:51:00 PM | 8/21/2024 11:42:55 AM | 1996-04-12 | 6/3/1998 6:51:42 PM | EMS |  | Concentration | mg/L | Y | (fl. conc.) |  | Fluorene (fl. conc.) | 86-73-7 | Polycyclic Aromatic Hydrocarbons (PAHs); | Polycyclic Aromatic Hydrocarbons (PAHs); EMS code: PA12 | Chemical |  | LAB
PA12 | ug/g | Fluorene | MP02 | GC/MS | GC/MS | 0.001 | Y | Y | Y | 523 | 3/28/2002 | 9/23/2021 3:24:00 AM | 10/18/2021 9:46:09 AM | 1996-04-12 | 6/3/1998 6:51:37 PM | EMS |  | MassPerMass | mg/kg | Y | (mass conc.) |  | Fluorene (mass conc.) | 86-73-7 | Polycyclic Aromatic Hydrocarbons (PAHs); | Polycyclic Aromatic Hydrocarbons (PAHs); EMS code: PA12 | Chemical |  | LAB
PA12 | mg/L | Fluorene | F010 | Extract disks,DCM extr:HRGC/LRMS | Water samples are spiked with a deuterated PAH Recovery Standards mix just prior to extraction with Empore extraction disks. Disks are eluted with Dichloromethane (DCM) and the extract concentrated. Deuterated PAH Internal Standards are added and the extract analyzed using High Resolution Gas Chromatography / Low Resolution Mass Spectrometry (HRGC/LRMS). MS data is acquired in Total Ion Mode (TIM) and quantitation is by multilevel Internal Standards. Recovery efficiency is calculated and reported but results are not recovery corrected. | 0.0002 | Y | Y | Y | 86 | 3/27/1997 4:24:00 PM | 3/11/2003 | 4/1/2004 1:13:26 AM | 1996-09-16 | 6/3/1998 6:51:44 PM | MPAK |  | Concentration | mg/L | Y | (fl. conc.) |  | Fluorene (fl. conc.) | 86-73-7 | Polycyclic Aromatic Hydrocarbons (PAHs); | Polycyclic Aromatic Hydrocarbons (PAHs); EMS code: PA12 | Chemical |  | LAB
PA13 | ug/g | Indeno(1;2;3-cd)pyrene | X214 | Mod.Proc21.001(AOAC,'84) | Mod.Proc21.001(AOAC,'84) | 0.02 | Y | Y | Y | 160 |  |  | 10/24/1996 2:26:15 PM | 1996-04-12 | 6/3/1998 6:51:45 PM | EMS |  | MassPerMass | mg/kg | Y | (mass conc.) |  | Indeno(1;2;3-cd)pyrene (mass conc.) | 193-39-5 | Polycyclic Aromatic Hydrocarbons (PAHs); | Polycyclic Aromatic Hydrocarbons (PAHs); EMS code: PA13 | Chemical |  | LAB
PA13 | mg/L | Indeno(1;2;3-cd)pyrene | F010 | Extract disks,DCM extr:HRGC/LRMS | Water samples are spiked with a deuterated PAH Recovery Standards mix just prior to extraction with Empore extraction disks. Disks are eluted with Dichloromethane (DCM) and the extract concentrated. Deuterated PAH Internal Standards are added and the extract analyzed using High Resolution Gas Chromatography / Low Resolution Mass Spectrometry (HRGC/LRMS). MS data is acquired in Total Ion Mode (TIM) and quantitation is by multilevel Internal Standards. Recovery efficiency is calculated and reported but results are not recovery corrected. | 0.005 | Y | Y | Y | 86 | 3/27/1997 4:24:00 PM | 3/11/2003 | 4/1/2004 1:13:26 AM | 1996-09-16 | 6/3/1998 6:51:44 PM | MPAK |  | Concentration | mg/L | Y | (fl. conc.) |  | Indeno(1;2;3-cd)pyrene (fl. conc.) | 193-39-5 | Polycyclic Aromatic Hydrocarbons (PAHs); | Polycyclic Aromatic Hydrocarbons (PAHs); EMS code: PA13 | Chemical |  | LAB
PA13 | mg/L | Indeno(1;2;3-cd)pyrene | X262 | CH2Cl2 Ext:Silica:GC/MS | The sample is placed in a separatory funnel, the pH of the sample is adjusted to alkaline and then acidic conditions, and extracted each time with dichloromethane to selectively remove compounds. The final extracts (basic and acidic) can be combined or run separately. The extract is cleaned by silica gel column chromatography, then analyzed by GC/MS. | 0.001 | Y | Y | Y | 126 | 6/12/2000 | 1/13/2015 | 6/7/2024 5:30:18 PM | 1996-04-12 | 6/3/1998 6:51:37 PM | EMS |  | Concentration | mg/L | Y | (fl. conc.) |  | Indeno(1;2;3-cd)pyrene (fl. conc.) | 193-39-5 | Polycyclic Aromatic Hydrocarbons (PAHs); | Polycyclic Aromatic Hydrocarbons (PAHs); EMS code: PA13 | Chemical |  | LAB
PA13 | mg/L | Indeno(1;2;3-cd)pyrene | MW02 | DCM extrn,GC/MS-SIM (PAH) | Water samples are liquid - liquid extracted with dichloromethane. The extract is then cleaned by silica gel column chromatography and analyzed by gas chromatography - mass spectroscopy. | 2e-05 | Y | Y | Y | 1018 | 2/5/2002 | 7/10/2024 5:51:00 PM | 8/21/2024 11:42:55 AM | 1996-04-12 | 6/3/1998 6:51:42 PM | EMS |  | Concentration | mg/L | Y | (fl. conc.) |  | Indeno(1;2;3-cd)pyrene (fl. conc.) | 193-39-5 | Polycyclic Aromatic Hydrocarbons (PAHs); | Polycyclic Aromatic Hydrocarbons (PAHs); EMS code: PA13 | Chemical |  | LAB
PA13 | ug/g | Indeno(1;2;3-cd)pyrene | MP02 | GC/MS | GC/MS | 0.001 | Y | Y | Y | 523 | 3/28/2002 | 9/23/2021 3:24:00 AM | 10/18/2021 9:46:09 AM | 1996-04-12 | 6/3/1998 6:51:37 PM | EMS |  | MassPerMass | mg/kg | Y | (mass conc.) |  | Indeno(1;2;3-cd)pyrene (mass conc.) | 193-39-5 | Polycyclic Aromatic Hydrocarbons (PAHs); | Polycyclic Aromatic Hydrocarbons (PAHs); EMS code: PA13 | Chemical |  | LAB
PA13 | ug/g | Indeno(1;2;3-cd)pyrene | MI01 | Microwave Extractn /GC-MS(PAH) | Soil samples are ground with anhydrous sodium sulphate and microwave extracted with dichloromethane. The extract is then cleaned up by silica gel column chromatography and analyzed by gc-ms in selective ion mode. | 0.001 | Y | Y | Y | 51 | 6/1/2000 11:42:00 AM | 9/19/2001 11:39:00 AM | 3/17/2003 1:22:53 PM | 1996-10-13 | 6/3/1998 6:51:46 PM | EMS |  | MassPerMass | mg/kg | Y | (mass conc.) |  | Indeno(1;2;3-cd)pyrene (mass conc.) | 193-39-5 | Polycyclic Aromatic Hydrocarbons (PAHs); | Polycyclic Aromatic Hydrocarbons (PAHs); EMS code: PA13 | Chemical |  | LAB
PA13 | ug/g wet | Indeno(1;2;3-cd)pyrene | X221 | Mod.Proc21.001(AOAC,'84) | Mod.Proc21.001(AOAC,'84) | 0.02 | Y | Y | Y | 74 |  |  | 10/24/1996 10:01:27 AM | 1996-04-12 | 6/3/1998 6:51:42 PM | EMS |  | MassPerMass | µg/g | Y | (mass conc.) | (wet) | Indeno(1;2;3-cd)pyrene (wet) (mass conc.) | 193-39-5 | Polycyclic Aromatic Hydrocarbons (PAHs); | Polycyclic Aromatic Hydrocarbons (PAHs); EMS code: PA13 | Chemical |  | LAB
PA13 | ug/m3 | Indeno(1;2;3-cd)pyrene | PAH1 | PUF:GC/MS | To understand the extent of human exposure to B[a]P and other PAHs, a reliable sampling and analytical method has been established which involves a combination quartz fibre filter/absorbent cartridge with subsequent extraction and analysis by gas chromatography (GC) with mass spectrometry (MS) detection (GC/MS). The analytical methods are a modification of EPA Test Method 625, Methods for Organic Chemical Analysis of Municipal and Industrial Wastewater, and Methods 8270, Test Methods for Evaluation of Solid Waste. PAH1 utilizes PUF as an adsorbant; PAH2 utilizes XAD-2 as an adsorbant. Note that isotopic surrogates may be added and are reported as % recovery. | 0.0001 | Y | Y | Y | 372 | 10/14/1997 2:50:00 PM | 4/22/1998 5:40:00 PM | 1/19/1999 6:45:10 PM | 1996-04-12 | 6/3/1998 6:51:37 PM | EMS |  | AirConcentration | µg/m³ | Y | (air conc.) |  | Indeno(1;2;3-cd)pyrene (air conc.) | 193-39-5 | Polycyclic Aromatic Hydrocarbons (PAHs); | Polycyclic Aromatic Hydrocarbons (PAHs); EMS code: PA13 | Chemical |  | LAB
PA13 | ug/g | Indeno(1;2;3-cd)pyrene | FS87 | DCM,Ac/Hex extr,PAM,HRGC/LRMS | DCM;Ac/Hex extr;PAM;HRGC/LRMS | 0.02 | Y | Y | Y | 250 | 4/12/2000 2:52:00 PM | 10/29/2022 8:18:00 AM | 1/11/2024 1:07:08 PM | 2001-02-23 |  | NPEPPIN |  | MassPerMass | mg/kg | Y | (mass conc.) |  | Indeno(1;2;3-cd)pyrene (mass conc.) | 193-39-5 | Polycyclic Aromatic Hydrocarbons (PAHs); | Polycyclic Aromatic Hydrocarbons (PAHs); EMS code: PA13 | Chemical |  | LAB
PA13 | ug/m3 | Indeno(1;2;3-cd)pyrene | PAH2 | XAD2: GC/MS | To understand the extent of human exposure to B[a]P, and other PAHs, a reliable sampling and analytical method has been established which involves a combination quartz fibre filter/adsorbent cartridge with subsequent extraction and analysis by gas chromatography (GC) with mass spectrometry (MS) detection (GC/MS). The analytical methods are a modification of EPA Test Method 625, Methods for Organic Chemical Analysis of Municipal and Industrial Wastewater, and Methods 8270, Test Methods for Evaluation of Solid Waste. PAH1 utilizes PUF as an adsorbant; PAH2 utilizes XAD-2 as an adsorbant. Note that isotopic surrogates may be added and are reported as % recovery. | 0.0001 | Y | Y | Y | 223 |  |  | 10/24/1996 2:07:37 PM | 1996-04-12 | 6/3/1998 6:51:37 PM | EMS |  | AirConcentration | µg/m³ | Y | (air conc.) |  | Indeno(1;2;3-cd)pyrene (air conc.) | 193-39-5 | Polycyclic Aromatic Hydrocarbons (PAHs); | Polycyclic Aromatic Hydrocarbons (PAHs); EMS code: PA13 | Chemical |  | LAB
PA13 | mg/L | Indeno(1;2;3-cd)pyrene | HEGC | Hexane extraction: GC/MS | Hexane extraction: GC/MS | 1e-05 | Y | Y | Y | 898 | 6/8/2015 | 8/15/2024 | 8/16/2024 1:36:03 PM | 2015-07-07 | 8/15/2019 11:27:47 AM | RDGIBSON | KATFRASE | Concentration | mg/L | Y | (fl. conc.) |  | Indeno(1;2;3-cd)pyrene (fl. conc.) | 193-39-5 | Polycyclic Aromatic Hydrocarbons (PAHs); | Polycyclic Aromatic Hydrocarbons (PAHs); EMS code: PA13 | Chemical |  | LAB
PA14 | ug/g | Naphthalene (C10H8) | MI01 | Microwave Extractn /GC-MS(PAH) | Soil samples are ground with anhydrous sodium sulphate and microwave extracted with dichloromethane. The extract is then cleaned up by silica gel column chromatography and analyzed by gc-ms in selective ion mode. | 0.002 | Y | Y | Y | 51 | 6/1/2000 11:42:00 AM | 9/19/2001 11:39:00 AM | 3/17/2003 1:22:54 PM | 1996-10-13 | 6/3/1998 6:51:46 PM | EMS |  | MassPerMass | mg/kg | Y | (mass conc.) |  | Naphthalene (C10H8) (mass conc.) | 91-20-3 | Polycyclic Aromatic Hydrocarbons (PAHs); | Polycyclic Aromatic Hydrocarbons (PAHs); EMS code: PA14 | Chemical |  | LAB
PA14 | ug/g wet | Naphthalene (C10H8) | X221 | Mod.Proc21.001(AOAC,'84) | Mod.Proc21.001(AOAC,'84) | 0.004 | Y | Y | Y | 74 |  |  | 10/24/1996 10:01:27 AM | 1996-04-12 | 6/3/1998 6:51:42 PM | EMS |  | MassPerMass | µg/g | Y | (mass conc.) | (wet) | Naphthalene (C10H8) (wet) (mass conc.) | 91-20-3 | Polycyclic Aromatic Hydrocarbons (PAHs); | Polycyclic Aromatic Hydrocarbons (PAHs); EMS code: PA14 | Chemical |  | LAB
PA14 | mg/L | Naphthalene (C10H8) | X384 | Purge and trap GC/MS | The sample is analyzed by purge and trap gas chromatography with detection by mass spectrometry. The procedure involves purging the volatile compounds from the sample with an inert gas, and trapping them on a solid sorbent. When purging of the sample is complete, the trap is heated and the compounds are transferred to the gas chromatographic column. Analysis is then accomplished by separation of the components by gas chromatography with detection by mass spectrometry. | 0.01 | Y | Y | Y | 2 | 8/5/2004 10:31:00 AM | 8/5/2004 10:31:00 AM | 11/28/2005 10:12:20 AM | 1996-04-12 | 6/3/1998 6:51:45 PM | EMS |  | Concentration | mg/L | Y | (fl. conc.) |  | Naphthalene (C10H8) (fl. conc.) | 91-20-3 | Polycyclic Aromatic Hydrocarbons (PAHs); | Polycyclic Aromatic Hydrocarbons (PAHs); EMS code: PA14 | Chemical |  | LAB
PA14 | mg/L | Naphthalene (C10H8) | MW02 | DCM extrn,GC/MS-SIM (PAH) | Water samples are liquid - liquid extracted with dichloromethane. The extract is then cleaned by silica gel column chromatography and analyzed by gas chromatography - mass spectroscopy. | 1e-05 | Y | Y | Y | 1106 | 2/5/2002 | 7/10/2024 5:51:00 PM | 8/21/2024 11:42:55 AM | 1996-04-12 | 6/3/1998 6:51:42 PM | EMS |  | Concentration | mg/L | Y | (fl. conc.) |  | Naphthalene (C10H8) (fl. conc.) | 91-20-3 | Polycyclic Aromatic Hydrocarbons (PAHs); | Polycyclic Aromatic Hydrocarbons (PAHs); EMS code: PA14 | Chemical |  | LAB
PA14 | ug/g | Naphthalene (C10H8) | X214 | Mod.Proc21.001(AOAC,'84) | Mod.Proc21.001(AOAC,'84) | 0.005 | Y | Y | Y | 160 |  |  | 10/24/1996 2:26:15 PM | 1996-04-12 | 6/3/1998 6:51:45 PM | EMS |  | MassPerMass | mg/kg | Y | (mass conc.) |  | Naphthalene (C10H8) (mass conc.) | 91-20-3 | Polycyclic Aromatic Hydrocarbons (PAHs); | Polycyclic Aromatic Hydrocarbons (PAHs); EMS code: PA14 | Chemical |  | LAB
PA14 | mg/L | Naphthalene (C10H8) | F010 | Extract disks,DCM extr:HRGC/LRMS | Water samples are spiked with a deuterated PAH Recovery Standards mix just prior to extraction with Empore extraction disks. Disks are eluted with Dichloromethane (DCM) and the extract concentrated. Deuterated PAH Internal Standards are added and the extract analyzed using High Resolution Gas Chromatography / Low Resolution Mass Spectrometry (HRGC/LRMS). MS data is acquired in Total Ion Mode (TIM) and quantitation is by multilevel Internal Standards. Recovery efficiency is calculated and reported but results are not recovery corrected. | 0.0002 | Y | Y | Y | 86 | 3/27/1997 4:24:00 PM | 3/11/2003 | 4/1/2004 1:13:26 AM | 1996-09-16 | 6/3/1998 6:51:46 PM | MPAK |  | Concentration | mg/L | Y | (fl. conc.) |  | Naphthalene (C10H8) (fl. conc.) | 91-20-3 | Polycyclic Aromatic Hydrocarbons (PAHs); | Polycyclic Aromatic Hydrocarbons (PAHs); EMS code: PA14 | Chemical |  | LAB
PA14 | mg/L | Naphthalene (C10H8) | MS01 | Purge&Trap GC/MS (VOC-W) | Purge&Trap GC/MS (VOC-W) | 0.005 | Y | Y | Y | 19 |  |  | 7/23/2024 11:41:50 AM | 2003-06-17 |  | UMULLER |  | Concentration | mg/L | Y | (fl. conc.) |  | Naphthalene (C10H8) (fl. conc.) | 91-20-3 | Polycyclic Aromatic Hydrocarbons (PAHs); | Polycyclic Aromatic Hydrocarbons (PAHs); EMS code: PA14 | Chemical |  | LAB
PA14 | ug/g | Naphthalene (C10H8) | MP02 | GC/MS | GC/MS | 0.001 | Y | Y | Y | 523 | 3/28/2002 | 9/23/2021 3:24:00 AM | 10/18/2021 9:46:09 AM | 1996-04-12 | 6/3/1998 6:51:37 PM | EMS |  | MassPerMass | mg/kg | Y | (mass conc.) |  | Naphthalene (C10H8) (mass conc.) | 91-20-3 | Polycyclic Aromatic Hydrocarbons (PAHs); | Polycyclic Aromatic Hydrocarbons (PAHs); EMS code: PA14 | Chemical |  | LAB
PA14 | ug/g | Naphthalene (C10H8) | FS87 | DCM,Ac/Hex extr,PAM,HRGC/LRMS | DCM;Ac/Hex extr;PAM;HRGC/LRMS | 0.02 | Y | Y | Y | 250 | 4/12/2000 2:52:00 PM | 10/29/2022 8:18:00 AM | 1/11/2024 1:07:08 PM | 2001-02-23 |  | NPEPPIN |  | MassPerMass | mg/kg | Y | (mass conc.) |  | Naphthalene (C10H8) (mass conc.) | 91-20-3 | Polycyclic Aromatic Hydrocarbons (PAHs); | Polycyclic Aromatic Hydrocarbons (PAHs); EMS code: PA14 | Chemical |  | LAB
PA14 | ug/m3 | Naphthalene (C10H8) | PAH1 | PUF:GC/MS | To understand the extent of human exposure to B[a]P and other PAHs, a reliable sampling and analytical method has been established which involves a combination quartz fibre filter/absorbent cartridge with subsequent extraction and analysis by gas chromatography (GC) with mass spectrometry (MS) detection (GC/MS). The analytical methods are a modification of EPA Test Method 625, Methods for Organic Chemical Analysis of Municipal and Industrial Wastewater, and Methods 8270, Test Methods for Evaluation of Solid Waste. PAH1 utilizes PUF as an adsorbant; PAH2 utilizes XAD-2 as an adsorbant. Note that isotopic surrogates may be added and are reported as % recovery. | 5e-05 | Y | Y | Y | 372 | 10/14/1997 2:50:00 PM | 4/22/1998 5:40:00 PM | 1/19/1999 6:45:11 PM | 1996-04-12 | 6/3/1998 6:51:37 PM | EMS |  | AirConcentration | µg/m³ | Y | (air conc.) |  | Naphthalene (C10H8) (air conc.) | 91-20-3 | Polycyclic Aromatic Hydrocarbons (PAHs); | Polycyclic Aromatic Hydrocarbons (PAHs); EMS code: PA14 | Chemical |  | LAB
PA14 | mg/L | Naphthalene (C10H8) | X262 | CH2Cl2 Ext:Silica:GC/MS | The sample is placed in a separatory funnel, the pH of the sample is adjusted to alkaline and then acidic conditions, and extracted each time with dichloromethane to selectively remove compounds. The final extracts (basic and acidic) can be combined or run separately. The extract is cleaned by silica gel column chromatography, then analyzed by GC/MS. | 0.001 | Y | Y | Y | 126 | 6/12/2000 | 1/13/2015 | 6/7/2024 5:30:18 PM | 1996-04-12 | 6/3/1998 6:51:37 PM | EMS |  | Concentration | mg/L | Y | (fl. conc.) |  | Naphthalene (C10H8) (fl. conc.) | 91-20-3 | Polycyclic Aromatic Hydrocarbons (PAHs); | Polycyclic Aromatic Hydrocarbons (PAHs); EMS code: PA14 | Chemical |  | LAB
PA14 | mg/L | Naphthalene (C10H8) | F007 | Static Headspace:HRGC/LRMS | Static headspace analysis is a method for the determination of volatile compounds in liquids and/or solids by analysing the vapour phase of a sample in a closed system; in this case a sealed vial. Static Headspace is sometimes referred to as Equilibrium Headspace because the vapour phase above the sample in a sealed vial is brought into thermodynamic equilibrium with the sample prior to analysis. Volatile organic compounds are thermally extracted from the sample matrix by Static Headspace sampling and introduced directly to a thick-film megabore GC column. Compounds eluting from the GC column are introduced into a Mass Spectrometer and target compounds are identified by a comparison of both retention times and mass spectra with those of reference standards stored in a computerized data system. Reference retention times and mass spectra for analytes are obtained by the measurement of calibration standards under precisely the same extraction conditions used for samples. The concentration of each identified component is measured by relating the response of the sample quantitation mass ions to the response of the quantitation mass ions of a series of external standards. A deuterated surrogate recovery standard is added to provide information on matrix effects. | 0.01 | Y | Y | Y | 1 |  |  | 8/11/2021 12:04:05 PM | 2010-11-16 |  | RDGIBSON |  | Concentration | mg/L | Y | (fl. conc.) |  | Naphthalene (C10H8) (fl. conc.) | 91-20-3 | Polycyclic Aromatic Hydrocarbons (PAHs); | Polycyclic Aromatic Hydrocarbons (PAHs); EMS code: PA14 | Chemical |  | LAB
PA14 | mg/L | Naphthalene (C10H8) | HEGC | Hexane extraction: GC/MS | Hexane extraction: GC/MS | 5e-05 | Y | Y | Y | 905 | 6/8/2015 | 8/15/2024 | 8/16/2024 1:36:03 PM | 2015-07-07 | 8/15/2019 11:28:08 AM | RDGIBSON | KATFRASE | Concentration | mg/L | Y | (fl. conc.) |  | Naphthalene (C10H8) (fl. conc.) | 91-20-3 | Polycyclic Aromatic Hydrocarbons (PAHs); | Polycyclic Aromatic Hydrocarbons (PAHs); EMS code: PA14 | Chemical |  | LAB
PA15 | ug/g | Phenanthrene | FS87 | DCM,Ac/Hex extr,PAM,HRGC/LRMS | DCM;Ac/Hex extr;PAM;HRGC/LRMS | 0.02 | Y | Y | Y | 250 | 4/12/2000 2:52:00 PM | 10/29/2022 8:18:00 AM | 1/11/2024 1:07:08 PM | 2001-02-23 |  | NPEPPIN |  | MassPerMass | mg/kg | Y | (mass conc.) |  | Phenanthrene (mass conc.) | 85-01-8 | Polycyclic Aromatic Hydrocarbons (PAHs); | Polycyclic Aromatic Hydrocarbons (PAHs); EMS code: PA15 | Chemical |  | LAB
PA15 | ug/m3 | Phenanthrene | PAH1 | PUF:GC/MS | To understand the extent of human exposure to B[a]P and other PAHs, a reliable sampling and analytical method has been established which involves a combination quartz fibre filter/absorbent cartridge with subsequent extraction and analysis by gas chromatography (GC) with mass spectrometry (MS) detection (GC/MS). The analytical methods are a modification of EPA Test Method 625, Methods for Organic Chemical Analysis of Municipal and Industrial Wastewater, and Methods 8270, Test Methods for Evaluation of Solid Waste. PAH1 utilizes PUF as an adsorbant; PAH2 utilizes XAD-2 as an adsorbant. Note that isotopic surrogates may be added and are reported as % recovery. | 5e-05 | Y | Y | Y | 372 | 10/14/1997 2:50:00 PM | 4/22/1998 5:40:00 PM | 1/19/1999 6:45:11 PM | 1996-04-12 | 6/3/1998 6:51:37 PM | EMS |  | AirConcentration | µg/m³ | Y | (air conc.) |  | Phenanthrene (air conc.) | 85-01-8 | Polycyclic Aromatic Hydrocarbons (PAHs); | Polycyclic Aromatic Hydrocarbons (PAHs); EMS code: PA15 | Chemical |  | LAB
PA15 | ug/g | Phenanthrene | MI01 | Microwave Extractn /GC-MS(PAH) | Soil samples are ground with anhydrous sodium sulphate and microwave extracted with dichloromethane. The extract is then cleaned up by silica gel column chromatography and analyzed by gc-ms in selective ion mode. | 0.006 | Y | Y | Y | 51 | 6/1/2000 11:42:00 AM | 9/19/2001 11:39:00 AM | 3/17/2003 1:22:54 PM | 1996-10-13 | 6/3/1998 6:51:46 PM | EMS |  | MassPerMass | mg/kg | Y | (mass conc.) |  | Phenanthrene (mass conc.) | 85-01-8 | Polycyclic Aromatic Hydrocarbons (PAHs); | Polycyclic Aromatic Hydrocarbons (PAHs); EMS code: PA15 | Chemical |  | LAB
PA15 | ug/g | Phenanthrene | MP02 | GC/MS | GC/MS | 0.001 | Y | Y | Y | 523 | 3/28/2002 | 9/23/2021 3:24:00 AM | 10/18/2021 9:46:09 AM | 1996-04-12 | 6/3/1998 6:51:37 PM | EMS |  | MassPerMass | mg/kg | Y | (mass conc.) |  | Phenanthrene (mass conc.) | 85-01-8 | Polycyclic Aromatic Hydrocarbons (PAHs); | Polycyclic Aromatic Hydrocarbons (PAHs); EMS code: PA15 | Chemical |  | LAB
PA15 | mg/L | Phenanthrene | HEGC | Hexane extraction: GC/MS | Hexane extraction: GC/MS | 2e-05 | Y | Y | Y | 905 | 6/8/2015 | 8/15/2024 | 8/16/2024 1:36:03 PM | 2015-07-07 | 8/15/2019 11:28:32 AM | RDGIBSON | KATFRASE | Concentration | mg/L | Y | (fl. conc.) |  | Phenanthrene (fl. conc.) | 85-01-8 | Polycyclic Aromatic Hydrocarbons (PAHs); | Polycyclic Aromatic Hydrocarbons (PAHs); EMS code: PA15 | Chemical |  | LAB
PA15 | mg/L | Phenanthrene | MW02 | DCM extrn,GC/MS-SIM (PAH) | Water samples are liquid - liquid extracted with dichloromethane. The extract is then cleaned by silica gel column chromatography and analyzed by gas chromatography - mass spectroscopy. | 1e-05 | Y | Y | Y | 1130 | 2/5/2002 | 7/10/2024 5:51:00 PM | 8/21/2024 11:42:55 AM | 1996-04-12 | 6/3/1998 6:51:42 PM | EMS |  | Concentration | mg/L | Y | (fl. conc.) |  | Phenanthrene (fl. conc.) | 85-01-8 | Polycyclic Aromatic Hydrocarbons (PAHs); | Polycyclic Aromatic Hydrocarbons (PAHs); EMS code: PA15 | Chemical |  | LAB
PA15 | mg/L | Phenanthrene | X262 | CH2Cl2 Ext:Silica:GC/MS | The sample is placed in a separatory funnel, the pH of the sample is adjusted to alkaline and then acidic conditions, and extracted each time with dichloromethane to selectively remove compounds. The final extracts (basic and acidic) can be combined or run separately. The extract is cleaned by silica gel column chromatography, then analyzed by GC/MS. | 0.001 | Y | Y | Y | 126 | 6/12/2000 | 1/13/2015 | 6/7/2024 5:30:18 PM | 1996-04-12 | 6/3/1998 6:51:45 PM | EMS |  | Concentration | mg/L | Y | (fl. conc.) |  | Phenanthrene (fl. conc.) | 85-01-8 | Polycyclic Aromatic Hydrocarbons (PAHs); | Polycyclic Aromatic Hydrocarbons (PAHs); EMS code: PA15 | Chemical |  | LAB
PA15 | mg/L | Phenanthrene | F010 | Extract disks,DCM extr:HRGC/LRMS | Water samples are spiked with a deuterated PAH Recovery Standards mix just prior to extraction with Empore extraction disks. Disks are eluted with Dichloromethane (DCM) and the extract concentrated. Deuterated PAH Internal Standards are added and the extract analyzed using High Resolution Gas Chromatography / Low Resolution Mass Spectrometry (HRGC/LRMS). MS data is acquired in Total Ion Mode (TIM) and quantitation is by multilevel Internal Standards. Recovery efficiency is calculated and reported but results are not recovery corrected. | 0.0002 | Y | Y | Y | 86 | 3/27/1997 4:24:00 PM | 3/11/2003 | 4/1/2004 1:13:26 AM | 1996-09-16 | 6/3/1998 6:51:44 PM | MPAK |  | Concentration | mg/L | Y | (fl. conc.) |  | Phenanthrene (fl. conc.) | 85-01-8 | Polycyclic Aromatic Hydrocarbons (PAHs); | Polycyclic Aromatic Hydrocarbons (PAHs); EMS code: PA15 | Chemical |  | LAB
PA15 | ug/g | Phenanthrene | X214 | Mod.Proc21.001(AOAC,'84) | Mod.Proc21.001(AOAC,'84) | 0.005 | Y | Y | Y | 160 |  |  | 10/24/1996 2:26:15 PM | 1996-04-12 | 6/3/1998 6:51:37 PM | EMS |  | MassPerMass | mg/kg | Y | (mass conc.) |  | Phenanthrene (mass conc.) | 85-01-8 | Polycyclic Aromatic Hydrocarbons (PAHs); | Polycyclic Aromatic Hydrocarbons (PAHs); EMS code: PA15 | Chemical |  | LAB
PA15 | ug/m3 | Phenanthrene | PAH2 | XAD2: GC/MS | To understand the extent of human exposure to B[a]P, and other PAHs, a reliable sampling and analytical method has been established which involves a combination quartz fibre filter/adsorbent cartridge with subsequent extraction and analysis by gas chromatography (GC) with mass spectrometry (MS) detection (GC/MS). The analytical methods are a modification of EPA Test Method 625, Methods for Organic Chemical Analysis of Municipal and Industrial Wastewater, and Methods 8270, Test Methods for Evaluation of Solid Waste. PAH1 utilizes PUF as an adsorbant; PAH2 utilizes XAD-2 as an adsorbant. Note that isotopic surrogates may be added and are reported as % recovery. | 5e-05 | Y | Y | Y | 223 |  |  | 10/24/1996 2:07:37 PM | 1996-04-12 | 6/3/1998 6:51:37 PM | EMS |  | AirConcentration | µg/m³ | Y | (air conc.) |  | Phenanthrene (air conc.) | 85-01-8 | Polycyclic Aromatic Hydrocarbons (PAHs); | Polycyclic Aromatic Hydrocarbons (PAHs); EMS code: PA15 | Chemical |  | LAB
PA15 | ug/g wet | Phenanthrene | X221 | Mod.Proc21.001(AOAC,'84) | Mod.Proc21.001(AOAC,'84) | 0.004 | Y | Y | Y | 74 |  |  | 10/24/1996 10:01:27 AM | 1996-04-12 | 6/3/1998 6:51:42 PM | EMS |  | MassPerMass | µg/g | Y | (mass conc.) | (wet) | Phenanthrene (wet) (mass conc.) | 85-01-8 | Polycyclic Aromatic Hydrocarbons (PAHs); | Polycyclic Aromatic Hydrocarbons (PAHs); EMS code: PA15 | Chemical |  | LAB
PA16 | ug/g | Pyrene | MP02 | GC/MS | GC/MS | 0.001 | Y | Y | Y | 523 | 3/28/2002 | 9/23/2021 3:24:00 AM | 10/18/2021 9:46:09 AM | 1996-04-12 | 6/3/1998 6:51:37 PM | EMS |  | MassPerMass | mg/kg | Y | (mass conc.) |  | Pyrene (mass conc.) | 129-00-0 | Polycyclic Aromatic Hydrocarbons (PAHs); | Polycyclic Aromatic Hydrocarbons (PAHs); EMS code: PA16 | Chemical |  | LAB
PA16 | mg/L | Pyrene | X262 | CH2Cl2 Ext:Silica:GC/MS | The sample is placed in a separatory funnel, the pH of the sample is adjusted to alkaline and then acidic conditions, and extracted each time with dichloromethane to selectively remove compounds. The final extracts (basic and acidic) can be combined or run separately. The extract is cleaned by silica gel column chromatography, then analyzed by GC/MS. | 0.001 | Y | Y | Y | 126 | 6/12/2000 | 1/13/2015 | 6/7/2024 5:30:18 PM | 1996-04-12 | 6/3/1998 6:51:37 PM | EMS |  | Concentration | mg/L | Y | (fl. conc.) |  | Pyrene (fl. conc.) | 129-00-0 | Polycyclic Aromatic Hydrocarbons (PAHs); | Polycyclic Aromatic Hydrocarbons (PAHs); EMS code: PA16 | Chemical |  | LAB
PA16 | ug/g | Pyrene | MI01 | Microwave Extractn /GC-MS(PAH) | Soil samples are ground with anhydrous sodium sulphate and microwave extracted with dichloromethane. The extract is then cleaned up by silica gel column chromatography and analyzed by gc-ms in selective ion mode. | 0.001 | Y | Y | Y | 51 | 6/1/2000 11:42:00 AM | 9/19/2001 11:39:00 AM | 3/17/2003 1:22:54 PM | 1996-10-13 | 6/3/1998 6:51:46 PM | EMS |  | MassPerMass | mg/kg | Y | (mass conc.) |  | Pyrene (mass conc.) | 129-00-0 | Polycyclic Aromatic Hydrocarbons (PAHs); | Polycyclic Aromatic Hydrocarbons (PAHs); EMS code: PA16 | Chemical |  | LAB
PA16 | mg/L | Pyrene | F010 | Extract disks,DCM extr:HRGC/LRMS | Water samples are spiked with a deuterated PAH Recovery Standards mix just prior to extraction with Empore extraction disks. Disks are eluted with Dichloromethane (DCM) and the extract concentrated. Deuterated PAH Internal Standards are added and the extract analyzed using High Resolution Gas Chromatography / Low Resolution Mass Spectrometry (HRGC/LRMS). MS data is acquired in Total Ion Mode (TIM) and quantitation is by multilevel Internal Standards. Recovery efficiency is calculated and reported but results are not recovery corrected. | 0.0002 | Y | Y | Y | 86 | 3/27/1997 4:24:00 PM | 3/11/2003 | 4/1/2004 1:13:27 AM | 1996-09-16 | 6/3/1998 6:51:44 PM | MPAK |  | Concentration | mg/L | Y | (fl. conc.) |  | Pyrene (fl. conc.) | 129-00-0 | Polycyclic Aromatic Hydrocarbons (PAHs); | Polycyclic Aromatic Hydrocarbons (PAHs); EMS code: PA16 | Chemical |  | LAB
PA16 | mg/L | Pyrene | MW02 | DCM extrn,GC/MS-SIM (PAH) | Water samples are liquid - liquid extracted with dichloromethane. The extract is then cleaned by silica gel column chromatography and analyzed by gas chromatography - mass spectroscopy. | 1e-05 | Y | Y | Y | 1127 | 2/5/2002 | 7/10/2024 5:51:00 PM | 8/21/2024 11:42:55 AM | 1996-04-12 | 6/3/1998 6:51:42 PM | EMS |  | Concentration | mg/L | Y | (fl. conc.) |  | Pyrene (fl. conc.) | 129-00-0 | Polycyclic Aromatic Hydrocarbons (PAHs); | Polycyclic Aromatic Hydrocarbons (PAHs); EMS code: PA16 | Chemical |  | LAB
PA16 | ug/g wet | Pyrene | X221 | Mod.Proc21.001(AOAC,'84) | Mod.Proc21.001(AOAC,'84) | 0.01 | Y | Y | Y | 74 |  |  | 10/24/1996 10:01:27 AM | 1996-04-12 | 6/3/1998 6:51:42 PM | EMS |  | MassPerMass | µg/g | Y | (mass conc.) | (wet) | Pyrene (wet) (mass conc.) | 129-00-0 | Polycyclic Aromatic Hydrocarbons (PAHs); | Polycyclic Aromatic Hydrocarbons (PAHs); EMS code: PA16 | Chemical |  | LAB
PA16 | ug/m3 | Pyrene | PAH1 | PUF:GC/MS | To understand the extent of human exposure to B[a]P and other PAHs, a reliable sampling and analytical method has been established which involves a combination quartz fibre filter/absorbent cartridge with subsequent extraction and analysis by gas chromatography (GC) with mass spectrometry (MS) detection (GC/MS). The analytical methods are a modification of EPA Test Method 625, Methods for Organic Chemical Analysis of Municipal and Industrial Wastewater, and Methods 8270, Test Methods for Evaluation of Solid Waste. PAH1 utilizes PUF as an adsorbant; PAH2 utilizes XAD-2 as an adsorbant. Note that isotopic surrogates may be added and are reported as % recovery. | 5e-05 | Y | Y | Y | 372 | 10/14/1997 2:50:00 PM | 4/22/1998 5:40:00 PM | 1/19/1999 6:45:11 PM | 1996-04-12 | 6/3/1998 6:51:37 PM | EMS |  | AirConcentration | µg/m³ | Y | (air conc.) |  | Pyrene (air conc.) | 129-00-0 | Polycyclic Aromatic Hydrocarbons (PAHs); | Polycyclic Aromatic Hydrocarbons (PAHs); EMS code: PA16 | Chemical |  | LAB
PA16 | mg/L | Pyrene | HEGC | Hexane extraction: GC/MS | Hexane extraction: GC/MS | 1e-05 | Y | Y | Y | 905 | 6/8/2015 | 8/15/2024 | 8/16/2024 1:36:03 PM | 2015-07-07 | 8/15/2019 11:28:57 AM | RDGIBSON | KATFRASE | Concentration | mg/L | Y | (fl. conc.) |  | Pyrene (fl. conc.) | 129-00-0 | Polycyclic Aromatic Hydrocarbons (PAHs); | Polycyclic Aromatic Hydrocarbons (PAHs); EMS code: PA16 | Chemical |  | LAB
PA16 | ug/m3 | Pyrene | PAH2 | XAD2: GC/MS | To understand the extent of human exposure to B[a]P, and other PAHs, a reliable sampling and analytical method has been established which involves a combination quartz fibre filter/adsorbent cartridge with subsequent extraction and analysis by gas chromatography (GC) with mass spectrometry (MS) detection (GC/MS). The analytical methods are a modification of EPA Test Method 625, Methods for Organic Chemical Analysis of Municipal and Industrial Wastewater, and Methods 8270, Test Methods for Evaluation of Solid Waste. PAH1 utilizes PUF as an adsorbant; PAH2 utilizes XAD-2 as an adsorbant. Note that isotopic surrogates may be added and are reported as % recovery. | 5e-05 | Y | Y | Y | 223 |  |  | 10/24/1996 2:07:37 PM | 1996-04-12 | 6/3/1998 6:51:42 PM | EMS |  | AirConcentration | µg/m³ | Y | (air conc.) |  | Pyrene (air conc.) | 129-00-0 | Polycyclic Aromatic Hydrocarbons (PAHs); | Polycyclic Aromatic Hydrocarbons (PAHs); EMS code: PA16 | Chemical |  | LAB
PA16 | ug/g | Pyrene | X214 | Mod.Proc21.001(AOAC,'84) | Mod.Proc21.001(AOAC,'84) | 0.01 | Y | Y | Y | 159 |  |  | 10/24/1996 2:26:15 PM | 1996-04-12 | 6/3/1998 6:51:37 PM | EMS |  | MassPerMass | mg/kg | Y | (mass conc.) |  | Pyrene (mass conc.) | 129-00-0 | Polycyclic Aromatic Hydrocarbons (PAHs); | Polycyclic Aromatic Hydrocarbons (PAHs); EMS code: PA16 | Chemical |  | LAB
PA16 | ug/g | Pyrene | FS87 | DCM,Ac/Hex extr,PAM,HRGC/LRMS | DCM;Ac/Hex extr;PAM;HRGC/LRMS | 0.02 | Y | Y | Y | 250 | 4/12/2000 2:52:00 PM | 10/29/2022 8:18:00 AM | 1/11/2024 1:07:08 PM | 2001-02-23 |  | NPEPPIN |  | MassPerMass | mg/kg | Y | (mass conc.) |  | Pyrene (mass conc.) | 129-00-0 | Polycyclic Aromatic Hydrocarbons (PAHs); | Polycyclic Aromatic Hydrocarbons (PAHs); EMS code: PA16 | Chemical |  | LAB
PA17 | ug/g | Benzo(b+k)fluoranthene | MW02 | DCM extrn,GC/MS-SIM (PAH) | Water samples are liquid - liquid extracted with dichloromethane. The extract is then cleaned by silica gel column chromatography and analyzed by gas chromatography - mass spectroscopy. | 1e-05 | Y | Y | Y | 23 |  |  | 10/24/1996 2:13:46 PM | 1996-04-12 | 6/3/1998 6:51:42 PM | EMS |  | MassPerMass | mg/kg | Y | (mass conc.) |  | Benzo(b+k)fluoranthene (mass conc.) | 205-99-2, 207-08-9 | Polycyclic Aromatic Hydrocarbons (PAHs); | Polycyclic Aromatic Hydrocarbons (PAHs); EMS code: PA17 | Chemical |  | LAB
PA17 | ug/g | Benzo(b+k)fluoranthene | MP02 | GC/MS | GC/MS | 0.001 | Y | Y | Y | 44 | 3/31/2006 | 4/4/2006 | 4/27/2006 1:02:37 AM | 1996-04-12 | 6/3/1998 6:51:37 PM | EMS |  | MassPerMass | mg/kg | Y | (mass conc.) |  | Benzo(b+k)fluoranthene (mass conc.) | 205-99-2, 207-08-9 | Polycyclic Aromatic Hydrocarbons (PAHs); | Polycyclic Aromatic Hydrocarbons (PAHs); EMS code: PA17 | Chemical |  | LAB
PA17 | ug/m3 | Benzo(b+k)fluoranthene | PAH1 | PUF:GC/MS | To understand the extent of human exposure to B[a]P and other PAHs, a reliable sampling and analytical method has been established which involves a combination quartz fibre filter/absorbent cartridge with subsequent extraction and analysis by gas chromatography (GC) with mass spectrometry (MS) detection (GC/MS). The analytical methods are a modification of EPA Test Method 625, Methods for Organic Chemical Analysis of Municipal and Industrial Wastewater, and Methods 8270, Test Methods for Evaluation of Solid Waste. PAH1 utilizes PUF as an adsorbant; PAH2 utilizes XAD-2 as an adsorbant. Note that isotopic surrogates may be added and are reported as % recovery. | 0.0001 | Y | Y | Y | 272 |  |  | 10/24/1996 1:01:23 PM | 1996-04-12 | 6/3/1998 6:51:37 PM | EMS |  | AirConcentration | µg/m³ | Y | (air conc.) |  | Benzo(b+k)fluoranthene (air conc.) | 205-99-2, 207-08-9 | Polycyclic Aromatic Hydrocarbons (PAHs); | Polycyclic Aromatic Hydrocarbons (PAHs); EMS code: PA17 | Chemical |  | LAB
PA17 | ug/m3 | Benzo(b+k)fluoranthene | PAH2 | XAD2: GC/MS | To understand the extent of human exposure to B[a]P, and other PAHs, a reliable sampling and analytical method has been established which involves a combination quartz fibre filter/adsorbent cartridge with subsequent extraction and analysis by gas chromatography (GC) with mass spectrometry (MS) detection (GC/MS). The analytical methods are a modification of EPA Test Method 625, Methods for Organic Chemical Analysis of Municipal and Industrial Wastewater, and Methods 8270, Test Methods for Evaluation of Solid Waste. PAH1 utilizes PUF as an adsorbant; PAH2 utilizes XAD-2 as an adsorbant. Note that isotopic surrogates may be added and are reported as % recovery. | 0.0001 | Y | Y | Y | 219 |  |  | 10/24/1996 1:48:35 PM | 1996-04-12 | 6/3/1998 6:51:37 PM | EMS |  | AirConcentration | µg/m³ | Y | (air conc.) |  | Benzo(b+k)fluoranthene (air conc.) | 205-99-2, 207-08-9 | Polycyclic Aromatic Hydrocarbons (PAHs); | Polycyclic Aromatic Hydrocarbons (PAHs); EMS code: PA17 | Chemical |  | LAB
PA18 | mg/L | Acridine | MW02 | DCM extrn,GC/MS-SIM (PAH) | Water samples are liquid - liquid extracted with dichloromethane. The extract is then cleaned by silica gel column chromatography and analyzed by gas chromatography - mass spectroscopy. | 5e-05 | Y | Y | Y | 1025 | 6/12/2000 | 7/10/2024 5:51:00 PM | 8/21/2024 11:42:55 AM | 1998-09-28 |  | NPEPPIN |  | Concentration | mg/L | Y | (fl. conc.) |  | Acridine (fl. conc.) | 260-94-6 | Polycyclic Aromatic Hydrocarbons (PAHs); | Polycyclic Aromatic Hydrocarbons (PAHs); EMS code: PA18 | Chemical |  | LAB
PA18 | ug/g | Acridine | FS87 | DCM,Ac/Hex extr,PAM,HRGC/LRMS | DCM;Ac/Hex extr;PAM;HRGC/LRMS | 0.01 | Y | Y | Y | 19 | 1/10/2022 | 9/9/2022 | 1/11/2024 1:07:08 PM | 2022-07-07 |  | EMS |  | MassPerMass | mg/kg | Y | (mass conc.) |  | Acridine (mass conc.) | 260-94-6 | Polycyclic Aromatic Hydrocarbons (PAHs); | Polycyclic Aromatic Hydrocarbons (PAHs); EMS code: PA18 | Chemical |  | LAB
PA18 | mg/L | Acridine | HEGC | Hexane extraction: GC/MS | Hexane extraction: GC/MS | 1e-05 | Y | Y | Y | 904 | 6/8/2015 | 8/15/2024 | 8/16/2024 1:36:03 PM | 2015-07-07 | 8/15/2019 11:29:22 AM | RDGIBSON | KATFRASE | Concentration | mg/L | Y | (fl. conc.) |  | Acridine (fl. conc.) | 260-94-6 | Polycyclic Aromatic Hydrocarbons (PAHs); | Polycyclic Aromatic Hydrocarbons (PAHs); EMS code: PA18 | Chemical |  | LAB
PA18 | mg/L | Acridine | X262 | CH2Cl2 Ext:Silica:GC/MS | The sample is placed in a separatory funnel, the pH of the sample is adjusted to alkaline and then acidic conditions, and extracted each time with dichloromethane to selectively remove compounds. The final extracts (basic and acidic) can be combined or run separately. The extract is cleaned by silica gel column chromatography, then analyzed by GC/MS. | 0.001 | Y | Y | Y | 100 | 8/27/2002 | 1/13/2015 | 3/12/2018 5:01:13 PM | 2003-06-03 |  | UMULLER |  | Concentration | mg/L | Y | (fl. conc.) |  | Acridine (fl. conc.) | 260-94-6 | Polycyclic Aromatic Hydrocarbons (PAHs); | Polycyclic Aromatic Hydrocarbons (PAHs); EMS code: PA18 | Chemical |  | LAB
PA19 | mg/L | Quinoline aka Benzo(b)pyridine | X262 | CH2Cl2 Ext:Silica:GC/MS | The sample is placed in a separatory funnel, the pH of the sample is adjusted to alkaline and then acidic conditions, and extracted each time with dichloromethane to selectively remove compounds. The final extracts (basic and acidic) can be combined or run separately. The extract is cleaned by silica gel column chromatography, then analyzed by GC/MS. | 0.001 | Y | Y | Y | 342 | 8/27/2002 | 1/9/2018 2:30:00 PM | 6/7/2024 5:30:17 PM | 2003-03-17 | 3/17/2003 11:10:19 AM | UMULLER | UMULLER | Concentration | mg/L | Y | (fl. conc.) |  | Quinoline aka Benzo(b)pyridine (fl. conc.) | 91-22-5 | Polycyclic Aromatic Hydrocarbons (PAHs); | Polycyclic Aromatic Hydrocarbons (PAHs); EMS code: PA19 | Chemical |  | LAB
PA19 | ug/g | Quinoline aka Benzo(b)pyridine | FS87 | DCM,Ac/Hex extr,PAM,HRGC/LRMS | DCM;Ac/Hex extr;PAM;HRGC/LRMS | 0.05 | Y | Y | Y | 67 | 12/22/2017 | 10/29/2022 8:18:00 AM | 1/11/2024 1:07:08 PM | 2018-01-10 |  | JKROGH |  | MassPerMass | mg/kg | Y | (mass conc.) |  | Quinoline aka Benzo(b)pyridine (mass conc.) | 91-22-5 | Polycyclic Aromatic Hydrocarbons (PAHs); | Polycyclic Aromatic Hydrocarbons (PAHs); EMS code: PA19 | Chemical |  | LAB
PA19 | mg/L | Quinoline aka Benzo(b)pyridine | HEGC | Hexane extraction: GC/MS | Hexane extraction: GC/MS | 5e-05 | Y | Y | Y | 904 | 6/8/2015 | 8/15/2024 | 8/16/2024 1:36:03 PM | 2015-07-07 | 8/15/2019 11:29:47 AM | RDGIBSON | KATFRASE | Concentration | mg/L | Y | (fl. conc.) |  | Quinoline aka Benzo(b)pyridine (fl. conc.) | 91-22-5 | Polycyclic Aromatic Hydrocarbons (PAHs); | Polycyclic Aromatic Hydrocarbons (PAHs); EMS code: PA19 | Chemical |  | LAB
PA19 | mg/L | Quinoline aka Benzo(b)pyridine | MW02 | DCM extrn,GC/MS-SIM (PAH) | Water samples are liquid - liquid extracted with dichloromethane. The extract is then cleaned by silica gel column chromatography and analyzed by gas chromatography - mass spectroscopy. | 0.05 | Y | Y | Y | 568 | 9/7/2014 9:36:00 PM | 7/10/2024 5:51:00 PM | 8/21/2024 11:42:55 AM | 2010-11-16 | 11/16/2010 11:26:08 AM | RDGIBSON | RDGIBSON | Concentration | mg/L | Y | (fl. conc.) |  | Quinoline aka Benzo(b)pyridine (fl. conc.) | 91-22-5 | Polycyclic Aromatic Hydrocarbons (PAHs); | Polycyclic Aromatic Hydrocarbons (PAHs); EMS code: PA19 | Chemical |  | LAB
PA22 | mg/L | Benzo(j)fluoranthene | MW02 | DCM extrn,GC/MS-SIM (PAH) | Water samples are liquid - liquid extracted with dichloromethane. The extract is then cleaned by silica gel column chromatography and analyzed by gas chromatography - mass spectroscopy. | 1e-05 | Y | Y | Y | 25 | 11/20/2018 9:23:00 PM | 7/10/2024 5:51:00 PM | 8/21/2024 11:42:55 AM | 1998-09-28 |  | NPEPPIN |  | Concentration | mg/L | Y | (fl. conc.) |  | Benzo(j)fluoranthene (fl. conc.) | 205-82-3 | Polycyclic Aromatic Hydrocarbons (PAHs); | Polycyclic Aromatic Hydrocarbons (PAHs); EMS code: PA22 | Chemical |  | LAB
PA23 | ug/g | 7;12-Dimethylbenz(a) | MP02 | GC/MS | GC/MS | 0.001 | Y | Y | Y | 227 |  |  | 4/19/2021 1:37:50 PM | 1996-04-12 | 6/3/1998 6:51:37 PM | EMS |  | MassPerMass | mg/kg | Y | (mass conc.) |  | 7;12-Dimethylbenz(a) (mass conc.) | 57-97-6 | Polycyclic Aromatic Hydrocarbons (PAHs); | Polycyclic Aromatic Hydrocarbons (PAHs); EMS code: PA23 | Chemical |  | LAB
PA28 | mg/L | 2-Methylnaphthalene | MW02 | DCM extrn,GC/MS-SIM (PAH) | Water samples are liquid - liquid extracted with dichloromethane. The extract is then cleaned by silica gel column chromatography and analyzed by gas chromatography - mass spectroscopy. | 1e-05 | Y | Y | Y | 878 | 2/5/2002 | 7/10/2024 5:51:00 PM | 8/21/2024 11:42:55 AM | 1998-09-28 |  | NPEPPIN |  | Concentration | mg/L | Y | (fl. conc.) |  | 2-Methylnaphthalene (fl. conc.) | 91-57-6 | Polycyclic Aromatic Hydrocarbons (PAHs); | Polycyclic Aromatic Hydrocarbons (PAHs); EMS code: PA28 | Chemical |  | LAB
PA28 | ug/g | 2-Methylnaphthalene | MI01 | Microwave Extractn /GC-MS(PAH) | Soil samples are ground with anhydrous sodium sulphate and microwave extracted with dichloromethane. The extract is then cleaned up by silica gel column chromatography and analyzed by gc-ms in selective ion mode. | 0.001 | Y | Y | Y | 1 | 12/19/2003 8:50:00 AM | 12/19/2003 8:50:00 AM | 3/24/2004 1:01:00 AM | 1998-09-28 |  | NPEPPIN |  | MassPerMass | mg/kg | Y | (mass conc.) |  | 2-Methylnaphthalene (mass conc.) | 91-57-6 | Polycyclic Aromatic Hydrocarbons (PAHs); | Polycyclic Aromatic Hydrocarbons (PAHs); EMS code: PA28 | Chemical |  | LAB
PA28 | mg/L | 2-Methylnaphthalene | HEGC | Hexane extraction: GC/MS | Hexane extraction: GC/MS | 5e-05 | Y | Y | Y | 567 | 8/18/2015 | 8/15/2024 | 8/16/2024 1:36:03 PM | 2016-08-17 |  | RRENTON |  | Concentration | mg/L | Y | (fl. conc.) |  | 2-Methylnaphthalene (fl. conc.) | 91-57-6 | Polycyclic Aromatic Hydrocarbons (PAHs); | Polycyclic Aromatic Hydrocarbons (PAHs); EMS code: PA28 | Chemical |  | LAB
PA28 | ug/g | 2-Methylnaphthalene | MP02 | GC/MS | GC/MS | 0.001 | Y | Y | Y | 322 | 3/3/2010 | 9/23/2021 3:24:00 AM | 10/18/2021 9:46:09 AM | 2009-04-28 |  | JCHUDAK |  | MassPerMass | mg/kg | Y | (mass conc.) |  | 2-Methylnaphthalene (mass conc.) | 91-57-6 | Polycyclic Aromatic Hydrocarbons (PAHs); | Polycyclic Aromatic Hydrocarbons (PAHs); EMS code: PA28 | Chemical |  | LAB
PA28 | ug/g | 2-Methylnaphthalene | FS87 | DCM,Ac/Hex extr,PAM,HRGC/LRMS | DCM;Ac/Hex extr;PAM;HRGC/LRMS | 0.05 | Y | Y | Y | 126 | 5/20/2015 | 10/29/2022 8:18:00 AM | 1/11/2024 1:07:08 PM | 2015-05-27 |  | RDGIBSON |  | MassPerMass | mg/kg | Y | (mass conc.) |  | 2-Methylnaphthalene (mass conc.) | 91-57-6 | Polycyclic Aromatic Hydrocarbons (PAHs); | Polycyclic Aromatic Hydrocarbons (PAHs); EMS code: PA28 | Chemical |  | LAB
PA28 | ug/g | 2-Methylnaphthalene | FS67 | NaOH4,MeOH/Cl2H2Cextr,HRGC/LRMS | The method consisted of first spiking a portion of the undried sediment samples with an aliquot of three different surrogate standard solutions: one for the analysis of alkanes, aliphatics and linear alkylbenzenes in the F1 extract; one for the analysis of PAHs in the F2 extract; and one for alcohols and plant sterols in the F3. Each tissue sample was ground with sodium sulphate, packed in a glass chromatographic column and eluted with methanol and dichloromethane. The eluate was backwashed with extracted water. The extract was concentrated and fractionated on silica gel into three fractions (F1, F2 and F3). F1 and F3 extracts were concentrated and ampouled. The F2 extract containing the PAH compounds of interest was washed with base to remove fatty acid interferences, concentrated and spiked with an aliquot of recovery standard solution prior to analysis by high resolution gas chromatography with low resolution (quadrupole) mass spectrometric detection (HRGC/LRMS). Results reported on a dry weight basis, based on moisture analysis of a separate aliquot. | 0.000706 | Y | Y | Y | 29 | 3/31/2006 | 11/11/2008 | 11/20/2008 5:09:03 PM | 1999-07-12 |  | NPEPPIN |  | MassPerMass | mg/kg | Y | (mass conc.) |  | 2-Methylnaphthalene (mass conc.) | 91-57-6 | Polycyclic Aromatic Hydrocarbons (PAHs); | Polycyclic Aromatic Hydrocarbons (PAHs); EMS code: PA28 | Chemical |  | LAB
PA29 | mg/L | 1-Methylnaphthalene | MW02 | DCM extrn,GC/MS-SIM (PAH) | Water samples are liquid - liquid extracted with dichloromethane. The extract is then cleaned by silica gel column chromatography and analyzed by gas chromatography - mass spectroscopy. | 1e-05 | Y | Y | Y | 516 | 12/3/2002 | 7/2/2024 9:48:00 PM | 7/22/2024 3:17:33 PM | 1998-09-28 |  | NPEPPIN |  | Concentration | mg/L | Y | (fl. conc.) |  | 1-Methylnaphthalene (fl. conc.) | 90-12-0 | Polycyclic Aromatic Hydrocarbons (PAHs); | Polycyclic Aromatic Hydrocarbons (PAHs); EMS code: PA29 | Chemical |  | LAB
PA29 | ug/g | 1-Methylnaphthalene | FS87 | DCM,Ac/Hex extr,PAM,HRGC/LRMS | DCM;Ac/Hex extr;PAM;HRGC/LRMS | 0.05 | Y | Y | Y | 67 | 12/22/2017 | 10/29/2022 8:18:00 AM | 1/11/2024 1:07:08 PM | 2018-01-10 |  | JKROGH |  | MassPerMass | mg/kg | Y | (mass conc.) |  | 1-Methylnaphthalene (mass conc.) | 90-12-0 | Polycyclic Aromatic Hydrocarbons (PAHs); | Polycyclic Aromatic Hydrocarbons (PAHs); EMS code: PA29 | Chemical |  | LAB
PA29 | mg/L | 1-Methylnaphthalene | HEGC | Hexane extraction: GC/MS | Hexane extraction: GC/MS | 5e-05 | Y | Y | Y | 567 | 8/18/2015 | 8/15/2024 | 8/16/2024 1:36:03 PM | 2016-08-17 |  | RRENTON |  | Concentration | mg/L | Y | (fl. conc.) |  | 1-Methylnaphthalene (fl. conc.) | 90-12-0 | Polycyclic Aromatic Hydrocarbons (PAHs); | Polycyclic Aromatic Hydrocarbons (PAHs); EMS code: PA29 | Chemical |  | LAB
PA33 | mg/L | Benzo(e)pyrene | MW02 | DCM extrn,GC/MS-SIM (PAH) | Water samples are liquid - liquid extracted with dichloromethane. The extract is then cleaned by silica gel column chromatography and analyzed by gas chromatography - mass spectroscopy. | 1e-05 | Y | Y | Y | 29 | 11/20/2018 9:23:00 PM | 7/10/2024 5:51:00 PM | 8/21/2024 11:42:55 AM | 1998-09-28 |  | NPEPPIN |  | Concentration | mg/L | Y | (fl. conc.) |  | Benzo(e)pyrene (fl. conc.) | 192-97-2 | Polycyclic Aromatic Hydrocarbons (PAHs); | Polycyclic Aromatic Hydrocarbons (PAHs); EMS code: PA33 | Chemical |  | LAB
PA33 | mg/L | Benzo(e)pyrene | HEGC | Hexane extraction: GC/MS | Hexane extraction: GC/MS | 1e-05 | Y | Y | Y | 9 | 5/17/2024 | 7/28/2024 | 8/6/2024 3:27:04 PM | 2024-05-28 |  | RDGIBSON |  | Concentration | mg/L | Y | (fl. conc.) |  | Benzo(e)pyrene (fl. conc.) | 192-97-2 | Polycyclic Aromatic Hydrocarbons (PAHs); | Polycyclic Aromatic Hydrocarbons (PAHs); EMS code: PA33 | Chemical |  | LAB
PA34 | mg/L | Perylene | HEGC | Hexane extraction: GC/MS | Hexane extraction: GC/MS | 1e-05 | Y | Y | Y | 9 | 5/17/2024 | 7/28/2024 | 8/6/2024 3:27:04 PM | 2024-05-28 |  | RDGIBSON |  | Concentration | mg/L | Y | (fl. conc.) |  | Perylene (fl. conc.) | 198-55-0 | Polycyclic Aromatic Hydrocarbons (PAHs); | Polycyclic Aromatic Hydrocarbons (PAHs); EMS code: PA34 | Chemical |  | LAB
PA34 | mg/L | Perylene | MW02 | DCM extrn,GC/MS-SIM (PAH) | Water samples are liquid - liquid extracted with dichloromethane. The extract is then cleaned by silica gel column chromatography and analyzed by gas chromatography - mass spectroscopy. | 1e-05 | Y | Y | Y | 12 | 12/3/2002 | 3/11/2003 | 6/7/2024 5:30:18 PM | 1998-09-28 |  | NPEPPIN |  | Concentration | mg/L | Y | (fl. conc.) |  | Perylene (fl. conc.) | 198-55-0 | Polycyclic Aromatic Hydrocarbons (PAHs); | Polycyclic Aromatic Hydrocarbons (PAHs); EMS code: PA34 | Chemical |  | LAB
PA35 | ug/g | 2;6-Dimethylnapthalene | MI01 | Microwave Extractn /GC-MS(PAH) | Soil samples are ground with anhydrous sodium sulphate and microwave extracted with dichloromethane. The extract is then cleaned up by silica gel column chromatography and analyzed by gc-ms in selective ion mode. | 0.01 | Y | Y | Y | 7 | 10/29/2002 4:15:00 PM | 9/22/2003 11:15:00 AM | 3/25/2004 1:00:40 AM | 1998-09-28 |  | NPEPPIN |  | MassPerMass | mg/kg | Y | (mass conc.) |  | 2;6-Dimethylnapthalene (mass conc.) | 581-42-0 | Organics; | Organics; EMS code: PA35 | Chemical |  | LAB
PA36 | ug/g | 1;6;7-Trimethylnapthalene | MI01 | Microwave Extractn /GC-MS(PAH) | Soil samples are ground with anhydrous sodium sulphate and microwave extracted with dichloromethane. The extract is then cleaned up by silica gel column chromatography and analyzed by gc-ms in selective ion mode. | 0.01 | Y | Y | Y | 7 | 10/29/2002 4:15:00 PM | 9/22/2003 11:15:00 AM | 3/25/2004 1:00:40 AM | 1998-09-28 |  | NPEPPIN |  | MassPerMass | mg/kg | Y | (mass conc.) |  | 1;6;7-Trimethylnapthalene (mass conc.) | 2245-38-7 | Organics; | Organics; EMS code: PA36 | Chemical |  | LAB
PA39 | ug/g | Benzo(b+j)fluoranthene | MP02 | GC/MS | GC/MS | 0.05 | Y | Y | Y | 274 | 8/13/2004 | 2/26/2015 8:13:00 PM | 2/27/2015 3:49:15 PM | 2004-11-15 |  | UMULLER |  | MassPerMass | mg/kg | Y | (mass conc.) |  | Benzo(b+j)fluoranthene (mass conc.) | 205-99-2, 205-82-3 | Polycyclic Aromatic Hydrocarbons (PAHs); | Polycyclic Aromatic Hydrocarbons (PAHs); EMS code: PA39 | Chemical |  | LAB
PA39 | mg/L | Benzo(b+j)fluoranthene | MW02 | DCM extrn,GC/MS-SIM (PAH) | Water samples are liquid - liquid extracted with dichloromethane. The extract is then cleaned by silica gel column chromatography and analyzed by gas chromatography - mass spectroscopy. | 1e-05 | Y | Y | Y | 833 | 8/18/2003 | 7/10/2024 5:51:00 PM | 8/21/2024 11:42:56 AM | 2004-09-09 |  | UMULLER |  | Concentration | mg/L | Y | (fl. conc.) |  | Benzo(b+j)fluoranthene (fl. conc.) | 205-99-2, 205-82-3 | Polycyclic Aromatic Hydrocarbons (PAHs); | Polycyclic Aromatic Hydrocarbons (PAHs); EMS code: PA39 | Chemical |  | LAB
PA39 | ug/g | Benzo(b+j)fluoranthene | MI01 | Microwave Extractn /GC-MS(PAH) | Soil samples are ground with anhydrous sodium sulphate and microwave extracted with dichloromethane. The extract is then cleaned up by silica gel column chromatography and analyzed by gc-ms in selective ion mode. | 0.01 | Y | Y | Y | 129 | 10/29/2002 4:15:00 PM | 9/23/2021 3:24:00 AM | 10/18/2021 9:46:09 AM | 1998-10-07 |  | NPEPPIN |  | MassPerMass | mg/kg | Y | (mass conc.) |  | Benzo(b+j)fluoranthene (mass conc.) | 205-99-2, 205-82-3 | Polycyclic Aromatic Hydrocarbons (PAHs); | Polycyclic Aromatic Hydrocarbons (PAHs); EMS code: PA39 | Chemical |  | LAB
PA39 | ug/g | Benzo(b+j)fluoranthene | FS87 | DCM,Ac/Hex extr,PAM,HRGC/LRMS | DCM;Ac/Hex extr;PAM;HRGC/LRMS | 0.01 | Y | Y | Y | 81 | 9/20/2017 | 10/29/2022 8:18:00 AM | 1/11/2024 1:07:08 PM | 2018-01-08 |  | JKROGH |  | MassPerMass | mg/kg | Y | (mass conc.) |  | Benzo(b+j)fluoranthene (mass conc.) | 205-99-2, 205-82-3 | Polycyclic Aromatic Hydrocarbons (PAHs); | Polycyclic Aromatic Hydrocarbons (PAHs); EMS code: PA39 | Chemical |  | LAB
PA39 | mg/L | Benzo(b+j)fluoranthene | HEGC | Hexane extraction: GC/MS | Hexane extraction: GC/MS | 1e-05 | Y | Y | Y | 626 | 8/4/2017 | 8/15/2024 | 8/16/2024 1:36:03 PM | 2018-01-08 |  | JKROGH |  | Concentration | mg/L | Y | (fl. conc.) |  | Benzo(b+j)fluoranthene (fl. conc.) | 205-99-2, 205-82-3 | Polycyclic Aromatic Hydrocarbons (PAHs); | Polycyclic Aromatic Hydrocarbons (PAHs); EMS code: PA39 | Chemical |  | LAB
PA40 | mg/L | Benzo(b;j;k)fluoranthene | HEGC | Hexane extraction: GC/MS | Hexane extraction: GC/MS | 1.5e-05 | Y | Y | Y | 504 | 5/2/2018 | 8/15/2024 | 8/16/2024 1:36:03 PM | 2018-07-11 |  | RDGIBSON |  | Concentration | mg/L | Y | (fl. conc.) |  | Benzo(b;j;k)fluoranthene (fl. conc.) | 205-99-2, 205-82-3, 207-08-9 | Polycyclic Aromatic Hydrocarbons (PAHs); | Polycyclic Aromatic Hydrocarbons (PAHs); EMS code: PA40 | Chemical |  | LAB
PA40 | ug/g | Benzo(b;j;k)fluoranthene | FS87 | DCM,Ac/Hex extr,PAM,HRGC/LRMS | DCM;Ac/Hex extr;PAM;HRGC/LRMS | 0.015 | Y | Y | Y | 78 | 9/20/2017 | 9/9/2022 | 1/11/2024 1:07:08 PM | 2018-01-08 |  | JKROGH |  | MassPerMass | mg/kg | Y | (mass conc.) |  | Benzo(b;j;k)fluoranthene (mass conc.) | 205-99-2, 205-82-3, 207-08-9 | Polycyclic Aromatic Hydrocarbons (PAHs); | Polycyclic Aromatic Hydrocarbons (PAHs); EMS code: PA40 | Chemical |  | LAB
PA40 | ug/g | Benzo(b;j;k)fluoranthene | MI01 | Microwave Extractn /GC-MS(PAH) | Soil samples are ground with anhydrous sodium sulphate and microwave extracted with dichloromethane. The extract is then cleaned up by silica gel column chromatography and analyzed by gc-ms in selective ion mode. | 0.01 | Y | Y | Y | 7 | 10/29/2002 4:15:00 PM | 9/22/2003 11:15:00 AM | 3/25/2004 1:00:40 AM | 1998-10-07 |  | NPEPPIN |  | MassPerMass | mg/kg | Y | (mass conc.) |  | Benzo(b;j;k)fluoranthene (mass conc.) | 205-99-2, 205-82-3, 207-08-9 | Polycyclic Aromatic Hydrocarbons (PAHs); | Polycyclic Aromatic Hydrocarbons (PAHs); EMS code: PA40 | Chemical |  | LAB
PA68 | ug/g wet | Dibenzofuran | FB20 | NaSO4,MeOH/Cl2H2Cextr,HRGC/LRMS | The method consisted of first spiking a portion of the undried tissue samples with an aliquot of three different surrogate standard solutions: one for the analysis of alkanes, aliphatics and linear alkylbenzenes in the F1 extract; one for the analysis of PAHs in the F2 extract; and one for alcohols and plant sterols in the F3. Each tissue sample was ground with sodium sulphate, packed in a glass chromatographic column and eluted with methanol and dichloromethane. The eluate was backwashed with extracted water. The extract was concentrated and fractionated on silica gel into three fractions (F1, F2 and F3). F1 and F3 extracts were concentrated and ampouled. The F2 extract containing the PAH compounds of interest was washed with base to remove fatty acid interferences, concentrated and spiked with an aliquot of recovery standard solution prior to analysis by high resolution gas chromatography with low resolution (quadrupole) mass spectrometric detection (HRGC/LRMS). Results reported on a wet weight basis. | 5e-05 | Y | Y | Y | 8 | 12/3/2002 | 3/11/2003 | 4/1/2004 1:13:27 AM | 1999-07-12 |  | NPEPPIN |  | MassPerMass | µg/g | Y | (mass conc.) | (wet) | Dibenzofuran (wet) (mass conc.) | 132-64-9 | Organics; | Organics; EMS code: PA68 | Chemical |  | LAB
PAH- | ug/g | Total Polyaromatic Hydrocarbons | CAL2 | Calculation (ug/g) | Calculation (ug/g) |  | Y | Y | Y | 322 | 3/4/2010 | 9/23/2021 | 10/18/2021 9:46:09 AM | 2009-10-06 |  | RDGIBSON |  | MassPerMass | mg/kg | Y | (mass conc.) |  | Total Polyaromatic Hydrocarbons (mass conc.) |  | Polycyclic Aromatic Hydrocarbons (PAHs); | Polycyclic Aromatic Hydrocarbons (PAHs); EMS code: PAH- | Chemical | Total | LAB
PAH- | mg/L | Total Polyaromatic Hydrocarbons | XCAL | Calculation (mg/L) | Calculation (mg/L) |  | Y | Y | Y | 410 | 5/13/2008 | 7/3/2024 | 7/22/2024 3:17:33 PM | 2009-10-06 | 10/6/2009 11:52:24 AM | RDGIBSON | RDGIBSON | Concentration | mg/L | Y | (fl. conc.) |  | Total Polyaromatic Hydrocarbons (fl. conc.) |  | Polycyclic Aromatic Hydrocarbons (PAHs); | Polycyclic Aromatic Hydrocarbons (PAHs); EMS code: PAH- | Chemical | Total | LAB
PAHH | ug/g | PAH's; High Molecular Weight | CAL2 | Calculation (ug/g) | Calculation (ug/g) |  | Y | Y | Y | 322 | 3/4/2010 | 9/23/2021 | 10/18/2021 9:46:09 AM | 2009-10-06 |  | RDGIBSON |  | MassPerMass | mg/kg | Y | (mass conc.) |  | PAH's; High Molecular Weight (mass conc.) |  | Polycyclic Aromatic Hydrocarbons (PAHs); | Polycyclic Aromatic Hydrocarbons (PAHs); EMS code: PAHH | Chemical |  | LAB
PAHH | mg/L | PAH's; High Molecular Weight | HEGC | Hexane extraction: GC/MS | Hexane extraction: GC/MS | 3e-05 | Y | Y | Y | 13 | 6/5/2023 | 7/28/2024 | 8/6/2024 3:27:04 PM | 2023-12-12 |  | EMS |  | Concentration | mg/L | Y | (fl. conc.) |  | PAH's; High Molecular Weight (fl. conc.) |  | Polycyclic Aromatic Hydrocarbons (PAHs); | Polycyclic Aromatic Hydrocarbons (PAHs); EMS code: PAHH | Chemical |  | LAB
PAHH | mg/L | PAH's; High Molecular Weight | XCAL | Calculation (mg/L) | Calculation (mg/L) |  | Y | Y | Y | 320 | 5/13/2008 | 7/3/2024 | 7/22/2024 3:17:33 PM | 2009-10-06 |  | RDGIBSON |  | Concentration | mg/L | Y | (fl. conc.) |  | PAH's; High Molecular Weight (fl. conc.) |  | Polycyclic Aromatic Hydrocarbons (PAHs); | Polycyclic Aromatic Hydrocarbons (PAHs); EMS code: PAHH | Chemical |  | LAB
PAHL | ug/g | PAH's; Low Molecular Weight | CAL2 | Calculation (ug/g) | Calculation (ug/g) |  | Y | Y | Y | 322 | 3/4/2010 | 9/23/2021 | 10/18/2021 9:46:09 AM | 2009-10-06 |  | RDGIBSON |  | MassPerMass | mg/kg | Y | (mass conc.) |  | PAH's; Low Molecular Weight (mass conc.) |  | Polycyclic Aromatic Hydrocarbons (PAHs); | Polycyclic Aromatic Hydrocarbons (PAHs); EMS code: PAHL | Chemical |  | LAB
PAHL | mg/L | PAH's; Low Molecular Weight | XCAL | Calculation (mg/L) | Calculation (mg/L) |  | Y | Y | Y | 320 | 5/13/2008 | 7/3/2024 | 7/22/2024 3:17:33 PM | 2009-10-06 |  | RDGIBSON |  | Concentration | mg/L | Y | (fl. conc.) |  | PAH's; Low Molecular Weight (fl. conc.) |  | Polycyclic Aromatic Hydrocarbons (PAHs); | Polycyclic Aromatic Hydrocarbons (PAHs); EMS code: PAHL | Chemical |  | LAB
PAHL | mg/L | PAH's; Low Molecular Weight | HEGC | Hexane extraction: GC/MS | Hexane extraction: GC/MS | 6e-05 | Y | Y | Y | 13 | 6/5/2023 | 7/28/2024 | 8/6/2024 3:27:04 PM | 2023-12-12 |  | EMS |  | Concentration | mg/L | Y | (fl. conc.) |  | PAH's; Low Molecular Weight (fl. conc.) |  | Polycyclic Aromatic Hydrocarbons (PAHs); | Polycyclic Aromatic Hydrocarbons (PAHs); EMS code: PAHL | Chemical |  | LAB
PAHM | mg/L | Methylnaphthalenes; 1+2- (PA28+PA29) | HEGC | Hexane extraction: GC/MS | Hexane extraction: GC/MS | 1.5e-05 | Y | Y | Y | 13 | 6/5/2023 | 7/28/2024 | 8/6/2024 3:27:04 PM | 2021-03-15 | 3/15/2021 3:05:00 PM | RDGIBSON | RDGIBSON | Concentration | mg/L | Y | (fl. conc.) |  | Methylnaphthalenes; 1+2- (PA28+PA29) (fl. conc.) | 90-12-0, 91-57-6 | Petroleum Hydrocarbons; | Petroleum Hydrocarbons; EMS code: PAHM | Chemical |  | LAB
PB-D | mg/L | Lead Dissolved | X028 | FA,HNO3,AA Analysis | FA;HNO3;AA Analysis | 0.001 | Y | Y | Y | 113 |  |  | 10/29/2001 11:34:43 AM | 1996-04-12 | 6/3/1998 6:51:38 PM | EMS |  | Concentration | mg/L | Y | (fl. conc.) |  | Lead Dissolved (fl. conc.) |  | Metals - Dissolved; | Metals - Dissolved; EMS code: PB-D | Chemical | Dissolved | LAB
PB-D | mg/L | Lead Dissolved | EQ01 | EQUIS (mg/L) | EQUIS (mg/L) | 0.001 | N | Y | Y | 7018 |  |  | 10/24/1996 3:16:55 PM | 1996-04-12 | 6/3/1998 6:51:38 PM | EMS |  | Concentration | mg/L | Y | (fl. conc.) |  | Lead Dissolved (fl. conc.) |  | Metals - Dissolved; | Metals - Dissolved; EMS code: PB-D | Chemical | Dissolved | LAB
PB-D | mg/L | Lead Dissolved | X544 | FF,FA,HNO3,ICPMS | Analytes nebulized followed by ion entrained in plasma gas and extracted by means of a differentially pumped vacuum interface into a mass spectrometer | 0.001 | Y | Y | Y | 90 |  |  | 10/30/2014 8:31:18 AM | 2001-09-18 | 9/18/2001 1:36:23 PM | NPEPPIN | NPEPPIN | Concentration | mg/L | Y | (fl. conc.) |  | Lead Dissolved (fl. conc.) |  | Metals - Dissolved; | Metals - Dissolved; EMS code: PB-D | Chemical | Dissolved | LAB
PB-D | mg/L | Lead Dissolved | F049 | Filt,HNO3 dig,conc'n,nebul'ICP-OES | This method provides an automated procedure for the Inductively Coupled Argon Plasma Optical Emission Spectrophotometric (ICP-OES) determination and quantitation of trace amounts of metals in surface, ground and precipitation waters after a manual acid digestion and concentration procedure and ultrasonic nebulisation. Organically-bound metals and metal complexes are converted to metal ions by digestion with nitric acid. To increase the signal detected by the ICP, the samples are concentrated 10 times. The liquid sample is pumped onto the piezoelectric transducer, where the ultrasonic action efficiently converts it to a fine, dense aerosol. The nebulizer carrier gas flow then transports the wet aerosol through a heated U-tube, where the solvent is vaporized. Subsequently, most of the solvent vapour is removed from the aerosol stream by passing the stream through a chilled condenser. The result is a dry, analyte-laden aerosol which is introduced to the plasma. There, the high temperature desolvates the aerosol and dissociates the sample material into individual atoms and ions, which are excited to emit light at wave lengths characteristic of the elements present. Dissolved samples are field filtered through 0.45 micron cellulose acetate filter prior to acid preservation. Samples for extractable analysis are acidified with 2 mL concentrated HNO3 per litre and allowed to stand a minimum of 24 hours before analysis. Clients providing these samples are advised to preserve samples at the collection site. | 0.0002 | Y | Y | Y | 5591 |  |  | 5/19/2015 3:10:36 PM | 1997-10-20 | 6/3/1998 6:51:47 PM | NPEPPIN |  | Concentration | mg/L | Y | (fl. conc.) |  | Lead Dissolved (fl. conc.) |  | Metals - Dissolved; | Metals - Dissolved; EMS code: PB-D | Chemical | Dissolved | LAB
PB-D | mg/L | Lead Dissolved | METB | MET-D-NP-BCMOE-MS-VA | Ultra trace metals in water are analyzed by CRC ICPMS, based on US EPA Method 6020B (July 2014). This procedure is intended for pristine field-filtered acid-preserved water samples. The detection limits (LOR) for this test are based on lab instrumental analysis only, not including filtration. ALS-supplied field filtration equipment does not support these LOR. Therefore, because of the high probability of false positives due to filtration, it is strongly recommended that a filtration blank be analysed to aid in data interpretation. BC Ministry of Environment version. | 5e-06 | Y | Y | Y | 596 | 6/11/2019 | 8/15/2024 | 8/23/2024 9:50:35 AM | 2019-08-20 |  | KATFRASE |  | Concentration | mg/L | Y | (fl. conc.) |  | Lead Dissolved (fl. conc.) |  | Metals - Dissolved; | Metals - Dissolved; EMS code: PB-D | Chemical | Dissolved | LAB
PB-D | mg/L | Lead Dissolved | 0051 | Graphite Furnace AA | Graphite Furnace | 0.001 | Y | Y | Y | 1694 |  |  | 1/5/2000 7:22:23 AM | 1996-04-12 | 6/3/1998 6:51:37 PM | EMS |  | Concentration | mg/L | Y | (fl. conc.) |  | Lead Dissolved (fl. conc.) |  | Metals - Dissolved; | Metals - Dissolved; EMS code: PB-D | Chemical | Dissolved | LAB
PB-D | mg/L | Lead Dissolved | F032 | Filt.,Acid,ICPMS | There is no acid digestion required for this analysis. A water sample is collected and preserved in 0.2 percent nitric acid. To analyse the extractable metals the sample is shaken, allowed to sit for at least twenty four hours, and a decanted portion is introduced to the spectrometer. When the dissolved metals are required, the samples must first be filtered through a 0.45 micron cellulose acetate filter paper and then followed by preservation in 0.2 percent nitric acid. It is important to note that different laboratories may have very different detection limits for some variables analyzed by ICPMS. | 0.0005 | Y | Y | Y | 77683 | 8/16/1996 2:35:00 PM | 8/26/2024 | 8/26/2024 4:53:05 PM | 1997-01-16 | 8/15/2019 10:58:35 AM | MPAK | KATFRASE | Concentration | mg/L | Y | (fl. conc.) |  | Lead Dissolved (fl. conc.) |  | Metals - Dissolved; | Metals - Dissolved; EMS code: PB-D | Chemical | Dissolved | LAB
PB-D | mg/L | Lead Dissolved | X556 | Metals in water, dissolved, "in-bottle" digestion, ICP-SFMS | The whole water sample is filtered through a 0.45 micron cellulose acetate filter paper and the filtrate is stored in a HDPE/Teflon bottle. The container is de-capped in a Class 100 workstation, acidified to 2% (v/v) HNO3 and 1% (v/v) HCl and recapped. The sample is "in-bottle" digested in a convection oven, set at 60 degrees C for 16 hours. The sample is cooled and introduced to an inductively coupled argon plasma-sector field mass spectrometer (ICP-SFMS) by pneumatic nebulization. Each respective element is measured at a specific mass to charge ratio, m/z value expressed in amu (atomic mass units), and at a specified intrument resolution. LR: Low Resolution ~300, MR: Mid Resolution ~4000 and HR: High Resolution ~10000. Ref: Method number 02-2003, National Laboratory for Environmental Testing, CCIW, Environment Canada, Burlington, Requested April 2003. | 5e-06 | Y | Y | Y | 4234 | 6/11/2003 | 12/6/2023 2:04:00 AM | 2/14/2024 9:05:32 AM | 2004-01-19 |  | UMULLER |  | Concentration | mg/L | Y | (fl. conc.) |  | Lead Dissolved (fl. conc.) |  | Metals - Dissolved; | Metals - Dissolved; EMS code: PB-D | Chemical | Dissolved | LAB
PB-D | mg/L | Lead Dissolved | X109 | LF, AA | LF; AA | 0.05 | Y | Y | Y | 25 |  |  | 10/24/1996 6:45:14 AM | 1996-04-12 | 6/3/1998 6:51:38 PM | EMS |  | Concentration | mg/L | Y | (fl. conc.) |  | Lead Dissolved (fl. conc.) |  | Metals - Dissolved; | Metals - Dissolved; EMS code: PB-D | Chemical | Dissolved | LAB
PB-D | mg/L | Lead Dissolved | X116 | FF,FA,HNO3,Graph.Furn,AA | Samples are field-filtered (0.45 um), field-acidified with HNO3; analysis by GFAA. | 0.001 | Y | Y | Y | 8592 |  |  | 8/26/2024 2:41:32 PM | 1996-04-12 | 6/3/1998 6:51:38 PM | EMS |  | Concentration | mg/L | Y | (fl. conc.) |  | Lead Dissolved (fl. conc.) |  | Metals - Dissolved; | Metals - Dissolved; EMS code: PB-D | Chemical | Dissolved | LAB
PB-D | mg/L | Lead Dissolved | X170 | ICP | This code is a catch-all for a large variety of ICP methods including water and also solids digested into solutions as an intermediate result. Different procedures will have greatly different detection limits. (a) For air particulates, preliminary results are reported in mg/L and loading results as mg/dm2/d. (b) For Hi Vol samples preliminiary results are same as reported in mg/L and loading results as ug/m3. (c) Note that for Rain samples there is no filtration nor digestion steps. | 0.1 | Y | Y | Y | 15330 | 4/11/1996 6:02:00 PM | 11/6/2018 | 4/12/2024 1:44:29 PM | 1996-04-12 | 6/3/1998 6:51:38 PM | EMS |  | Concentration | mg/L | Y | (fl. conc.) |  | Lead Dissolved (fl. conc.) |  | Metals - Dissolved; | Metals - Dissolved; EMS code: PB-D | Chemical | Dissolved | LAB
PB-D | mg/L | Lead Dissolved | X085 | LF,LA,HNO3,AA | Samples are lab-filtered (0.45 um), and lab-acidified (HNO3); analysis by AA. | 0.001 | Y | Y | Y | 756 | 12/12/1996 | 9/18/2015 | 10/28/2015 9:54:16 AM | 1996-04-12 | 6/3/1998 6:51:38 PM | EMS |  | Concentration | mg/L | Y | (fl. conc.) |  | Lead Dissolved (fl. conc.) |  | Metals - Dissolved; | Metals - Dissolved; EMS code: PB-D | Chemical | Dissolved | LAB
PB-D | mg/L | Lead Dissolved | 0050 | Graphite Furnace AA | Graphite Furnace AA | 0.001 | Y | Y | Y | 2858 | 8/28/2002 | 6/5/2003 | 6/9/2003 10:20:23 AM | 1996-04-12 | 6/3/1998 6:51:37 PM | EMS |  | Concentration | mg/L | Y | (fl. conc.) |  | Lead Dissolved (fl. conc.) |  | Metals - Dissolved; | Metals - Dissolved; EMS code: PB-D | Chemical | Dissolved | LAB
PB-D | mg/L | Lead Dissolved | F068 | HNO3dig,TeflonTube/Autoclave,ICP-MS | This method is applicable to determiniation of dissolved, extractable, and total recoverable metals in ground waters, surface waters, drinking waters, and wastewaters. Dissolved, extractable, and total recoverable sample fractions are acidified to pH <1.5 with analytical grade nitric acid (HNO3) to prevent the metals from adsorbing to the container material or precipitating. Extractable and total recoverable water samples are collected in 250 ml acid washed polyethylene bottles and are preserved with Nitric acid (0.5 ml acid for 100 mLs of sample). All preserved samples are stored at room temperature and can be held for up to 6 months. Unfiltered, un-preserved samples are stored at 4C for no more than 30 days. For dissolved and extractable water fractions, no sample preparation is necessary. Samples are simply poured in autosampler tubes (i.e. 15 ml centrifuge tubes), and placed on the autosampler and analyzed. Samples for total recoverable metals are digested with Nitric acid in special Teflon tubes in an autoclave prior to analyses. | 1e-05 | Y | Y | Y | 1818 | 6/1/1999 1:21:00 PM | 4/17/2015 1:06:00 PM | 7/18/2024 7:48:39 AM | 1999-06-28 |  | NPEPPIN |  | Concentration | mg/L | Y | (fl. conc.) |  | Lead Dissolved (fl. conc.) |  | Metals - Dissolved; | Metals - Dissolved; EMS code: PB-D | Chemical | Dissolved | LAB
PB-D | mg/L | Lead Dissolved | F039 | Field filt + acid, ICP-AES | Dissolved and extractable water samples are ready for analysis once collected. Total samples are autoclaved with HNO3 and hydrochloric acid (HCI) in capped Teflon vials. Analysis is performed via Inductively Coupled Argon Plasma Atomic Emission Spectrometry (ICP-AES)(simultaneous multielement analysis), Graphite Furnace Atomic Absorption Spectrometry (GF-AAS)(low level analysis) and/or Flame Atomic Absorption Spectrometry (F-AAS). | 0.05 | Y | Y | Y | 6488 | 7/29/1996 6:03:00 PM | 7/5/2019 1:21:00 PM | 9/11/2019 8:47:50 AM | 1997-05-08 | 6/3/1998 6:51:46 PM | MPAK |  | Concentration | mg/L | Y | (fl. conc.) |  | Lead Dissolved (fl. conc.) |  | Metals - Dissolved; | Metals - Dissolved; EMS code: PB-D | Chemical | Dissolved | LAB
PB-D | mg/L | Lead Dissolved | X095 | FF,FA,HNO3,ICP | FF;FA;HNO3;ICP | 0.001 | Y | Y | Y | 214 |  |  | 10/24/1996 12:20:00 PM | 1996-04-12 | 6/3/1998 6:51:38 PM | EMS |  | Concentration | mg/L | Y | (fl. conc.) |  | Lead Dissolved (fl. conc.) |  | Metals - Dissolved; | Metals - Dissolved; EMS code: PB-D | Chemical | Dissolved | LAB
PB-D | mg/L | Lead Dissolved | F023 | HNO3,GFAA | Dissolved and extractable water samples are ready for analysis once collected. Total samples are autoclaved with HNO3 and hydrochloric acid (HCl) in capped Teflon vials. Analysis is performed via Inductively Coupled Argon Plasma Atomic Emission Spectrometry (ICP-AES) (simultaneous multielement analysis), Graphite Furnace Atomic Absorption Spectrometry (GF-AAS) (low level analysis) and/or Flame Atomic Absorption Spectrometry (F-AAS). Note that extractable samples are acidified with 0.5 mL HNO3 per 100 mL of sample at time of sample collection. | 0.0005 | Y | Y | Y | 487 | 8/6/1996 7:16:00 PM | 9/16/2003 3:18:00 PM | 11/28/2003 3:22:14 PM | 1996-12-11 | 6/3/1998 6:51:46 PM | MPAK |  | Concentration | mg/L | Y | (fl. conc.) |  | Lead Dissolved (fl. conc.) |  | Metals - Dissolved; | Metals - Dissolved; EMS code: PB-D | Chemical | Dissolved | LAB
PB-D | mg/L | Lead Dissolved | 0090 | Ext.(APDC):AA Anal. | Ext.(APDC):AA Anal. | 0.001 | Y | Y | Y | 40 |  |  | 10/24/1996 11:35:00 AM | 1996-04-12 | 6/3/1998 6:51:37 PM | EMS |  | Concentration | mg/L | Y | (fl. conc.) |  | Lead Dissolved (fl. conc.) |  | Metals - Dissolved; | Metals - Dissolved; EMS code: PB-D | Chemical | Dissolved | LAB
PB-D | mg/L | Lead Dissolved | X285 | HNO3 Dig:ICP Analysis | HNO3 Dig:ICP Analysis | 0.005 | Y | Y | Y | 156 | 1/30/2002 | 11/13/2002 | 3/18/2004 10:28:34 AM | 1998-05-22 | 6/3/1998 6:51:49 PM | NPEPPIN |  | Concentration | mg/L | Y | (fl. conc.) |  | Lead Dissolved (fl. conc.) |  | Metals - Dissolved; | Metals - Dissolved; EMS code: PB-D | Chemical | Dissolved | LAB
PB-D | mg/L | Lead Dissolved | X072 | FA,HNO3,Graph.Furnace,AA | Samples are field-acidified with HNO3; HNO3 digestion; analysis by GFAA. | 0.001 | Y | Y | Y | 156 |  |  | 10/24/1996 1:00:47 PM | 1996-04-12 | 6/3/1998 6:51:41 PM | EMS |  | Concentration | mg/L | Y | (fl. conc.) |  | Lead Dissolved (fl. conc.) |  | Metals - Dissolved; | Metals - Dissolved; EMS code: PB-D | Chemical | Dissolved | LAB
PB-D | mg/L | Lead Dissolved | X562 | Dissolved metals: FF, FA, HNO3, HR-ICPMS | Ultra trace metals in water are analyzed by high resolution inductively coupled plasma mass spectrometry (HR-ICPMS), based on US EPA Method 200.8, (Oct. 1999). This procedure is intended for pristine field-filtered acid-preserved water samples. The detection limits (LOR) for this test are based on lab instrumental analysis only, not including filtration. ALS-supplied field filtration equipment does not support these LOR. Therefore, because of the high probability of false positives due to filtration, it is strongly recommended that a filtration blank be analyzed to aid in data interpretation. | 5e-06 | Y | Y | Y | 1264 | 6/8/2015 | 10/19/2023 | 11/9/2023 7:41:27 AM | 2015-10-22 |  | RDGIBSON |  | Concentration | mg/L | Y | (fl. conc.) |  | Lead Dissolved (fl. conc.) |  | Metals - Dissolved; | Metals - Dissolved; EMS code: PB-D | Chemical | Dissolved | LAB
PB-D | mg/L | Lead Dissolved | X113 | HNO3/HCl/Perchlor,AA | HNO3/HCl/Perchlor;AA | 0.001 | Y | Y | Y | 2776 |  |  | 11/15/2000 1:53:55 PM | 1996-04-12 | 6/3/1998 6:51:41 PM | EMS |  | Concentration | mg/L | Y | (fl. conc.) |  | Lead Dissolved (fl. conc.) |  | Metals - Dissolved; | Metals - Dissolved; EMS code: PB-D | Chemical | Dissolved | LAB
PB-D | mg/L | Lead Dissolved | X356 | LF,LA:HNO3:ICP | Samples are lab-filtered (0.45 um); lab-acidified with HNO3; analysis by ICP/AES. | 0.02 | Y | Y | Y | 523 | 10/17/2000 | 5/1/2002 | 3/23/2010 5:08:09 PM | 1996-04-12 | 6/3/1998 6:51:38 PM | EMS |  | Concentration | mg/L | Y | (fl. conc.) |  | Lead Dissolved (fl. conc.) |  | Metals - Dissolved; | Metals - Dissolved; EMS code: PB-D | Chemical | Dissolved | LAB
PB-D | mg/L | Lead Dissolved | X027 | Ext,APDC,AA | Ext;APDC;AA | 0.001 | Y | Y | Y | 9 |  |  | 10/24/1996 6:45:08 AM | 1996-04-12 | 6/3/1998 6:51:38 PM | EMS |  | Concentration | mg/L | Y | (fl. conc.) |  | Lead Dissolved (fl. conc.) |  | Metals - Dissolved; | Metals - Dissolved; EMS code: PB-D | Chemical | Dissolved | LAB
PB-D | mg/L | Lead Dissolved | 0020 | HNO3 Dig:AA Analysis | HNO3 Dig:AA Analysis | 0.1 | Y | Y | Y | 2 |  |  | 10/24/1996 11:30:41 AM | 1996-04-12 | 6/3/1998 6:51:37 PM | EMS |  | Concentration | mg/L | Y | (fl. conc.) |  | Lead Dissolved (fl. conc.) |  | Metals - Dissolved; | Metals - Dissolved; EMS code: PB-D | Chemical | Dissolved | LAB
PB-D | mg/L | Lead Dissolved | MWID | Metals in Water by ICPMS (digested) | Metals in Water by ICPMS (digested) Based on EPA 200.8 | 0.0002 | Y | Y | Y | 31784 | 10/26/2018 5:20:00 PM | 10/26/2018 5:20:00 PM | 7/31/2024 1:46:41 PM | 2009-12-24 |  | RDGIBSON |  | Concentration | mg/L | Y | (fl. conc.) |  | Lead Dissolved (fl. conc.) |  | Metals - Dissolved; | Metals - Dissolved; EMS code: PB-D | Chemical | Dissolved | LAB
PB-D | t/d | Lead Dissolved | CALC | Calculation (t/d) | Calculation (t/d) | 0 | Y | Y | Y | 3755 |  |  | 7/3/2024 12:59:26 PM | 1996-04-12 | 6/3/1998 6:51:41 PM | EMS |  | MassFlow | t/d | Y | (mass flow) |  | Lead Dissolved (mass flow) |  | Metals - Dissolved; | Metals - Dissolved; EMS code: PB-D | Chemical | Dissolved | LAB
PB-D | mg/L | Lead Dissolved | MSEA | Metals in Seawater (undigested) by (CRC) ICP-MS | This method describes the multi-elemental determination of trace elements by Collision/Reaction Cell (CRC) ICP-MS. The method measures ions produced by a radio-frequency inductively coupled plasma. Analyte species originating in a liquid are pneumatically nebulized and the resulting aerosol transported by argon gas into the plasma torch. The ions produced are entrained in the plasma gas and extracted, by means of a differentially pumped vacuum interface, into a mass spectrometer. The ions flow through a Collision/Reaction Cell (CRC) where depending on the mode of analysis either the ions passes through the cell unhindered (no gas - classical ICP-MS mode), or collisions/reactions chemistry takes place between the ions and a neutral gas such as hydrogen or helium. Any polyatomic ions which give interference on the analytes are removed. The ions are then sorted according to their mass-to-charge ratios by a quadrupole mass spectrometer. The ions are then sorted according to their mass-to-charge ratios by a quadrupole mass spectrometer. The ions transmitted through the quadrupole are quantified by a channel electron multiplier and the ion information processed by a data handling system. Although most polyatomic interference is removed, any additional interference must be assessed and valid corrections applied or the data flagged to indicate problems. Interference correction must include compensation for background ions contributed by the plasma gas, reagents, and constituents of the sample matrix. Instrument drift as well as suppressions or enhancements of instrument response caused by the sample matrix must be corrected for by the use of internal standards. | 1e-05 | Y | Y | Y | 4 | 4/12/2023 | 4/29/2024 | 5/16/2024 2:00:14 PM | 2009-09-08 |  | RDGIBSON |  | Concentration | mg/L | Y | (fl. conc.) |  | Lead Dissolved (fl. conc.) |  | Metals - Dissolved; | Metals - Dissolved; EMS code: PB-D | Chemical | Dissolved | LAB
PB-D | mg/L | Lead Dissolved | METC | MET-D-NP-CCMS-VA | Ultra trace metals in water are analyzed by CRC ICPMS, based on US EPA Method 6020B (July 2014). This procedure is intended for pristine field-filtered acid-preserved water samples. The detection limits (LOR) for this test are based on lab instrumental analysis only, not including filtration. ALS-supplied field filtration equipment does not support these LOR. Therefore, because of the high probability of false positives due to filtration, it is strongly recommended that a filtration blank be analysed to aid in data interpretation. | 5e-06 | Y | Y | Y | 1179 | 4/3/2012 | 4/9/2024 | 4/29/2024 4:53:06 PM | 2021-04-14 |  | RDGIBSON |  | Concentration | mg/L | Y | (fl. conc.) |  | Lead Dissolved (fl. conc.) |  | Metals - Dissolved; | Metals - Dissolved; EMS code: PB-D | Chemical | Dissolved | LAB
PB-E | mg/L | Lead Extractable | F069 | SEM extr,6N HCl,ICP/OES | 6.0 N Hydrochloric acid (HCI) is added to an aliquot of the sediment sample. The extract produced from the addition of the acid is then analyzed for simultaneously extracted metals (SEM) using inductively coupled plasma - optical emission spectrophotometry. The detection limits depend on the metal parameters in question and the instrumentation used. | 0.05 | Y | Y | Y | 10 | 3/26/2012 3:01:00 AM | 3/26/2012 3:01:00 AM | 11/25/2021 1:50:15 PM | 1999-07-09 |  | NPEPPIN |  | Concentration | mg/L | Y | (fl. conc.) |  | Lead Extractable (fl. conc.) |  | Metals - Extractable; | Metals - Extractable; EMS code: PB-E | Chemical | Extractable | LAB
PB-E | ug/g | Lead Extractable | X226 | Hom,HNO3/HCl,AA | Hom;HNO3/HCl;AA | 2.5 | Y | Y | Y | 70 |  |  | 10/24/1996 2:25:28 PM | 1996-04-12 | 6/3/1998 6:51:38 PM | EMS |  | MassPerMass | mg/kg | Y | (mass conc.) |  | Lead Extractable (mass conc.) |  | Metals - Extractable; | Metals - Extractable; EMS code: PB-E | Chemical | Extractable | LAB
PB-E | mg/L | Lead Extractable | F040 | Field Acid, lab filt, ICP-AES | Dissolved and extractable water samples are ready for analysis once collected. Total samples are autoclaved with HNO3 and hydrochloric acid (HCI) in capped Teflon vials. Analysis is performed via Inductively Coupled Argon Plasma Atomic Emission Spectrometry (ICP-AES)(simultaneous multielement analysis), Graphite Furnace Atomic Absorption Spectrometry (GF-AAS)(low level analysis) and/or Flame Atomic Absorption Spectrometry (F-AAS). | 0.05 | Y | Y | Y | 2361 | 8/5/1997 6:06:00 PM | 3/14/2008 | 9/11/2019 8:47:41 AM | 1997-05-08 | 6/3/1998 6:51:46 PM | MPAK |  | Concentration | mg/L | Y | (fl. conc.) |  | Lead Extractable (fl. conc.) |  | Metals - Extractable; | Metals - Extractable; EMS code: PB-E | Chemical | Extractable | LAB
PB-E | umol/g | Lead Extractable | FSEM | AVS extraction, SEM Analysis ICPMS | Determination of simultaneous extractable metals (SEM) in sediment after sample extracted with hydrochloric acid for acid volatile sulfide (AVS) | 0.2 | Y | Y | Y | 19 | 3/12/2020 | 3/24/2020 | 2/22/2023 12:51:03 PM | 2010-11-16 |  | RDGIBSON |  | ParticlesPerMass | µmol/g | Y | (part. /m) |  | Lead Extractable (part. /m) |  | Metals - Extractable; | Metals - Extractable; EMS code: PB-E | Chemical | Extractable | LAB
PB-E | mg/L | Lead Extractable | X553 | Metals in water, extractable, ICP-SFMS | The whole water sample is collected in a HDPE/Teflon bottle. In a Class 100 workstation it is preserved with concentrated Nitric Acid at concentration of 0.2 percent (v/v). The acid preservative extracts the metals from material in the sample and a decanted portion is introduced directly to an inductively coupled argon plasma-sector field mass spectrometer (ICP-SFMS) by pneumatic nebulization. Each respective element is measured at a specific mass to charge ratio, m/z value expressed in amu (atomic mass units), and at a specified instrument resolution. LR: Low Resolution ~300, MR: Mid Resolution ~4000 and HR: High Resolution ~10000. Ref: Method number 02-2061, National Laboratory for Environmental Testing, CCIW, Environment Canada, Burlington. Requested April 2003. | 5e-06 | Y | Y | Y | 2597 | 5/22/2003 | 9/21/2007 | 11/28/2023 2:11:53 PM | 2004-01-19 |  | UMULLER |  | Concentration | mg/L | Y | (fl. conc.) |  | Lead Extractable (fl. conc.) |  | Metals - Extractable; | Metals - Extractable; EMS code: PB-E | Chemical | Extractable | LAB
PB-E | mg/L | Lead Extractable | F047 | 24hr HNO3 dil'n,ICP-MS | There is no acid digestion required for this analysis. A water sample is collected and preserved in 0.2 percent nitric acid. To analyse the extractable metals the sample is shaken, allowed to sit for at least twenty four hours, and a decanted portion is introduced to the spectrometer. When the dissolved metals are required, the samples must first be filtered through a 0.45 micron cellulose acetate filter paper and then followed by preservation in 0.2 percent nitric acid. It is important to note that different laboratories may have very different detection limits for some variables analyzed by ICP/MS. | 1e-05 | Y | Y | Y | 2680 | 4/16/1998 10:21:00 AM | 12/13/2011 12:02:00 PM | 9/11/2019 8:46:18 AM | 1997-10-20 | 6/3/1998 6:51:47 PM | NPEPPIN |  | Concentration | mg/L | Y | (fl. conc.) |  | Lead Extractable (fl. conc.) |  | Metals - Extractable; | Metals - Extractable; EMS code: PB-E | Chemical | Extractable | LAB
PB-E | mg/L | Lead Extractable | X512 | HNO3,pHto7.5,MIK extr'n,AA | HNO3 Digestion;pH to 7.5; MIK solvent extraction;AA | 0.001 | Y | Y | Y | 875 |  |  | 8/23/1999 3:43:46 PM | 1999-04-22 |  | NPEPPIN |  | Concentration | mg/L | Y | (fl. conc.) |  | Lead Extractable (fl. conc.) |  | Metals - Extractable; | Metals - Extractable; EMS code: PB-E | Chemical | Extractable | LAB
PB-I | mg/dm2/d | Lead-Insoluble | 5039 | N/P Dig, ICP | Aliquot of prepared sample is digested first with nitric acid then with perchloric acid. Analysis by ICP. (a) For air particulates, preliminary results are reported in mg/L and loading results as mg/dm2/d. (b) For Hi Vol samples preliminary results are same as reported in mg/L and loading results as ug/m3. | 0.0005 | Y | Y | Y | 198 | 12/5/1997 | 5/15/2000 | 11/18/2000 1:30:13 AM | 1996-04-12 | 6/3/1998 6:51:38 PM | EMS |  | MassFlux | mg/dm²/d | Y | (mass flux) |  | Lead-Insoluble (mass flux) |  | Metals - Insoluble; | Metals - Insoluble; EMS code: PB-I | Chemical | Insoluble | LAB
PB-I | mg/dm2/d | Lead-Insoluble | 5037 | N/P Dig, HGA Analysis | N/P Dig; HGA Analysis | 0.0001 | Y | Y | Y | 6 |  |  | 10/24/1996 6:47:55 AM | 1996-04-12 | 6/3/1998 6:51:38 PM | EMS |  | MassFlux | mg/dm²/d | Y | (mass flux) |  | Lead-Insoluble (mass flux) |  | Metals - Insoluble; | Metals - Insoluble; EMS code: PB-I | Chemical | Insoluble | LAB
PB-I | mg/dm2/d | Lead-Insoluble | EQ15 | EQUIS (mg/dm2/d) | EQUIS (mg/dm2/d) | 0.001 | N | Y | Y | 169 |  |  | 10/22/1996 10:18:48 PM | 1996-04-12 | 6/3/1998 6:51:38 PM | EMS |  | MassFlux | mg/dm²/d | Y | (mass flux) |  | Lead-Insoluble (mass flux) |  | Metals - Insoluble; | Metals - Insoluble; EMS code: PB-I | Chemical | Insoluble | LAB
PB-L | mg/L | Leachable Lead after EPA TCLP extraction of solid sample | 1312 | EPA 1311/6010B (ICPMS) | This analysis is carried out in accordance with the extraction procedure outlined in "Test Methods for Evaluating Solid Waste - Physical/Chemical Methods Volume 1C" SW-846 EPA Method 1311, published by the United States Environmental Protection Agency (EPA). In summary, the sample is extracted at a 20:1 liquid to solids ratio for 16 to 20 hours using either extraction fluid #1 (glacial acetic acid, water and sodium hydroxide) or extraction fluid #2 (glacial acetic acid), depending on the pH of the original sample. The extract is then filtered through a 0.6 to 0.8 micron glass fibre filter and analysed using inductively coupled plasma - optical emission spectrophotometry (EPA Method 6010B) | 0.25 | Y | Y | Y | 16 |  |  | 8/24/2016 2:59:13 PM | 2011-04-27 |  | RDGIBSON |  | Concentration | mg/L | Y | (fl. conc.) |  | Leachable Lead after EPA TCLP extraction of solid sample (fl. conc.) |  | Metals - Extractable; | Metals - Extractable; EMS code: PB-L | Chemical | Leachable | LAB
PB-S | mg/dm2/d | Lead-Soluble | 0050 | Graphite Furnace AA | Graphite Furnace AA | 2e-05 | Y | Y | Y | 36 |  |  | 10/24/1996 6:48:05 AM | 1996-04-12 | 6/3/1998 6:51:41 PM | EMS |  | MassFlux | mg/dm²/d | Y | (mass flux) |  | Lead-Soluble (mass flux) |  | Metals - Soluble; | Metals - Soluble; EMS code: PB-S | Chemical | Soluble | LAB
PB-S | umol/g | Lead-Soluble | FSEM | AVS extraction, SEM Analysis ICPMS | Determination of simultaneous extractable metals (SEM) in sediment after sample extracted with hydrochloric acid for acid volatile sulfide (AVS) | 0.2 | Y | Y | Y | 0 |  |  |  | 2024-02-02 |  | EMS |  | ParticlesPerMass | µmol/g | Y | (part. /m) |  | Lead-Soluble (part. /m) |  | Metals - Soluble; | Metals - Soluble; EMS code: PB-S | Chemical | Soluble | LAB
PB-S | mg/dm2/d | Lead-Soluble | EQ15 | EQUIS (mg/dm2/d) | EQUIS (mg/dm2/d) | 0.001 | N | Y | Y | 174 |  |  | 10/22/1996 10:18:48 PM | 1996-04-12 | 6/3/1998 6:51:38 PM | EMS |  | MassFlux | mg/dm²/d | Y | (mass flux) |  | Lead-Soluble (mass flux) |  | Metals - Soluble; | Metals - Soluble; EMS code: PB-S | Chemical | Soluble | LAB
PB-S | mg/dm2/d | Lead-Soluble | 5045 | ICP Analysis | ICP a) For Air particulates analyzed without an acid digestion step, preliminary results are reported in mg/L and loading results as mg/dm2/d. b) For HiVol samples, preliminary results are reported as mg/L and loading results as ug/m3. | 0.01 | Y | Y | Y | 39 | 11/28/1997 | 5/15/2000 | 11/18/2000 1:30:13 AM | 1998-10-28 |  | NPEPPIN |  | MassFlux | mg/dm²/d | Y | (mass flux) |  | Lead-Soluble (mass flux) |  | Metals - Soluble; | Metals - Soluble; EMS code: PB-S | Chemical | Soluble | LAB
PB-S | mg/dm2/d | Lead-Soluble | X170 | ICP | This code is a catch-all for a large variety of ICP methods including water and also solids digested into solutions as an intermediate result. Different procedures will have greatly different detection limits. (a) For air particulates, preliminary results are reported in mg/L and loading results as mg/dm2/d. (b) For Hi Vol samples preliminiary results are same as reported in mg/L and loading results as ug/m3. (c) Note that for Rain samples there is no filtration nor digestion steps. | 0.01 | Y | Y | Y | 143 |  |  | 2/13/1998 2:28:23 AM | 1996-04-12 | 6/3/1998 6:51:38 PM | EMS |  | MassFlux | mg/dm²/d | Y | (mass flux) |  | Lead-Soluble (mass flux) |  | Metals - Soluble; | Metals - Soluble; EMS code: PB-S | Chemical | Soluble | LAB
PB-T | mg/L | Lead Total | 1349 | Tot Metals in Seawater by CRC ICPMS (BC) | Tot Metals in Seawater by CRC ICPMS (BC) | 5e-05 | Y | Y | Y | 98 | 8/16/2019 | 12/5/2021 | 12/13/2021 12:34:16 PM | 2019-09-30 |  | KATFRASE |  | Concentration | mg/L | Y | (fl. conc.) |  | Lead Total (fl. conc.) |  | Metals - Total; | Metals - Total; EMS code: PB-T | Chemical | Total | LAB
PB-T | mg/L | Lead Total | 0042 | GI HNO3 Dig: ICP | Sample aliquot taken from unfiltered unpreserved General Ions bottle; HNO3 digestion; ICP analysis. | 0.03 | Y | Y | Y | 6243 | 9/20/1996 | 7/5/2019 3:21:00 PM | 7/12/2019 6:49:13 PM | 1996-04-12 | 6/3/1998 6:51:38 PM | EMS |  | Concentration | mg/L | Y | (fl. conc.) |  | Lead Total (fl. conc.) |  | Metals - Total; | Metals - Total; EMS code: PB-T | Chemical | Total | LAB
PB-T | mg/L | Lead Total | TCLP | Toxicity Characteristic Leaching Procedure | The TCLP procedure is designed to simulate the climatic leaching action | 0.1 | Y | Y | Y | 22 | 12/8/2009 7:49:00 PM | 9/28/2011 5:53:00 PM | 11/3/2011 11:10:21 PM | 2009-03-06 |  | JCHUDAK |  | Concentration | mg/L | Y | (fl. conc.) |  | Lead Total (fl. conc.) |  | Metals - Total; | Metals - Total; EMS code: PB-T | Chemical | Total | LAB
PB-T | mg/L | Lead Total | 0060 | HNO3 Dig: HGA Analysis | HNO3 Dig: HGA Analysis | 0.001 | Y | Y | Y | 4887 |  |  | 2/4/2000 9:38:27 AM | 1996-04-12 | 6/3/1998 6:51:41 PM | EMS |  | Concentration | mg/L | Y | (fl. conc.) |  | Lead Total (fl. conc.) |  | Metals - Total; | Metals - Total; EMS code: PB-T | Chemical | Total | LAB
PB-T | mg/L | Lead Total | MSEA | Metals in Seawater (undigested) by (CRC) ICP-MS | This method describes the multi-elemental determination of trace elements by Collision/Reaction Cell (CRC) ICP-MS. The method measures ions produced by a radio-frequency inductively coupled plasma. Analyte species originating in a liquid are pneumatically nebulized and the resulting aerosol transported by argon gas into the plasma torch. The ions produced are entrained in the plasma gas and extracted, by means of a differentially pumped vacuum interface, into a mass spectrometer. The ions flow through a Collision/Reaction Cell (CRC) where depending on the mode of analysis either the ions passes through the cell unhindered (no gas - classical ICP-MS mode), or collisions/reactions chemistry takes place between the ions and a neutral gas such as hydrogen or helium. Any polyatomic ions which give interference on the analytes are removed. The ions are then sorted according to their mass-to-charge ratios by a quadrupole mass spectrometer. The ions are then sorted according to their mass-to-charge ratios by a quadrupole mass spectrometer. The ions transmitted through the quadrupole are quantified by a channel electron multiplier and the ion information processed by a data handling system. Although most polyatomic interference is removed, any additional interference must be assessed and valid corrections applied or the data flagged to indicate problems. Interference correction must include compensation for background ions contributed by the plasma gas, reagents, and constituents of the sample matrix. Instrument drift as well as suppressions or enhancements of instrument response caused by the sample matrix must be corrected for by the use of internal standards. | 1e-05 | Y | Y | Y | 432 | 8/12/2009 12:08:00 AM | 4/29/2024 | 6/7/2024 10:03:16 AM | 2009-09-08 |  | RDGIBSON |  | Concentration | mg/L | Y | (fl. conc.) |  | Lead Total (fl. conc.) |  | Metals - Total; | Metals - Total; EMS code: PB-T | Chemical | Total | LAB
PB-T | t/d | Lead Total | CALC | Calculation (t/d) | Calculation (t/d) | 0 | Y | Y | Y | 5776 |  |  | 7/3/2024 12:59:29 PM | 1996-04-12 | 6/3/1998 6:51:38 PM | EMS |  | MassFlow | t/d | Y | (mass flow) |  | Lead Total (mass flow) |  | Metals - Total; | Metals - Total; EMS code: PB-T | Chemical | Total | LAB
PB-T | None | Lead Total | AISP | Special Analysis | Special Analysis | 0 | N | Y | N | 8 |  |  | 10/24/1996 11:23:43 AM | 1996-04-12 | 6/3/1998 6:51:41 PM | EMS |  | None | None | N | (unitless) |  | Lead Total (unitless) |  | Metals - Total; | Metals - Total; EMS code: PB-T | Chemical | Total | LAB
PB-T | ug/g | Lead Total | 0196 | D&G: LTA/HCl: ICP | D&G: LTA/HCl: ICP | 1 | Y | Y | Y | 19 |  |  | 10/24/1996 2:08:13 PM | 1996-04-12 | 6/3/1998 6:51:41 PM | EMS |  | MassPerMass | mg/kg | Y | (mass conc.) |  | Lead Total (mass conc.) |  | Metals - Total; | Metals - Total; EMS code: PB-T | Chemical | Total | LAB
PB-T | ug/g | Lead Total | 0190 | D&G:N/P Dig:ICP Anal. | Sample is dried and grind-mixed before nitric perchloric digestion followed by ICP analysis. | 10 | Y | Y | Y | 920 |  |  | 10/24/1996 3:05:58 PM | 1996-04-12 | 6/3/1998 6:51:41 PM | EMS |  | MassPerMass | mg/kg | Y | (mass conc.) |  | Lead Total (mass conc.) |  | Metals - Total; | Metals - Total; EMS code: PB-T | Chemical | Total | LAB
PB-T | mg/L | Lead Total | METT | MET-T-NP-BCMOE-MS-VA | Ultra trace metals in water are analyzed by CRC ICPMS, based on US EPA Method 6020B (July 2014). The detection limits provided can only be met for undigested samples. This procedure is intended for colorless, non-turbid, acid-preserved water samples (i.e. pristine water samples), having turbidity < 1 NTU and no odor. Where turbidity exceeds 1 NTU, and/or the sample is colored and has an odor, results may be biased low compared to true Total Metals concentrations. ALS recommends that turbidity analysis be requested on samples submitted for this test to aid with interpretation of results. Where turbidity is <1NTU, undigested metals are equivalent to total metals concentrations. BC Ministry of Environment version. | 5e-06 | Y | Y | Y | 2589 | 6/5/2019 | 8/18/2024 | 8/26/2024 1:03:06 PM | 2019-08-20 |  | KATFRASE |  | Concentration | mg/L | Y | (fl. conc.) |  | Lead Total (fl. conc.) |  | Metals - Total; | Metals - Total; EMS code: PB-T | Chemical | Total | LAB
PB-T | mg/L | Lead Total | X072 | FA,HNO3,Graph.Furnace,AA | Samples are field-acidified with HNO3; HNO3 digestion; analysis by GFAA. | 0.001 | Y | Y | Y | 7985 |  |  | 8/26/2024 2:41:32 PM | 1996-04-12 | 6/3/1998 6:51:38 PM | EMS |  | Concentration | mg/L | Y | (fl. conc.) |  | Lead Total (fl. conc.) |  | Metals - Total; | Metals - Total; EMS code: PB-T | Chemical | Total | LAB
PB-T | mg/L | Lead Total | 0020 | HNO3 Dig:AA Analysis | HNO3 Dig:AA Analysis | 0.1 | Y | Y | Y | 92 | 6/25/2010 3:15:00 PM | 12/1/2010 | 12/9/2010 3:55:27 PM | 1996-04-12 | 6/3/1998 6:51:41 PM | EMS |  | Concentration | mg/L | Y | (fl. conc.) |  | Lead Total (fl. conc.) |  | Metals - Total; | Metals - Total; EMS code: PB-T | Chemical | Total | LAB
PB-T | mg/L | Lead Total | MWID | Metals in Water by ICPMS (digested) | Metals in Water by ICPMS (digested) Based on EPA 200.8 | 0.0002 | Y | Y | Y | 36976 |  |  | 7/31/2024 1:46:41 PM | 2009-12-24 |  | RDGIBSON |  | Concentration | mg/L | Y | (fl. conc.) |  | Lead Total (fl. conc.) |  | Metals - Total; | Metals - Total; EMS code: PB-T | Chemical | Total | LAB
PB-T | mg/L | Lead Total | X113 | HNO3/HCl/Perchlor,AA | HNO3/HCl/Perchlor;AA | 0.001 | Y | Y | Y | 1376 |  |  | 11/15/2000 1:51:19 PM | 1996-04-12 | 6/3/1998 6:51:38 PM | EMS |  | Concentration | mg/L | Y | (fl. conc.) |  | Lead Total (fl. conc.) |  | Metals - Total; | Metals - Total; EMS code: PB-T | Chemical | Total | LAB
PB-T | ppm (S) | Lead Total | 6041 | Aqua Regia by ICPMS | These were sediment/soils digested with Aqua-Regia (Nitric Acid & Hydrochloric Acids) and then analysis by ICPMS. Results would be reported in mg/kg and/or %. | 0.005 | Y | Y | Y | 21 | 10/13/2019 | 8/5/2020 | 8/10/2020 3:52:06 PM | 2019-09-06 | 10/21/2019 11:55:53 AM | KATFRASE | KATFRASE | DimensionlessRatio | ppm | N | (ratio) |  | Lead Total (ratio) |  | Metals - Total; | Metals - Total; EMS code: PB-T | Chemical | Total | LAB
PB-T | mg/L | Lead Total | 0040 | HNO3 Dig: ICP Analysis | HNO3 Dig: ICP Analysis | 0.1 | Y | Y | Y | 10700 | 10/17/2000 | 10/31/2003 | 10/18/2023 8:55:31 AM | 1996-04-12 | 6/3/1998 6:51:38 PM | EMS |  | Concentration | mg/L | Y | (fl. conc.) |  | Lead Total (fl. conc.) |  | Metals - Total; | Metals - Total; EMS code: PB-T | Chemical | Total | LAB
PB-T | mg/dm2/d | Lead Total | EQ15 | EQUIS (mg/dm2/d) | EQUIS (mg/dm2/d) | 0.001 | N | Y | Y | 206 |  |  | 10/22/1996 10:18:48 PM | 1996-04-12 | 6/3/1998 6:51:38 PM | EMS |  | MassFlux | mg/dm²/d | Y | (mass flux) |  | Lead Total (mass flux) |  | Metals - Total; | Metals - Total; EMS code: PB-T | Chemical | Total | LAB
PB-T | ug/g | Lead Total | ICPM | Metals in Soil by CRC ICPMS | This analysis is carried out using procedures from CSR Analytical Method: "Strong Acid Leachable Metals (SALM) in Soil", BC Ministry of Environment, 26 June 2009, and procedures adapted from EPA Method 200.2. The sample is manually homogenized, dried at 60 degrees Celsius, sieved through a 63 um (230 mesh) sieve, and a representative subsample of the dry material is weighed. The sample is then digested at 95 degrees Celsius for 2 hours by block digester using concentrated nitric and hydrochloric acids. Instrumental analysis is by collision cell inductively coupled plasma - mass spectrometry (modifed from EPA Method 6020A). Method Limitation: This method is not a total digestion technique. It is a very strong acid digestion that is intended to dissolve those metals that may be environmentally available. By design, elements bound in silicate structures are not normally dissolved by this procedure as they are not usually mobile in the environment. | 0.1 | Y | Y | Y | 103 | 9/19/2017 | 9/24/2023 | 1/12/2024 11:16:52 AM | 2018-01-09 |  | JKROGH |  | MassPerMass | mg/kg | Y | (mass conc.) |  | Lead Total (mass conc.) |  | Metals - Total; | Metals - Total; EMS code: PB-T | Chemical | Total | LAB
PB-T | ug/m3 | Lead Total | F088 | Aliquot,HNO3/HCl dig, ICPMS,loading | Aliquot microwave digested with 3:1 HNO3/HCL. Analysis by ICPMS. For HiVol samples preliminary result are reported in mg/L and loading results are reported in ug/m3 |  | Y | Y | Y | 565 | 8/17/2000 10:46:00 AM | 12/14/2006 | 12/30/2006 1:01:37 AM | 2000-08-16 |  | NPEPPIN |  | AirConcentration | µg/m³ | Y | (air conc.) |  | Lead Total (air conc.) |  | Metals - Total; | Metals - Total; EMS code: PB-T | Chemical | Total | LAB
PB-T | ug/m3 | Lead Total | EQ40 | EQUIS (mg/m3) | EQUIS (mg/m3) | 0 | N | Y | Y | 3204 |  |  | 10/23/1996 12:01:37 AM | 1996-04-12 | 6/3/1998 6:51:38 PM | EMS |  | AirConcentration | µg/m³ | Y | (air conc.) |  | Lead Total (air conc.) |  | Metals - Total; | Metals - Total; EMS code: PB-T | Chemical | Total | LAB
PB-T | mg/L | Lead Total | ULTR | Total Metals in Water by ICPMS (Ultra) | Total Metals in Water by ICPMS (Ultra), EPA SW-846 3005A/6020A, MET-TOT-ULTRA-MS-VA This analysis is carried out using procedures adapted from "Standard Methods for the Examination of Water and Wastewater" published by the American Public Health Association, and with procedures adapted from "Test Methods for Evaluating Solid Waste" SW-846 published by the United States Environmental Protection Agency (EPA). The procedures may involve preliminary sample treatment by acid digestion, using either hotblock or microwave oven, or filtration (EPA Method 3005A). Instrumental analysis is by inductively coupled plasma - mass spectrometry (EPA Method 6020A). | 3e-05 | Y | Y | Y | 1 | 10/4/2017 | 10/4/2017 | 10/18/2017 3:31:11 PM | 2017-10-18 |  | JKROGH |  | Concentration | mg/L | Y | (fl. conc.) |  | Lead Total (fl. conc.) |  | Metals - Total; | Metals - Total; EMS code: PB-T | Chemical | Total | LAB
PB-T | mg/L | Lead Total | X285 | HNO3 Dig:ICP Analysis | HNO3 Dig:ICP Analysis | 0.005 | Y | Y | Y | 507 |  |  | 4/30/2004 1:59:45 PM | 1996-04-12 | 6/3/1998 6:51:41 PM | EMS |  | Concentration | mg/L | Y | (fl. conc.) |  | Lead Total (fl. conc.) |  | Metals - Total; | Metals - Total; EMS code: PB-T | Chemical | Total | LAB
PB-T | ug/m3 | Lead Total | X136 | HNO3/HCl/Perchlor,AA | HNO3/HCl/Perchlor;AA | 0 | Y | Y | Y | 755 |  |  | 2/19/2002 10:06:11 AM | 1996-04-12 | 6/3/1998 6:51:41 PM | EMS |  | AirConcentration | µg/m³ | Y | (air conc.) |  | Lead Total (air conc.) |  | Metals - Total; | Metals - Total; EMS code: PB-T | Chemical | Total | LAB
PB-T | mg/L | Lead Total | 0050 | Graphite Furnace AA | Graphite Furnace AA | 0.001 | Y | Y | Y | 43 | 8/28/2002 | 6/5/2003 | 6/9/2003 10:18:25 AM | 1996-04-12 | 6/3/1998 6:51:38 PM | EMS |  | Concentration | mg/L | Y | (fl. conc.) |  | Lead Total (fl. conc.) |  | Metals - Total; | Metals - Total; EMS code: PB-T | Chemical | Total | LAB
PB-T | ug/m3 | Lead Total | 5014 | N/P Dig,HGA Analysis | N/P Dig;HGA Analysis | 0.01 | Y | Y | Y | 18 |  |  | 10/24/1996 6:48:04 AM | 1996-04-12 | 6/3/1998 6:51:38 PM | EMS |  | AirConcentration | µg/m³ | Y | (air conc.) |  | Lead Total (air conc.) |  | Metals - Total; | Metals - Total; EMS code: PB-T | Chemical | Total | LAB
PB-T | ug/g | Lead Total | FS01 | D&G,Nitric/HCl:ICP | Samples are dried at 60 C, then either sieved to -100mesh (150 um) and rolled to homogenize, or finely ground with a mortar and pestle or a shatter box laboratory mill and rolled (depending on the submitter's requirements). The samples are then weighed into Teflon digestion liners and digested with HNO3 and HCl in a microwave oven. Samples are cooled, volumized, and allowed to settle overnight before being decanted. Analysis is then performed via Inductively Coupled Argon Plasma Atomic Emission Spectrometry (ICP-AES) (Simultaneous multi element analysis). In ICP Spectrometry, the liquid sample is pumped to the nebulizer, where it is convertd into an aerosol. the nebulizer carrier gas flow than transports the wet aerosol past the impact bead, where the large droplets condense out and flow to the waste container. The remaining fine aerosol is carried into the plasma. There, the high temperature desolvates the aerosol and dissociates the sample material into individual atoms and ions, which are excited to emit light at wavelengths characteristic of the elements present. A diffraction grating in the spectrometer diffracts the various wavelengths of light through exit slits to the photo-multiplier tubes, which convert the light energy to electronic signals. | 8 | Y | Y | Y | 2642 | 8/30/1996 1:52:00 PM | 10/10/2007 | 5/24/2008 1:15:03 AM | 1996-11-20 | 6/3/1998 6:51:46 PM | MPAK |  | MassPerMass | mg/kg | Y | (mass conc.) |  | Lead Total (mass conc.) |  | Metals - Total; | Metals - Total; EMS code: PB-T | Chemical | Total | LAB
PB-T | ug/g wet | Lead Total | X552 | Metals in Biota (wet), total, acid digestion, ICP-SFMS | One gram of wet, homogenized biota is weighed in a microwave TFM vessel, and 8mL of Nitric Acid and 1mL of Hydrogen Peroxide are added. The sample is digested at 200 degrees C for 15 minutes in a closed vessel high pressure microwave oven. The sample is cooled, the volume adjusted to 100mL and introduced to an inductively coupled argon plasma-sector field mass spectrometer (ICP-SFMS) by pneumatic nebulization. Each respective element is measured at a specific mass to charge ratio, m/z value expressed in amu (atomic mass units), and at a specified instrument resolution. LR: Low Resolution ~300, MR: Mid Resolution ~4000. Ref: 02-2704, National Laboratory for Environmental Testing, CCIW, Environment Canada, Burlington. Requested April 2003. | 0.001 | Y | Y | Y | 37 | 8/9/2016 | 10/16/2023 | 11/28/2023 2:21:33 PM | 2004-01-19 | 4/19/2021 1:22:00 PM | UMULLER | EMS | MassPerMass | µg/g | Y | (mass conc.) | (wet) | Lead Total (wet) (mass conc.) |  | Metals - Total; | Metals - Total; EMS code: PB-T | Chemical | Total | LAB
PB-T | ug/g | Lead Total | 0195 | Hom, N/P Dig, HGA Anal | Hom; N/P Dig; HGA Anal | 1 | Y | Y | Y | 504 |  |  | 10/24/1996 1:45:56 PM | 1996-04-12 | 6/3/1998 6:51:38 PM | EMS |  | MassPerMass | mg/kg | Y | (mass conc.) |  | Lead Total (mass conc.) |  | Metals - Total; | Metals - Total; EMS code: PB-T | Chemical | Total | LAB
PB-T | ppm (S) | Lead Total | MS61 | ALS Minerals ME-MS61 | ALS Minerals ME-MS61. A prepared sample is digested with HNO3; HF; and HClO4 in a graphite heating block to near dryness. After cooling; the sample is boiled in diluted HCl. Bulk to volume with deionized water; mixed and analyzed by ICP-AES and ICP-MS for all elements. | 0.5 | Y | Y | Y | 0 |  |  |  | 2024-01-29 |  | EMS |  | DimensionlessRatio | ppm | N | (ratio) |  | Lead Total (ratio) |  | Metals - Total; | Metals - Total; EMS code: PB-T | Chemical | Total | LAB
PB-T | mg/L | Lead Total | X170 | ICP | This code is a catch-all for a large variety of ICP methods including water and also solids digested into solutions as an intermediate result. Different procedures will have greatly different detection limits. (a) For air particulates, preliminary results are reported in mg/L and loading results as mg/dm2/d. (b) For Hi Vol samples preliminiary results are same as reported in mg/L and loading results as ug/m3. (c) Note that for Rain samples there is no filtration nor digestion steps. | 0.02 | Y | Y | Y | 11110 | 3/28/1996 6:03:00 PM | 11/7/2018 | 7/25/2024 11:22:07 AM | 1996-04-12 | 6/3/1998 6:51:38 PM | EMS |  | Concentration | mg/L | Y | (fl. conc.) |  | Lead Total (fl. conc.) |  | Metals - Total; | Metals - Total; EMS code: PB-T | Chemical | Total | LAB
PB-T | mg/dm2/d | Lead Total | CS15 | System Calculation | System Calculation (mg/dm2/d) |  | N | Y | N | 44 |  |  | 11/18/2000 1:30:25 AM | 1996-04-12 | 6/3/1998 6:51:44 PM | EMS |  | MassFlux | mg/dm²/d | Y | (mass flux) |  | Lead Total (mass flux) |  | Metals - Total; | Metals - Total; EMS code: PB-T | Chemical | Total | LAB
PB-T | ug/g | Lead Total | PB02 | PBM,HNO3/HCl dig,dry60C,100mesh, | Performance Based Method. Samples are dried at 60 C and sieved using a .149 mm (100 mesh) sieve. Solid samples are digested with the combination of nitric and hydrochloric acids with a 1:1 ratio of acids. For environmental purposes, strong acid leaches are used routinely since they have the advantage of speed, low cost, and superio reproducibility. All metals trapped within the refractory matrix are unlikely to have environmental consequences. Instrumental analysis of extracts produced using this digestion method can be performed by a variety of techniques. |  | Y | Y | Y | 52 | 5/24/2001 3:42:00 PM | 11/26/2001 10:23:00 AM | 5/26/2003 2:28:17 PM | 2001-10-11 |  | NPEPPIN |  | MassPerMass | mg/kg | Y | (mass conc.) |  | Lead Total (mass conc.) |  | Metals - Total; | Metals - Total; EMS code: PB-T | Chemical | Total | LAB
PB-T | ug/g | Lead Total | FB04 | Freeze Dry,Hom,Acid Dig,ICP | Sample moisture is determined by weighing samples before and after freeze-drying. Freeze-dried samples are ground in a Braun blender before being weighed into Teflon PFA liners and digested with concentrated HNO3, concentrated HCl, and 50% H2O2 in a microwave oven; which results in a complete sample dissolution. Samples are collected and volumized with Type I water. Analysis is performed via Inductively Coupled Argon Plasma Atomic Emission Spectrometry (ICP-AES) (simultaneous multielement analysis), Graphite Furnace Atomic Absorption Spectrometry (GF-AAS) (low level Cd and Pb analysis), Cold Vapor Atomic Fluorescence Spectrometry (CV-AFS) (Mercury analysis) and/or Flame Atomic Absorption Spectrometry (F-AAS) (K analysis). | 4 | Y | Y | Y | 241 | 8/22/1996 6:11:00 PM | 2/1/2002 2:30:00 PM | 6/23/2003 1:15:33 PM | 1997-02-27 | 6/3/1998 6:51:46 PM | MPAK |  | MassPerMass | mg/kg | Y | (mass conc.) |  | Lead Total (mass conc.) |  | Metals - Total; | Metals - Total; EMS code: PB-T | Chemical | Total | LAB
PB-T | mg/L | Lead Total | METN | MET-T-NP-CCMS-VA | Ultra trace metals in water are analyzed by CRC ICPMS, based on US EPA Method 6020B (July 2014). The detection limits provided can only be met for undigested samples. This procedure is intended for colorless, non-turbid, acid-preserved water samples (i.e. pristine water samples), having turbidity < 1 NTU and no odor. Where turbidity exceeds 1 NTU, and/or the sample is colored and has an odor, results may be biased low compared to true Total Metals concentrations. ALS recommends that turbidity analysis be requested on samples submitted for this test to aid with interpretation of results. Where turbidity is <1NTU, undigested metals are equivalent to total metals concentrations. | 5e-06 | Y | Y | Y | 1059 | 4/3/2012 | 4/4/2024 | 4/29/2024 4:53:06 PM | 2021-04-14 |  | RDGIBSON |  | Concentration | mg/L | Y | (fl. conc.) |  | Lead Total (fl. conc.) |  | Metals - Total; | Metals - Total; EMS code: PB-T | Chemical | Total | LAB
PB-T | ug/g wet | Lead Total | BPSY | MET-WET-BPSY-CCMS-VA | Metals in Animal Tissue and Vegetation (Biota) - Prescriptive. EPA 200.3/6020B This method is conducted following British Columbia Lab Manual method "Metals in Animal Tissue and Vegetation (Biota) - Prescriptive". Tissue samples are digested with nitric and hydrochloric acids; in combination with addition of hydrogen peroxide. Instrumental analysis is by collision cell inductively coupled plasma - mass spectrometry (modified from EPA Method 6020B). | 0.01 | Y | Y | Y | 2 | 11/7/2019 | 11/7/2019 | 6/10/2021 1:57:10 PM | 2021-06-04 |  | EMS |  | MassPerMass | µg/g | Y | (mass conc.) | (wet) | Lead Total (wet) (mass conc.) |  | Metals - Total; | Metals - Total; EMS code: PB-T | Chemical | Total | LAB
PB-T | mg/dm2/d | Lead Total | CAL1 | Calculation (mg/dm2/d) | Calculation (mg/dm2/d) | 0.0003 | Y | Y | Y | 3046 |  |  | 7/3/2024 12:59:27 PM | 1996-04-12 | 6/3/1998 6:51:42 PM | EMS |  | MassFlux | mg/dm²/d | Y | (mass flux) |  | Lead Total (mass flux) |  | Metals - Total; | Metals - Total; EMS code: PB-T | Chemical | Total | LAB
PB-T | ug/g | Lead Total | X208 | Hom,Nitr/Perchl,Grph Fur | Hom;Nitr/Perchl;Grph Fur | 0.05 | Y | Y | Y | 36 |  |  | 10/24/1996 2:26:14 PM | 1996-04-12 | 6/3/1998 6:51:38 PM | EMS |  | MassPerMass | mg/kg | Y | (mass conc.) |  | Lead Total (mass conc.) |  | Metals - Total; | Metals - Total; EMS code: PB-T | Chemical | Total | LAB
PB-T | ug/g | Lead Total | X226 | Hom,HNO3/HCl,AA | Hom;HNO3/HCl;AA | 2.5 | Y | Y | Y | 183 |  |  | 3/19/1998 11:44:20 AM | 1996-04-12 | 6/3/1998 6:51:38 PM | EMS |  | MassPerMass | mg/kg | Y | (mass conc.) |  | Lead Total (mass conc.) |  | Metals - Total; | Metals - Total; EMS code: PB-T | Chemical | Total | LAB
PB-T | ug/m3 | Lead Total | 5012 | N/P Dig:HGA Analysis | N/P Dig:HGA Analysis | 0.1 | N | Y | Y | 801 |  |  | 10/24/1996 1:49:24 PM | 1996-04-12 | 6/3/1998 6:51:38 PM | EMS |  | AirConcentration | µg/m³ | Y | (air conc.) |  | Lead Total (air conc.) |  | Metals - Total; | Metals - Total; EMS code: PB-T | Chemical | Total | LAB
PB-T | mg/L | Lead Total | 0210 | N/P Dig:HGA Analysis | N/P Dig:HGA Analysis | 0.01 | Y | Y | Y | 14 |  |  | 10/24/1996 2:10:08 PM | 1996-04-12 | 6/3/1998 6:51:38 PM | EMS |  | Concentration | mg/L | Y | (fl. conc.) |  | Lead Total (fl. conc.) |  | Metals - Total; | Metals - Total; EMS code: PB-T | Chemical | Total | LAB
PB-T | mg/L | Lead Total | F082 | Aliquot,HNO3/HCL dig,ICPMS | Aliquot microwave digested with 3:1 HNO3/HCL. Analysis by ICPMS. For HiVol samples preliminary result are reported in mg/L and loading results are reported in ug/m3 | 0.0005 | Y | Y | Y | 28020 | 1/24/2011 | 8/23/2024 | 8/26/2024 4:53:05 PM | 2000-06-13 | 8/15/2019 11:10:32 AM | NPEPPIN | KATFRASE | Concentration | mg/L | Y | (fl. conc.) |  | Lead Total (fl. conc.) |  | Metals - Total; | Metals - Total; EMS code: PB-T | Chemical | Total | LAB
PB-T | ug/g | Lead Total | 0192 | Hom, Nitric/Perchlor,ICP | Sample already in a dry state so does not require drying before homogenization and nitric perchlorid digestion followed by ICP analysis. | 10 | Y | Y | Y | 357 | 2/13/2002 | 5/29/2002 | 1/14/2004 1:00:44 AM | 1996-04-12 | 6/3/1998 6:51:38 PM | EMS |  | MassPerMass | mg/kg | Y | (mass conc.) |  | Lead Total (mass conc.) |  | Metals - Total; | Metals - Total; EMS code: PB-T | Chemical | Total | LAB
PB-T | ug/g | Lead Total | X230 | Hom,HNO3/HCl,ICP | The sample is dried, ground to pass 100 mesh, and digested with a nitric-hydrochloric acid mixture to solubilize the solid matter and to remove the organic material by oxidation and volatilization. The completed digests are submitted to metals instrumentation personnel for metals analysis by Inductively Coupled Atomic Emission Spectrophotometry (ICP-AES), or hydride analysis by atomic absorption. | 1 | Y | Y | Y | 9 | 12/30/2005 | 12/30/2005 | 2/3/2006 1:13:13 AM | 1998-11-05 |  | NPEPPIN |  | MassPerMass | mg/kg | Y | (mass conc.) |  | Lead Total (mass conc.) |  | Metals - Total; | Metals - Total; EMS code: PB-T | Chemical | Total | LAB
PB-T | mg/L | Lead Total | X078 | Aqua Reg.Dig,Hydr.gen,AA | Aqua Reg.Dig;Hydr.gen;AA | 0.001 | Y | Y | Y | 257 |  |  | 10/29/2001 11:34:43 AM | 1996-04-12 | 6/3/1998 6:51:38 PM | EMS |  | Concentration | mg/L | Y | (fl. conc.) |  | Lead Total (fl. conc.) |  | Metals - Total; | Metals - Total; EMS code: PB-T | Chemical | Total | LAB
PB-T | mg/L | Lead Total | X145 | Dig, ICP | Dig; ICP | 0.001 | Y | Y | Y | 13 |  |  | 10/24/1996 2:38:05 PM | 1996-04-12 | 6/3/1998 6:51:38 PM | EMS |  | Concentration | mg/L | Y | (fl. conc.) |  | Lead Total (fl. conc.) |  | Metals - Total; | Metals - Total; EMS code: PB-T | Chemical | Total | LAB
PB-T | ug/g wet | Lead Total | X560 | Metals in Biota (wet), total, acid digestion, ICP-MS | Five grams of wet, homogenized biota is weighed in a centrifuge tube, and 2.5 mL of Nitric Acid and 2.5 mL of Hydrochloric acid are added. The sample is digested at room temperature for one hour and then digested at 90-95 degrees C in a water bath or heat block for one more hour. The sample is cooled and the volume is made to 50 mL with DI. The digestate is then diluted at ten fold and introduced to a collision/reaction cell inductively coupled argon plasma mass spectrometer (CRC ICPMS) by pneumatic nebulization. Each respective element is measured at a specific mass to charge ratio, m/z value expressed in amu (atomic mass units). - Maxxam Analytics (BC). Added to EMS Dec 10, 2008. | 0.01 | Y | Y | Y | 541 | 11/15/2008 4:00:00 AM | 10/30/2023 | 11/28/2023 2:21:33 PM | 2008-12-11 | 3/11/2015 4:17:10 PM | RDGIBSON | EMS | MassPerMass | µg/g | Y | (mass conc.) | (wet) | Lead Total (wet) (mass conc.) |  | Metals - Total; | Metals - Total; EMS code: PB-T | Chemical | Total | LAB
PB-T | ug/m3 | Lead Total | 5312 | Tef:N/P Dig:ICP Analysis | A nitric/perchloric acid digestion if used to bring the metals into solution prior to ICP analysis. For Hi-Vol samples preliminary results are reported as mg/L and 24 hour loading results as ug/m3. | 0.3 | Y | Y | Y | 1719 | 11/7/1997 | 3/29/2000 | 9/28/2000 1:34:43 PM | 1996-04-12 | 6/3/1998 6:51:38 PM | EMS |  | AirConcentration | µg/m³ | Y | (air conc.) |  | Lead Total (air conc.) |  | Metals - Total; | Metals - Total; EMS code: PB-T | Chemical | Total | LAB
PB-T | mg/L | Lead Total | F068 | HNO3dig,TeflonTube/Autoclave,ICP-MS | This method is applicable to determiniation of dissolved, extractable, and total recoverable metals in ground waters, surface waters, drinking waters, and wastewaters. Dissolved, extractable, and total recoverable sample fractions are acidified to pH <1.5 with analytical grade nitric acid (HNO3) to prevent the metals from adsorbing to the container material or precipitating. Extractable and total recoverable water samples are collected in 250 ml acid washed polyethylene bottles and are preserved with Nitric acid (0.5 ml acid for 100 mLs of sample). All preserved samples are stored at room temperature and can be held for up to 6 months. Unfiltered, un-preserved samples are stored at 4C for no more than 30 days. For dissolved and extractable water fractions, no sample preparation is necessary. Samples are simply poured in autosampler tubes (i.e. 15 ml centrifuge tubes), and placed on the autosampler and analyzed. Samples for total recoverable metals are digested with Nitric acid in special Teflon tubes in an autoclave prior to analyses. | 1e-05 | Y | Y | Y | 4078 | 6/1/1999 1:33:00 PM | 7/4/2008 9:48:00 AM | 7/30/2024 6:32:12 AM | 1999-06-28 |  | NPEPPIN |  | Concentration | mg/L | Y | (fl. conc.) |  | Lead Total (fl. conc.) |  | Metals - Total; | Metals - Total; EMS code: PB-T | Chemical | Total | LAB
PB-T | mg/L | Lead Total | F108 | SPR,IDA Chelation,ICPMS | SPR;IDA Chelation;ICPMS |  | Y | Y | Y | 4 | 11/30/2000 10:32:00 AM | 11/30/2000 10:32:00 AM | 9/10/2002 9:29:49 AM | 2001-03-08 |  | NPEPPIN |  | Concentration | mg/L | Y | (fl. conc.) |  | Lead Total (fl. conc.) |  | Metals - Total; | Metals - Total; EMS code: PB-T | Chemical | Total | LAB
PB-T | mg/L | Lead Total | F033 | Acid Dig,Centrifuge,ICPMS | For the determination of total recoverable elements in water or waste water, take 100 ml aliquot from a well mixed, acid preserved, sample containing not more than 0.25% (w/v) total solids and transfer to a 250 ml Griffin beaker. (If total solids are greater than 0.25% reduce the size of the aliquot by a proportionate amount). Add 1 ml of conc. nitric acid and 0.5 ml conc. hydrochloric acid. Heat on a hot plate until the volume has been reduced to 15-20ml, ensuring that the sample does not boil. Allow to cool and transfer to a 50 ml volumetric flask. Dilute to volume with ASTM type I water. Centrifuge the sample or allow to stand overnight to separate insoluble material. Prior analysis, pipette 20 ml into a 50 ml volumetric flask and dilute to volume with ASTM type I water. Analysis should be performed within seventy two hours of the completed preparation. Note that different laboratories may have very different field preservation, digestion procedures and detection limits for this method, so these should be documented. | 0.0001 | Y | Y | Y | 38658 | 10/15/1996 9:27:00 AM | 8/22/2024 3:24:00 AM | 8/25/2024 9:47:36 PM | 1997-01-16 | 6/3/1998 6:51:46 PM | MPAK |  | Concentration | mg/L | Y | (fl. conc.) |  | Lead Total (fl. conc.) |  | Metals - Total; | Metals - Total; EMS code: PB-T | Chemical | Total | LAB
PB-T | ug/g | Lead Total | FS82 | PBM,dry60C,10mesh,SALM-ICP-MS | PBM;dry60C;10mesh;SALM-ICP-MS | 0.1 | Y | Y | Y | 3506 | 5/21/2002 | 7/24/2024 | 7/30/2024 10:00:14 AM | 2002-11-12 |  | UMULLER |  | MassPerMass | mg/kg | Y | (mass conc.) |  | Lead Total (mass conc.) |  | Metals - Total; | Metals - Total; EMS code: PB-T | Chemical | Total | LAB
PB-T | mg/L | Lead Total | 0100 | Dig:Ext.(APDC):AA Anal. | Dig:Ext.(APDC):AA Anal. | 0.001 | Y | Y | Y | 1063 |  |  | 6/16/2000 4:19:25 PM | 1996-04-12 | 6/3/1998 6:51:38 PM | EMS |  | Concentration | mg/L | Y | (fl. conc.) |  | Lead Total (fl. conc.) |  | Metals - Total; | Metals - Total; EMS code: PB-T | Chemical | Total | LAB
PB-T | ug/m3 | Lead Total | 6040 | Aqua Regia Dig:ICP | Aqua regia digestion is used to bring the metals into solution prior to ICP analysis. For Hi-Vol samples, preliminary results are reported as mg/L and24 hr loadings as ug/m3 | 0.3 | Y | Y | Y | 437 | 2/13/2002 | 2/9/2005 | 2/23/2005 1:13:47 AM | 1998-05-05 | 6/3/1998 6:51:48 PM | NPEPPIN |  | AirConcentration | µg/m³ | Y | (air conc.) |  | Lead Total (air conc.) |  | Metals - Total; | Metals - Total; EMS code: PB-T | Chemical | Total | LAB
PB-T | ug/g | Lead Total | X551 | Metals in Biota (dry), total, acid digestion, ICP-SFMS | 0.1 gram of dried, homogenized biota is weighed in a microwave TFM vessel, and 8mL of Nitric Acid and 1mL of Hydrogen Peroxide are added. The sample is digested at 200 degrees C for 15 minutes in a closed vessel high pressure microwave oven. The sample is cooled, the volume adjusted to 100mL and introduced to an iductively coupled argon plasma-sector field mass spectrometer (ICP-SFMS) by pneumatic nebulization. Each respective element is measured at a specific mass to charge ratio, m/z value expressed in amu (atomic mass units), and at a specified instrument resolution. LR: Low Resolution ~300, MR: Mid Resolution ~4000 and HR: Hich Resolution ~10000. Ref: Method number 02-2705, National Laboratory for Environmental Testing, CCIW, Environment Canada, Burlington. Requested April 2003. | 0.01 | Y | Y | Y | 49 | 8/11/2015 | 10/30/2023 | 11/28/2023 2:21:33 PM | 2004-01-19 |  | UMULLER |  | MassPerMass | mg/kg | Y | (mass conc.) |  | Lead Total (mass conc.) |  | Metals - Total; | Metals - Total; EMS code: PB-T | Chemical | Total | LAB
PB-T | mg/L | Lead Total | F048 | Acid dig,conc'n,nebul.,ICP-OES | This method provides an automated procedure for the Inductively Coupled Argon Plasma Optical Emission Spectrophotometric (ICP-OES) determination and quantitation of trace amounts of metals in surface, ground and precipitation waters after a manual acid digestion and concentration procedure and ultrasonic nebulisation. Organically-bound metals and metal complexes are converted to metal ions by digestion with nitric acid [method details says HCl]. To increase the signal detected by the ICP, the samples are concentrated 10 times. The liquid sample is pumped onto the piezoelectric transducer, where the ultrasonic action efficiently converts it to a fine, dense aerosol. The nebulizer carrier gas flow then transports the wet aerosol through a heated U-tube, where the solvent is vaporized. Subsequently, most of the solvent vapour is removed from the aerosol stream by passing the stream through a chilled condenser. The result is a dry, analyte-laden aerosol which is introduced to the plasma. There, the high temperature desolvates the aerosol and dissociates the sample material into individual atoms and ions, which are excited to emit light at wave lengths characteristic of the elements present. Dissolved samples are field filtered through 0.45 micron cellulose acetate filter prior to acid preservation. Samples for extractable analysis are acidified with 2 mL concentrated HNO3 per litre and allowed to stand a minimum of 24 hours before analysis. Clients providing these samples are advised to preserve samples at the collection site. | 0.0002 | Y | Y | Y | 6159 | 7/4/1997 2:36:00 PM | 4/23/2003 | 9/11/2019 8:46:38 AM | 1997-10-20 | 6/3/1998 6:51:47 PM | NPEPPIN |  | Concentration | mg/L | Y | (fl. conc.) |  | Lead Total (fl. conc.) |  | Metals - Total; | Metals - Total; EMS code: PB-T | Chemical | Total | LAB
PB-T | ug/g | Lead Total | FS89 | Homo,HNO3 dig,ICP-AES | Homo;HNO3 dig;ICP-AES | 0.4 | Y | Y | Y | 4 | 1/5/2001 | 8/20/2001 12:00:00 PM | 10/23/2002 1:42:50 PM | 2001-03-08 |  | NPEPPIN |  | MassPerMass | mg/kg | Y | (mass conc.) |  | Lead Total (mass conc.) |  | Metals - Total; | Metals - Total; EMS code: PB-T | Chemical | Total | LAB
PB-T | mg/L | Lead Total | 5078 | N/P Digestion/ ICP Anal. | N/P Digestion/ ICP Anal. | 0.1 | N | Y | Y | 18 |  |  | 10/24/1996 2:10:29 PM | 1996-04-12 | 6/3/1998 6:51:41 PM | EMS |  | Concentration | mg/L | Y | (fl. conc.) |  | Lead Total (fl. conc.) |  | Metals - Total; | Metals - Total; EMS code: PB-T | Chemical | Total | LAB
PB-T | ug/g | Lead Total | EQ02 | EQUIS (ug/g) | EQUIS (ug/g) | 1 | N | Y | Y | 900 |  |  | 10/23/1996 12:08:28 AM | 1996-04-12 | 6/3/1998 6:51:38 PM | EMS |  | MassPerMass | mg/kg | Y | (mass conc.) |  | Lead Total (mass conc.) |  | Metals - Total; | Metals - Total; EMS code: PB-T | Chemical | Total | LAB
PB-T | mg/L | Lead Total | F024 | HNO3/HCl,GFAA | Dissolved and extractable water samples are ready for analysis once collected. Total samples are autoclaved with HNO3 and hydrochloric acid (HCl) in capped Teflon vials. Analysis is performed via Inductively Coupled Argon Plasma Atomic Emission Spectrometry (ICP-AES) (simultaneous multielement analysis), Graphite Furnace Atomic Absorption Spectrometry (GF-AAS) (low level analysis) and/or Flame Atomic Absorption Spectrometry (F-AAS). | 0.0005 | Y | Y | Y | 613 | 8/6/1996 7:16:00 PM | 9/12/2003 3:16:00 PM | 12/1/2003 5:30:23 PM | 1996-12-11 | 6/3/1998 6:51:46 PM | MPAK |  | Concentration | mg/L | Y | (fl. conc.) |  | Lead Total (fl. conc.) |  | Metals - Total; | Metals - Total; EMS code: PB-T | Chemical | Total | LAB
PB-T | ug/g | Lead Total | 0031 | ICP | ICP | 10 | Y | Y | Y | 98 | 7/26/2005 7:04:00 AM | 7/18/2012 7:03:00 AM | 7/26/2012 12:28:00 PM | 1996-10-11 | 6/3/1998 6:51:46 PM | NPEPPIN |  | MassPerMass | mg/kg | Y | (mass conc.) |  | Lead Total (mass conc.) |  | Metals - Total; | Metals - Total; EMS code: PB-T | Chemical | Total | LAB
PB-T | ug/g | Lead Total | 0230 | D&G,N/P Dig:HGA Analysis | D&G;N/P Dig:HGA Analysis | 1 | Y | Y | Y | 19 |  |  | 10/24/1996 3:06:16 PM | 1996-04-12 | 6/3/1998 6:51:38 PM | EMS |  | MassPerMass | mg/kg | Y | (mass conc.) |  | Lead Total (mass conc.) |  | Metals - Total; | Metals - Total; EMS code: PB-T | Chemical | Total | LAB
PB-T | mg/L | Lead Total | 0105 | 5%AR Dig:APDC Ext:AA Anl | 5%AR Dig:APDC Ext:AA Anl | 0.001 | Y | Y | Y | 472 |  |  | 10/24/1996 2:13:46 PM | 1996-04-12 | 6/3/1998 6:51:38 PM | EMS |  | Concentration | mg/L | Y | (fl. conc.) |  | Lead Total (fl. conc.) |  | Metals - Total; | Metals - Total; EMS code: PB-T | Chemical | Total | LAB
PB-T | mg/L | Lead Total | X545 | FA,HNO3,ICPMS | Analytes nebulized followed by ion entrained in plasma gas and extracted by means of a differentially pumped vacuum interface into a mass spectrometer | 0.001 | Y | Y | Y | 30 |  |  | 10/30/2014 8:31:18 AM | 2001-09-18 | 9/18/2001 1:36:23 PM | NPEPPIN | NPEPPIN | Concentration | mg/L | Y | (fl. conc.) |  | Lead Total (fl. conc.) |  | Metals - Total; | Metals - Total; EMS code: PB-T | Chemical | Total | LAB
PB-T | ug/m3 | Lead Total | X180 | LVOL:Dig,AA | LVOL:Dig;AA | 0 | Y | Y | Y | 1832 |  |  | 3/9/1998 10:35:51 AM | 1996-04-12 | 6/3/1998 6:51:41 PM | EMS |  | AirConcentration | µg/m³ | Y | (air conc.) |  | Lead Total (air conc.) |  | Metals - Total; | Metals - Total; EMS code: PB-T | Chemical | Total | LAB
PB-T | mg/L | Lead Total | F121 | Acid,ICPMS | Acidified, no digestion, ICPMS | 1e-05 | Y | Y | Y | 8460 | 3/12/2010 7:57:00 PM | 8/15/2024 9:10:00 PM | 8/26/2024 8:06:03 AM | 2002-07-22 | 7/16/2012 3:28:49 PM | UMULLER | RDGIBSON | Concentration | mg/L | Y | (fl. conc.) |  | Lead Total (fl. conc.) |  | Metals - Total; | Metals - Total; EMS code: PB-T | Chemical | Total | LAB
PB-T | mg/L | Lead Total | 0200 | 20%A.R.Dig:AA Anal. | 20%A.R.Dig:AA Anal. | 0.1 | Y | Y | Y | 10 |  |  | 10/24/1996 4:16:44 AM | 1996-04-12 | 6/3/1998 6:51:38 PM | EMS |  | Concentration | mg/L | Y | (fl. conc.) |  | Lead Total (fl. conc.) |  | Metals - Total; | Metals - Total; EMS code: PB-T | Chemical | Total | LAB
PB-T | mg/L | Lead Total | X561 | Total metals: FA, HNO3, HR-ICPMS (No Prep)Prep) | Ultra trace metals in water are analyzed by high resolution inductively coupled plasma mass spectrometry (HR-ICPMS), modified from US EPA 200.8, (Revision 5.5). The detection limits provided can only be met for undigested samples. This procedure is intended for colorless, non-turbid, acid-preserved water samples (i.e. pristine water samples), having turbidity < 1 NTU and no odor. Where turbidity exceeds 1 NTU, and/or the sample is colored and has an odor, results may be biased low compared to true Total Metals concentrations. ALS recommends that turbidity analysis be requested on samples submitted for this test to aid with interpretation of results. | 5e-06 | Y | Y | Y | 2756 | 6/8/2015 | 10/5/2023 | 1/31/2024 2:37:41 PM | 2015-10-22 |  | RDGIBSON |  | Concentration | mg/L | Y | (fl. conc.) |  | Lead Total (fl. conc.) |  | Metals - Total; | Metals - Total; EMS code: PB-T | Chemical | Total | LAB
PB-T | ppm (S) | Lead Total | 6042 | Aqua Regia by ICPMS | These were sediment/soils digested with Aqua-Regia (Nitric Acid & Hydrochloric Acids) and then analysis by ICPMS. Results would be reported in ppm s. | 0.005 | Y | Y | Y | 71 | 5/14/2022 | 7/27/2024 | 7/29/2024 9:13:08 AM | 2022-06-28 |  | EMS |  | DimensionlessRatio | ppm | N | (ratio) |  | Lead Total (ratio) |  | Metals - Total; | Metals - Total; EMS code: PB-T | Chemical | Total | LAB
PB-T | mg/L | Lead Total | EQ01 | EQUIS (mg/L) | EQUIS (mg/L) | 0.001 | N | Y | Y | 7873 |  |  | 10/24/1996 3:16:18 PM | 1996-04-12 | 6/3/1998 6:51:38 PM | EMS |  | Concentration | mg/L | Y | (fl. conc.) |  | Lead Total (fl. conc.) |  | Metals - Total; | Metals - Total; EMS code: PB-T | Chemical | Total | LAB
PB-T | mg/L | Lead Total | 0061 | HNO3 Dig: GFAA Analysis | HNO3 Dig: GFAA Analysis | 0.003 | Y | Y | Y | 1847 | 8/11/1993 | 6/11/1998 2:00:00 PM | 3/8/2001 9:35:18 AM | 1996-04-12 | 6/3/1998 6:51:38 PM | EMS |  | Concentration | mg/L | Y | (fl. conc.) |  | Lead Total (fl. conc.) |  | Metals - Total; | Metals - Total; EMS code: PB-T | Chemical | Total | LAB
PB-T | ug/g wet | Lead Total | MTIW | MET-WET-CCMS-MID-VA | Metals in Tissue by CRC ICPMS (WET) Methodology: This method is conducted following British Columbia Lab Manual method "Metals in Animal Tissue and Vegetation (Biota) - Prescriptive". Tissue samples are homogenized and sub-sampled prior to hotblock digestion with nitric and hydrochloric acids, in combination with addition of hydrogen peroxide. Instrumental analysis is by collision cell inductively coupled plasma - mass spectrometry (modified from EPA Method 6020A). Method Limitation: This method employs a strong acid/peroxide digestion, and is intended to provide a conservative estimate of bio-available metals. Near complete recoveries are achieved for most toxicologically important metals, but elements associated with recalcitrant minerals may be only partially recovered. | 0.01 | Y | Y | Y | 24 | 11/25/2018 | 12/6/2018 | 7/21/2021 2:55:08 PM | 2018-12-24 |  | EMS |  | MassPerMass | µg/g | Y | (mass conc.) | (wet) | Lead Total (wet) (mass conc.) |  | Metals - Total; | Metals - Total; EMS code: PB-T | Chemical | Total | LAB
PB-T | mg/dm2/d | Lead Total | DUST | HNO3 dig.(microwave or hot block),ICPMS | HNO3 dig.(microwave or hot block);ICPMS |  | Y | Y | Y | 855 | 7/19/2016 1:11:00 PM | 7/31/2024 | 7/31/2024 11:23:12 AM | 2008-02-07 |  | RDGIBSON |  | MassFlux | mg/dm²/d | Y | (mass flux) |  | Lead Total (mass flux) |  | Metals - Total; | Metals - Total; EMS code: PB-T | Chemical | Total | LAB
PB-T | ppm (S) | Lead Total | MS41 | ALS Minerals ME-MS41 | ALS Minerals ME-MS41. A prepared sample is digested with aqua regia in a graphite heating block. After cooling; the resulting solution is diluted to with deionized water; mixed and analyzed by ICP-AES. Samples are then analyzed by ICP-MS for the remaining suite of elements. | 0.2 | Y | Y | Y | 0 |  |  |  | 2024-01-29 |  | EMS |  | DimensionlessRatio | ppm | N | (ratio) |  | Lead Total (ratio) |  | Metals - Total; | Metals - Total; EMS code: PB-T | Chemical | Total | LAB
PB-T | mg/L | Lead Total | F038 | HNO3/HCl dig,ICP-AES | Dissolved and extractable water samples are ready for analysis once collected. Total samples are autoclaved with HNO3 and hydrochloric acid (HCI) in capped Teflon vials. Analysis is performed via Inductively Coupled Argon Plasma Atomic Emission Spectrometry (ICP-AES)(simultaneous multielement analysis), Graphite Furnace Atomic Absorption Spectrometry (GF-AAS) (low level analysis) and/or Flame Atomic Absorption Spectrometry (F-AAS). | 0.06 | Y | Y | Y | 5426 | 8/1/1996 6:04:00 PM | 5/6/2008 | 9/11/2019 8:47:42 AM | 1997-05-08 | 6/3/1998 6:51:46 PM | MPAK |  | Concentration | mg/L | Y | (fl. conc.) |  | Lead Total (fl. conc.) |  | Metals - Total; | Metals - Total; EMS code: PB-T | Chemical | Total | LAB
PB-T | mg/L | Lead Total | X095 | FF,FA,HNO3,ICP | FF;FA;HNO3;ICP | 0.001 | Y | Y | Y | 19 |  |  | 10/24/1996 12:17:00 PM | 1996-04-12 | 6/3/1998 6:51:41 PM | EMS |  | Concentration | mg/L | Y | (fl. conc.) |  | Lead Total (fl. conc.) |  | Metals - Total; | Metals - Total; EMS code: PB-T | Chemical | Total | LAB
PB-T | ug/m3 | Lead Total | 5097 | N/P Dig(Filt+Wash),ICP | N/P Dig(Filt+Wash);ICP | 0.001 | Y | Y | Y | 10963 |  |  | 7/3/2024 12:59:27 PM | 1996-04-12 | 6/3/1998 6:51:38 PM | EMS |  | AirConcentration | µg/m³ | Y | (air conc.) |  | Lead Total (air conc.) |  | Metals - Total; | Metals - Total; EMS code: PB-T | Chemical | Total | LAB
PB-T | ug/g wet | Lead Total | X215 | Hom,Nitr/Perchl,Grph Fur | Hom;Nitr/Perchl;Grph Fur | 0.01 | Y | Y | Y | 74 |  |  | 10/24/1996 10:01:27 AM | 1996-04-12 | 6/3/1998 6:51:42 PM | EMS |  | MassPerMass | µg/g | Y | (mass conc.) | (wet) | Lead Total (wet) (mass conc.) |  | Metals - Total; | Metals - Total; EMS code: PB-T | Chemical | Total | LAB
PB-T | mg/L | Lead Total | X512 | HNO3,pHto7.5,MIK extr'n,AA | HNO3 Digestion;pH to 7.5; MIK solvent extraction;AA | 0.001 | Y | Y | Y | 2979 |  |  | 8/23/1999 4:52:44 PM | 1999-04-22 |  | NPEPPIN |  | Concentration | mg/L | Y | (fl. conc.) |  | Lead Total (fl. conc.) |  | Metals - Total; | Metals - Total; EMS code: PB-T | Chemical | Total | LAB
PB-T | ug/g | Lead Total | MTID | MET-DRY-CCMS-VA | Metals in Tissue by CRC ICPMS (DRY) Methodology: This method is conducted following British Columbia Lab Manual method "Metals in Animal Tissue and Vegetation (Biota) - Prescriptive". Tissue samples are homogenized and sub-sampled prior to hotblock digestion with nitric and hydrochloric acids, in combination with addition of hydrogen peroxide. Instrumental analysis is by collision cell inductively coupled plasma - mass spectrometry (modified from EPA Method 6020A). Method Limitation: This method employs a strong acid/peroxide digestion, and is intended to provide a conservative estimate of bio-available metals. Near complete recoveries are achieved for most toxicologically important metals, but elements associated with recalcitrant minerals may be only partially recovered. | 0.02 | Y | Y | Y | 122 | 8/13/2014 | 9/13/2022 | 10/26/2022 7:14:15 AM | 2015-03-09 | 3/11/2015 4:12:05 PM | RDGIBSON | EMS | MassPerMass | mg/kg | Y | (mass conc.) |  | Lead Total (mass conc.) |  | Metals - Total; | Metals - Total; EMS code: PB-T | Chemical | Total | LAB
PB-T | mg/L | Lead Total | X555 | Metals in water, total, "in-bottle" digestion, ICP-SFMS | The whole water sample collected in a HDPE/Teflon bottle, is de-capped in a Class 100 workstation, acidified to 2% (v/v) HNO3 and 1% (v/v) HCl and recapped. The sample is "in-bottle" digested in a convection oven, set at 60 degrees C for 16 hours. The sample is cooled and introduced to an inductively coupled argon plasma-sector field mass spectrometer (ICP-SFMS) by pneumatic nebulization. Each respective element is measured at a specific mass to charge ratio, m/z value expressed in amu (atomic mass units), and at resolution. LR: Low Resolution ~300, MR: Mid Resolution ~4000 and HR: High a specified instrument Resolution ~10000. Ref: Method number 02-2003, National Laboratory for Environmental Testing, CCIW, Environment Canada, Burlington. Requested April 2003. | 5e-06 | Y | Y | Y | 17336 | 5/20/2003 | 12/6/2023 2:08:00 AM | 2/14/2024 9:05:32 AM | 2004-01-19 |  | UMULLER |  | Concentration | mg/L | Y | (fl. conc.) |  | Lead Total (fl. conc.) |  | Metals - Total; | Metals - Total; EMS code: PB-T | Chemical | Total | LAB
PB-T | ug/g | Lead Total | FS02 | D&G,Nitric/HCl:GF | Samples are dried at 60 C, then either sieved to -100 mesh (150 um) and rolled to homogenize, or finely ground with a mortar and pestle or a shatter box laboratory mill and rolled (depending on the submitter's requirements). The samples are then weighed into Teflon digestion liners and digested with HNO3 and HCl in a microwave oven. Samples are cooled, volumized, and allowed to settle overnight before being decanted. Analysis is then performed via Graphite Furnace Atomic Absorption Spectrometry (GF-AAS) (low level analysis). GFAA Spectrometry involves heating the sample in three or more stages. First, a low current heats the graphite tube to dry the sample. The second, or charring stage, destroys the organic matter and volatilizes other components at an intermediate temperature. Finally, a high current heats the tube to incandescence and, in an inert atmosphere, atomizes the element being determined. Additional stages may be added to drying and charring, and to cool and clean the tubes in between samples. The resultant ground-state atomic vapour absorbs monochromatic radiation from the source. A photoelectric detector measures the intensity of the transmitted radiation characteristic of the element being determined, which is inversely proportional to the concentration of ground-state atoms in the optical path over a limited range. Low level Cadmium and Lead are determined using Deuterium lamp background correction system whereas Arsenic and Selenium are determined by using Zeeman background correction system. | 0.4 | Y | Y | Y | 22 | 2/20/1998 6:02:00 PM | 12/28/2000 1:29:00 PM | 5/30/2001 1:30:01 AM | 1996-06-19 | 6/3/1998 6:51:44 PM | MPAK |  | MassPerMass | mg/kg | Y | (mass conc.) |  | Lead Total (mass conc.) |  | Metals - Total; | Metals - Total; EMS code: PB-T | Chemical | Total | LAB
PB-T | mg/L | Lead Total | X179 | LA,HNO3,Graphite Furnace | Samples are lab-acidified with HNO3; HNO3 digestion; analysis by GFAA. | 0.001 | Y | Y | Y | 186 | 2/11/1998 | 7/18/2002 | 8/6/2002 5:43:22 PM | 1996-04-12 | 6/3/1998 6:51:38 PM | EMS |  | Concentration | mg/L | Y | (fl. conc.) |  | Lead Total (fl. conc.) |  | Metals - Total; | Metals - Total; EMS code: PB-T | Chemical | Total | LAB
PB-T | ug/g | Lead Total | X492 | Homo,HCl-HNO3 dig,ICP-OES,dry calc | Homogenize samples and take sub-samples. For Total Recoverable Metals Digest with Aqua Regia. Analyses by ICPOES (Inductively Coupled Plasma Optical Emission Spectroscopy). Perform moisture analysis on separate sub-sample and calculate dry value for trace elements. | 50 | Y | Y | Y | 1 |  |  | 8/24/2016 2:59:13 PM | 1998-11-13 |  | NPEPPIN |  | MassPerMass | mg/kg | Y | (mass conc.) |  | Lead Total (mass conc.) |  | Metals - Total; | Metals - Total; EMS code: PB-T | Chemical | Total | LAB
PB-T | ug/g | Lead Total | FB02 | Freeze Dry,Hom,Acid Dig,GF | Sample moisture is determined by weighing samples before and after freeze-drying. Freeze-dried samples are ground in a Braun blender before being weighed into Teflon PFA liners and digested with concentrated HNO3, concentrated HCl, and 50% H2O2 in a microwave oven; which results in a complete sample dissolution. Samples are cooled and volumized with Type I water. Analysis is performed via Inductively Coupled Argon Plasma Atomic Emission Spectrometry (ICP-AES) (simultaneous multielement analysis), Graphite Furnace Atomic Absorption Spectrometry (GF-AAS) (low level Cd and Pb analysis), Cold Vapor Atomic Fluorescence Spectrometry (CV-AFS) (Mercury analysis) and/or Flame Atomic Absorption Spectrometry (F-AAS) (K analysis). | 0.2 | Y | Y | Y | 63 | 7/17/2000 1:47:00 PM | 8/29/2001 9:41:00 AM | 11/20/2002 1:41:08 PM | 1996-06-19 | 6/3/1998 6:51:44 PM | MPAK |  | MassPerMass | mg/kg | Y | (mass conc.) |  | Lead Total (mass conc.) |  | Metals - Total; | Metals - Total; EMS code: PB-T | Chemical | Total | LAB
PC42 | ug/g | PCB 1242 | X213 | Hex/AcExt,Part,Chrom,GC | Hex/AcExt;Part;Chrom;GC | 0.01 | Y | Y | Y | 78 |  |  | 10/24/1996 2:26:15 PM | 1996-04-12 | 6/3/1998 6:51:42 PM | EMS |  | MassPerMass | mg/kg | Y | (mass conc.) |  | PCB 1242 (mass conc.) | 53469-21-9 | Polychlorinated biphenylds (PCBs); | Polychlorinated biphenylds (PCBs); EMS code: PC42 | Chemical |  | LAB
PC48 | ug/g | PCB 1248 | X213 | Hex/AcExt,Part,Chrom,GC | Hex/AcExt;Part;Chrom;GC | 0.01 | Y | Y | Y | 78 |  |  | 10/24/1996 2:26:15 PM | 1996-04-12 | 6/3/1998 6:51:42 PM | EMS |  | MassPerMass | mg/kg | Y | (mass conc.) |  | PCB 1248 (mass conc.) | 12672-29-6 | Polychlorinated biphenylds (PCBs); | Polychlorinated biphenylds (PCBs); EMS code: PC48 | Chemical |  | LAB
PC54 | ug/g | PCB 1254 | X213 | Hex/AcExt,Part,Chrom,GC | Hex/AcExt;Part;Chrom;GC | 0.01 | Y | Y | Y | 78 |  |  | 10/24/1996 2:26:15 PM | 1996-04-12 | 6/3/1998 6:51:42 PM | EMS |  | MassPerMass | mg/kg | Y | (mass conc.) |  | PCB 1254 (mass conc.) | 11097-69-1 | Polychlorinated biphenylds (PCBs); | Polychlorinated biphenylds (PCBs); EMS code: PC54 | Chemical |  | LAB
PC60 | ug/g | PCB 1260 | X213 | Hex/AcExt,Part,Chrom,GC | Hex/AcExt;Part;Chrom;GC | 0.01 | Y | Y | Y | 78 |  |  | 10/24/1996 2:26:15 PM | 1996-04-12 | 6/3/1998 6:51:42 PM | EMS |  | MassPerMass | mg/kg | Y | (mass conc.) |  | PCB 1260 (mass conc.) | 11096-82-5 | Polychlorinated biphenylds (PCBs); | Polychlorinated biphenylds (PCBs); EMS code: PC60 | Chemical |  | LAB
PCNB | ug/g | Pentachloronitrobenzene CAS: 82-68-8 | ZN03 | DCM+Ac.Extr:GC/MS | DCM+Ac.Extr:GC/MS EPA 8270E (mod) | 0.005 | Y | Y | Y | 2 | 11/14/2022 | 11/14/2022 | 12/15/2022 10:50:04 AM | 2022-06-23 |  | EMS |  | MassPerMass | mg/kg | Y | (mass conc.) |  | Pentachloronitrobenzene (mass conc.) | 82-68-8 | Pesticides; | Pesticides; EMS code: PCNB | Chemical |  | LAB
PCNB | mg/L | Pentachloronitrobenzene CAS: 82-68-8 | X268 | CH2Cl2 ext:GC/MS | CH2Cl2 ext:GC/MS |  | Y | Y | Y | 16 | 12/29/2021 | 9/2/2022 | 11/10/2022 9:54:07 AM | 2022-04-04 |  | EMS |  | Concentration | mg/L | Y | (fl. conc.) |  | Pentachloronitrobenzene (fl. conc.) | 82-68-8 | Pesticides; | Pesticides; EMS code: PCNB | Chemical |  | LAB
PCSP | None | Special | PCSP | Special Analysis by IC | Special Analysis by IC | 0 | N | N | N | 1 |  |  | 10/24/1996 7:09:27 AM | 1996-04-12 | 6/3/1998 6:51:37 PM | EMS |  | None | None | N | (unitless) |  | Special (unitless) |  |  | EMS code: PCSP | Chemical |  | LAB
PD-D | mg/L | Palladium Dissolved | X556 | Metals in water, dissolved, "in-bottle" digestion, ICP-SFMS | The whole water sample is filtered through a 0.45 micron cellulose acetate filter paper and the filtrate is stored in a HDPE/Teflon bottle. The container is de-capped in a Class 100 workstation, acidified to 2% (v/v) HNO3 and 1% (v/v) HCl and recapped. The sample is "in-bottle" digested in a convection oven, set at 60 degrees C for 16 hours. The sample is cooled and introduced to an inductively coupled argon plasma-sector field mass spectrometer (ICP-SFMS) by pneumatic nebulization. Each respective element is measured at a specific mass to charge ratio, m/z value expressed in amu (atomic mass units), and at a specified intrument resolution. LR: Low Resolution ~300, MR: Mid Resolution ~4000 and HR: High Resolution ~10000. Ref: Method number 02-2003, National Laboratory for Environmental Testing, CCIW, Environment Canada, Burlington, Requested April 2003. | 1e-06 | Y | Y | Y | 913 | 4/29/2021 12:42:00 AM | 12/6/2023 2:04:00 AM | 2/14/2024 9:05:32 AM | 2023-02-14 |  | EMS |  | Concentration | mg/L | Y | (fl. conc.) |  | Palladium Dissolved (fl. conc.) |  | Metals - Dissolved; | Metals - Dissolved; EMS code: PD-D | Chemical | Dissolved | LAB
PD-T | ppm (S) | Palladium | 6042 | Aqua Regia by ICPMS | These were sediment/soils digested with Aqua-Regia (Nitric Acid & Hydrochloric Acids) and then analysis by ICPMS. Results would be reported in ppm s. | 0.001 | Y | Y | Y | 71 | 5/14/2022 | 7/27/2024 | 7/29/2024 9:13:08 AM | 2022-06-28 |  | EMS |  | DimensionlessRatio | ppm | N | (ratio) |  | Palladium (ratio) |  | Metals; | Metals; EMS code: PD-T | Chemical |  | LAB
PD-T | ppm (S) | Palladium | 6041 | Aqua Regia by ICPMS | These were sediment/soils digested with Aqua-Regia (Nitric Acid & Hydrochloric Acids) and then analysis by ICPMS. Results would be reported in mg/kg and/or %. | 0.001 | Y | Y | Y | 21 | 10/13/2019 | 8/5/2020 | 8/10/2020 3:52:06 PM | 2019-09-06 | 10/21/2019 11:55:56 AM | KATFRASE | KATFRASE | DimensionlessRatio | ppm | N | (ratio) |  | Palladium (ratio) |  | Metals; | Metals; EMS code: PD-T | Chemical |  | LAB
PENA | mg/L | 1-Pentanol, n-Pentanol, CAS: 71-41-0 | E-8015 | EPA 8015D | EPA Method 8015D (SW-846): Nonhalogenated Organics Using GC/FID |  | Y | Y | Y | 14 | 12/29/2021 | 1/12/2022 | 9/2/2022 6:21:54 PM | 2022-04-05 |  | EMS |  | Concentration | mg/L | Y | (fl. conc.) |  | 1-Pentanol, n-Pentanol (fl. conc.) | 71-41-0 | Organics; | Organics; EMS code: PENA | Chemical |  | LAB
PERY | % | Perylene d12 | SURR | Surrogate Recovery | Surrogate Recovery | 30 | Y | Y | Y | 502 | 8/18/2003 | 10/30/2004 | 1/26/2005 1:01:44 AM | 1996-04-12 | 6/3/1998 6:51:38 PM | EMS |  | DimensionlessRatio | % | N | (ratio) |  | Perylene d12 (ratio) |  | Polycyclic Aromatic Hydrocarbons (PAHs); | Polycyclic Aromatic Hydrocarbons (PAHs); EMS code: PERY | Chemical |  | SURROGATE_RESULT
PERY | % | Perylene d12 | PAH1 | PUF:GC/MS | To understand the extent of human exposure to B[a]P and other PAHs, a reliable sampling and analytical method has been established which involves a combination quartz fibre filter/absorbent cartridge with subsequent extraction and analysis by gas chromatography (GC) with mass spectrometry (MS) detection (GC/MS). The analytical methods are a modification of EPA Test Method 625, Methods for Organic Chemical Analysis of Municipal and Industrial Wastewater, and Methods 8270, Test Methods for Evaluation of Solid Waste. PAH1 utilizes PUF as an adsorbant; PAH2 utilizes XAD-2 as an adsorbant. Note that isotopic surrogates may be added and are reported as % recovery. | 0 | Y | Y | Y | 100 | 10/14/1997 2:50:00 PM | 4/22/1998 5:40:00 PM | 1/19/1999 6:45:11 PM | 1996-04-12 | 6/3/1998 6:51:38 PM | EMS |  | DimensionlessRatio | % | N | (ratio) |  | Perylene d12 (ratio) |  | Polycyclic Aromatic Hydrocarbons (PAHs); | Polycyclic Aromatic Hydrocarbons (PAHs); EMS code: PERY | Chemical |  | SURROGATE_RESULT
PERY | mg/L | Perylene d12 | MW02 | DCM extrn,GC/MS-SIM (PAH) | Water samples are liquid - liquid extracted with dichloromethane. The extract is then cleaned by silica gel column chromatography and analyzed by gas chromatography - mass spectroscopy. | 1e-05 | Y | Y | Y | 111 | 2/5/2002 | 7/10/2024 5:51:00 PM | 8/21/2024 11:42:56 AM | 1996-04-12 | 6/3/1998 6:51:42 PM | EMS |  | Concentration | mg/L | Y | (fl. conc.) |  | Perylene d12 (fl. conc.) |  | Polycyclic Aromatic Hydrocarbons (PAHs); | Polycyclic Aromatic Hydrocarbons (PAHs); EMS code: PERY | Chemical |  | SURROGATE_RESULT
PERY | ug/g | Perylene d12 | MP02 | GC/MS | GC/MS | 0.001 | Y | Y | Y | 101 | 3/28/2002 | 11/29/2004 2:05:00 PM | 1/6/2005 1:01:58 AM | 1996-04-12 | 6/3/1998 6:51:41 PM | EMS |  | MassPerMass | mg/kg | Y | (mass conc.) |  | Perylene d12 (mass conc.) |  | Polycyclic Aromatic Hydrocarbons (PAHs); | Polycyclic Aromatic Hydrocarbons (PAHs); EMS code: PERY | Chemical |  | SURROGATE_RESULT
PF3ON | mg/L | Hexadecafluoro-3-oxanonane-1-sulfonic acid, 9-chloro- | E3533 | MECP E3533 | PFAs in water by direct injection LC/MS/MS | 2e-05 | Y | Y | Y | 2 | 6/3/2024 | 6/3/2024 | 8/21/2024 3:32:03 PM | 2023-11-15 | 6/27/2024 5:06:31 PM | EMS | EMS | Concentration | mg/L | Y | (fl. conc.) |  | Hexadecafluoro-3-oxanonane-1-sulfonic acid, 9-chloro- (fl. conc.) |  |  | EMS code: PF3ON | Chemical |  | LAB
PF3OU | mg/L | Eicosafluoro-3-oxaundecane-1-sulfonic acid, 11-chloro- | E3533 | MECP E3533 | PFAs in water by direct injection LC/MS/MS | 2e-05 | Y | Y | Y | 2 | 6/3/2024 | 6/3/2024 | 8/21/2024 3:32:03 PM | 2023-10-06 | 6/27/2024 5:06:31 PM | EMS | EMS | Concentration | mg/L | Y | (fl. conc.) |  | Eicosafluoro-3-oxaundecane-1-sulfonic acid, 11-chloro- (fl. conc.) |  |  | EMS code: PF3OU | Chemical |  | LAB
PFBA | mg/L | Perfluorobutanoic acid (PFBA) CAS: 375-22-4 | E3533 | MECP E3533 | PFAs in water by direct injection LC/MS/MS | 0.0001 | Y | Y | Y | 2 | 6/3/2024 | 6/3/2024 | 8/21/2024 3:32:03 PM | 2023-11-15 | 6/27/2024 5:06:31 PM | EMS | EMS | Concentration | mg/L | Y | (fl. conc.) |  | Perfluorobutanoic acid (PFBA) (fl. conc.) | 375-22-4 | Per- and polyfluoroalkyl substances (PFAS); | Per- and polyfluoroalkyl substances (PFAS); EMS code: PFBA | Chemical |  | LAB
PFBS | ug/g | Perfluorobutane sulfonic acid (PFBS) (CAS: 375-73-5) | 3506 | MOECC E3506 | MOECC E3506 | 0.0001 | Y | Y | Y | 8 | 3/18/2020 | 3/19/2020 | 2/21/2023 2:10:31 PM | 2023-02-06 |  | EMS |  | MassPerMass | mg/kg | Y | (mass conc.) |  | Perfluorobutane sulfonic acid (PFBS) (mass conc.) | 375-73-5 | Per- and polyfluoroalkyl substances (PFAS); | Per- and polyfluoroalkyl substances (PFAS); EMS code: PFBS | Chemical |  | LAB
PFBS | mg/L | Perfluorobutane sulfonic acid (PFBS) (CAS: 375-73-5) | E3533 | MECP E3533 | PFAs in water by direct injection LC/MS/MS | 1e-05 | Y | Y | Y | 2 | 6/3/2024 | 6/3/2024 | 8/21/2024 3:32:03 PM | 2023-11-15 | 6/27/2024 5:06:31 PM | EMS | EMS | Concentration | mg/L | Y | (fl. conc.) |  | Perfluorobutane sulfonic acid (PFBS) (fl. conc.) | 375-73-5 | Per- and polyfluoroalkyl substances (PFAS); | Per- and polyfluoroalkyl substances (PFAS); EMS code: PFBS | Chemical |  | LAB
PFD1 | mg/L | Perfluorododecanesulfonic acid (PFDoS) | E3533 | MECP E3533 | PFAs in water by direct injection LC/MS/MS | 2e-05 | Y | Y | Y | 2 | 6/3/2024 | 6/3/2024 | 8/21/2024 3:32:03 PM | 2023-11-15 | 6/27/2024 5:06:31 PM | EMS | EMS | Concentration | mg/L | Y | (fl. conc.) |  | Perfluorododecane sulfonic acid (PFDoS) (fl. conc.) | 335-77-3 | Per- and polyfluoroalkyl substances (PFAS); | Per- and polyfluoroalkyl substances (PFAS); EMS code: PFD1 | Chemical |  | LAB
PFD2 | mg/L | Perfluorododecanoic acid (PFDoA) CAS: 307-55-1 | E3533 | MECP E3533 | PFAs in water by direct injection LC/MS/MS | 2e-05 | Y | Y | Y | 2 | 6/3/2024 | 6/3/2024 | 8/21/2024 3:32:03 PM | 2023-11-15 | 6/27/2024 5:06:31 PM | EMS | EMS | Concentration | mg/L | Y | (fl. conc.) |  | Perfluorododecanoic acid (PFDoA) (fl. conc.) | 307-55-1 | Per- and polyfluoroalkyl substances (PFAS); | Per- and polyfluoroalkyl substances (PFAS); EMS code: PFD2 | Chemical |  | LAB
PFDA | mg/L | Perfluorodecanoic acid (PFDA) CAS: 335-76-2 | E3533 | MECP E3533 | PFAs in water by direct injection LC/MS/MS | 2e-05 | Y | Y | Y | 2 | 6/3/2024 | 6/3/2024 | 8/21/2024 3:32:03 PM | 2023-11-15 | 6/27/2024 5:06:31 PM | EMS | EMS | Concentration | mg/L | Y | (fl. conc.) |  | Perfluorodecanoic acid (PFDA) (fl. conc.) | 335-76-2 | Per- and polyfluoroalkyl substances (PFAS); | Per- and polyfluoroalkyl substances (PFAS); EMS code: PFDA | Chemical |  | LAB
PFDS | mg/L | Perfluorodecane sulfonic acid (PFDS) CAS: 335-77-3 | E3533 | MECP E3533 | PFAs in water by direct injection LC/MS/MS | 2e-05 | Y | Y | Y | 2 | 6/3/2024 | 6/3/2024 | 8/21/2024 3:32:03 PM | 2023-11-15 | 6/27/2024 5:06:31 PM | EMS | EMS | Concentration | mg/L | Y | (fl. conc.) |  | Perfluorodecane sulfonic acid (PFDS) (fl. conc.) | 335-77-3 | Per- and polyfluoroalkyl substances (PFAS); | Per- and polyfluoroalkyl substances (PFAS); EMS code: PFDS | Chemical |  | LAB
PFEES | mg/L | perfluoro(2-ethoxyethane)sulfonic acid (PFEESA) | E3533 | MECP E3533 | PFAs in water by direct injection LC/MS/MS | 2e-05 | Y | Y | Y | 2 | 6/3/2024 | 6/3/2024 | 8/21/2024 3:32:03 PM | 2023-11-15 | 6/27/2024 5:06:31 PM | EMS | EMS | Concentration | mg/L | Y | (fl. conc.) |  | perfluoro(2-ethoxyethane)sulfonic acid (PFEESA) (fl. conc.) | 113507-82-7 | Per- and polyfluoroalkyl substances (PFAS); | Per- and polyfluoroalkyl substances (PFAS); EMS code: PFEES | Chemical |  | LAB
PFEHS | mg/L | Perfluoroethylcyclohexanesulfonic acid (PFecHS) | E3533 | MECP E3533 | PFAs in water by direct injection LC/MS/MS | 2e-05 | Y | Y | Y | 2 | 6/3/2024 | 6/3/2024 | 8/21/2024 3:32:03 PM | 2023-11-15 | 6/27/2024 5:06:31 PM | EMS | EMS | Concentration | mg/L | Y | (fl. conc.) |  | Perfluoroethylcyclohexanesulfonic acid (PFecHS) (fl. conc.) | 1135682-88-0 | Per- and polyfluoroalkyl substances (PFAS); | Per- and polyfluoroalkyl substances (PFAS); EMS code: PFEHS | Chemical |  | LAB
PFHA | mg/L | Perfluorohexanoic acid (PFHxA) CAS: 307-24-4 | E3533 | MECP E3533 | PFAs in water by direct injection LC/MS/MS | 2e-05 | Y | Y | Y | 2 | 6/3/2024 | 6/3/2024 | 8/21/2024 3:32:03 PM | 2023-11-15 | 6/27/2024 5:06:31 PM | EMS | EMS | Concentration | mg/L | Y | (fl. conc.) |  | Perfluorohexanoic acid (PFHxA) (fl. conc.) | 307-24-4 | Per- and polyfluoroalkyl substances (PFAS); | Per- and polyfluoroalkyl substances (PFAS); EMS code: PFHA | Chemical |  | LAB
PFHDA | mg/L | Perfluorohexadecanoic acid (PFHxDA) | E3533 | MECP E3533 | PFAs in water by direct injection LC/MS/MS | 2e-05 | Y | Y | Y | 2 | 6/3/2024 | 6/3/2024 | 8/21/2024 3:32:03 PM | 2023-11-15 | 6/27/2024 5:06:31 PM | EMS | EMS | Concentration | mg/L | Y | (fl. conc.) |  | Perfluorohexadecanoic acid (PFHxDA) (fl. conc.) | 67905-19-5 | Per- and polyfluoroalkyl substances (PFAS); | Per- and polyfluoroalkyl substances (PFAS); EMS code: PFHDA | Chemical |  | LAB
PFHPA | mg/L | Perfluoroheptanoic acid (PFHpA) CAS: 375-85-9 | E3533 | MECP E3533 | PFAs in water by direct injection LC/MS/MS | 2e-05 | Y | Y | Y | 2 | 6/3/2024 | 6/3/2024 | 8/21/2024 3:32:03 PM | 2023-11-15 | 6/27/2024 5:06:31 PM | EMS | EMS | Concentration | mg/L | Y | (fl. conc.) |  | Perfluoroheptanoic acid (PFHpA) (fl. conc.) | 375-85-9 | Per- and polyfluoroalkyl substances (PFAS); | Per- and polyfluoroalkyl substances (PFAS); EMS code: PFHPA | Chemical |  | LAB
PFHPS | mg/L | Perfluoroheptanesulfonic acid (PFHpS) CAS: 375-92-8 | E3533 | MECP E3533 | PFAs in water by direct injection LC/MS/MS | 2e-05 | Y | Y | Y | 2 | 6/3/2024 | 6/3/2024 | 8/21/2024 3:32:03 PM | 2023-11-15 | 6/27/2024 5:06:31 PM | EMS | EMS | Concentration | mg/L | Y | (fl. conc.) |  | Perfluoroheptanesulfonic acid (PFHpS) (fl. conc.) | 375-92-8 | Per- and polyfluoroalkyl substances (PFAS); | Per- and polyfluoroalkyl substances (PFAS); EMS code: PFHPS | Chemical |  | LAB
PFHS | mg/L | Perfluorohexanesulfonic acid (PFHxS) CAS: 355-46-4 | E3533 | MECP E3533 | PFAs in water by direct injection LC/MS/MS | 2e-05 | Y | Y | Y | 2 | 6/3/2024 | 6/3/2024 | 8/21/2024 3:32:03 PM | 2023-11-15 | 6/27/2024 5:06:31 PM | EMS | EMS | Concentration | mg/L | Y | (fl. conc.) |  | Perfluorohexanesulfonic acid (PFHxS) (fl. conc.) | 355-46-4 | Per- and polyfluoroalkyl substances (PFAS); | Per- and polyfluoroalkyl substances (PFAS); EMS code: PFHS | Chemical |  | LAB
PFLT | None | Paint filter liquids test | 9095B | EPA 9095B (mod) | EPA 9095B (mod) Free Liquids are determined by the Paint Filter Test, where a specified amount of material is placed in a paint filter. If any portion of the test material passes through the filter in the five minute test period, the material is deemed to contain free liquids. The absence of any amount of free liquid classifies the sample as a solid. |  | Y | Y | Y | 0 |  |  |  | 2024-01-24 |  | RDGIBSON |  | None | None | N | (unitless) |  | Paint filter liquids test (unitless) |  | Physical Properties; | Physical Properties; EMS code: PFLT | Physical |  | LAB
PFMB | mg/L | Perfluoro-4-methoxybutanoic acid (PFMBA) | E3533 | MECP E3533 | PFAs in water by direct injection LC/MS/MS | 2e-05 | Y | Y | Y | 2 | 6/3/2024 | 6/3/2024 | 8/21/2024 3:32:03 PM | 2023-11-15 | 6/27/2024 5:06:31 PM | EMS | EMS | Concentration | mg/L | Y | (fl. conc.) |  | Perfluoro-4-methoxybutanoic acid (PFMBA) (fl. conc.) |  | Per- and polyfluoroalkyl substances (PFAS); | Per- and polyfluoroalkyl substances (PFAS); EMS code: PFMB | Chemical |  | LAB
PFMP | mg/L | Perfluoro-3-methoxypropanoic acid (PFMPA) | E3533 | MECP E3533 | PFAs in water by direct injection LC/MS/MS | 2e-05 | Y | Y | Y | 2 | 6/3/2024 | 6/3/2024 | 8/21/2024 3:32:03 PM | 2023-11-15 | 6/27/2024 5:06:31 PM | EMS | EMS | Concentration | mg/L | Y | (fl. conc.) |  | Perfluoro-3-methoxypropanoic acid (PFMPA) (fl. conc.) |  | Per- and polyfluoroalkyl substances (PFAS); | Per- and polyfluoroalkyl substances (PFAS); EMS code: PFMP | Chemical |  | LAB
PFNA | mg/L | Perfluorononanoic acid (PFNA) CAS: 375-95-1 | E3533 | MECP E3533 | PFAs in water by direct injection LC/MS/MS | 2e-05 | Y | Y | Y | 2 | 6/3/2024 | 6/3/2024 | 8/21/2024 3:32:03 PM | 2023-11-15 | 6/27/2024 5:06:31 PM | EMS | EMS | Concentration | mg/L | Y | (fl. conc.) |  | Perfluorononanoic acid (PFNA) (fl. conc.) | 375-95-1 | Per- and polyfluoroalkyl substances (PFAS); | Per- and polyfluoroalkyl substances (PFAS); EMS code: PFNA | Chemical |  | LAB
PFNS | mg/L | Perfluorononanesulfonic acid (PFNS) CAS: 68259-12-1 | E3533 | MECP E3533 | PFAs in water by direct injection LC/MS/MS | 2e-05 | Y | Y | Y | 2 | 6/3/2024 | 6/3/2024 | 8/21/2024 3:32:03 PM | 2023-11-15 | 6/27/2024 5:06:31 PM | EMS | EMS | Concentration | mg/L | Y | (fl. conc.) |  | Perfluorononanesulfonic acid (PFNS) (fl. conc.) | 68259-12-1 | Per- and polyfluoroalkyl substances (PFAS); | Per- and polyfluoroalkyl substances (PFAS); EMS code: PFNS | Chemical |  | LAB
PFOA | ug/g | Perfluorooctanoic acid (PFOA) (CAS: 335-67-1) | 3506 | MOECC E3506 | MOECC E3506 | 0.0001 | Y | Y | Y | 8 | 3/18/2020 | 3/19/2020 | 2/21/2023 2:10:31 PM | 2023-02-06 |  | EMS |  | MassPerMass | mg/kg | Y | (mass conc.) |  | Perfluorooctanoic acid (PFOA) (mass conc.) | 335-67-1 | Per- and polyfluoroalkyl substances (PFAS); | Per- and polyfluoroalkyl substances (PFAS); EMS code: PFOA | Chemical |  | LAB
PFOA | mg/L | Perfluorooctanoic acid (PFOA) (CAS: 335-67-1) | E3533 | MECP E3533 | PFAs in water by direct injection LC/MS/MS | 2e-05 | Y | Y | Y | 2 | 6/3/2024 | 6/3/2024 | 8/21/2024 3:32:03 PM | 2023-11-15 | 6/27/2024 5:06:31 PM | EMS | EMS | Concentration | mg/L | Y | (fl. conc.) |  | Perfluorooctanoic acid (PFOA) (fl. conc.) | 335-67-1 | Per- and polyfluoroalkyl substances (PFAS); | Per- and polyfluoroalkyl substances (PFAS); EMS code: PFOA | Chemical |  | LAB
PFOS | mg/L | Perfluorooctanesulfonic acid (PFOS) (CAS:1763-23-1) | E3533 | MECP E3533 | PFAs in water by direct injection LC/MS/MS | 2e-05 | Y | Y | Y | 2 | 6/3/2024 | 6/3/2024 | 8/21/2024 3:32:03 PM | 2023-11-15 | 6/27/2024 5:06:31 PM | EMS | EMS | Concentration | mg/L | Y | (fl. conc.) |  | Perfluorooctanesulfonic acid (PFOS) (fl. conc.) | 1763-23-1 | Per- and polyfluoroalkyl substances (PFAS); | Per- and polyfluoroalkyl substances (PFAS); EMS code: PFOS | Chemical |  | LAB
PFOS | ug/g | Perfluorooctanesulfonic acid (PFOS) (CAS:1763-23-1) | 3506 | MOECC E3506 | MOECC E3506 | 0.0005 | Y | Y | Y | 8 | 3/18/2020 | 3/19/2020 | 2/21/2023 2:10:31 PM | 2023-02-06 |  | EMS |  | MassPerMass | mg/kg | Y | (mass conc.) |  | Perfluorooctanesulfonic acid (PFOS) (mass conc.) | 1763-23-1 | Per- and polyfluoroalkyl substances (PFAS); | Per- and polyfluoroalkyl substances (PFAS); EMS code: PFOS | Chemical |  | LAB
PFOSA | mg/L | Perfluorooctanesulfonamide (PFOSA) CAS: 754-91-6 | E3533 | MECP E3533 | PFAs in water by direct injection LC/MS/MS | 2e-05 | Y | Y | Y | 2 | 6/3/2024 | 6/3/2024 | 8/21/2024 3:32:03 PM | 2023-11-15 | 6/27/2024 5:06:31 PM | EMS | EMS | Concentration | mg/L | Y | (fl. conc.) |  | Perfluorooctanesulfonamide (PFOSA) (fl. conc.) | 754-91-6 | Per- and polyfluoroalkyl substances (PFAS); | Per- and polyfluoroalkyl substances (PFAS); EMS code: PFOSA | Chemical |  | LAB
PFOSC | % | Perfluorooctanesulfonic acid (13C8-PFOS) CAS: 2265893-05-6 | E3533 | MECP E3533 | PFAs in water by direct injection LC/MS/MS | 1e-05 | Y | Y | Y | 0 |  |  |  | 2023-11-15 | 8/20/2024 12:03:10 PM | EMS | EMS | DimensionlessRatio | % | N | (ratio) |  | Perfluorooctanesulfonic acid (13C8-PFOS) (ratio) | 2265893-05-6 | Per- and polyfluoroalkyl substances (PFAS); | Per- and polyfluoroalkyl substances (PFAS); EMS code: PFOSC | Chemical |  | SURROGATE_RESULT
PFOSC | % | Perfluorooctanesulfonic acid (13C8-PFOS) CAS: 2265893-05-6 | SURR | Surrogate Recovery | Surrogate Recovery | 1e-05 | Y | Y | Y | 2 | 6/3/2024 | 6/3/2024 | 8/21/2024 3:32:03 PM | 2024-06-17 | 8/20/2024 12:03:00 PM | RDGIBSON | EMS | DimensionlessRatio | % | N | (ratio) |  | Perfluorooctanesulfonic acid (13C8-PFOS) (ratio) | 2265893-05-6 | Per- and polyfluoroalkyl substances (PFAS); | Per- and polyfluoroalkyl substances (PFAS); EMS code: PFOSC | Chemical |  | SURROGATE_RESULT
PFPA | mg/L | Perfluoropentanoic acid (PFPeA) CAS: 2706-90-3 | E3533 | MECP E3533 | PFAs in water by direct injection LC/MS/MS | 2e-05 | Y | Y | Y | 2 | 6/3/2024 | 6/3/2024 | 8/21/2024 3:32:03 PM | 2023-11-15 | 6/27/2024 5:06:31 PM | EMS | EMS | Concentration | mg/L | Y | (fl. conc.) |  | Perfluoropentanoic acid (PFPeA) (fl. conc.) | 2706-90-3 | Per- and polyfluoroalkyl substances (PFAS); | Per- and polyfluoroalkyl substances (PFAS); EMS code: PFPA | Chemical |  | LAB
PFPS | mg/L | Perfluoropentanesulfonic acid (PFPeS) CAS: 2706-91-4 | E3533 | MECP E3533 | PFAs in water by direct injection LC/MS/MS | 2e-05 | Y | Y | Y | 2 | 6/3/2024 | 6/3/2024 | 8/21/2024 3:32:03 PM | 2023-11-15 | 6/27/2024 5:06:31 PM | EMS | EMS | Concentration | mg/L | Y | (fl. conc.) |  | Perfluoropentanesulfonic acid (PFPeS) (fl. conc.) | 2706-91-4 | Per- and polyfluoroalkyl substances (PFAS); | Per- and polyfluoroalkyl substances (PFAS); EMS code: PFPS | Chemical |  | LAB
PFTED | mg/L | Perfluorotetradecanoic acid (PFTeDA) CAS: 376-06-7 | E3533 | MECP E3533 | PFAs in water by direct injection LC/MS/MS | 2e-05 | Y | Y | Y | 2 | 6/3/2024 | 6/3/2024 | 8/21/2024 3:32:03 PM | 2023-11-15 | 6/27/2024 5:06:31 PM | EMS | EMS | Concentration | mg/L | Y | (fl. conc.) |  | Perfluorotetradecanoic acid (PFTeDA) (fl. conc.) | 376-06-7 | Per- and polyfluoroalkyl substances (PFAS); | Per- and polyfluoroalkyl substances (PFAS); EMS code: PFTED | Chemical |  | LAB
PFTRD | mg/L | Perfluorotridecanoic acid (PFTrDA) CAS: 72629-94-8 | E3533 | MECP E3533 | PFAs in water by direct injection LC/MS/MS | 2e-05 | Y | Y | Y | 2 | 6/3/2024 | 6/3/2024 | 8/21/2024 3:32:03 PM | 2023-11-15 | 6/27/2024 5:06:31 PM | EMS | EMS | Concentration | mg/L | Y | (fl. conc.) |  | Perfluorotridecanoic acid (PFTrDA) (fl. conc.) | 72629-94-8 | Per- and polyfluoroalkyl substances (PFAS); | Per- and polyfluoroalkyl substances (PFAS); EMS code: PFTRD | Chemical |  | LAB
PFTRDS | mg/L | Perfluorotridecanesulfonic acid (PFTrDS) | E3533 | MECP E3533 | PFAs in water by direct injection LC/MS/MS | 2e-05 | Y | Y | Y | 2 | 6/3/2024 | 6/3/2024 | 8/21/2024 3:32:03 PM | 2023-11-15 | 6/27/2024 5:06:31 PM | EMS | EMS | Concentration | mg/L | Y | (fl. conc.) |  | Perfluorotridecanesulfonic acid (PFTrDS) (fl. conc.) | 791563-89-8 | Per- and polyfluoroalkyl substances (PFAS); | Per- and polyfluoroalkyl substances (PFAS); EMS code: PFTRDS | Chemical |  | LAB
PFUDS | mg/L | Perfluoroundecanesulfonic acid (PFUdS) | E3533 | MECP E3533 | PFAs in water by direct injection LC/MS/MS | 2e-05 | Y | Y | Y | 2 | 6/3/2024 | 6/3/2024 | 8/21/2024 3:32:03 PM | 2023-11-15 | 6/27/2024 5:06:31 PM | EMS | EMS | Concentration | mg/L | Y | (fl. conc.) |  | Perfluoroundecanesulfonic acid (PFUdS) (fl. conc.) | 749786-16-1 | Per- and polyfluoroalkyl substances (PFAS); | Per- and polyfluoroalkyl substances (PFAS); EMS code: PFUDS | Chemical |  | LAB
PFUN | mg/L | Perfluoroundecanoic acid (PFUnA) CAS: 2058-94-8 | E3533 | MECP E3533 | PFAs in water by direct injection LC/MS/MS | 2e-05 | Y | Y | Y | 2 | 6/3/2024 | 6/3/2024 | 8/21/2024 3:32:03 PM | 2023-11-15 | 6/27/2024 5:06:31 PM | EMS | EMS | Concentration | mg/L | Y | (fl. conc.) |  | Perfluoroundecanoic acid (PFUnA) (fl. conc.) | 2058-94-8 | Per- and polyfluoroalkyl substances (PFAS); | Per- and polyfluoroalkyl substances (PFAS); EMS code: PFUN | Chemical |  | LAB
PH-- | pH units | pH (Rain) | EQ05 | EQUIS (pH units) | EQUIS (pH units) | 1 | N | Y | Y | 122 |  |  | 10/22/1996 10:26:38 PM | 1996-04-12 | 6/3/1998 6:51:41 PM | EMS |  | Acidity | pH Units | Y | (acidity) |  | pH (Rain) (acidity) |  | Physical Properties; | Physical Properties; EMS code: PH-- | Physical |  | LAB
PH-- | pH units | pH (Rain) | 5065 | Meter: Glass/Ref Low Ionic Strength | Low ionic strength glass electrode (Radiometer PHC 2701 or equivalent). Calibrate with low ionic strength buffers. Measurement accuracy and precision to be +- 0.01 pH unit or better. Measurements made at 25 C or results are temperature corrected to 25 C. Code "PH--" must be used for precipitation and code "0004" must be used for surface waters. | 0 | Y | Y | Y | 5627 | 4/30/1997 3:50:00 PM | 4/10/2008 | 5/28/2008 2:08:44 AM | 1996-04-12 | 6/3/1998 6:51:41 PM | EMS |  | Acidity | pH Units | Y | (acidity) |  | pH (Rain) (acidity) |  | Physical Properties; | Physical Properties; EMS code: PH-- | Physical |  | LAB
PH-F | pH units | pH-Field | EXOSO | EXO Sonde YSI | EXO Sonde YSI |  | Y | Y | Y | 359 |  |  | 7/15/2024 3:01:47 PM | 2023-06-12 |  | EMS |  | Acidity | pH Units | Y | (acidity) |  | pH-Field (acidity) |  | Physical Properties; | Physical Properties; EMS code: PH-F | Physical |  | FIELD_RESULT | EXO Sonde YSI
PH-F | pH units | pH-Field | FLD | Field sampling | Field sampling method | 0.1 | Y | Y | Y | 8074 | 10/28/2015 | 8/26/2024 | 8/26/2024 4:51:04 PM | 2019-07-10 | 6/30/2020 4:45:03 PM | KATFRASE | RDGIBSON | Acidity | pH Units | Y | (acidity) |  | pH-Field (acidity) |  | Physical Properties; | Physical Properties; EMS code: PH-F | Physical |  | FIELD_RESULT | Field sampling
PH-F | pH units | pH-Field | XM00 | Meter-in situ | Meter-in situ | 1 | Y | Y | Y | 13132 | 4/15/2005 10:00:00 AM | 8/13/2024 4:38:00 PM | 8/26/2024 8:06:03 AM | 2012-10-12 | 12/20/2019 11:02:20 AM | RDGIBSON | RDGIBSON | Acidity | pH Units | Y | (acidity) |  | pH-Field (acidity) |  | Physical Properties; | Physical Properties; EMS code: PH-F | Physical |  | FIELD_RESULT | Meter-in situ
PH-LE | pH units | Final pH of solvent in leachability methods | X330 | Meter | The conductance (= conductivity) of a sample is measured by use of a self-contained conductivity meter. pH is measured with a combined glass/reference (Ag/AgCl) electrode. For both conductivity and pH the measurements are made at 25 C, or the results are temperature-corrected to 25 C. Note that laboratory and field measurements are distinguished by the EMS Analytical Agency codes. Ex situ and in situ measurements are distinguished by the computer codes. | 0.1 | Y | Y | Y | 15 | 1/10/2022 | 2/4/2022 | 1/11/2024 1:07:08 PM | 2022-06-20 |  | EMS |  | Acidity | pH Units | Y | (acidity) |  | Final pH of solvent in leachability methods (acidity) |  | Physical Properties; | Physical Properties; EMS code: PH-LE | Chemical |  | LAB
PH33 | mg/L | 2-Methylphenol | F028 | Acid,Extr.MeCl,HRGC/LRMS:TIM,SIM | A measured volume of sample is acidified and extracted with Methylene Chloride. For samples which have a high organic content that potentially produce chromatographic interferences an initial pre-extraction under basic conditions is performed. The sample extract is then evaporated to small volume and analysed by High Resolution Gas Chromatography / Low Resolution Mass Spectrometry (HRGC/LRMS). LRMS data is acquired in Total Ion Mode (TIM) and quantitation is performed using a single mass ion. Positive Identification of the various target compounds is made by comparison of absolute retention times to those of standards combined with computer based mass spectral mass/intensity library searching and matching routines. | 0.0005 | Y | Y | Y | 24 | 8/20/2009 | 4/14/2015 12:26:00 AM | 4/16/2015 6:00:11 PM | 2009-12-23 |  | RDGIBSON |  | Concentration | mg/L | Y | (fl. conc.) |  | 2-Methylphenol (fl. conc.) | 95-48-7 | Organics; | Organics; EMS code: PH33 | Chemical |  | LAB
PH37 | mg/L | 3+4-Methylphenol | F028 | Acid,Extr.MeCl,HRGC/LRMS:TIM,SIM | A measured volume of sample is acidified and extracted with Methylene Chloride. For samples which have a high organic content that potentially produce chromatographic interferences an initial pre-extraction under basic conditions is performed. The sample extract is then evaporated to small volume and analysed by High Resolution Gas Chromatography / Low Resolution Mass Spectrometry (HRGC/LRMS). LRMS data is acquired in Total Ion Mode (TIM) and quantitation is performed using a single mass ion. Positive Identification of the various target compounds is made by comparison of absolute retention times to those of standards combined with computer based mass spectral mass/intensity library searching and matching routines. | 0.0005 | Y | Y | Y | 24 | 8/20/2009 | 4/14/2015 12:26:00 AM | 4/16/2015 6:00:11 PM | 2009-12-23 |  | RDGIBSON |  | Concentration | mg/L | Y | (fl. conc.) |  | 3+4-Methylphenol (fl. conc.) | 1319-77-3, 108-39-4, 106-44-5 | Organics; | Organics; EMS code: PH37 | Chemical |  | LAB
PHEN | % | Phenanthrene d10 | PAH1 | PUF:GC/MS | To understand the extent of human exposure to B[a]P and other PAHs, a reliable sampling and analytical method has been established which involves a combination quartz fibre filter/absorbent cartridge with subsequent extraction and analysis by gas chromatography (GC) with mass spectrometry (MS) detection (GC/MS). The analytical methods are a modification of EPA Test Method 625, Methods for Organic Chemical Analysis of Municipal and Industrial Wastewater, and Methods 8270, Test Methods for Evaluation of Solid Waste. PAH1 utilizes PUF as an adsorbant; PAH2 utilizes XAD-2 as an adsorbant. Note that isotopic surrogates may be added and are reported as % recovery. | 0 | Y | Y | Y | 100 | 10/14/1997 2:50:00 PM | 4/22/1998 5:40:00 PM | 1/19/1999 6:45:11 PM | 1996-04-12 | 6/3/1998 6:51:41 PM | EMS |  | DimensionlessRatio | % | N | (ratio) |  | Phenanthrene d10 (ratio) | 1517-22-2 | Polycyclic Aromatic Hydrocarbons (PAHs); | Polycyclic Aromatic Hydrocarbons (PAHs); EMS code: PHEN | Chemical |  | SURROGATE_RESULT
PHEN | % | Phenanthrene d10 | SURR | Surrogate Recovery | Surrogate Recovery | 30 | Y | Y | Y | 1466 | 8/18/2003 | 8/15/2024 | 8/16/2024 1:36:03 PM | 1996-04-12 | 6/3/1998 6:51:38 PM | EMS |  | DimensionlessRatio | % | N | (ratio) |  | Phenanthrene d10 (ratio) | 1517-22-2 | Polycyclic Aromatic Hydrocarbons (PAHs); | Polycyclic Aromatic Hydrocarbons (PAHs); EMS code: PHEN | Chemical |  | SURROGATE_RESULT
PHEN | mg/L | Phenanthrene d10 | MW02 | DCM extrn,GC/MS-SIM (PAH) | Water samples are liquid - liquid extracted with dichloromethane. The extract is then cleaned by silica gel column chromatography and analyzed by gas chromatography - mass spectroscopy. | 1e-05 | Y | Y | Y | 72 | 2/5/2002 | 8/16/2004 | 9/21/2004 1:03:57 AM | 1996-04-12 | 6/3/1998 6:51:42 PM | EMS |  | Concentration | mg/L | Y | (fl. conc.) |  | Phenanthrene d10 (fl. conc.) | 1517-22-2 | Polycyclic Aromatic Hydrocarbons (PAHs); | Polycyclic Aromatic Hydrocarbons (PAHs); EMS code: PHEN | Chemical |  | SURROGATE_RESULT
PHEN | ug/g | Phenanthrene d10 | MP02 | GC/MS | GC/MS | 0.001 | Y | Y | Y | 101 | 3/28/2002 | 3/28/2011 | 5/9/2011 5:20:47 PM | 1996-04-12 | 6/3/1998 6:51:38 PM | EMS |  | MassPerMass | mg/kg | Y | (mass conc.) |  | Phenanthrene d10 (mass conc.) | 1517-22-2 | Polycyclic Aromatic Hydrocarbons (PAHs); | Polycyclic Aromatic Hydrocarbons (PAHs); EMS code: PHEN | Chemical |  | SURROGATE_RESULT
PHEN | % | Phenanthrene d10 | X262 | CH2Cl2 Ext:Silica:GC/MS | The sample is placed in a separatory funnel, the pH of the sample is adjusted to alkaline and then acidic conditions, and extracted each time with dichloromethane to selectively remove compounds. The final extracts (basic and acidic) can be combined or run separately. The extract is cleaned by silica gel column chromatography, then analyzed by GC/MS. | 1 | Y | Y | Y | 1 | 8/28/2014 | 8/28/2014 | 3/11/2015 3:57:19 PM | 2015-03-11 |  | RDGIBSON |  | DimensionlessRatio | % | N | (ratio) |  | Phenanthrene d10 (ratio) | 1517-22-2 | Polycyclic Aromatic Hydrocarbons (PAHs); | Polycyclic Aromatic Hydrocarbons (PAHs); EMS code: PHEN | Chemical |  | SURROGATE_RESULT
PHEN | % | Phenanthrene d10 | FS87 | DCM,Ac/Hex extr,PAM,HRGC/LRMS | DCM;Ac/Hex extr;PAM;HRGC/LRMS | 1 | Y | Y | Y | 43 | 5/20/2015 | 5/23/2015 | 7/3/2015 3:19:04 PM | 2015-05-27 |  | RDGIBSON |  | DimensionlessRatio | % | N | (ratio) |  | Phenanthrene d10 (ratio) | 1517-22-2 | Polycyclic Aromatic Hydrocarbons (PAHs); | Polycyclic Aromatic Hydrocarbons (PAHs); EMS code: PHEN | Chemical |  | SURROGATE_RESULT
PHES | pH units | pH of the Extraction Fluid used for EPA TCLP extraction | 1311 | EPA 1311/meter | This analysis is carried out in accordance with the extraction procedure outlined in "Test Methods for Evaluating Solid Waste - Physical/Chemical Methods Volume 1C" SW-846 EPA Method 1311, published by the United States Environmental Protection Agency (EPA). Thisple is extracted at a 20:1 liquid to solids ratio for 16 to 20 hours using either extraction fluid #1 (glacial acetic acid, water and sodium hydroxide) or extraction fluid #2 (glacial acetic acid), depending on the pH of the original sample. The extract is then filtered through a 0.6 to 0.8 micron glass fibre filter and analysed using inductively coupled plasma - optical emission spectrophotometry (EPA Method 6010B) | 0.1 | Y | Y | Y | 194 |  |  | 5/29/2024 2:43:20 PM | 2011-04-27 |  | RDGIBSON |  | Acidity | pH Units | Y | (acidity) |  | pH of the Extraction Fluid used for EPA TCLP extraction (acidity) |  | Toxicity; | Toxicity; EMS code: PHES | Chemical |  | LAB
PICL | mg/L | Picloram | F052 | Acid,DCM extr'n,methn,Florisil,HRGC-EC | Effluent and water samples are acidified with Sulphuric Acid and extracted with Dichloromethane. After concentration, the extract is methylated using ethereal Diazomethane. Extracts are cleaned up as required using Florisil. The extracts are analysed for Acid Extractable Herbicides using High Resolution Gas Chromatography with Electron Capture detection. Positive results may be verified by use of a secondary confirmation GC column. | 5e-05 | Y | Y | Y | 33 | 12/8/1998 3:39:00 PM | 12/31/2001 12:00:00 PM | 6/18/2003 2:53:52 PM | 1997-10-20 | 6/3/1998 6:51:47 PM | NPEPPIN |  | Concentration | mg/L | Y | (fl. conc.) |  | Picloram (fl. conc.) | 1918-02-1 | Pesticides; | Pesticides; EMS code: PICL | Chemical |  | LAB
PICL | mg/L | Picloram | AEWA | H2SO4 acid,CHCl2 extr.florisil. GC/MS | Effluent and water samples are acidified with Sulphuric Acid and extracted with Dichloromethane. After concentration, the extract is methylated using ethereal Diazomethane. Extracts are cleaned up as required using Florisil. The extracts are analyzed for Acid Extractable Herbicides using GC/MS. | 0.0002 | Y | Y | Y | 6 | 6/16/2003 | 7/29/2005 7:10:00 AM | 6/2/2009 12:01:11 PM | 1996-04-12 | 6/3/1998 6:51:42 PM | EMS |  | Concentration | mg/L | Y | (fl. conc.) |  | Picloram (fl. conc.) | 1918-02-1 | Pesticides; | Pesticides; EMS code: PICL | Chemical |  | LAB
PICL | mg/L | Picloram | P015 | Phenoxyacid Herbicides by LC-MS/MS | Phenoxyacid Herbicides by LC-MS/MS | 0.5 | Y | Y | Y | 71 | 10/9/2019 | 9/9/2022 | 12/8/2022 12:57:05 PM | 2019-10-17 |  | KATFRASE |  | Concentration | mg/L | Y | (fl. conc.) |  | Picloram (fl. conc.) | 1918-02-1 | Pesticides; | Pesticides; EMS code: PICL | Chemical |  | LAB
PICL | ug/g | Picloram | ZN03 | DCM+Ac.Extr:GC/MS | DCM+Ac.Extr:GC/MS EPA 8270E (mod) | 0.001 | Y | Y | Y | 15 | 1/14/2022 | 2/2/2022 | 1/11/2024 1:07:08 PM | 2022-07-27 |  | EMS |  | MassPerMass | mg/kg | Y | (mass conc.) |  | Picloram (mass conc.) | 1918-02-1 | Pesticides; | Pesticides; EMS code: PICL | Chemical |  | LAB
PICL | ug/g | Picloram | FS15 | DCM/MEOH acid extr'n,DCM,methn,Florisil,HRGC-EC | Sediment samples are shaken with a mixture of Dichloromethane, Methanol and Sulphuric Acid, filtered, then extracted with Dichloromethane. After concentration, the extract is methylated using ethereal Diazomethane. Extracts are cleaned up as required using Florisil. The extracts are analysed for Acid Extractable Herbicides using High Resolution Gas Chromatography with Electron Capture detection. Positive results may be verified by use of a secondary confirmation GC column. | 0.05 | Y | Y | Y | 35 | 4/24/1997 4:15:00 PM | 6/24/2003 | 2/26/2004 1:05:30 AM | 1997-10-20 | 6/3/1998 6:51:47 PM | NPEPPIN |  | MassPerMass | mg/kg | Y | (mass conc.) |  | Picloram (mass conc.) | 1918-02-1 | Pesticides; | Pesticides; EMS code: PICL | Chemical |  | LAB
PM02 | ug/m3 | Particulate <2.5u | PTFEO | [Partisol] Teflon Filter Equipped with O-ring, Gravimetric | [Partisol] Pure Teflon Filter Equipped with O-ring 2um opening - 47mm in diameter | 2 | Y | Y | Y | 13915 | 3/15/2007 | 5/27/2024 8:00:00 AM | 6/28/2024 1:06:55 PM | 2006-06-28 | 6/28/2006 1:46:01 PM | UMULLER | UMULLER | AirConcentration | µg/m³ | Y | (air conc.) |  | Particulate <2.5u (air conc.) |  | Particulates; | Particulates; EMS code: PM02 | Chemical |  | LAB
PM02 | ug/m3 | Particulate <2.5u | 5309 | Tef,47mm,Gravimetric | Teflon;47mm;Gravimetric. Used for PM2.5 and PM10 | 2 | Y | Y | Y | 449 | 7/11/2023 | 8/15/2024 | 8/16/2024 11:33:06 AM | 2001-04-06 |  | NPEPPIN |  | AirConcentration | µg/m³ | Y | (air conc.) |  | Particulate <2.5u (air conc.) |  | Particulates; | Particulates; EMS code: PM02 | Chemical |  | LAB
PM02 | ug/m3 | Particulate <2.5u | X048 | Gravimetric | Gravimetric | 0.5 | N | Y | Y | 90 |  |  | 10/24/1996 9:18:47 AM | 1996-04-12 | 6/3/1998 6:51:41 PM | EMS |  | AirConcentration | µg/m³ | Y | (air conc.) |  | Particulate <2.5u (air conc.) |  | Particulates; | Particulates; EMS code: PM02 | Chemical |  | LAB
PM02 | ug/m3 | Particulate <2.5u | 5305 | Tef:Tot Part HV 0.3um | PM10 is the designation for particulate matter in the atmosphere that has an anerodynamic diameter of 10 micrometers (um) or less for method PM10 5305. A high volume (HV) PM10 sampler draws a known volume of ambient air at a constant flow rate through a size selective inlet and through one or more filters. Particles in the PM10 size range are then collect on the 8X10 filter(s) during the specified 24 hour sampling period. Each sample filter is weighted before and after sampling to determine the net weight (mass) gain of the collected PM10 sample. For method TP-T 5305, a high volume (HV) sampler draws a known volume of ambient air at a constant flow rate through one or more filters. Particles are then collected on the 8X10 filter(s) during the specified 24 hour sampling period. Each sample filter is weighed before and after sampling to determine the net weight (mass) gain of the collected total particulate sample. For method PM02 5305, a measured volume of ambient air is drawn through an inlet that passes only particles less than 2.5 um. The particulate which is collected on a 0.3 micron teflon coated borosilicate 8X10 glass filter constitutes PM02 particulate. | 2 | Y | Y | Y | 81 | 8/11/2000 9:54:00 AM | 12/17/2001 3:32:00 PM | 1/4/2002 1:59:28 AM | 1998-01-13 | 6/3/1998 6:51:47 PM | NPEPPIN |  | AirConcentration | µg/m³ | Y | (air conc.) |  | Particulate <2.5u (air conc.) |  | Particulates; | Particulates; EMS code: PM02 | Chemical |  | LAB
PM02 | ug/m3 | Particulate <2.5u | 5142 | Gravimetric | PM10 and PM2.5 | 2 | Y | Y | Y | 358 | 3/31/2009 | 4/19/2012 | 4/20/2012 3:47:58 PM | 2001-03-13 | 3/13/2001 11:21:30 AM | NPEPPIN | NPEPPIN | AirConcentration | µg/m³ | Y | (air conc.) |  | Particulate <2.5u (air conc.) |  | Particulates; | Particulates; EMS code: PM02 | Chemical |  | LAB
PM02 | ug/m3 | Particulate <2.5u | 5306 | 47mm 0.3um Teflon-coated Filter | For methods TP-T 5306 and PM10 5306, a measured volume of air is drawn through a 47 mm filter rather than through 8x10 filters as for TP-T 5305 and PM10 5305. Air particulate is trapped on a pre-weighed Teflon filter. The weight change after sampling is used to calculate the particulate in ug/m3. | 2 | Y | Y | Y | 7996 | 7/3/1997 4:30:00 PM | 2/4/2009 1:05:00 PM | 10/29/2015 8:19:16 AM | 1996-04-12 | 4/6/2001 2:07:09 PM | EMS | NPEPPIN | AirConcentration | µg/m³ | Y | (air conc.) |  | Particulate <2.5u (air conc.) |  | Particulates; | Particulates; EMS code: PM02 | Chemical |  | LAB
PM02 | ug/m3 | Particulate <2.5u | DICH | Gravimetric, PTFEO for Dichot | Gravimetric method, PTFEO Teflon filter for Dichotomous Sampler | 2 | Y | Y | Y | 1191 | 4/27/2009 | 3/15/2022 | 5/9/2022 1:15:09 PM | 2009-09-02 |  | RDGIBSON |  | AirConcentration | µg/m³ | Y | (air conc.) |  | Particulate <2.5u (air conc.) |  | Particulates; | Particulates; EMS code: PM02 | Chemical |  | LAB
PM10 | ug/m3 | Particulate <10u | 5305 | Tef:Tot Part HV 0.3um | PM10 is the designation for particulate matter in the atmosphere that has an anerodynamic diameter of 10 micrometers (um) or less for method PM10 5305. A high volume (HV) PM10 sampler draws a known volume of ambient air at a constant flow rate through a size selective inlet and through one or more filters. Particles in the PM10 size range are then collect on the 8X10 filter(s) during the specified 24 hour sampling period. Each sample filter is weighted before and after sampling to determine the net weight (mass) gain of the collected PM10 sample. For method TP-T 5305, a high volume (HV) sampler draws a known volume of ambient air at a constant flow rate through one or more filters. Particles are then collected on the 8X10 filter(s) during the specified 24 hour sampling period. Each sample filter is weighed before and after sampling to determine the net weight (mass) gain of the collected total particulate sample. For method PM02 5305, a measured volume of ambient air is drawn through an inlet that passes only particles less than 2.5 um. The particulate which is collected on a 0.3 micron teflon coated borosilicate 8X10 glass filter constitutes PM02 particulate. | 2 | Y | Y | Y | 30974 | 7/3/1997 2:45:00 PM | 4/6/2009 10:18:00 AM | 6/25/2009 10:06:40 AM | 1996-04-12 | 6/3/1998 6:51:37 PM | EMS |  | AirConcentration | µg/m³ | Y | (air conc.) |  | Particulate <10u (air conc.) |  | Particulates; | Particulates; EMS code: PM10 | Chemical |  | LAB
PM10 | ug/m3 | Particulate <10u | 5309 | Tef,47mm,Gravimetric | Teflon;47mm;Gravimetric. Used for PM2.5 and PM10 | 2 | Y | Y | Y | 102 | 7/11/2023 | 8/15/2024 | 8/16/2024 11:33:06 AM | 2001-04-06 |  | NPEPPIN |  | AirConcentration | µg/m³ | Y | (air conc.) |  | Particulate <10u (air conc.) |  | Particulates; | Particulates; EMS code: PM10 | Chemical |  | LAB
PM10 | ug/m3 | Particulate <10u | 5205 | Tot Part HiVol 0.3 um | Tot Part HiVol 0.3 um | 2 | Y | Y | Y | 228 |  |  | 10/29/2015 8:19:16 AM | 1996-04-12 | 12/7/2000 11:30:12 AM | EMS | NPEPPIN | AirConcentration | µg/m³ | Y | (air conc.) |  | Particulate <10u (air conc.) |  | Particulates; | Particulates; EMS code: PM10 | Chemical |  | LAB
PM10 | ug/m3 | Particulate <10u | DICH | Gravimetric, PTFEO for Dichot | Gravimetric method, PTFEO Teflon filter for Dichotomous Sampler | 2 | Y | Y | Y | 102 | 5/8/2007 | 3/15/2022 | 5/9/2022 1:15:09 PM | 2009-09-02 |  | RDGIBSON |  | AirConcentration | µg/m³ | Y | (air conc.) |  | Particulate <10u (air conc.) |  | Particulates; | Particulates; EMS code: PM10 | Chemical |  | LAB
PM10 | ug/m3 | Particulate <10u | PTFEO | [Partisol] Teflon Filter Equipped with O-ring, Gravimetric | [Partisol] Pure Teflon Filter Equipped with O-ring 2um opening - 47mm in diameter | 2 | Y | Y | Y | 7204 | 3/21/2007 | 5/3/2024 | 5/7/2024 9:34:05 AM | 2006-06-28 | 6/28/2006 1:45:45 PM | UMULLER | UMULLER | AirConcentration | µg/m³ | Y | (air conc.) |  | Particulate <10u (air conc.) |  | Particulates; | Particulates; EMS code: PM10 | Chemical |  | LAB
PM10 | ug/m3 | Particulate <10u | 5306 | 47mm 0.3um Teflon-coated Filter | For methods TP-T 5306 and PM10 5306, a measured volume of air is drawn through a 47 mm filter rather than through 8x10 filters as for TP-T 5305 and PM10 5305. Air particulate is trapped on a pre-weighed Teflon filter. The weight change after sampling is used to calculate the particulate in ug/m3. | 2 | Y | Y | Y | 3306 | 10/27/1997 2:35:00 PM | 2/23/2007 | 10/29/2015 8:19:16 AM | 1996-04-12 | 6/3/1998 6:51:37 PM | EMS |  | AirConcentration | µg/m³ | Y | (air conc.) |  | Particulate <10u (air conc.) |  | Particulates; | Particulates; EMS code: PM10 | Chemical |  | LAB
PM10 | ug/m3 | Particulate <10u | 5005 | Gravimetric Anal-ConEnvR | Gravimetric Anal-ConEnvR | 2 | Y | Y | Y | 5463 | 4/4/2005 | 11/16/2010 | 1/11/2011 1:33:13 PM | 1996-04-12 | 6/3/1998 6:51:37 PM | EMS |  | AirConcentration | µg/m³ | Y | (air conc.) |  | Particulate <10u (air conc.) |  | Particulates; | Particulates; EMS code: PM10 | Chemical |  | LAB
PM10 | ug/m3 | Particulate <10u | X048 | Gravimetric | Gravimetric | 2 | N | Y | Y | 430 |  |  | 10/24/1996 2:01:05 PM | 1996-04-12 | 12/7/2000 11:30:12 AM | EMS | NPEPPIN | AirConcentration | µg/m³ | Y | (air conc.) |  | Particulate <10u (air conc.) |  | Particulates; | Particulates; EMS code: PM10 | Chemical |  | LAB
PM10 | ug/m3 | Particulate <10u | 5142 | Gravimetric | PM10 and PM2.5 | 2 | Y | Y | Y | 7157 |  |  | 7/3/2024 12:59:27 PM | 1996-04-12 | 6/3/1998 6:51:37 PM | EMS |  | AirConcentration | µg/m³ | Y | (air conc.) |  | Particulate <10u (air conc.) |  | Particulates; | Particulates; EMS code: PM10 | Chemical |  | LAB
PMC | ug/m3 | Total coarse particulates in air | 5309 | Tef,47mm,Gravimetric | Teflon;47mm;Gravimetric. Used for PM2.5 and PM10 | 2 | Y | Y | Y | 102 | 7/11/2023 | 8/15/2024 | 8/16/2024 11:33:06 AM | 2023-08-23 |  | EMS |  | AirConcentration | µg/m³ | Y | (air conc.) |  | Total coarse particulates in air (air conc.) |  | Particulates; | Particulates; EMS code: PMC | Physical | Total | LAB
PMCO | ug/m3 | Mass of particulates between PM2.5 to PM10 from Dichot sampl | DICH | Gravimetric, PTFEO for Dichot | Gravimetric method, PTFEO Teflon filter for Dichotomous Sampler | 2 | Y | Y | Y | 1472 | 10/15/2005 | 3/15/2022 | 5/9/2022 1:15:09 PM | 2009-09-02 |  | RDGIBSON |  | AirConcentration | µg/m³ | Y | (air conc.) |  | Mass of particulates between PM2.5 to PM10 from Dichot sampl (air conc.) |  | Particulates; | Particulates; EMS code: PMCO | Chemical |  | LAB
PMF | ug/m3 | Total fine particulates in air | 5309 | Tef,47mm,Gravimetric | Teflon;47mm;Gravimetric. Used for PM2.5 and PM10 | 2 | Y | Y | Y | 102 | 7/11/2023 | 8/15/2024 | 8/16/2024 11:33:06 AM | 2023-08-23 |  | EMS |  | AirConcentration | µg/m³ | Y | (air conc.) |  | Total fine particulates in air (air conc.) |  | Particulates; | Particulates; EMS code: PMF | Physical | Total | LAB
PO4- | mg/dm2/d | Phosphate-Soluble | EQ15 | EQUIS (mg/dm2/d) | EQUIS (mg/dm2/d) | 0.0002 | N | Y | Y | 3 |  |  | 10/22/1996 8:20:30 PM | 1996-04-12 | 6/3/1998 6:51:37 PM | EMS |  | MassFlux | mg/dm²/d | Y | (mass flux) |  | Phosphate-Soluble (mass flux) | 14265-44-2 | Nutrients; | Nutrients; EMS code: PO4- | Chemical | Soluble | LAB
PO4- | mg/dm2/d | Phosphate-Soluble | 5057 | Aut.Col/ Ascorbic Acid | Aut.Col/ Ascorbic Acid | 0.0002 | Y | Y | Y | 5 |  |  | 10/23/1996 10:06:33 PM | 1996-04-12 | 6/3/1998 6:51:37 PM | EMS |  | MassFlux | mg/dm²/d | Y | (mass flux) |  | Phosphate-Soluble (mass flux) | 14265-44-2 | Nutrients; | Nutrients; EMS code: PO4- | Chemical | Soluble | LAB
PO4- | ug/m3 | Phosphate-Soluble | EQ40 | EQUIS (mg/m3) | EQUIS (mg/m3) | 0.01 | N | Y | Y | 12 |  |  | 10/22/1996 2:45:22 PM | 1996-04-12 | 6/3/1998 6:51:37 PM | EMS |  | AirConcentration | µg/m³ | Y | (air conc.) |  | Phosphate-Soluble (air conc.) | 14265-44-2 | Nutrients; | Nutrients; EMS code: PO4- | Chemical | Soluble | LAB
PO4- | mg/L | Phosphate-Soluble | 5076 | Auto Col-Ascorbic Acd Rd | Auto Col-Ascorbic Acd Rd | 0.009 | Y | Y | Y | 209 | 8/7/2003 | 8/7/2003 | 11/27/2003 1:32:13 PM | 1996-04-12 | 6/3/1998 6:51:37 PM | EMS |  | Concentration | mg/L | Y | (fl. conc.) |  | Phosphate-Soluble (fl. conc.) | 14265-44-2 | Nutrients; | Nutrients; EMS code: PO4- | Chemical | Soluble | LAB
PO4- | mg/L | Phosphate-Soluble | 5068 | Ion Chromatography-Anion | (a) For the Soluble Particulates Pkg., a portion of filtrate from the Particulate-Total procedure made up to 500 mL is analyzed by Ion Chromatography. Intermediate results are in mg/L. Loading results are then calculated with units of mg/dm2/d. (b) For precipitation samples, samples and standards are spiked with concentrated carbonate/bicarbonate solution to give the sample the same background as the eluent used (avoiding a "water dip" effect). The anions of interest are separated through an anion "guard" and anion "separator" column. An anion micro membrane suppressor following the separator columns is used to reduce the background eluent conductivity by converting the carbonate and bicarbonate species to carbonic acid, while enhancing the conductivity of the ions of interest by converting these ions to their corresponding acids. A conductivity detector senses the same species in direct proportion to their initial concentration. (c) For HiVol samples, discs are cut from the HiVol filter, extracted with DI water, and anions analyzed by ion chromatograph, Preliminary results are reported in mg/L and intermediate HiVol loading results are reported as ug/m3. (*) Note that the air-convention is to report on basis of ion weight not element, e.g. as NO3 for nitrate, and as PO4 for phosphate. | 0.01 | Y | Y | Y | 276 |  |  | 6/18/2018 9:20:13 AM | 1996-04-12 | 6/3/1998 6:51:37 PM | EMS |  | Concentration | mg/L | Y | (fl. conc.) |  | Phosphate-Soluble (fl. conc.) | 14265-44-2 | Nutrients; | Nutrients; EMS code: PO4- | Chemical | Soluble | LAB
PO4-E | ug/g | Ortho-Phosphate (as PO4) Extractable | CAL2 | Calculation (ug/g) | Calculation (ug/g) | 0.01 | Y | Y | Y | 15 | 1/14/2022 | 2/8/2022 | 1/11/2024 1:07:08 PM | 2022-07-06 |  | EMS |  | MassPerMass | mg/kg | Y | (mass conc.) |  | Ortho-Phosphate (as PO4) Extractable (mass conc.) | 14265-44-2 | Nutrients; | Nutrients; EMS code: PO4-E | Chemical | Extractable | LAB
PO4-E | mg/L | Ortho-Phosphate (as PO4) Extractable | A4500P | APHA 4500-P | See APHA 4500-P for metod description. |  | Y | Y | Y | 15 | 1/13/2022 | 2/6/2022 | 1/11/2024 1:07:08 PM | 2022-07-07 |  | EMS |  | Concentration | mg/L | Y | (fl. conc.) |  | Ortho-Phosphate (as PO4) Extractable (fl. conc.) | 14265-44-2 | Nutrients; | Nutrients; EMS code: PO4-E | Chemical | Extractable | LAB
PO4-S | mg/L | Phosphate; ortho- (as PO4); soluble ion content | A4500P | APHA 4500-P | See APHA 4500-P for metod description. |  | Y | Y | Y | 0 |  |  |  | 2024-02-02 |  | EMS |  | Concentration | mg/L | Y | (fl. conc.) |  | Phosphate; ortho- (as PO4); soluble ion content (fl. conc.) |  | Nutrients; | Nutrients; EMS code: PO4-S | Chemical |  | LAB
PO4-S | ug/g | Phosphate; ortho- (as PO4); soluble ion content | CAL2 | Calculation (ug/g) | Calculation (ug/g) | 0.01 | Y | Y | Y | 0 |  |  |  | 2024-02-02 |  | EMS |  | MassPerMass | mg/kg | Y | (mass conc.) |  | Phosphate; ortho- (as PO4); soluble ion content (mass conc.) |  | Nutrients; | Nutrients; EMS code: PO4-S | Chemical |  | LAB
PP-I | mg/dm2/d | Phosphorus Tot Insoluble | 5132 | Dig,Auto Color As Acd A | For air particulate a 10.4 cm cannister containing a collection medium is exposed to the ambient air for a period of approximately 30 days. The prepared sample undergoes a sulfuric acid persulfate digestion to oxidize organically bound phosphorus to phosphate and also to hydrolyse polyphosphates to ortho-phosphate, which is then determined colorimetrically as heteropoly molybdophosphoric acid. Intermediate results are reported with units of mg/L and loading results with units of mg/dm2/d. Note that the air convention is to report as ion weight, ie as PO4 not as P. | 6e-05 | Y | Y | Y | 520 |  |  | 4/2/1998 8:39:05 AM | 1996-04-12 | 6/3/1998 6:51:37 PM | EMS |  | MassFlux | mg/dm²/d | Y | (mass flux) |  | Phosphorus Tot Insoluble (mass flux) | 7723-14-0 | Nutrients; | Nutrients; EMS code: PP-I | Chemical | Total | LAB
PP-I | mg/dm2/d | Phosphorus Tot Insoluble | EQ15 | EQUIS (mg/dm2/d) | EQUIS (mg/dm2/d) | 0.0001 | N | Y | Y | 5 |  |  | 10/22/1996 8:20:30 PM | 1996-04-12 | 6/3/1998 6:51:45 PM | EMS |  | MassFlux | mg/dm²/d | Y | (mass flux) |  | Phosphorus Tot Insoluble (mass flux) | 7723-14-0 | Nutrients; | Nutrients; EMS code: PP-I | Chemical | Total | LAB
PP-S | mg/dm2/d | Phosphorus Tot Soluble | 5135 | Dig, Auto Colour As Acd | For air particulate a 10.4 cm cannister containing a collection medium is exposed to the ambient air for a period of approximately 30 days. The prepared sample undergoes a sulfuric acid persulfate digestion to oxidize organically bound phosphorus to phosphate and also to hydrolyse polyphosphates to ortho-phosphate, which is then determinied colorimetrically as heteropoly molybdophosphoric acid. Intermediate results are reported with units of mg/L and loading results with units of mg/dm2/d. Note that the air convention is to report as ion-weight, ie PO4 not as P. | 6e-05 | Y | Y | Y | 532 |  |  | 4/2/1998 8:40:00 AM | 1996-04-12 | 6/3/1998 6:51:37 PM | EMS |  | MassFlux | mg/dm²/d | Y | (mass flux) |  | Phosphorus Tot Soluble (mass flux) | 7723-14-0 | Nutrients; | Nutrients; EMS code: PP-S | Chemical | Soluble | LAB
PP-S | mg/dm2/d | Phosphorus Tot Soluble | EQ15 | EQUIS (mg/dm2/d) | EQUIS (mg/dm2/d) | 0 | N | Y | Y | 5 |  |  | 10/22/1996 8:20:30 PM | 1996-04-12 | 6/3/1998 6:51:37 PM | EMS |  | MassFlux | mg/dm²/d | Y | (mass flux) |  | Phosphorus Tot Soluble (mass flux) | 7723-14-0 | Nutrients; | Nutrients; EMS code: PP-S | Chemical | Soluble | LAB
PP-S | mg/L | Phosphorus Tot Soluble | 5139 | Dig, Auto Color As Acd A | For air particulate a 10.4 cm cannister containing a collection medium is exposed to the ambient air for a period of approximately 30 days. The prepared sample undergoes a sulfuric acid persulfate digestion to oxidize organically bound phosphorus to phosphate and also to hydrolyse polyphosphates to orthophosphate, which is then determined colorimetrically as heteropoly molybdophosphoric acid. Intermediate results are reported with units of mg/L and loading results with units of mg/dm2/d. Note that the air convention is to report as ion weight, ie as PO4 not as P. | 0.003 | Y | Y | Y | 40 |  |  | 3/19/2006 4:43:31 PM | 1998-02-17 | 6/3/1998 6:51:48 PM | NPEPPIN |  | Concentration | mg/L | Y | (fl. conc.) |  | Phosphorus Tot Soluble (fl. conc.) | 7723-14-0 | Nutrients; | Nutrients; EMS code: PP-S | Chemical | Soluble | LAB
PP-T | mg/dm2/d | Phosphorus Total A | CAL1 | Calculation (mg/dm2/d) | Calculation (mg/dm2/d) | 0.0003 | Y | Y | Y | 11 |  |  | 4/2/1998 8:38:10 AM | 1996-04-12 | 6/3/1998 6:51:42 PM | EMS |  | MassFlux | mg/dm²/d | Y | (mass flux) |  | Phosphorus Total A (mass flux) | 7723-14-0 | Nutrients; | Nutrients; EMS code: PP-T | Chemical | Total | LAB
PQ-8 | mg/L | Copper Quinolate | IC01 | 2ChlmethaneExt:IonChroma | 2ChlmethaneExt:IonChroma | 5 | Y | Y | Y | 4 |  |  | 10/24/1996 3:34:59 AM | 1996-04-12 | 6/3/1998 6:51:37 PM | EMS |  | Concentration | mg/L | Y | (fl. conc.) |  | Copper Quinolate (fl. conc.) | 10380-28-6 | Organometallics; | Organometallics; EMS code: PQ-8 | Chemical |  | LAB
PR-T | ppm (S) | Praseodymium | MS81 | ALS Minerals: ME-MS81 | ALS Minerals: ME-MS81. A prepared sample (0.100 g) is added to lithium metaborate/lithium tetraborate flux; mixed well and fused in a furnace at 1025°C. The resulting melt is then cooled and dissolved in an acid mixture containing nitric; hydrochloric and hydrofluoric acids. This solution is then analyzed by inductively coupled plasma - mass spectrometry. | 0.02 | Y | Y | Y | 0 |  |  |  | 2024-01-29 |  | EMS |  | DimensionlessRatio | ppm | N | (ratio) |  | Praseodymium (ratio) |  | Metals; | Metals; EMS code: PR-T | Chemical |  | LAB
PRED | m | Precipit'n Depth-Spl | 5066 | Calculation/Sample Volum | Calculation/Sample Volum | 0.0001 | Y | Y | Y | 6244 | 4/30/1997 3:50:00 PM | 4/3/2008 | 5/28/2008 2:08:45 AM | 1996-04-12 | 6/3/1998 6:51:37 PM | EMS |  | Length | m | Y | (len.) |  | Precipit'n Depth-Spl (len.) |  | Physical Properties; | Physical Properties; EMS code: PRED | Physical |  | LAB
PROH | mg/L | 1-Propanol, n-Propanol, CAS: 71-23-8 | E-8015 | EPA 8015D | EPA Method 8015D (SW-846): Nonhalogenated Organics Using GC/FID |  | Y | Y | Y | 14 | 12/29/2021 | 1/12/2022 | 9/2/2022 6:21:54 PM | 2022-04-06 |  | EMS |  | Concentration | mg/L | Y | (fl. conc.) |  | 1-Propanol, n-Propanol (fl. conc.) | 71-23-8 | Organics; | Organics; EMS code: PROH | Chemical |  | LAB
PROH | ug/g | 1-Propanol, n-Propanol, CAS: 71-23-8 | E8015S | EPA 8015 D in soil | See EPA 8015D | 10 | Y | Y | Y | 15 | 1/12/2022 | 2/1/2022 | 1/11/2024 1:07:08 PM | 2022-06-23 |  | EMS |  | MassPerMass | mg/kg | Y | (mass conc.) |  | 1-Propanol, n-Propanol (mass conc.) | 71-23-8 | Organics; | Organics; EMS code: PROH | Chemical |  | LAB
PRPA | mg/L | Propionic Acid | VFAD | Volatile Fatty Acids by GC/FID | Volatile Fatty Acids by GC/FID; Standard Methods for the Examination of Water and Wastewater; 21st Edition, Method 5560 D; pages 5-59 to 5-61. 5560 D. Gas Chromatographic Method Principle: This gas chromatographic procedure may be usedto determine the individual concentrations of the following fattyacids: acetic; propionic; butyric; isobutyric; valeric; andisovaleric acids in digester sludge samples. Caproic and heptanoicacids also may be determined; but with less certaintybecause of their reduced solubility in water. Volatile fatty acidsare generally soluble in digester sludge and exist in either theprotonated or unprotonated form; depending on the pH of thesludge sample. Acetic and propionic acids usually are in higherconcentrations than the other fatty acids in digester sludge. Allacids must be converted to the protonated (volatile) form beforeinjection into the gas chromatograph (GC) so vaporization mayoccur. The fatty acids must also be separated from the solidsmaterial in the sample before introduction into the GC to minimizedegradation of the GC column. Digester sludge samples areprepared by acidification; centrifugation; and filtration. The sampleis analyzed by direct injection into a gas chromatographequipped with a flame ionization detector after both centrifugationand filtration.b. Interference: The blank amount may be high relative tosample analyte. High blanks may be caused by a buildup ofcontaminants in the injector and guard column or by samplecarry-over. The analyst must be aware of the blank levels andcorrect unacceptably high blank levels that are above the reportinglimits.c. Minimum detectable concentration: The method detectionlevel has been determined by the USEPA method1 in reagentwater to be about 3 mg/L for acetic acid and about 1 mg/L for allother target compounds. | 2 | Y | Y | Y | 5 | 3/14/2017 10:30:00 AM | 11/19/2017 5:10:00 PM | 6/20/2018 9:02:11 AM | 2018-06-19 |  | RDGIBSON |  | Concentration | mg/L | Y | (fl. conc.) |  | Propionic Acid (fl. conc.) | 79-09-4 | Organics; | Organics; EMS code: PRPA | Chemical |  | LAB
PS04 | %(W/W) | Particle Size 2.00mm | FS22 | Dry Sieve 2.0mm | Dry Sieve 2mm | 0 | Y | Y | Y | 524 | 10/27/1998 1:11:00 PM | 8/29/2021 12:03:00 AM | 9/12/2021 3:13:22 PM | 1998-07-31 |  | NPEPPIN |  | DimensionlessRatio | % | N | (ratio) | (wet) | Particle Size 2.00mm (wet) (ratio) |  | Sediment Size; | Sediment Size; EMS code: PS04 | Physical |  | LAB
PS06 | %(W/W) | Particle Size 1.0mm | FS24 | Wet Sieve 1.0mm | Wet Sieve 1.0mm | 0 | Y | Y | Y | 58 | 5/29/2000 10:58:00 AM | 11/15/2001 8:02:00 AM | 11/26/2002 11:02:00 AM | 1998-07-31 |  | NPEPPIN |  | DimensionlessRatio | % | N | (ratio) | (wet) | Particle Size 1.0mm (wet) (ratio) |  | Sediment Size; | Sediment Size; EMS code: PS06 | Physical |  | LAB
PS08 | %(W/W) | Particle Size 0.500mm | FS26 | Wet sieve 0.500mm | Wet sieve 0.500mm | 0 | Y | Y | Y | 58 | 5/29/2000 10:58:00 AM | 11/15/2001 8:02:00 AM | 11/26/2002 11:02:00 AM | 1998-07-31 |  | NPEPPIN |  | DimensionlessRatio | % | N | (ratio) | (wet) | Particle Size 0.500mm (wet) (ratio) |  | Sediment Size; | Sediment Size; EMS code: PS08 | Physical |  | LAB
PS10 | %(W/W) | Particle Size 0.250mm | FS28 | Wet Sieve 0.250mm | Wet Sieve 0.250mm | 0 | Y | Y | Y | 508 | 10/27/1998 1:11:00 PM | 12/19/2001 12:25:00 PM | 5/26/2003 1:56:31 PM | 1998-07-31 |  | NPEPPIN |  | DimensionlessRatio | % | N | (ratio) | (wet) | Particle Size 0.250mm (wet) (ratio) |  | Sediment Size; | Sediment Size; EMS code: PS10 | Physical |  | LAB
PS12 | %(W/W) | Particle Size 0.125mm | FS30 | Wet Sieve 0.125mm | Wet Sieve 0.125mm | 0 | Y | Y | Y | 509 | 10/27/1998 1:11:00 PM | 12/19/2001 12:25:00 PM | 5/26/2003 1:56:32 PM | 1998-07-31 |  | NPEPPIN |  | DimensionlessRatio | % | N | (ratio) | (wet) | Particle Size 0.125mm (wet) (ratio) |  | Sediment Size; | Sediment Size; EMS code: PS12 | Physical |  | LAB
PS15 | %(W/W) | Particle Size 0.063mm | FS36 | Pipette 0.063mm | Pipette 0.063mm | 0 | Y | Y | Y | 64 | 10/27/1998 1:12:00 PM | 11/15/2001 8:02:00 AM | 11/26/2002 11:02:01 AM | 1998-07-31 |  | NPEPPIN |  | DimensionlessRatio | % | N | (ratio) | (wet) | Particle Size 0.063mm (wet) (ratio) |  | Sediment Size; | Sediment Size; EMS code: PS15 | Chemical |  | LAB
PS16 | %(W/W) | Particle Size 0.053mm | FS37 | Pipette 0.053MM | Pipette 0.053mm | 0 | Y | Y | Y | 445 | 9/7/2000 12:17:00 PM | 12/19/2001 12:25:00 PM | 5/26/2003 1:56:32 PM | 1998-07-31 |  | NPEPPIN |  | DimensionlessRatio | % | N | (ratio) | (wet) | Particle Size 0.053mm (wet) (ratio) |  | Sediment Size; | Sediment Size; EMS code: PS16 | Physical |  | LAB
PS18 | %(W/W) | Particle Size 0.004mm | FS38 | Pipette 0.004mm | Pipette 0.004mm | 0 | Y | Y | Y | 66 | 10/27/1998 1:12:00 PM | 11/15/2001 8:02:00 AM | 11/26/2002 11:02:02 AM | 1998-07-31 |  | NPEPPIN |  | DimensionlessRatio | % | N | (ratio) | (wet) | Particle Size 0.004mm (wet) (ratio) |  | Sediment Size; | Sediment Size; EMS code: PS18 | Chemical |  | LAB
PS19 | %(W/W) | Particle Size 0.002mm | FS39 | Pipette 0.002mm | Pipette 0.002mm | 0 | Y | Y | Y | 501 | 5/29/2000 10:58:00 AM | 12/19/2001 12:25:00 PM | 5/26/2003 1:56:32 PM | 1998-07-31 |  | NPEPPIN |  | DimensionlessRatio | % | N | (ratio) | (wet) | Particle Size 0.002mm (wet) (ratio) |  | Sediment Size; | Sediment Size; EMS code: PS19 | Physical |  | LAB
PS20 | %(W/W) | Particle Size >2.00mm | FS40 | Dry Sieve >2.00mm | Dry Sieve >2.00mm | 0 | Y | Y | Y | 533 | 10/27/1998 1:12:00 PM | 12/12/2019 | 12/20/2019 10:51:07 AM | 1998-07-31 |  | NPEPPIN |  | DimensionlessRatio | % | N | (ratio) | (wet) | Particle Size >2.00mm (wet) (ratio) |  | Sediment Size; | Sediment Size; EMS code: PS20 | Physical |  | LAB
PS20 | % | Particle Size >2.00mm | PSP | PSA Pipet Detail | Particle size distribution is determined by a combination of techniques. Dry sieving is performed for coarse particles, wet sieving for sand particles and the pipette sedimentation method for clay particles. Burt, R. (2009). Soil Survey Field and Laboratory Methods Manual. Soil Survey Investigations Report No. 5. Method 3.2.1.2.2. United States Department of Agriculture Natural Resources Conservation Service. | 1 | Y | Y | Y | 155 | 9/16/2017 | 9/22/2023 | 1/12/2024 11:16:52 AM | 2018-01-09 |  | JKROGH |  | DimensionlessRatio | % | N | (ratio) |  | Particle Size >2.00mm (ratio) |  | Sediment Size; | Sediment Size; EMS code: PS20 | Physical |  | LAB
PS20 | % | Particle Size >2.00mm | FSDS | Dry Sieve | Dry Sieve | 0.1 | Y | Y | Y | 7 | 9/18/2022 | 9/18/2022 | 3/9/2023 4:55:51 PM | 2023-02-28 |  | EMS |  | DimensionlessRatio | % | N | (ratio) |  | Particle Size >2.00mm (ratio) |  | Sediment Size; | Sediment Size; EMS code: PS20 | Physical |  | LAB
PS21 | %(W/W) | Particle Size <2.00mm->0.053mm | FS41 | Wet Sieve <2.00mm->0.053mm | Wet Sieve <2.00mm->0.053mm | 0 | Y | Y | Y | 469 | 9/7/2000 12:17:00 PM | 12/12/2019 | 12/20/2019 10:51:07 AM | 1998-07-31 |  | NPEPPIN |  | DimensionlessRatio | % | N | (ratio) | (wet) | Particle Size <2.00mm->0.053mm (wet) (ratio) |  | Sediment Size; | Sediment Size; EMS code: PS21 | Chemical |  | LAB
PS21 | % | Particle Size <2.00mm->0.053mm | PSP | PSA Pipet Detail | Particle size distribution is determined by a combination of techniques. Dry sieving is performed for coarse particles, wet sieving for sand particles and the pipette sedimentation method for clay particles. Burt, R. (2009). Soil Survey Field and Laboratory Methods Manual. Soil Survey Investigations Report No. 5. Method 3.2.1.2.2. United States Department of Agriculture Natural Resources Conservation Service. | 1 | Y | Y | Y | 120 | 3/17/2020 | 9/22/2023 | 1/12/2024 11:16:52 AM | 2018-01-10 |  | JKROGH |  | DimensionlessRatio | % | N | (ratio) |  | Particle Size <2.00mm->0.053mm (ratio) |  | Sediment Size; | Sediment Size; EMS code: PS21 | Chemical |  | LAB
PS22 | %(W/W) | Particle Size <0.063mm->0.004mm | FS43 | Pipette <0.063mm->0.004mm | Pipette <0.063mm->0.004mm | 0 | Y | Y | Y | 82 | 5/29/2000 10:58:00 AM | 12/12/2019 | 12/20/2019 10:51:07 AM | 1998-07-31 |  | NPEPPIN |  | DimensionlessRatio | % | N | (ratio) | (wet) | Particle Size <0.063mm->0.004mm (wet) (ratio) |  | Sediment Size; | Sediment Size; EMS code: PS22 | Chemical |  | LAB
PS22 | % | Particle Size <0.063mm->0.004mm | PSP | PSA Pipet Detail | Particle size distribution is determined by a combination of techniques. Dry sieving is performed for coarse particles, wet sieving for sand particles and the pipette sedimentation method for clay particles. Burt, R. (2009). Soil Survey Field and Laboratory Methods Manual. Soil Survey Investigations Report No. 5. Method 3.2.1.2.2. United States Department of Agriculture Natural Resources Conservation Service. | 1 | Y | Y | Y | 120 | 3/17/2020 | 9/22/2023 | 1/12/2024 11:16:52 AM | 2018-01-10 |  | JKROGH |  | DimensionlessRatio | % | N | (ratio) |  | Particle Size <0.063mm->0.004mm (ratio) |  | Sediment Size; | Sediment Size; EMS code: PS22 | Chemical |  | LAB
PS24 | %(W/W) | Particle Size <0.053mm - >0.002mm | FS45 | Pipette <0.053mm->0.002mm | Pipette <0.053mm->0.002mm | 0 | Y | Y | Y | 451 | 10/27/1998 1:18:00 PM | 12/19/2001 12:25:00 PM | 5/26/2003 1:56:33 PM | 1998-08-04 |  | NPEPPIN |  | DimensionlessRatio | % | N | (ratio) | (wet) | Particle Size <0.053mm - >0.002mm (wet) (ratio) |  | Sediment Size; | Sediment Size; EMS code: PS24 | Physical |  | LAB
PS26 | %(W/W) | Particle Size <0.002mm | FS48 | Pipette <0.002mm | Pipette <0.002mm | 0 | Y | Y | Y | 445 | 9/7/2000 12:17:00 PM | 12/19/2001 12:25:00 PM | 5/26/2003 1:56:34 PM | 1998-08-04 |  | NPEPPIN |  | DimensionlessRatio | % | N | (ratio) | (wet) | Particle Size <0.002mm (wet) (ratio) |  | Sediment Size; | Sediment Size; EMS code: PS26 | Physical |  | LAB
PS27 | %(W/W) | Particle Size>6.3 mm | FS49 | Dry Sieve 6.3mm | Dry Sieve 6.3mm | 0 | Y | Y | Y | 59 | 2/23/1998 8:57:00 AM | 2/23/1998 10:26:00 AM | 12/9/1998 6:12:50 PM | 1998-08-04 |  | NPEPPIN |  | DimensionlessRatio | % | N | (ratio) | (wet) | Particle Size>6.3 mm (wet) (ratio) |  | Sediment Size; | Sediment Size; EMS code: PS27 | Chemical |  | LAB
PS28 | %(W/W) | Particle Size<6.3mm->4mm | FS50 | Dry Sieve <6.3->4mm | Dry Sieve <6.3->4mm | 0 | Y | Y | Y | 59 | 2/23/1998 8:57:00 AM | 2/23/1998 10:25:00 AM | 12/9/1998 6:12:50 PM | 1998-08-04 |  | NPEPPIN |  | DimensionlessRatio | % | N | (ratio) | (wet) | Particle Size<6.3mm->4mm (wet) (ratio) |  | Sediment Size; | Sediment Size; EMS code: PS28 | Chemical |  | LAB
PS29 | %(W/W) | Particle Size <4mm->2.8mm | FS51 | Dry Sieve <4mm->2.8mm | Dry sieve 4mm->2.8mm | 0 | Y | Y | Y | 59 | 2/23/1998 8:57:00 AM | 2/23/1998 10:25:00 AM | 12/9/1998 6:12:50 PM | 1998-08-04 |  | NPEPPIN |  | DimensionlessRatio | % | N | (ratio) | (wet) | Particle Size <4mm->2.8mm (wet) (ratio) |  | Sediment Size; | Sediment Size; EMS code: PS29 | Chemical |  | LAB
PS30 | %(W/W) | Particle Size <2.8mm->2mm | FS52 | Dry Sieve <2.8mm->2mm | Dry Sieve <2.8mm->2mm | 0 | Y | Y | Y | 59 | 2/23/1998 8:57:00 AM | 2/23/1998 10:25:00 AM | 12/9/1998 6:12:50 PM | 1998-08-04 |  | NPEPPIN |  | DimensionlessRatio | % | N | (ratio) | (wet) | Particle Size <2.8mm->2mm (wet) (ratio) |  | Sediment Size; | Sediment Size; EMS code: PS30 | Chemical |  | LAB
PS31 | %(W/W) | Particle Size <2mm->0.50mm | FS53 | Wet Sieve <2mm->0.50mm | Wet Sieve <2mm->0.50mm | 0 | Y | Y | Y | 59 | 2/23/1998 8:57:00 AM | 2/23/1998 10:25:00 AM | 12/9/1998 6:12:50 PM | 1998-08-04 |  | NPEPPIN |  | DimensionlessRatio | % | N | (ratio) | (wet) | Particle Size <2mm->0.50mm (wet) (ratio) |  | Sediment Size; | Sediment Size; EMS code: PS31 | Chemical |  | LAB
PS32 | %(W/W) | Particle Size <0.50mm->0.25mm | FS54 | Wet Sieve <0.50mm->0.25mm | Wet Sieve <0.50mm->0.25mm | 0 | Y | Y | Y | 59 | 2/23/1998 8:57:00 AM | 2/23/1998 10:25:00 AM | 12/9/1998 6:12:50 PM | 1998-08-04 |  | NPEPPIN |  | DimensionlessRatio | % | N | (ratio) | (wet) | Particle Size <0.50mm->0.25mm (wet) (ratio) |  | Sediment Size; | Sediment Size; EMS code: PS32 | Chemical |  | LAB
PS34 | %(W/W) | Particle Size <0.25mm->0.125mm | FS56 | Wet Sieve <0.25mm->.125mm | Wet Sieve <0.25mm->125mm | 0 | Y | Y | Y | 59 | 2/23/1998 9:22:00 AM | 2/23/1998 9:27:00 AM | 12/9/1998 6:12:50 PM | 1998-08-04 |  | NPEPPIN |  | DimensionlessRatio | % | N | (ratio) | (wet) | Particle Size <0.25mm->0.125mm (wet) (ratio) |  | Sediment Size; | Sediment Size; EMS code: PS34 | Chemical |  | LAB
PS36 | %(W/W) | Particle Size <0.125mm->0.063mm | FS58 | Wet Sieve <0.125mm->0.063mm | Wet Sieve <0.125mm->0.063mm | 0 | Y | Y | Y | 59 | 2/23/1998 9:27:00 AM | 2/23/1998 9:33:00 AM | 12/9/1998 6:12:50 PM | 1998-08-04 |  | NPEPPIN |  | DimensionlessRatio | % | N | (ratio) | (wet) | Particle Size <0.125mm->0.063mm (wet) (ratio) |  | Sediment Size; | Sediment Size; EMS code: PS36 | Chemical |  | LAB
PS37 | %(W/W) | Particle Size <0.063mm | FS59 | Wet Sieve <0.063mm | Wet Sieve <0.063mm | 0 | Y | Y | Y | 59 | 2/23/1998 8:57:00 AM | 5/11/1998 9:51:00 AM | 12/9/1998 6:12:50 PM | 1998-08-04 |  | NPEPPIN |  | DimensionlessRatio | % | N | (ratio) | (wet) | Particle Size <0.063mm (wet) (ratio) |  | Sediment Size; | Sediment Size; EMS code: PS37 | Chemical |  | LAB
PS38 | % | Particle Size <0.004mm | PSP | PSA Pipet Detail | Particle size distribution is determined by a combination of techniques. Dry sieving is performed for coarse particles, wet sieving for sand particles and the pipette sedimentation method for clay particles. Burt, R. (2009). Soil Survey Field and Laboratory Methods Manual. Soil Survey Investigations Report No. 5. Method 3.2.1.2.2. United States Department of Agriculture Natural Resources Conservation Service. | 1 | Y | Y | Y | 155 | 9/16/2017 | 9/22/2023 | 1/12/2024 11:16:52 AM | 2018-01-09 |  | JKROGH |  | DimensionlessRatio | % | N | (ratio) |  | Particle Size <0.004mm (ratio) |  | Sediment Size; | Sediment Size; EMS code: PS38 | Physical |  | LAB
PS38 | %(W/W) | Particle Size <0.004mm | FS47 | Pipette <0.004mm | Pipette <0.004mm | 0 | Y | Y | Y | 88 | 10/27/1998 1:18:00 PM | 12/12/2019 | 12/20/2019 10:51:07 AM | 1998-09-28 |  | NPEPPIN |  | DimensionlessRatio | % | N | (ratio) | (wet) | Particle Size <0.004mm (wet) (ratio) |  | Sediment Size; | Sediment Size; EMS code: PS38 | Physical |  | LAB
PS41 | % | Particle Size >4.75mm | FSDS | Dry Sieve | Dry Sieve | 0.1 | Y | Y | Y | 4 | 9/18/2022 | 9/18/2022 | 3/9/2023 4:55:50 PM | 2022-09-29 |  | EMS |  | DimensionlessRatio | % | N | (ratio) |  | Particle Size >4.75mm (ratio) |  | Sediment Size; | Sediment Size; EMS code: PS41 | Physical |  | LAB
PS45 | % | Particle Size >1.18mm | FSDS | Dry Sieve | Dry Sieve | 0.1 | Y | Y | Y | 7 | 9/18/2022 | 9/18/2022 | 3/9/2023 4:55:51 PM | 2022-09-29 |  | EMS |  | DimensionlessRatio | % | N | (ratio) |  | Particle Size >1.18mm (ratio) |  | Sediment Size; | Sediment Size; EMS code: PS45 | Physical |  | LAB
PS46 | % | Particle Size >0.15mm | FSDS | Dry Sieve | Dry Sieve | 0.1 | Y | Y | Y | 7 | 9/18/2022 | 9/18/2022 | 3/9/2023 4:55:51 PM | 2022-09-29 |  | EMS |  | DimensionlessRatio | % | N | (ratio) |  | Particle Size >0.15mm (ratio) |  | Sediment Size; | Sediment Size; EMS code: PS46 | Physical |  | LAB
PS47 | % | Particle Size >2.36mm | FSDS | Dry Sieve | Dry Sieve | 0.1 | Y | Y | Y | 7 | 9/18/2022 | 9/18/2022 | 3/9/2023 4:55:51 PM | 2022-09-29 |  | EMS |  | DimensionlessRatio | % | N | (ratio) |  | Particle Size >2.36mm (ratio) |  | Sediment Size; | Sediment Size; EMS code: PS47 | Physical |  | LAB
PS48 | % | Particle Size >0.3mm | FSDS | Dry Sieve | Dry Sieve | 0.1 | Y | Y | Y | 7 | 9/18/2022 | 9/18/2022 | 3/9/2023 4:55:51 PM | 2022-09-29 |  | EMS |  | DimensionlessRatio | % | N | (ratio) |  | Particle Size >0.3mm (ratio) |  | Sediment Size; | Sediment Size; EMS code: PS48 | Physical |  | LAB
PS49 | % | Particle Size >0.425mm | FSDS | Dry Sieve | Dry Sieve | 0.1 | Y | Y | Y | 7 | 9/18/2022 | 9/18/2022 | 3/9/2023 4:55:51 PM | 2022-09-29 |  | EMS |  | DimensionlessRatio | % | N | (ratio) |  | Particle Size >0.425mm (ratio) |  | Sediment Size; | Sediment Size; EMS code: PS49 | Physical |  | LAB
PS50 | % | Particle Size >0.6mm | FSDS | Dry Sieve | Dry Sieve | 0.1 | Y | Y | Y | 7 | 9/18/2022 | 9/18/2022 | 3/9/2023 4:55:51 PM | 2022-09-29 |  | EMS |  | DimensionlessRatio | % | N | (ratio) |  | Particle Size >0.6mm (ratio) |  | Sediment Size; | Sediment Size; EMS code: PS50 | Physical |  | LAB
PS51 | % | Passing (19mm) | PSP | PSA Pipet Detail | Particle size distribution is determined by a combination of techniques. Dry sieving is performed for coarse particles, wet sieving for sand particles and the pipette sedimentation method for clay particles. Burt, R. (2009). Soil Survey Field and Laboratory Methods Manual. Soil Survey Investigations Report No. 5. Method 3.2.1.2.2. United States Department of Agriculture Natural Resources Conservation Service. | 1 | Y | Y | Y | 70 | 9/21/2023 | 9/21/2023 | 1/12/2024 11:16:52 AM | 2024-01-08 | 1/8/2024 11:10:16 AM | EMS | EMS | DimensionlessRatio | % | N | (ratio) |  | Passing (19mm) (ratio) |  | Sediment Size; | Sediment Size; EMS code: PS51 | Physical |  | LAB
PS52 | % | Passing (2.0mm) | PSP | PSA Pipet Detail | Particle size distribution is determined by a combination of techniques. Dry sieving is performed for coarse particles, wet sieving for sand particles and the pipette sedimentation method for clay particles. Burt, R. (2009). Soil Survey Field and Laboratory Methods Manual. Soil Survey Investigations Report No. 5. Method 3.2.1.2.2. United States Department of Agriculture Natural Resources Conservation Service. | 1 | Y | Y | Y | 70 | 9/21/2023 | 9/21/2023 | 1/12/2024 11:16:52 AM | 2024-01-08 | 1/8/2024 11:10:16 AM | EMS | EMS | DimensionlessRatio | % | N | (ratio) |  | Passing (2.0mm) (ratio) |  | Sediment Size; | Sediment Size; EMS code: PS52 | Physical |  | LAB
PS53 | % | Passing (25.4mm) | PSP | PSA Pipet Detail | Particle size distribution is determined by a combination of techniques. Dry sieving is performed for coarse particles, wet sieving for sand particles and the pipette sedimentation method for clay particles. Burt, R. (2009). Soil Survey Field and Laboratory Methods Manual. Soil Survey Investigations Report No. 5. Method 3.2.1.2.2. United States Department of Agriculture Natural Resources Conservation Service. | 1 | Y | Y | Y | 70 | 9/21/2023 | 9/21/2023 | 1/12/2024 11:16:52 AM | 2024-01-08 | 1/8/2024 11:10:16 AM | EMS | EMS | DimensionlessRatio | % | N | (ratio) |  | Passing (25.4mm) (ratio) |  | Sediment Size; | Sediment Size; EMS code: PS53 | Physical |  | LAB
PS54 | % | Passing (38.1mm) | PSP | PSA Pipet Detail | Particle size distribution is determined by a combination of techniques. Dry sieving is performed for coarse particles, wet sieving for sand particles and the pipette sedimentation method for clay particles. Burt, R. (2009). Soil Survey Field and Laboratory Methods Manual. Soil Survey Investigations Report No. 5. Method 3.2.1.2.2. United States Department of Agriculture Natural Resources Conservation Service. | 1 | Y | Y | Y | 70 | 9/21/2023 | 9/21/2023 | 1/12/2024 11:16:52 AM | 2024-01-08 | 1/8/2024 11:10:16 AM | EMS | EMS | DimensionlessRatio | % | N | (ratio) |  | Passing (38.1mm) (ratio) |  | Sediment Size; | Sediment Size; EMS code: PS54 | Physical |  | LAB
PS55 | % | Passing (4.75mm) | PSP | PSA Pipet Detail | Particle size distribution is determined by a combination of techniques. Dry sieving is performed for coarse particles, wet sieving for sand particles and the pipette sedimentation method for clay particles. Burt, R. (2009). Soil Survey Field and Laboratory Methods Manual. Soil Survey Investigations Report No. 5. Method 3.2.1.2.2. United States Department of Agriculture Natural Resources Conservation Service. | 1 | Y | Y | Y | 70 | 9/21/2023 | 9/21/2023 | 1/12/2024 11:16:52 AM | 2024-01-08 | 1/8/2024 11:10:16 AM | EMS | EMS | DimensionlessRatio | % | N | (ratio) |  | Passing (4.75mm) (ratio) |  | Sediment Size; | Sediment Size; EMS code: PS55 | Physical |  | LAB
PS56 | % | Passing (50.8mm) | PSP | PSA Pipet Detail | Particle size distribution is determined by a combination of techniques. Dry sieving is performed for coarse particles, wet sieving for sand particles and the pipette sedimentation method for clay particles. Burt, R. (2009). Soil Survey Field and Laboratory Methods Manual. Soil Survey Investigations Report No. 5. Method 3.2.1.2.2. United States Department of Agriculture Natural Resources Conservation Service. | 1 | Y | Y | Y | 70 | 9/21/2023 | 9/21/2023 | 1/12/2024 11:16:52 AM | 2024-01-08 | 1/8/2024 11:10:16 AM | EMS | EMS | DimensionlessRatio | % | N | (ratio) |  | Passing (50.8mm) (ratio) |  | Sediment Size; | Sediment Size; EMS code: PS56 | Physical |  | LAB
PS57 | % | Passing (76.2mm) | PSP | PSA Pipet Detail | Particle size distribution is determined by a combination of techniques. Dry sieving is performed for coarse particles, wet sieving for sand particles and the pipette sedimentation method for clay particles. Burt, R. (2009). Soil Survey Field and Laboratory Methods Manual. Soil Survey Investigations Report No. 5. Method 3.2.1.2.2. United States Department of Agriculture Natural Resources Conservation Service. | 1 | Y | Y | Y | 70 | 9/21/2023 | 9/21/2023 | 1/12/2024 11:16:52 AM | 2024-01-08 | 1/8/2024 11:10:16 AM | EMS | EMS | DimensionlessRatio | % | N | (ratio) |  | Passing (76.2mm) (ratio) |  | Sediment Size; | Sediment Size; EMS code: PS57 | Physical |  | LAB
PS58 | % | Passing (9.5mm) | PSP | PSA Pipet Detail | Particle size distribution is determined by a combination of techniques. Dry sieving is performed for coarse particles, wet sieving for sand particles and the pipette sedimentation method for clay particles. Burt, R. (2009). Soil Survey Field and Laboratory Methods Manual. Soil Survey Investigations Report No. 5. Method 3.2.1.2.2. United States Department of Agriculture Natural Resources Conservation Service. | 1 | Y | Y | Y | 70 | 9/21/2023 | 9/21/2023 | 1/12/2024 11:16:52 AM | 2024-01-08 | 1/8/2024 11:10:16 AM | EMS | EMS | DimensionlessRatio | % | N | (ratio) |  | Passing (9.5mm) (ratio) |  | Sediment Size; | Sediment Size; EMS code: PS58 | Physical |  | LAB
PS59 | % | Passing (0.05mm) | PSP | PSA Pipet Detail | Particle size distribution is determined by a combination of techniques. Dry sieving is performed for coarse particles, wet sieving for sand particles and the pipette sedimentation method for clay particles. Burt, R. (2009). Soil Survey Field and Laboratory Methods Manual. Soil Survey Investigations Report No. 5. Method 3.2.1.2.2. United States Department of Agriculture Natural Resources Conservation Service. | 1 | Y | Y | Y | 70 | 9/21/2023 | 9/21/2023 | 1/12/2024 11:16:52 AM | 2024-01-08 | 1/8/2024 11:10:16 AM | EMS | EMS | DimensionlessRatio | % | N | (ratio) |  | Passing (0.05mm) (ratio) |  | Sediment Size; | Sediment Size; EMS code: PS59 | Physical |  | LAB
PS60 | % | Passing (0.063mm) | PSP | PSA Pipet Detail | Particle size distribution is determined by a combination of techniques. Dry sieving is performed for coarse particles, wet sieving for sand particles and the pipette sedimentation method for clay particles. Burt, R. (2009). Soil Survey Field and Laboratory Methods Manual. Soil Survey Investigations Report No. 5. Method 3.2.1.2.2. United States Department of Agriculture Natural Resources Conservation Service. | 1 | Y | Y | Y | 70 | 9/21/2023 | 9/21/2023 | 1/12/2024 11:16:52 AM | 2024-01-08 | 1/8/2024 11:10:16 AM | EMS | EMS | DimensionlessRatio | % | N | (ratio) |  | Passing (0.063mm) (ratio) |  | Sediment Size; | Sediment Size; EMS code: PS60 | Physical |  | LAB
PS61 | % | Passing (0.075mm) | PSP | PSA Pipet Detail | Particle size distribution is determined by a combination of techniques. Dry sieving is performed for coarse particles, wet sieving for sand particles and the pipette sedimentation method for clay particles. Burt, R. (2009). Soil Survey Field and Laboratory Methods Manual. Soil Survey Investigations Report No. 5. Method 3.2.1.2.2. United States Department of Agriculture Natural Resources Conservation Service. | 1 | Y | Y | Y | 70 | 9/21/2023 | 9/21/2023 | 1/12/2024 11:16:52 AM | 2024-01-08 | 1/8/2024 11:10:16 AM | EMS | EMS | DimensionlessRatio | % | N | (ratio) |  | Passing (0.075mm) (ratio) |  | Sediment Size; | Sediment Size; EMS code: PS61 | Physical |  | LAB
PS62 | % | Passing (0.125mm) | PSP | PSA Pipet Detail | Particle size distribution is determined by a combination of techniques. Dry sieving is performed for coarse particles, wet sieving for sand particles and the pipette sedimentation method for clay particles. Burt, R. (2009). Soil Survey Field and Laboratory Methods Manual. Soil Survey Investigations Report No. 5. Method 3.2.1.2.2. United States Department of Agriculture Natural Resources Conservation Service. | 1 | Y | Y | Y | 70 | 9/21/2023 | 9/21/2023 | 1/12/2024 11:16:52 AM | 2024-01-08 | 1/8/2024 11:10:16 AM | EMS | EMS | DimensionlessRatio | % | N | (ratio) |  | Passing (0.125mm) (ratio) |  | Sediment Size; | Sediment Size; EMS code: PS62 | Physical |  | LAB
PS63 | % | Passing (0.149mm) | PSP | PSA Pipet Detail | Particle size distribution is determined by a combination of techniques. Dry sieving is performed for coarse particles, wet sieving for sand particles and the pipette sedimentation method for clay particles. Burt, R. (2009). Soil Survey Field and Laboratory Methods Manual. Soil Survey Investigations Report No. 5. Method 3.2.1.2.2. United States Department of Agriculture Natural Resources Conservation Service. | 1 | Y | Y | Y | 70 | 9/21/2023 | 9/21/2023 | 1/12/2024 11:16:52 AM | 2024-01-08 | 1/8/2024 11:10:16 AM | EMS | EMS | DimensionlessRatio | % | N | (ratio) |  | Passing (0.149mm) (ratio) |  | Sediment Size; | Sediment Size; EMS code: PS63 | Physical |  | LAB
PS64 | % | Passing (0.250mm) | PSP | PSA Pipet Detail | Particle size distribution is determined by a combination of techniques. Dry sieving is performed for coarse particles, wet sieving for sand particles and the pipette sedimentation method for clay particles. Burt, R. (2009). Soil Survey Field and Laboratory Methods Manual. Soil Survey Investigations Report No. 5. Method 3.2.1.2.2. United States Department of Agriculture Natural Resources Conservation Service. | 1 | Y | Y | Y | 70 | 9/21/2023 | 9/21/2023 | 1/12/2024 11:16:52 AM | 2024-01-08 | 1/8/2024 11:10:16 AM | EMS | EMS | DimensionlessRatio | % | N | (ratio) |  | Passing (0.250mm) (ratio) |  | Sediment Size; | Sediment Size; EMS code: PS64 | Physical |  | LAB
PS65 | % | Passing (0.420mm) | PSP | PSA Pipet Detail | Particle size distribution is determined by a combination of techniques. Dry sieving is performed for coarse particles, wet sieving for sand particles and the pipette sedimentation method for clay particles. Burt, R. (2009). Soil Survey Field and Laboratory Methods Manual. Soil Survey Investigations Report No. 5. Method 3.2.1.2.2. United States Department of Agriculture Natural Resources Conservation Service. | 1 | Y | Y | Y | 70 | 9/21/2023 | 9/21/2023 | 1/12/2024 11:16:52 AM | 2024-01-08 | 1/8/2024 11:10:16 AM | EMS | EMS | DimensionlessRatio | % | N | (ratio) |  | Passing (0.420mm) (ratio) |  | Sediment Size; | Sediment Size; EMS code: PS65 | Physical |  | LAB
PS66 | % | Passing (0.50mm) | PSP | PSA Pipet Detail | Particle size distribution is determined by a combination of techniques. Dry sieving is performed for coarse particles, wet sieving for sand particles and the pipette sedimentation method for clay particles. Burt, R. (2009). Soil Survey Field and Laboratory Methods Manual. Soil Survey Investigations Report No. 5. Method 3.2.1.2.2. United States Department of Agriculture Natural Resources Conservation Service. | 1 | Y | Y | Y | 70 | 9/21/2023 | 9/21/2023 | 1/12/2024 11:16:52 AM | 2024-01-08 | 1/8/2024 11:10:16 AM | EMS | EMS | DimensionlessRatio | % | N | (ratio) |  | Passing (0.50mm) (ratio) |  | Sediment Size; | Sediment Size; EMS code: PS66 | Physical |  | LAB
PS67 | % | Passing (0.841mm) | PSP | PSA Pipet Detail | Particle size distribution is determined by a combination of techniques. Dry sieving is performed for coarse particles, wet sieving for sand particles and the pipette sedimentation method for clay particles. Burt, R. (2009). Soil Survey Field and Laboratory Methods Manual. Soil Survey Investigations Report No. 5. Method 3.2.1.2.2. United States Department of Agriculture Natural Resources Conservation Service. | 1 | Y | Y | Y | 70 | 9/21/2023 | 9/21/2023 | 1/12/2024 11:16:52 AM | 2024-01-08 | 1/8/2024 11:10:16 AM | EMS | EMS | DimensionlessRatio | % | N | (ratio) |  | Passing (0.841mm) (ratio) |  | Sediment Size; | Sediment Size; EMS code: PS67 | Physical |  | LAB
PS68 | % | Passing (1.0mm) | PSP | PSA Pipet Detail | Particle size distribution is determined by a combination of techniques. Dry sieving is performed for coarse particles, wet sieving for sand particles and the pipette sedimentation method for clay particles. Burt, R. (2009). Soil Survey Field and Laboratory Methods Manual. Soil Survey Investigations Report No. 5. Method 3.2.1.2.2. United States Department of Agriculture Natural Resources Conservation Service. | 1 | Y | Y | Y | 70 | 9/21/2023 | 9/21/2023 | 1/12/2024 11:16:52 AM | 2024-01-08 | 1/8/2024 11:10:16 AM | EMS | EMS | DimensionlessRatio | % | N | (ratio) |  | Passing (1.0mm) (ratio) |  | Sediment Size; | Sediment Size; EMS code: PS68 | Physical |  | LAB
PS69 | % | Passing (0.002mm) | PSP | PSA Pipet Detail | Particle size distribution is determined by a combination of techniques. Dry sieving is performed for coarse particles, wet sieving for sand particles and the pipette sedimentation method for clay particles. Burt, R. (2009). Soil Survey Field and Laboratory Methods Manual. Soil Survey Investigations Report No. 5. Method 3.2.1.2.2. United States Department of Agriculture Natural Resources Conservation Service. | 1 | Y | Y | Y | 70 | 9/21/2023 | 9/21/2023 | 1/12/2024 11:16:52 AM | 2024-01-08 | 1/8/2024 11:10:16 AM | EMS | EMS | DimensionlessRatio | % | N | (ratio) |  | Passing (0.002mm) (ratio) |  | Sediment Size; | Sediment Size; EMS code: PS69 | Physical |  | LAB
PS70 | % | Passing (0.004mm) | PSP | PSA Pipet Detail | Particle size distribution is determined by a combination of techniques. Dry sieving is performed for coarse particles, wet sieving for sand particles and the pipette sedimentation method for clay particles. Burt, R. (2009). Soil Survey Field and Laboratory Methods Manual. Soil Survey Investigations Report No. 5. Method 3.2.1.2.2. United States Department of Agriculture Natural Resources Conservation Service. | 1 | Y | Y | Y | 70 | 9/21/2023 | 9/21/2023 | 1/12/2024 11:16:52 AM | 2024-01-08 | 1/8/2024 11:10:16 AM | EMS | EMS | DimensionlessRatio | % | N | (ratio) |  | Passing (0.004mm) (ratio) |  | Sediment Size; | Sediment Size; EMS code: PS70 | Physical |  | LAB
PS71 | % | Passing (0.005mm) | PSP | PSA Pipet Detail | Particle size distribution is determined by a combination of techniques. Dry sieving is performed for coarse particles, wet sieving for sand particles and the pipette sedimentation method for clay particles. Burt, R. (2009). Soil Survey Field and Laboratory Methods Manual. Soil Survey Investigations Report No. 5. Method 3.2.1.2.2. United States Department of Agriculture Natural Resources Conservation Service. | 1 | Y | Y | Y | 70 | 9/21/2023 | 9/21/2023 | 1/12/2024 11:16:52 AM | 2024-01-08 | 1/8/2024 11:10:16 AM | EMS | EMS | DimensionlessRatio | % | N | (ratio) |  | Passing (0.005mm) (ratio) |  | Sediment Size; | Sediment Size; EMS code: PS71 | Physical |  | LAB
PS72 | % | Passing (0.020mm) | PSP | PSA Pipet Detail | Particle size distribution is determined by a combination of techniques. Dry sieving is performed for coarse particles, wet sieving for sand particles and the pipette sedimentation method for clay particles. Burt, R. (2009). Soil Survey Field and Laboratory Methods Manual. Soil Survey Investigations Report No. 5. Method 3.2.1.2.2. United States Department of Agriculture Natural Resources Conservation Service. | 1 | Y | Y | Y | 70 | 9/21/2023 | 9/21/2023 | 1/12/2024 11:16:52 AM | 2024-01-08 | 1/8/2024 11:10:16 AM | EMS | EMS | DimensionlessRatio | % | N | (ratio) |  | Passing (0.020mm) (ratio) |  | Sediment Size; | Sediment Size; EMS code: PS72 | Physical |  | LAB
PS73 | % | Passing (0.0312mm) | PSP | PSA Pipet Detail | Particle size distribution is determined by a combination of techniques. Dry sieving is performed for coarse particles, wet sieving for sand particles and the pipette sedimentation method for clay particles. Burt, R. (2009). Soil Survey Field and Laboratory Methods Manual. Soil Survey Investigations Report No. 5. Method 3.2.1.2.2. United States Department of Agriculture Natural Resources Conservation Service. | 1 | Y | Y | Y | 70 | 9/21/2023 | 9/21/2023 | 1/12/2024 11:16:52 AM | 2024-01-08 | 1/8/2024 11:10:16 AM | EMS | EMS | DimensionlessRatio | % | N | (ratio) |  | Passing (0.0312mm) (ratio) |  | Sediment Size; | Sediment Size; EMS code: PS73 | Physical |  | LAB
PSEU | CFU/100mL | Pseudomonas aeruginosa | X022 | Membrane Filter Count | Non-turbid water samples are passed through a 0.45 um membrane filter which is placed on a selective medium for 18 - 24 hours incubation at 35 degrees C for growth of total coliform colonies, of fecal coliform colonies, fecal streptococci colonies, enterococci colonies, or of E. coli colonies. | 2 | Y | Y | Y | 619 | 7/2/2002 | 5/28/2008 | 2/3/2011 11:45:13 AM | 1996-04-12 | 6/3/1998 6:51:37 PM | EMS |  | MicrobialFormation | CFU/100mL | N | (microb.) |  | Pseudomonas aeruginosa (microb.) | 68583-29-9 | Bacteriological; | Bacteriological; EMS code: PSEU | Chemical |  | LAB
PSTA | None | Particle Size - Textural Analysis | FS65 | USDA Class. Scheme | USDA Classification Scheme |  | Y | Y | Y | 468 | 9/7/2000 12:17:00 PM | 12/12/2019 | 12/20/2019 10:51:07 AM | 1999-04-08 |  | NPEPPIN |  | None | None | N | (unitless) |  | Particle Size - Textural Analysis (unitless) |  | Sediment Size; | Sediment Size; EMS code: PSTA | Physical |  | LAB
PSTA | None | Particle Size - Textural Analysis | HYDR | Hydrometer Analysis | Hydrometer Analysis |  | Y | Y | Y | 12 | 9/22/2022 | 11/1/2022 | 3/9/2023 4:55:51 PM | 2022-06-30 |  | EMS |  | None | None | N | (unitless) |  | Particle Size - Textural Analysis (unitless) |  | Sediment Size; | Sediment Size; EMS code: PSTA | Physical |  | LAB
PSTA | None | Particle Size - Textural Analysis | PSP | PSA Pipet Detail | Particle size distribution is determined by a combination of techniques. Dry sieving is performed for coarse particles, wet sieving for sand particles and the pipette sedimentation method for clay particles. Burt, R. (2009). Soil Survey Field and Laboratory Methods Manual. Soil Survey Investigations Report No. 5. Method 3.2.1.2.2. United States Department of Agriculture Natural Resources Conservation Service. |  | Y | Y | Y | 35 | 9/16/2017 | 10/4/2018 | 10/10/2018 2:46:10 PM | 2018-01-10 |  | JKROGH |  | None | None | N | (unitless) |  | Particle Size - Textural Analysis (unitless) |  | Sediment Size; | Sediment Size; EMS code: PSTA | Physical |  | LAB
PT-D | mg/L | Platinum Dissolved | X556 | Metals in water, dissolved, "in-bottle" digestion, ICP-SFMS | The whole water sample is filtered through a 0.45 micron cellulose acetate filter paper and the filtrate is stored in a HDPE/Teflon bottle. The container is de-capped in a Class 100 workstation, acidified to 2% (v/v) HNO3 and 1% (v/v) HCl and recapped. The sample is "in-bottle" digested in a convection oven, set at 60 degrees C for 16 hours. The sample is cooled and introduced to an inductively coupled argon plasma-sector field mass spectrometer (ICP-SFMS) by pneumatic nebulization. Each respective element is measured at a specific mass to charge ratio, m/z value expressed in amu (atomic mass units), and at a specified intrument resolution. LR: Low Resolution ~300, MR: Mid Resolution ~4000 and HR: High Resolution ~10000. Ref: Method number 02-2003, National Laboratory for Environmental Testing, CCIW, Environment Canada, Burlington, Requested April 2003. |  | Y | Y | Y | 4158 | 6/27/2004 | 12/6/2023 2:04:00 AM | 2/14/2024 9:05:32 AM | 2008-06-12 |  | RDGIBSON |  | Concentration | mg/L | Y | (fl. conc.) |  | Platinum Dissolved (fl. conc.) | 7440-06-4 | Metals - Dissolved; | Metals - Dissolved; EMS code: PT-D | Chemical | Dissolved | LAB
PT-E | mg/L | Platinum Extractable | X553 | Metals in water, extractable, ICP-SFMS | The whole water sample is collected in a HDPE/Teflon bottle. In a Class 100 workstation it is preserved with concentrated Nitric Acid at concentration of 0.2 percent (v/v). The acid preservative extracts the metals from material in the sample and a decanted portion is introduced directly to an inductively coupled argon plasma-sector field mass spectrometer (ICP-SFMS) by pneumatic nebulization. Each respective element is measured at a specific mass to charge ratio, m/z value expressed in amu (atomic mass units), and at a specified instrument resolution. LR: Low Resolution ~300, MR: Mid Resolution ~4000 and HR: High Resolution ~10000. Ref: Method number 02-2061, National Laboratory for Environmental Testing, CCIW, Environment Canada, Burlington. Requested April 2003. | 0.001 | Y | Y | Y | 1715 | 6/27/2004 | 9/21/2007 | 4/18/2018 9:28:02 AM | 2009-05-14 | 5/15/2009 11:25:44 AM | RDGIBSON | RDGIBSON | Concentration | mg/L | Y | (fl. conc.) |  | Platinum Extractable (fl. conc.) | 7440-06-4 | Metals - Extractable; | Metals - Extractable; EMS code: PT-E | Chemical | Extractable | LAB
PT-T | ppm (S) | Platinum Total | 6041 | Aqua Regia by ICPMS | These were sediment/soils digested with Aqua-Regia (Nitric Acid & Hydrochloric Acids) and then analysis by ICPMS. Results would be reported in mg/kg and/or %. | 0.002 | Y | Y | Y | 21 | 10/13/2019 | 8/5/2020 | 8/10/2020 3:52:06 PM | 2019-09-06 | 10/21/2019 11:56:00 AM | KATFRASE | KATFRASE | DimensionlessRatio | ppm | N | (ratio) |  | Platinum Total (ratio) | 7440-06-4 | Metals - Total; | Metals - Total; EMS code: PT-T | Chemical | Total | LAB
PT-T | ppm (S) | Platinum Total | 6042 | Aqua Regia by ICPMS | These were sediment/soils digested with Aqua-Regia (Nitric Acid & Hydrochloric Acids) and then analysis by ICPMS. Results would be reported in ppm s. | 0.002 | Y | Y | Y | 71 | 5/14/2022 | 7/27/2024 | 7/29/2024 9:13:08 AM | 2022-06-28 |  | EMS |  | DimensionlessRatio | ppm | N | (ratio) |  | Platinum Total (ratio) | 7440-06-4 | Metals - Total; | Metals - Total; EMS code: PT-T | Chemical | Total | LAB
PT-T | mg/L | Platinum Total | X555 | Metals in water, total, "in-bottle" digestion, ICP-SFMS | The whole water sample collected in a HDPE/Teflon bottle, is de-capped in a Class 100 workstation, acidified to 2% (v/v) HNO3 and 1% (v/v) HCl and recapped. The sample is "in-bottle" digested in a convection oven, set at 60 degrees C for 16 hours. The sample is cooled and introduced to an inductively coupled argon plasma-sector field mass spectrometer (ICP-SFMS) by pneumatic nebulization. Each respective element is measured at a specific mass to charge ratio, m/z value expressed in amu (atomic mass units), and at resolution. LR: Low Resolution ~300, MR: Mid Resolution ~4000 and HR: High a specified instrument Resolution ~10000. Ref: Method number 02-2003, National Laboratory for Environmental Testing, CCIW, Environment Canada, Burlington. Requested April 2003. | 1e-06 | Y | Y | Y | 15774 | 6/27/2004 | 12/6/2023 2:08:00 AM | 2/14/2024 9:05:32 AM | 2009-03-31 |  | JCHUDAK |  | Concentration | mg/L | Y | (fl. conc.) |  | Platinum Total (fl. conc.) | 7440-06-4 | Metals - Total; | Metals - Total; EMS code: PT-T | Chemical | Total | LAB
PYCS | mg/L | Pyraclostrobin | E751 | E751D Pesticides in Water by LC-MS/MS (POS1 List) | E751D Pesticides in Water by LC-MS/MS (POS1 List) (MOECC E3501) | 1e-05 | Y | Y | Y | 39 | 12/9/2021 | 12/31/2021 | 2/15/2022 9:50:06 AM | 2021-12-14 |  | EMS |  | Concentration | mg/L | Y | (fl. conc.) |  | Pyraclostrobin (fl. conc.) | 175013-18-0 | Pesticides; | Pesticides; EMS code: PYCS | Chemical |  | LAB
PYRE | % | Pyrene d10 | SURR | Surrogate Recovery | Surrogate Recovery | 30 | Y | Y | Y | 4 |  |  | 6/7/2024 5:30:18 PM | 1996-04-12 | 6/3/1998 6:51:38 PM | EMS |  | DimensionlessRatio | % | N | (ratio) |  | Pyrene d10 (ratio) | 1718-52-1 | Polycyclic Aromatic Hydrocarbons (PAHs); | Polycyclic Aromatic Hydrocarbons (PAHs); EMS code: PYRE | Chemical |  | SURROGATE_RESULT
Q001 | ug/g | Quintozene | ZN03 | DCM+Ac.Extr:GC/MS | DCM+Ac.Extr:GC/MS EPA 8270E (mod) | 0.05 | Y | Y | Y | 2 | 11/2/2022 12:53:00 AM | 11/2/2022 4:41:00 AM | 1/7/2023 3:12:11 PM | 1996-04-12 | 11/16/1999 2:27:40 PM | EMS | NPEPPIN | MassPerMass | mg/kg | Y | (mass conc.) |  | Quintozene (mass conc.) | 82-68-8 | Pesticides; | Pesticides; EMS code: Q001 | Chemical |  | LAB
Q001 | mg/L | Quintozene | ZN01 | DCM+Ac.Extr:GC/MS | DCM+Ac.Extr:GC/MS | 0.0005 | Y | Y | Y | 28 | 7/28/2005 7:03:00 AM | 3/29/2012 10:15:00 PM | 5/3/2012 9:27:52 AM | 1996-04-12 | 11/16/1999 4:06:28 PM | EMS | NPEPPIN | Concentration | mg/L | Y | (fl. conc.) |  | Quintozene (fl. conc.) | 82-68-8 | Pesticides; | Pesticides; EMS code: Q001 | Chemical |  | LAB
Q002 | mg/L | Quinalophos | ZN01 | DCM+Ac.Extr:GC/MS | DCM+Ac.Extr:GC/MS | 0.0003 | Y | Y | Y | 28 | 7/29/2005 7:10:00 AM | 3/29/2012 10:15:00 PM | 5/3/2012 9:27:52 AM | 1996-04-12 | 11/16/1999 4:06:28 PM | EMS | NPEPPIN | Concentration | mg/L | Y | (fl. conc.) |  | Quinalophos (fl. conc.) | 13593-03-8 | Pesticides; | Pesticides; EMS code: Q002 | Chemical |  | LAB
Q100 | mg/L | Quizalofop-ethyl | ZN01 | DCM+Ac.Extr:GC/MS | DCM+Ac.Extr:GC/MS | 0.5 | Y | Y | Y | 1 | 7/28/2005 7:03:00 AM | 7/28/2005 7:03:00 AM | 6/2/2009 12:01:11 PM | 2009-02-20 |  | JCHUDAK |  | Concentration | mg/L | Y | (fl. conc.) |  | Quizalofop-ethyl (fl. conc.) | 76578-14-8 | Pesticides; | Pesticides; EMS code: Q100 | Chemical |  | LAB
QUIN6 | mg/L | 6PPD-Quinone | UTORG | Ultra Trace Organics-LC/MS/MS | A water sample (whole bottle) is passed through a solid phase extraction (SPE) cartridge. The final extract is analyzed by LC/MS/MS | 1e-06 | Y | Y | Y | 33 | 10/4/2023 | 3/4/2024 | 4/12/2024 2:01:06 PM | 2023-12-13 |  | EMS |  | Concentration | mg/L | Y | (fl. conc.) |  | 6PPD-Quinone (fl. conc.) | 793-24-8 | Organics; | Organics; EMS code: QUIN6 | Chemical |  | LAB
RA-D | Bq/L | Radium Dissolved | EQ16 | EQUIS (Bq/L) | EQUIS (Bq/L) | 0.01 | N | Y | Y | 9 |  |  | 10/22/1996 8:45:51 PM | 1996-04-12 | 6/3/1998 6:51:38 PM | EMS |  | Radioactivity | Bq/L | Y | (rad./vol.) |  | Radium Dissolved (rad./vol.) | 7440-14-4 | Metals - Dissolved; | Metals - Dissolved; EMS code: RA-D | Chemical | Dissolved | LAB
RA-D | Bq/L | Radium Dissolved | 0560 | BaSO4 Co-ppt: A-Scin | BaSO4 Co-ppt: A-Scin | 0.01 | Y | Y | Y | 1 |  |  | 10/23/1996 8:20:17 PM | 1996-04-12 | 6/3/1998 6:51:38 PM | EMS |  | Radioactivity | Bq/L | Y | (rad./vol.) |  | Radium Dissolved (rad./vol.) | 7440-14-4 | Metals - Dissolved; | Metals - Dissolved; EMS code: RA-D | Chemical | Dissolved | LAB
RA-T | Bq/L | Radium Total | 0560 | BaSO4 Co-ppt: A-Scin | BaSO4 Co-ppt: A-Scin | 0.01 | Y | Y | Y | 5 |  |  | 10/24/1996 1:08:24 PM | 1996-04-12 | 6/3/1998 6:51:38 PM | EMS |  | Radioactivity | Bq/L | Y | (rad./vol.) |  | Radium Total (rad./vol.) | 7440-14-4 | Metals - Total; | Metals - Total; EMS code: RA-T | Chemical | Total | LAB
RA-T | Bq/L | Radium Total | EQ16 | EQUIS (Bq/L) | EQUIS (Bq/L) | 0.01 | N | Y | Y | 18 |  |  | 10/22/1996 8:45:51 PM | 1996-04-12 | 6/3/1998 6:51:38 PM | EMS |  | Radioactivity | Bq/L | Y | (rad./vol.) |  | Radium Total (rad./vol.) | 7440-14-4 | Metals - Total; | Metals - Total; EMS code: RA-T | Chemical | Total | LAB
RA26 | Bq/L | Radium-226 | X240 | Gamma Spectrum | Gamma Spectrum | 0.001 | Y | Y | Y | 1167 | 5/2/2007 | 8/23/2024 | 8/26/2024 3:51:07 PM | 1996-04-12 | 6/3/1998 6:51:38 PM | EMS |  | Radioactivity | Bq/L | Y | (rad./vol.) |  | Radium-226 (rad./vol.) | 13982-63-3 | Radionuclide; | Radionuclide; EMS code: RA26 | Chemical |  | LAB
RA26 | Bq/L | Radium-226 | 9031 | EPA 903.1 Ra-226 using Emanation | EPA 903.1 Ra-226 using Emanation | 0.01 | Y | Y | Y | 511 | 3/2/2015 | 8/31/2022 | 1/23/2024 11:13:56 AM | 2015-03-17 |  | RDGIBSON |  | Radioactivity | Bq/L | Y | (rad./vol.) |  | Radium-226 (rad./vol.) | 13982-63-3 | Radionuclide; | Radionuclide; EMS code: RA26 | Chemical |  | LAB
RAWW | mg/L | Resin Acids; Total (BC WWQG) | E677 | Resins & Fatty Acids are methylated then analyzed by GCMS | EPA 8270E; AB Env. 129; NCASI RA/FA-85.02 (mod) Resins & Fatty Acids are methylated then analyzed by GCMS. Total fatty acids is the sum of arachidic acid; behenic acid; lauric acid; lignoceric acid; linoleic acid; linolenic acid; myristic acid; oleic acid; palmitic acid; and stearic acid. Total resin acids is the sum of abietic acid; 12-chlorodehydroabietic acid; 14-chlorodehydroabietic acid; dehydroabietic acid; 12;14-dichlorodehydroabietic acid; isopimaric acid; levopimaric acid; neoabietic acid; palustric acid; pimaric acid and sandaracopimaric acid. Total resin acids (BC WWQG) is the sum of abietic acid; neoabietic acid; pimaric acid; isopimaric acid; and sandaracopimaric acid. | 0.0025 | Y | Y | Y | 12 | 11/28/2022 | 6/15/2024 | 7/8/2024 11:10:02 AM | 2021-12-30 |  | EMS |  | Concentration | mg/L | Y | (fl. conc.) |  | Resin Acids; Total (BC WWQG) (fl. conc.) |  |  | EMS code: RAWW | Chemical | Total | LAB
RB-D | mg/L | Rubidium Dissolved | METC | MET-D-NP-CCMS-VA | Ultra trace metals in water are analyzed by CRC ICPMS, based on US EPA Method 6020B (July 2014). This procedure is intended for pristine field-filtered acid-preserved water samples. The detection limits (LOR) for this test are based on lab instrumental analysis only, not including filtration. ALS-supplied field filtration equipment does not support these LOR. Therefore, because of the high probability of false positives due to filtration, it is strongly recommended that a filtration blank be analysed to aid in data interpretation. | 5e-06 | Y | Y | Y | 394 | 3/30/2020 | 4/9/2024 | 4/29/2024 4:53:06 PM | 2022-10-07 |  | EMS |  | Concentration | mg/L | Y | (fl. conc.) |  | Rubidium Dissolved (fl. conc.) | 7440-17-7 | Metals - Dissolved; | Metals - Dissolved; EMS code: RB-D | Chemical | Dissolved | LAB
RB-D | mg/L | Rubidium Dissolved | MWID | Metals in Water by ICPMS (digested) | Metals in Water by ICPMS (digested) Based on EPA 200.8 | 0.0002 | Y | Y | Y | 0 |  |  |  | 2023-10-10 |  | EMS |  | Concentration | mg/L | Y | (fl. conc.) |  | Rubidium Dissolved (fl. conc.) | 7440-17-7 | Metals - Dissolved; | Metals - Dissolved; EMS code: RB-D | Chemical | Dissolved | LAB
RB-D | mg/L | Rubidium Dissolved | F032 | Filt.,Acid,ICPMS | There is no acid digestion required for this analysis. A water sample is collected and preserved in 0.2 percent nitric acid. To analyse the extractable metals the sample is shaken, allowed to sit for at least twenty four hours, and a decanted portion is introduced to the spectrometer. When the dissolved metals are required, the samples must first be filtered through a 0.45 micron cellulose acetate filter paper and then followed by preservation in 0.2 percent nitric acid. It is important to note that different laboratories may have very different detection limits for some variables analyzed by ICPMS. | 0.001 | Y | Y | Y | 9441 | 4/3/2015 | 8/26/2024 | 8/26/2024 4:51:04 PM | 1997-10-20 | 6/22/1998 3:38:30 PM | NPEPPIN | NPEPPIN | Concentration | mg/L | Y | (fl. conc.) |  | Rubidium Dissolved (fl. conc.) | 7440-17-7 | Metals - Dissolved; | Metals - Dissolved; EMS code: RB-D | Chemical | Dissolved | LAB
RB-D | mg/L | Rubidium Dissolved | X562 | Dissolved metals: FF, FA, HNO3, HR-ICPMS | Ultra trace metals in water are analyzed by high resolution inductively coupled plasma mass spectrometry (HR-ICPMS), based on US EPA Method 200.8, (Oct. 1999). This procedure is intended for pristine field-filtered acid-preserved water samples. The detection limits (LOR) for this test are based on lab instrumental analysis only, not including filtration. ALS-supplied field filtration equipment does not support these LOR. Therefore, because of the high probability of false positives due to filtration, it is strongly recommended that a filtration blank be analyzed to aid in data interpretation. | 2e-05 | Y | Y | Y | 10 | 10/20/2015 | 10/24/2018 | 10/24/2018 12:26:13 PM | 2016-03-24 |  | RDGIBSON |  | Concentration | mg/L | Y | (fl. conc.) |  | Rubidium Dissolved (fl. conc.) | 7440-17-7 | Metals - Dissolved; | Metals - Dissolved; EMS code: RB-D | Chemical | Dissolved | LAB
RB-D | mg/L | Rubidium Dissolved | MSEA | Metals in Seawater (undigested) by (CRC) ICP-MS | This method describes the multi-elemental determination of trace elements by Collision/Reaction Cell (CRC) ICP-MS. The method measures ions produced by a radio-frequency inductively coupled plasma. Analyte species originating in a liquid are pneumatically nebulized and the resulting aerosol transported by argon gas into the plasma torch. The ions produced are entrained in the plasma gas and extracted, by means of a differentially pumped vacuum interface, into a mass spectrometer. The ions flow through a Collision/Reaction Cell (CRC) where depending on the mode of analysis either the ions passes through the cell unhindered (no gas - classical ICP-MS mode), or collisions/reactions chemistry takes place between the ions and a neutral gas such as hydrogen or helium. Any polyatomic ions which give interference on the analytes are removed. The ions are then sorted according to their mass-to-charge ratios by a quadrupole mass spectrometer. The ions are then sorted according to their mass-to-charge ratios by a quadrupole mass spectrometer. The ions transmitted through the quadrupole are quantified by a channel electron multiplier and the ion information processed by a data handling system. Although most polyatomic interference is removed, any additional interference must be assessed and valid corrections applied or the data flagged to indicate problems. Interference correction must include compensation for background ions contributed by the plasma gas, reagents, and constituents of the sample matrix. Instrument drift as well as suppressions or enhancements of instrument response caused by the sample matrix must be corrected for by the use of internal standards. | 0.005 | Y | Y | Y | 4 | 4/12/2023 | 4/29/2024 | 5/16/2024 2:00:14 PM | 2023-05-26 |  | EMS |  | Concentration | mg/L | Y | (fl. conc.) |  | Rubidium Dissolved (fl. conc.) | 7440-17-7 | Metals - Dissolved; | Metals - Dissolved; EMS code: RB-D | Chemical | Dissolved | LAB
RB-D | mg/L | Rubidium Dissolved | X556 | Metals in water, dissolved, "in-bottle" digestion, ICP-SFMS | The whole water sample is filtered through a 0.45 micron cellulose acetate filter paper and the filtrate is stored in a HDPE/Teflon bottle. The container is de-capped in a Class 100 workstation, acidified to 2% (v/v) HNO3 and 1% (v/v) HCl and recapped. The sample is "in-bottle" digested in a convection oven, set at 60 degrees C for 16 hours. The sample is cooled and introduced to an inductively coupled argon plasma-sector field mass spectrometer (ICP-SFMS) by pneumatic nebulization. Each respective element is measured at a specific mass to charge ratio, m/z value expressed in amu (atomic mass units), and at a specified intrument resolution. LR: Low Resolution ~300, MR: Mid Resolution ~4000 and HR: High Resolution ~10000. Ref: Method number 02-2003, National Laboratory for Environmental Testing, CCIW, Environment Canada, Burlington, Requested April 2003. | 1e-05 | Y | Y | Y | 4285 | 6/11/2003 | 12/6/2023 2:04:00 AM | 2/14/2024 9:05:32 AM | 2004-01-19 |  | UMULLER |  | Concentration | mg/L | Y | (fl. conc.) |  | Rubidium Dissolved (fl. conc.) | 7440-17-7 | Metals - Dissolved; | Metals - Dissolved; EMS code: RB-D | Chemical | Dissolved | LAB
RB-E | mg/L | Rubidium Extractable | F047 | 24hr HNO3 dil'n,ICP-MS | There is no acid digestion required for this analysis. A water sample is collected and preserved in 0.2 percent nitric acid. To analyse the extractable metals the sample is shaken, allowed to sit for at least twenty four hours, and a decanted portion is introduced to the spectrometer. When the dissolved metals are required, the samples must first be filtered through a 0.45 micron cellulose acetate filter paper and then followed by preservation in 0.2 percent nitric acid. It is important to note that different laboratories may have very different detection limits for some variables analyzed by ICP/MS. | 2e-06 | Y | Y | Y | 2676 | 4/16/1998 10:21:00 AM | 6/26/2003 | 9/11/2019 8:46:18 AM | 1997-10-20 | 6/3/1998 6:51:47 PM | NPEPPIN |  | Concentration | mg/L | Y | (fl. conc.) |  | Rubidium Extractable (fl. conc.) | 7440-17-7 | Metals - Extractable; | Metals - Extractable; EMS code: RB-E | Chemical | Extractable | LAB
RB-E | mg/L | Rubidium Extractable | X553 | Metals in water, extractable, ICP-SFMS | The whole water sample is collected in a HDPE/Teflon bottle. In a Class 100 workstation it is preserved with concentrated Nitric Acid at concentration of 0.2 percent (v/v). The acid preservative extracts the metals from material in the sample and a decanted portion is introduced directly to an inductively coupled argon plasma-sector field mass spectrometer (ICP-SFMS) by pneumatic nebulization. Each respective element is measured at a specific mass to charge ratio, m/z value expressed in amu (atomic mass units), and at a specified instrument resolution. LR: Low Resolution ~300, MR: Mid Resolution ~4000 and HR: High Resolution ~10000. Ref: Method number 02-2061, National Laboratory for Environmental Testing, CCIW, Environment Canada, Burlington. Requested April 2003. | 1e-05 | Y | Y | Y | 2583 | 5/22/2003 | 9/21/2007 | 4/18/2018 9:28:02 AM | 2004-01-19 |  | UMULLER |  | Concentration | mg/L | Y | (fl. conc.) |  | Rubidium Extractable (fl. conc.) | 7440-17-7 | Metals - Extractable; | Metals - Extractable; EMS code: RB-E | Chemical | Extractable | LAB
RB-T | ppm (S) | Rubidium Total | MS81 | ALS Minerals: ME-MS81 | ALS Minerals: ME-MS81. A prepared sample (0.100 g) is added to lithium metaborate/lithium tetraborate flux; mixed well and fused in a furnace at 1025°C. The resulting melt is then cooled and dissolved in an acid mixture containing nitric; hydrochloric and hydrofluoric acids. This solution is then analyzed by inductively coupled plasma - mass spectrometry. | 0.2 | Y | Y | Y | 0 |  |  |  | 2024-01-29 |  | EMS |  | DimensionlessRatio | ppm | N | (ratio) |  | Rubidium Total (ratio) | 7440-17-7 | Metals - Total; | Metals - Total; EMS code: RB-T | Chemical | Total | LAB
RB-T | mg/L | Rubidium Total | F082 | Aliquot,HNO3/HCL dig,ICPMS | Aliquot microwave digested with 3:1 HNO3/HCL. Analysis by ICPMS. For HiVol samples preliminary result are reported in mg/L and loading results are reported in ug/m3 | 0.0002 | Y | Y | Y | 9162 | 4/6/2015 | 8/23/2024 | 8/26/2024 3:51:07 PM | 2015-05-06 |  | RDGIBSON |  | Concentration | mg/L | Y | (fl. conc.) |  | Rubidium Total (fl. conc.) | 7440-17-7 | Metals - Total; | Metals - Total; EMS code: RB-T | Chemical | Total | LAB
RB-T | mg/L | Rubidium Total | MSEA | Metals in Seawater (undigested) by (CRC) ICP-MS | This method describes the multi-elemental determination of trace elements by Collision/Reaction Cell (CRC) ICP-MS. The method measures ions produced by a radio-frequency inductively coupled plasma. Analyte species originating in a liquid are pneumatically nebulized and the resulting aerosol transported by argon gas into the plasma torch. The ions produced are entrained in the plasma gas and extracted, by means of a differentially pumped vacuum interface, into a mass spectrometer. The ions flow through a Collision/Reaction Cell (CRC) where depending on the mode of analysis either the ions passes through the cell unhindered (no gas - classical ICP-MS mode), or collisions/reactions chemistry takes place between the ions and a neutral gas such as hydrogen or helium. Any polyatomic ions which give interference on the analytes are removed. The ions are then sorted according to their mass-to-charge ratios by a quadrupole mass spectrometer. The ions are then sorted according to their mass-to-charge ratios by a quadrupole mass spectrometer. The ions transmitted through the quadrupole are quantified by a channel electron multiplier and the ion information processed by a data handling system. Although most polyatomic interference is removed, any additional interference must be assessed and valid corrections applied or the data flagged to indicate problems. Interference correction must include compensation for background ions contributed by the plasma gas, reagents, and constituents of the sample matrix. Instrument drift as well as suppressions or enhancements of instrument response caused by the sample matrix must be corrected for by the use of internal standards. | 0.005 | Y | Y | Y | 20 | 4/13/2023 | 4/29/2024 | 6/7/2024 10:03:16 AM | 2023-05-26 |  | EMS |  | Concentration | mg/L | Y | (fl. conc.) |  | Rubidium Total (fl. conc.) | 7440-17-7 | Metals - Total; | Metals - Total; EMS code: RB-T | Chemical | Total | LAB
RB-T | ppm (S) | Rubidium Total | MS61 | ALS Minerals ME-MS61 | ALS Minerals ME-MS61. A prepared sample is digested with HNO3; HF; and HClO4 in a graphite heating block to near dryness. After cooling; the sample is boiled in diluted HCl. Bulk to volume with deionized water; mixed and analyzed by ICP-AES and ICP-MS for all elements. | 0.1 | Y | Y | Y | 0 |  |  |  | 2024-01-29 |  | EMS |  | DimensionlessRatio | ppm | N | (ratio) |  | Rubidium Total (ratio) | 7440-17-7 | Metals - Total; | Metals - Total; EMS code: RB-T | Chemical | Total | LAB
RB-T | ug/g | Rubidium Total | X551 | Metals in Biota (dry), total, acid digestion, ICP-SFMS | 0.1 gram of dried, homogenized biota is weighed in a microwave TFM vessel, and 8mL of Nitric Acid and 1mL of Hydrogen Peroxide are added. The sample is digested at 200 degrees C for 15 minutes in a closed vessel high pressure microwave oven. The sample is cooled, the volume adjusted to 100mL and introduced to an iductively coupled argon plasma-sector field mass spectrometer (ICP-SFMS) by pneumatic nebulization. Each respective element is measured at a specific mass to charge ratio, m/z value expressed in amu (atomic mass units), and at a specified instrument resolution. LR: Low Resolution ~300, MR: Mid Resolution ~4000 and HR: Hich Resolution ~10000. Ref: Method number 02-2705, National Laboratory for Environmental Testing, CCIW, Environment Canada, Burlington. Requested April 2003. | 0.01 | Y | Y | Y | 43 | 8/11/2015 | 10/30/2023 | 11/28/2023 2:21:33 PM | 2004-01-19 |  | UMULLER |  | MassPerMass | mg/kg | Y | (mass conc.) |  | Rubidium Total (mass conc.) | 7440-17-7 | Metals - Total; | Metals - Total; EMS code: RB-T | Chemical | Total | LAB
RB-T | ppm (S) | Rubidium Total | 6041 | Aqua Regia by ICPMS | These were sediment/soils digested with Aqua-Regia (Nitric Acid & Hydrochloric Acids) and then analysis by ICPMS. Results would be reported in mg/kg and/or %. | 0.005 | Y | Y | Y | 21 | 10/13/2019 | 8/5/2020 | 8/10/2020 3:52:06 PM | 2019-09-06 | 10/21/2019 11:56:08 AM | KATFRASE | KATFRASE | DimensionlessRatio | ppm | N | (ratio) |  | Rubidium Total (ratio) | 7440-17-7 | Metals - Total; | Metals - Total; EMS code: RB-T | Chemical | Total | LAB
RB-T | ppm (S) | Rubidium Total | MS41 | ALS Minerals ME-MS41 | ALS Minerals ME-MS41. A prepared sample is digested with aqua regia in a graphite heating block. After cooling; the resulting solution is diluted to with deionized water; mixed and analyzed by ICP-AES. Samples are then analyzed by ICP-MS for the remaining suite of elements. | 0.1 | Y | Y | Y | 0 |  |  |  | 2024-01-29 |  | EMS |  | DimensionlessRatio | ppm | N | (ratio) |  | Rubidium Total (ratio) | 7440-17-7 | Metals - Total; | Metals - Total; EMS code: RB-T | Chemical | Total | LAB
RB-T | mg/L | Rubidium Total | X561 | Total metals: FA, HNO3, HR-ICPMS (No Prep)Prep) | Ultra trace metals in water are analyzed by high resolution inductively coupled plasma mass spectrometry (HR-ICPMS), modified from US EPA 200.8, (Revision 5.5). The detection limits provided can only be met for undigested samples. This procedure is intended for colorless, non-turbid, acid-preserved water samples (i.e. pristine water samples), having turbidity < 1 NTU and no odor. Where turbidity exceeds 1 NTU, and/or the sample is colored and has an odor, results may be biased low compared to true Total Metals concentrations. ALS recommends that turbidity analysis be requested on samples submitted for this test to aid with interpretation of results. | 2e-05 | Y | Y | Y | 19 | 10/20/2015 | 10/24/2018 | 10/24/2018 12:26:13 PM | 2016-03-24 |  | RDGIBSON |  | Concentration | mg/L | Y | (fl. conc.) |  | Rubidium Total (fl. conc.) | 7440-17-7 | Metals - Total; | Metals - Total; EMS code: RB-T | Chemical | Total | LAB
RB-T | mg/L | Rubidium Total | 1349 | Tot Metals in Seawater by CRC ICPMS (BC) | Tot Metals in Seawater by CRC ICPMS (BC) | 0.005 | Y | Y | Y | 98 | 8/16/2019 | 12/5/2021 | 12/13/2021 12:34:16 PM | 2019-09-30 |  | KATFRASE |  | Concentration | mg/L | Y | (fl. conc.) |  | Rubidium Total (fl. conc.) | 7440-17-7 | Metals - Total; | Metals - Total; EMS code: RB-T | Chemical | Total | LAB
RB-T | ppm (S) | Rubidium Total | 6042 | Aqua Regia by ICPMS | These were sediment/soils digested with Aqua-Regia (Nitric Acid & Hydrochloric Acids) and then analysis by ICPMS. Results would be reported in ppm s. | 0.005 | Y | Y | Y | 71 | 5/14/2022 | 7/27/2024 | 7/29/2024 9:13:08 AM | 2022-06-28 |  | EMS |  | DimensionlessRatio | ppm | N | (ratio) |  | Rubidium Total (ratio) | 7440-17-7 | Metals - Total; | Metals - Total; EMS code: RB-T | Chemical | Total | LAB
RB-T | mg/L | Rubidium Total | METN | MET-T-NP-CCMS-VA | Ultra trace metals in water are analyzed by CRC ICPMS, based on US EPA Method 6020B (July 2014). The detection limits provided can only be met for undigested samples. This procedure is intended for colorless, non-turbid, acid-preserved water samples (i.e. pristine water samples), having turbidity < 1 NTU and no odor. Where turbidity exceeds 1 NTU, and/or the sample is colored and has an odor, results may be biased low compared to true Total Metals concentrations. ALS recommends that turbidity analysis be requested on samples submitted for this test to aid with interpretation of results. Where turbidity is <1NTU, undigested metals are equivalent to total metals concentrations. | 5e-06 | Y | Y | Y | 280 | 8/6/2019 | 4/4/2024 | 4/29/2024 4:53:06 PM | 2021-06-16 |  | EMS |  | Concentration | mg/L | Y | (fl. conc.) |  | Rubidium Total (fl. conc.) | 7440-17-7 | Metals - Total; | Metals - Total; EMS code: RB-T | Chemical | Total | LAB
RB-T | ug/g | Rubidium Total | MTID | MET-DRY-CCMS-VA | Metals in Tissue by CRC ICPMS (DRY) Methodology: This method is conducted following British Columbia Lab Manual method "Metals in Animal Tissue and Vegetation (Biota) - Prescriptive". Tissue samples are homogenized and sub-sampled prior to hotblock digestion with nitric and hydrochloric acids, in combination with addition of hydrogen peroxide. Instrumental analysis is by collision cell inductively coupled plasma - mass spectrometry (modified from EPA Method 6020A). Method Limitation: This method employs a strong acid/peroxide digestion, and is intended to provide a conservative estimate of bio-available metals. Near complete recoveries are achieved for most toxicologically important metals, but elements associated with recalcitrant minerals may be only partially recovered. | 0.05 | Y | Y | Y | 122 | 8/13/2014 | 9/8/2022 | 10/26/2022 7:14:15 AM | 2015-03-09 | 3/11/2015 4:12:05 PM | RDGIBSON | EMS | MassPerMass | mg/kg | Y | (mass conc.) |  | Rubidium Total (mass conc.) | 7440-17-7 | Metals - Total; | Metals - Total; EMS code: RB-T | Chemical | Total | LAB
RB-T | ug/g wet | Rubidium Total | BPSY | MET-WET-BPSY-CCMS-VA | Metals in Animal Tissue and Vegetation (Biota) - Prescriptive. EPA 200.3/6020B This method is conducted following British Columbia Lab Manual method "Metals in Animal Tissue and Vegetation (Biota) - Prescriptive". Tissue samples are digested with nitric and hydrochloric acids; in combination with addition of hydrogen peroxide. Instrumental analysis is by collision cell inductively coupled plasma - mass spectrometry (modified from EPA Method 6020B). | 0.01 | Y | Y | Y | 2 | 11/7/2019 | 11/7/2019 | 6/10/2021 1:57:10 PM | 2021-06-04 |  | EMS |  | MassPerMass | µg/g | Y | (mass conc.) | (wet) | Rubidium Total (wet) (mass conc.) | 7440-17-7 | Metals - Total; | Metals - Total; EMS code: RB-T | Chemical | Total | LAB
RB-T | mg/L | Rubidium Total | X555 | Metals in water, total, "in-bottle" digestion, ICP-SFMS | The whole water sample collected in a HDPE/Teflon bottle, is de-capped in a Class 100 workstation, acidified to 2% (v/v) HNO3 and 1% (v/v) HCl and recapped. The sample is "in-bottle" digested in a convection oven, set at 60 degrees C for 16 hours. The sample is cooled and introduced to an inductively coupled argon plasma-sector field mass spectrometer (ICP-SFMS) by pneumatic nebulization. Each respective element is measured at a specific mass to charge ratio, m/z value expressed in amu (atomic mass units), and at resolution. LR: Low Resolution ~300, MR: Mid Resolution ~4000 and HR: High a specified instrument Resolution ~10000. Ref: Method number 02-2003, National Laboratory for Environmental Testing, CCIW, Environment Canada, Burlington. Requested April 2003. | 1e-05 | Y | Y | Y | 17404 | 5/20/2003 | 12/6/2023 2:08:00 AM | 2/14/2024 9:05:32 AM | 2004-01-19 |  | UMULLER |  | Concentration | mg/L | Y | (fl. conc.) |  | Rubidium Total (fl. conc.) | 7440-17-7 | Metals - Total; | Metals - Total; EMS code: RB-T | Chemical | Total | LAB
RB-T | ug/g wet | Rubidium Total | MTIW | MET-WET-CCMS-MID-VA | Metals in Tissue by CRC ICPMS (WET) Methodology: This method is conducted following British Columbia Lab Manual method "Metals in Animal Tissue and Vegetation (Biota) - Prescriptive". Tissue samples are homogenized and sub-sampled prior to hotblock digestion with nitric and hydrochloric acids, in combination with addition of hydrogen peroxide. Instrumental analysis is by collision cell inductively coupled plasma - mass spectrometry (modified from EPA Method 6020A). Method Limitation: This method employs a strong acid/peroxide digestion, and is intended to provide a conservative estimate of bio-available metals. Near complete recoveries are achieved for most toxicologically important metals, but elements associated with recalcitrant minerals may be only partially recovered. | 0.01 | Y | Y | Y | 24 | 11/25/2018 | 12/6/2018 | 7/21/2021 2:55:08 PM | 2018-12-24 |  | EMS |  | MassPerMass | µg/g | Y | (mass conc.) | (wet) | Rubidium Total (wet) (mass conc.) | 7440-17-7 | Metals - Total; | Metals - Total; EMS code: RB-T | Chemical | Total | LAB
RB-T | ug/g wet | Rubidium Total | X552 | Metals in Biota (wet), total, acid digestion, ICP-SFMS | One gram of wet, homogenized biota is weighed in a microwave TFM vessel, and 8mL of Nitric Acid and 1mL of Hydrogen Peroxide are added. The sample is digested at 200 degrees C for 15 minutes in a closed vessel high pressure microwave oven. The sample is cooled, the volume adjusted to 100mL and introduced to an inductively coupled argon plasma-sector field mass spectrometer (ICP-SFMS) by pneumatic nebulization. Each respective element is measured at a specific mass to charge ratio, m/z value expressed in amu (atomic mass units), and at a specified instrument resolution. LR: Low Resolution ~300, MR: Mid Resolution ~4000. Ref: 02-2704, National Laboratory for Environmental Testing, CCIW, Environment Canada, Burlington. Requested April 2003. | 0.001 | Y | Y | Y | 37 | 8/9/2016 | 10/16/2023 | 11/28/2023 2:21:33 PM | 2004-01-19 | 4/19/2021 1:22:00 PM | UMULLER | EMS | MassPerMass | µg/g | Y | (mass conc.) | (wet) | Rubidium Total (wet) (mass conc.) | 7440-17-7 | Metals - Total; | Metals - Total; EMS code: RB-T | Chemical | Total | LAB
RB-T | ug/g wet | Rubidium Total | X560 | Metals in Biota (wet), total, acid digestion, ICP-MS | Five grams of wet, homogenized biota is weighed in a centrifuge tube, and 2.5 mL of Nitric Acid and 2.5 mL of Hydrochloric acid are added. The sample is digested at room temperature for one hour and then digested at 90-95 degrees C in a water bath or heat block for one more hour. The sample is cooled and the volume is made to 50 mL with DI. The digestate is then diluted at ten fold and introduced to a collision/reaction cell inductively coupled argon plasma mass spectrometer (CRC ICPMS) by pneumatic nebulization. Each respective element is measured at a specific mass to charge ratio, m/z value expressed in amu (atomic mass units). - Maxxam Analytics (BC). Added to EMS Dec 10, 2008. | 0.01 | Y | Y | Y | 179 | 8/13/2014 | 10/30/2023 | 11/28/2023 2:21:33 PM | 2015-03-11 | 3/11/2015 4:17:10 PM | RDGIBSON | EMS | MassPerMass | µg/g | Y | (mass conc.) | (wet) | Rubidium Total (wet) (mass conc.) | 7440-17-7 | Metals - Total; | Metals - Total; EMS code: RB-T | Chemical | Total | LAB
RE-D | mg/L | Rhenium Dissolved | METC | MET-D-NP-CCMS-VA | Ultra trace metals in water are analyzed by CRC ICPMS, based on US EPA Method 6020B (July 2014). This procedure is intended for pristine field-filtered acid-preserved water samples. The detection limits (LOR) for this test are based on lab instrumental analysis only, not including filtration. ALS-supplied field filtration equipment does not support these LOR. Therefore, because of the high probability of false positives due to filtration, it is strongly recommended that a filtration blank be analysed to aid in data interpretation. | 5e-06 | Y | Y | Y | 394 | 3/30/2020 | 4/9/2024 | 4/29/2024 4:53:06 PM | 2022-10-07 |  | EMS |  | Concentration | mg/L | Y | (fl. conc.) |  | Rhenium Dissolved (fl. conc.) | 7440-15-5 | Metals - Dissolved; | Metals - Dissolved; EMS code: RE-D | Chemical | Dissolved | LAB
RE-D | mg/L | Rhenium Dissolved | X562 | Dissolved metals: FF, FA, HNO3, HR-ICPMS | Ultra trace metals in water are analyzed by high resolution inductively coupled plasma mass spectrometry (HR-ICPMS), based on US EPA Method 200.8, (Oct. 1999). This procedure is intended for pristine field-filtered acid-preserved water samples. The detection limits (LOR) for this test are based on lab instrumental analysis only, not including filtration. ALS-supplied field filtration equipment does not support these LOR. Therefore, because of the high probability of false positives due to filtration, it is strongly recommended that a filtration blank be analyzed to aid in data interpretation. | 5e-06 | Y | Y | Y | 10 | 10/20/2015 | 10/24/2018 | 10/24/2018 12:26:13 PM | 2016-03-24 |  | RDGIBSON |  | Concentration | mg/L | Y | (fl. conc.) |  | Rhenium Dissolved (fl. conc.) | 7440-15-5 | Metals - Dissolved; | Metals - Dissolved; EMS code: RE-D | Chemical | Dissolved | LAB
RE-D | mg/L | Rhenium Dissolved | MSEA | Metals in Seawater (undigested) by (CRC) ICP-MS | This method describes the multi-elemental determination of trace elements by Collision/Reaction Cell (CRC) ICP-MS. The method measures ions produced by a radio-frequency inductively coupled plasma. Analyte species originating in a liquid are pneumatically nebulized and the resulting aerosol transported by argon gas into the plasma torch. The ions produced are entrained in the plasma gas and extracted, by means of a differentially pumped vacuum interface, into a mass spectrometer. The ions flow through a Collision/Reaction Cell (CRC) where depending on the mode of analysis either the ions passes through the cell unhindered (no gas - classical ICP-MS mode), or collisions/reactions chemistry takes place between the ions and a neutral gas such as hydrogen or helium. Any polyatomic ions which give interference on the analytes are removed. The ions are then sorted according to their mass-to-charge ratios by a quadrupole mass spectrometer. The ions are then sorted according to their mass-to-charge ratios by a quadrupole mass spectrometer. The ions transmitted through the quadrupole are quantified by a channel electron multiplier and the ion information processed by a data handling system. Although most polyatomic interference is removed, any additional interference must be assessed and valid corrections applied or the data flagged to indicate problems. Interference correction must include compensation for background ions contributed by the plasma gas, reagents, and constituents of the sample matrix. Instrument drift as well as suppressions or enhancements of instrument response caused by the sample matrix must be corrected for by the use of internal standards. | 0.0005 | Y | Y | Y | 4 | 4/12/2023 | 4/29/2024 | 5/16/2024 2:00:14 PM | 2023-05-26 |  | EMS |  | Concentration | mg/L | Y | (fl. conc.) |  | Rhenium Dissolved (fl. conc.) | 7440-15-5 | Metals - Dissolved; | Metals - Dissolved; EMS code: RE-D | Chemical | Dissolved | LAB
RE-D | mg/L | Rhenium Dissolved | F032 | Filt.,Acid,ICPMS | There is no acid digestion required for this analysis. A water sample is collected and preserved in 0.2 percent nitric acid. To analyse the extractable metals the sample is shaken, allowed to sit for at least twenty four hours, and a decanted portion is introduced to the spectrometer. When the dissolved metals are required, the samples must first be filtered through a 0.45 micron cellulose acetate filter paper and then followed by preservation in 0.2 percent nitric acid. It is important to note that different laboratories may have very different detection limits for some variables analyzed by ICPMS. | 0.001 | Y | Y | Y | 48 | 7/30/2020 | 12/19/2023 | 1/2/2024 4:32:04 PM | 1998-06-22 |  | NPEPPIN |  | Concentration | mg/L | Y | (fl. conc.) |  | Rhenium Dissolved (fl. conc.) | 7440-15-5 | Metals - Dissolved; | Metals - Dissolved; EMS code: RE-D | Chemical | Dissolved | LAB
RE-T | mg/L | Rhenium Total | F082 | Aliquot,HNO3/HCL dig,ICPMS | Aliquot microwave digested with 3:1 HNO3/HCL. Analysis by ICPMS. For HiVol samples preliminary result are reported in mg/L and loading results are reported in ug/m3 | 5e-06 | Y | Y | Y | 69 | 7/30/2020 | 8/16/2024 | 8/21/2024 11:26:07 AM | 2021-11-03 |  | RDGIBSON |  | Concentration | mg/L | Y | (fl. conc.) |  | Rhenium Total (fl. conc.) | 7440-15-5 | Metals - Total; | Metals - Total; EMS code: RE-T | Chemical | Total | LAB
RE-T | ppm (S) | Rhenium Total | MS41 | ALS Minerals ME-MS41 | ALS Minerals ME-MS41. A prepared sample is digested with aqua regia in a graphite heating block. After cooling; the resulting solution is diluted to with deionized water; mixed and analyzed by ICP-AES. Samples are then analyzed by ICP-MS for the remaining suite of elements. | 0.001 | Y | Y | Y | 0 |  |  |  | 2024-01-29 |  | EMS |  | DimensionlessRatio | ppm | N | (ratio) |  | Rhenium Total (ratio) | 7440-15-5 | Metals - Total; | Metals - Total; EMS code: RE-T | Chemical | Total | LAB
RE-T | mg/L | Rhenium Total | MSEA | Metals in Seawater (undigested) by (CRC) ICP-MS | This method describes the multi-elemental determination of trace elements by Collision/Reaction Cell (CRC) ICP-MS. The method measures ions produced by a radio-frequency inductively coupled plasma. Analyte species originating in a liquid are pneumatically nebulized and the resulting aerosol transported by argon gas into the plasma torch. The ions produced are entrained in the plasma gas and extracted, by means of a differentially pumped vacuum interface, into a mass spectrometer. The ions flow through a Collision/Reaction Cell (CRC) where depending on the mode of analysis either the ions passes through the cell unhindered (no gas - classical ICP-MS mode), or collisions/reactions chemistry takes place between the ions and a neutral gas such as hydrogen or helium. Any polyatomic ions which give interference on the analytes are removed. The ions are then sorted according to their mass-to-charge ratios by a quadrupole mass spectrometer. The ions are then sorted according to their mass-to-charge ratios by a quadrupole mass spectrometer. The ions transmitted through the quadrupole are quantified by a channel electron multiplier and the ion information processed by a data handling system. Although most polyatomic interference is removed, any additional interference must be assessed and valid corrections applied or the data flagged to indicate problems. Interference correction must include compensation for background ions contributed by the plasma gas, reagents, and constituents of the sample matrix. Instrument drift as well as suppressions or enhancements of instrument response caused by the sample matrix must be corrected for by the use of internal standards. | 0.0005 | Y | Y | Y | 20 | 4/13/2023 | 4/29/2024 | 6/7/2024 10:03:16 AM | 2023-05-26 |  | EMS |  | Concentration | mg/L | Y | (fl. conc.) |  | Rhenium Total (fl. conc.) | 7440-15-5 | Metals - Total; | Metals - Total; EMS code: RE-T | Chemical | Total | LAB
RE-T | ppm (S) | Rhenium Total | 6042 | Aqua Regia by ICPMS | These were sediment/soils digested with Aqua-Regia (Nitric Acid & Hydrochloric Acids) and then analysis by ICPMS. Results would be reported in ppm s. | 0.001 | Y | Y | Y | 71 | 5/14/2022 | 7/27/2024 | 7/29/2024 9:13:08 AM | 2022-06-28 |  | EMS |  | DimensionlessRatio | ppm | N | (ratio) |  | Rhenium Total (ratio) | 7440-15-5 | Metals - Total; | Metals - Total; EMS code: RE-T | Chemical | Total | LAB
RE-T | ppm (S) | Rhenium Total | 6041 | Aqua Regia by ICPMS | These were sediment/soils digested with Aqua-Regia (Nitric Acid & Hydrochloric Acids) and then analysis by ICPMS. Results would be reported in mg/kg and/or %. | 0.0002 | Y | Y | Y | 21 | 10/13/2019 | 8/5/2020 | 8/10/2020 3:52:06 PM | 2019-09-06 | 10/21/2019 11:56:13 AM | KATFRASE | KATFRASE | DimensionlessRatio | ppm | N | (ratio) |  | Rhenium Total (ratio) | 7440-15-5 | Metals - Total; | Metals - Total; EMS code: RE-T | Chemical | Total | LAB
RE-T | mg/L | Rhenium Total | 1349 | Tot Metals in Seawater by CRC ICPMS (BC) | Tot Metals in Seawater by CRC ICPMS (BC) | 0.0005 | Y | Y | Y | 98 | 8/16/2019 | 12/5/2021 | 12/13/2021 12:34:16 PM | 2019-09-30 |  | KATFRASE |  | Concentration | mg/L | Y | (fl. conc.) |  | Rhenium Total (fl. conc.) | 7440-15-5 | Metals - Total; | Metals - Total; EMS code: RE-T | Chemical | Total | LAB
RE-T | ppm (S) | Rhenium Total | MS61 | ALS Minerals ME-MS61 | ALS Minerals ME-MS61. A prepared sample is digested with HNO3; HF; and HClO4 in a graphite heating block to near dryness. After cooling; the sample is boiled in diluted HCl. Bulk to volume with deionized water; mixed and analyzed by ICP-AES and ICP-MS for all elements. | 0.002 | Y | Y | Y | 0 |  |  |  | 2024-01-29 |  | EMS |  | DimensionlessRatio | ppm | N | (ratio) |  | Rhenium Total (ratio) | 7440-15-5 | Metals - Total; | Metals - Total; EMS code: RE-T | Chemical | Total | LAB
RE-T | mg/L | Rhenium Total | METN | MET-T-NP-CCMS-VA | Ultra trace metals in water are analyzed by CRC ICPMS, based on US EPA Method 6020B (July 2014). The detection limits provided can only be met for undigested samples. This procedure is intended for colorless, non-turbid, acid-preserved water samples (i.e. pristine water samples), having turbidity < 1 NTU and no odor. Where turbidity exceeds 1 NTU, and/or the sample is colored and has an odor, results may be biased low compared to true Total Metals concentrations. ALS recommends that turbidity analysis be requested on samples submitted for this test to aid with interpretation of results. Where turbidity is <1NTU, undigested metals are equivalent to total metals concentrations. | 5e-06 | Y | Y | Y | 280 | 8/6/2019 | 4/4/2024 | 4/29/2024 4:53:06 PM | 2021-06-16 |  | EMS |  | Concentration | mg/L | Y | (fl. conc.) |  | Rhenium Total (fl. conc.) | 7440-15-5 | Metals - Total; | Metals - Total; EMS code: RE-T | Chemical | Total | LAB
RE-T | mg/L | Rhenium Total | X561 | Total metals: FA, HNO3, HR-ICPMS (No Prep)Prep) | Ultra trace metals in water are analyzed by high resolution inductively coupled plasma mass spectrometry (HR-ICPMS), modified from US EPA 200.8, (Revision 5.5). The detection limits provided can only be met for undigested samples. This procedure is intended for colorless, non-turbid, acid-preserved water samples (i.e. pristine water samples), having turbidity < 1 NTU and no odor. Where turbidity exceeds 1 NTU, and/or the sample is colored and has an odor, results may be biased low compared to true Total Metals concentrations. ALS recommends that turbidity analysis be requested on samples submitted for this test to aid with interpretation of results. | 5e-06 | Y | Y | Y | 19 | 10/20/2015 | 10/24/2018 | 10/24/2018 12:26:13 PM | 2016-03-24 |  | RDGIBSON |  | Concentration | mg/L | Y | (fl. conc.) |  | Rhenium Total (fl. conc.) | 7440-15-5 | Metals - Total; | Metals - Total; EMS code: RE-T | Chemical | Total | LAB
RF-F | mg/L | Residue: Fixed Filterable (TDS; Fixed) | CS01 | System Calculation | System Calculation (mg/L) |  | N | Y | N | 45 |  |  | 6/27/2024 3:13:06 PM | 1996-04-12 | 6/3/1998 6:51:42 PM | EMS |  | Concentration | mg/L | Y | (fl. conc.) |  | Residue: Fixed Filterable (TDS; Fixed) (fl. conc.) |  | Physical Properties; | Physical Properties; EMS code: RF-F | Physical |  | LAB
RF-F | mg/L | Residue: Fixed Filterable (TDS; Fixed) | F012 | 1.2 um GF filt:Grav. 550C | A homogeneous sample is filtered through a 5.5 cm Whatman GF/C glass fibre filter (1.2 um particle retention) that has been muffled at 550 C for 20 minutes and pre-weighed. For fixed non-filterable residue (FNFR) analysis, the filter with the residue is then dried at 103 C for one hour and weighed once again to constant weight. The filter with dried and weighed residue is then ignited in a muffle at 550 C to determine FNFR. For fixed filterable residue (FFR) analysis, evaporated overnight at 90-95C, then dried for 1 hour at 103C and weighed for total filterable residue (FR), then ignited for 1 hour at 550C to determine FFR. Volatile residue is the difference in weights after drying at 103 C and asking at 550 C. | 10 | Y | Y | Y | 2193 |  |  | 9/23/2003 9:46:15 PM | 1996-06-19 | 6/3/1998 6:51:44 PM | MPAK |  | Concentration | mg/L | Y | (fl. conc.) |  | Residue: Fixed Filterable (TDS; Fixed) (fl. conc.) |  | Physical Properties; | Physical Properties; EMS code: RF-F | Physical |  | LAB
RF-F | mg/L | Residue: Fixed Filterable (TDS; Fixed) | X237 | Gravimetric | This code is a catch-all for a variety of gravimetric procedures. | 10 | Y | Y | Y | 4 |  |  | 9/11/1997 1:50:45 PM | 1996-04-12 | 6/3/1998 6:51:38 PM | EMS |  | Concentration | mg/L | Y | (fl. conc.) |  | Residue: Fixed Filterable (TDS; Fixed) (fl. conc.) |  | Physical Properties; | Physical Properties; EMS code: RF-F | Physical |  | LAB
ROCK | % | Whole rock; total | ICP06 | ALS Minerals ME-ICP06 | ALS Minerals ME-ICP06. A prepared sample is fused with a lithium tetraborate/lithium metaborate flux and poured into a graphite crucible and fused in a furnace at 1000°C. The resulting melt is then cooled and dissolved in 100 mL of 4% HNO3/2% HCl followed by ICP-AES. | 0.01 | Y | Y | Y | 0 |  |  |  | 2024-01-29 |  | EMS |  | DimensionlessRatio | % | N | (ratio) |  | Whole rock; total (ratio) |  | Physical Properties; | Physical Properties; EMS code: ROCK | Physical |  | LAB
ROTH | None | Rept Prod-Other Spls | 5104 | Report Prod'n -Other | Report Prod'n -Other | 0 | N | Y | Y | 126 |  |  | 10/24/1996 9:18:47 AM | 1996-04-12 | 6/3/1998 6:51:38 PM | EMS |  | None | None | N | (unitless) |  | Rept Prod-Other Spls (unitless) |  |  | EMS code: ROTH | Chemical |  | LAB
ROXT | mg/L | Roxithromycin CAS: 80214-83-1 | XMSP | HPLC and Tri. Quadpole MS/MS | EPA 1694 An aliquot of filtered sample containing 5% organic is injected directly. Each compound is separated by reversed phase HPLC. Analyzed by Electro Spray Ionization MS/MS detection using Triple Quadrupole MS/MS detector. Quantiitation with internal or external standard. |  | Y | Y | Y | 12 | 5/8/2020 | 6/4/2024 | 8/21/2024 3:32:03 PM | 2022-05-04 |  | EMS |  | Concentration | mg/L | Y | (fl. conc.) |  | Roxithromycin (fl. conc.) | 80214-83-1 | Pharmaceuticals and Personal Care Products (PPCP); | Pharmaceuticals and Personal Care Products (PPCP); EMS code: ROXT | Chemical |  | LAB
S | % | Sulfur; elemental | GRA07 | ALS Minerals S-GRA07 | ALS Minerals S-GRA07. Carbon disulfide is added to a prepared sample and the mixture is shaken for 30 minutes. An aliquot of the solution is taken and evaporated until dry. The residue is then weighed as elemental sulfur. | 0.01 | Y | Y | Y | 0 |  |  |  | 2024-01-29 |  | EMS |  | DimensionlessRatio | % | N | (ratio) |  | Sulfur; elemental (ratio) |  |  | EMS code: S | Chemical |  | LAB
S--D | mg/L | Sulfur Dissolved | METC | MET-D-NP-CCMS-VA | Ultra trace metals in water are analyzed by CRC ICPMS, based on US EPA Method 6020B (July 2014). This procedure is intended for pristine field-filtered acid-preserved water samples. The detection limits (LOR) for this test are based on lab instrumental analysis only, not including filtration. ALS-supplied field filtration equipment does not support these LOR. Therefore, because of the high probability of false positives due to filtration, it is strongly recommended that a filtration blank be analysed to aid in data interpretation. | 0.5 | Y | Y | Y | 394 | 3/30/2020 | 4/9/2024 | 4/29/2024 4:53:06 PM | 2022-10-07 |  | EMS |  | Concentration | mg/L | Y | (fl. conc.) |  | Sulfur Dissolved (fl. conc.) | 7704-34-9 | Metals - Dissolved; | Metals - Dissolved; EMS code: S--D | Chemical | Dissolved | LAB
S--D | mg/L | Sulfur Dissolved | F039 | Field filt + acid, ICP-AES | Dissolved and extractable water samples are ready for analysis once collected. Total samples are autoclaved with HNO3 and hydrochloric acid (HCI) in capped Teflon vials. Analysis is performed via Inductively Coupled Argon Plasma Atomic Emission Spectrometry (ICP-AES)(simultaneous multielement analysis), Graphite Furnace Atomic Absorption Spectrometry (GF-AAS)(low level analysis) and/or Flame Atomic Absorption Spectrometry (F-AAS). | 0.05 | Y | Y | Y | 11159 | 12/12/1996 6:02:00 PM | 7/5/2019 1:21:00 PM | 12/11/2019 4:35:32 PM | 1997-06-23 | 6/3/1998 6:51:47 PM | NPEPPIN |  | Concentration | mg/L | Y | (fl. conc.) |  | Sulfur Dissolved (fl. conc.) | 7704-34-9 | Metals - Dissolved; | Metals - Dissolved; EMS code: S--D | Chemical | Dissolved | LAB
S--D | mg/L | Sulfur Dissolved | X257 | Colorimetric | Colorimetric. General use for several parameters. E.g.Total Sulfide is determined by spectrophotometer using the methlyene blue colourimetric method. The H2S calculation converts Total Sulphide as (S2-) and reports it as Total Sulphide as (H2S). | 0.02 | Y | Y | Y | 271 | 1/5/2007 | 8/15/2024 | 8/16/2024 4:05:21 PM | 2012-06-25 |  | RDGIBSON |  | Concentration | mg/L | Y | (fl. conc.) |  | Sulfur Dissolved (fl. conc.) | 7704-34-9 | Metals - Dissolved; | Metals - Dissolved; EMS code: S--D | Chemical | Dissolved | LAB
S--D | mg/L | Sulfur Dissolved | F032 | Filt.,Acid,ICPMS | There is no acid digestion required for this analysis. A water sample is collected and preserved in 0.2 percent nitric acid. To analyse the extractable metals the sample is shaken, allowed to sit for at least twenty four hours, and a decanted portion is introduced to the spectrometer. When the dissolved metals are required, the samples must first be filtered through a 0.45 micron cellulose acetate filter paper and then followed by preservation in 0.2 percent nitric acid. It is important to note that different laboratories may have very different detection limits for some variables analyzed by ICPMS. | 2 | Y | Y | Y | 50651 | 4/8/2010 12:25:00 AM | 8/26/2024 | 8/26/2024 4:51:04 PM | 2010-05-17 | 7/16/2012 3:34:51 PM | RDGIBSON | RDGIBSON | Concentration | mg/L | Y | (fl. conc.) |  | Sulfur Dissolved (fl. conc.) | 7704-34-9 | Metals - Dissolved; | Metals - Dissolved; EMS code: S--D | Chemical | Dissolved | LAB
S--D | mg/L | Sulfur Dissolved | X170 | ICP | This code is a catch-all for a large variety of ICP methods including water and also solids digested into solutions as an intermediate result. Different procedures will have greatly different detection limits. (a) For air particulates, preliminary results are reported in mg/L and loading results as mg/dm2/d. (b) For Hi Vol samples preliminiary results are same as reported in mg/L and loading results as ug/m3. (c) Note that for Rain samples there is no filtration nor digestion steps. | 0.03 | Y | Y | Y | 1977 | 12/3/1996 | 9/22/2016 | 6/13/2024 3:58:21 PM | 1996-04-12 | 6/3/1998 6:51:38 PM | EMS |  | Concentration | mg/L | Y | (fl. conc.) |  | Sulfur Dissolved (fl. conc.) | 7704-34-9 | Metals - Dissolved; | Metals - Dissolved; EMS code: S--D | Chemical | Dissolved | LAB
S--D | mg/L | Sulfur Dissolved | MSEA | Metals in Seawater (undigested) by (CRC) ICP-MS | This method describes the multi-elemental determination of trace elements by Collision/Reaction Cell (CRC) ICP-MS. The method measures ions produced by a radio-frequency inductively coupled plasma. Analyte species originating in a liquid are pneumatically nebulized and the resulting aerosol transported by argon gas into the plasma torch. The ions produced are entrained in the plasma gas and extracted, by means of a differentially pumped vacuum interface, into a mass spectrometer. The ions flow through a Collision/Reaction Cell (CRC) where depending on the mode of analysis either the ions passes through the cell unhindered (no gas - classical ICP-MS mode), or collisions/reactions chemistry takes place between the ions and a neutral gas such as hydrogen or helium. Any polyatomic ions which give interference on the analytes are removed. The ions are then sorted according to their mass-to-charge ratios by a quadrupole mass spectrometer. The ions are then sorted according to their mass-to-charge ratios by a quadrupole mass spectrometer. The ions transmitted through the quadrupole are quantified by a channel electron multiplier and the ion information processed by a data handling system. Although most polyatomic interference is removed, any additional interference must be assessed and valid corrections applied or the data flagged to indicate problems. Interference correction must include compensation for background ions contributed by the plasma gas, reagents, and constituents of the sample matrix. Instrument drift as well as suppressions or enhancements of instrument response caused by the sample matrix must be corrected for by the use of internal standards. | 6000 | Y | Y | Y | 4 | 4/12/2023 | 4/29/2024 | 5/16/2024 2:00:14 PM | 2009-09-08 |  | RDGIBSON |  | Concentration | mg/L | Y | (fl. conc.) |  | Sulfur Dissolved (fl. conc.) | 7704-34-9 | Metals - Dissolved; | Metals - Dissolved; EMS code: S--D | Chemical | Dissolved | LAB
S--D | mg/L | Sulfur Dissolved | X562 | Dissolved metals: FF, FA, HNO3, HR-ICPMS | Ultra trace metals in water are analyzed by high resolution inductively coupled plasma mass spectrometry (HR-ICPMS), based on US EPA Method 200.8, (Oct. 1999). This procedure is intended for pristine field-filtered acid-preserved water samples. The detection limits (LOR) for this test are based on lab instrumental analysis only, not including filtration. ALS-supplied field filtration equipment does not support these LOR. Therefore, because of the high probability of false positives due to filtration, it is strongly recommended that a filtration blank be analyzed to aid in data interpretation. | 0.5 | Y | Y | Y | 15 | 5/29/2018 | 10/19/2023 | 11/9/2023 7:41:27 AM | 2018-06-06 |  | RDGIBSON |  | Concentration | mg/L | Y | (fl. conc.) |  | Sulfur Dissolved (fl. conc.) | 7704-34-9 | Metals - Dissolved; | Metals - Dissolved; EMS code: S--D | Chemical | Dissolved | LAB
S--D | mg/L | Sulfur Dissolved | X285 | HNO3 Dig:ICP Analysis | HNO3 Dig:ICP Analysis | 0.1 | Y | Y | Y | 519 | 4/26/2012 3:12:00 PM | 4/26/2012 3:17:00 PM | 5/4/2012 2:22:51 PM | 1998-05-22 | 6/3/1998 6:51:49 PM | NPEPPIN |  | Concentration | mg/L | Y | (fl. conc.) |  | Sulfur Dissolved (fl. conc.) | 7704-34-9 | Metals - Dissolved; | Metals - Dissolved; EMS code: S--D | Chemical | Dissolved | LAB
S--D | mg/L | Sulfur Dissolved | MWID | Metals in Water by ICPMS (digested) | Metals in Water by ICPMS (digested) Based on EPA 200.8 | 60 | Y | Y | Y | 451 |  |  | 1/27/2023 10:37:54 AM | 2009-12-24 |  | RDGIBSON |  | Concentration | mg/L | Y | (fl. conc.) |  | Sulfur Dissolved (fl. conc.) | 7704-34-9 | Metals - Dissolved; | Metals - Dissolved; EMS code: S--D | Chemical | Dissolved | LAB
S--E | mg/L | Sulfur Extractable | F040 | Field Acid, lab filt, ICP-AES | Dissolved and extractable water samples are ready for analysis once collected. Total samples are autoclaved with HNO3 and hydrochloric acid (HCI) in capped Teflon vials. Analysis is performed via Inductively Coupled Argon Plasma Atomic Emission Spectrometry (ICP-AES)(simultaneous multielement analysis), Graphite Furnace Atomic Absorption Spectrometry (GF-AAS)(low level analysis) and/or Flame Atomic Absorption Spectrometry (F-AAS). | 0.05 | Y | Y | Y | 2350 | 8/5/1997 6:06:00 PM | 3/14/2008 | 9/11/2019 8:47:41 AM | 1997-06-23 | 6/3/1998 6:51:47 PM | NPEPPIN |  | Concentration | mg/L | Y | (fl. conc.) |  | Sulfur Extractable (fl. conc.) | 7704-34-9 | Metals - Extractable; | Metals - Extractable; EMS code: S--E | Chemical | Extractable | LAB
S--T | ug/g | Sulfur Total | EQ02 | EQUIS (ug/g) | EQUIS (ug/g) | 0 | N | Y | Y | 174 |  |  | 10/22/1996 10:22:16 PM | 1996-04-12 | 6/3/1998 6:51:38 PM | EMS |  | MassPerMass | mg/kg | Y | (mass conc.) |  | Sulfur Total (mass conc.) | 7704-34-9 |  | EMS code: S--T | Chemical | Total | LAB
S--T | ug/g | Sulfur Total | FS01 | D&G,Nitric/HCl:ICP | Samples are dried at 60 C, then either sieved to -100mesh (150 um) and rolled to homogenize, or finely ground with a mortar and pestle or a shatter box laboratory mill and rolled (depending on the submitter's requirements). The samples are then weighed into Teflon digestion liners and digested with HNO3 and HCl in a microwave oven. Samples are cooled, volumized, and allowed to settle overnight before being decanted. Analysis is then performed via Inductively Coupled Argon Plasma Atomic Emission Spectrometry (ICP-AES) (Simultaneous multi element analysis). In ICP Spectrometry, the liquid sample is pumped to the nebulizer, where it is convertd into an aerosol. the nebulizer carrier gas flow than transports the wet aerosol past the impact bead, where the large droplets condense out and flow to the waste container. The remaining fine aerosol is carried into the plasma. There, the high temperature desolvates the aerosol and dissociates the sample material into individual atoms and ions, which are excited to emit light at wavelengths characteristic of the elements present. A diffraction grating in the spectrometer diffracts the various wavelengths of light through exit slits to the photo-multiplier tubes, which convert the light energy to electronic signals. | 8 | Y | Y | Y | 2189 | 1/29/1997 6:04:00 PM | 2/23/2004 | 4/1/2004 1:19:35 AM | 1997-10-17 | 6/3/1998 6:51:47 PM | NPEPPIN |  | MassPerMass | mg/kg | Y | (mass conc.) |  | Sulfur Total (mass conc.) | 7704-34-9 |  | EMS code: S--T | Chemical | Total | LAB
S--T | % | Sulfur Total | IR08 | ALS Minerals ME-IR08 | ALS Minerals ME-IR08. A prepared sample is heated to approximately 1350°C in an induction furnace with oxygen stream. Sulfur and carbon dioxide released from the sample are measured by IR (Leco analyzer) | 0.01 | Y | Y | Y | 0 |  |  |  | 2024-01-29 |  | RDGIBSON |  | DimensionlessRatio | % | N | (ratio) |  | Sulfur Total (ratio) | 7704-34-9 |  | EMS code: S--T | Chemical | Total | LAB
S--T | mg/L | Sulfur Total | 1349 | Tot Metals in Seawater by CRC ICPMS (BC) | Tot Metals in Seawater by CRC ICPMS (BC) | 5 | Y | Y | Y | 98 | 8/16/2019 | 12/5/2021 | 12/13/2021 12:34:16 PM | 2019-09-30 |  | KATFRASE |  | Concentration | mg/L | Y | (fl. conc.) |  | Sulfur Total (fl. conc.) | 7704-34-9 |  | EMS code: S--T | Chemical | Total | LAB
S--T | ug/g | Sulfur Total | X551 | Metals in Biota (dry), total, acid digestion, ICP-SFMS | 0.1 gram of dried, homogenized biota is weighed in a microwave TFM vessel, and 8mL of Nitric Acid and 1mL of Hydrogen Peroxide are added. The sample is digested at 200 degrees C for 15 minutes in a closed vessel high pressure microwave oven. The sample is cooled, the volume adjusted to 100mL and introduced to an iductively coupled argon plasma-sector field mass spectrometer (ICP-SFMS) by pneumatic nebulization. Each respective element is measured at a specific mass to charge ratio, m/z value expressed in amu (atomic mass units), and at a specified instrument resolution. LR: Low Resolution ~300, MR: Mid Resolution ~4000 and HR: Hich Resolution ~10000. Ref: Method number 02-2705, National Laboratory for Environmental Testing, CCIW, Environment Canada, Burlington. Requested April 2003. | 20 | Y | Y | Y | 6 | 12/22/2015 | 12/7/2016 | 9/28/2018 3:56:05 PM | 2016-09-29 |  | RRENTON |  | MassPerMass | mg/kg | Y | (mass conc.) |  | Sulfur Total (mass conc.) | 7704-34-9 |  | EMS code: S--T | Chemical | Total | LAB
S--T | ug/g | Sulfur Total | FB04 | Freeze Dry,Hom,Acid Dig,ICP | Sample moisture is determined by weighing samples before and after freeze-drying. Freeze-dried samples are ground in a Braun blender before being weighed into Teflon PFA liners and digested with concentrated HNO3, concentrated HCl, and 50% H2O2 in a microwave oven; which results in a complete sample dissolution. Samples are collected and volumized with Type I water. Analysis is performed via Inductively Coupled Argon Plasma Atomic Emission Spectrometry (ICP-AES) (simultaneous multielement analysis), Graphite Furnace Atomic Absorption Spectrometry (GF-AAS) (low level Cd and Pb analysis), Cold Vapor Atomic Fluorescence Spectrometry (CV-AFS) (Mercury analysis) and/or Flame Atomic Absorption Spectrometry (F-AAS) (K analysis). | 4 | Y | Y | Y | 235 | 4/16/1997 6:01:00 PM | 2/1/2002 2:30:00 PM | 6/23/2003 1:15:33 PM | 1997-04-02 | 6/3/1998 6:51:46 PM | MPAK |  | MassPerMass | mg/kg | Y | (mass conc.) |  | Sulfur Total (mass conc.) | 7704-34-9 |  | EMS code: S--T | Chemical | Total | LAB
S--T | ug/g | Sulfur Total | 0630 | Sulfur Anal.Leco | Sulfur Anal.Leco | 100 | Y | Y | Y | 213 | 9/22/2021 2:54:00 PM | 9/30/2021 1:19:00 PM | 10/18/2021 9:46:09 AM | 1996-04-12 | 6/3/1998 6:51:45 PM | EMS |  | MassPerMass | mg/kg | Y | (mass conc.) |  | Sulfur Total (mass conc.) | 7704-34-9 |  | EMS code: S--T | Chemical | Total | LAB
S--T | mg/dm2/d | Sulfur Total | EQ15 | EQUIS (mg/dm2/d) | EQUIS (mg/dm2/d) | 0 | N | Y | Y | 6 |  |  | 10/22/1996 10:10:50 PM | 1996-04-12 | 6/3/1998 6:51:38 PM | EMS |  | MassFlux | mg/dm²/d | Y | (mass flux) |  | Sulfur Total (mass flux) | 7704-34-9 |  | EMS code: S--T | Chemical | Total | LAB
S--T | mg/L | Sulfur Total | X349 | FA:HNO3:ICP | Samples are field-acidified with HNO3; HNO3 digestion; analysis by ICP/AES. | 0.1 | Y | Y | Y | 30 |  |  | 10/30/2014 8:31:18 AM | 1998-05-22 | 6/3/1998 6:51:49 PM | NPEPPIN |  | Concentration | mg/L | Y | (fl. conc.) |  | Sulfur Total (fl. conc.) | 7704-34-9 |  | EMS code: S--T | Chemical | Total | LAB
S--T | mg/L | Sulfur Total | F082 | Aliquot,HNO3/HCL dig,ICPMS | Aliquot microwave digested with 3:1 HNO3/HCL. Analysis by ICPMS. For HiVol samples preliminary result are reported in mg/L and loading results are reported in ug/m3 | 0.5 | Y | Y | Y | 17508 | 4/6/2015 | 8/23/2024 | 8/26/2024 3:51:07 PM | 2015-07-30 |  | RDGIBSON |  | Concentration | mg/L | Y | (fl. conc.) |  | Sulfur Total (fl. conc.) | 7704-34-9 |  | EMS code: S--T | Chemical | Total | LAB
S--T | mg/L | Sulfur Total | F038 | HNO3/HCl dig,ICP-AES | Dissolved and extractable water samples are ready for analysis once collected. Total samples are autoclaved with HNO3 and hydrochloric acid (HCI) in capped Teflon vials. Analysis is performed via Inductively Coupled Argon Plasma Atomic Emission Spectrometry (ICP-AES)(simultaneous multielement analysis), Graphite Furnace Atomic Absorption Spectrometry (GF-AAS) (low level analysis) and/or Flame Atomic Absorption Spectrometry (F-AAS). | 0.06 | Y | Y | Y | 6331 | 12/16/1996 6:03:00 PM | 6/3/2008 9:44:00 AM | 9/11/2019 8:47:42 AM | 1997-06-23 | 6/3/1998 6:51:47 PM | NPEPPIN |  | Concentration | mg/L | Y | (fl. conc.) |  | Sulfur Total (fl. conc.) | 7704-34-9 |  | EMS code: S--T | Chemical | Total | LAB
S--T | ug/g | Sulfur Total | FS89 | Homo,HNO3 dig,ICP-AES | Homo;HNO3 dig;ICP-AES | 8 | Y | Y | Y | 4 | 1/5/2001 | 8/20/2001 12:00:00 PM | 10/23/2002 1:42:51 PM | 2002-09-10 | 9/10/2002 9:36:24 AM | UMULLER | UMULLER | MassPerMass | mg/kg | Y | (mass conc.) |  | Sulfur Total (mass conc.) | 7704-34-9 |  | EMS code: S--T | Chemical | Total | LAB
S--T | % | Sulfur Total | X471 | Recovery Calculation (%) | Recovery Calculation (%) | 0 | Y | Y | Y | 18 |  |  | 10/29/2012 9:36:49 AM | 1996-12-09 | 6/3/1998 6:51:46 PM | MPAK |  | DimensionlessRatio | % | N | (ratio) |  | Sulfur Total (ratio) | 7704-34-9 |  | EMS code: S--T | Chemical | Total | LAB
S--T | mg/L | Sulfur Total | F033 | Acid Dig,Centrifuge,ICPMS | For the determination of total recoverable elements in water or waste water, take 100 ml aliquot from a well mixed, acid preserved, sample containing not more than 0.25% (w/v) total solids and transfer to a 250 ml Griffin beaker. (If total solids are greater than 0.25% reduce the size of the aliquot by a proportionate amount). Add 1 ml of conc. nitric acid and 0.5 ml conc. hydrochloric acid. Heat on a hot plate until the volume has been reduced to 15-20ml, ensuring that the sample does not boil. Allow to cool and transfer to a 50 ml volumetric flask. Dilute to volume with ASTM type I water. Centrifuge the sample or allow to stand overnight to separate insoluble material. Prior analysis, pipette 20 ml into a 50 ml volumetric flask and dilute to volume with ASTM type I water. Analysis should be performed within seventy two hours of the completed preparation. Note that different laboratories may have very different field preservation, digestion procedures and detection limits for this method, so these should be documented. | 2 | Y | Y | Y | 21383 | 4/8/2010 8:00:00 AM | 8/22/2024 3:24:00 AM | 8/25/2024 9:47:36 PM | 2010-05-17 | 7/16/2012 3:35:15 PM | RDGIBSON | RDGIBSON | Concentration | mg/L | Y | (fl. conc.) |  | Sulfur Total (fl. conc.) | 7704-34-9 |  | EMS code: S--T | Chemical | Total | LAB
S--T | ug/g | Sulfur Total | PB02 | PBM,HNO3/HCl dig,dry60C,100mesh, | Performance Based Method. Samples are dried at 60 C and sieved using a .149 mm (100 mesh) sieve. Solid samples are digested with the combination of nitric and hydrochloric acids with a 1:1 ratio of acids. For environmental purposes, strong acid leaches are used routinely since they have the advantage of speed, low cost, and superio reproducibility. All metals trapped within the refractory matrix are unlikely to have environmental consequences. Instrumental analysis of extracts produced using this digestion method can be performed by a variety of techniques. |  | Y | Y | Y | 10 | 9/13/2001 | 9/13/2001 | 11/6/2002 3:52:39 PM | 2001-07-18 |  | NPEPPIN |  | MassPerMass | mg/kg | Y | (mass conc.) |  | Sulfur Total (mass conc.) | 7704-34-9 |  | EMS code: S--T | Chemical | Total | LAB
S--T | ug/g | Sulfur Total | TSCM | Tot S by combustion | The sample is ignited in a combustion analyzer where sulfur in the reduced SO2 gas is determined using a thermal conductivity detector. | 500 | Y | Y | Y | 21 | 9/26/2017 | 10/5/2018 | 10/10/2018 2:46:10 PM | 2018-01-09 |  | JKROGH |  | MassPerMass | mg/kg | Y | (mass conc.) |  | Sulfur Total (mass conc.) | 7704-34-9 |  | EMS code: S--T | Chemical | Total | LAB
S--T | % | Sulfur Total | MS61 | ALS Minerals ME-MS61 | ALS Minerals ME-MS61. A prepared sample is digested with HNO3; HF; and HClO4 in a graphite heating block to near dryness. After cooling; the sample is boiled in diluted HCl. Bulk to volume with deionized water; mixed and analyzed by ICP-AES and ICP-MS for all elements. | 0.01 | Y | Y | Y | 0 |  |  |  | 2024-01-29 |  | EMS |  | DimensionlessRatio | % | N | (ratio) |  | Sulfur Total (ratio) | 7704-34-9 |  | EMS code: S--T | Chemical | Total | LAB
S--T | mg/L | Sulfur Total | MWID | Metals in Water by ICPMS (digested) | Metals in Water by ICPMS (digested) Based on EPA 200.8 | 60 | Y | Y | Y | 350 |  |  | 1/19/2023 12:03:00 PM | 2009-12-24 |  | RDGIBSON |  | Concentration | mg/L | Y | (fl. conc.) |  | Sulfur Total (fl. conc.) | 7704-34-9 |  | EMS code: S--T | Chemical | Total | LAB
S--T | t/d | Sulfur Total | CALC | Calculation (t/d) | Calculation (t/d) | 0 | Y | Y | Y | 4 | 12/1/2010 | 12/1/2010 | 12/9/2010 3:55:27 PM | 1997-01-15 | 6/3/1998 6:51:46 PM | MPAK |  | MassFlow | t/d | Y | (mass flow) |  | Sulfur Total (mass flow) | 7704-34-9 |  | EMS code: S--T | Chemical | Total | LAB
S--T | mg/L | Sulfur Total | X145 | Dig, ICP | Dig; ICP | 0.5 | Y | Y | Y | 437 | 5/14/2013 | 9/22/2016 | 6/13/2024 3:58:21 PM | 2015-03-11 |  | RDGIBSON |  | Concentration | mg/L | Y | (fl. conc.) |  | Sulfur Total (fl. conc.) | 7704-34-9 |  | EMS code: S--T | Chemical | Total | LAB
S--T | % | Sulfur Total | 6041 | Aqua Regia by ICPMS | These were sediment/soils digested with Aqua-Regia (Nitric Acid & Hydrochloric Acids) and then analysis by ICPMS. Results would be reported in mg/kg and/or %. | 0.01 | Y | Y | Y | 92 | 10/13/2019 | 7/27/2024 | 7/29/2024 9:13:08 AM | 2019-09-06 |  | KATFRASE |  | DimensionlessRatio | % | N | (ratio) |  | Sulfur Total (ratio) | 7704-34-9 |  | EMS code: S--T | Chemical | Total | LAB
S--T | ug/g | Sulfur Total | MTID | MET-DRY-CCMS-VA | Metals in Tissue by CRC ICPMS (DRY) Methodology: This method is conducted following British Columbia Lab Manual method "Metals in Animal Tissue and Vegetation (Biota) - Prescriptive". Tissue samples are homogenized and sub-sampled prior to hotblock digestion with nitric and hydrochloric acids, in combination with addition of hydrogen peroxide. Instrumental analysis is by collision cell inductively coupled plasma - mass spectrometry (modified from EPA Method 6020A). Method Limitation: This method employs a strong acid/peroxide digestion, and is intended to provide a conservative estimate of bio-available metals. Near complete recoveries are achieved for most toxicologically important metals, but elements associated with recalcitrant minerals may be only partially recovered. | 100 | Y | Y | Y | 20 | 9/1/2018 | 9/1/2018 | 9/5/2018 11:41:04 AM | 2018-04-05 |  | RDGIBSON |  | MassPerMass | mg/kg | Y | (mass conc.) |  | Sulfur Total (mass conc.) | 7704-34-9 |  | EMS code: S--T | Chemical | Total | LAB
S--T | % | Sulfur Total | S-IR08 | ALS Minerals S-IR08 | ALS Minerals S-IR08. A prepared sample is heated to approximately 1350°C in an induction furnace with oxygen stream. Sulfur dioxide released from the sample are measured by IR (Leco analyzer) and the Total Sulfur result is provided. | 0.01 | Y | Y | Y | 0 |  |  |  | 2024-01-29 |  | EMS |  | DimensionlessRatio | % | N | (ratio) |  | Sulfur Total (ratio) | 7704-34-9 |  | EMS code: S--T | Chemical | Total | LAB
S--T | ug/g | Sulfur Total | X230 | Hom,HNO3/HCl,ICP | The sample is dried, ground to pass 100 mesh, and digested with a nitric-hydrochloric acid mixture to solubilize the solid matter and to remove the organic material by oxidation and volatilization. The completed digests are submitted to metals instrumentation personnel for metals analysis by Inductively Coupled Atomic Emission Spectrophotometry (ICP-AES), or hydride analysis by atomic absorption. | 1 | Y | Y | Y | 36 | 2/12/2002 | 5/29/2002 | 1/14/2004 1:00:44 AM | 1998-11-05 |  | NPEPPIN |  | MassPerMass | mg/kg | Y | (mass conc.) |  | Sulfur Total (mass conc.) | 7704-34-9 |  | EMS code: S--T | Chemical | Total | LAB
S--T | mg/L | Sulfur Total | 0040 | HNO3 Dig: ICP Analysis | HNO3 Dig: ICP Analysis | 0.03 | Y | Y | Y | 1109 | 12/7/2009 3:52:00 PM | 5/14/2014 2:05:00 AM | 6/9/2014 3:03:09 PM | 1996-04-12 | 6/3/1998 6:51:45 PM | EMS |  | Concentration | mg/L | Y | (fl. conc.) |  | Sulfur Total (fl. conc.) | 7704-34-9 |  | EMS code: S--T | Chemical | Total | LAB
S--T | ug/g | Sulfur Total | FS82 | PBM,dry60C,10mesh,SALM-ICP-MS | PBM;dry60C;10mesh;SALM-ICP-MS | 1000 | Y | Y | Y | 238 | 3/9/2010 11:41:00 AM | 7/24/2024 | 7/29/2024 5:17:09 PM | 2010-04-19 |  | RDGIBSON |  | MassPerMass | mg/kg | Y | (mass conc.) |  | Sulfur Total (mass conc.) | 7704-34-9 |  | EMS code: S--T | Chemical | Total | LAB
S--T | mg/L | Sulfur Total | X170 | ICP | This code is a catch-all for a large variety of ICP methods including water and also solids digested into solutions as an intermediate result. Different procedures will have greatly different detection limits. (a) For air particulates, preliminary results are reported in mg/L and loading results as mg/dm2/d. (b) For Hi Vol samples preliminiary results are same as reported in mg/L and loading results as ug/m3. (c) Note that for Rain samples there is no filtration nor digestion steps. | 0.03 | Y | Y | Y | 2829 | 12/16/1996 6:07:00 PM | 2/12/2015 | 10/24/2022 10:58:25 AM | 1996-04-12 | 6/3/1998 6:51:38 PM | EMS |  | Concentration | mg/L | Y | (fl. conc.) |  | Sulfur Total (fl. conc.) | 7704-34-9 |  | EMS code: S--T | Chemical | Total | LAB
S--T | mg/L | Sulfur Total | 0042 | GI HNO3 Dig: ICP | Sample aliquot taken from unfiltered unpreserved General Ions bottle; HNO3 digestion; ICP analysis. | 0.1 | Y | Y | Y | 6605 | 9/20/1996 | 7/5/2019 3:21:00 PM | 7/12/2019 6:49:13 PM | 1996-04-12 | 6/3/1998 6:51:38 PM | EMS |  | Concentration | mg/L | Y | (fl. conc.) |  | Sulfur Total (fl. conc.) | 7704-34-9 |  | EMS code: S--T | Chemical | Total | LAB
S--T | mg/L | Sulfur Total | X561 | Total metals: FA, HNO3, HR-ICPMS (No Prep)Prep) | Ultra trace metals in water are analyzed by high resolution inductively coupled plasma mass spectrometry (HR-ICPMS), modified from US EPA 200.8, (Revision 5.5). The detection limits provided can only be met for undigested samples. This procedure is intended for colorless, non-turbid, acid-preserved water samples (i.e. pristine water samples), having turbidity < 1 NTU and no odor. Where turbidity exceeds 1 NTU, and/or the sample is colored and has an odor, results may be biased low compared to true Total Metals concentrations. ALS recommends that turbidity analysis be requested on samples submitted for this test to aid with interpretation of results. | 0.5 | Y | Y | Y | 6 | 5/29/2018 | 10/5/2023 | 1/31/2024 2:37:41 PM | 2018-06-06 |  | RDGIBSON |  | Concentration | mg/L | Y | (fl. conc.) |  | Sulfur Total (fl. conc.) | 7704-34-9 |  | EMS code: S--T | Chemical | Total | LAB
S--T | mg/L | Sulfur Total | F121 | Acid,ICPMS | Acidified, no digestion, ICPMS | 1 | Y | Y | Y | 4408 | 10/4/2010 | 8/16/2024 | 8/26/2024 8:06:03 AM | 2011-09-09 |  | RDGIBSON |  | Concentration | mg/L | Y | (fl. conc.) |  | Sulfur Total (fl. conc.) | 7704-34-9 |  | EMS code: S--T | Chemical | Total | LAB
S--T | mg/L | Sulfur Total | X285 | HNO3 Dig:ICP Analysis | HNO3 Dig:ICP Analysis | 0.1 | Y | Y | Y | 561 | 8/10/1993 | 10/5/1993 | 4/30/2004 11:32:38 AM | 1998-05-22 | 6/3/1998 6:51:49 PM | NPEPPIN |  | Concentration | mg/L | Y | (fl. conc.) |  | Sulfur Total (fl. conc.) | 7704-34-9 |  | EMS code: S--T | Chemical | Total | LAB
S--T | ug/g | Sulfur Total | 0190 | D&G:N/P Dig:ICP Anal. | Sample is dried and grind-mixed before nitric perchloric digestion followed by ICP analysis. | 5 | Y | Y | Y | 2 |  |  | 10/24/1996 3:06:16 PM | 1996-04-12 | 6/3/1998 6:51:38 PM | EMS |  | MassPerMass | mg/kg | Y | (mass conc.) |  | Sulfur Total (mass conc.) | 7704-34-9 |  | EMS code: S--T | Chemical | Total | LAB
S--T | % | Sulfur Total | MS41 | ALS Minerals ME-MS41 | ALS Minerals ME-MS41. A prepared sample is digested with aqua regia in a graphite heating block. After cooling; the resulting solution is diluted to with deionized water; mixed and analyzed by ICP-AES. Samples are then analyzed by ICP-MS for the remaining suite of elements. | 0.01 | Y | Y | Y | 0 |  |  |  | 2024-01-29 |  | EMS |  | DimensionlessRatio | % | N | (ratio) |  | Sulfur Total (ratio) | 7704-34-9 |  | EMS code: S--T | Chemical | Total | LAB
S--T | mg/L | Sulfur Total | METN | MET-T-NP-CCMS-VA | Ultra trace metals in water are analyzed by CRC ICPMS, based on US EPA Method 6020B (July 2014). The detection limits provided can only be met for undigested samples. This procedure is intended for colorless, non-turbid, acid-preserved water samples (i.e. pristine water samples), having turbidity < 1 NTU and no odor. Where turbidity exceeds 1 NTU, and/or the sample is colored and has an odor, results may be biased low compared to true Total Metals concentrations. ALS recommends that turbidity analysis be requested on samples submitted for this test to aid with interpretation of results. Where turbidity is <1NTU, undigested metals are equivalent to total metals concentrations. | 0.5 | Y | Y | Y | 280 | 8/6/2019 | 4/4/2024 | 4/29/2024 4:53:06 PM | 2021-06-16 |  | EMS |  | Concentration | mg/L | Y | (fl. conc.) |  | Sulfur Total (fl. conc.) | 7704-34-9 |  | EMS code: S--T | Chemical | Total | LAB
S--T | ug/m3 | Sulfur Total | EQ40 | EQUIS (mg/m3) | EQUIS (mg/m3) | 0 | N | Y | Y | 49 |  |  | 10/22/1996 10:10:52 PM | 1996-04-12 | 6/3/1998 6:51:38 PM | EMS |  | AirConcentration | µg/m³ | Y | (air conc.) |  | Sulfur Total (air conc.) | 7704-34-9 |  | EMS code: S--T | Chemical | Total | LAB
S--T | mg/L | Sulfur Total | MSEA | Metals in Seawater (undigested) by (CRC) ICP-MS | This method describes the multi-elemental determination of trace elements by Collision/Reaction Cell (CRC) ICP-MS. The method measures ions produced by a radio-frequency inductively coupled plasma. Analyte species originating in a liquid are pneumatically nebulized and the resulting aerosol transported by argon gas into the plasma torch. The ions produced are entrained in the plasma gas and extracted, by means of a differentially pumped vacuum interface, into a mass spectrometer. The ions flow through a Collision/Reaction Cell (CRC) where depending on the mode of analysis either the ions passes through the cell unhindered (no gas - classical ICP-MS mode), or collisions/reactions chemistry takes place between the ions and a neutral gas such as hydrogen or helium. Any polyatomic ions which give interference on the analytes are removed. The ions are then sorted according to their mass-to-charge ratios by a quadrupole mass spectrometer. The ions are then sorted according to their mass-to-charge ratios by a quadrupole mass spectrometer. The ions transmitted through the quadrupole are quantified by a channel electron multiplier and the ion information processed by a data handling system. Although most polyatomic interference is removed, any additional interference must be assessed and valid corrections applied or the data flagged to indicate problems. Interference correction must include compensation for background ions contributed by the plasma gas, reagents, and constituents of the sample matrix. Instrument drift as well as suppressions or enhancements of instrument response caused by the sample matrix must be corrected for by the use of internal standards. | 6000 | Y | Y | Y | 432 | 8/11/2009 11:57:00 PM | 4/29/2024 | 6/7/2024 10:03:16 AM | 2009-09-08 |  | RDGIBSON |  | Concentration | mg/L | Y | (fl. conc.) |  | Sulfur Total (fl. conc.) | 7704-34-9 |  | EMS code: S--T | Chemical | Total | LAB
S001 | ug/g | Simazine | ZN03 | DCM+Ac.Extr:GC/MS | DCM+Ac.Extr:GC/MS EPA 8270E (mod) | 0.05 | Y | Y | Y | 16 | 1/25/2022 | 11/2/2022 4:41:00 AM | 1/7/2023 3:12:11 PM | 1996-04-12 | 11/16/1999 2:27:40 PM | EMS | NPEPPIN | MassPerMass | mg/kg | Y | (mass conc.) |  | Simazine (mass conc.) | 122-34-9 | Pesticides; | Pesticides; EMS code: S001 | Chemical |  | LAB
S001 | mg/L | Simazine | F056 | DCM extr,HRGC-NPD | Effluent samples are extracted with Dichloromethane. The extract is concentrated and analysed using High Resolution Gas Chromatography with a Thermionic Specific Detector (TSD), alternately referred to as a Nitrogen Phosphorus Detector (NPD). | 0.0005 | Y | Y | Y | 12 | 4/3/2000 10:51:00 AM | 2/16/2001 3:36:00 PM | 9/11/2002 9:34:57 AM | 1998-10-02 |  | NPEPPIN |  | Concentration | mg/L | Y | (fl. conc.) |  | Simazine (fl. conc.) | 122-34-9 | Pesticides; | Pesticides; EMS code: S001 | Chemical |  | LAB
S001 | ug/g | Simazine | FS61 | DCM +MeOH extr,HRGC-NPD | Sediment samples are shaken with a mixture of Dichloromethane and Methanol, filtered then re-extracted with Dichloromethane. The extract is concentrated and analysed using High Resolution Gas Chromatography with a Thermionic Specific Detector (TSD), alternately referred to as a Nitrogen Phosphorus Detector (NPD). | 0.01 | Y | Y | Y | 34 | 4/20/2002 3:40:00 PM | 7/16/2002 3:20:00 PM | 2/26/2004 1:05:31 AM | 1998-10-02 |  | NPEPPIN |  | MassPerMass | mg/kg | Y | (mass conc.) |  | Simazine (mass conc.) | 122-34-9 | Pesticides; | Pesticides; EMS code: S001 | Chemical |  | LAB
S001 | mg/L | Simazine | ZN01 | DCM+Ac.Extr:GC/MS | DCM+Ac.Extr:GC/MS | 0.0005 | Y | Y | Y | 28 | 7/28/2005 7:03:00 AM | 3/29/2012 10:15:00 PM | 5/3/2012 9:27:52 AM | 1996-04-12 | 11/16/1999 4:06:28 PM | EMS | NPEPPIN | Concentration | mg/L | Y | (fl. conc.) |  | Simazine (fl. conc.) | 122-34-9 | Pesticides; | Pesticides; EMS code: S001 | Chemical |  | LAB
S001 | mg/L | Simazine | E751 | E751D Pesticides in Water by LC-MS/MS (POS1 List) | E751D Pesticides in Water by LC-MS/MS (POS1 List) (MOECC E3501) | 0.0001 | Y | Y | Y | 39 | 12/9/2021 | 12/31/2021 | 2/15/2022 9:50:06 AM | 2021-12-14 |  | EMS |  | Concentration | mg/L | Y | (fl. conc.) |  | Simazine (fl. conc.) | 122-34-9 | Pesticides; | Pesticides; EMS code: S001 | Chemical |  | LAB
S001 | mg/L | Simazine | E660 | EPA 8270E (mod) | EPA 8270E (mod) Miscellaneous Pesticides by GC-MS | 0.0001 | Y | Y | Y | 68 | 12/7/2021 | 9/12/2022 | 12/8/2022 12:57:06 PM | 2021-12-09 | 12/9/2021 3:22:12 PM | EMS | EMS | Concentration | mg/L | Y | (fl. conc.) |  | Simazine (fl. conc.) | 122-34-9 | Pesticides; | Pesticides; EMS code: S001 | Chemical |  | LAB
S003 | ug/g | Sulfotep | ZN03 | DCM+Ac.Extr:GC/MS | DCM+Ac.Extr:GC/MS EPA 8270E (mod) | 0.01 | Y | Y | Y | 21 | 4/20/2002 10:25:00 AM | 11/2/2022 4:41:00 AM | 1/7/2023 3:12:11 PM | 1996-04-12 | 11/16/1999 2:27:40 PM | EMS | NPEPPIN | MassPerMass | mg/kg | Y | (mass conc.) |  | Sulfotep (mass conc.) | 3689-24-5 | Pesticides; | Pesticides; EMS code: S003 | Chemical |  | LAB
S003 | mg/L | Sulfotep | ZN01 | DCM+Ac.Extr:GC/MS | DCM+Ac.Extr:GC/MS | 0.0001 | Y | Y | Y | 52 | 10/3/2002 | 3/29/2012 10:15:00 PM | 5/3/2012 9:27:52 AM | 1996-04-12 | 11/16/1999 4:06:28 PM | EMS | NPEPPIN | Concentration | mg/L | Y | (fl. conc.) |  | Sulfotep (fl. conc.) | 3689-24-5 | Pesticides; | Pesticides; EMS code: S003 | Chemical |  | LAB
S004 | % | Silt (0.004mm - 0.0312mm) | PSP | PSA Pipet Detail | Particle size distribution is determined by a combination of techniques. Dry sieving is performed for coarse particles, wet sieving for sand particles and the pipette sedimentation method for clay particles. Burt, R. (2009). Soil Survey Field and Laboratory Methods Manual. Soil Survey Investigations Report No. 5. Method 3.2.1.2.2. United States Department of Agriculture Natural Resources Conservation Service. | 1 | Y | Y | Y | 35 | 9/16/2017 | 10/4/2018 | 10/10/2018 2:46:10 PM | 2018-01-09 |  | JKROGH |  | DimensionlessRatio | % | N | (ratio) |  | Silt (0.004mm - 0.0312mm) (ratio) |  | Sediment Size; | Sediment Size; EMS code: S004 | Physical |  | LAB
S006 | mg/L | Sandaracopimaric Acid | F009 | pH 9,Diethyl Ether:HRGC/LRMS | Effluent samples are first adjusted to pH 9 and then extracted with Diethyl Ether. The extracts are rotary evaporated and the resulting residues, containing the Resin Acids and other solvent soluble organics are then redissolved in Methanol, transferred to small glass vials, methylated with Diazomethane then evaporated to dryness under Nitrogen. [Methylation Reaction R-COOH(Resin/Fatty Acids) + CH2N2(Diazomethane) = R-COOCH3(Methyl ester) + N2) The residue is re-dissolved in Methyl t-Butyl Ether (MTBE) and injected on a High Resolution Gas Chromatograph/Low Resolution Mass Spectrometer (HRGC/LRMS) system for analysis. LRMS data is acquired in Total Ion Mode (TIM) and quantitation is performed using selected mass ions. Positive identification of the various Resin Acids is made on the basis of a comparison of absolute retention times to those of the external standards and computer based mass spectral mass/intensity library matching routines. | 0.005 | Y | Y | Y | 198 | 12/27/1996 4:16:00 PM | 5/13/2003 | 4/1/2004 1:08:47 AM | 1996-09-16 | 6/3/1998 6:51:44 PM | MPAK |  | Concentration | mg/L | Y | (fl. conc.) |  | Sandaracopimaric Acid (fl. conc.) | 471-74-9 | Organics; | Organics; EMS code: S006 | Chemical |  | LAB
S006 | mg/L | Sandaracopimaric Acid | P035 | DCM-DIAZO-FLOR | DCM-DIAZO-FLOR | 0.0005 | Y | Y | Y | 22 |  |  | 10/24/1996 12:44:24 PM | 1996-04-12 | 6/3/1998 6:51:38 PM | EMS |  | Concentration | mg/L | Y | (fl. conc.) |  | Sandaracopimaric Acid (fl. conc.) | 471-74-9 | Organics; | Organics; EMS code: S006 | Chemical |  | LAB
S006 | ug/g | Sandaracopimaric Acid | FS05 | MeOH,DCM extr:HRGC/LRMS | Sediment samples are extracted by shaking with Methanol in a screw cap centrifuge bottle. The extracts are rotary evaporated and the resulting residues, containing the Resin Acids and other solvent soluble organics, are then redissolved in Methanol, transferred to small glass vials, methylated with Diazomethane then evaporated to dryness under Nitrogen. [Methylation Reaction R-COOH(Resin/Fatty Acids) + CH2N2 (Diazomethane) = R-COOCH3 (Methyl ester) + N2] The residue is re-dissolved in Methyl t-Butyl Ether (MTBE) and injected on a High Resolution Gas Chromatograph / Low Resolution Mass Spectrometer (HRGC/LRMS) system for analysis. LRMS data is acquired in Total Ion Mode (TIM) and quantitation is performed using selected mass ions. Positive identification of the various Resin Acids is made on the basis of a comparison of absolute retention times to those of the external standards and computer based mass spectral mass/intensity library matching routines. | 0.2 | Y | Y | Y | 5 | 9/14/2001 3:29:00 PM | 12/3/2002 | 3/27/2004 1:09:21 AM | 1996-09-16 | 6/3/1998 6:51:44 PM | MPAK |  | MassPerMass | mg/kg | Y | (mass conc.) |  | Sandaracopimaric Acid (mass conc.) | 471-74-9 | Organics; | Organics; EMS code: S006 | Chemical |  | LAB
S006 | mg/L | Sandaracopimaric Acid | RFAS | Separation of resin and fatty acids (RFAs) in water by GC/MS | Scientific Principle The separation of resin and fatty acids (RFAs) in water samples is performed using gas chromatography/mass spectrometry (GC/MS) in the Selected Ion Monitoring (SIM) mode. The separation is based on compound size; boiling point and polarity. Analyte concentration is calculated using the internal standard method. RFA compounds are extracted from water samples using a liquid-liquid solvent extraction with DCM. Extraction occurs after the sample is brought to a pH of ~ 2 and placed in a rotary tumbler (at approx. 40 rpm) for 1 hour. The resulting extract is transferred to a vial; dried with acidified baked sodium sulfate; methylated with diazomethane and concentrated to 1mL under nitrogen prior to analysis. The resulting extracts are vialed and fortified with labeled internal standard compounds prior to analysis by gas chromatography with a (50%-Phenyl)-methylpolysiloxane (DB-17) column and a mass spectrometer detector in SIM Mode. | 0.006 | Y | Y | Y | 1 | 9/17/2021 7:49:00 PM | 9/17/2021 7:49:00 PM | 2/1/2022 2:38:14 PM | 2022-01-19 |  | EMS |  | Concentration | mg/L | Y | (fl. conc.) |  | Sandaracopimaric Acid (fl. conc.) | 471-74-9 | Organics; | Organics; EMS code: S006 | Chemical |  | LAB
S006 | mg/L | Sandaracopimaric Acid | X239 | Gas Chromatography | Gas Chromatography | 0.001 | Y | Y | Y | 12 |  |  | 10/23/1996 11:56:37 PM | 1996-04-12 | 6/3/1998 6:51:38 PM | EMS |  | Concentration | mg/L | Y | (fl. conc.) |  | Sandaracopimaric Acid (fl. conc.) | 471-74-9 | Organics; | Organics; EMS code: S006 | Chemical |  | LAB
S006 | mg/L | Sandaracopimaric Acid | P030 | Resin Acids | The water sample is acidified and extracted with dichloro-methane. The raw extract is concentrated and derivatized with diazomethane to produce the corresponding methyl ester derivatives. (Suitable substitute derivatives may be produced by other available techniques.) If required, the extracts are cleaned up by Florisil column chromatography. The derivatives are analyzed by gas chromatography with flame ionization detection. | 0.001 | Y | Y | Y | 476 | 8/26/2014 | 8/15/2017 | 8/15/2017 4:37:09 PM | 1996-04-12 | 6/3/1998 6:51:38 PM | EMS |  | Concentration | mg/L | Y | (fl. conc.) |  | Sandaracopimaric Acid (fl. conc.) | 471-74-9 | Organics; | Organics; EMS code: S006 | Chemical |  | LAB
S006 | mg/L | Sandaracopimaric Acid | P033 | Resin&Fatty Acids,GC/MS | Resin&Fatty Acids;GC/MS | 0.001 | Y | Y | Y | 34 |  |  | 6/7/2024 5:30:18 PM | 1996-04-12 | 6/3/1998 6:51:38 PM | EMS |  | Concentration | mg/L | Y | (fl. conc.) |  | Sandaracopimaric Acid (fl. conc.) | 471-74-9 | Organics; | Organics; EMS code: S006 | Chemical |  | LAB
S006 | mg/L | Sandaracopimaric Acid | E677 | Resins & Fatty Acids are methylated then analyzed by GCMS | EPA 8270E; AB Env. 129; NCASI RA/FA-85.02 (mod) Resins & Fatty Acids are methylated then analyzed by GCMS. Total fatty acids is the sum of arachidic acid; behenic acid; lauric acid; lignoceric acid; linoleic acid; linolenic acid; myristic acid; oleic acid; palmitic acid; and stearic acid. Total resin acids is the sum of abietic acid; 12-chlorodehydroabietic acid; 14-chlorodehydroabietic acid; dehydroabietic acid; 12;14-dichlorodehydroabietic acid; isopimaric acid; levopimaric acid; neoabietic acid; palustric acid; pimaric acid and sandaracopimaric acid. Total resin acids (BC WWQG) is the sum of abietic acid; neoabietic acid; pimaric acid; isopimaric acid; and sandaracopimaric acid. | 0.001 | Y | Y | Y | 12 | 11/28/2022 | 6/15/2024 | 7/8/2024 11:10:02 AM | 2021-12-30 |  | EMS |  | Concentration | mg/L | Y | (fl. conc.) |  | Sandaracopimaric Acid (fl. conc.) | 471-74-9 | Organics; | Organics; EMS code: S006 | Chemical |  | LAB
S010 | mg/L | Styrene | F105 | Static Headspace/GCMS | Static Headspace/GCMS | 0.0002 | Y | Y | Y | 299 | 8/23/2000 2:29:00 PM | 7/6/2024 9:57:00 PM | 7/22/2024 3:17:33 PM | 2001-02-23 |  | NPEPPIN |  | Concentration | mg/L | Y | (fl. conc.) |  | Styrene (fl. conc.) | 100-42-5 | Volatile Organic Compounds (VOCs); | Volatile Organic Compounds (VOCs); EMS code: S010 | Chemical |  | LAB
S010 | ug/g | Styrene | X263 | Purge and trap GC/MS | Purge and trap GC/MS | 0.001 | Y | Y | Y | 31 | 11/19/2004 | 4/4/2006 | 4/27/2006 1:02:37 AM | 1996-04-12 | 6/3/1998 6:51:42 PM | EMS |  | MassPerMass | mg/kg | Y | (mass conc.) |  | Styrene (mass conc.) | 100-42-5 | Volatile Organic Compounds (VOCs); | Volatile Organic Compounds (VOCs); EMS code: S010 | Chemical |  | LAB
S010 | mg/L | Styrene | MS04 | Static Headspace GC/MS (Water) | (Water) A static headspace procedure is used as a means of driving the analytes out of the matrix into a headspace. Once equilibrium is achieved, a fixed volume of headspeace is taken for subsequent determination by Gas Chromatography - Mass Spectroscopy (GC/MS) using Selected Ion Monitoring. | 0.0005 | Y | Y | Y | 75 | 8/13/2009 | 6/28/2024 12:29:00 AM | 7/9/2024 3:31:02 PM | 2009-12-23 |  | RDGIBSON |  | Concentration | mg/L | Y | (fl. conc.) |  | Styrene (fl. conc.) | 100-42-5 | Volatile Organic Compounds (VOCs); | Volatile Organic Compounds (VOCs); EMS code: S010 | Chemical |  | LAB
S010 | mg/L | Styrene | F007 | Static Headspace:HRGC/LRMS | Static headspace analysis is a method for the determination of volatile compounds in liquids and/or solids by analysing the vapour phase of a sample in a closed system; in this case a sealed vial. Static Headspace is sometimes referred to as Equilibrium Headspace because the vapour phase above the sample in a sealed vial is brought into thermodynamic equilibrium with the sample prior to analysis. Volatile organic compounds are thermally extracted from the sample matrix by Static Headspace sampling and introduced directly to a thick-film megabore GC column. Compounds eluting from the GC column are introduced into a Mass Spectrometer and target compounds are identified by a comparison of both retention times and mass spectra with those of reference standards stored in a computerized data system. Reference retention times and mass spectra for analytes are obtained by the measurement of calibration standards under precisely the same extraction conditions used for samples. The concentration of each identified component is measured by relating the response of the sample quantitation mass ions to the response of the quantitation mass ions of a series of external standards. A deuterated surrogate recovery standard is added to provide information on matrix effects. | 0.0005 | Y | Y | Y | 994 | 4/15/2011 | 8/14/2024 | 8/16/2024 4:05:21 PM | 2012-03-28 |  | RDGIBSON |  | Concentration | mg/L | Y | (fl. conc.) |  | Styrene (fl. conc.) | 100-42-5 | Volatile Organic Compounds (VOCs); | Volatile Organic Compounds (VOCs); EMS code: S010 | Chemical |  | LAB
S010 | mg/L | Styrene | X384 | Purge and trap GC/MS | The sample is analyzed by purge and trap gas chromatography with detection by mass spectrometry. The procedure involves purging the volatile compounds from the sample with an inert gas, and trapping them on a solid sorbent. When purging of the sample is complete, the trap is heated and the compounds are transferred to the gas chromatographic column. Analysis is then accomplished by separation of the components by gas chromatography with detection by mass spectrometry. | 0.0005 | Y | Y | Y | 153 | 3/6/2002 | 7/6/2010 7:02:00 AM | 7/19/2010 1:33:14 PM | 1996-04-12 | 6/3/1998 6:51:38 PM | EMS |  | Concentration | mg/L | Y | (fl. conc.) |  | Styrene (fl. conc.) | 100-42-5 | Volatile Organic Compounds (VOCs); | Volatile Organic Compounds (VOCs); EMS code: S010 | Chemical |  | LAB
S010 | ug/g | Styrene | MS02 | Purge&Trap GC/MS (VOC-S) | Purge&Trap GC/MS (VOC-S) | 0.05 | Y | Y | Y | 16 | 1/14/2022 | 10/28/2022 5:04:00 PM | 1/11/2024 1:07:08 PM | 2022-07-27 |  | EMS |  | MassPerMass | mg/kg | Y | (mass conc.) |  | Styrene (mass conc.) | 100-42-5 | Volatile Organic Compounds (VOCs); | Volatile Organic Compounds (VOCs); EMS code: S010 | Chemical |  | LAB
S010 | mg/L | Styrene | MS01 | Purge&Trap GC/MS (VOC-W) | Purge&Trap GC/MS (VOC-W) | 0.0001 | Y | Y | Y | 1019 | 4/15/2002 | 6/4/2024 10:24:00 AM | 6/13/2024 5:01:27 PM | 2002-08-12 | 8/12/2002 4:13:56 PM | NPEPPIN | NPEPPIN | Concentration | mg/L | Y | (fl. conc.) |  | Styrene (fl. conc.) | 100-42-5 | Volatile Organic Compounds (VOCs); | Volatile Organic Compounds (VOCs); EMS code: S010 | Chemical |  | LAB
S010 | ug/g | Styrene | MEOH | MeOH extrac,GC/FID,nC6-nC10 | MeOH extrac;GC/FID;nC6-nC10. EPA8260B; 5021; 5035; BC MOE | 0.05 | Y | Y | Y | 49 | 5/21/2015 | 8/11/2019 | 8/13/2019 10:56:08 AM | 2015-05-26 |  | RDGIBSON |  | MassPerMass | mg/kg | Y | (mass conc.) |  | Styrene (mass conc.) | 100-42-5 | Volatile Organic Compounds (VOCs); | Volatile Organic Compounds (VOCs); EMS code: S010 | Chemical |  | LAB
S010 | ug/g | Styrene | FS04 | Static Headspace:HRGC/LRMS | Static Headspace anaylsis is a method for the determination of volatile compounds in liquids and/or solids by analysing the vapour phase of a sample in a closed system; in this case a sealed vial. Static Headspace is sometimes referred to as Equilibrium Headspace because the vapour phase above the sample in a sealed vial is brought into thermodynamic equilibrium with the sample prior to analysis. Volatile organic compounds are thermally extracted from the sample matrix by Static Headspace sampling and introduced directly to a thick-film megabore GC column. Compounds eluting from the GC column are introduced into a Mass Spectrometer and target compounds are identified by a comparison of both retention times and mass spectra with those of reference standards stored in a computerized data system. Reference retention times and mass spectra for analytes are obtained by the measurement of calibration standards under precisely the same extraction conditions used for samples. The concentration of each identified component is measured by relating the response of the sample quantitation mass ions to the response of the quantitation mass ions of a series of external standards. A deuterated surrogate recovery standard is added to provide information on matrix effects. | 0.1 | Y | Y | Y | 18 | 3/4/2013 6:06:00 PM | 3/7/2013 3:50:00 AM | 3/19/2013 1:15:56 PM | 2010-11-16 |  | RDGIBSON |  | MassPerMass | mg/kg | Y | (mass conc.) |  | Styrene (mass conc.) | 100-42-5 | Volatile Organic Compounds (VOCs); | Volatile Organic Compounds (VOCs); EMS code: S010 | Chemical |  | LAB
S031 | % | Silt (0.0312mm - 0.063mm) | PSP | PSA Pipet Detail | Particle size distribution is determined by a combination of techniques. Dry sieving is performed for coarse particles, wet sieving for sand particles and the pipette sedimentation method for clay particles. Burt, R. (2009). Soil Survey Field and Laboratory Methods Manual. Soil Survey Investigations Report No. 5. Method 3.2.1.2.2. United States Department of Agriculture Natural Resources Conservation Service. | 1 | Y | Y | Y | 35 | 9/16/2017 | 10/4/2018 | 10/10/2018 2:46:10 PM | 2018-01-09 |  | JKROGH |  | DimensionlessRatio | % | N | (ratio) |  | Silt (0.0312mm - 0.063mm) (ratio) |  | Sediment Size; | Sediment Size; EMS code: S031 | Physical |  | LAB
S063 | % | Sand (0.0063mm - 0.125mm) | PSP | PSA Pipet Detail | Particle size distribution is determined by a combination of techniques. Dry sieving is performed for coarse particles, wet sieving for sand particles and the pipette sedimentation method for clay particles. Burt, R. (2009). Soil Survey Field and Laboratory Methods Manual. Soil Survey Investigations Report No. 5. Method 3.2.1.2.2. United States Department of Agriculture Natural Resources Conservation Service. | 1 | Y | Y | Y | 35 | 9/16/2017 | 10/4/2018 | 10/10/2018 2:46:10 PM | 2018-01-09 |  | JKROGH |  | DimensionlessRatio | % | N | (ratio) |  | Sand (0.0063mm - 0.125mm) (ratio) |  | Sediment Size; | Sediment Size; EMS code: S063 | Chemical |  | LAB
S101 | % | Sand (1.00mm - 2.00mm) | PSP | PSA Pipet Detail | Particle size distribution is determined by a combination of techniques. Dry sieving is performed for coarse particles, wet sieving for sand particles and the pipette sedimentation method for clay particles. Burt, R. (2009). Soil Survey Field and Laboratory Methods Manual. Soil Survey Investigations Report No. 5. Method 3.2.1.2.2. United States Department of Agriculture Natural Resources Conservation Service. | 1 | Y | Y | Y | 35 | 9/16/2017 | 10/4/2018 | 10/10/2018 2:46:10 PM | 2018-01-09 |  | JKROGH |  | DimensionlessRatio | % | N | (ratio) |  | Sand (1.00mm - 2.00mm) (ratio) |  | Sediment Size; | Sediment Size; EMS code: S101 | Chemical |  | LAB
S125 | % | Sand (0.125mm - 0.250mm) | PSP | PSA Pipet Detail | Particle size distribution is determined by a combination of techniques. Dry sieving is performed for coarse particles, wet sieving for sand particles and the pipette sedimentation method for clay particles. Burt, R. (2009). Soil Survey Field and Laboratory Methods Manual. Soil Survey Investigations Report No. 5. Method 3.2.1.2.2. United States Department of Agriculture Natural Resources Conservation Service. | 1 | Y | Y | Y | 35 | 9/16/2017 | 10/4/2018 | 10/10/2018 2:46:10 PM | 2018-01-09 |  | JKROGH |  | DimensionlessRatio | % | N | (ratio) |  | Sand (0.125mm - 0.250mm) (ratio) |  | Sediment Size; | Sediment Size; EMS code: S125 | Chemical |  | LAB
S250 | % | Sand (0.250mm - 0.500mm) | PSP | PSA Pipet Detail | Particle size distribution is determined by a combination of techniques. Dry sieving is performed for coarse particles, wet sieving for sand particles and the pipette sedimentation method for clay particles. Burt, R. (2009). Soil Survey Field and Laboratory Methods Manual. Soil Survey Investigations Report No. 5. Method 3.2.1.2.2. United States Department of Agriculture Natural Resources Conservation Service. | 1 | Y | Y | Y | 35 | 9/16/2017 | 10/4/2018 | 10/10/2018 2:46:10 PM | 2018-01-09 |  | JKROGH |  | DimensionlessRatio | % | N | (ratio) |  | Sand (0.250mm - 0.500mm) (ratio) |  | Sediment Size; | Sediment Size; EMS code: S250 | Chemical |  | LAB
S500 | % | Sand (0.500mm - 1.00mm) | PSP | PSA Pipet Detail | Particle size distribution is determined by a combination of techniques. Dry sieving is performed for coarse particles, wet sieving for sand particles and the pipette sedimentation method for clay particles. Burt, R. (2009). Soil Survey Field and Laboratory Methods Manual. Soil Survey Investigations Report No. 5. Method 3.2.1.2.2. United States Department of Agriculture Natural Resources Conservation Service. | 1 | Y | Y | Y | 35 | 9/16/2017 | 10/4/2018 | 10/10/2018 2:46:10 PM | 2018-01-09 |  | JKROGH |  | DimensionlessRatio | % | N | (ratio) |  | Sand (0.500mm - 1.00mm) (ratio) |  | Sediment Size; | Sediment Size; EMS code: S500 | Chemical |  | LAB
SALM | MPN/100mL | Salmonella | X015 | MPN Confirmed | Multiple-tube fermentation technique is the examination of replicate tubes and dilutions reported in terms of Most Probable Number (MPN) of organisms present. This number, based on certain probability formulas, is an estimate of the mean density of bacteria in the sample. MPN method is not influenced by turbidity. | 2 | Y | Y | Y | 498 |  |  | 2/16/1999 9:15:04 AM | 1996-04-12 | 6/3/1998 6:51:38 PM | EMS |  | MicrobialFormation | MPN/100mL | N | (microb.) |  | Salmonella (microb.) |  | Bacteriological; | Bacteriological; EMS code: SALM | Chemical |  | LAB
SAND | % | % sand by hydrometer | HYDR | Hydrometer Analysis | Hydrometer Analysis | 2 | Y | Y | Y | 144 | 10/4/2013 3:54:00 PM | 11/1/2022 | 1/11/2024 1:07:08 PM | 2010-11-16 |  | RDGIBSON |  | DimensionlessRatio | % | N | (ratio) |  | % sand by hydrometer (ratio) |  | Sediment Size; | Sediment Size; EMS code: SAND | Chemical |  | LAB
SAR | None | CCME Sodium Adsorption Ratio (SAR) | SAR2 | Sodium Adsorption Ratio | CCME Sodium Adsorption Ratio (SAR) | 0.1 | Y | Y | Y | 19 | 1/11/2022 | 2/7/2022 | 6/7/2024 5:30:17 PM | 2022-06-27 |  | EMS |  | None | None | N | (unitless) |  | CCME Sodium Adsorption Ratio (SAR) (unitless) |  |  | EMS code: SAR | Chemical |  | LAB
SATPR | % | % saturation | SATP | Saturation Percentage | Saturation Percentage (SP) is the total volume of water present in a saturated paste (in mL) divided by the dry weight of the sample (in grams), expressed as a percentage, as described in "Soil Sampling and Methods of Analysis" by M. Carter. | 1 | Y | Y | Y | 16 | 1/10/2022 | 2/7/2023 | 1/11/2024 1:07:08 PM | 2018-01-09 |  | JKROGH |  | DimensionlessRatio | % | N | (ratio) |  | % saturation (ratio) |  | Physical Properties; | Physical Properties; EMS code: SATPR | Chemical |  | LAB
SB-D | mg/L | Antimony Dissolved | F068 | HNO3dig,TeflonTube/Autoclave,ICP-MS | This method is applicable to determiniation of dissolved, extractable, and total recoverable metals in ground waters, surface waters, drinking waters, and wastewaters. Dissolved, extractable, and total recoverable sample fractions are acidified to pH <1.5 with analytical grade nitric acid (HNO3) to prevent the metals from adsorbing to the container material or precipitating. Extractable and total recoverable water samples are collected in 250 ml acid washed polyethylene bottles and are preserved with Nitric acid (0.5 ml acid for 100 mLs of sample). All preserved samples are stored at room temperature and can be held for up to 6 months. Unfiltered, un-preserved samples are stored at 4C for no more than 30 days. For dissolved and extractable water fractions, no sample preparation is necessary. Samples are simply poured in autosampler tubes (i.e. 15 ml centrifuge tubes), and placed on the autosampler and analyzed. Samples for total recoverable metals are digested with Nitric acid in special Teflon tubes in an autoclave prior to analyses. | 5e-06 | Y | Y | Y | 1398 | 6/1/1999 1:22:00 PM | 4/17/2015 1:06:00 PM | 8/11/2021 12:04:11 PM | 1999-06-28 | 6/30/1999 1:05:24 PM | NPEPPIN | NPEPPIN | Concentration | mg/L | Y | (fl. conc.) |  | Antimony Dissolved (fl. conc.) | 7440-36-0 | Metals - Dissolved; | Metals - Dissolved; EMS code: SB-D | Chemical | Dissolved | LAB
SB-D | mg/L | Antimony Dissolved | METC | MET-D-NP-CCMS-VA | Ultra trace metals in water are analyzed by CRC ICPMS, based on US EPA Method 6020B (July 2014). This procedure is intended for pristine field-filtered acid-preserved water samples. The detection limits (LOR) for this test are based on lab instrumental analysis only, not including filtration. ALS-supplied field filtration equipment does not support these LOR. Therefore, because of the high probability of false positives due to filtration, it is strongly recommended that a filtration blank be analysed to aid in data interpretation. | 2e-05 | Y | Y | Y | 1776 | 4/3/2012 | 8/15/2024 | 8/23/2024 9:50:35 AM | 2019-08-20 |  | KATFRASE |  | Concentration | mg/L | Y | (fl. conc.) |  | Antimony Dissolved (fl. conc.) | 7440-36-0 | Metals - Dissolved; | Metals - Dissolved; EMS code: SB-D | Chemical | Dissolved | LAB
SB-D | mg/L | Antimony Dissolved | X170 | ICP | This code is a catch-all for a large variety of ICP methods including water and also solids digested into solutions as an intermediate result. Different procedures will have greatly different detection limits. (a) For air particulates, preliminary results are reported in mg/L and loading results as mg/dm2/d. (b) For Hi Vol samples preliminiary results are same as reported in mg/L and loading results as ug/m3. (c) Note that for Rain samples there is no filtration nor digestion steps. | 0.015 | Y | Y | Y | 13486 | 4/11/1996 6:02:00 PM | 11/6/2018 | 8/26/2024 2:41:32 PM | 1996-04-12 | 6/3/1998 6:51:38 PM | EMS |  | Concentration | mg/L | Y | (fl. conc.) |  | Antimony Dissolved (fl. conc.) | 7440-36-0 | Metals - Dissolved; | Metals - Dissolved; EMS code: SB-D | Chemical | Dissolved | LAB
SB-D | mg/L | Antimony Dissolved | MWID | Metals in Water by ICPMS (digested) | Metals in Water by ICPMS (digested) Based on EPA 200.8 | 0.0005 | Y | Y | Y | 31870 | 10/26/2018 5:20:00 PM | 10/26/2018 5:20:00 PM | 7/31/2024 1:46:41 PM | 2009-12-24 |  | RDGIBSON |  | Concentration | mg/L | Y | (fl. conc.) |  | Antimony Dissolved (fl. conc.) | 7440-36-0 | Metals - Dissolved; | Metals - Dissolved; EMS code: SB-D | Chemical | Dissolved | LAB
SB-D | mg/L | Antimony Dissolved | F023 | HNO3,GFAA | Dissolved and extractable water samples are ready for analysis once collected. Total samples are autoclaved with HNO3 and hydrochloric acid (HCl) in capped Teflon vials. Analysis is performed via Inductively Coupled Argon Plasma Atomic Emission Spectrometry (ICP-AES) (simultaneous multielement analysis), Graphite Furnace Atomic Absorption Spectrometry (GF-AAS) (low level analysis) and/or Flame Atomic Absorption Spectrometry (F-AAS). Note that extractable samples are acidified with 0.5 mL HNO3 per 100 mL of sample at time of sample collection. | 0.001 | Y | Y | Y | 90 | 4/8/1997 3:55:00 PM | 9/16/2003 3:18:00 PM | 11/28/2003 3:22:15 PM | 1997-04-02 | 6/3/1998 6:51:46 PM | MPAK |  | Concentration | mg/L | Y | (fl. conc.) |  | Antimony Dissolved (fl. conc.) | 7440-36-0 | Metals - Dissolved; | Metals - Dissolved; EMS code: SB-D | Chemical | Dissolved | LAB
SB-D | mg/L | Antimony Dissolved | X562 | Dissolved metals: FF, FA, HNO3, HR-ICPMS | Ultra trace metals in water are analyzed by high resolution inductively coupled plasma mass spectrometry (HR-ICPMS), based on US EPA Method 200.8, (Oct. 1999). This procedure is intended for pristine field-filtered acid-preserved water samples. The detection limits (LOR) for this test are based on lab instrumental analysis only, not including filtration. ALS-supplied field filtration equipment does not support these LOR. Therefore, because of the high probability of false positives due to filtration, it is strongly recommended that a filtration blank be analyzed to aid in data interpretation. | 2e-05 | Y | Y | Y | 1264 | 6/8/2015 | 10/19/2023 | 11/9/2023 7:41:27 AM | 2015-10-22 |  | RDGIBSON |  | Concentration | mg/L | Y | (fl. conc.) |  | Antimony Dissolved (fl. conc.) | 7440-36-0 | Metals - Dissolved; | Metals - Dissolved; EMS code: SB-D | Chemical | Dissolved | LAB
SB-D | mg/L | Antimony Dissolved | MSEA | Metals in Seawater (undigested) by (CRC) ICP-MS | This method describes the multi-elemental determination of trace elements by Collision/Reaction Cell (CRC) ICP-MS. The method measures ions produced by a radio-frequency inductively coupled plasma. Analyte species originating in a liquid are pneumatically nebulized and the resulting aerosol transported by argon gas into the plasma torch. The ions produced are entrained in the plasma gas and extracted, by means of a differentially pumped vacuum interface, into a mass spectrometer. The ions flow through a Collision/Reaction Cell (CRC) where depending on the mode of analysis either the ions passes through the cell unhindered (no gas - classical ICP-MS mode), or collisions/reactions chemistry takes place between the ions and a neutral gas such as hydrogen or helium. Any polyatomic ions which give interference on the analytes are removed. The ions are then sorted according to their mass-to-charge ratios by a quadrupole mass spectrometer. The ions are then sorted according to their mass-to-charge ratios by a quadrupole mass spectrometer. The ions transmitted through the quadrupole are quantified by a channel electron multiplier and the ion information processed by a data handling system. Although most polyatomic interference is removed, any additional interference must be assessed and valid corrections applied or the data flagged to indicate problems. Interference correction must include compensation for background ions contributed by the plasma gas, reagents, and constituents of the sample matrix. Instrument drift as well as suppressions or enhancements of instrument response caused by the sample matrix must be corrected for by the use of internal standards. | 2e-05 | Y | Y | Y | 4 | 4/12/2023 | 4/29/2024 | 5/16/2024 2:00:14 PM | 2009-09-08 |  | RDGIBSON |  | Concentration | mg/L | Y | (fl. conc.) |  | Antimony Dissolved (fl. conc.) | 7440-36-0 | Metals - Dissolved; | Metals - Dissolved; EMS code: SB-D | Chemical | Dissolved | LAB
SB-D | mg/L | Antimony Dissolved | 0185 | Hydr.At.Absorp.Spec(H2O) | Hydr.At.Absorp.Spec(H2O) | 0.5 | Y | Y | Y | 6 |  |  | 2/13/1998 12:49:06 AM | 1996-04-12 | 6/3/1998 6:51:38 PM | EMS |  | Concentration | mg/L | Y | (fl. conc.) |  | Antimony Dissolved (fl. conc.) | 7440-36-0 | Metals - Dissolved; | Metals - Dissolved; EMS code: SB-D | Chemical | Dissolved | LAB
SB-D | mg/L | Antimony Dissolved | F032 | Filt.,Acid,ICPMS | There is no acid digestion required for this analysis. A water sample is collected and preserved in 0.2 percent nitric acid. To analyse the extractable metals the sample is shaken, allowed to sit for at least twenty four hours, and a decanted portion is introduced to the spectrometer. When the dissolved metals are required, the samples must first be filtered through a 0.45 micron cellulose acetate filter paper and then followed by preservation in 0.2 percent nitric acid. It is important to note that different laboratories may have very different detection limits for some variables analyzed by ICPMS. | 0.0005 | Y | Y | Y | 81113 | 3/10/1998 2:03:00 PM | 8/26/2024 | 8/26/2024 4:53:05 PM | 1998-05-14 | 8/15/2019 10:59:08 AM | NPEPPIN | KATFRASE | Concentration | mg/L | Y | (fl. conc.) |  | Antimony Dissolved (fl. conc.) | 7440-36-0 | Metals - Dissolved; | Metals - Dissolved; EMS code: SB-D | Chemical | Dissolved | LAB
SB-D | mg/L | Antimony Dissolved | 0050 | Graphite Furnace AA | Graphite Furnace AA | 0.005 | Y | Y | Y | 296 |  |  | 10/24/1996 3:06:41 PM | 1996-04-12 | 6/3/1998 6:51:38 PM | EMS |  | Concentration | mg/L | Y | (fl. conc.) |  | Antimony Dissolved (fl. conc.) | 7440-36-0 | Metals - Dissolved; | Metals - Dissolved; EMS code: SB-D | Chemical | Dissolved | LAB
SB-D | mg/L | Antimony Dissolved | X027 | Ext,APDC,AA | Ext;APDC;AA | 0.005 | Y | Y | Y | 106 |  |  | 10/23/1996 10:13:11 PM | 1996-04-12 | 6/3/1998 6:51:38 PM | EMS |  | Concentration | mg/L | Y | (fl. conc.) |  | Antimony Dissolved (fl. conc.) | 7440-36-0 | Metals - Dissolved; | Metals - Dissolved; EMS code: SB-D | Chemical | Dissolved | LAB
SB-D | mg/L | Antimony Dissolved | X285 | HNO3 Dig:ICP Analysis | HNO3 Dig:ICP Analysis | 0.06 | Y | Y | Y | 670 | 1/30/2002 | 11/13/2002 | 3/18/2004 10:28:34 AM | 1998-05-22 | 6/3/1998 6:51:49 PM | NPEPPIN |  | Concentration | mg/L | Y | (fl. conc.) |  | Antimony Dissolved (fl. conc.) | 7440-36-0 | Metals - Dissolved; | Metals - Dissolved; EMS code: SB-D | Chemical | Dissolved | LAB
SB-D | mg/L | Antimony Dissolved | X544 | FF,FA,HNO3,ICPMS | Analytes nebulized followed by ion entrained in plasma gas and extracted by means of a differentially pumped vacuum interface into a mass spectrometer | 0.001 | Y | Y | Y | 30 |  |  | 10/30/2014 8:31:18 AM | 2001-09-18 | 9/18/2001 1:36:23 PM | NPEPPIN | NPEPPIN | Concentration | mg/L | Y | (fl. conc.) |  | Antimony Dissolved (fl. conc.) | 7440-36-0 | Metals - Dissolved; | Metals - Dissolved; EMS code: SB-D | Chemical | Dissolved | LAB
SB-D | mg/L | Antimony Dissolved | EQ01 | EQUIS (mg/L) | EQUIS (mg/L) | 0.005 | N | Y | Y | 430 |  |  | 10/22/1996 10:25:33 PM | 1996-04-12 | 6/3/1998 6:51:38 PM | EMS |  | Concentration | mg/L | Y | (fl. conc.) |  | Antimony Dissolved (fl. conc.) | 7440-36-0 | Metals - Dissolved; | Metals - Dissolved; EMS code: SB-D | Chemical | Dissolved | LAB
SB-D | mg/L | Antimony Dissolved | X556 | Metals in water, dissolved, "in-bottle" digestion, ICP-SFMS | The whole water sample is filtered through a 0.45 micron cellulose acetate filter paper and the filtrate is stored in a HDPE/Teflon bottle. The container is de-capped in a Class 100 workstation, acidified to 2% (v/v) HNO3 and 1% (v/v) HCl and recapped. The sample is "in-bottle" digested in a convection oven, set at 60 degrees C for 16 hours. The sample is cooled and introduced to an inductively coupled argon plasma-sector field mass spectrometer (ICP-SFMS) by pneumatic nebulization. Each respective element is measured at a specific mass to charge ratio, m/z value expressed in amu (atomic mass units), and at a specified intrument resolution. LR: Low Resolution ~300, MR: Mid Resolution ~4000 and HR: High Resolution ~10000. Ref: Method number 02-2003, National Laboratory for Environmental Testing, CCIW, Environment Canada, Burlington, Requested April 2003. | 1e-06 | Y | Y | Y | 4232 | 6/11/2003 | 12/6/2023 2:04:00 AM | 2/14/2024 9:05:32 AM | 2004-01-19 |  | UMULLER |  | Concentration | mg/L | Y | (fl. conc.) |  | Antimony Dissolved (fl. conc.) | 7440-36-0 | Metals - Dissolved; | Metals - Dissolved; EMS code: SB-D | Chemical | Dissolved | LAB
SB-D | mg/L | Antimony Dissolved | X113 | HNO3/HCl/Perchlor,AA | HNO3/HCl/Perchlor;AA | 0.0005 | Y | Y | Y | 492 |  |  | 12/9/1996 10:18:15 PM | 1996-04-12 | 6/3/1998 6:51:38 PM | EMS |  | Concentration | mg/L | Y | (fl. conc.) |  | Antimony Dissolved (fl. conc.) | 7440-36-0 | Metals - Dissolved; | Metals - Dissolved; EMS code: SB-D | Chemical | Dissolved | LAB
SB-D | mg/L | Antimony Dissolved | X350 | FF,FA:HNO3:ICP | Samples are field-filtered (0.45 um); field-acidified with HNO3; analysis by ICP/AES. | 0.015 | Y | Y | Y | 765 |  |  | 10/1/2003 9:41:33 AM | 1996-04-12 | 6/3/1998 6:51:38 PM | EMS |  | Concentration | mg/L | Y | (fl. conc.) |  | Antimony Dissolved (fl. conc.) | 7440-36-0 | Metals - Dissolved; | Metals - Dissolved; EMS code: SB-D | Chemical | Dissolved | LAB
SB-D | mg/L | Antimony Dissolved | F039 | Field filt + acid, ICP-AES | Dissolved and extractable water samples are ready for analysis once collected. Total samples are autoclaved with HNO3 and hydrochloric acid (HCI) in capped Teflon vials. Analysis is performed via Inductively Coupled Argon Plasma Atomic Emission Spectrometry (ICP-AES)(simultaneous multielement analysis), Graphite Furnace Atomic Absorption Spectrometry (GF-AAS)(low level analysis) and/or Flame Atomic Absorption Spectrometry (F-AAS). | 0.05 | Y | Y | Y | 5937 | 7/29/1996 6:03:00 PM | 7/5/2019 1:21:00 PM | 9/11/2019 8:47:50 AM | 1997-05-08 | 6/3/1998 6:51:46 PM | MPAK |  | Concentration | mg/L | Y | (fl. conc.) |  | Antimony Dissolved (fl. conc.) | 7440-36-0 | Metals - Dissolved; | Metals - Dissolved; EMS code: SB-D | Chemical | Dissolved | LAB
SB-D | mg/L | Antimony Dissolved | X071 | FA,HNO3,Hydr.Gen,AA | FA;HNO3;Hydr.Gen;AA | 0.005 | Y | Y | Y | 538 |  |  | 10/24/1996 12:45:23 PM | 1996-04-12 | 6/3/1998 6:51:38 PM | EMS |  | Concentration | mg/L | Y | (fl. conc.) |  | Antimony Dissolved (fl. conc.) | 7440-36-0 | Metals - Dissolved; | Metals - Dissolved; EMS code: SB-D | Chemical | Dissolved | LAB
SB-D | mg/L | Antimony Dissolved | X356 | LF,LA:HNO3:ICP | Samples are lab-filtered (0.45 um); lab-acidified with HNO3; analysis by ICP/AES. | 0.015 | Y | Y | Y | 397 |  |  | 3/23/2010 5:08:09 PM | 1996-04-12 | 6/3/1998 6:51:38 PM | EMS |  | Concentration | mg/L | Y | (fl. conc.) |  | Antimony Dissolved (fl. conc.) | 7440-36-0 | Metals - Dissolved; | Metals - Dissolved; EMS code: SB-D | Chemical | Dissolved | LAB
SB-D | mg/L | Antimony Dissolved | X085 | LF,LA,HNO3,AA | Samples are lab-filtered (0.45 um), and lab-acidified (HNO3); analysis by AA. | 0.2 | Y | Y | Y | 14 | 9/29/2000 | 12/1/2010 | 12/9/2010 3:55:27 PM | 1996-04-12 | 6/3/1998 6:51:41 PM | EMS |  | Concentration | mg/L | Y | (fl. conc.) |  | Antimony Dissolved (fl. conc.) | 7440-36-0 | Metals - Dissolved; | Metals - Dissolved; EMS code: SB-D | Chemical | Dissolved | LAB
SB-E | ug/g | Antimony-Extractable | X230 | Hom,HNO3/HCl,ICP | The sample is dried, ground to pass 100 mesh, and digested with a nitric-hydrochloric acid mixture to solubilize the solid matter and to remove the organic material by oxidation and volatilization. The completed digests are submitted to metals instrumentation personnel for metals analysis by Inductively Coupled Atomic Emission Spectrophotometry (ICP-AES), or hydride analysis by atomic absorption. | 15 | Y | Y | Y | 74 | 12/13/2011 12:02:00 PM | 12/13/2011 12:02:00 PM | 2/9/2012 12:15:57 PM | 1996-04-12 | 6/3/1998 6:51:38 PM | EMS |  | MassPerMass | mg/kg | Y | (mass conc.) |  | Antimony-Extractable (mass conc.) | 7440-36-0 | Metals - Extractable; | Metals - Extractable; EMS code: SB-E | Chemical | Extractable | LAB
SB-E | mg/L | Antimony-Extractable | X553 | Metals in water, extractable, ICP-SFMS | The whole water sample is collected in a HDPE/Teflon bottle. In a Class 100 workstation it is preserved with concentrated Nitric Acid at concentration of 0.2 percent (v/v). The acid preservative extracts the metals from material in the sample and a decanted portion is introduced directly to an inductively coupled argon plasma-sector field mass spectrometer (ICP-SFMS) by pneumatic nebulization. Each respective element is measured at a specific mass to charge ratio, m/z value expressed in amu (atomic mass units), and at a specified instrument resolution. LR: Low Resolution ~300, MR: Mid Resolution ~4000 and HR: High Resolution ~10000. Ref: Method number 02-2061, National Laboratory for Environmental Testing, CCIW, Environment Canada, Burlington. Requested April 2003. | 1e-06 | Y | Y | Y | 2587 | 5/22/2003 | 9/21/2007 | 11/25/2021 1:50:15 PM | 2004-01-19 |  | UMULLER |  | Concentration | mg/L | Y | (fl. conc.) |  | Antimony-Extractable (fl. conc.) | 7440-36-0 | Metals - Extractable; | Metals - Extractable; EMS code: SB-E | Chemical | Extractable | LAB
SB-E | mg/L | Antimony-Extractable | F069 | SEM extr,6N HCl,ICP/OES | 6.0 N Hydrochloric acid (HCI) is added to an aliquot of the sediment sample. The extract produced from the addition of the acid is then analyzed for simultaneously extracted metals (SEM) using inductively coupled plasma - optical emission spectrophotometry. The detection limits depend on the metal parameters in question and the instrumentation used. | 0.2 | Y | Y | Y | 10 | 3/26/2012 3:01:00 AM | 3/26/2012 3:01:00 AM | 11/25/2021 1:50:15 PM | 1999-07-09 |  | NPEPPIN |  | Concentration | mg/L | Y | (fl. conc.) |  | Antimony-Extractable (fl. conc.) | 7440-36-0 | Metals - Extractable; | Metals - Extractable; EMS code: SB-E | Chemical | Extractable | LAB
SB-E | mg/L | Antimony-Extractable | F040 | Field Acid, lab filt, ICP-AES | Dissolved and extractable water samples are ready for analysis once collected. Total samples are autoclaved with HNO3 and hydrochloric acid (HCI) in capped Teflon vials. Analysis is performed via Inductively Coupled Argon Plasma Atomic Emission Spectrometry (ICP-AES)(simultaneous multielement analysis), Graphite Furnace Atomic Absorption Spectrometry (GF-AAS)(low level analysis) and/or Flame Atomic Absorption Spectrometry (F-AAS). | 0.05 | Y | Y | Y | 2350 | 8/5/1997 6:06:00 PM | 3/14/2008 | 9/11/2019 8:47:41 AM | 1997-05-08 | 6/3/1998 6:51:46 PM | MPAK |  | Concentration | mg/L | Y | (fl. conc.) |  | Antimony-Extractable (fl. conc.) | 7440-36-0 | Metals - Extractable; | Metals - Extractable; EMS code: SB-E | Chemical | Extractable | LAB
SB-T | ug/g wet | Antimony Total | BPSY | MET-WET-BPSY-CCMS-VA | Metals in Animal Tissue and Vegetation (Biota) - Prescriptive. EPA 200.3/6020B This method is conducted following British Columbia Lab Manual method "Metals in Animal Tissue and Vegetation (Biota) - Prescriptive". Tissue samples are digested with nitric and hydrochloric acids; in combination with addition of hydrogen peroxide. Instrumental analysis is by collision cell inductively coupled plasma - mass spectrometry (modified from EPA Method 6020B). | 0.004 | Y | Y | Y | 2 | 11/7/2019 | 11/7/2019 | 6/10/2021 1:57:10 PM | 2021-06-04 |  | EMS |  | MassPerMass | µg/g | Y | (mass conc.) | (wet) | Antimony Total (wet) (mass conc.) | 7440-36-0 | Metals - Total; | Metals - Total; EMS code: SB-T | Chemical | Total | LAB
SB-T | ug/g | Antimony Total | X551 | Metals in Biota (dry), total, acid digestion, ICP-SFMS | 0.1 gram of dried, homogenized biota is weighed in a microwave TFM vessel, and 8mL of Nitric Acid and 1mL of Hydrogen Peroxide are added. The sample is digested at 200 degrees C for 15 minutes in a closed vessel high pressure microwave oven. The sample is cooled, the volume adjusted to 100mL and introduced to an iductively coupled argon plasma-sector field mass spectrometer (ICP-SFMS) by pneumatic nebulization. Each respective element is measured at a specific mass to charge ratio, m/z value expressed in amu (atomic mass units), and at a specified instrument resolution. LR: Low Resolution ~300, MR: Mid Resolution ~4000 and HR: Hich Resolution ~10000. Ref: Method number 02-2705, National Laboratory for Environmental Testing, CCIW, Environment Canada, Burlington. Requested April 2003. | 0.01 | Y | Y | Y | 49 | 8/11/2015 | 10/30/2023 | 11/28/2023 2:21:33 PM | 2004-01-19 |  | UMULLER |  | MassPerMass | mg/kg | Y | (mass conc.) |  | Antimony Total (mass conc.) | 7440-36-0 | Metals - Total; | Metals - Total; EMS code: SB-T | Chemical | Total | LAB
SB-T | mg/L | Antimony Total | 0040 | HNO3 Dig: ICP Analysis | HNO3 Dig: ICP Analysis | 0.015 | Y | Y | Y | 1739 | 6/25/2010 3:15:00 PM | 12/1/2010 | 12/9/2010 3:55:27 PM | 1996-04-12 | 6/3/1998 6:51:38 PM | EMS |  | Concentration | mg/L | Y | (fl. conc.) |  | Antimony Total (fl. conc.) | 7440-36-0 | Metals - Total; | Metals - Total; EMS code: SB-T | Chemical | Total | LAB
SB-T | mg/L | Antimony Total | METN | MET-T-NP-CCMS-VA | Ultra trace metals in water are analyzed by CRC ICPMS, based on US EPA Method 6020B (July 2014). The detection limits provided can only be met for undigested samples. This procedure is intended for colorless, non-turbid, acid-preserved water samples (i.e. pristine water samples), having turbidity < 1 NTU and no odor. Where turbidity exceeds 1 NTU, and/or the sample is colored and has an odor, results may be biased low compared to true Total Metals concentrations. ALS recommends that turbidity analysis be requested on samples submitted for this test to aid with interpretation of results. Where turbidity is <1NTU, undigested metals are equivalent to total metals concentrations. | 2e-05 | Y | Y | Y | 3648 | 4/3/2012 | 8/19/2024 | 8/26/2024 1:03:06 PM | 2019-08-20 |  | KATFRASE |  | Concentration | mg/L | Y | (fl. conc.) |  | Antimony Total (fl. conc.) | 7440-36-0 | Metals - Total; | Metals - Total; EMS code: SB-T | Chemical | Total | LAB
SB-T | ug/m3 | Antimony Total | X138 | 535nm,Col'metric -Air | 535nm;Col'metric -Air | 0 | Y | Y | Y | 459 |  |  | 3/9/1998 8:54:23 AM | 1996-04-12 | 6/3/1998 6:51:41 PM | EMS |  | AirConcentration | µg/m³ | Y | (air conc.) |  | Antimony Total (air conc.) | 7440-36-0 | Metals - Total; | Metals - Total; EMS code: SB-T | Chemical | Total | LAB
SB-T | mg/L | Antimony Total | X555 | Metals in water, total, "in-bottle" digestion, ICP-SFMS | The whole water sample collected in a HDPE/Teflon bottle, is de-capped in a Class 100 workstation, acidified to 2% (v/v) HNO3 and 1% (v/v) HCl and recapped. The sample is "in-bottle" digested in a convection oven, set at 60 degrees C for 16 hours. The sample is cooled and introduced to an inductively coupled argon plasma-sector field mass spectrometer (ICP-SFMS) by pneumatic nebulization. Each respective element is measured at a specific mass to charge ratio, m/z value expressed in amu (atomic mass units), and at resolution. LR: Low Resolution ~300, MR: Mid Resolution ~4000 and HR: High a specified instrument Resolution ~10000. Ref: Method number 02-2003, National Laboratory for Environmental Testing, CCIW, Environment Canada, Burlington. Requested April 2003. | 1e-06 | Y | Y | Y | 17336 | 5/20/2003 | 12/6/2023 2:08:00 AM | 2/14/2024 9:05:32 AM | 2004-01-19 |  | UMULLER |  | Concentration | mg/L | Y | (fl. conc.) |  | Antimony Total (fl. conc.) | 7440-36-0 | Metals - Total; | Metals - Total; EMS code: SB-T | Chemical | Total | LAB
SB-T | mg/L | Antimony Total | X170 | ICP | This code is a catch-all for a large variety of ICP methods including water and also solids digested into solutions as an intermediate result. Different procedures will have greatly different detection limits. (a) For air particulates, preliminary results are reported in mg/L and loading results as mg/dm2/d. (b) For Hi Vol samples preliminiary results are same as reported in mg/L and loading results as ug/m3. (c) Note that for Rain samples there is no filtration nor digestion steps. | 0.015 | Y | Y | Y | 10423 | 3/28/1996 6:03:00 PM | 11/7/2018 | 8/26/2024 2:41:32 PM | 1996-04-12 | 6/3/1998 6:51:38 PM | EMS |  | Concentration | mg/L | Y | (fl. conc.) |  | Antimony Total (fl. conc.) | 7440-36-0 | Metals - Total; | Metals - Total; EMS code: SB-T | Chemical | Total | LAB
SB-T | ug/g | Antimony Total | FS82 | PBM,dry60C,10mesh,SALM-ICP-MS | PBM;dry60C;10mesh;SALM-ICP-MS | 0.1 | Y | Y | Y | 3648 | 5/21/2002 | 7/24/2024 | 7/29/2024 5:17:09 PM | 2002-11-12 |  | UMULLER |  | MassPerMass | mg/kg | Y | (mass conc.) |  | Antimony Total (mass conc.) | 7440-36-0 | Metals - Total; | Metals - Total; EMS code: SB-T | Chemical | Total | LAB
SB-T | ppm (S) | Antimony Total | MS61 | ALS Minerals ME-MS61 | ALS Minerals ME-MS61. A prepared sample is digested with HNO3; HF; and HClO4 in a graphite heating block to near dryness. After cooling; the sample is boiled in diluted HCl. Bulk to volume with deionized water; mixed and analyzed by ICP-AES and ICP-MS for all elements. | 0.05 | Y | Y | Y | 0 |  |  |  | 2024-01-29 |  | EMS |  | DimensionlessRatio | ppm | N | (ratio) |  | Antimony Total (ratio) | 7440-36-0 | Metals - Total; | Metals - Total; EMS code: SB-T | Chemical | Total | LAB
SB-T | mg/L | Antimony Total | X113 | HNO3/HCl/Perchlor,AA | HNO3/HCl/Perchlor;AA | 0.005 | Y | Y | Y | 573 |  |  | 12/9/1996 10:18:15 PM | 1996-04-12 | 6/3/1998 6:51:38 PM | EMS |  | Concentration | mg/L | Y | (fl. conc.) |  | Antimony Total (fl. conc.) | 7440-36-0 | Metals - Total; | Metals - Total; EMS code: SB-T | Chemical | Total | LAB
SB-T | mg/L | Antimony Total | F121 | Acid,ICPMS | Acidified, no digestion, ICPMS | 2e-05 | Y | Y | Y | 8436 | 3/12/2010 7:57:00 PM | 8/15/2024 9:10:00 PM | 8/26/2024 8:06:03 AM | 2002-07-22 | 7/16/2012 3:28:13 PM | UMULLER | RDGIBSON | Concentration | mg/L | Y | (fl. conc.) |  | Antimony Total (fl. conc.) | 7440-36-0 | Metals - Total; | Metals - Total; EMS code: SB-T | Chemical | Total | LAB
SB-T | mg/L | Antimony Total | F082 | Aliquot,HNO3/HCL dig,ICPMS | Aliquot microwave digested with 3:1 HNO3/HCL. Analysis by ICPMS. For HiVol samples preliminary result are reported in mg/L and loading results are reported in ug/m3 | 0.0005 | Y | Y | Y | 27253 | 1/24/2011 | 8/23/2024 | 8/26/2024 4:53:05 PM | 2000-06-13 | 8/15/2019 11:10:58 AM | NPEPPIN | KATFRASE | Concentration | mg/L | Y | (fl. conc.) |  | Antimony Total (fl. conc.) | 7440-36-0 | Metals - Total; | Metals - Total; EMS code: SB-T | Chemical | Total | LAB
SB-T | mg/L | Antimony Total | MWID | Metals in Water by ICPMS (digested) | Metals in Water by ICPMS (digested) Based on EPA 200.8 | 0.0005 | Y | Y | Y | 36811 |  |  | 7/31/2024 1:46:41 PM | 2009-12-24 |  | RDGIBSON |  | Concentration | mg/L | Y | (fl. conc.) |  | Antimony Total (fl. conc.) | 7440-36-0 | Metals - Total; | Metals - Total; EMS code: SB-T | Chemical | Total | LAB
SB-T | mg/L | Antimony Total | F048 | Acid dig,conc'n,nebul.,ICP-OES | This method provides an automated procedure for the Inductively Coupled Argon Plasma Optical Emission Spectrophotometric (ICP-OES) determination and quantitation of trace amounts of metals in surface, ground and precipitation waters after a manual acid digestion and concentration procedure and ultrasonic nebulisation. Organically-bound metals and metal complexes are converted to metal ions by digestion with nitric acid [method details says HCl]. To increase the signal detected by the ICP, the samples are concentrated 10 times. The liquid sample is pumped onto the piezoelectric transducer, where the ultrasonic action efficiently converts it to a fine, dense aerosol. The nebulizer carrier gas flow then transports the wet aerosol through a heated U-tube, where the solvent is vaporized. Subsequently, most of the solvent vapour is removed from the aerosol stream by passing the stream through a chilled condenser. The result is a dry, analyte-laden aerosol which is introduced to the plasma. There, the high temperature desolvates the aerosol and dissociates the sample material into individual atoms and ions, which are excited to emit light at wave lengths characteristic of the elements present. Dissolved samples are field filtered through 0.45 micron cellulose acetate filter prior to acid preservation. Samples for extractable analysis are acidified with 2 mL concentrated HNO3 per litre and allowed to stand a minimum of 24 hours before analysis. Clients providing these samples are advised to preserve samples at the collection site. | 0.2 | Y | Y | Y | 29 |  |  | 7/6/2001 11:10:13 AM | 1998-11-13 |  | NPEPPIN |  | Concentration | mg/L | Y | (fl. conc.) |  | Antimony Total (fl. conc.) | 7440-36-0 | Metals - Total; | Metals - Total; EMS code: SB-T | Chemical | Total | LAB
SB-T | mg/L | Antimony Total | F068 | HNO3dig,TeflonTube/Autoclave,ICP-MS | This method is applicable to determiniation of dissolved, extractable, and total recoverable metals in ground waters, surface waters, drinking waters, and wastewaters. Dissolved, extractable, and total recoverable sample fractions are acidified to pH <1.5 with analytical grade nitric acid (HNO3) to prevent the metals from adsorbing to the container material or precipitating. Extractable and total recoverable water samples are collected in 250 ml acid washed polyethylene bottles and are preserved with Nitric acid (0.5 ml acid for 100 mLs of sample). All preserved samples are stored at room temperature and can be held for up to 6 months. Unfiltered, un-preserved samples are stored at 4C for no more than 30 days. For dissolved and extractable water fractions, no sample preparation is necessary. Samples are simply poured in autosampler tubes (i.e. 15 ml centrifuge tubes), and placed on the autosampler and analyzed. Samples for total recoverable metals are digested with Nitric acid in special Teflon tubes in an autoclave prior to analyses. | 5e-06 | Y | Y | Y | 1663 | 6/1/1999 1:33:00 PM | 2/12/2002 2:06:00 PM | 7/30/2024 6:32:12 AM | 1999-06-28 | 6/30/1999 1:05:24 PM | NPEPPIN | NPEPPIN | Concentration | mg/L | Y | (fl. conc.) |  | Antimony Total (fl. conc.) | 7440-36-0 | Metals - Total; | Metals - Total; EMS code: SB-T | Chemical | Total | LAB
SB-T | mg/L | Antimony Total | TCLP | Toxicity Characteristic Leaching Procedure | The TCLP procedure is designed to simulate the climatic leaching action | 0.1 | Y | Y | Y | 22 | 12/8/2009 7:49:00 PM | 9/28/2011 5:53:00 PM | 11/3/2011 11:10:21 PM | 2009-03-05 |  | JCHUDAK |  | Concentration | mg/L | Y | (fl. conc.) |  | Antimony Total (fl. conc.) | 7440-36-0 | Metals - Total; | Metals - Total; EMS code: SB-T | Chemical | Total | LAB
SB-T | mg/L | Antimony Total | MSEA | Metals in Seawater (undigested) by (CRC) ICP-MS | This method describes the multi-elemental determination of trace elements by Collision/Reaction Cell (CRC) ICP-MS. The method measures ions produced by a radio-frequency inductively coupled plasma. Analyte species originating in a liquid are pneumatically nebulized and the resulting aerosol transported by argon gas into the plasma torch. The ions produced are entrained in the plasma gas and extracted, by means of a differentially pumped vacuum interface, into a mass spectrometer. The ions flow through a Collision/Reaction Cell (CRC) where depending on the mode of analysis either the ions passes through the cell unhindered (no gas - classical ICP-MS mode), or collisions/reactions chemistry takes place between the ions and a neutral gas such as hydrogen or helium. Any polyatomic ions which give interference on the analytes are removed. The ions are then sorted according to their mass-to-charge ratios by a quadrupole mass spectrometer. The ions are then sorted according to their mass-to-charge ratios by a quadrupole mass spectrometer. The ions transmitted through the quadrupole are quantified by a channel electron multiplier and the ion information processed by a data handling system. Although most polyatomic interference is removed, any additional interference must be assessed and valid corrections applied or the data flagged to indicate problems. Interference correction must include compensation for background ions contributed by the plasma gas, reagents, and constituents of the sample matrix. Instrument drift as well as suppressions or enhancements of instrument response caused by the sample matrix must be corrected for by the use of internal standards. | 2e-05 | Y | Y | Y | 441 | 8/12/2009 12:08:00 AM | 4/29/2024 | 6/7/2024 10:03:16 AM | 2009-09-08 |  | RDGIBSON |  | Concentration | mg/L | Y | (fl. conc.) |  | Antimony Total (fl. conc.) | 7440-36-0 | Metals - Total; | Metals - Total; EMS code: SB-T | Chemical | Total | LAB
SB-T | ug/m3 | Antimony Total | 6040 | Aqua Regia Dig:ICP | Aqua regia digestion is used to bring the metals into solution prior to ICP analysis. For Hi-Vol samples, preliminary results are reported as mg/L and24 hr loadings as ug/m3 | 0.09 | Y | Y | Y | 437 | 2/13/2002 | 2/9/2005 | 2/23/2005 1:13:47 AM | 1998-05-05 | 6/3/1998 6:51:48 PM | NPEPPIN |  | AirConcentration | µg/m³ | Y | (air conc.) |  | Antimony Total (air conc.) | 7440-36-0 | Metals - Total; | Metals - Total; EMS code: SB-T | Chemical | Total | LAB
SB-T | ug/g wet | Antimony Total | X560 | Metals in Biota (wet), total, acid digestion, ICP-MS | Five grams of wet, homogenized biota is weighed in a centrifuge tube, and 2.5 mL of Nitric Acid and 2.5 mL of Hydrochloric acid are added. The sample is digested at room temperature for one hour and then digested at 90-95 degrees C in a water bath or heat block for one more hour. The sample is cooled and the volume is made to 50 mL with DI. The digestate is then diluted at ten fold and introduced to a collision/reaction cell inductively coupled argon plasma mass spectrometer (CRC ICPMS) by pneumatic nebulization. Each respective element is measured at a specific mass to charge ratio, m/z value expressed in amu (atomic mass units). - Maxxam Analytics (BC). Added to EMS Dec 10, 2008. | 0.1 | Y | Y | Y | 541 | 11/15/2008 4:00:00 AM | 10/30/2023 | 11/28/2023 2:21:33 PM | 2008-12-11 | 3/11/2015 4:17:10 PM | RDGIBSON | EMS | MassPerMass | µg/g | Y | (mass conc.) | (wet) | Antimony Total (wet) (mass conc.) | 7440-36-0 | Metals - Total; | Metals - Total; EMS code: SB-T | Chemical | Total | LAB
SB-T | ug/g | Antimony Total | FB04 | Freeze Dry,Hom,Acid Dig,ICP | Sample moisture is determined by weighing samples before and after freeze-drying. Freeze-dried samples are ground in a Braun blender before being weighed into Teflon PFA liners and digested with concentrated HNO3, concentrated HCl, and 50% H2O2 in a microwave oven; which results in a complete sample dissolution. Samples are collected and volumized with Type I water. Analysis is performed via Inductively Coupled Argon Plasma Atomic Emission Spectrometry (ICP-AES) (simultaneous multielement analysis), Graphite Furnace Atomic Absorption Spectrometry (GF-AAS) (low level Cd and Pb analysis), Cold Vapor Atomic Fluorescence Spectrometry (CV-AFS) (Mercury analysis) and/or Flame Atomic Absorption Spectrometry (F-AAS) (K analysis). | 4 | Y | Y | Y | 241 | 8/22/1996 6:11:00 PM | 2/1/2002 2:30:00 PM | 6/23/2003 1:15:33 PM | 1996-06-19 | 6/3/1998 6:51:44 PM | MPAK |  | MassPerMass | mg/kg | Y | (mass conc.) |  | Antimony Total (mass conc.) | 7440-36-0 | Metals - Total; | Metals - Total; EMS code: SB-T | Chemical | Total | LAB
SB-T | mg/L | Antimony Total | 1349 | Tot Metals in Seawater by CRC ICPMS (BC) | Tot Metals in Seawater by CRC ICPMS (BC) | 0.001 | Y | Y | Y | 98 | 8/16/2019 | 12/5/2021 | 12/13/2021 12:34:16 PM | 2019-09-30 |  | KATFRASE |  | Concentration | mg/L | Y | (fl. conc.) |  | Antimony Total (fl. conc.) | 7440-36-0 | Metals - Total; | Metals - Total; EMS code: SB-T | Chemical | Total | LAB
SB-T | ug/g wet | Antimony Total | X552 | Metals in Biota (wet), total, acid digestion, ICP-SFMS | One gram of wet, homogenized biota is weighed in a microwave TFM vessel, and 8mL of Nitric Acid and 1mL of Hydrogen Peroxide are added. The sample is digested at 200 degrees C for 15 minutes in a closed vessel high pressure microwave oven. The sample is cooled, the volume adjusted to 100mL and introduced to an inductively coupled argon plasma-sector field mass spectrometer (ICP-SFMS) by pneumatic nebulization. Each respective element is measured at a specific mass to charge ratio, m/z value expressed in amu (atomic mass units), and at a specified instrument resolution. LR: Low Resolution ~300, MR: Mid Resolution ~4000. Ref: 02-2704, National Laboratory for Environmental Testing, CCIW, Environment Canada, Burlington. Requested April 2003. | 0.001 | Y | Y | Y | 37 | 8/9/2016 | 10/16/2023 | 11/28/2023 2:21:33 PM | 2004-01-19 | 4/19/2021 1:22:00 PM | UMULLER | EMS | MassPerMass | µg/g | Y | (mass conc.) | (wet) | Antimony Total (wet) (mass conc.) | 7440-36-0 | Metals - Total; | Metals - Total; EMS code: SB-T | Chemical | Total | LAB
SB-T | mg/L | Antimony Total | F033 | Acid Dig,Centrifuge,ICPMS | For the determination of total recoverable elements in water or waste water, take 100 ml aliquot from a well mixed, acid preserved, sample containing not more than 0.25% (w/v) total solids and transfer to a 250 ml Griffin beaker. (If total solids are greater than 0.25% reduce the size of the aliquot by a proportionate amount). Add 1 ml of conc. nitric acid and 0.5 ml conc. hydrochloric acid. Heat on a hot plate until the volume has been reduced to 15-20ml, ensuring that the sample does not boil. Allow to cool and transfer to a 50 ml volumetric flask. Dilute to volume with ASTM type I water. Centrifuge the sample or allow to stand overnight to separate insoluble material. Prior analysis, pipette 20 ml into a 50 ml volumetric flask and dilute to volume with ASTM type I water. Analysis should be performed within seventy two hours of the completed preparation. Note that different laboratories may have very different field preservation, digestion procedures and detection limits for this method, so these should be documented. | 0.005 | Y | Y | Y | 40307 | 11/20/1997 9:30:00 AM | 8/22/2024 3:24:00 AM | 8/25/2024 9:47:36 PM | 1998-05-14 | 6/3/1998 6:51:49 PM | NPEPPIN |  | Concentration | mg/L | Y | (fl. conc.) |  | Antimony Total (fl. conc.) | 7440-36-0 | Metals - Total; | Metals - Total; EMS code: SB-T | Chemical | Total | LAB
SB-T | ug/m3 | Antimony Total | F088 | Aliquot,HNO3/HCl dig, ICPMS,loading | Aliquot microwave digested with 3:1 HNO3/HCL. Analysis by ICPMS. For HiVol samples preliminary result are reported in mg/L and loading results are reported in ug/m3 |  | Y | Y | Y | 559 | 8/17/2000 10:46:00 AM | 12/14/2006 | 12/30/2006 1:01:37 AM | 2000-08-16 |  | NPEPPIN |  | AirConcentration | µg/m³ | Y | (air conc.) |  | Antimony Total (air conc.) | 7440-36-0 | Metals - Total; | Metals - Total; EMS code: SB-T | Chemical | Total | LAB
SB-T | ug/g | Antimony Total | MTID | MET-DRY-CCMS-VA | Metals in Tissue by CRC ICPMS (DRY) Methodology: This method is conducted following British Columbia Lab Manual method "Metals in Animal Tissue and Vegetation (Biota) - Prescriptive". Tissue samples are homogenized and sub-sampled prior to hotblock digestion with nitric and hydrochloric acids, in combination with addition of hydrogen peroxide. Instrumental analysis is by collision cell inductively coupled plasma - mass spectrometry (modified from EPA Method 6020A). Method Limitation: This method employs a strong acid/peroxide digestion, and is intended to provide a conservative estimate of bio-available metals. Near complete recoveries are achieved for most toxicologically important metals, but elements associated with recalcitrant minerals may be only partially recovered. | 0.01 | Y | Y | Y | 122 | 8/13/2014 | 9/8/2022 | 10/26/2022 7:14:15 AM | 2015-03-09 | 3/11/2015 4:12:05 PM | RDGIBSON | EMS | MassPerMass | mg/kg | Y | (mass conc.) |  | Antimony Total (mass conc.) | 7440-36-0 | Metals - Total; | Metals - Total; EMS code: SB-T | Chemical | Total | LAB
SB-T | mg/L | Antimony Total | EQ01 | EQUIS (mg/L) | EQUIS (mg/L) | 0.005 | N | Y | Y | 356 |  |  | 10/22/1996 10:24:55 PM | 1996-04-12 | 6/3/1998 6:51:38 PM | EMS |  | Concentration | mg/L | Y | (fl. conc.) |  | Antimony Total (fl. conc.) | 7440-36-0 | Metals - Total; | Metals - Total; EMS code: SB-T | Chemical | Total | LAB
SB-T | ug/m3 | Antimony Total | 5097 | N/P Dig(Filt+Wash),ICP | N/P Dig(Filt+Wash);ICP | 0.001 | Y | Y | Y | 829 |  |  | 7/3/2024 12:59:21 PM | 1998-04-27 | 6/3/1998 6:51:48 PM | NPEPPIN |  | AirConcentration | µg/m³ | Y | (air conc.) |  | Antimony Total (air conc.) | 7440-36-0 | Metals - Total; | Metals - Total; EMS code: SB-T | Chemical | Total | LAB
SB-T | ug/g | Antimony Total | PB02 | PBM,HNO3/HCl dig,dry60C,100mesh, | Performance Based Method. Samples are dried at 60 C and sieved using a .149 mm (100 mesh) sieve. Solid samples are digested with the combination of nitric and hydrochloric acids with a 1:1 ratio of acids. For environmental purposes, strong acid leaches are used routinely since they have the advantage of speed, low cost, and superio reproducibility. All metals trapped within the refractory matrix are unlikely to have environmental consequences. Instrumental analysis of extracts produced using this digestion method can be performed by a variety of techniques. |  | Y | Y | Y | 52 | 5/24/2001 3:42:00 PM | 11/26/2001 10:23:00 AM | 5/26/2003 2:28:21 PM | 2001-07-18 |  | NPEPPIN |  | MassPerMass | mg/kg | Y | (mass conc.) |  | Antimony Total (mass conc.) | 7440-36-0 | Metals - Total; | Metals - Total; EMS code: SB-T | Chemical | Total | LAB
SB-T | ug/g | Antimony Total | FS01 | D&G,Nitric/HCl:ICP | Samples are dried at 60 C, then either sieved to -100mesh (150 um) and rolled to homogenize, or finely ground with a mortar and pestle or a shatter box laboratory mill and rolled (depending on the submitter's requirements). The samples are then weighed into Teflon digestion liners and digested with HNO3 and HCl in a microwave oven. Samples are cooled, volumized, and allowed to settle overnight before being decanted. Analysis is then performed via Inductively Coupled Argon Plasma Atomic Emission Spectrometry (ICP-AES) (Simultaneous multi element analysis). In ICP Spectrometry, the liquid sample is pumped to the nebulizer, where it is convertd into an aerosol. the nebulizer carrier gas flow than transports the wet aerosol past the impact bead, where the large droplets condense out and flow to the waste container. The remaining fine aerosol is carried into the plasma. There, the high temperature desolvates the aerosol and dissociates the sample material into individual atoms and ions, which are excited to emit light at wavelengths characteristic of the elements present. A diffraction grating in the spectrometer diffracts the various wavelengths of light through exit slits to the photo-multiplier tubes, which convert the light energy to electronic signals. | 8 | Y | Y | Y | 2642 | 8/30/1996 1:52:00 PM | 10/10/2007 | 5/24/2008 1:15:03 AM | 1996-11-20 | 6/3/1998 6:51:46 PM | MPAK |  | MassPerMass | mg/kg | Y | (mass conc.) |  | Antimony Total (mass conc.) | 7440-36-0 | Metals - Total; | Metals - Total; EMS code: SB-T | Chemical | Total | LAB
SB-T | ug/g | Antimony Total | 0190 | D&G:N/P Dig:ICP Anal. | Sample is dried and grind-mixed before nitric perchloric digestion followed by ICP analysis. | 1 | Y | Y | Y | 16 |  |  | 10/24/1996 11:32:22 AM | 1996-04-12 | 6/3/1998 6:51:38 PM | EMS |  | MassPerMass | mg/kg | Y | (mass conc.) |  | Antimony Total (mass conc.) | 7440-36-0 | Metals - Total; | Metals - Total; EMS code: SB-T | Chemical | Total | LAB
SB-T | mg/L | Antimony Total | 0185 | Hydr.At.Absorp.Spec(H2O) | Hydr.At.Absorp.Spec(H2O) | 0.5 | Y | Y | Y | 83 | 12/6/1996 | 12/6/1996 | 3/8/2001 9:35:19 AM | 1996-04-12 | 6/3/1998 6:51:38 PM | EMS |  | Concentration | mg/L | Y | (fl. conc.) |  | Antimony Total (fl. conc.) | 7440-36-0 | Metals - Total; | Metals - Total; EMS code: SB-T | Chemical | Total | LAB
SB-T | mg/L | Antimony Total | F038 | HNO3/HCl dig,ICP-AES | Dissolved and extractable water samples are ready for analysis once collected. Total samples are autoclaved with HNO3 and hydrochloric acid (HCI) in capped Teflon vials. Analysis is performed via Inductively Coupled Argon Plasma Atomic Emission Spectrometry (ICP-AES)(simultaneous multielement analysis), Graphite Furnace Atomic Absorption Spectrometry (GF-AAS) (low level analysis) and/or Flame Atomic Absorption Spectrometry (F-AAS). | 0.06 | Y | Y | Y | 5393 | 8/1/1996 6:04:00 PM | 5/6/2008 | 9/11/2019 8:47:43 AM | 1997-05-08 | 6/3/1998 6:51:46 PM | MPAK |  | Concentration | mg/L | Y | (fl. conc.) |  | Antimony Total (fl. conc.) | 7440-36-0 | Metals - Total; | Metals - Total; EMS code: SB-T | Chemical | Total | LAB
SB-T | ug/m3 | Antimony Total | 5312 | Tef:N/P Dig:ICP Analysis | A nitric/perchloric acid digestion if used to bring the metals into solution prior to ICP analysis. For Hi-Vol samples preliminary results are reported as mg/L and 24 hour loading results as ug/m3. | 0.09 | Y | Y | Y | 1186 | 11/7/1997 | 3/29/2000 | 9/28/2000 1:34:44 PM | 1996-10-13 | 6/3/1998 6:51:46 PM | EMS |  | AirConcentration | µg/m³ | Y | (air conc.) |  | Antimony Total (air conc.) | 7440-36-0 | Metals - Total; | Metals - Total; EMS code: SB-T | Chemical | Total | LAB
SB-T | mg/L | Antimony Total | X351 | LA:HNO3:AA | Samples are lab-acidified (HNO3); digested HNO3; analysis by AA. | 0.2 | Y | Y | Y | 82 |  |  | 8/26/2004 11:46:08 AM | 1996-04-12 | 6/3/1998 6:51:41 PM | EMS |  | Concentration | mg/L | Y | (fl. conc.) |  | Antimony Total (fl. conc.) | 7440-36-0 | Metals - Total; | Metals - Total; EMS code: SB-T | Chemical | Total | LAB
SB-T | ug/m3 | Antimony Total | EQ40 | EQUIS (mg/m3) | EQUIS (mg/m3) | 0 | N | Y | Y | 19 |  |  | 10/22/1996 11:59:06 PM | 1996-04-12 | 6/3/1998 6:51:41 PM | EMS |  | AirConcentration | µg/m³ | Y | (air conc.) |  | Antimony Total (air conc.) | 7440-36-0 | Metals - Total; | Metals - Total; EMS code: SB-T | Chemical | Total | LAB
SB-T | ug/g | Antimony Total | ICPM | Metals in Soil by CRC ICPMS | This analysis is carried out using procedures from CSR Analytical Method: "Strong Acid Leachable Metals (SALM) in Soil", BC Ministry of Environment, 26 June 2009, and procedures adapted from EPA Method 200.2. The sample is manually homogenized, dried at 60 degrees Celsius, sieved through a 63 um (230 mesh) sieve, and a representative subsample of the dry material is weighed. The sample is then digested at 95 degrees Celsius for 2 hours by block digester using concentrated nitric and hydrochloric acids. Instrumental analysis is by collision cell inductively coupled plasma - mass spectrometry (modifed from EPA Method 6020A). Method Limitation: This method is not a total digestion technique. It is a very strong acid digestion that is intended to dissolve those metals that may be environmentally available. By design, elements bound in silicate structures are not normally dissolved by this procedure as they are not usually mobile in the environment. | 0.1 | Y | Y | Y | 103 | 9/19/2017 | 9/24/2023 | 1/12/2024 11:16:52 AM | 2018-01-09 |  | JKROGH |  | MassPerMass | mg/kg | Y | (mass conc.) |  | Antimony Total (mass conc.) | 7440-36-0 | Metals - Total; | Metals - Total; EMS code: SB-T | Chemical | Total | LAB
SB-T | ug/g wet | Antimony Total | MTIW | MET-WET-CCMS-MID-VA | Metals in Tissue by CRC ICPMS (WET) Methodology: This method is conducted following British Columbia Lab Manual method "Metals in Animal Tissue and Vegetation (Biota) - Prescriptive". Tissue samples are homogenized and sub-sampled prior to hotblock digestion with nitric and hydrochloric acids, in combination with addition of hydrogen peroxide. Instrumental analysis is by collision cell inductively coupled plasma - mass spectrometry (modified from EPA Method 6020A). Method Limitation: This method employs a strong acid/peroxide digestion, and is intended to provide a conservative estimate of bio-available metals. Near complete recoveries are achieved for most toxicologically important metals, but elements associated with recalcitrant minerals may be only partially recovered. | 0.002 | Y | Y | Y | 24 | 11/25/2018 | 12/6/2018 | 7/21/2021 2:55:08 PM | 2018-12-24 |  | EMS |  | MassPerMass | µg/g | Y | (mass conc.) | (wet) | Antimony Total (wet) (mass conc.) | 7440-36-0 | Metals - Total; | Metals - Total; EMS code: SB-T | Chemical | Total | LAB
SB-T | mg/L | Antimony Total | X072 | FA,HNO3,Graph.Furnace,AA | Samples are field-acidified with HNO3; HNO3 digestion; analysis by GFAA. | 0.003 | Y | Y | Y | 49 |  |  | 10/24/1996 1:51:39 PM | 1996-04-12 | 6/3/1998 6:51:38 PM | EMS |  | Concentration | mg/L | Y | (fl. conc.) |  | Antimony Total (fl. conc.) | 7440-36-0 | Metals - Total; | Metals - Total; EMS code: SB-T | Chemical | Total | LAB
SB-T | ppm (S) | Antimony Total | MS41 | ALS Minerals ME-MS41 | ALS Minerals ME-MS41. A prepared sample is digested with aqua regia in a graphite heating block. After cooling; the resulting solution is diluted to with deionized water; mixed and analyzed by ICP-AES. Samples are then analyzed by ICP-MS for the remaining suite of elements. | 0.05 | Y | Y | Y | 0 |  |  |  | 2024-01-29 |  | EMS |  | DimensionlessRatio | ppm | N | (ratio) |  | Antimony Total (ratio) | 7440-36-0 | Metals - Total; | Metals - Total; EMS code: SB-T | Chemical | Total | LAB
SB-T | mg/L | Antimony Total | X545 | FA,HNO3,ICPMS | Analytes nebulized followed by ion entrained in plasma gas and extracted by means of a differentially pumped vacuum interface into a mass spectrometer | 0.001 | Y | Y | Y | 30 |  |  | 10/30/2014 8:31:18 AM | 2001-09-18 | 9/18/2001 1:36:23 PM | NPEPPIN | NPEPPIN | Concentration | mg/L | Y | (fl. conc.) |  | Antimony Total (fl. conc.) | 7440-36-0 | Metals - Total; | Metals - Total; EMS code: SB-T | Chemical | Total | LAB
SB-T | ppm (S) | Antimony Total | 6041 | Aqua Regia by ICPMS | These were sediment/soils digested with Aqua-Regia (Nitric Acid & Hydrochloric Acids) and then analysis by ICPMS. Results would be reported in mg/kg and/or %. | 0.005 | Y | Y | Y | 21 | 10/13/2019 | 8/5/2020 | 8/10/2020 3:52:07 PM | 2019-09-06 | 10/21/2019 11:56:17 AM | KATFRASE | KATFRASE | DimensionlessRatio | ppm | N | (ratio) |  | Antimony Total (ratio) | 7440-36-0 | Metals - Total; | Metals - Total; EMS code: SB-T | Chemical | Total | LAB
SB-T | ug/g | Antimony Total | EQ02 | EQUIS (ug/g) | EQUIS (ug/g) | 1 | N | Y | Y | 19 |  |  | 10/22/1996 11:00:36 PM | 1996-04-12 | 6/3/1998 6:51:38 PM | EMS |  | MassPerMass | mg/kg | Y | (mass conc.) |  | Antimony Total (mass conc.) | 7440-36-0 | Metals - Total; | Metals - Total; EMS code: SB-T | Chemical | Total | LAB
SB-T | mg/L | Antimony Total | X349 | FA:HNO3:ICP | Samples are field-acidified with HNO3; HNO3 digestion; analysis by ICP/AES. | 0.015 | Y | Y | Y | 749 |  |  | 10/1/2003 9:41:34 AM | 1996-04-12 | 6/3/1998 6:51:38 PM | EMS |  | Concentration | mg/L | Y | (fl. conc.) |  | Antimony Total (fl. conc.) | 7440-36-0 | Metals - Total; | Metals - Total; EMS code: SB-T | Chemical | Total | LAB
SB-T | mg/L | Antimony Total | 0042 | GI HNO3 Dig: ICP | Sample aliquot taken from unfiltered unpreserved General Ions bottle; HNO3 digestion; ICP analysis. | 0.02 | Y | Y | Y | 6149 | 9/20/1996 | 7/5/2019 3:21:00 PM | 7/12/2019 6:49:13 PM | 1996-04-12 | 6/3/1998 6:51:38 PM | EMS |  | Concentration | mg/L | Y | (fl. conc.) |  | Antimony Total (fl. conc.) | 7440-36-0 | Metals - Total; | Metals - Total; EMS code: SB-T | Chemical | Total | LAB
SB-T | mg/L | Antimony Total | F024 | HNO3/HCl,GFAA | Dissolved and extractable water samples are ready for analysis once collected. Total samples are autoclaved with HNO3 and hydrochloric acid (HCl) in capped Teflon vials. Analysis is performed via Inductively Coupled Argon Plasma Atomic Emission Spectrometry (ICP-AES) (simultaneous multielement analysis), Graphite Furnace Atomic Absorption Spectrometry (GF-AAS) (low level analysis) and/or Flame Atomic Absorption Spectrometry (F-AAS). | 0.001 | Y | Y | Y | 115 | 2/13/1998 5:55:00 PM | 9/12/2003 3:16:00 PM | 12/1/2003 5:30:23 PM | 1997-04-02 | 6/3/1998 6:51:46 PM | MPAK |  | Concentration | mg/L | Y | (fl. conc.) |  | Antimony Total (fl. conc.) | 7440-36-0 | Metals - Total; | Metals - Total; EMS code: SB-T | Chemical | Total | LAB
SB-T | mg/L | Antimony Total | X073 | FA,HNO3,Direct Flame,AA | Samples are field-acidified with HNO3; HNO3 digestion; analysis by AA. | 0.5 | Y | Y | Y | 8 |  |  | 10/11/2001 10:17:28 AM | 1996-04-12 | 6/3/1998 6:51:38 PM | EMS |  | Concentration | mg/L | Y | (fl. conc.) |  | Antimony Total (fl. conc.) | 7440-36-0 | Metals - Total; | Metals - Total; EMS code: SB-T | Chemical | Total | LAB
SB-T | mg/dm2/d | Antimony Total | DUST | HNO3 dig.(microwave or hot block),ICPMS | HNO3 dig.(microwave or hot block);ICPMS |  | Y | Y | Y | 420 | 7/19/2016 1:11:00 PM | 7/31/2024 | 7/31/2024 11:23:12 AM | 2008-02-07 |  | RDGIBSON |  | MassFlux | mg/dm²/d | Y | (mass flux) |  | Antimony Total (mass flux) | 7440-36-0 | Metals - Total; | Metals - Total; EMS code: SB-T | Chemical | Total | LAB
SB-T | ug/g | Antimony Total | X230 | Hom,HNO3/HCl,ICP | The sample is dried, ground to pass 100 mesh, and digested with a nitric-hydrochloric acid mixture to solubilize the solid matter and to remove the organic material by oxidation and volatilization. The completed digests are submitted to metals instrumentation personnel for metals analysis by Inductively Coupled Atomic Emission Spectrophotometry (ICP-AES), or hydride analysis by atomic absorption. | 1 | Y | Y | Y | 91 | 2/12/2002 | 12/30/2005 | 2/3/2006 1:13:13 AM | 1996-04-12 | 6/3/1998 6:51:38 PM | EMS |  | MassPerMass | mg/kg | Y | (mass conc.) |  | Antimony Total (mass conc.) | 7440-36-0 | Metals - Total; | Metals - Total; EMS code: SB-T | Chemical | Total | LAB
SB-T | mg/L | Antimony Total | 0050 | Graphite Furnace AA | Graphite Furnace AA | 0.005 | Y | Y | Y | 414 |  |  | 10/24/1996 11:57:35 AM | 1996-04-12 | 6/3/1998 6:51:41 PM | EMS |  | Concentration | mg/L | Y | (fl. conc.) |  | Antimony Total (fl. conc.) | 7440-36-0 | Metals - Total; | Metals - Total; EMS code: SB-T | Chemical | Total | LAB
SB-T | mg/L | Antimony Total | X285 | HNO3 Dig:ICP Analysis | HNO3 Dig:ICP Analysis | 0.06 | Y | Y | Y | 595 | 8/10/1993 | 6/25/2010 3:15:00 PM | 7/20/2010 12:10:05 PM | 1998-05-22 | 6/3/1998 6:51:49 PM | NPEPPIN |  | Concentration | mg/L | Y | (fl. conc.) |  | Antimony Total (fl. conc.) | 7440-36-0 | Metals - Total; | Metals - Total; EMS code: SB-T | Chemical | Total | LAB
SB-T | ppm (S) | Antimony Total | 6042 | Aqua Regia by ICPMS | These were sediment/soils digested with Aqua-Regia (Nitric Acid & Hydrochloric Acids) and then analysis by ICPMS. Results would be reported in ppm s. | 0.005 | Y | Y | Y | 71 | 5/14/2022 | 7/27/2024 | 7/29/2024 9:13:08 AM | 2022-06-28 |  | EMS |  | DimensionlessRatio | ppm | N | (ratio) |  | Antimony Total (ratio) | 7440-36-0 | Metals - Total; | Metals - Total; EMS code: SB-T | Chemical | Total | LAB
SB-T | ug/m3 | Antimony Total | 5008 | N/P/S Dig,Hydride ICP | N/P/S Dig;Hydride ICP | 0.002 | Y | Y | Y | 20 |  |  | 10/24/1996 4:45:35 AM | 1996-04-12 | 6/3/1998 6:51:38 PM | EMS |  | AirConcentration | µg/m³ | Y | (air conc.) |  | Antimony Total (air conc.) | 7440-36-0 | Metals - Total; | Metals - Total; EMS code: SB-T | Chemical | Total | LAB
SB-T | mg/L | Antimony Total | X561 | Total metals: FA, HNO3, HR-ICPMS (No Prep)Prep) | Ultra trace metals in water are analyzed by high resolution inductively coupled plasma mass spectrometry (HR-ICPMS), modified from US EPA 200.8, (Revision 5.5). The detection limits provided can only be met for undigested samples. This procedure is intended for colorless, non-turbid, acid-preserved water samples (i.e. pristine water samples), having turbidity < 1 NTU and no odor. Where turbidity exceeds 1 NTU, and/or the sample is colored and has an odor, results may be biased low compared to true Total Metals concentrations. ALS recommends that turbidity analysis be requested on samples submitted for this test to aid with interpretation of results. | 2e-05 | Y | Y | Y | 2756 | 6/8/2015 | 10/5/2023 | 1/31/2024 2:37:41 PM | 2015-10-22 |  | RDGIBSON |  | Concentration | mg/L | Y | (fl. conc.) |  | Antimony Total (fl. conc.) | 7440-36-0 | Metals - Total; | Metals - Total; EMS code: SB-T | Chemical | Total | LAB
SB25 | Bq/L | Antimony-125 | X240 | Gamma Spectrum | Gamma Spectrum | 0.001 | Y | Y | Y | 125 |  |  | 10/24/1996 7:20:16 AM | 1996-04-12 | 6/3/1998 6:51:38 PM | EMS |  | Radioactivity | Bq/L | Y | (rad./vol.) |  | Antimony-125 (rad./vol.) | 14234-35-6 | Radionuclide; | Radionuclide; EMS code: SB25 | Chemical |  | LAB
SC-D | mg/L | Scandium Dissolved | F032 | Filt.,Acid,ICPMS | There is no acid digestion required for this analysis. A water sample is collected and preserved in 0.2 percent nitric acid. To analyse the extractable metals the sample is shaken, allowed to sit for at least twenty four hours, and a decanted portion is introduced to the spectrometer. When the dissolved metals are required, the samples must first be filtered through a 0.45 micron cellulose acetate filter paper and then followed by preservation in 0.2 percent nitric acid. It is important to note that different laboratories may have very different detection limits for some variables analyzed by ICPMS. | 0.001 | Y | Y | Y | 68 | 5/9/2003 9:40:00 AM | 9/16/2003 3:18:00 PM | 11/28/2003 3:22:15 PM | 1998-06-22 |  | NPEPPIN |  | Concentration | mg/L | Y | (fl. conc.) |  | Scandium Dissolved (fl. conc.) | 7440-20-2 | Metals - Dissolved; | Metals - Dissolved; EMS code: SC-D | Chemical | Dissolved | LAB
SC-F | uS/cm | Specific Conductivity-Field | FLD | Field sampling | Field sampling method | 2 | Y | Y | Y | 7310 | 2/16/2016 3:00:00 PM | 8/22/2024 | 8/26/2024 4:51:04 PM | 2019-07-10 | 6/30/2020 4:45:17 PM | KATFRASE | RDGIBSON | ElectricConductivity | µS/cm | Y | (elec. cond.) |  | Specific Conductivity-Field (elec. cond.) |  | Physical Properties; | Physical Properties; EMS code: SC-F | Physical |  | FIELD_RESULT | Field sampling
SC-F | uS/cm | Specific Conductivity-Field | EXOSO | EXO Sonde YSI | EXO Sonde YSI |  | Y | Y | Y | 359 |  |  | 7/15/2024 3:01:47 PM | 2023-06-12 |  | EMS |  | ElectricConductivity | µS/cm | Y | (elec. cond.) |  | Specific Conductivity-Field (elec. cond.) |  | Physical Properties; | Physical Properties; EMS code: SC-F | Physical |  | FIELD_RESULT | EXO Sonde YSI
SC-F | uS/cm | Specific Conductivity-Field | XM00 | Meter-in situ | Meter-in situ | 1 | Y | Y | Y | 58370 | 4/15/2005 10:00:00 AM | 7/26/2024 3:57:00 PM | 8/20/2024 7:35:41 AM | 2012-10-12 | 12/20/2019 11:02:20 AM | RDGIBSON | RDGIBSON | ElectricConductivity | µS/cm | Y | (elec. cond.) |  | Specific Conductivity-Field (elec. cond.) |  | Physical Properties; | Physical Properties; EMS code: SC-F | Physical |  | FIELD_RESULT | Meter-in situ
SC-T | mg/L | Scandium Total | F033 | Acid Dig,Centrifuge,ICPMS | For the determination of total recoverable elements in water or waste water, take 100 ml aliquot from a well mixed, acid preserved, sample containing not more than 0.25% (w/v) total solids and transfer to a 250 ml Griffin beaker. (If total solids are greater than 0.25% reduce the size of the aliquot by a proportionate amount). Add 1 ml of conc. nitric acid and 0.5 ml conc. hydrochloric acid. Heat on a hot plate until the volume has been reduced to 15-20ml, ensuring that the sample does not boil. Allow to cool and transfer to a 50 ml volumetric flask. Dilute to volume with ASTM type I water. Centrifuge the sample or allow to stand overnight to separate insoluble material. Prior analysis, pipette 20 ml into a 50 ml volumetric flask and dilute to volume with ASTM type I water. Analysis should be performed within seventy two hours of the completed preparation. Note that different laboratories may have very different field preservation, digestion procedures and detection limits for this method, so these should be documented. | 0.001 | Y | Y | Y | 68 | 5/8/2003 9:40:00 AM | 9/12/2003 3:16:00 PM | 12/1/2003 5:30:23 PM | 1998-06-22 |  | NPEPPIN |  | Concentration | mg/L | Y | (fl. conc.) |  | Scandium Total (fl. conc.) | 7440-20-2 | Metals - Total; | Metals - Total; EMS code: SC-T | Chemical | Total | LAB
SC-T | ppm (S) | Scandium Total | 6042 | Aqua Regia by ICPMS | These were sediment/soils digested with Aqua-Regia (Nitric Acid & Hydrochloric Acids) and then analysis by ICPMS. Results would be reported in ppm s. | 0.005 | Y | Y | Y | 71 | 5/14/2022 | 7/27/2024 | 7/29/2024 9:13:08 AM | 2022-06-28 |  | EMS |  | DimensionlessRatio | ppm | N | (ratio) |  | Scandium Total (ratio) | 7440-20-2 | Metals - Total; | Metals - Total; EMS code: SC-T | Chemical | Total | LAB
SC-T | ppm (S) | Scandium Total | MS81 | ALS Minerals: ME-MS81 | ALS Minerals: ME-MS81. A prepared sample (0.100 g) is added to lithium metaborate/lithium tetraborate flux; mixed well and fused in a furnace at 1025°C. The resulting melt is then cooled and dissolved in an acid mixture containing nitric; hydrochloric and hydrofluoric acids. This solution is then analyzed by inductively coupled plasma - mass spectrometry. | 0.5 | Y | Y | Y | 0 |  |  |  | 2024-01-29 |  | EMS |  | DimensionlessRatio | ppm | N | (ratio) |  | Scandium Total (ratio) | 7440-20-2 | Metals - Total; | Metals - Total; EMS code: SC-T | Chemical | Total | LAB
SC-T | ppm (S) | Scandium Total | MS41 | ALS Minerals ME-MS41 | ALS Minerals ME-MS41. A prepared sample is digested with aqua regia in a graphite heating block. After cooling; the resulting solution is diluted to with deionized water; mixed and analyzed by ICP-AES. Samples are then analyzed by ICP-MS for the remaining suite of elements. | 0.1 | Y | Y | Y | 0 |  |  |  | 2024-01-29 |  | EMS |  | DimensionlessRatio | ppm | N | (ratio) |  | Scandium Total (ratio) | 7440-20-2 | Metals - Total; | Metals - Total; EMS code: SC-T | Chemical | Total | LAB
SC-T | ppm (S) | Scandium Total | 6041 | Aqua Regia by ICPMS | These were sediment/soils digested with Aqua-Regia (Nitric Acid & Hydrochloric Acids) and then analysis by ICPMS. Results would be reported in mg/kg and/or %. | 0.005 | Y | Y | Y | 21 | 10/13/2019 | 8/5/2020 | 8/10/2020 3:52:07 PM | 2019-09-06 | 10/21/2019 11:56:20 AM | KATFRASE | KATFRASE | DimensionlessRatio | ppm | N | (ratio) |  | Scandium Total (ratio) | 7440-20-2 | Metals - Total; | Metals - Total; EMS code: SC-T | Chemical | Total | LAB
SC-T | ppm (S) | Scandium Total | MS61 | ALS Minerals ME-MS61 | ALS Minerals ME-MS61. A prepared sample is digested with HNO3; HF; and HClO4 in a graphite heating block to near dryness. After cooling; the sample is boiled in diluted HCl. Bulk to volume with deionized water; mixed and analyzed by ICP-AES and ICP-MS for all elements. | 0.1 | Y | Y | Y | 0 |  |  |  | 2024-01-29 |  | EMS |  | DimensionlessRatio | ppm | N | (ratio) |  | Scandium Total (ratio) | 7440-20-2 | Metals - Total; | Metals - Total; EMS code: SC-T | Chemical | Total | LAB
SCAL | % | Sulfate (as S); carbonate leach | GRA06 | ALS Minerals S-GRA06 | ALS Minerals S-GRA06. A preppared sample is heated with diluted Na₂CO₃ and then filtered. The sulfate in the filtrate is precipitated with BaCl2 in a dilute HCl. The BaSO4 precipitate is filtered; ignited; weighed and the Sulfate Sulfur is determined (as S). | 0.01 | Y | Y | Y | 0 |  |  |  | 2024-01-29 |  | EMS |  | DimensionlessRatio | % | N | (ratio) |  | Sulfate (as S); carbonate leach (ratio) |  |  | EMS code: SCAL | Chemical |  | LAB
SD35 | % | Stearic Acid-d35 | SURR | Surrogate Recovery | Surrogate Recovery | 0.001 | Y | Y | Y | 12 | 11/28/2022 | 6/15/2024 | 7/8/2024 11:10:02 AM | 2022-01-18 | 1/3/2023 11:51:02 AM | RDGIBSON | EMS | DimensionlessRatio | % | N | (ratio) |  | Stearic Acid-d35 (ratio) | 17660-51-4 |  | EMS code: SD35 | Chemical |  | SURROGATE_RESULT
SE-D | mg/L | Selenium Dissolved | X271 | Hydride:AA | Hydride:AA | 0.05 | Y | Y | Y | 2116 |  |  | 5/19/2015 3:10:36 PM | 1996-04-12 | 6/3/1998 6:51:41 PM | EMS |  | Concentration | mg/L | Y | (fl. conc.) |  | Selenium Dissolved (fl. conc.) | 7782-49-2 | Metals - Dissolved; | Metals - Dissolved; EMS code: SE-D | Chemical | Dissolved | LAB
SE-D | mg/L | Selenium Dissolved | X556 | Metals in water, dissolved, "in-bottle" digestion, ICP-SFMS | The whole water sample is filtered through a 0.45 micron cellulose acetate filter paper and the filtrate is stored in a HDPE/Teflon bottle. The container is de-capped in a Class 100 workstation, acidified to 2% (v/v) HNO3 and 1% (v/v) HCl and recapped. The sample is "in-bottle" digested in a convection oven, set at 60 degrees C for 16 hours. The sample is cooled and introduced to an inductively coupled argon plasma-sector field mass spectrometer (ICP-SFMS) by pneumatic nebulization. Each respective element is measured at a specific mass to charge ratio, m/z value expressed in amu (atomic mass units), and at a specified intrument resolution. LR: Low Resolution ~300, MR: Mid Resolution ~4000 and HR: High Resolution ~10000. Ref: Method number 02-2003, National Laboratory for Environmental Testing, CCIW, Environment Canada, Burlington, Requested April 2003. | 5e-05 | Y | Y | Y | 4234 | 6/11/2003 | 12/6/2023 2:04:00 AM | 2/14/2024 9:05:32 AM | 2004-01-19 |  | UMULLER |  | Concentration | mg/L | Y | (fl. conc.) |  | Selenium Dissolved (fl. conc.) | 7782-49-2 | Metals - Dissolved; | Metals - Dissolved; EMS code: SE-D | Chemical | Dissolved | LAB
SE-D | mg/L | Selenium Dissolved | X348 | FF,FA:HCL:HVAAS | Samples are field-filtered (0.45 um) and field-acidified (HCl); analysis by HVAAS. | 0.0005 | Y | Y | Y | 5860 |  |  | 8/28/2018 3:14:04 PM | 1996-04-12 | 6/3/1998 6:51:38 PM | EMS |  | Concentration | mg/L | Y | (fl. conc.) |  | Selenium Dissolved (fl. conc.) | 7782-49-2 | Metals - Dissolved; | Metals - Dissolved; EMS code: SE-D | Chemical | Dissolved | LAB
SE-D | mg/L | Selenium Dissolved | X359 | LF,LA:HNO3:HVAAS | Samples are lab-filtered (0.45 um) and lab-acidified (HNO3); analysis by HVAAS. | 0.0005 | Y | Y | Y | 322 | 3/31/2000 | 1/3/2007 | 11/22/2007 11:47:51 AM | 1996-04-12 | 6/3/1998 6:51:38 PM | EMS |  | Concentration | mg/L | Y | (fl. conc.) |  | Selenium Dissolved (fl. conc.) | 7782-49-2 | Metals - Dissolved; | Metals - Dissolved; EMS code: SE-D | Chemical | Dissolved | LAB
SE-D | mg/L | Selenium Dissolved | X085 | LF,LA,HNO3,AA | Samples are lab-filtered (0.45 um), and lab-acidified (HNO3); analysis by AA. | 0.2 | Y | Y | Y | 29 | 6/12/2000 | 9/21/2001 | 10/2/2001 1:53:12 PM | 1996-04-12 | 6/3/1998 6:51:38 PM | EMS |  | Concentration | mg/L | Y | (fl. conc.) |  | Selenium Dissolved (fl. conc.) | 7782-49-2 | Metals - Dissolved; | Metals - Dissolved; EMS code: SE-D | Chemical | Dissolved | LAB
SE-D | mg/L | Selenium Dissolved | X562 | Dissolved metals: FF, FA, HNO3, HR-ICPMS | Ultra trace metals in water are analyzed by high resolution inductively coupled plasma mass spectrometry (HR-ICPMS), based on US EPA Method 200.8, (Oct. 1999). This procedure is intended for pristine field-filtered acid-preserved water samples. The detection limits (LOR) for this test are based on lab instrumental analysis only, not including filtration. ALS-supplied field filtration equipment does not support these LOR. Therefore, because of the high probability of false positives due to filtration, it is strongly recommended that a filtration blank be analyzed to aid in data interpretation. | 4e-05 | Y | Y | Y | 1264 | 6/8/2015 | 10/19/2023 | 11/9/2023 7:41:27 AM | 2015-10-22 |  | RDGIBSON |  | Concentration | mg/L | Y | (fl. conc.) |  | Selenium Dissolved (fl. conc.) | 7782-49-2 | Metals - Dissolved; | Metals - Dissolved; EMS code: SE-D | Chemical | Dissolved | LAB
SE-D | mg/L | Selenium Dissolved | MWID | Metals in Water by ICPMS (digested) | Metals in Water by ICPMS (digested) Based on EPA 200.8 | 0.0008 | Y | Y | Y | 35694 | 10/26/2018 5:20:00 PM | 7/29/2024 | 7/31/2024 1:46:41 PM | 2009-12-24 |  | RDGIBSON |  | Concentration | mg/L | Y | (fl. conc.) |  | Selenium Dissolved (fl. conc.) | 7782-49-2 | Metals - Dissolved; | Metals - Dissolved; EMS code: SE-D | Chemical | Dissolved | LAB
SE-D | mg/L | Selenium Dissolved | MSEA | Metals in Seawater (undigested) by (CRC) ICP-MS | This method describes the multi-elemental determination of trace elements by Collision/Reaction Cell (CRC) ICP-MS. The method measures ions produced by a radio-frequency inductively coupled plasma. Analyte species originating in a liquid are pneumatically nebulized and the resulting aerosol transported by argon gas into the plasma torch. The ions produced are entrained in the plasma gas and extracted, by means of a differentially pumped vacuum interface, into a mass spectrometer. The ions flow through a Collision/Reaction Cell (CRC) where depending on the mode of analysis either the ions passes through the cell unhindered (no gas - classical ICP-MS mode), or collisions/reactions chemistry takes place between the ions and a neutral gas such as hydrogen or helium. Any polyatomic ions which give interference on the analytes are removed. The ions are then sorted according to their mass-to-charge ratios by a quadrupole mass spectrometer. The ions are then sorted according to their mass-to-charge ratios by a quadrupole mass spectrometer. The ions transmitted through the quadrupole are quantified by a channel electron multiplier and the ion information processed by a data handling system. Although most polyatomic interference is removed, any additional interference must be assessed and valid corrections applied or the data flagged to indicate problems. Interference correction must include compensation for background ions contributed by the plasma gas, reagents, and constituents of the sample matrix. Instrument drift as well as suppressions or enhancements of instrument response caused by the sample matrix must be corrected for by the use of internal standards. | 6e-05 | Y | Y | Y | 4 | 4/12/2023 | 4/29/2024 | 5/16/2024 2:00:14 PM | 2009-09-08 |  | RDGIBSON |  | Concentration | mg/L | Y | (fl. conc.) |  | Selenium Dissolved (fl. conc.) | 7782-49-2 | Metals - Dissolved; | Metals - Dissolved; EMS code: SE-D | Chemical | Dissolved | LAB
SE-D | mg/L | Selenium Dissolved | F037 | K2S2O8/HCl dig, Hydride,ICP | The water sample is digested with potassium persulphate and concentrated hydrochloric acid. Samples, blanks and standards are placed in an automatic sampler and analyzed by emission spectroscopy utilizing induction coupled argon plasma as an excitation source. All forms of arsenic, selenium and antimony are broken down with acidic persulphate to arsenate, selenate and antimonate. Sodium borohydride is required to reduce the arsenate to arsenite. Hydrochloric acid reduces the selenate to selenite. Potassium iodide is required to reduce the antimonate to antimonite. Hydrides of arsenic, selenium and antimony are formed by the action of sodium borohydride in an acidic medium, and analysed by inductively coupled argon plasma optical emission spectrometry. | 0.0001 | Y | Y | Y | 222 | 8/10/1998 2:45:00 PM | 1/16/2004 | 4/18/2018 9:26:32 AM | 1997-04-02 | 6/3/1998 6:51:46 PM | MPAK |  | Concentration | mg/L | Y | (fl. conc.) |  | Selenium Dissolved (fl. conc.) | 7782-49-2 | Metals - Dissolved; | Metals - Dissolved; EMS code: SE-D | Chemical | Dissolved | LAB
SE-D | mg/L | Selenium Dissolved | METC | MET-D-NP-CCMS-VA | Ultra trace metals in water are analyzed by CRC ICPMS, based on US EPA Method 6020B (July 2014). This procedure is intended for pristine field-filtered acid-preserved water samples. The detection limits (LOR) for this test are based on lab instrumental analysis only, not including filtration. ALS-supplied field filtration equipment does not support these LOR. Therefore, because of the high probability of false positives due to filtration, it is strongly recommended that a filtration blank be analysed to aid in data interpretation. | 4e-05 | Y | Y | Y | 1776 | 4/3/2012 | 8/15/2024 | 8/23/2024 9:50:35 AM | 2019-08-20 |  | KATFRASE |  | Concentration | mg/L | Y | (fl. conc.) |  | Selenium Dissolved (fl. conc.) | 7782-49-2 | Metals - Dissolved; | Metals - Dissolved; EMS code: SE-D | Chemical | Dissolved | LAB
SE-D | mg/L | Selenium Dissolved | F068 | HNO3dig,TeflonTube/Autoclave,ICP-MS | This method is applicable to determiniation of dissolved, extractable, and total recoverable metals in ground waters, surface waters, drinking waters, and wastewaters. Dissolved, extractable, and total recoverable sample fractions are acidified to pH <1.5 with analytical grade nitric acid (HNO3) to prevent the metals from adsorbing to the container material or precipitating. Extractable and total recoverable water samples are collected in 250 ml acid washed polyethylene bottles and are preserved with Nitric acid (0.5 ml acid for 100 mLs of sample). All preserved samples are stored at room temperature and can be held for up to 6 months. Unfiltered, un-preserved samples are stored at 4C for no more than 30 days. For dissolved and extractable water fractions, no sample preparation is necessary. Samples are simply poured in autosampler tubes (i.e. 15 ml centrifuge tubes), and placed on the autosampler and analyzed. Samples for total recoverable metals are digested with Nitric acid in special Teflon tubes in an autoclave prior to analyses. | 0.0002 | Y | Y | Y | 1388 | 6/1/1999 1:22:00 PM | 4/17/2015 1:06:00 PM | 1/29/2021 1:41:30 PM | 1999-06-28 |  | NPEPPIN |  | Concentration | mg/L | Y | (fl. conc.) |  | Selenium Dissolved (fl. conc.) | 7782-49-2 | Metals - Dissolved; | Metals - Dissolved; EMS code: SE-D | Chemical | Dissolved | LAB
SE-D | mg/L | Selenium Dissolved | X285 | HNO3 Dig:ICP Analysis | HNO3 Dig:ICP Analysis | 0.2 | Y | Y | Y | 678 | 2/5/2001 | 11/13/2002 | 10/27/2023 10:45:18 AM | 1998-05-22 | 6/3/1998 6:51:49 PM | NPEPPIN |  | Concentration | mg/L | Y | (fl. conc.) |  | Selenium Dissolved (fl. conc.) | 7782-49-2 | Metals - Dissolved; | Metals - Dissolved; EMS code: SE-D | Chemical | Dissolved | LAB
SE-D | mg/L | Selenium Dissolved | F032 | Filt.,Acid,ICPMS | There is no acid digestion required for this analysis. A water sample is collected and preserved in 0.2 percent nitric acid. To analyse the extractable metals the sample is shaken, allowed to sit for at least twenty four hours, and a decanted portion is introduced to the spectrometer. When the dissolved metals are required, the samples must first be filtered through a 0.45 micron cellulose acetate filter paper and then followed by preservation in 0.2 percent nitric acid. It is important to note that different laboratories may have very different detection limits for some variables analyzed by ICPMS. | 5e-05 | Y | Y | Y | 76562 | 3/10/1998 2:07:00 PM | 8/26/2024 | 8/26/2024 4:53:05 PM | 1998-05-14 | 8/15/2019 10:59:42 AM | NPEPPIN | KATFRASE | Concentration | mg/L | Y | (fl. conc.) |  | Selenium Dissolved (fl. conc.) | 7782-49-2 | Metals - Dissolved; | Metals - Dissolved; EMS code: SE-D | Chemical | Dissolved | LAB
SE-D | mg/L | Selenium Dissolved | X170 | ICP | This code is a catch-all for a large variety of ICP methods including water and also solids digested into solutions as an intermediate result. Different procedures will have greatly different detection limits. (a) For air particulates, preliminary results are reported in mg/L and loading results as mg/dm2/d. (b) For Hi Vol samples preliminiary results are same as reported in mg/L and loading results as ug/m3. (c) Note that for Rain samples there is no filtration nor digestion steps. | 0.03 | Y | Y | Y | 6769 | 4/11/1996 6:02:00 PM | 11/6/2018 | 6/12/2020 1:21:12 PM | 1996-04-12 | 6/3/1998 6:51:38 PM | EMS |  | Concentration | mg/L | Y | (fl. conc.) |  | Selenium Dissolved (fl. conc.) | 7782-49-2 | Metals - Dissolved; | Metals - Dissolved; EMS code: SE-D | Chemical | Dissolved | LAB
SE-D | mg/L | Selenium Dissolved | EQ01 | EQUIS (mg/L) | EQUIS (mg/L) | 0.001 | N | Y | Y | 42 |  |  | 10/24/1996 3:15:52 PM | 1996-04-12 | 6/3/1998 6:51:38 PM | EMS |  | Concentration | mg/L | Y | (fl. conc.) |  | Selenium Dissolved (fl. conc.) | 7782-49-2 | Metals - Dissolved; | Metals - Dissolved; EMS code: SE-D | Chemical | Dissolved | LAB
SE-D | mg/L | Selenium Dissolved | X116 | FF,FA,HNO3,Graph.Furn,AA | Samples are field-filtered (0.45 um), field-acidified with HNO3; analysis by GFAA. | 0.002 | Y | Y | Y | 51 |  |  | 10/29/2001 11:34:43 AM | 1996-04-12 | 6/3/1998 6:51:41 PM | EMS |  | Concentration | mg/L | Y | (fl. conc.) |  | Selenium Dissolved (fl. conc.) | 7782-49-2 | Metals - Dissolved; | Metals - Dissolved; EMS code: SE-D | Chemical | Dissolved | LAB
SE-D | mg/L | Selenium Dissolved | X544 | FF,FA,HNO3,ICPMS | Analytes nebulized followed by ion entrained in plasma gas and extracted by means of a differentially pumped vacuum interface into a mass spectrometer | 0.001 | Y | Y | Y | 30 |  |  | 10/30/2014 8:31:18 AM | 2001-09-18 |  | NPEPPIN |  | Concentration | mg/L | Y | (fl. conc.) |  | Selenium Dissolved (fl. conc.) | 7782-49-2 | Metals - Dissolved; | Metals - Dissolved; EMS code: SE-D | Chemical | Dissolved | LAB
SE-D | mg/L | Selenium Dissolved | X356 | LF,LA:HNO3:ICP | Samples are lab-filtered (0.45 um); lab-acidified with HNO3; analysis by ICP/AES. | 0.03 | Y | Y | Y | 397 |  |  | 3/23/2010 5:08:09 PM | 1996-04-12 | 6/3/1998 6:51:38 PM | EMS |  | Concentration | mg/L | Y | (fl. conc.) |  | Selenium Dissolved (fl. conc.) | 7782-49-2 | Metals - Dissolved; | Metals - Dissolved; EMS code: SE-D | Chemical | Dissolved | LAB
SE-D | mg/L | Selenium Dissolved | F039 | Field filt + acid, ICP-AES | Dissolved and extractable water samples are ready for analysis once collected. Total samples are autoclaved with HNO3 and hydrochloric acid (HCI) in capped Teflon vials. Analysis is performed via Inductively Coupled Argon Plasma Atomic Emission Spectrometry (ICP-AES)(simultaneous multielement analysis), Graphite Furnace Atomic Absorption Spectrometry (GF-AAS)(low level analysis) and/or Flame Atomic Absorption Spectrometry (F-AAS). | 0.05 | Y | Y | Y | 5945 | 7/29/1996 6:03:00 PM | 7/5/2019 1:21:00 PM | 9/11/2019 8:47:50 AM | 1997-05-08 | 6/3/1998 6:51:46 PM | MPAK |  | Concentration | mg/L | Y | (fl. conc.) |  | Selenium Dissolved (fl. conc.) | 7782-49-2 | Metals - Dissolved; | Metals - Dissolved; EMS code: SE-D | Chemical | Dissolved | LAB
SE-D | mg/L | Selenium Dissolved | F023 | HNO3,GFAA | Dissolved and extractable water samples are ready for analysis once collected. Total samples are autoclaved with HNO3 and hydrochloric acid (HCl) in capped Teflon vials. Analysis is performed via Inductively Coupled Argon Plasma Atomic Emission Spectrometry (ICP-AES) (simultaneous multielement analysis), Graphite Furnace Atomic Absorption Spectrometry (GF-AAS) (low level analysis) and/or Flame Atomic Absorption Spectrometry (F-AAS). Note that extractable samples are acidified with 0.5 mL HNO3 per 100 mL of sample at time of sample collection. | 0.001 | Y | Y | Y | 158 | 2/26/1997 4:56:00 PM | 9/16/2003 3:18:00 PM | 11/28/2003 3:22:15 PM | 1996-09-17 | 6/3/1998 6:51:45 PM | MPAK |  | Concentration | mg/L | Y | (fl. conc.) |  | Selenium Dissolved (fl. conc.) | 7782-49-2 | Metals - Dissolved; | Metals - Dissolved; EMS code: SE-D | Chemical | Dissolved | LAB
SE-D | mg/L | Selenium Dissolved | 0185 | Hydr.At.Absorp.Spec(H2O) | Hydr.At.Absorp.Spec(H2O) | 0.5 | Y | Y | Y | 168 | 1/19/2007 | 4/18/2007 | 11/15/2007 1:07:50 PM | 1996-04-12 | 6/3/1998 6:51:38 PM | EMS |  | Concentration | mg/L | Y | (fl. conc.) |  | Selenium Dissolved (fl. conc.) | 7782-49-2 | Metals - Dissolved; | Metals - Dissolved; EMS code: SE-D | Chemical | Dissolved | LAB
SE-D | mg/L | Selenium Dissolved | X202 | FF:FA,HNO3,Hydrid Gen,AA | Samples are field-filtered (0.45 um) and field-preserved (HNO3); analyssis by gaseous hydride AA. | 0.0005 | Y | Y | Y | 7004 |  |  | 8/26/2024 2:41:32 PM | 1996-04-12 | 6/3/1998 6:51:38 PM | EMS |  | Concentration | mg/L | Y | (fl. conc.) |  | Selenium Dissolved (fl. conc.) | 7782-49-2 | Metals - Dissolved; | Metals - Dissolved; EMS code: SE-D | Chemical | Dissolved | LAB
SE-E | ug/g | Selenium Extractable | X228 | Hom,HNO3/HCl,HVAAS | Sample is not pre-dried, but is homogenized and digested with HNO3/HCl prior to analysis by hydride vapour AA Spectrophotometry (HVAAS). Moisture determined on separate aliquot is used to calculate dry weight result. | 0.2 | Y | Y | Y | 70 |  |  | 10/24/1996 2:25:28 PM | 1996-04-12 | 6/3/1998 6:51:41 PM | EMS |  | MassPerMass | mg/kg | Y | (mass conc.) |  | Selenium Extractable (mass conc.) | 7782-49-2 | Metals - Extractable; | Metals - Extractable; EMS code: SE-E | Chemical | Extractable | LAB
SE-E | mg/L | Selenium Extractable | F069 | SEM extr,6N HCl,ICP/OES | 6.0 N Hydrochloric acid (HCI) is added to an aliquot of the sediment sample. The extract produced from the addition of the acid is then analyzed for simultaneously extracted metals (SEM) using inductively coupled plasma - optical emission spectrophotometry. The detection limits depend on the metal parameters in question and the instrumentation used. | 0.2 | Y | Y | Y | 36 | 3/26/2012 3:01:00 AM | 3/26/2012 3:01:00 AM | 2/28/2022 4:37:31 PM | 1999-07-09 |  | NPEPPIN |  | Concentration | mg/L | Y | (fl. conc.) |  | Selenium Extractable (fl. conc.) | 7782-49-2 | Metals - Extractable; | Metals - Extractable; EMS code: SE-E | Chemical | Extractable | LAB
SE-E | mg/L | Selenium Extractable | F068 | HNO3dig,TeflonTube/Autoclave,ICP-MS | This method is applicable to determiniation of dissolved, extractable, and total recoverable metals in ground waters, surface waters, drinking waters, and wastewaters. Dissolved, extractable, and total recoverable sample fractions are acidified to pH <1.5 with analytical grade nitric acid (HNO3) to prevent the metals from adsorbing to the container material or precipitating. Extractable and total recoverable water samples are collected in 250 ml acid washed polyethylene bottles and are preserved with Nitric acid (0.5 ml acid for 100 mLs of sample). All preserved samples are stored at room temperature and can be held for up to 6 months. Unfiltered, un-preserved samples are stored at 4C for no more than 30 days. For dissolved and extractable water fractions, no sample preparation is necessary. Samples are simply poured in autosampler tubes (i.e. 15 ml centrifuge tubes), and placed on the autosampler and analyzed. Samples for total recoverable metals are digested with Nitric acid in special Teflon tubes in an autoclave prior to analyses. | 0.0002 | Y | Y | Y | 4 | 12/13/2011 12:02:00 PM | 12/13/2011 12:02:00 PM | 2/9/2012 12:15:57 PM | 1999-06-28 |  | NPEPPIN |  | Concentration | mg/L | Y | (fl. conc.) |  | Selenium Extractable (fl. conc.) | 7782-49-2 | Metals - Extractable; | Metals - Extractable; EMS code: SE-E | Chemical | Extractable | LAB
SE-E | mg/L | Selenium Extractable | X129 | Extr,Flameless AA | Extr;Flameless AA | 0.0001 | Y | Y | Y | 1333 |  |  | 9/9/1999 12:46:49 PM | 1999-08-06 | 8/6/1999 9:36:08 AM | NPEPPIN | NPEPPIN | Concentration | mg/L | Y | (fl. conc.) |  | Selenium Extractable (fl. conc.) | 7782-49-2 | Metals - Extractable; | Metals - Extractable; EMS code: SE-E | Chemical | Extractable | LAB
SE-E | mg/L | Selenium Extractable | X553 | Metals in water, extractable, ICP-SFMS | The whole water sample is collected in a HDPE/Teflon bottle. In a Class 100 workstation it is preserved with concentrated Nitric Acid at concentration of 0.2 percent (v/v). The acid preservative extracts the metals from material in the sample and a decanted portion is introduced directly to an inductively coupled argon plasma-sector field mass spectrometer (ICP-SFMS) by pneumatic nebulization. Each respective element is measured at a specific mass to charge ratio, m/z value expressed in amu (atomic mass units), and at a specified instrument resolution. LR: Low Resolution ~300, MR: Mid Resolution ~4000 and HR: High Resolution ~10000. Ref: Method number 02-2061, National Laboratory for Environmental Testing, CCIW, Environment Canada, Burlington. Requested April 2003. | 5e-05 | Y | Y | Y | 2877 | 5/22/2003 | 9/21/2007 | 5/29/2024 2:43:20 PM | 2004-01-19 |  | UMULLER |  | Concentration | mg/L | Y | (fl. conc.) |  | Selenium Extractable (fl. conc.) | 7782-49-2 | Metals - Extractable; | Metals - Extractable; EMS code: SE-E | Chemical | Extractable | LAB
SE-E | mg/L | Selenium Extractable | F040 | Field Acid, lab filt, ICP-AES | Dissolved and extractable water samples are ready for analysis once collected. Total samples are autoclaved with HNO3 and hydrochloric acid (HCI) in capped Teflon vials. Analysis is performed via Inductively Coupled Argon Plasma Atomic Emission Spectrometry (ICP-AES)(simultaneous multielement analysis), Graphite Furnace Atomic Absorption Spectrometry (GF-AAS)(low level analysis) and/or Flame Atomic Absorption Spectrometry (F-AAS). | 0.05 | Y | Y | Y | 2361 | 8/5/1997 6:06:00 PM | 3/14/2008 | 9/11/2019 8:47:41 AM | 1997-05-08 | 6/3/1998 6:51:46 PM | MPAK |  | Concentration | mg/L | Y | (fl. conc.) |  | Selenium Extractable (fl. conc.) | 7782-49-2 | Metals - Extractable; | Metals - Extractable; EMS code: SE-E | Chemical | Extractable | LAB
SE-T | mg/L | Selenium Total | 0185 | Hydr.At.Absorp.Spec(H2O) | Hydr.At.Absorp.Spec(H2O) | 0.5 | Y | Y | Y | 208 | 9/24/1996 | 4/19/2007 | 11/15/2007 1:07:50 PM | 1996-04-12 | 6/3/1998 6:51:38 PM | EMS |  | Concentration | mg/L | Y | (fl. conc.) |  | Selenium Total (fl. conc.) | 7782-49-2 | Metals - Total; | Metals - Total; EMS code: SE-T | Chemical | Total | LAB
SE-T | ug/g wet | Selenium Total | X552 | Metals in Biota (wet), total, acid digestion, ICP-SFMS | One gram of wet, homogenized biota is weighed in a microwave TFM vessel, and 8mL of Nitric Acid and 1mL of Hydrogen Peroxide are added. The sample is digested at 200 degrees C for 15 minutes in a closed vessel high pressure microwave oven. The sample is cooled, the volume adjusted to 100mL and introduced to an inductively coupled argon plasma-sector field mass spectrometer (ICP-SFMS) by pneumatic nebulization. Each respective element is measured at a specific mass to charge ratio, m/z value expressed in amu (atomic mass units), and at a specified instrument resolution. LR: Low Resolution ~300, MR: Mid Resolution ~4000. Ref: 02-2704, National Laboratory for Environmental Testing, CCIW, Environment Canada, Burlington. Requested April 2003. | 0.01 | Y | Y | Y | 37 | 8/9/2016 | 10/16/2023 | 11/28/2023 2:21:33 PM | 2004-01-19 | 4/19/2021 1:22:00 PM | UMULLER | EMS | MassPerMass | µg/g | Y | (mass conc.) | (wet) | Selenium Total (wet) (mass conc.) | 7782-49-2 | Metals - Total; | Metals - Total; EMS code: SE-T | Chemical | Total | LAB
SE-T | mg/L | Selenium Total | X354 | LA:HNO3:HVAAS | Samples are lab-acidified with HNO3; HNO3 digestion; analysis by HVAAS. | 0.0005 | Y | Y | Y | 75 | 3/20/1998 | 7/26/2006 | 11/22/2007 11:47:51 AM | 1996-04-12 | 6/3/1998 6:51:38 PM | EMS |  | Concentration | mg/L | Y | (fl. conc.) |  | Selenium Total (fl. conc.) | 7782-49-2 | Metals - Total; | Metals - Total; EMS code: SE-T | Chemical | Total | LAB
SE-T | ug/g wet | Selenium Total | MTIW | MET-WET-CCMS-MID-VA | Metals in Tissue by CRC ICPMS (WET) Methodology: This method is conducted following British Columbia Lab Manual method "Metals in Animal Tissue and Vegetation (Biota) - Prescriptive". Tissue samples are homogenized and sub-sampled prior to hotblock digestion with nitric and hydrochloric acids, in combination with addition of hydrogen peroxide. Instrumental analysis is by collision cell inductively coupled plasma - mass spectrometry (modified from EPA Method 6020A). Method Limitation: This method employs a strong acid/peroxide digestion, and is intended to provide a conservative estimate of bio-available metals. Near complete recoveries are achieved for most toxicologically important metals, but elements associated with recalcitrant minerals may be only partially recovered. | 0.02 | Y | Y | Y | 24 | 11/25/2018 | 12/6/2018 | 7/21/2021 2:55:08 PM | 2018-12-24 |  | EMS |  | MassPerMass | µg/g | Y | (mass conc.) | (wet) | Selenium Total (wet) (mass conc.) | 7782-49-2 | Metals - Total; | Metals - Total; EMS code: SE-T | Chemical | Total | LAB
SE-T | mg/L | Selenium Total | X072 | FA,HNO3,Graph.Furnace,AA | Samples are field-acidified with HNO3; HNO3 digestion; analysis by GFAA. | 0.002 | Y | Y | Y | 245 |  |  | 10/29/2001 11:34:43 AM | 1996-04-12 | 6/3/1998 6:51:38 PM | EMS |  | Concentration | mg/L | Y | (fl. conc.) |  | Selenium Total (fl. conc.) | 7782-49-2 | Metals - Total; | Metals - Total; EMS code: SE-T | Chemical | Total | LAB
SE-T | mg/L | Selenium Total | F048 | Acid dig,conc'n,nebul.,ICP-OES | This method provides an automated procedure for the Inductively Coupled Argon Plasma Optical Emission Spectrophotometric (ICP-OES) determination and quantitation of trace amounts of metals in surface, ground and precipitation waters after a manual acid digestion and concentration procedure and ultrasonic nebulisation. Organically-bound metals and metal complexes are converted to metal ions by digestion with nitric acid [method details says HCl]. To increase the signal detected by the ICP, the samples are concentrated 10 times. The liquid sample is pumped onto the piezoelectric transducer, where the ultrasonic action efficiently converts it to a fine, dense aerosol. The nebulizer carrier gas flow then transports the wet aerosol through a heated U-tube, where the solvent is vaporized. Subsequently, most of the solvent vapour is removed from the aerosol stream by passing the stream through a chilled condenser. The result is a dry, analyte-laden aerosol which is introduced to the plasma. There, the high temperature desolvates the aerosol and dissociates the sample material into individual atoms and ions, which are excited to emit light at wave lengths characteristic of the elements present. Dissolved samples are field filtered through 0.45 micron cellulose acetate filter prior to acid preservation. Samples for extractable analysis are acidified with 2 mL concentrated HNO3 per litre and allowed to stand a minimum of 24 hours before analysis. Clients providing these samples are advised to preserve samples at the collection site. | 0.2 | Y | Y | Y | 29 |  |  | 7/6/2001 11:10:13 AM | 1998-11-13 |  | NPEPPIN |  | Concentration | mg/L | Y | (fl. conc.) |  | Selenium Total (fl. conc.) | 7782-49-2 | Metals - Total; | Metals - Total; EMS code: SE-T | Chemical | Total | LAB
SE-T | mg/L | Selenium Total | X073 | FA,HNO3,Direct Flame,AA | Samples are field-acidified with HNO3; HNO3 digestion; analysis by AA. | 0.2 | Y | Y | Y | 74 |  |  | 6/25/2003 9:22:36 AM | 1996-04-12 | 6/3/1998 6:51:38 PM | EMS |  | Concentration | mg/L | Y | (fl. conc.) |  | Selenium Total (fl. conc.) | 7782-49-2 | Metals - Total; | Metals - Total; EMS code: SE-T | Chemical | Total | LAB
SE-T | mg/L | Selenium Total | X270 | HNO3:Hydride:AA | HNO3:Hydride:AA | 0.05 | Y | Y | Y | 4 |  |  | 10/24/1996 1:51:33 PM | 1996-04-12 | 6/3/1998 6:51:38 PM | EMS |  | Concentration | mg/L | Y | (fl. conc.) |  | Selenium Total (fl. conc.) | 7782-49-2 | Metals - Total; | Metals - Total; EMS code: SE-T | Chemical | Total | LAB
SE-T | mg/L | Selenium Total | F068 | HNO3dig,TeflonTube/Autoclave,ICP-MS | This method is applicable to determiniation of dissolved, extractable, and total recoverable metals in ground waters, surface waters, drinking waters, and wastewaters. Dissolved, extractable, and total recoverable sample fractions are acidified to pH <1.5 with analytical grade nitric acid (HNO3) to prevent the metals from adsorbing to the container material or precipitating. Extractable and total recoverable water samples are collected in 250 ml acid washed polyethylene bottles and are preserved with Nitric acid (0.5 ml acid for 100 mLs of sample). All preserved samples are stored at room temperature and can be held for up to 6 months. Unfiltered, un-preserved samples are stored at 4C for no more than 30 days. For dissolved and extractable water fractions, no sample preparation is necessary. Samples are simply poured in autosampler tubes (i.e. 15 ml centrifuge tubes), and placed on the autosampler and analyzed. Samples for total recoverable metals are digested with Nitric acid in special Teflon tubes in an autoclave prior to analyses. | 0.0002 | Y | Y | Y | 1779 | 6/1/1999 1:33:00 PM | 7/4/2008 9:48:00 AM | 7/30/2024 6:32:12 AM | 1999-06-28 |  | NPEPPIN |  | Concentration | mg/L | Y | (fl. conc.) |  | Selenium Total (fl. conc.) | 7782-49-2 | Metals - Total; | Metals - Total; EMS code: SE-T | Chemical | Total | LAB
SE-T | ug/g wet | Selenium Total | BPSY | MET-WET-BPSY-CCMS-VA | Metals in Animal Tissue and Vegetation (Biota) - Prescriptive. EPA 200.3/6020B This method is conducted following British Columbia Lab Manual method "Metals in Animal Tissue and Vegetation (Biota) - Prescriptive". Tissue samples are digested with nitric and hydrochloric acids; in combination with addition of hydrogen peroxide. Instrumental analysis is by collision cell inductively coupled plasma - mass spectrometry (modified from EPA Method 6020B). | 0.02 | Y | Y | Y | 2 | 11/7/2019 | 11/7/2019 | 6/10/2021 1:57:10 PM | 2021-06-04 |  | EMS |  | MassPerMass | µg/g | Y | (mass conc.) | (wet) | Selenium Total (wet) (mass conc.) | 7782-49-2 | Metals - Total; | Metals - Total; EMS code: SE-T | Chemical | Total | LAB
SE-T | mg/L | Selenium Total | EQ01 | EQUIS (mg/L) | EQUIS (mg/L) | 0 | N | Y | Y | 63 |  |  | 10/22/1996 10:41:01 PM | 1996-04-12 | 6/3/1998 6:51:38 PM | EMS |  | Concentration | mg/L | Y | (fl. conc.) |  | Selenium Total (fl. conc.) | 7782-49-2 | Metals - Total; | Metals - Total; EMS code: SE-T | Chemical | Total | LAB
SE-T | mg/L | Selenium Total | METN | MET-T-NP-CCMS-VA | Ultra trace metals in water are analyzed by CRC ICPMS, based on US EPA Method 6020B (July 2014). The detection limits provided can only be met for undigested samples. This procedure is intended for colorless, non-turbid, acid-preserved water samples (i.e. pristine water samples), having turbidity < 1 NTU and no odor. Where turbidity exceeds 1 NTU, and/or the sample is colored and has an odor, results may be biased low compared to true Total Metals concentrations. ALS recommends that turbidity analysis be requested on samples submitted for this test to aid with interpretation of results. Where turbidity is <1NTU, undigested metals are equivalent to total metals concentrations. | 4e-05 | Y | Y | Y | 3648 | 4/3/2012 | 8/19/2024 | 8/26/2024 1:03:06 PM | 2019-08-20 |  | KATFRASE |  | Concentration | mg/L | Y | (fl. conc.) |  | Selenium Total (fl. conc.) | 7782-49-2 | Metals - Total; | Metals - Total; EMS code: SE-T | Chemical | Total | LAB
SE-T | ug/g | Selenium Total | 0196 | D&G: LTA/HCl: ICP | D&G: LTA/HCl: ICP | 10 | Y | Y | Y | 27 |  |  | 10/24/1996 2:08:13 PM | 1996-04-12 | 6/3/1998 6:51:38 PM | EMS |  | MassPerMass | mg/kg | Y | (mass conc.) |  | Selenium Total (mass conc.) | 7782-49-2 | Metals - Total; | Metals - Total; EMS code: SE-T | Chemical | Total | LAB
SE-T | mg/L | Selenium Total | F038 | HNO3/HCl dig,ICP-AES | Dissolved and extractable water samples are ready for analysis once collected. Total samples are autoclaved with HNO3 and hydrochloric acid (HCI) in capped Teflon vials. Analysis is performed via Inductively Coupled Argon Plasma Atomic Emission Spectrometry (ICP-AES)(simultaneous multielement analysis), Graphite Furnace Atomic Absorption Spectrometry (GF-AAS) (low level analysis) and/or Flame Atomic Absorption Spectrometry (F-AAS). | 0.06 | Y | Y | Y | 5424 | 8/1/1996 6:04:00 PM | 6/25/2010 3:15:00 PM | 9/11/2019 8:47:43 AM | 1997-05-08 | 6/3/1998 6:51:46 PM | MPAK |  | Concentration | mg/L | Y | (fl. conc.) |  | Selenium Total (fl. conc.) | 7782-49-2 | Metals - Total; | Metals - Total; EMS code: SE-T | Chemical | Total | LAB
SE-T | ug/m3 | Selenium Total | F088 | Aliquot,HNO3/HCl dig, ICPMS,loading | Aliquot microwave digested with 3:1 HNO3/HCL. Analysis by ICPMS. For HiVol samples preliminary result are reported in mg/L and loading results are reported in ug/m3 |  | Y | Y | Y | 565 | 8/17/2000 10:46:00 AM | 12/14/2006 | 12/30/2006 1:01:38 AM | 2000-08-16 |  | NPEPPIN |  | AirConcentration | µg/m³ | Y | (air conc.) |  | Selenium Total (air conc.) | 7782-49-2 | Metals - Total; | Metals - Total; EMS code: SE-T | Chemical | Total | LAB
SE-T | ug/g | Selenium Total | 0031 | ICP | ICP | 50 | Y | Y | Y | 3 |  |  | 10/23/1996 11:03:49 PM | 1996-10-11 | 6/3/1998 6:51:46 PM | NPEPPIN |  | MassPerMass | mg/kg | Y | (mass conc.) |  | Selenium Total (mass conc.) | 7782-49-2 | Metals - Total; | Metals - Total; EMS code: SE-T | Chemical | Total | LAB
SE-T | mg/L | Selenium Total | F024 | HNO3/HCl,GFAA | Dissolved and extractable water samples are ready for analysis once collected. Total samples are autoclaved with HNO3 and hydrochloric acid (HCl) in capped Teflon vials. Analysis is performed via Inductively Coupled Argon Plasma Atomic Emission Spectrometry (ICP-AES) (simultaneous multielement analysis), Graphite Furnace Atomic Absorption Spectrometry (GF-AAS) (low level analysis) and/or Flame Atomic Absorption Spectrometry (F-AAS). | 0.001 | Y | Y | Y | 2138 | 3/5/1997 5:03:00 PM | 9/12/2003 3:16:00 PM | 4/18/2016 11:43:13 AM | 1996-09-17 | 6/3/1998 6:51:45 PM | MPAK |  | Concentration | mg/L | Y | (fl. conc.) |  | Selenium Total (fl. conc.) | 7782-49-2 | Metals - Total; | Metals - Total; EMS code: SE-T | Chemical | Total | LAB
SE-T | ug/g | Selenium Total | FS01 | D&G,Nitric/HCl:ICP | Samples are dried at 60 C, then either sieved to -100mesh (150 um) and rolled to homogenize, or finely ground with a mortar and pestle or a shatter box laboratory mill and rolled (depending on the submitter's requirements). The samples are then weighed into Teflon digestion liners and digested with HNO3 and HCl in a microwave oven. Samples are cooled, volumized, and allowed to settle overnight before being decanted. Analysis is then performed via Inductively Coupled Argon Plasma Atomic Emission Spectrometry (ICP-AES) (Simultaneous multi element analysis). In ICP Spectrometry, the liquid sample is pumped to the nebulizer, where it is convertd into an aerosol. the nebulizer carrier gas flow than transports the wet aerosol past the impact bead, where the large droplets condense out and flow to the waste container. The remaining fine aerosol is carried into the plasma. There, the high temperature desolvates the aerosol and dissociates the sample material into individual atoms and ions, which are excited to emit light at wavelengths characteristic of the elements present. A diffraction grating in the spectrometer diffracts the various wavelengths of light through exit slits to the photo-multiplier tubes, which convert the light energy to electronic signals. | 8 | Y | Y | Y | 2643 | 10/21/1996 4:34:00 PM | 10/10/2007 | 5/24/2008 1:15:03 AM | 1997-04-28 | 6/3/1998 6:51:46 PM | MPAK |  | MassPerMass | mg/kg | Y | (mass conc.) |  | Selenium Total (mass conc.) | 7782-49-2 | Metals - Total; | Metals - Total; EMS code: SE-T | Chemical | Total | LAB
SE-T | ug/g wet | Selenium Total | X560 | Metals in Biota (wet), total, acid digestion, ICP-MS | Five grams of wet, homogenized biota is weighed in a centrifuge tube, and 2.5 mL of Nitric Acid and 2.5 mL of Hydrochloric acid are added. The sample is digested at room temperature for one hour and then digested at 90-95 degrees C in a water bath or heat block for one more hour. The sample is cooled and the volume is made to 50 mL with DI. The digestate is then diluted at ten fold and introduced to a collision/reaction cell inductively coupled argon plasma mass spectrometer (CRC ICPMS) by pneumatic nebulization. Each respective element is measured at a specific mass to charge ratio, m/z value expressed in amu (atomic mass units). - Maxxam Analytics (BC). Added to EMS Dec 10, 2008. | 0.01 | Y | Y | Y | 542 | 11/15/2008 4:00:00 AM | 10/30/2023 | 11/28/2023 2:21:33 PM | 2008-12-11 | 3/11/2015 4:17:10 PM | RDGIBSON | EMS | MassPerMass | µg/g | Y | (mass conc.) | (wet) | Selenium Total (wet) (mass conc.) | 7782-49-2 | Metals - Total; | Metals - Total; EMS code: SE-T | Chemical | Total | LAB
SE-T | mg/dm2/d | Selenium Total | DUST | HNO3 dig.(microwave or hot block),ICPMS | HNO3 dig.(microwave or hot block);ICPMS |  | Y | Y | Y | 420 | 7/19/2016 1:11:00 PM | 7/31/2024 | 7/31/2024 11:23:12 AM | 2008-02-07 |  | RDGIBSON |  | MassFlux | mg/dm²/d | Y | (mass flux) |  | Selenium Total (mass flux) | 7782-49-2 | Metals - Total; | Metals - Total; EMS code: SE-T | Chemical | Total | LAB
SE-T | ug/g | Selenium Total | FS02 | D&G,Nitric/HCl:GF | Samples are dried at 60 C, then either sieved to -100 mesh (150 um) and rolled to homogenize, or finely ground with a mortar and pestle or a shatter box laboratory mill and rolled (depending on the submitter's requirements). The samples are then weighed into Teflon digestion liners and digested with HNO3 and HCl in a microwave oven. Samples are cooled, volumized, and allowed to settle overnight before being decanted. Analysis is then performed via Graphite Furnace Atomic Absorption Spectrometry (GF-AAS) (low level analysis). GFAA Spectrometry involves heating the sample in three or more stages. First, a low current heats the graphite tube to dry the sample. The second, or charring stage, destroys the organic matter and volatilizes other components at an intermediate temperature. Finally, a high current heats the tube to incandescence and, in an inert atmosphere, atomizes the element being determined. Additional stages may be added to drying and charring, and to cool and clean the tubes in between samples. The resultant ground-state atomic vapour absorbs monochromatic radiation from the source. A photoelectric detector measures the intensity of the transmitted radiation characteristic of the element being determined, which is inversely proportional to the concentration of ground-state atoms in the optical path over a limited range. Low level Cadmium and Lead are determined using Deuterium lamp background correction system whereas Arsenic and Selenium are determined by using Zeeman background correction system. | 0.5 | Y | Y | Y | 17 | 2/19/1998 3:00:00 PM | 6/6/2001 3:14:00 PM | 8/4/2001 1:01:01 AM | 1999-04-01 |  | NPEPPIN |  | MassPerMass | mg/kg | Y | (mass conc.) |  | Selenium Total (mass conc.) | 7782-49-2 | Metals - Total; | Metals - Total; EMS code: SE-T | Chemical | Total | LAB
SE-T | ppm (S) | Selenium Total | MS41 | ALS Minerals ME-MS41 | ALS Minerals ME-MS41. A prepared sample is digested with aqua regia in a graphite heating block. After cooling; the resulting solution is diluted to with deionized water; mixed and analyzed by ICP-AES. Samples are then analyzed by ICP-MS for the remaining suite of elements. | 0.2 | Y | Y | Y | 0 |  |  |  | 2024-01-29 |  | EMS |  | DimensionlessRatio | ppm | N | (ratio) |  | Selenium Total (ratio) | 7782-49-2 | Metals - Total; | Metals - Total; EMS code: SE-T | Chemical | Total | LAB
SE-T | mg/L | Selenium Total | 1349 | Tot Metals in Seawater by CRC ICPMS (BC) | Tot Metals in Seawater by CRC ICPMS (BC) | 0.0005 | Y | Y | Y | 98 | 8/16/2019 | 12/5/2021 | 12/13/2021 12:34:16 PM | 2019-09-30 |  | KATFRASE |  | Concentration | mg/L | Y | (fl. conc.) |  | Selenium Total (fl. conc.) | 7782-49-2 | Metals - Total; | Metals - Total; EMS code: SE-T | Chemical | Total | LAB
SE-T | mg/L | Selenium Total | F037 | K2S2O8/HCl dig, Hydride,ICP | The water sample is digested with potassium persulphate and concentrated hydrochloric acid. Samples, blanks and standards are placed in an automatic sampler and analyzed by emission spectroscopy utilizing induction coupled argon plasma as an excitation source. All forms of arsenic, selenium and antimony are broken down with acidic persulphate to arsenate, selenate and antimonate. Sodium borohydride is required to reduce the arsenate to arsenite. Hydrochloric acid reduces the selenate to selenite. Potassium iodide is required to reduce the antimonate to antimonite. Hydrides of arsenic, selenium and antimony are formed by the action of sodium borohydride in an acidic medium, and analysed by inductively coupled argon plasma optical emission spectrometry. | 0.0001 | Y | Y | Y | 9737 | 7/4/1997 2:35:00 PM | 1/16/2004 | 9/11/2019 8:46:38 AM | 1997-04-02 | 6/3/1998 6:51:46 PM | MPAK |  | Concentration | mg/L | Y | (fl. conc.) |  | Selenium Total (fl. conc.) | 7782-49-2 | Metals - Total; | Metals - Total; EMS code: SE-T | Chemical | Total | LAB
SE-T | ug/m3 | Selenium Total | 6040 | Aqua Regia Dig:ICP | Aqua regia digestion is used to bring the metals into solution prior to ICP analysis. For Hi-Vol samples, preliminary results are reported as mg/L and24 hr loadings as ug/m3 | 0.01 | Y | Y | Y | 437 | 2/13/2002 | 2/9/2005 | 2/23/2005 1:13:48 AM | 1998-05-05 | 6/3/1998 6:51:48 PM | NPEPPIN |  | AirConcentration | µg/m³ | Y | (air conc.) |  | Selenium Total (air conc.) | 7782-49-2 | Metals - Total; | Metals - Total; EMS code: SE-T | Chemical | Total | LAB
SE-T | mg/L | Selenium Total | F033 | Acid Dig,Centrifuge,ICPMS | For the determination of total recoverable elements in water or waste water, take 100 ml aliquot from a well mixed, acid preserved, sample containing not more than 0.25% (w/v) total solids and transfer to a 250 ml Griffin beaker. (If total solids are greater than 0.25% reduce the size of the aliquot by a proportionate amount). Add 1 ml of conc. nitric acid and 0.5 ml conc. hydrochloric acid. Heat on a hot plate until the volume has been reduced to 15-20ml, ensuring that the sample does not boil. Allow to cool and transfer to a 50 ml volumetric flask. Dilute to volume with ASTM type I water. Centrifuge the sample or allow to stand overnight to separate insoluble material. Prior analysis, pipette 20 ml into a 50 ml volumetric flask and dilute to volume with ASTM type I water. Analysis should be performed within seventy two hours of the completed preparation. Note that different laboratories may have very different field preservation, digestion procedures and detection limits for this method, so these should be documented. | 0.002 | Y | Y | Y | 37079 | 11/20/1997 9:32:00 AM | 8/22/2024 12:35:00 PM | 8/25/2024 9:47:36 PM | 1998-05-14 | 6/3/1998 6:51:49 PM | NPEPPIN |  | Concentration | mg/L | Y | (fl. conc.) |  | Selenium Total (fl. conc.) | 7782-49-2 | Metals - Total; | Metals - Total; EMS code: SE-T | Chemical | Total | LAB
SE-T | mg/L | Selenium Total | X545 | FA,HNO3,ICPMS | Analytes nebulized followed by ion entrained in plasma gas and extracted by means of a differentially pumped vacuum interface into a mass spectrometer | 0.001 | Y | Y | Y | 30 |  |  | 10/30/2014 8:31:18 AM | 2001-09-18 | 9/18/2001 1:36:23 PM | NPEPPIN | NPEPPIN | Concentration | mg/L | Y | (fl. conc.) |  | Selenium Total (fl. conc.) | 7782-49-2 | Metals - Total; | Metals - Total; EMS code: SE-T | Chemical | Total | LAB
SE-T | ug/g | Selenium Total | FB02 | Freeze Dry,Hom,Acid Dig,GF | Sample moisture is determined by weighing samples before and after freeze-drying. Freeze-dried samples are ground in a Braun blender before being weighed into Teflon PFA liners and digested with concentrated HNO3, concentrated HCl, and 50% H2O2 in a microwave oven; which results in a complete sample dissolution. Samples are cooled and volumized with Type I water. Analysis is performed via Inductively Coupled Argon Plasma Atomic Emission Spectrometry (ICP-AES) (simultaneous multielement analysis), Graphite Furnace Atomic Absorption Spectrometry (GF-AAS) (low level Cd and Pb analysis), Cold Vapor Atomic Fluorescence Spectrometry (CV-AFS) (Mercury analysis) and/or Flame Atomic Absorption Spectrometry (F-AAS) (K analysis). | 0.1 | Y | Y | Y | 17 | 6/7/2001 1:41:00 PM | 9/9/2001 1:27:00 PM | 10/5/2001 1:21:26 AM | 1999-04-01 |  | NPEPPIN |  | MassPerMass | mg/kg | Y | (mass conc.) |  | Selenium Total (mass conc.) | 7782-49-2 | Metals - Total; | Metals - Total; EMS code: SE-T | Chemical | Total | LAB
SE-T | ug/g | Selenium Total | X228 | Hom,HNO3/HCl,HVAAS | Sample is not pre-dried, but is homogenized and digested with HNO3/HCl prior to analysis by hydride vapour AA Spectrophotometry (HVAAS). Moisture determined on separate aliquot is used to calculate dry weight result. | 0.05 | Y | Y | Y | 99 |  |  | 10/24/1996 2:26:14 PM | 1996-04-12 | 6/3/1998 6:51:38 PM | EMS |  | MassPerMass | mg/kg | Y | (mass conc.) |  | Selenium Total (mass conc.) | 7782-49-2 | Metals - Total; | Metals - Total; EMS code: SE-T | Chemical | Total | LAB
SE-T | ppm (S) | Selenium Total | MS61 | ALS Minerals ME-MS61 | ALS Minerals ME-MS61. A prepared sample is digested with HNO3; HF; and HClO4 in a graphite heating block to near dryness. After cooling; the sample is boiled in diluted HCl. Bulk to volume with deionized water; mixed and analyzed by ICP-AES and ICP-MS for all elements. | 1 | Y | Y | Y | 0 |  |  |  | 2024-01-29 |  | EMS |  | DimensionlessRatio | ppm | N | (ratio) |  | Selenium Total (ratio) | 7782-49-2 | Metals - Total; | Metals - Total; EMS code: SE-T | Chemical | Total | LAB
SE-T | mg/L | Selenium Total | 0181 | HCl/K2S2O8, Hydride, ICP | HCl/K2S2O8, Hydride; ICP | 0.002 | Y | Y | Y | 59 |  |  | 10/24/1996 1:20:19 PM | 1996-04-12 | 6/3/1998 6:51:41 PM | EMS |  | Concentration | mg/L | Y | (fl. conc.) |  | Selenium Total (fl. conc.) | 7782-49-2 | Metals - Total; | Metals - Total; EMS code: SE-T | Chemical | Total | LAB
SE-T | mg/L | Selenium Total | MWID | Metals in Water by ICPMS (digested) | Metals in Water by ICPMS (digested) Based on EPA 200.8 | 0.0008 | Y | Y | Y | 44598 |  |  | 7/31/2024 1:46:41 PM | 2009-12-24 |  | RDGIBSON |  | Concentration | mg/L | Y | (fl. conc.) |  | Selenium Total (fl. conc.) | 7782-49-2 | Metals - Total; | Metals - Total; EMS code: SE-T | Chemical | Total | LAB
SE-T | mg/L | Selenium Total | 0040 | HNO3 Dig: ICP Analysis | HNO3 Dig: ICP Analysis | 0.03 | Y | Y | Y | 2316 | 1/25/2001 | 4/16/2001 | 1/30/2013 3:34:51 PM | 1996-04-12 | 6/3/1998 6:51:38 PM | EMS |  | Concentration | mg/L | Y | (fl. conc.) |  | Selenium Total (fl. conc.) | 7782-49-2 | Metals - Total; | Metals - Total; EMS code: SE-T | Chemical | Total | LAB
SE-T | mg/L | Selenium Total | X285 | HNO3 Dig:ICP Analysis | HNO3 Dig:ICP Analysis | 0.2 | Y | Y | Y | 653 | 8/12/1993 | 10/6/1993 | 4/30/2004 11:32:38 AM | 1996-04-12 | 6/3/1998 6:51:38 PM | EMS |  | Concentration | mg/L | Y | (fl. conc.) |  | Selenium Total (fl. conc.) | 7782-49-2 | Metals - Total; | Metals - Total; EMS code: SE-T | Chemical | Total | LAB
SE-T | ug/g | Selenium Total | FB09 | Homo,Freeze Dry,Acid Dig,Graphite Furnace AA | Sample moisture is determined by weighing samples before and after freeze-drying. Freeze-dried samples are ground in a Braun blender before being weighed into Teflon PFA liners and digested with concentrated nitric, concentrated hydrochloric and 30% hydrogen peroxide in a microwave oven; which results in a complete sample dissolution. Samples are cooled and volumized with Type I water. The digestion in a mixture of acids using microwave digestion techniques result in the complete extraction of selenium from tissues. Analysis is performed via a graphite furnace atomic absorption (GFAA) technique, using Zeeman background correction system. GFAA Spectroscopy involves heating the sample in 3 or more stages. First, a low current heats the graphite tube to dry the sample. The second, or charring stage, destroys organic matter and volatilizes other components at an intermediate temperature. Finally, a high current heats the tube to incandescence and, in an inert atmosphere, atomizes the element being determined. Additional stages may be added to aid in drying and charring, and to clean and cool the tube between samples. The resultant ground-state atomic vapour absorbs monochromatic radiation from the source. A photoelectric detector measures the intensity of transmitted radiation characteristic of the Se or Sb absorption wavelength, which is inversely proportional to the quantity (concentration) of ground-state atoms in the optical path over a limited range. | 0.1 | N | Y | Y | 1 | 9/15/1997 2:02:00 PM | 9/15/1997 2:02:00 PM | 4/9/1998 6:24:05 PM | 1997-02-27 | 4/1/1999 3:32:08 PM | MPAK | NPEPPIN | MassPerMass | mg/kg | Y | (mass conc.) |  | Selenium Total (mass conc.) | 7782-49-2 | Metals - Total; | Metals - Total; EMS code: SE-T | Chemical | Total | LAB
SE-T | ug/g | Selenium Total | EQ02 | EQUIS (ug/g) | EQUIS (ug/g) | 10 | N | Y | Y | 743 |  |  | 10/23/1996 12:08:28 AM | 1996-04-12 | 6/3/1998 6:51:41 PM | EMS |  | MassPerMass | mg/kg | Y | (mass conc.) |  | Selenium Total (mass conc.) | 7782-49-2 | Metals - Total; | Metals - Total; EMS code: SE-T | Chemical | Total | LAB
SE-T | ug/g | Selenium Total | 0192 | Hom, Nitric/Perchlor,ICP | Sample already in a dry state so does not require drying before homogenization and nitric perchlorid digestion followed by ICP analysis. | 10 | Y | Y | Y | 744 | 2/13/2002 | 5/29/2002 | 1/14/2004 1:00:44 AM | 1996-04-12 | 6/3/1998 6:51:41 PM | EMS |  | MassPerMass | mg/kg | Y | (mass conc.) |  | Selenium Total (mass conc.) | 7782-49-2 | Metals - Total; | Metals - Total; EMS code: SE-T | Chemical | Total | LAB
SE-T | mg/L | Selenium Total | F121 | Acid,ICPMS | Acidified, no digestion, ICPMS | 4e-05 | Y | Y | Y | 8440 | 3/12/2010 7:57:00 PM | 8/15/2024 9:10:00 PM | 8/26/2024 8:06:03 AM | 2002-07-22 | 7/16/2012 3:30:16 PM | UMULLER | RDGIBSON | Concentration | mg/L | Y | (fl. conc.) |  | Selenium Total (fl. conc.) | 7782-49-2 | Metals - Total; | Metals - Total; EMS code: SE-T | Chemical | Total | LAB
SE-T | mg/L | Selenium Total | X561 | Total metals: FA, HNO3, HR-ICPMS (No Prep)Prep) | Ultra trace metals in water are analyzed by high resolution inductively coupled plasma mass spectrometry (HR-ICPMS), modified from US EPA 200.8, (Revision 5.5). The detection limits provided can only be met for undigested samples. This procedure is intended for colorless, non-turbid, acid-preserved water samples (i.e. pristine water samples), having turbidity < 1 NTU and no odor. Where turbidity exceeds 1 NTU, and/or the sample is colored and has an odor, results may be biased low compared to true Total Metals concentrations. ALS recommends that turbidity analysis be requested on samples submitted for this test to aid with interpretation of results. | 4e-05 | Y | Y | Y | 2756 | 6/8/2015 | 10/5/2023 | 1/31/2024 2:37:41 PM | 2015-10-22 |  | RDGIBSON |  | Concentration | mg/L | Y | (fl. conc.) |  | Selenium Total (fl. conc.) | 7782-49-2 | Metals - Total; | Metals - Total; EMS code: SE-T | Chemical | Total | LAB
SE-T | mg/L | Selenium Total | X345 | FA:HCL:HVAAS | Samples are field-acidified with HCl; HCl digestion; analysis by HVAAS. | 0.0005 | Y | Y | Y | 3327 |  |  | 8/28/2018 3:14:04 PM | 1996-04-12 | 6/3/1998 6:51:38 PM | EMS |  | Concentration | mg/L | Y | (fl. conc.) |  | Selenium Total (fl. conc.) | 7782-49-2 | Metals - Total; | Metals - Total; EMS code: SE-T | Chemical | Total | LAB
SE-T | ug/g | Selenium Total | FS82 | PBM,dry60C,10mesh,SALM-ICP-MS | PBM;dry60C;10mesh;SALM-ICP-MS | 0.5 | Y | Y | Y | 3309 | 5/21/2002 | 7/24/2024 | 7/30/2024 10:00:14 AM | 2002-11-12 |  | UMULLER |  | MassPerMass | mg/kg | Y | (mass conc.) |  | Selenium Total (mass conc.) | 7782-49-2 | Metals - Total; | Metals - Total; EMS code: SE-T | Chemical | Total | LAB
SE-T | mg/L | Selenium Total | 0042 | GI HNO3 Dig: ICP | Sample aliquot taken from unfiltered unpreserved General Ions bottle; HNO3 digestion; ICP analysis. | 0.03 | Y | Y | Y | 6024 | 1/14/2002 | 7/5/2019 3:21:00 PM | 7/12/2019 6:49:13 PM | 1996-04-12 | 6/3/1998 6:51:38 PM | EMS |  | Concentration | mg/L | Y | (fl. conc.) |  | Selenium Total (fl. conc.) | 7782-49-2 | Metals - Total; | Metals - Total; EMS code: SE-T | Chemical | Total | LAB
SE-T | ug/g | Selenium Total | FB04 | Freeze Dry,Hom,Acid Dig,ICP | Sample moisture is determined by weighing samples before and after freeze-drying. Freeze-dried samples are ground in a Braun blender before being weighed into Teflon PFA liners and digested with concentrated HNO3, concentrated HCl, and 50% H2O2 in a microwave oven; which results in a complete sample dissolution. Samples are collected and volumized with Type I water. Analysis is performed via Inductively Coupled Argon Plasma Atomic Emission Spectrometry (ICP-AES) (simultaneous multielement analysis), Graphite Furnace Atomic Absorption Spectrometry (GF-AAS) (low level Cd and Pb analysis), Cold Vapor Atomic Fluorescence Spectrometry (CV-AFS) (Mercury analysis) and/or Flame Atomic Absorption Spectrometry (F-AAS) (K analysis). | 4 | Y | Y | Y | 245 | 8/22/1996 6:11:00 PM | 12/1/2010 | 12/9/2010 3:55:27 PM | 1997-02-27 | 6/3/1998 6:51:46 PM | MPAK |  | MassPerMass | mg/kg | Y | (mass conc.) |  | Selenium Total (mass conc.) | 7782-49-2 | Metals - Total; | Metals - Total; EMS code: SE-T | Chemical | Total | LAB
SE-T | ug/g | Selenium Total | MTID | MET-DRY-CCMS-VA | Metals in Tissue by CRC ICPMS (DRY) Methodology: This method is conducted following British Columbia Lab Manual method "Metals in Animal Tissue and Vegetation (Biota) - Prescriptive". Tissue samples are homogenized and sub-sampled prior to hotblock digestion with nitric and hydrochloric acids, in combination with addition of hydrogen peroxide. Instrumental analysis is by collision cell inductively coupled plasma - mass spectrometry (modified from EPA Method 6020A). Method Limitation: This method employs a strong acid/peroxide digestion, and is intended to provide a conservative estimate of bio-available metals. Near complete recoveries are achieved for most toxicologically important metals, but elements associated with recalcitrant minerals may be only partially recovered. | 0.05 | Y | Y | Y | 122 | 8/13/2014 | 9/8/2022 | 10/26/2022 7:14:15 AM | 2015-03-09 | 3/11/2015 4:12:05 PM | RDGIBSON | EMS | MassPerMass | mg/kg | Y | (mass conc.) |  | Selenium Total (mass conc.) | 7782-49-2 | Metals - Total; | Metals - Total; EMS code: SE-T | Chemical | Total | LAB
SE-T | mg/L | Selenium Total | 0180 | HNO3 Dig:Hydride ICP | HNO3 Dig:Hydride ICP | 0.005 | Y | Y | Y | 114 |  |  | 10/24/1996 1:43:19 PM | 1996-04-12 | 6/3/1998 6:51:38 PM | EMS |  | Concentration | mg/L | Y | (fl. conc.) |  | Selenium Total (fl. conc.) | 7782-49-2 | Metals - Total; | Metals - Total; EMS code: SE-T | Chemical | Total | LAB
SE-T | ug/g | Selenium Total | ICPM | Metals in Soil by CRC ICPMS | This analysis is carried out using procedures from CSR Analytical Method: "Strong Acid Leachable Metals (SALM) in Soil", BC Ministry of Environment, 26 June 2009, and procedures adapted from EPA Method 200.2. The sample is manually homogenized, dried at 60 degrees Celsius, sieved through a 63 um (230 mesh) sieve, and a representative subsample of the dry material is weighed. The sample is then digested at 95 degrees Celsius for 2 hours by block digester using concentrated nitric and hydrochloric acids. Instrumental analysis is by collision cell inductively coupled plasma - mass spectrometry (modifed from EPA Method 6020A). Method Limitation: This method is not a total digestion technique. It is a very strong acid digestion that is intended to dissolve those metals that may be environmentally available. By design, elements bound in silicate structures are not normally dissolved by this procedure as they are not usually mobile in the environment. | 0.1 | Y | Y | Y | 103 | 9/19/2017 | 9/24/2023 | 1/12/2024 11:16:52 AM | 2018-01-09 |  | JKROGH |  | MassPerMass | mg/kg | Y | (mass conc.) |  | Selenium Total (mass conc.) | 7782-49-2 | Metals - Total; | Metals - Total; EMS code: SE-T | Chemical | Total | LAB
SE-T | mg/L | Selenium Total | TCLP | Toxicity Characteristic Leaching Procedure | The TCLP procedure is designed to simulate the climatic leaching action | 0.1 | Y | Y | Y | 38 | 12/8/2009 7:49:00 PM | 9/28/2011 5:53:00 PM | 8/24/2016 2:59:13 PM | 2009-03-06 |  | JCHUDAK |  | Concentration | mg/L | Y | (fl. conc.) |  | Selenium Total (fl. conc.) | 7782-49-2 | Metals - Total; | Metals - Total; EMS code: SE-T | Chemical | Total | LAB
SE-T | ug/g | Selenium Total | X551 | Metals in Biota (dry), total, acid digestion, ICP-SFMS | 0.1 gram of dried, homogenized biota is weighed in a microwave TFM vessel, and 8mL of Nitric Acid and 1mL of Hydrogen Peroxide are added. The sample is digested at 200 degrees C for 15 minutes in a closed vessel high pressure microwave oven. The sample is cooled, the volume adjusted to 100mL and introduced to an iductively coupled argon plasma-sector field mass spectrometer (ICP-SFMS) by pneumatic nebulization. Each respective element is measured at a specific mass to charge ratio, m/z value expressed in amu (atomic mass units), and at a specified instrument resolution. LR: Low Resolution ~300, MR: Mid Resolution ~4000 and HR: Hich Resolution ~10000. Ref: Method number 02-2705, National Laboratory for Environmental Testing, CCIW, Environment Canada, Burlington. Requested April 2003. | 0.1 | Y | Y | Y | 4354 | 8/11/2015 | 10/30/2023 | 8/26/2024 11:43:29 AM | 2004-01-19 |  | UMULLER |  | MassPerMass | mg/kg | Y | (mass conc.) |  | Selenium Total (mass conc.) | 7782-49-2 | Metals - Total; | Metals - Total; EMS code: SE-T | Chemical | Total | LAB
SE-T | mg/L | Selenium Total | F082 | Aliquot,HNO3/HCL dig,ICPMS | Aliquot microwave digested with 3:1 HNO3/HCL. Analysis by ICPMS. For HiVol samples preliminary result are reported in mg/L and loading results are reported in ug/m3 | 5e-05 | Y | Y | Y | 29122 | 7/26/2005 7:04:00 AM | 8/23/2024 | 8/26/2024 4:53:05 PM | 2000-06-13 | 8/15/2019 11:11:25 AM | NPEPPIN | KATFRASE | Concentration | mg/L | Y | (fl. conc.) |  | Selenium Total (fl. conc.) | 7782-49-2 | Metals - Total; | Metals - Total; EMS code: SE-T | Chemical | Total | LAB
SE-T | ug/g | Selenium Total | X230 | Hom,HNO3/HCl,ICP | The sample is dried, ground to pass 100 mesh, and digested with a nitric-hydrochloric acid mixture to solubilize the solid matter and to remove the organic material by oxidation and volatilization. The completed digests are submitted to metals instrumentation personnel for metals analysis by Inductively Coupled Atomic Emission Spectrophotometry (ICP-AES), or hydride analysis by atomic absorption. | 10 | Y | Y | Y | 5 | 12/30/2005 | 12/30/2005 | 2/3/2006 1:13:13 AM | 1996-04-12 | 6/3/1998 6:51:38 PM | EMS |  | MassPerMass | mg/kg | Y | (mass conc.) |  | Selenium Total (mass conc.) | 7782-49-2 | Metals - Total; | Metals - Total; EMS code: SE-T | Chemical | Total | LAB
SE-T | mg/L | Selenium Total | X355 | LA:HCL:HVAAS | Samples are lab-acidified with HCl; HCl digestion; analysis by HVAAS. | 0.0005 | Y | Y | Y | 2179 |  |  | 5/23/2013 1:08:49 PM | 1996-04-12 | 6/3/1998 6:51:38 PM | EMS |  | Concentration | mg/L | Y | (fl. conc.) |  | Selenium Total (fl. conc.) | 7782-49-2 | Metals - Total; | Metals - Total; EMS code: SE-T | Chemical | Total | LAB
SE-T | ppm (S) | Selenium Total | 6042 | Aqua Regia by ICPMS | These were sediment/soils digested with Aqua-Regia (Nitric Acid & Hydrochloric Acids) and then analysis by ICPMS. Results would be reported in ppm s. | 0.1 | Y | Y | Y | 71 | 5/14/2022 | 7/27/2024 | 7/29/2024 9:13:08 AM | 2022-06-28 |  | EMS |  | DimensionlessRatio | ppm | N | (ratio) |  | Selenium Total (ratio) | 7782-49-2 | Metals - Total; | Metals - Total; EMS code: SE-T | Chemical | Total | LAB
SE-T | mg/L | Selenium Total | MSEA | Metals in Seawater (undigested) by (CRC) ICP-MS | This method describes the multi-elemental determination of trace elements by Collision/Reaction Cell (CRC) ICP-MS. The method measures ions produced by a radio-frequency inductively coupled plasma. Analyte species originating in a liquid are pneumatically nebulized and the resulting aerosol transported by argon gas into the plasma torch. The ions produced are entrained in the plasma gas and extracted, by means of a differentially pumped vacuum interface, into a mass spectrometer. The ions flow through a Collision/Reaction Cell (CRC) where depending on the mode of analysis either the ions passes through the cell unhindered (no gas - classical ICP-MS mode), or collisions/reactions chemistry takes place between the ions and a neutral gas such as hydrogen or helium. Any polyatomic ions which give interference on the analytes are removed. The ions are then sorted according to their mass-to-charge ratios by a quadrupole mass spectrometer. The ions are then sorted according to their mass-to-charge ratios by a quadrupole mass spectrometer. The ions transmitted through the quadrupole are quantified by a channel electron multiplier and the ion information processed by a data handling system. Although most polyatomic interference is removed, any additional interference must be assessed and valid corrections applied or the data flagged to indicate problems. Interference correction must include compensation for background ions contributed by the plasma gas, reagents, and constituents of the sample matrix. Instrument drift as well as suppressions or enhancements of instrument response caused by the sample matrix must be corrected for by the use of internal standards. | 6e-05 | Y | Y | Y | 441 | 8/12/2009 12:08:00 AM | 4/29/2024 | 6/7/2024 10:03:16 AM | 2009-09-08 |  | RDGIBSON |  | Concentration | mg/L | Y | (fl. conc.) |  | Selenium Total (fl. conc.) | 7782-49-2 | Metals - Total; | Metals - Total; EMS code: SE-T | Chemical | Total | LAB
SE-T | ppm (S) | Selenium Total | 6041 | Aqua Regia by ICPMS | These were sediment/soils digested with Aqua-Regia (Nitric Acid & Hydrochloric Acids) and then analysis by ICPMS. Results would be reported in mg/kg and/or %. | 0.003 | Y | Y | Y | 21 | 10/13/2019 | 8/5/2020 | 8/10/2020 3:52:07 PM | 2019-09-06 | 10/21/2019 11:56:25 AM | KATFRASE | KATFRASE | DimensionlessRatio | ppm | N | (ratio) |  | Selenium Total (ratio) | 7782-49-2 | Metals - Total; | Metals - Total; EMS code: SE-T | Chemical | Total | LAB
SE-T | mg/L | Selenium Total | X170 | ICP | This code is a catch-all for a large variety of ICP methods including water and also solids digested into solutions as an intermediate result. Different procedures will have greatly different detection limits. (a) For air particulates, preliminary results are reported in mg/L and loading results as mg/dm2/d. (b) For Hi Vol samples preliminiary results are same as reported in mg/L and loading results as ug/m3. (c) Note that for Rain samples there is no filtration nor digestion steps. | 0.03 | Y | Y | Y | 5690 | 3/28/1996 6:03:00 PM | 11/7/2018 | 10/24/2022 10:58:25 AM | 1996-04-12 | 6/3/1998 6:51:38 PM | EMS |  | Concentration | mg/L | Y | (fl. conc.) |  | Selenium Total (fl. conc.) | 7782-49-2 | Metals - Total; | Metals - Total; EMS code: SE-T | Chemical | Total | LAB
SE-T | mg/L | Selenium Total | X071 | FA,HNO3,Hydr.Gen,AA | FA;HNO3;Hydr.Gen;AA | 0.0005 | Y | Y | Y | 6525 |  |  | 8/26/2024 2:41:32 PM | 1996-04-12 | 6/3/1998 6:51:42 PM | EMS |  | Concentration | mg/L | Y | (fl. conc.) |  | Selenium Total (fl. conc.) | 7782-49-2 | Metals - Total; | Metals - Total; EMS code: SE-T | Chemical | Total | LAB
SE-T | ug/m3 | Selenium Total | 5312 | Tef:N/P Dig:ICP Analysis | A nitric/perchloric acid digestion if used to bring the metals into solution prior to ICP analysis. For Hi-Vol samples preliminary results are reported as mg/L and 24 hour loading results as ug/m3. | 0.01 | Y | Y | Y | 1185 | 11/7/1997 | 3/29/2000 | 9/28/2000 1:34:44 PM | 1996-10-13 | 6/3/1998 6:51:46 PM | EMS |  | AirConcentration | µg/m³ | Y | (air conc.) |  | Selenium Total (air conc.) | 7782-49-2 | Metals - Total; | Metals - Total; EMS code: SE-T | Chemical | Total | LAB
SE-T | ug/g | Selenium Total | X492 | Homo,HCl-HNO3 dig,ICP-OES,dry calc | Homogenize samples and take sub-samples. For Total Recoverable Metals Digest with Aqua Regia. Analyses by ICPOES (Inductively Coupled Plasma Optical Emission Spectroscopy). Perform moisture analysis on separate sub-sample and calculate dry value for trace elements. | 50 | Y | Y | Y | 1 |  |  | 8/24/2016 2:59:13 PM | 1998-11-13 |  | NPEPPIN |  | MassPerMass | mg/kg | Y | (mass conc.) |  | Selenium Total (mass conc.) | 7782-49-2 | Metals - Total; | Metals - Total; EMS code: SE-T | Chemical | Total | LAB
SE-T | ug/g | Selenium Total | PB02 | PBM,HNO3/HCl dig,dry60C,100mesh, | Performance Based Method. Samples are dried at 60 C and sieved using a .149 mm (100 mesh) sieve. Solid samples are digested with the combination of nitric and hydrochloric acids with a 1:1 ratio of acids. For environmental purposes, strong acid leaches are used routinely since they have the advantage of speed, low cost, and superio reproducibility. All metals trapped within the refractory matrix are unlikely to have environmental consequences. Instrumental analysis of extracts produced using this digestion method can be performed by a variety of techniques. |  | Y | Y | Y | 52 | 5/24/2001 3:42:00 PM | 11/26/2001 10:23:00 AM | 5/26/2003 2:28:26 PM | 2001-07-18 |  | NPEPPIN |  | MassPerMass | mg/kg | Y | (mass conc.) |  | Selenium Total (mass conc.) | 7782-49-2 | Metals - Total; | Metals - Total; EMS code: SE-T | Chemical | Total | LAB
SE-T | mg/L | Selenium Total | X555 | Metals in water, total, "in-bottle" digestion, ICP-SFMS | The whole water sample collected in a HDPE/Teflon bottle, is de-capped in a Class 100 workstation, acidified to 2% (v/v) HNO3 and 1% (v/v) HCl and recapped. The sample is "in-bottle" digested in a convection oven, set at 60 degrees C for 16 hours. The sample is cooled and introduced to an inductively coupled argon plasma-sector field mass spectrometer (ICP-SFMS) by pneumatic nebulization. Each respective element is measured at a specific mass to charge ratio, m/z value expressed in amu (atomic mass units), and at resolution. LR: Low Resolution ~300, MR: Mid Resolution ~4000 and HR: High a specified instrument Resolution ~10000. Ref: Method number 02-2003, National Laboratory for Environmental Testing, CCIW, Environment Canada, Burlington. Requested April 2003. | 5e-05 | Y | Y | Y | 17342 | 5/20/2003 | 12/6/2023 2:08:00 AM | 2/14/2024 9:05:32 AM | 2004-01-19 |  | UMULLER |  | Concentration | mg/L | Y | (fl. conc.) |  | Selenium Total (fl. conc.) | 7782-49-2 | Metals - Total; | Metals - Total; EMS code: SE-T | Chemical | Total | LAB
SE-T | ug/g | Selenium Total | 0190 | D&G:N/P Dig:ICP Anal. | Sample is dried and grind-mixed before nitric perchloric digestion followed by ICP analysis. | 10 | Y | Y | Y | 944 |  |  | 10/24/1996 3:06:16 PM | 1996-04-12 | 6/3/1998 6:51:38 PM | EMS |  | MassPerMass | mg/kg | Y | (mass conc.) |  | Selenium Total (mass conc.) | 7782-49-2 | Metals - Total; | Metals - Total; EMS code: SE-T | Chemical | Total | LAB
SECN | mg/L | selenocyanate SeCN(-1) | X564 | IC-ICP-CRC-MS | Ion chromatography inductively coupled plasma collision reaction cell mass spectrometry | 5e-06 | Y | Y | Y | 6112 | 12/13/2021 | 8/13/2024 | 8/19/2024 12:19:03 PM | 2016-11-16 | 3/23/2017 9:37:36 AM | RRENTON | RRENTON | Concentration | mg/L | Y | (fl. conc.) |  | selenocyanate SeCN(-1) (fl. conc.) | 5749-48-4 | Organometallics; | Organometallics; EMS code: SECN | Chemical |  | LAB
SEME | mg/L | selenomethionine CH3SeCH2CH2CH(NH2)CO2H | E540 | Miekeley Spect. Acta B 60 (2005) 633-641 | E540 and E540B: Instrumental analysis is by Anion Exchange HPLC-ICPMS. Reported Total Unknown Se Species is semi-quantitative in nature and consists of a sum of all the unknown peaks observed in the chromatogram. Field filtration is recommended. | 0.0001 | Y | Y | Y | 176 | 12/13/2021 | 8/11/2024 | 8/19/2024 12:19:03 PM | 2022-01-18 |  | RDGIBSON |  | Concentration | mg/L | Y | (fl. conc.) |  | selenomethionine CH3SeCH2CH2CH(NH2)CO2H (fl. conc.) | 1464-42-2 | Organometallics; | Organometallics; EMS code: SEME | Chemical |  | LAB
SEME | mg/L | selenomethionine CH3SeCH2CH2CH(NH2)CO2H | X564 | IC-ICP-CRC-MS | Ion chromatography inductively coupled plasma collision reaction cell mass spectrometry | 5e-06 | Y | Y | Y | 5902 |  |  | 7/30/2024 12:34:59 PM | 2016-11-16 | 3/23/2017 9:36:32 AM | RRENTON | RRENTON | Concentration | mg/L | Y | (fl. conc.) |  | selenomethionine CH3SeCH2CH2CH(NH2)CO2H (fl. conc.) | 1464-42-2 | Organometallics; | Organometallics; EMS code: SEME | Chemical |  | LAB
SEO3 | mg/L | selenite SeO3(-2) | X564 | IC-ICP-CRC-MS | Ion chromatography inductively coupled plasma collision reaction cell mass spectrometry | 5e-06 | Y | Y | Y | 6546 | 12/13/2021 | 8/13/2024 | 8/19/2024 12:19:03 PM | 2016-11-16 | 3/23/2017 9:35:41 AM | RRENTON | RRENTON | Concentration | mg/L | Y | (fl. conc.) |  | selenite SeO3(-2) (fl. conc.) | 14124-67-5 | Organometallics; | Organometallics; EMS code: SEO3 | Chemical |  | LAB
SEO4 | mg/L | selenate SeO4(-2) | X564 | IC-ICP-CRC-MS | Ion chromatography inductively coupled plasma collision reaction cell mass spectrometry | 5e-06 | Y | Y | Y | 6546 | 12/13/2021 | 8/13/2024 | 8/19/2024 12:19:03 PM | 2016-11-16 | 3/23/2017 9:36:13 AM | RRENTON | RRENTON | Concentration | mg/L | Y | (fl. conc.) |  | selenate SeO4(-2) (fl. conc.) | 14124-68-6 | Organometallics; | Organometallics; EMS code: SEO4 | Chemical |  | LAB
SESO | mg/L | Selenosulfate | X564 | IC-ICP-CRC-MS | Ion chromatography inductively coupled plasma collision reaction cell mass spectrometry | 5e-06 | Y | Y | Y | 5901 |  |  | 7/30/2024 12:34:59 PM | 2016-11-16 | 3/23/2017 9:38:03 AM | RRENTON | RRENTON | Concentration | mg/L | Y | (fl. conc.) |  | Selenosulfate (fl. conc.) | 16450-48-9 | Organometallics; | Organometallics; EMS code: SESO | Chemical |  | LAB
SEUN | mg/L | Unknown Selenium Species | X564 | IC-ICP-CRC-MS | Ion chromatography inductively coupled plasma collision reaction cell mass spectrometry | 5e-06 | Y | Y | Y | 6111 | 12/13/2021 | 8/13/2024 | 8/19/2024 12:19:03 PM | 2016-11-16 | 3/23/2017 9:38:27 AM | RRENTON | RRENTON | Concentration | mg/L | Y | (fl. conc.) |  | Unknown Selenium Species (fl. conc.) |  | Metals; | Metals; EMS code: SEUN | Chemical |  | LAB
SGWL | m | Static Waterlevel-Depth of reading below surface | XM18 | Wetted Tape | Wetted Tape |  | Y | Y | Y | 18 |  |  | 6/10/2021 12:19:29 PM | 2003-11-07 |  | UMULLER |  | Length | m | Y | (len.) |  | Static Waterlevel-Depth of reading below surface (len.) |  | Physical Properties; | Physical Properties; EMS code: SGWL | Physical |  | LAB
SGWL | m | Static Waterlevel-Depth of reading below surface | XG05 | Staff Gauge | Staff Gauge |  | Y | Y | Y | 550 |  |  | 4/30/2021 4:41:02 PM | 2003-11-07 |  | UMULLER |  | Length | m | Y | (len.) |  | Static Waterlevel-Depth of reading below surface (len.) |  | Physical Properties; | Physical Properties; EMS code: SGWL | Physical |  | LAB
SHCLL | % | Sulfate (as S); HCl leach | GRA06A | ALS Minerals S-GRA06A | ALS Minerals S-GRA06A. A prepared sample is heated with dilute HCl and then filtered. The sulfate in the filtrate is precipitated with BaCl2 in a dilute HCl. The BaSO4 precipitate is filtered; ignited; weighed and the Sulfate Sulfur is determined (as S). | 0.01 | Y | Y | Y | 0 |  |  |  | 2024-01-29 |  | EMS |  | DimensionlessRatio | % | N | (ratio) |  | Sulfate (as S); HCl leach (ratio) |  |  | EMS code: SHCLL | Chemical |  | LAB
SI-D | mg/L | Silicon Dissolved | MWID | Metals in Water by ICPMS (digested) | Metals in Water by ICPMS (digested) Based on EPA 200.8 | 2 | Y | Y | Y | 918 |  |  | 7/31/2024 1:46:41 PM | 2009-12-24 |  | RDGIBSON |  | Concentration | mg/L | Y | (fl. conc.) |  | Silicon Dissolved (fl. conc.) | 7440-21-3 | Metals - Dissolved; | Metals - Dissolved; EMS code: SI-D | Chemical | Dissolved | LAB
SI-D | mg/L | Silicon Dissolved | F049 | Filt,HNO3 dig,conc'n,nebul'ICP-OES | This method provides an automated procedure for the Inductively Coupled Argon Plasma Optical Emission Spectrophotometric (ICP-OES) determination and quantitation of trace amounts of metals in surface, ground and precipitation waters after a manual acid digestion and concentration procedure and ultrasonic nebulisation. Organically-bound metals and metal complexes are converted to metal ions by digestion with nitric acid. To increase the signal detected by the ICP, the samples are concentrated 10 times. The liquid sample is pumped onto the piezoelectric transducer, where the ultrasonic action efficiently converts it to a fine, dense aerosol. The nebulizer carrier gas flow then transports the wet aerosol through a heated U-tube, where the solvent is vaporized. Subsequently, most of the solvent vapour is removed from the aerosol stream by passing the stream through a chilled condenser. The result is a dry, analyte-laden aerosol which is introduced to the plasma. There, the high temperature desolvates the aerosol and dissociates the sample material into individual atoms and ions, which are excited to emit light at wave lengths characteristic of the elements present. Dissolved samples are field filtered through 0.45 micron cellulose acetate filter prior to acid preservation. Samples for extractable analysis are acidified with 2 mL concentrated HNO3 per litre and allowed to stand a minimum of 24 hours before analysis. Clients providing these samples are advised to preserve samples at the collection site. | 0.05 | Y | Y | Y | 82 | 3/10/1998 2:07:00 PM | 2/25/1999 2:08:00 PM | 10/20/1999 2:21:47 AM | 1998-11-17 |  | NPEPPIN |  | Concentration | mg/L | Y | (fl. conc.) |  | Silicon Dissolved (fl. conc.) | 7440-21-3 | Metals - Dissolved; | Metals - Dissolved; EMS code: SI-D | Chemical | Dissolved | LAB
SI-D | mg/L | Silicon Dissolved | X511 | FF,FA,Dir.Asp.,ICP | FF;FA;Direct Aspiration;ICP | 0.008 | Y | Y | Y | 6182 |  |  | 5/8/2000 4:03:53 PM | 1999-04-22 |  | NPEPPIN |  | Concentration | mg/L | Y | (fl. conc.) |  | Silicon Dissolved (fl. conc.) | 7440-21-3 | Metals - Dissolved; | Metals - Dissolved; EMS code: SI-D | Chemical | Dissolved | LAB
SI-D | mg/L | Silicon Dissolved | F039 | Field filt + acid, ICP-AES | Dissolved and extractable water samples are ready for analysis once collected. Total samples are autoclaved with HNO3 and hydrochloric acid (HCI) in capped Teflon vials. Analysis is performed via Inductively Coupled Argon Plasma Atomic Emission Spectrometry (ICP-AES)(simultaneous multielement analysis), Graphite Furnace Atomic Absorption Spectrometry (GF-AAS)(low level analysis) and/or Flame Atomic Absorption Spectrometry (F-AAS). | 0.05 | Y | Y | Y | 6380 | 7/29/1996 6:03:00 PM | 7/5/2019 1:21:00 PM | 5/16/2022 4:21:33 PM | 1997-05-08 | 6/3/1998 6:51:46 PM | MPAK |  | Concentration | mg/L | Y | (fl. conc.) |  | Silicon Dissolved (fl. conc.) | 7440-21-3 | Metals - Dissolved; | Metals - Dissolved; EMS code: SI-D | Chemical | Dissolved | LAB
SI-D | mg/L | Silicon Dissolved | X170 | ICP | This code is a catch-all for a large variety of ICP methods including water and also solids digested into solutions as an intermediate result. Different procedures will have greatly different detection limits. (a) For air particulates, preliminary results are reported in mg/L and loading results as mg/dm2/d. (b) For Hi Vol samples preliminiary results are same as reported in mg/L and loading results as ug/m3. (c) Note that for Rain samples there is no filtration nor digestion steps. | 0.03 | Y | Y | Y | 16798 | 4/11/1996 6:02:00 PM | 11/6/2018 | 8/26/2024 2:41:32 PM | 1996-04-12 | 6/3/1998 6:51:38 PM | EMS |  | Concentration | mg/L | Y | (fl. conc.) |  | Silicon Dissolved (fl. conc.) | 7440-21-3 | Metals - Dissolved; | Metals - Dissolved; EMS code: SI-D | Chemical | Dissolved | LAB
SI-D | mg/L | Silicon Dissolved | F032 | Filt.,Acid,ICPMS | There is no acid digestion required for this analysis. A water sample is collected and preserved in 0.2 percent nitric acid. To analyse the extractable metals the sample is shaken, allowed to sit for at least twenty four hours, and a decanted portion is introduced to the spectrometer. When the dissolved metals are required, the samples must first be filtered through a 0.45 micron cellulose acetate filter paper and then followed by preservation in 0.2 percent nitric acid. It is important to note that different laboratories may have very different detection limits for some variables analyzed by ICPMS. | 0.05 | Y | Y | Y | 58007 | 5/7/1998 | 8/26/2024 | 8/26/2024 4:53:05 PM | 1998-05-14 | 6/3/1998 6:51:49 PM | NPEPPIN |  | Concentration | mg/L | Y | (fl. conc.) |  | Silicon Dissolved (fl. conc.) | 7440-21-3 | Metals - Dissolved; | Metals - Dissolved; EMS code: SI-D | Chemical | Dissolved | LAB
SI-D | mg/L | Silicon Dissolved | X285 | HNO3 Dig:ICP Analysis | HNO3 Dig:ICP Analysis | 0.05 | Y | Y | Y | 651 | 4/16/2001 | 11/13/2002 | 2/19/2004 4:22:43 PM | 1998-05-22 | 6/3/1998 6:51:49 PM | NPEPPIN |  | Concentration | mg/L | Y | (fl. conc.) |  | Silicon Dissolved (fl. conc.) | 7440-21-3 | Metals - Dissolved; | Metals - Dissolved; EMS code: SI-D | Chemical | Dissolved | LAB
SI-D | mg/L | Silicon Dissolved | METC | MET-D-NP-CCMS-VA | Ultra trace metals in water are analyzed by CRC ICPMS, based on US EPA Method 6020B (July 2014). This procedure is intended for pristine field-filtered acid-preserved water samples. The detection limits (LOR) for this test are based on lab instrumental analysis only, not including filtration. ALS-supplied field filtration equipment does not support these LOR. Therefore, because of the high probability of false positives due to filtration, it is strongly recommended that a filtration blank be analysed to aid in data interpretation. | 0.05 | Y | Y | Y | 1156 | 11/13/2017 | 8/15/2024 | 8/23/2024 9:50:35 AM | 2019-08-20 |  | KATFRASE |  | Concentration | mg/L | Y | (fl. conc.) |  | Silicon Dissolved (fl. conc.) | 7440-21-3 | Metals - Dissolved; | Metals - Dissolved; EMS code: SI-D | Chemical | Dissolved | LAB
SI-D | mg/L | Silicon Dissolved | X562 | Dissolved metals: FF, FA, HNO3, HR-ICPMS | Ultra trace metals in water are analyzed by high resolution inductively coupled plasma mass spectrometry (HR-ICPMS), based on US EPA Method 200.8, (Oct. 1999). This procedure is intended for pristine field-filtered acid-preserved water samples. The detection limits (LOR) for this test are based on lab instrumental analysis only, not including filtration. ALS-supplied field filtration equipment does not support these LOR. Therefore, because of the high probability of false positives due to filtration, it is strongly recommended that a filtration blank be analyzed to aid in data interpretation. | 0.05 | Y | Y | Y | 1259 | 6/8/2015 | 10/19/2023 | 11/9/2023 7:41:27 AM | 2015-10-22 |  | RDGIBSON |  | Concentration | mg/L | Y | (fl. conc.) |  | Silicon Dissolved (fl. conc.) | 7440-21-3 | Metals - Dissolved; | Metals - Dissolved; EMS code: SI-D | Chemical | Dissolved | LAB
SI-E | mg/L | Silicon Extractable | F040 | Field Acid, lab filt, ICP-AES | Dissolved and extractable water samples are ready for analysis once collected. Total samples are autoclaved with HNO3 and hydrochloric acid (HCI) in capped Teflon vials. Analysis is performed via Inductively Coupled Argon Plasma Atomic Emission Spectrometry (ICP-AES)(simultaneous multielement analysis), Graphite Furnace Atomic Absorption Spectrometry (GF-AAS)(low level analysis) and/or Flame Atomic Absorption Spectrometry (F-AAS). | 0.05 | Y | Y | Y | 1055 | 8/5/1997 6:06:00 PM | 2/7/2002 | 5/26/2003 3:19:22 PM | 1997-05-08 | 6/3/1998 6:51:46 PM | MPAK |  | Concentration | mg/L | Y | (fl. conc.) |  | Silicon Extractable (fl. conc.) | 7440-21-3 | Metals - Extractable; | Metals - Extractable; EMS code: SI-E | Chemical | Extractable | LAB
SI-E | mg/L | Silicon Extractable | F055 | No Pres/No Filtr/No Dig,ICP | Shortcut method which excludes preservation, filtration and digestion steps. Results are used for estimation of water hardness. | 0.05 | Y | Y | Y | 3662 | 8/23/1996 6:04:00 PM | 3/14/2008 | 9/11/2019 8:47:41 AM | 1998-04-21 | 6/3/1998 6:51:48 PM | NPEPPIN |  | Concentration | mg/L | Y | (fl. conc.) |  | Silicon Extractable (fl. conc.) | 7440-21-3 | Metals - Extractable; | Metals - Extractable; EMS code: SI-E | Chemical | Extractable | LAB
SI-E | mg/L | Silicon Extractable | F069 | SEM extr,6N HCl,ICP/OES | 6.0 N Hydrochloric acid (HCI) is added to an aliquot of the sediment sample. The extract produced from the addition of the acid is then analyzed for simultaneously extracted metals (SEM) using inductively coupled plasma - optical emission spectrophotometry. The detection limits depend on the metal parameters in question and the instrumentation used. | 0.05 | Y | Y | Y | 9 | 3/23/2012 3:11:00 AM | 3/23/2012 3:11:00 AM | 3/28/2012 10:45:30 AM | 1999-07-09 |  | NPEPPIN |  | Concentration | mg/L | Y | (fl. conc.) |  | Silicon Extractable (fl. conc.) | 7440-21-3 | Metals - Extractable; | Metals - Extractable; EMS code: SI-E | Chemical | Extractable | LAB
SI-L | mg/L | Silicon Leachable | MEND | MEND (mod) EPA 300.1 (mod) | MEND (mod) EPA 300.1 (mod). Inorganic anions are analyzed by obtaining an extract from a soil sample that has been added in a 3:1 ratio of deionized water to soil; then leached and the extract filtered prior to analysis; which is then analyzed by Ion Chromatography with conductivity and/or UV detection. | 0.05 | Y | Y | Y | 0 |  |  |  | 2024-01-29 |  | RDGIBSON |  | Concentration | mg/L | Y | (fl. conc.) |  | Silicon Leachable (fl. conc.) |  |  | EMS code: SI-L | Chemical |  | LAB
SI-O2 | % | Silicon (as SiO2) | ICP06 | ALS Minerals ME-ICP06 | ALS Minerals ME-ICP06. A prepared sample is fused with a lithium tetraborate/lithium metaborate flux and poured into a graphite crucible and fused in a furnace at 1000°C. The resulting melt is then cooled and dissolved in 100 mL of 4% HNO3/2% HCl followed by ICP-AES. | 0.01 | Y | Y | Y | 0 |  |  |  | 2024-01-29 |  | EMS |  | DimensionlessRatio | % | N | (ratio) |  | Silicon (as SiO2) (ratio) |  |  | EMS code: SI-O2 | Chemical |  | LAB
SI-T | mg/L | Silicon Total | F033 | Acid Dig,Centrifuge,ICPMS | For the determination of total recoverable elements in water or waste water, take 100 ml aliquot from a well mixed, acid preserved, sample containing not more than 0.25% (w/v) total solids and transfer to a 250 ml Griffin beaker. (If total solids are greater than 0.25% reduce the size of the aliquot by a proportionate amount). Add 1 ml of conc. nitric acid and 0.5 ml conc. hydrochloric acid. Heat on a hot plate until the volume has been reduced to 15-20ml, ensuring that the sample does not boil. Allow to cool and transfer to a 50 ml volumetric flask. Dilute to volume with ASTM type I water. Centrifuge the sample or allow to stand overnight to separate insoluble material. Prior analysis, pipette 20 ml into a 50 ml volumetric flask and dilute to volume with ASTM type I water. Analysis should be performed within seventy two hours of the completed preparation. Note that different laboratories may have very different field preservation, digestion procedures and detection limits for this method, so these should be documented. | 0.05 | Y | Y | Y | 26010 | 5/7/1998 | 8/22/2024 3:24:00 AM | 8/25/2024 9:47:36 PM | 1998-05-14 | 6/3/1998 6:51:49 PM | NPEPPIN |  | Concentration | mg/L | Y | (fl. conc.) |  | Silicon Total (fl. conc.) | 7440-21-3 | Metals - Total; | Metals - Total; EMS code: SI-T | Chemical | Total | LAB
SI-T | mg/L | Silicon Total | METN | MET-T-NP-CCMS-VA | Ultra trace metals in water are analyzed by CRC ICPMS, based on US EPA Method 6020B (July 2014). The detection limits provided can only be met for undigested samples. This procedure is intended for colorless, non-turbid, acid-preserved water samples (i.e. pristine water samples), having turbidity < 1 NTU and no odor. Where turbidity exceeds 1 NTU, and/or the sample is colored and has an odor, results may be biased low compared to true Total Metals concentrations. ALS recommends that turbidity analysis be requested on samples submitted for this test to aid with interpretation of results. Where turbidity is <1NTU, undigested metals are equivalent to total metals concentrations. | 0.05 | Y | Y | Y | 3028 | 11/14/2017 | 8/19/2024 | 8/26/2024 1:03:06 PM | 2019-08-20 |  | KATFRASE |  | Concentration | mg/L | Y | (fl. conc.) |  | Silicon Total (fl. conc.) | 7440-21-3 | Metals - Total; | Metals - Total; EMS code: SI-T | Chemical | Total | LAB
SI-T | ug/g | Silicon Total | FS01 | D&G,Nitric/HCl:ICP | Samples are dried at 60 C, then either sieved to -100mesh (150 um) and rolled to homogenize, or finely ground with a mortar and pestle or a shatter box laboratory mill and rolled (depending on the submitter's requirements). The samples are then weighed into Teflon digestion liners and digested with HNO3 and HCl in a microwave oven. Samples are cooled, volumized, and allowed to settle overnight before being decanted. Analysis is then performed via Inductively Coupled Argon Plasma Atomic Emission Spectrometry (ICP-AES) (Simultaneous multi element analysis). In ICP Spectrometry, the liquid sample is pumped to the nebulizer, where it is convertd into an aerosol. the nebulizer carrier gas flow than transports the wet aerosol past the impact bead, where the large droplets condense out and flow to the waste container. The remaining fine aerosol is carried into the plasma. There, the high temperature desolvates the aerosol and dissociates the sample material into individual atoms and ions, which are excited to emit light at wavelengths characteristic of the elements present. A diffraction grating in the spectrometer diffracts the various wavelengths of light through exit slits to the photo-multiplier tubes, which convert the light energy to electronic signals. | 8 | Y | Y | Y | 1907 | 8/30/1996 1:52:00 PM | 2/5/2002 | 5/28/2003 2:34:25 PM | 1996-06-19 | 6/3/1998 6:51:44 PM | MPAK |  | MassPerMass | mg/kg | Y | (mass conc.) |  | Silicon Total (mass conc.) | 7440-21-3 | Metals - Total; | Metals - Total; EMS code: SI-T | Chemical | Total | LAB
SI-T | mg/L | Silicon Total | X349 | FA:HNO3:ICP | Samples are field-acidified with HNO3; HNO3 digestion; analysis by ICP/AES. | 0.06 | Y | Y | Y | 30 |  |  | 10/30/2014 8:31:18 AM | 1998-05-22 | 6/3/1998 6:51:49 PM | NPEPPIN |  | Concentration | mg/L | Y | (fl. conc.) |  | Silicon Total (fl. conc.) | 7440-21-3 | Metals - Total; | Metals - Total; EMS code: SI-T | Chemical | Total | LAB
SI-T | mg/L | Silicon Total | F048 | Acid dig,conc'n,nebul.,ICP-OES | This method provides an automated procedure for the Inductively Coupled Argon Plasma Optical Emission Spectrophotometric (ICP-OES) determination and quantitation of trace amounts of metals in surface, ground and precipitation waters after a manual acid digestion and concentration procedure and ultrasonic nebulisation. Organically-bound metals and metal complexes are converted to metal ions by digestion with nitric acid [method details says HCl]. To increase the signal detected by the ICP, the samples are concentrated 10 times. The liquid sample is pumped onto the piezoelectric transducer, where the ultrasonic action efficiently converts it to a fine, dense aerosol. The nebulizer carrier gas flow then transports the wet aerosol through a heated U-tube, where the solvent is vaporized. Subsequently, most of the solvent vapour is removed from the aerosol stream by passing the stream through a chilled condenser. The result is a dry, analyte-laden aerosol which is introduced to the plasma. There, the high temperature desolvates the aerosol and dissociates the sample material into individual atoms and ions, which are excited to emit light at wave lengths characteristic of the elements present. Dissolved samples are field filtered through 0.45 micron cellulose acetate filter prior to acid preservation. Samples for extractable analysis are acidified with 2 mL concentrated HNO3 per litre and allowed to stand a minimum of 24 hours before analysis. Clients providing these samples are advised to preserve samples at the collection site. | 0.05 | Y | Y | Y | 231 | 11/20/1997 9:32:00 AM | 7/17/2012 7:03:00 AM | 7/26/2012 12:28:00 PM | 1998-11-13 |  | NPEPPIN |  | Concentration | mg/L | Y | (fl. conc.) |  | Silicon Total (fl. conc.) | 7440-21-3 | Metals - Total; | Metals - Total; EMS code: SI-T | Chemical | Total | LAB
SI-T | mg/L | Silicon Total | MSEA | Metals in Seawater (undigested) by (CRC) ICP-MS | This method describes the multi-elemental determination of trace elements by Collision/Reaction Cell (CRC) ICP-MS. The method measures ions produced by a radio-frequency inductively coupled plasma. Analyte species originating in a liquid are pneumatically nebulized and the resulting aerosol transported by argon gas into the plasma torch. The ions produced are entrained in the plasma gas and extracted, by means of a differentially pumped vacuum interface, into a mass spectrometer. The ions flow through a Collision/Reaction Cell (CRC) where depending on the mode of analysis either the ions passes through the cell unhindered (no gas - classical ICP-MS mode), or collisions/reactions chemistry takes place between the ions and a neutral gas such as hydrogen or helium. Any polyatomic ions which give interference on the analytes are removed. The ions are then sorted according to their mass-to-charge ratios by a quadrupole mass spectrometer. The ions are then sorted according to their mass-to-charge ratios by a quadrupole mass spectrometer. The ions transmitted through the quadrupole are quantified by a channel electron multiplier and the ion information processed by a data handling system. Although most polyatomic interference is removed, any additional interference must be assessed and valid corrections applied or the data flagged to indicate problems. Interference correction must include compensation for background ions contributed by the plasma gas, reagents, and constituents of the sample matrix. Instrument drift as well as suppressions or enhancements of instrument response caused by the sample matrix must be corrected for by the use of internal standards. | 0.003 | Y | Y | Y | 421 | 8/12/2009 12:08:00 AM | 11/30/2023 12:44:00 AM | 12/1/2023 9:45:27 AM | 2009-09-08 |  | RDGIBSON |  | Concentration | mg/L | Y | (fl. conc.) |  | Silicon Total (fl. conc.) | 7440-21-3 | Metals - Total; | Metals - Total; EMS code: SI-T | Chemical | Total | LAB
SI-T | mg/L | Silicon Total | F082 | Aliquot,HNO3/HCL dig,ICPMS | Aliquot microwave digested with 3:1 HNO3/HCL. Analysis by ICPMS. For HiVol samples preliminary result are reported in mg/L and loading results are reported in ug/m3 | 0.05 | Y | Y | Y | 23576 | 1/24/2011 | 8/23/2024 | 8/26/2024 4:53:05 PM | 2015-04-15 |  | RDGIBSON |  | Concentration | mg/L | Y | (fl. conc.) |  | Silicon Total (fl. conc.) | 7440-21-3 | Metals - Total; | Metals - Total; EMS code: SI-T | Chemical | Total | LAB
SI-T | ug/m3 | Silicon Total | 5312 | Tef:N/P Dig:ICP Analysis | A nitric/perchloric acid digestion if used to bring the metals into solution prior to ICP analysis. For Hi-Vol samples preliminary results are reported as mg/L and 24 hour loading results as ug/m3. | 0.07 | Y | Y | Y | 1185 | 11/7/1997 | 3/29/2000 | 9/28/2000 1:34:45 PM | 1996-10-13 | 6/3/1998 6:51:46 PM | EMS |  | AirConcentration | µg/m³ | Y | (air conc.) |  | Silicon Total (air conc.) | 7440-21-3 | Metals - Total; | Metals - Total; EMS code: SI-T | Chemical | Total | LAB
SI-T | ug/g | Silicon Total | FS82 | PBM,dry60C,10mesh,SALM-ICP-MS | PBM;dry60C;10mesh;SALM-ICP-MS | 500 | Y | Y | Y | 302 | 12/7/2009 3:52:00 PM | 2/10/2017 5:29:00 PM | 2/17/2017 5:01:07 PM | 2010-03-29 |  | EMS |  | MassPerMass | mg/kg | Y | (mass conc.) |  | Silicon Total (mass conc.) | 7440-21-3 | Metals - Total; | Metals - Total; EMS code: SI-T | Chemical | Total | LAB
SI-T | mg/L | Silicon Total | X145 | Dig, ICP | Dig; ICP | 0.05 | Y | Y | Y | 406 | 5/14/2013 | 9/22/2016 | 6/13/2024 3:58:21 PM | 2015-03-11 |  | RDGIBSON |  | Concentration | mg/L | Y | (fl. conc.) |  | Silicon Total (fl. conc.) | 7440-21-3 | Metals - Total; | Metals - Total; EMS code: SI-T | Chemical | Total | LAB
SI-T | ug/g wet | Silicon Total | X560 | Metals in Biota (wet), total, acid digestion, ICP-MS | Five grams of wet, homogenized biota is weighed in a centrifuge tube, and 2.5 mL of Nitric Acid and 2.5 mL of Hydrochloric acid are added. The sample is digested at room temperature for one hour and then digested at 90-95 degrees C in a water bath or heat block for one more hour. The sample is cooled and the volume is made to 50 mL with DI. The digestate is then diluted at ten fold and introduced to a collision/reaction cell inductively coupled argon plasma mass spectrometer (CRC ICPMS) by pneumatic nebulization. Each respective element is measured at a specific mass to charge ratio, m/z value expressed in amu (atomic mass units). - Maxxam Analytics (BC). Added to EMS Dec 10, 2008. | 2 | Y | Y | Y | 322 | 10/19/2009 | 6/19/2014 4:57:00 PM | 6/30/2014 12:56:35 PM | 2010-02-01 | 3/11/2015 4:17:10 PM | RDGIBSON | EMS | MassPerMass | µg/g | Y | (mass conc.) | (wet) | Silicon Total (wet) (mass conc.) | 7440-21-3 | Metals - Total; | Metals - Total; EMS code: SI-T | Chemical | Total | LAB
SI-T | mg/L | Silicon Total | 0042 | GI HNO3 Dig: ICP | Sample aliquot taken from unfiltered unpreserved General Ions bottle; HNO3 digestion; ICP analysis. | 0.8 | Y | Y | Y | 3961 | 9/20/1996 | 7/5/2019 3:21:00 PM | 7/12/2019 6:49:13 PM | 1996-04-12 | 6/3/1998 6:51:38 PM | EMS |  | Concentration | mg/L | Y | (fl. conc.) |  | Silicon Total (fl. conc.) | 7440-21-3 | Metals - Total; | Metals - Total; EMS code: SI-T | Chemical | Total | LAB
SI-T | mg/L | Silicon Total | X170 | ICP | This code is a catch-all for a large variety of ICP methods including water and also solids digested into solutions as an intermediate result. Different procedures will have greatly different detection limits. (a) For air particulates, preliminary results are reported in mg/L and loading results as mg/dm2/d. (b) For Hi Vol samples preliminiary results are same as reported in mg/L and loading results as ug/m3. (c) Note that for Rain samples there is no filtration nor digestion steps. | 0.03 | Y | Y | Y | 13222 | 3/28/1996 6:03:00 PM | 11/7/2018 | 8/26/2024 2:41:32 PM | 1996-04-12 | 6/3/1998 6:51:38 PM | EMS |  | Concentration | mg/L | Y | (fl. conc.) |  | Silicon Total (fl. conc.) | 7440-21-3 | Metals - Total; | Metals - Total; EMS code: SI-T | Chemical | Total | LAB
SI-T | mg/L | Silicon Total | X561 | Total metals: FA, HNO3, HR-ICPMS (No Prep)Prep) | Ultra trace metals in water are analyzed by high resolution inductively coupled plasma mass spectrometry (HR-ICPMS), modified from US EPA 200.8, (Revision 5.5). The detection limits provided can only be met for undigested samples. This procedure is intended for colorless, non-turbid, acid-preserved water samples (i.e. pristine water samples), having turbidity < 1 NTU and no odor. Where turbidity exceeds 1 NTU, and/or the sample is colored and has an odor, results may be biased low compared to true Total Metals concentrations. ALS recommends that turbidity analysis be requested on samples submitted for this test to aid with interpretation of results. | 0.05 | Y | Y | Y | 2745 | 6/8/2015 | 10/5/2023 | 1/31/2024 2:37:41 PM | 2015-10-22 |  | RDGIBSON |  | Concentration | mg/L | Y | (fl. conc.) |  | Silicon Total (fl. conc.) | 7440-21-3 | Metals - Total; | Metals - Total; EMS code: SI-T | Chemical | Total | LAB
SI-T | ug/g | Silicon Total | FB04 | Freeze Dry,Hom,Acid Dig,ICP | Sample moisture is determined by weighing samples before and after freeze-drying. Freeze-dried samples are ground in a Braun blender before being weighed into Teflon PFA liners and digested with concentrated HNO3, concentrated HCl, and 50% H2O2 in a microwave oven; which results in a complete sample dissolution. Samples are collected and volumized with Type I water. Analysis is performed via Inductively Coupled Argon Plasma Atomic Emission Spectrometry (ICP-AES) (simultaneous multielement analysis), Graphite Furnace Atomic Absorption Spectrometry (GF-AAS) (low level Cd and Pb analysis), Cold Vapor Atomic Fluorescence Spectrometry (CV-AFS) (Mercury analysis) and/or Flame Atomic Absorption Spectrometry (F-AAS) (K analysis). | 4 | Y | Y | Y | 241 | 8/22/1996 6:11:00 PM | 2/1/2002 2:30:00 PM | 6/23/2003 1:15:34 PM | 1996-06-19 | 6/3/1998 6:51:44 PM | MPAK |  | MassPerMass | mg/kg | Y | (mass conc.) |  | Silicon Total (mass conc.) | 7440-21-3 | Metals - Total; | Metals - Total; EMS code: SI-T | Chemical | Total | LAB
SI-T | ug/g | Silicon Total | FS89 | Homo,HNO3 dig,ICP-AES | Homo;HNO3 dig;ICP-AES | 8 | Y | Y | Y | 4 | 1/5/2001 | 8/20/2001 12:00:00 PM | 10/23/2002 1:42:51 PM | 2001-03-08 |  | NPEPPIN |  | MassPerMass | mg/kg | Y | (mass conc.) |  | Silicon Total (mass conc.) | 7440-21-3 | Metals - Total; | Metals - Total; EMS code: SI-T | Chemical | Total | LAB
SI-T | ug/g | Silicon Total | EQ02 | EQUIS (ug/g) | EQUIS (ug/g) | 1 | N | Y | Y | 174 |  |  | 10/23/1996 12:08:28 AM | 1996-04-12 | 6/3/1998 6:51:38 PM | EMS |  | MassPerMass | mg/kg | Y | (mass conc.) |  | Silicon Total (mass conc.) | 7440-21-3 | Metals - Total; | Metals - Total; EMS code: SI-T | Chemical | Total | LAB
SI-T | mg/L | Silicon Total | F121 | Acid,ICPMS | Acidified, no digestion, ICPMS | 0.1 | Y | Y | Y | 4854 | 3/12/2010 7:57:00 PM | 8/15/2024 9:10:00 PM | 8/26/2024 8:06:03 AM | 2002-07-22 | 7/16/2012 3:30:25 PM | UMULLER | RDGIBSON | Concentration | mg/L | Y | (fl. conc.) |  | Silicon Total (fl. conc.) | 7440-21-3 | Metals - Total; | Metals - Total; EMS code: SI-T | Chemical | Total | LAB
SI-T | ug/g | Silicon Total | 0190 | D&G:N/P Dig:ICP Anal. | Sample is dried and grind-mixed before nitric perchloric digestion followed by ICP analysis. | 1 | Y | Y | Y | 3 |  |  | 10/24/1996 2:06:02 PM | 1996-04-12 | 6/3/1998 6:51:38 PM | EMS |  | MassPerMass | mg/kg | Y | (mass conc.) |  | Silicon Total (mass conc.) | 7440-21-3 | Metals - Total; | Metals - Total; EMS code: SI-T | Chemical | Total | LAB
SI-T | mg/dm2/d | Silicon Total | DUST | HNO3 dig.(microwave or hot block),ICPMS | HNO3 dig.(microwave or hot block);ICPMS |  | Y | Y | Y | 121 | 2/4/2020 | 7/31/2024 | 7/31/2024 11:23:12 AM | 2022-09-01 |  | EMS |  | MassFlux | mg/dm²/d | Y | (mass flux) |  | Silicon Total (mass flux) | 7440-21-3 | Metals - Total; | Metals - Total; EMS code: SI-T | Chemical | Total | LAB
SI-T | mg/L | Silicon Total | X285 | HNO3 Dig:ICP Analysis | HNO3 Dig:ICP Analysis | 0.05 | Y | Y | Y | 539 | 8/10/1993 | 10/5/1993 | 3/20/2000 2:06:59 PM | 1998-05-22 | 6/3/1998 6:51:49 PM | NPEPPIN |  | Concentration | mg/L | Y | (fl. conc.) |  | Silicon Total (fl. conc.) | 7440-21-3 | Metals - Total; | Metals - Total; EMS code: SI-T | Chemical | Total | LAB
SI-T | mg/L | Silicon Total | MWID | Metals in Water by ICPMS (digested) | Metals in Water by ICPMS (digested) Based on EPA 200.8 | 2 | Y | Y | Y | 826 |  |  | 7/31/2024 1:46:41 PM | 2009-12-24 |  | RDGIBSON |  | Concentration | mg/L | Y | (fl. conc.) |  | Silicon Total (fl. conc.) | 7440-21-3 | Metals - Total; | Metals - Total; EMS code: SI-T | Chemical | Total | LAB
SI-T | mg/L | Silicon Total | 1349 | Tot Metals in Seawater by CRC ICPMS (BC) | Tot Metals in Seawater by CRC ICPMS (BC) | 1 | Y | Y | Y | 18 | 4/13/2023 | 4/29/2024 | 6/7/2024 10:03:16 AM | 2023-05-26 |  | EMS |  | Concentration | mg/L | Y | (fl. conc.) |  | Silicon Total (fl. conc.) | 7440-21-3 | Metals - Total; | Metals - Total; EMS code: SI-T | Chemical | Total | LAB
SI-T | mg/L | Silicon Total | 0040 | HNO3 Dig: ICP Analysis | HNO3 Dig: ICP Analysis | 0.03 | Y | Y | Y | 1722 | 6/25/2010 3:15:00 PM | 6/25/2010 3:15:00 PM | 7/22/2010 2:42:37 PM | 1996-04-12 | 6/3/1998 6:51:41 PM | EMS |  | Concentration | mg/L | Y | (fl. conc.) |  | Silicon Total (fl. conc.) | 7440-21-3 | Metals - Total; | Metals - Total; EMS code: SI-T | Chemical | Total | LAB
SI-T | mg/L | Silicon Total | F038 | HNO3/HCl dig,ICP-AES | Dissolved and extractable water samples are ready for analysis once collected. Total samples are autoclaved with HNO3 and hydrochloric acid (HCI) in capped Teflon vials. Analysis is performed via Inductively Coupled Argon Plasma Atomic Emission Spectrometry (ICP-AES)(simultaneous multielement analysis), Graphite Furnace Atomic Absorption Spectrometry (GF-AAS) (low level analysis) and/or Flame Atomic Absorption Spectrometry (F-AAS). | 0.06 | Y | Y | Y | 7149 | 8/1/1996 6:04:00 PM | 7/25/2012 7:11:00 AM | 9/11/2019 8:47:43 AM | 1997-05-08 | 6/3/1998 6:51:46 PM | MPAK |  | Concentration | mg/L | Y | (fl. conc.) |  | Silicon Total (fl. conc.) | 7440-21-3 | Metals - Total; | Metals - Total; EMS code: SI-T | Chemical | Total | LAB
SI-T | ug/g | Silicon Total | PB02 | PBM,HNO3/HCl dig,dry60C,100mesh, | Performance Based Method. Samples are dried at 60 C and sieved using a .149 mm (100 mesh) sieve. Solid samples are digested with the combination of nitric and hydrochloric acids with a 1:1 ratio of acids. For environmental purposes, strong acid leaches are used routinely since they have the advantage of speed, low cost, and superio reproducibility. All metals trapped within the refractory matrix are unlikely to have environmental consequences. Instrumental analysis of extracts produced using this digestion method can be performed by a variety of techniques. |  | Y | Y | Y | 10 | 9/13/2001 | 9/13/2001 | 11/6/2002 3:52:40 PM | 2001-07-18 |  | NPEPPIN |  | MassPerMass | mg/kg | Y | (mass conc.) |  | Silicon Total (mass conc.) | 7440-21-3 | Metals - Total; | Metals - Total; EMS code: SI-T | Chemical | Total | LAB
SIF4 | ug/m3 | Silicon Tetra-F | X194 | Specific Ion Electr- Stack | Specific Ion Electr- Stack | 0 | Y | Y | Y | 102 |  |  | 10/24/1996 11:05:14 AM | 1996-04-12 | 6/3/1998 6:51:41 PM | EMS |  | AirConcentration | µg/m³ | Y | (air conc.) |  | Silicon Tetra-F (air conc.) | 7783-61-1 | Volatile Organic Compounds (VOCs); | Volatile Organic Compounds (VOCs); EMS code: SIF4 | Chemical |  | LAB
SIF4 | ug/m3 | Silicon Tetra-F | EQ40 | EQUIS (mg/m3) | EQUIS (mg/m3) | 0 | N | Y | Y | 121 |  |  | 10/22/1996 11:53:54 PM | 1996-04-12 | 6/3/1998 6:51:38 PM | EMS |  | AirConcentration | µg/m³ | Y | (air conc.) |  | Silicon Tetra-F (air conc.) | 7783-61-1 | Volatile Organic Compounds (VOCs); | Volatile Organic Compounds (VOCs); EMS code: SIF4 | Chemical |  | LAB
SILT | % | % silt by hydrometer | HYDR | Hydrometer Analysis | Hydrometer Analysis | 2 | Y | Y | Y | 144 | 10/4/2013 3:54:00 PM | 11/1/2022 | 1/11/2024 1:07:08 PM | 2010-11-16 |  | RDGIBSON |  | DimensionlessRatio | % | N | (ratio) |  | % silt by hydrometer (ratio) |  | Sediment Size; | Sediment Size; EMS code: SILT | Physical |  | LAB
SM-T | ppm (S) | Samarium | MS81 | ALS Minerals: ME-MS81 | ALS Minerals: ME-MS81. A prepared sample (0.100 g) is added to lithium metaborate/lithium tetraborate flux; mixed well and fused in a furnace at 1025°C. The resulting melt is then cooled and dissolved in an acid mixture containing nitric; hydrochloric and hydrofluoric acids. This solution is then analyzed by inductively coupled plasma - mass spectrometry. | 0.03 | Y | Y | Y | 0 |  |  |  | 2024-01-29 |  | EMS |  | DimensionlessRatio | ppm | N | (ratio) |  | Samarium (ratio) |  |  | EMS code: SM-T | Chemical |  | LAB
SN-D | mg/L | Tin Dissolved | X562 | Dissolved metals: FF, FA, HNO3, HR-ICPMS | Ultra trace metals in water are analyzed by high resolution inductively coupled plasma mass spectrometry (HR-ICPMS), based on US EPA Method 200.8, (Oct. 1999). This procedure is intended for pristine field-filtered acid-preserved water samples. The detection limits (LOR) for this test are based on lab instrumental analysis only, not including filtration. ALS-supplied field filtration equipment does not support these LOR. Therefore, because of the high probability of false positives due to filtration, it is strongly recommended that a filtration blank be analyzed to aid in data interpretation. | 1e-05 | Y | Y | Y | 1264 | 6/8/2015 | 10/19/2023 | 11/9/2023 7:41:27 AM | 2015-10-22 |  | RDGIBSON |  | Concentration | mg/L | Y | (fl. conc.) |  | Tin Dissolved (fl. conc.) | 7440-31-5 | Metals - Dissolved; | Metals - Dissolved; EMS code: SN-D | Chemical | Dissolved | LAB
SN-D | mg/L | Tin Dissolved | METC | MET-D-NP-CCMS-VA | Ultra trace metals in water are analyzed by CRC ICPMS, based on US EPA Method 6020B (July 2014). This procedure is intended for pristine field-filtered acid-preserved water samples. The detection limits (LOR) for this test are based on lab instrumental analysis only, not including filtration. ALS-supplied field filtration equipment does not support these LOR. Therefore, because of the high probability of false positives due to filtration, it is strongly recommended that a filtration blank be analysed to aid in data interpretation. | 5e-05 | Y | Y | Y | 1776 | 4/3/2012 | 8/15/2024 | 8/23/2024 9:50:35 AM | 2019-08-20 |  | KATFRASE |  | Concentration | mg/L | Y | (fl. conc.) |  | Tin Dissolved (fl. conc.) | 7440-31-5 | Metals - Dissolved; | Metals - Dissolved; EMS code: SN-D | Chemical | Dissolved | LAB
SN-D | mg/L | Tin Dissolved | X285 | HNO3 Dig:ICP Analysis | HNO3 Dig:ICP Analysis | 0.3 | Y | Y | Y | 650 | 1/30/2002 | 11/13/2002 | 2/19/2004 4:22:43 PM | 1998-05-22 | 6/3/1998 6:51:49 PM | NPEPPIN |  | Concentration | mg/L | Y | (fl. conc.) |  | Tin Dissolved (fl. conc.) | 7440-31-5 | Metals - Dissolved; | Metals - Dissolved; EMS code: SN-D | Chemical | Dissolved | LAB
SN-D | mg/L | Tin Dissolved | F039 | Field filt + acid, ICP-AES | Dissolved and extractable water samples are ready for analysis once collected. Total samples are autoclaved with HNO3 and hydrochloric acid (HCI) in capped Teflon vials. Analysis is performed via Inductively Coupled Argon Plasma Atomic Emission Spectrometry (ICP-AES)(simultaneous multielement analysis), Graphite Furnace Atomic Absorption Spectrometry (GF-AAS)(low level analysis) and/or Flame Atomic Absorption Spectrometry (F-AAS). | 0.05 | Y | Y | Y | 5967 | 7/29/1996 6:03:00 PM | 7/5/2019 1:21:00 PM | 9/11/2019 8:47:50 AM | 1997-05-08 | 6/3/1998 6:51:46 PM | MPAK |  | Concentration | mg/L | Y | (fl. conc.) |  | Tin Dissolved (fl. conc.) | 7440-31-5 | Metals - Dissolved; | Metals - Dissolved; EMS code: SN-D | Chemical | Dissolved | LAB
SN-D | mg/L | Tin Dissolved | F032 | Filt.,Acid,ICPMS | There is no acid digestion required for this analysis. A water sample is collected and preserved in 0.2 percent nitric acid. To analyse the extractable metals the sample is shaken, allowed to sit for at least twenty four hours, and a decanted portion is introduced to the spectrometer. When the dissolved metals are required, the samples must first be filtered through a 0.45 micron cellulose acetate filter paper and then followed by preservation in 0.2 percent nitric acid. It is important to note that different laboratories may have very different detection limits for some variables analyzed by ICPMS. | 0.0005 | Y | Y | Y | 74373 | 3/10/1998 2:07:00 PM | 8/26/2024 | 8/26/2024 4:53:05 PM | 1998-05-14 | 8/15/2019 11:00:06 AM | NPEPPIN | KATFRASE | Concentration | mg/L | Y | (fl. conc.) |  | Tin Dissolved (fl. conc.) | 7440-31-5 | Metals - Dissolved; | Metals - Dissolved; EMS code: SN-D | Chemical | Dissolved | LAB
SN-D | mg/L | Tin Dissolved | MSEA | Metals in Seawater (undigested) by (CRC) ICP-MS | This method describes the multi-elemental determination of trace elements by Collision/Reaction Cell (CRC) ICP-MS. The method measures ions produced by a radio-frequency inductively coupled plasma. Analyte species originating in a liquid are pneumatically nebulized and the resulting aerosol transported by argon gas into the plasma torch. The ions produced are entrained in the plasma gas and extracted, by means of a differentially pumped vacuum interface, into a mass spectrometer. The ions flow through a Collision/Reaction Cell (CRC) where depending on the mode of analysis either the ions passes through the cell unhindered (no gas - classical ICP-MS mode), or collisions/reactions chemistry takes place between the ions and a neutral gas such as hydrogen or helium. Any polyatomic ions which give interference on the analytes are removed. The ions are then sorted according to their mass-to-charge ratios by a quadrupole mass spectrometer. The ions are then sorted according to their mass-to-charge ratios by a quadrupole mass spectrometer. The ions transmitted through the quadrupole are quantified by a channel electron multiplier and the ion information processed by a data handling system. Although most polyatomic interference is removed, any additional interference must be assessed and valid corrections applied or the data flagged to indicate problems. Interference correction must include compensation for background ions contributed by the plasma gas, reagents, and constituents of the sample matrix. Instrument drift as well as suppressions or enhancements of instrument response caused by the sample matrix must be corrected for by the use of internal standards. | 5e-05 | Y | Y | Y | 4 | 4/12/2023 | 4/29/2024 | 5/16/2024 2:00:14 PM | 2009-09-08 |  | RDGIBSON |  | Concentration | mg/L | Y | (fl. conc.) |  | Tin Dissolved (fl. conc.) | 7440-31-5 | Metals - Dissolved; | Metals - Dissolved; EMS code: SN-D | Chemical | Dissolved | LAB
SN-D | mg/L | Tin Dissolved | MWID | Metals in Water by ICPMS (digested) | Metals in Water by ICPMS (digested) Based on EPA 200.8 | 0.005 | Y | Y | Y | 31680 | 10/26/2018 5:20:00 PM | 10/26/2018 5:20:00 PM | 7/31/2024 1:46:41 PM | 2009-12-24 |  | RDGIBSON |  | Concentration | mg/L | Y | (fl. conc.) |  | Tin Dissolved (fl. conc.) | 7440-31-5 | Metals - Dissolved; | Metals - Dissolved; EMS code: SN-D | Chemical | Dissolved | LAB
SN-D | mg/L | Tin Dissolved | X170 | ICP | This code is a catch-all for a large variety of ICP methods including water and also solids digested into solutions as an intermediate result. Different procedures will have greatly different detection limits. (a) For air particulates, preliminary results are reported in mg/L and loading results as mg/dm2/d. (b) For Hi Vol samples preliminiary results are same as reported in mg/L and loading results as ug/m3. (c) Note that for Rain samples there is no filtration nor digestion steps. | 0.3 | Y | Y | Y | 7257 | 4/11/1996 6:02:00 PM | 11/6/2018 | 6/12/2020 1:21:12 PM | 1996-04-12 | 6/3/1998 6:51:38 PM | EMS |  | Concentration | mg/L | Y | (fl. conc.) |  | Tin Dissolved (fl. conc.) | 7440-31-5 | Metals - Dissolved; | Metals - Dissolved; EMS code: SN-D | Chemical | Dissolved | LAB
SN-D | mg/L | Tin Dissolved | X556 | Metals in water, dissolved, "in-bottle" digestion, ICP-SFMS | The whole water sample is filtered through a 0.45 micron cellulose acetate filter paper and the filtrate is stored in a HDPE/Teflon bottle. The container is de-capped in a Class 100 workstation, acidified to 2% (v/v) HNO3 and 1% (v/v) HCl and recapped. The sample is "in-bottle" digested in a convection oven, set at 60 degrees C for 16 hours. The sample is cooled and introduced to an inductively coupled argon plasma-sector field mass spectrometer (ICP-SFMS) by pneumatic nebulization. Each respective element is measured at a specific mass to charge ratio, m/z value expressed in amu (atomic mass units), and at a specified intrument resolution. LR: Low Resolution ~300, MR: Mid Resolution ~4000 and HR: High Resolution ~10000. Ref: Method number 02-2003, National Laboratory for Environmental Testing, CCIW, Environment Canada, Burlington, Requested April 2003. | 5e-06 | Y | Y | Y | 4158 | 6/27/2004 | 12/6/2023 2:04:00 AM | 2/14/2024 9:05:32 AM | 2008-06-12 |  | RDGIBSON |  | Concentration | mg/L | Y | (fl. conc.) |  | Tin Dissolved (fl. conc.) | 7440-31-5 | Metals - Dissolved; | Metals - Dissolved; EMS code: SN-D | Chemical | Dissolved | LAB
SN-D | mg/L | Tin Dissolved | F068 | HNO3dig,TeflonTube/Autoclave,ICP-MS | This method is applicable to determiniation of dissolved, extractable, and total recoverable metals in ground waters, surface waters, drinking waters, and wastewaters. Dissolved, extractable, and total recoverable sample fractions are acidified to pH <1.5 with analytical grade nitric acid (HNO3) to prevent the metals from adsorbing to the container material or precipitating. Extractable and total recoverable water samples are collected in 250 ml acid washed polyethylene bottles and are preserved with Nitric acid (0.5 ml acid for 100 mLs of sample). All preserved samples are stored at room temperature and can be held for up to 6 months. Unfiltered, un-preserved samples are stored at 4C for no more than 30 days. For dissolved and extractable water fractions, no sample preparation is necessary. Samples are simply poured in autosampler tubes (i.e. 15 ml centrifuge tubes), and placed on the autosampler and analyzed. Samples for total recoverable metals are digested with Nitric acid in special Teflon tubes in an autoclave prior to analyses. | 1e-05 | Y | Y | Y | 1398 | 6/1/1999 1:22:00 PM | 4/17/2015 1:06:00 PM | 8/11/2021 12:04:11 PM | 1999-06-28 |  | NPEPPIN |  | Concentration | mg/L | Y | (fl. conc.) |  | Tin Dissolved (fl. conc.) | 7440-31-5 | Metals - Dissolved; | Metals - Dissolved; EMS code: SN-D | Chemical | Dissolved | LAB
SN-D | mg/L | Tin Dissolved | X085 | LF,LA,HNO3,AA | Samples are lab-filtered (0.45 um), and lab-acidified (HNO3); analysis by AA. | 0.5 | Y | Y | Y | 8 | 9/29/2000 | 12/1/2010 | 12/9/2010 3:55:27 PM | 1996-04-12 | 6/3/1998 6:51:38 PM | EMS |  | Concentration | mg/L | Y | (fl. conc.) |  | Tin Dissolved (fl. conc.) | 7440-31-5 | Metals - Dissolved; | Metals - Dissolved; EMS code: SN-D | Chemical | Dissolved | LAB
SN-E | mg/L | Tin Extractable | F040 | Field Acid, lab filt, ICP-AES | Dissolved and extractable water samples are ready for analysis once collected. Total samples are autoclaved with HNO3 and hydrochloric acid (HCI) in capped Teflon vials. Analysis is performed via Inductively Coupled Argon Plasma Atomic Emission Spectrometry (ICP-AES)(simultaneous multielement analysis), Graphite Furnace Atomic Absorption Spectrometry (GF-AAS)(low level analysis) and/or Flame Atomic Absorption Spectrometry (F-AAS). | 0.05 | Y | Y | Y | 2350 | 8/5/1997 6:06:00 PM | 3/14/2008 | 9/11/2019 8:47:41 AM | 1997-05-08 | 6/3/1998 6:51:46 PM | MPAK |  | Concentration | mg/L | Y | (fl. conc.) |  | Tin Extractable (fl. conc.) | 7440-31-5 | Metals - Extractable; | Metals - Extractable; EMS code: SN-E | Chemical | Extractable | LAB
SN-E | mg/L | Tin Extractable | X553 | Metals in water, extractable, ICP-SFMS | The whole water sample is collected in a HDPE/Teflon bottle. In a Class 100 workstation it is preserved with concentrated Nitric Acid at concentration of 0.2 percent (v/v). The acid preservative extracts the metals from material in the sample and a decanted portion is introduced directly to an inductively coupled argon plasma-sector field mass spectrometer (ICP-SFMS) by pneumatic nebulization. Each respective element is measured at a specific mass to charge ratio, m/z value expressed in amu (atomic mass units), and at a specified instrument resolution. LR: Low Resolution ~300, MR: Mid Resolution ~4000 and HR: High Resolution ~10000. Ref: Method number 02-2061, National Laboratory for Environmental Testing, CCIW, Environment Canada, Burlington. Requested April 2003. | 0.005 | Y | Y | Y | 1715 | 6/27/2004 | 9/21/2007 | 4/18/2018 9:28:02 AM | 2009-05-14 | 5/15/2009 11:26:11 AM | RDGIBSON | RDGIBSON | Concentration | mg/L | Y | (fl. conc.) |  | Tin Extractable (fl. conc.) | 7440-31-5 | Metals - Extractable; | Metals - Extractable; EMS code: SN-E | Chemical | Extractable | LAB
SN-L | mg/L | Tin Leachable | MEND | MEND (mod) EPA 300.1 (mod) | MEND (mod) EPA 300.1 (mod). Inorganic anions are analyzed by obtaining an extract from a soil sample that has been added in a 3:1 ratio of deionized water to soil; then leached and the extract filtered prior to analysis; which is then analyzed by Ion Chromatography with conductivity and/or UV detection. | 0.0005 | Y | Y | Y | 0 |  |  |  | 2024-01-29 |  | RDGIBSON |  | Concentration | mg/L | Y | (fl. conc.) |  | Tin Leachable (fl. conc.) | 7440-31-5 | Metals - Extractable; | Metals - Extractable; EMS code: SN-L | Chemical |  | LAB
SN-T | mg/L | Tin Total | X555 | Metals in water, total, "in-bottle" digestion, ICP-SFMS | The whole water sample collected in a HDPE/Teflon bottle, is de-capped in a Class 100 workstation, acidified to 2% (v/v) HNO3 and 1% (v/v) HCl and recapped. The sample is "in-bottle" digested in a convection oven, set at 60 degrees C for 16 hours. The sample is cooled and introduced to an inductively coupled argon plasma-sector field mass spectrometer (ICP-SFMS) by pneumatic nebulization. Each respective element is measured at a specific mass to charge ratio, m/z value expressed in amu (atomic mass units), and at resolution. LR: Low Resolution ~300, MR: Mid Resolution ~4000 and HR: High a specified instrument Resolution ~10000. Ref: Method number 02-2003, National Laboratory for Environmental Testing, CCIW, Environment Canada, Burlington. Requested April 2003. | 5e-06 | Y | Y | Y | 15773 | 6/27/2004 | 12/6/2023 2:08:00 AM | 2/14/2024 9:05:32 AM | 2009-03-31 |  | JCHUDAK |  | Concentration | mg/L | Y | (fl. conc.) |  | Tin Total (fl. conc.) | 7440-31-5 | Metals - Total; | Metals - Total; EMS code: SN-T | Chemical | Total | LAB
SN-T | ug/g wet | Tin Total | X552 | Metals in Biota (wet), total, acid digestion, ICP-SFMS | One gram of wet, homogenized biota is weighed in a microwave TFM vessel, and 8mL of Nitric Acid and 1mL of Hydrogen Peroxide are added. The sample is digested at 200 degrees C for 15 minutes in a closed vessel high pressure microwave oven. The sample is cooled, the volume adjusted to 100mL and introduced to an inductively coupled argon plasma-sector field mass spectrometer (ICP-SFMS) by pneumatic nebulization. Each respective element is measured at a specific mass to charge ratio, m/z value expressed in amu (atomic mass units), and at a specified instrument resolution. LR: Low Resolution ~300, MR: Mid Resolution ~4000. Ref: 02-2704, National Laboratory for Environmental Testing, CCIW, Environment Canada, Burlington. Requested April 2003. | 0.01 | Y | Y | Y | 37 | 8/9/2016 | 10/16/2023 | 11/28/2023 2:21:33 PM | 2004-01-19 | 4/19/2021 1:22:00 PM | UMULLER | EMS | MassPerMass | µg/g | Y | (mass conc.) | (wet) | Tin Total (wet) (mass conc.) | 7440-31-5 | Metals - Total; | Metals - Total; EMS code: SN-T | Chemical | Total | LAB
SN-T | ug/m3 | Tin Total | F088 | Aliquot,HNO3/HCl dig, ICPMS,loading | Aliquot microwave digested with 3:1 HNO3/HCL. Analysis by ICPMS. For HiVol samples preliminary result are reported in mg/L and loading results are reported in ug/m3 |  | Y | Y | Y | 544 | 8/17/2000 10:46:00 AM | 10/12/2006 | 11/7/2006 1:15:13 AM | 2000-08-16 |  | NPEPPIN |  | AirConcentration | µg/m³ | Y | (air conc.) |  | Tin Total (air conc.) | 7440-31-5 | Metals - Total; | Metals - Total; EMS code: SN-T | Chemical | Total | LAB
SN-T | ug/g wet | Tin Total | X560 | Metals in Biota (wet), total, acid digestion, ICP-MS | Five grams of wet, homogenized biota is weighed in a centrifuge tube, and 2.5 mL of Nitric Acid and 2.5 mL of Hydrochloric acid are added. The sample is digested at room temperature for one hour and then digested at 90-95 degrees C in a water bath or heat block for one more hour. The sample is cooled and the volume is made to 50 mL with DI. The digestate is then diluted at ten fold and introduced to a collision/reaction cell inductively coupled argon plasma mass spectrometer (CRC ICPMS) by pneumatic nebulization. Each respective element is measured at a specific mass to charge ratio, m/z value expressed in amu (atomic mass units). - Maxxam Analytics (BC). Added to EMS Dec 10, 2008. | 0.1 | Y | Y | Y | 541 | 11/15/2008 4:00:00 AM | 10/30/2023 | 11/28/2023 2:21:33 PM | 2008-12-11 | 3/11/2015 4:17:10 PM | RDGIBSON | EMS | MassPerMass | µg/g | Y | (mass conc.) | (wet) | Tin Total (wet) (mass conc.) | 7440-31-5 | Metals - Total; | Metals - Total; EMS code: SN-T | Chemical | Total | LAB
SN-T | mg/L | Tin Total | MWID | Metals in Water by ICPMS (digested) | Metals in Water by ICPMS (digested) Based on EPA 200.8 | 0.005 | Y | Y | Y | 36810 |  |  | 7/31/2024 1:46:41 PM | 2009-12-24 |  | RDGIBSON |  | Concentration | mg/L | Y | (fl. conc.) |  | Tin Total (fl. conc.) | 7440-31-5 | Metals - Total; | Metals - Total; EMS code: SN-T | Chemical | Total | LAB
SN-T | ug/g | Tin Total | FB04 | Freeze Dry,Hom,Acid Dig,ICP | Sample moisture is determined by weighing samples before and after freeze-drying. Freeze-dried samples are ground in a Braun blender before being weighed into Teflon PFA liners and digested with concentrated HNO3, concentrated HCl, and 50% H2O2 in a microwave oven; which results in a complete sample dissolution. Samples are collected and volumized with Type I water. Analysis is performed via Inductively Coupled Argon Plasma Atomic Emission Spectrometry (ICP-AES) (simultaneous multielement analysis), Graphite Furnace Atomic Absorption Spectrometry (GF-AAS) (low level Cd and Pb analysis), Cold Vapor Atomic Fluorescence Spectrometry (CV-AFS) (Mercury analysis) and/or Flame Atomic Absorption Spectrometry (F-AAS) (K analysis). | 4 | Y | Y | Y | 241 | 8/22/1996 6:11:00 PM | 2/1/2002 2:30:00 PM | 6/23/2003 1:15:34 PM | 1996-06-19 | 6/3/1998 6:51:44 PM | MPAK |  | MassPerMass | mg/kg | Y | (mass conc.) |  | Tin Total (mass conc.) | 7440-31-5 | Metals - Total; | Metals - Total; EMS code: SN-T | Chemical | Total | LAB
SN-T | ug/g | Tin Total | 0192 | Hom, Nitric/Perchlor,ICP | Sample already in a dry state so does not require drying before homogenization and nitric perchlorid digestion followed by ICP analysis. | 5 | Y | Y | Y | 236 |  |  | 10/24/1996 11:23:42 AM | 1996-04-12 | 6/3/1998 6:51:38 PM | EMS |  | MassPerMass | mg/kg | Y | (mass conc.) |  | Tin Total (mass conc.) | 7440-31-5 | Metals - Total; | Metals - Total; EMS code: SN-T | Chemical | Total | LAB
SN-T | ug/g wet | Tin Total | MTIW | MET-WET-CCMS-MID-VA | Metals in Tissue by CRC ICPMS (WET) Methodology: This method is conducted following British Columbia Lab Manual method "Metals in Animal Tissue and Vegetation (Biota) - Prescriptive". Tissue samples are homogenized and sub-sampled prior to hotblock digestion with nitric and hydrochloric acids, in combination with addition of hydrogen peroxide. Instrumental analysis is by collision cell inductively coupled plasma - mass spectrometry (modified from EPA Method 6020A). Method Limitation: This method employs a strong acid/peroxide digestion, and is intended to provide a conservative estimate of bio-available metals. Near complete recoveries are achieved for most toxicologically important metals, but elements associated with recalcitrant minerals may be only partially recovered. | 0.02 | Y | Y | Y | 24 | 11/25/2018 | 12/6/2018 | 7/21/2021 2:55:08 PM | 2018-12-24 |  | EMS |  | MassPerMass | µg/g | Y | (mass conc.) | (wet) | Tin Total (wet) (mass conc.) | 7440-31-5 | Metals - Total; | Metals - Total; EMS code: SN-T | Chemical | Total | LAB
SN-T | ug/g | Tin Total | FS82 | PBM,dry60C,10mesh,SALM-ICP-MS | PBM;dry60C;10mesh;SALM-ICP-MS | 0.1 | Y | Y | Y | 3647 | 5/21/2002 | 7/24/2024 | 7/29/2024 5:17:09 PM | 2002-11-12 |  | UMULLER |  | MassPerMass | mg/kg | Y | (mass conc.) |  | Tin Total (mass conc.) | 7440-31-5 | Metals - Total; | Metals - Total; EMS code: SN-T | Chemical | Total | LAB
SN-T | mg/L | Tin Total | ULTR | Total Metals in Water by ICPMS (Ultra) | Total Metals in Water by ICPMS (Ultra), EPA SW-846 3005A/6020A, MET-TOT-ULTRA-MS-VA This analysis is carried out using procedures adapted from "Standard Methods for the Examination of Water and Wastewater" published by the American Public Health Association, and with procedures adapted from "Test Methods for Evaluating Solid Waste" SW-846 published by the United States Environmental Protection Agency (EPA). The procedures may involve preliminary sample treatment by acid digestion, using either hotblock or microwave oven, or filtration (EPA Method 3005A). Instrumental analysis is by inductively coupled plasma - mass spectrometry (EPA Method 6020A). | 5e-05 | Y | Y | Y | 3 | 10/4/2017 | 3/9/2019 | 3/21/2019 11:46:03 AM | 2017-10-18 |  | JKROGH |  | Concentration | mg/L | Y | (fl. conc.) |  | Tin Total (fl. conc.) | 7440-31-5 | Metals - Total; | Metals - Total; EMS code: SN-T | Chemical | Total | LAB
SN-T | mg/L | Tin Total | F033 | Acid Dig,Centrifuge,ICPMS | For the determination of total recoverable elements in water or waste water, take 100 ml aliquot from a well mixed, acid preserved, sample containing not more than 0.25% (w/v) total solids and transfer to a 250 ml Griffin beaker. (If total solids are greater than 0.25% reduce the size of the aliquot by a proportionate amount). Add 1 ml of conc. nitric acid and 0.5 ml conc. hydrochloric acid. Heat on a hot plate until the volume has been reduced to 15-20ml, ensuring that the sample does not boil. Allow to cool and transfer to a 50 ml volumetric flask. Dilute to volume with ASTM type I water. Centrifuge the sample or allow to stand overnight to separate insoluble material. Prior analysis, pipette 20 ml into a 50 ml volumetric flask and dilute to volume with ASTM type I water. Analysis should be performed within seventy two hours of the completed preparation. Note that different laboratories may have very different field preservation, digestion procedures and detection limits for this method, so these should be documented. | 0.001 | Y | Y | Y | 36946 | 11/20/1997 9:32:00 AM | 8/22/2024 3:24:00 AM | 8/25/2024 9:47:36 PM | 1998-05-14 | 6/3/1998 6:51:49 PM | NPEPPIN |  | Concentration | mg/L | Y | (fl. conc.) |  | Tin Total (fl. conc.) | 7440-31-5 | Metals - Total; | Metals - Total; EMS code: SN-T | Chemical | Total | LAB
SN-T | ppm (S) | Tin Total | 6041 | Aqua Regia by ICPMS | These were sediment/soils digested with Aqua-Regia (Nitric Acid & Hydrochloric Acids) and then analysis by ICPMS. Results would be reported in mg/kg and/or %. | 0.01 | Y | Y | Y | 21 | 10/13/2019 | 8/5/2020 | 8/10/2020 3:52:07 PM | 2019-09-06 | 10/21/2019 11:56:28 AM | KATFRASE | KATFRASE | DimensionlessRatio | ppm | N | (ratio) |  | Tin Total (ratio) | 7440-31-5 | Metals - Total; | Metals - Total; EMS code: SN-T | Chemical | Total | LAB
SN-T | mg/L | Tin Total | F068 | HNO3dig,TeflonTube/Autoclave,ICP-MS | This method is applicable to determiniation of dissolved, extractable, and total recoverable metals in ground waters, surface waters, drinking waters, and wastewaters. Dissolved, extractable, and total recoverable sample fractions are acidified to pH <1.5 with analytical grade nitric acid (HNO3) to prevent the metals from adsorbing to the container material or precipitating. Extractable and total recoverable water samples are collected in 250 ml acid washed polyethylene bottles and are preserved with Nitric acid (0.5 ml acid for 100 mLs of sample). All preserved samples are stored at room temperature and can be held for up to 6 months. Unfiltered, un-preserved samples are stored at 4C for no more than 30 days. For dissolved and extractable water fractions, no sample preparation is necessary. Samples are simply poured in autosampler tubes (i.e. 15 ml centrifuge tubes), and placed on the autosampler and analyzed. Samples for total recoverable metals are digested with Nitric acid in special Teflon tubes in an autoclave prior to analyses. | 1e-05 | Y | Y | Y | 1779 | 6/1/1999 1:33:00 PM | 7/4/2008 9:48:00 AM | 7/30/2024 6:32:12 AM | 1999-06-28 | 7/6/1999 3:59:45 PM | NPEPPIN | NPEPPIN | Concentration | mg/L | Y | (fl. conc.) |  | Tin Total (fl. conc.) | 7440-31-5 | Metals - Total; | Metals - Total; EMS code: SN-T | Chemical | Total | LAB
SN-T | ppm (S) | Tin Total | MS61 | ALS Minerals ME-MS61 | ALS Minerals ME-MS61. A prepared sample is digested with HNO3; HF; and HClO4 in a graphite heating block to near dryness. After cooling; the sample is boiled in diluted HCl. Bulk to volume with deionized water; mixed and analyzed by ICP-AES and ICP-MS for all elements. | 0.2 | Y | Y | Y | 0 |  |  |  | 2024-01-29 |  | EMS |  | DimensionlessRatio | ppm | N | (ratio) |  | Tin Total (ratio) | 7440-31-5 | Metals - Total; | Metals - Total; EMS code: SN-T | Chemical | Total | LAB
SN-T | mg/L | Tin Total | F082 | Aliquot,HNO3/HCL dig,ICPMS | Aliquot microwave digested with 3:1 HNO3/HCL. Analysis by ICPMS. For HiVol samples preliminary result are reported in mg/L and loading results are reported in ug/m3 | 0.0005 | Y | Y | Y | 26969 | 7/26/2005 7:04:00 AM | 8/23/2024 | 8/26/2024 4:53:05 PM | 2000-06-13 | 8/15/2019 11:11:50 AM | NPEPPIN | KATFRASE | Concentration | mg/L | Y | (fl. conc.) |  | Tin Total (fl. conc.) | 7440-31-5 | Metals - Total; | Metals - Total; EMS code: SN-T | Chemical | Total | LAB
SN-T | ug/m3 | Tin Total | 6040 | Aqua Regia Dig:ICP | Aqua regia digestion is used to bring the metals into solution prior to ICP analysis. For Hi-Vol samples, preliminary results are reported as mg/L and24 hr loadings as ug/m3 | 0.02 | Y | Y | Y | 437 | 2/13/2002 | 2/9/2005 | 2/23/2005 1:13:48 AM | 1998-05-05 | 6/3/1998 6:51:48 PM | NPEPPIN |  | AirConcentration | µg/m³ | Y | (air conc.) |  | Tin Total (air conc.) | 7440-31-5 | Metals - Total; | Metals - Total; EMS code: SN-T | Chemical | Total | LAB
SN-T | mg/L | Tin Total | F121 | Acid,ICPMS | Acidified, no digestion, ICPMS | 1e-05 | Y | Y | Y | 8413 | 3/12/2010 7:57:00 PM | 8/15/2024 9:10:00 PM | 8/26/2024 8:06:03 AM | 2002-07-22 | 7/16/2012 3:31:29 PM | UMULLER | RDGIBSON | Concentration | mg/L | Y | (fl. conc.) |  | Tin Total (fl. conc.) | 7440-31-5 | Metals - Total; | Metals - Total; EMS code: SN-T | Chemical | Total | LAB
SN-T | mg/L | Tin Total | X285 | HNO3 Dig:ICP Analysis | HNO3 Dig:ICP Analysis | 0.3 | Y | Y | Y | 612 | 8/10/1993 | 9/12/2003 3:16:00 PM | 4/30/2004 11:32:38 AM | 1996-04-12 | 6/3/1998 6:51:40 PM | EMS |  | Concentration | mg/L | Y | (fl. conc.) |  | Tin Total (fl. conc.) | 7440-31-5 | Metals - Total; | Metals - Total; EMS code: SN-T | Chemical | Total | LAB
SN-T | mg/L | Tin Total | X073 | FA,HNO3,Direct Flame,AA | Samples are field-acidified with HNO3; HNO3 digestion; analysis by AA. | 0.5 | Y | Y | Y | 69 |  |  | 6/25/2003 9:22:38 AM | 1996-04-12 | 6/3/1998 6:51:40 PM | EMS |  | Concentration | mg/L | Y | (fl. conc.) |  | Tin Total (fl. conc.) | 7440-31-5 | Metals - Total; | Metals - Total; EMS code: SN-T | Chemical | Total | LAB
SN-T | ug/g | Tin Total | FS01 | D&G,Nitric/HCl:ICP | Samples are dried at 60 C, then either sieved to -100mesh (150 um) and rolled to homogenize, or finely ground with a mortar and pestle or a shatter box laboratory mill and rolled (depending on the submitter's requirements). The samples are then weighed into Teflon digestion liners and digested with HNO3 and HCl in a microwave oven. Samples are cooled, volumized, and allowed to settle overnight before being decanted. Analysis is then performed via Inductively Coupled Argon Plasma Atomic Emission Spectrometry (ICP-AES) (Simultaneous multi element analysis). In ICP Spectrometry, the liquid sample is pumped to the nebulizer, where it is convertd into an aerosol. the nebulizer carrier gas flow than transports the wet aerosol past the impact bead, where the large droplets condense out and flow to the waste container. The remaining fine aerosol is carried into the plasma. There, the high temperature desolvates the aerosol and dissociates the sample material into individual atoms and ions, which are excited to emit light at wavelengths characteristic of the elements present. A diffraction grating in the spectrometer diffracts the various wavelengths of light through exit slits to the photo-multiplier tubes, which convert the light energy to electronic signals. | 8 | Y | Y | Y | 2642 | 8/30/1996 1:52:00 PM | 10/10/2007 | 5/24/2008 1:15:03 AM | 1996-06-19 | 6/3/1998 6:51:44 PM | MPAK |  | MassPerMass | mg/kg | Y | (mass conc.) |  | Tin Total (mass conc.) | 7440-31-5 | Metals - Total; | Metals - Total; EMS code: SN-T | Chemical | Total | LAB
SN-T | mg/L | Tin Total | 0040 | HNO3 Dig: ICP Analysis | HNO3 Dig: ICP Analysis | 0.02 | Y | Y | Y | 1733 |  |  | 3/23/2010 5:08:09 PM | 1996-04-12 | 6/3/1998 6:51:38 PM | EMS |  | Concentration | mg/L | Y | (fl. conc.) |  | Tin Total (fl. conc.) | 7440-31-5 | Metals - Total; | Metals - Total; EMS code: SN-T | Chemical | Total | LAB
SN-T | mg/L | Tin Total | X349 | FA:HNO3:ICP | Samples are field-acidified with HNO3; HNO3 digestion; analysis by ICP/AES. | 0.06 | Y | Y | Y | 30 |  |  | 10/30/2014 8:31:18 AM | 1998-05-22 | 6/3/1998 6:51:49 PM | NPEPPIN |  | Concentration | mg/L | Y | (fl. conc.) |  | Tin Total (fl. conc.) | 7440-31-5 | Metals - Total; | Metals - Total; EMS code: SN-T | Chemical | Total | LAB
SN-T | mg/L | Tin Total | X561 | Total metals: FA, HNO3, HR-ICPMS (No Prep)Prep) | Ultra trace metals in water are analyzed by high resolution inductively coupled plasma mass spectrometry (HR-ICPMS), modified from US EPA 200.8, (Revision 5.5). The detection limits provided can only be met for undigested samples. This procedure is intended for colorless, non-turbid, acid-preserved water samples (i.e. pristine water samples), having turbidity < 1 NTU and no odor. Where turbidity exceeds 1 NTU, and/or the sample is colored and has an odor, results may be biased low compared to true Total Metals concentrations. ALS recommends that turbidity analysis be requested on samples submitted for this test to aid with interpretation of results. | 1e-05 | Y | Y | Y | 2756 | 6/8/2015 | 10/5/2023 | 1/31/2024 2:37:41 PM | 2015-10-22 |  | RDGIBSON |  | Concentration | mg/L | Y | (fl. conc.) |  | Tin Total (fl. conc.) | 7440-31-5 | Metals - Total; | Metals - Total; EMS code: SN-T | Chemical | Total | LAB
SN-T | mg/L | Tin Total | MSEA | Metals in Seawater (undigested) by (CRC) ICP-MS | This method describes the multi-elemental determination of trace elements by Collision/Reaction Cell (CRC) ICP-MS. The method measures ions produced by a radio-frequency inductively coupled plasma. Analyte species originating in a liquid are pneumatically nebulized and the resulting aerosol transported by argon gas into the plasma torch. The ions produced are entrained in the plasma gas and extracted, by means of a differentially pumped vacuum interface, into a mass spectrometer. The ions flow through a Collision/Reaction Cell (CRC) where depending on the mode of analysis either the ions passes through the cell unhindered (no gas - classical ICP-MS mode), or collisions/reactions chemistry takes place between the ions and a neutral gas such as hydrogen or helium. Any polyatomic ions which give interference on the analytes are removed. The ions are then sorted according to their mass-to-charge ratios by a quadrupole mass spectrometer. The ions are then sorted according to their mass-to-charge ratios by a quadrupole mass spectrometer. The ions transmitted through the quadrupole are quantified by a channel electron multiplier and the ion information processed by a data handling system. Although most polyatomic interference is removed, any additional interference must be assessed and valid corrections applied or the data flagged to indicate problems. Interference correction must include compensation for background ions contributed by the plasma gas, reagents, and constituents of the sample matrix. Instrument drift as well as suppressions or enhancements of instrument response caused by the sample matrix must be corrected for by the use of internal standards. | 5e-05 | Y | Y | Y | 441 | 8/12/2009 12:08:00 AM | 4/29/2024 | 6/7/2024 10:03:16 AM | 2009-09-08 |  | RDGIBSON |  | Concentration | mg/L | Y | (fl. conc.) |  | Tin Total (fl. conc.) | 7440-31-5 | Metals - Total; | Metals - Total; EMS code: SN-T | Chemical | Total | LAB
SN-T | ug/g | Tin Total | EQ02 | EQUIS (ug/g) | EQUIS (ug/g) | 5 | N | Y | Y | 743 |  |  | 10/23/1996 12:08:28 AM | 1996-04-12 | 6/3/1998 6:51:40 PM | EMS |  | MassPerMass | mg/kg | Y | (mass conc.) |  | Tin Total (mass conc.) | 7440-31-5 | Metals - Total; | Metals - Total; EMS code: SN-T | Chemical | Total | LAB
SN-T | mg/dm2/d | Tin Total | DUST | HNO3 dig.(microwave or hot block),ICPMS | HNO3 dig.(microwave or hot block);ICPMS |  | Y | Y | Y | 420 | 7/19/2016 1:11:00 PM | 7/31/2024 | 7/31/2024 11:23:12 AM | 2008-02-07 |  | RDGIBSON |  | MassFlux | mg/dm²/d | Y | (mass flux) |  | Tin Total (mass flux) | 7440-31-5 | Metals - Total; | Metals - Total; EMS code: SN-T | Chemical | Total | LAB
SN-T | ug/g | Tin Total | PB02 | PBM,HNO3/HCl dig,dry60C,100mesh, | Performance Based Method. Samples are dried at 60 C and sieved using a .149 mm (100 mesh) sieve. Solid samples are digested with the combination of nitric and hydrochloric acids with a 1:1 ratio of acids. For environmental purposes, strong acid leaches are used routinely since they have the advantage of speed, low cost, and superio reproducibility. All metals trapped within the refractory matrix are unlikely to have environmental consequences. Instrumental analysis of extracts produced using this digestion method can be performed by a variety of techniques. |  | Y | Y | Y | 52 | 5/24/2001 3:42:00 PM | 11/26/2001 10:23:00 AM | 5/26/2003 2:28:33 PM | 2001-07-18 |  | NPEPPIN |  | MassPerMass | mg/kg | Y | (mass conc.) |  | Tin Total (mass conc.) | 7440-31-5 | Metals - Total; | Metals - Total; EMS code: SN-T | Chemical | Total | LAB
SN-T | ppm (S) | Tin Total | MS41 | ALS Minerals ME-MS41 | ALS Minerals ME-MS41. A prepared sample is digested with aqua regia in a graphite heating block. After cooling; the resulting solution is diluted to with deionized water; mixed and analyzed by ICP-AES. Samples are then analyzed by ICP-MS for the remaining suite of elements. | 0.2 | Y | Y | Y | 0 |  |  |  | 2024-01-29 |  | EMS |  | DimensionlessRatio | ppm | N | (ratio) |  | Tin Total (ratio) | 7440-31-5 | Metals - Total; | Metals - Total; EMS code: SN-T | Chemical | Total | LAB
SN-T | ug/g | Tin Total | X551 | Metals in Biota (dry), total, acid digestion, ICP-SFMS | 0.1 gram of dried, homogenized biota is weighed in a microwave TFM vessel, and 8mL of Nitric Acid and 1mL of Hydrogen Peroxide are added. The sample is digested at 200 degrees C for 15 minutes in a closed vessel high pressure microwave oven. The sample is cooled, the volume adjusted to 100mL and introduced to an iductively coupled argon plasma-sector field mass spectrometer (ICP-SFMS) by pneumatic nebulization. Each respective element is measured at a specific mass to charge ratio, m/z value expressed in amu (atomic mass units), and at a specified instrument resolution. LR: Low Resolution ~300, MR: Mid Resolution ~4000 and HR: Hich Resolution ~10000. Ref: Method number 02-2705, National Laboratory for Environmental Testing, CCIW, Environment Canada, Burlington. Requested April 2003. | 0.1 | Y | Y | Y | 49 | 8/11/2015 | 10/30/2023 | 11/28/2023 2:21:33 PM | 2004-01-19 |  | UMULLER |  | MassPerMass | mg/kg | Y | (mass conc.) |  | Tin Total (mass conc.) | 7440-31-5 | Metals - Total; | Metals - Total; EMS code: SN-T | Chemical | Total | LAB
SN-T | ug/g | Tin Total | ICPM | Metals in Soil by CRC ICPMS | This analysis is carried out using procedures from CSR Analytical Method: "Strong Acid Leachable Metals (SALM) in Soil", BC Ministry of Environment, 26 June 2009, and procedures adapted from EPA Method 200.2. The sample is manually homogenized, dried at 60 degrees Celsius, sieved through a 63 um (230 mesh) sieve, and a representative subsample of the dry material is weighed. The sample is then digested at 95 degrees Celsius for 2 hours by block digester using concentrated nitric and hydrochloric acids. Instrumental analysis is by collision cell inductively coupled plasma - mass spectrometry (modifed from EPA Method 6020A). Method Limitation: This method is not a total digestion technique. It is a very strong acid digestion that is intended to dissolve those metals that may be environmentally available. By design, elements bound in silicate structures are not normally dissolved by this procedure as they are not usually mobile in the environment. | 0.1 | Y | Y | Y | 103 | 9/19/2017 | 9/24/2023 | 1/12/2024 11:16:52 AM | 2018-01-09 |  | JKROGH |  | MassPerMass | mg/kg | Y | (mass conc.) |  | Tin Total (mass conc.) | 7440-31-5 | Metals - Total; | Metals - Total; EMS code: SN-T | Chemical | Total | LAB
SN-T | ug/g | Tin Total | X230 | Hom,HNO3/HCl,ICP | The sample is dried, ground to pass 100 mesh, and digested with a nitric-hydrochloric acid mixture to solubilize the solid matter and to remove the organic material by oxidation and volatilization. The completed digests are submitted to metals instrumentation personnel for metals analysis by Inductively Coupled Atomic Emission Spectrophotometry (ICP-AES), or hydride analysis by atomic absorption. | 30 | Y | Y | Y | 119 | 2/12/2002 | 12/30/2005 | 2/3/2006 1:13:13 AM | 1996-04-12 | 6/3/1998 6:51:40 PM | EMS |  | MassPerMass | mg/kg | Y | (mass conc.) |  | Tin Total (mass conc.) | 7440-31-5 | Metals - Total; | Metals - Total; EMS code: SN-T | Chemical | Total | LAB
SN-T | ug/g | Tin Total | MTID | MET-DRY-CCMS-VA | Metals in Tissue by CRC ICPMS (DRY) Methodology: This method is conducted following British Columbia Lab Manual method "Metals in Animal Tissue and Vegetation (Biota) - Prescriptive". Tissue samples are homogenized and sub-sampled prior to hotblock digestion with nitric and hydrochloric acids, in combination with addition of hydrogen peroxide. Instrumental analysis is by collision cell inductively coupled plasma - mass spectrometry (modified from EPA Method 6020A). Method Limitation: This method employs a strong acid/peroxide digestion, and is intended to provide a conservative estimate of bio-available metals. Near complete recoveries are achieved for most toxicologically important metals, but elements associated with recalcitrant minerals may be only partially recovered. | 0.1 | Y | Y | Y | 122 | 8/13/2014 | 9/8/2022 | 10/26/2022 7:14:15 AM | 2015-03-09 | 3/11/2015 4:12:05 PM | RDGIBSON | EMS | MassPerMass | mg/kg | Y | (mass conc.) |  | Tin Total (mass conc.) | 7440-31-5 | Metals - Total; | Metals - Total; EMS code: SN-T | Chemical | Total | LAB
SN-T | ppm (S) | Tin Total | 6042 | Aqua Regia by ICPMS | These were sediment/soils digested with Aqua-Regia (Nitric Acid & Hydrochloric Acids) and then analysis by ICPMS. Results would be reported in ppm s. | 0.01 | Y | Y | Y | 71 | 5/14/2022 | 7/27/2024 | 7/29/2024 9:13:08 AM | 2022-06-28 |  | EMS |  | DimensionlessRatio | ppm | N | (ratio) |  | Tin Total (ratio) | 7440-31-5 | Metals - Total; | Metals - Total; EMS code: SN-T | Chemical | Total | LAB
SN-T | mg/L | Tin Total | F038 | HNO3/HCl dig,ICP-AES | Dissolved and extractable water samples are ready for analysis once collected. Total samples are autoclaved with HNO3 and hydrochloric acid (HCI) in capped Teflon vials. Analysis is performed via Inductively Coupled Argon Plasma Atomic Emission Spectrometry (ICP-AES)(simultaneous multielement analysis), Graphite Furnace Atomic Absorption Spectrometry (GF-AAS) (low level analysis) and/or Flame Atomic Absorption Spectrometry (F-AAS). | 0.06 | Y | Y | Y | 5399 | 8/1/1996 6:04:00 PM | 12/1/2010 | 9/11/2019 8:47:43 AM | 1997-05-08 | 6/3/1998 6:51:46 PM | MPAK |  | Concentration | mg/L | Y | (fl. conc.) |  | Tin Total (fl. conc.) | 7440-31-5 | Metals - Total; | Metals - Total; EMS code: SN-T | Chemical | Total | LAB
SN-T | ug/m3 | Tin Total | 5312 | Tef:N/P Dig:ICP Analysis | A nitric/perchloric acid digestion if used to bring the metals into solution prior to ICP analysis. For Hi-Vol samples preliminary results are reported as mg/L and 24 hour loading results as ug/m3. | 0.02 | Y | Y | Y | 1191 | 11/7/1997 | 3/29/2000 | 9/28/2000 1:34:45 PM | 1996-10-13 | 6/3/1998 6:51:46 PM | EMS |  | AirConcentration | µg/m³ | Y | (air conc.) |  | Tin Total (air conc.) | 7440-31-5 | Metals - Total; | Metals - Total; EMS code: SN-T | Chemical | Total | LAB
SN-T | mg/L | Tin Total | 1349 | Tot Metals in Seawater by CRC ICPMS (BC) | Tot Metals in Seawater by CRC ICPMS (BC) | 0.001 | Y | Y | Y | 98 | 8/16/2019 | 12/5/2021 | 12/13/2021 12:34:16 PM | 2019-09-30 |  | KATFRASE |  | Concentration | mg/L | Y | (fl. conc.) |  | Tin Total (fl. conc.) | 7440-31-5 | Metals - Total; | Metals - Total; EMS code: SN-T | Chemical | Total | LAB
SN-T | mg/L | Tin Total | F048 | Acid dig,conc'n,nebul.,ICP-OES | This method provides an automated procedure for the Inductively Coupled Argon Plasma Optical Emission Spectrophotometric (ICP-OES) determination and quantitation of trace amounts of metals in surface, ground and precipitation waters after a manual acid digestion and concentration procedure and ultrasonic nebulisation. Organically-bound metals and metal complexes are converted to metal ions by digestion with nitric acid [method details says HCl]. To increase the signal detected by the ICP, the samples are concentrated 10 times. The liquid sample is pumped onto the piezoelectric transducer, where the ultrasonic action efficiently converts it to a fine, dense aerosol. The nebulizer carrier gas flow then transports the wet aerosol through a heated U-tube, where the solvent is vaporized. Subsequently, most of the solvent vapour is removed from the aerosol stream by passing the stream through a chilled condenser. The result is a dry, analyte-laden aerosol which is introduced to the plasma. There, the high temperature desolvates the aerosol and dissociates the sample material into individual atoms and ions, which are excited to emit light at wave lengths characteristic of the elements present. Dissolved samples are field filtered through 0.45 micron cellulose acetate filter prior to acid preservation. Samples for extractable analysis are acidified with 2 mL concentrated HNO3 per litre and allowed to stand a minimum of 24 hours before analysis. Clients providing these samples are advised to preserve samples at the collection site. | 0.03 | Y | Y | Y | 29 |  |  | 7/6/2001 11:10:14 AM | 1998-11-13 |  | NPEPPIN |  | Concentration | mg/L | Y | (fl. conc.) |  | Tin Total (fl. conc.) | 7440-31-5 | Metals - Total; | Metals - Total; EMS code: SN-T | Chemical | Total | LAB
SN-T | ug/g | Tin Total | 0190 | D&G:N/P Dig:ICP Anal. | Sample is dried and grind-mixed before nitric perchloric digestion followed by ICP analysis. | 5 | Y | Y | Y | 144 |  |  | 10/24/1996 6:10:56 AM | 1996-04-12 | 6/3/1998 6:51:41 PM | EMS |  | MassPerMass | mg/kg | Y | (mass conc.) |  | Tin Total (mass conc.) | 7440-31-5 | Metals - Total; | Metals - Total; EMS code: SN-T | Chemical | Total | LAB
SN-T | mg/L | Tin Total | 0042 | GI HNO3 Dig: ICP | Sample aliquot taken from unfiltered unpreserved General Ions bottle; HNO3 digestion; ICP analysis. | 0.02 | Y | Y | Y | 6253 | 9/20/1996 | 7/5/2019 3:21:00 PM | 7/12/2019 6:49:13 PM | 1996-04-12 | 6/3/1998 6:51:38 PM | EMS |  | Concentration | mg/L | Y | (fl. conc.) |  | Tin Total (fl. conc.) | 7440-31-5 | Metals - Total; | Metals - Total; EMS code: SN-T | Chemical | Total | LAB
SN-T | mg/L | Tin Total | METN | MET-T-NP-CCMS-VA | Ultra trace metals in water are analyzed by CRC ICPMS, based on US EPA Method 6020B (July 2014). The detection limits provided can only be met for undigested samples. This procedure is intended for colorless, non-turbid, acid-preserved water samples (i.e. pristine water samples), having turbidity < 1 NTU and no odor. Where turbidity exceeds 1 NTU, and/or the sample is colored and has an odor, results may be biased low compared to true Total Metals concentrations. ALS recommends that turbidity analysis be requested on samples submitted for this test to aid with interpretation of results. Where turbidity is <1NTU, undigested metals are equivalent to total metals concentrations. | 5e-05 | Y | Y | Y | 3648 | 4/3/2012 | 8/19/2024 | 8/26/2024 1:03:06 PM | 2019-08-20 |  | KATFRASE |  | Concentration | mg/L | Y | (fl. conc.) |  | Tin Total (fl. conc.) | 7440-31-5 | Metals - Total; | Metals - Total; EMS code: SN-T | Chemical | Total | LAB
SN-T | ug/g wet | Tin Total | BPSY | MET-WET-BPSY-CCMS-VA | Metals in Animal Tissue and Vegetation (Biota) - Prescriptive. EPA 200.3/6020B This method is conducted following British Columbia Lab Manual method "Metals in Animal Tissue and Vegetation (Biota) - Prescriptive". Tissue samples are digested with nitric and hydrochloric acids; in combination with addition of hydrogen peroxide. Instrumental analysis is by collision cell inductively coupled plasma - mass spectrometry (modified from EPA Method 6020B). | 0.02 | Y | Y | Y | 2 | 11/7/2019 | 11/7/2019 | 6/10/2021 1:57:10 PM | 2021-06-04 |  | EMS |  | MassPerMass | µg/g | Y | (mass conc.) | (wet) | Tin Total (wet) (mass conc.) | 7440-31-5 | Metals - Total; | Metals - Total; EMS code: SN-T | Chemical | Total | LAB
SN-T | ppm (S) | Tin Total | MS81 | ALS Minerals: ME-MS81 | ALS Minerals: ME-MS81. A prepared sample (0.100 g) is added to lithium metaborate/lithium tetraborate flux; mixed well and fused in a furnace at 1025°C. The resulting melt is then cooled and dissolved in an acid mixture containing nitric; hydrochloric and hydrofluoric acids. This solution is then analyzed by inductively coupled plasma - mass spectrometry. | 1 | Y | Y | Y | 0 |  |  |  | 2024-01-29 |  | EMS |  | DimensionlessRatio | ppm | N | (ratio) |  | Tin Total (ratio) | 7440-31-5 | Metals - Total; | Metals - Total; EMS code: SN-T | Chemical | Total | LAB
SN-T | mg/L | Tin Total | X170 | ICP | This code is a catch-all for a large variety of ICP methods including water and also solids digested into solutions as an intermediate result. Different procedures will have greatly different detection limits. (a) For air particulates, preliminary results are reported in mg/L and loading results as mg/dm2/d. (b) For Hi Vol samples preliminiary results are same as reported in mg/L and loading results as ug/m3. (c) Note that for Rain samples there is no filtration nor digestion steps. | 0.02 | Y | Y | Y | 3636 | 3/28/1996 6:03:00 PM | 11/7/2018 | 10/24/2022 10:58:25 AM | 1996-04-12 | 6/3/1998 6:51:41 PM | EMS |  | Concentration | mg/L | Y | (fl. conc.) |  | Tin Total (fl. conc.) | 7440-31-5 | Metals - Total; | Metals - Total; EMS code: SN-T | Chemical | Total | LAB
SO2- | ug/m3 | Sulfur Dioxide | EQ40 | EQUIS (mg/m3) | EQUIS (mg/m3) | 0 | N | Y | Y | 119 |  |  | 10/23/1996 12:01:37 AM | 1996-04-12 | 6/3/1998 6:51:38 PM | EMS |  | AirConcentration | µg/m³ | Y | (air conc.) |  | Sulfur Dioxide (air conc.) | 7446-09-5 | Inorganics; | Inorganics; EMS code: SO2- | Chemical |  | LAB
SO2- | ug/m3 | Sulfur Dioxide | X198 | Dynascience-Cont. Anal. | Dynascience-Cont. Anal. | 0 | Y | Y | Y | 9349 |  |  | 5/21/1998 2:11:35 PM | 1996-04-12 | 6/3/1998 6:51:46 PM | EMS |  | AirConcentration | µg/m³ | Y | (air conc.) |  | Sulfur Dioxide (air conc.) | 7446-09-5 | Inorganics; | Inorganics; EMS code: SO2- | Chemical |  | LAB
SO2- | ug/m3 | Sulfur Dioxide | X239 | Gas Chromatography | Gas Chromatography | 0 | Y | Y | Y | 423 | 2/21/2014 8:00:00 AM | 6/18/2024 8:00:00 AM | 7/30/2024 10:08:09 AM | 1996-04-12 | 6/3/1998 6:51:38 PM | EMS |  | AirConcentration | µg/m³ | Y | (air conc.) |  | Sulfur Dioxide (air conc.) | 7446-09-5 | Inorganics; | Inorganics; EMS code: SO2- | Chemical |  | LAB
SO2- | t/d | Sulfur Dioxide | CALC | Calculation (t/d) | Calculation (t/d) | 0 | Y | Y | Y | 1 |  |  | 11/22/1996 12:40:20 PM | 1996-04-12 | 6/3/1998 6:51:38 PM | EMS |  | MassFlow | t/d | Y | (mass flow) |  | Sulfur Dioxide (mass flow) | 7446-09-5 | Inorganics; | Inorganics; EMS code: SO2- | Chemical |  | LAB
SO2- | ug/m3 | Sulfur Dioxide | 5098 | Ion Chrom-Anions Stack | Ion Chrom-Anions Stack | 0.001 | Y | Y | Y | 2 |  |  | 10/23/1996 11:06:05 PM | 1996-04-12 | 6/3/1998 6:51:38 PM | EMS |  | AirConcentration | µg/m³ | Y | (air conc.) |  | Sulfur Dioxide (air conc.) | 7446-09-5 | Inorganics; | Inorganics; EMS code: SO2- | Chemical |  | LAB
SO2- | ug/m3 | Sulfur Dioxide | X193 | HCl,Vol'metric Titr.-Air | HCl;Vol'metric Titr.-Air | 0 | Y | Y | Y | 289 |  |  | 2/19/2002 10:06:21 AM | 1996-04-12 | 6/3/1998 6:51:38 PM | EMS |  | AirConcentration | µg/m³ | Y | (air conc.) |  | Sulfur Dioxide (air conc.) | 7446-09-5 | Inorganics; | Inorganics; EMS code: SO2- | Chemical |  | LAB
SO2- | t/d | Sulfur Dioxide | EQ30 | EQUIS (T/d) | EQUIS (T/d) | 0 | N | Y | Y | 9 |  |  | 10/22/1996 7:36:35 PM | 1996-04-12 | 6/3/1998 6:51:38 PM | EMS |  | MassFlow | t/d | Y | (mass flow) |  | Sulfur Dioxide (mass flow) | 7446-09-5 | Inorganics; | Inorganics; EMS code: SO2- | Chemical |  | LAB
SO2- | % | Sulfur Dioxide | EQ34 | EQUIS (%) | EQUIS (%) | 0 | N | Y | Y | 4158 |  |  | 10/23/1996 12:01:33 AM | 1996-04-12 | 6/3/1998 6:51:38 PM | EMS |  | DimensionlessRatio | % | N | (ratio) |  | Sulfur Dioxide (ratio) | 7446-09-5 | Inorganics; | Inorganics; EMS code: SO2- | Chemical |  | LAB
SO2- | ug/m3 | Sulfur Dioxide | CAL6 | Calculation (mg/m3) | Calculation (mg/m3) | 1 | Y | Y | Y | 34884 |  |  | 7/3/2024 12:59:29 PM | 1997-01-16 | 6/3/1998 6:51:46 PM | MPAK |  | AirConcentration | µg/m³ | Y | (air conc.) |  | Sulfur Dioxide (air conc.) | 7446-09-5 | Inorganics; | Inorganics; EMS code: SO2- | Chemical |  | LAB
SO2- | ppb A | Sulfur Dioxide | DIOXR | Dioxides by Radiello | Dioxides in air by Radiello | 1 | Y | Y | Y | 64 | 8/25/2023 | 8/21/2024 | 8/22/2024 2:48:06 PM | 2023-08-29 |  | EMS |  | DimensionlessRatio | ppb | N | (ratio) |  | Sulfur Dioxide (ratio) | 7446-09-5 | Inorganics; | Inorganics; EMS code: SO2- | Chemical |  | LAB
SO2- | ug/m3 | Sulfur Dioxide | 5232 | Thorin Titration | Thorin Titration | 0.001 | Y | Y | Y | 7 |  |  | 8/1/2024 3:52:42 PM | 1996-04-12 | 6/3/1998 6:51:38 PM | EMS |  | AirConcentration | µg/m³ | Y | (air conc.) |  | Sulfur Dioxide (air conc.) | 7446-09-5 | Inorganics; | Inorganics; EMS code: SO2- | Chemical |  | LAB
SO4- | ug/m3 | Sulfate-Soluble | 5322 | Tef:Sulfate-IC soluble | Tef:Sulfate-IC soluble | 0.01 | Y | Y | Y | 794 | 10/15/1999 7:20:00 PM | 8/30/2005 | 11/5/2005 1:11:14 AM | 1996-04-12 | 6/3/1998 6:51:38 PM | EMS |  | AirConcentration | µg/m³ | Y | (air conc.) |  | Sulfate-Soluble (air conc.) | 14808-79-8 |  | EMS code: SO4- | Chemical | Soluble | LAB
SO4- | mg/L | Sulfate-Soluble | 5074 | Auto Color- MTB Method | Auto Color- MTB Method | 0.5 | Y | Y | Y | 217 |  |  | 5/26/1998 2:55:02 PM | 1996-04-12 | 6/3/1998 6:51:38 PM | EMS |  | Concentration | mg/L | Y | (fl. conc.) |  | Sulfate-Soluble (fl. conc.) | 14808-79-8 |  | EMS code: SO4- | Chemical | Soluble | LAB
SO4- | mg/L | Sulfate-Soluble | 5068 | Ion Chromatography-Anion | (a) For the Soluble Particulates Pkg., a portion of filtrate from the Particulate-Total procedure made up to 500 mL is analyzed by Ion Chromatography. Intermediate results are in mg/L. Loading results are then calculated with units of mg/dm2/d. (b) For precipitation samples, samples and standards are spiked with concentrated carbonate/bicarbonate solution to give the sample the same background as the eluent used (avoiding a "water dip" effect). The anions of interest are separated through an anion "guard" and anion "separator" column. An anion micro membrane suppressor following the separator columns is used to reduce the background eluent conductivity by converting the carbonate and bicarbonate species to carbonic acid, while enhancing the conductivity of the ions of interest by converting these ions to their corresponding acids. A conductivity detector senses the same species in direct proportion to their initial concentration. (c) For HiVol samples, discs are cut from the HiVol filter, extracted with DI water, and anions analyzed by ion chromatograph, Preliminary results are reported in mg/L and intermediate HiVol loading results are reported as ug/m3. (*) Note that the air-convention is to report on basis of ion weight not element, e.g. as NO3 for nitrate, and as PO4 for phosphate. | 0.01 | Y | Y | Y | 7007 | 4/30/1997 3:50:00 PM | 12/10/2011 12:01:00 PM | 2/9/2012 12:15:57 PM | 1996-04-12 | 6/3/1998 6:51:38 PM | EMS |  | Concentration | mg/L | Y | (fl. conc.) |  | Sulfate-Soluble (fl. conc.) | 14808-79-8 |  | EMS code: SO4- | Chemical | Soluble | LAB
SO4- | mg/dm2/d | Sulfate-Soluble | 5049 | Soluble Particulates Pkg. Ion Chrom. Anions | For the Soluble Particulates Pkg., a portion of filtrate from the Particulate-Total procedure made up to 500 mL is analyzed by Ion Chromatography. Intermediate results are in mg/L. Loading results are then calculated with units of mg/dm2/d. Note that the air-convention is to report on basis of ion weight not element, e.g. as NO3 for nitrate, and as PO4 for phosphate. | 0.1 | Y | Y | Y | 128 |  |  | 10/24/1996 6:07:34 AM | 1996-04-12 | 6/3/1998 6:51:38 PM | EMS |  | MassFlux | mg/dm²/d | Y | (mass flux) |  | Sulfate-Soluble (mass flux) | 14808-79-8 |  | EMS code: SO4- | Chemical | Soluble | LAB
SO4- | ug/m3 | Sulfate-Soluble | EQ40 | EQUIS (mg/m3) | EQUIS (mg/m3) | 0 | N | Y | Y | 2961 |  |  | 10/22/1996 10:21:01 PM | 1996-04-12 | 6/3/1998 6:51:38 PM | EMS |  | AirConcentration | µg/m³ | Y | (air conc.) |  | Sulfate-Soluble (air conc.) | 14808-79-8 |  | EMS code: SO4- | Chemical | Soluble | LAB
SO4- | mg/dm2/d | Sulfate-Soluble | EQ15 | EQUIS (mg/dm2/d) | EQUIS (mg/dm2/d) | 0.002 | N | Y | Y | 582 |  |  | 10/22/1996 10:27:18 PM | 1996-04-12 | 6/3/1998 6:51:38 PM | EMS |  | MassFlux | mg/dm²/d | Y | (mass flux) |  | Sulfate-Soluble (mass flux) | 14808-79-8 |  | EMS code: SO4- | Chemical | Soluble | LAB
SO4- | mg/L | Sulfate-Soluble | X064 | BaCl2,Turbimetric | Sulfate ion is converted to a barium sulfate suspension under controlled conditions. The resulting turbidity is determined using a nephelometer or spectrophotometer and compared to a curve prepared from standard sulfate solutions. | 0.001 | Y | Y | Y | 136 |  |  | 10/25/1999 2:46:13 PM | 1996-04-12 | 6/3/1998 6:51:38 PM | EMS |  | Concentration | mg/L | Y | (fl. conc.) |  | Sulfate-Soluble (fl. conc.) | 14808-79-8 |  | EMS code: SO4- | Chemical | Soluble | LAB
SO4- | mg/L | Sulfate-Soluble | EQ01 | EQUIS (mg/L) | EQUIS (mg/L) | 0.05 | N | Y | Y | 132 |  |  | 10/22/1996 11:12:48 PM | 1996-04-12 | 6/3/1998 6:51:38 PM | EMS |  | Concentration | mg/L | Y | (fl. conc.) |  | Sulfate-Soluble (fl. conc.) | 14808-79-8 |  | EMS code: SO4- | Chemical | Soluble | LAB
SO4- | ug/m3 | Sulfate-Soluble | 5022 | Water Extr,Ion Chr-Anion | Discs cut from the HiVol filter are extracted with deionized water and the resulting anions are separated an measured using an ion chromatograph. (See Anions-Ion Chromatography-Precipitation, methods NO3-5068, S04-5068, CL-S5068). Values are then converted to units of ug/m3. | 5 | Y | Y | Y | 1723 | 7/10/2000 10:05:00 AM | 1/14/2002 11:59:00 AM | 8/11/2003 2:38:42 PM | 1996-04-12 | 6/3/1998 6:51:38 PM | EMS |  | AirConcentration | µg/m³ | Y | (air conc.) |  | Sulfate-Soluble (air conc.) | 14808-79-8 |  | EMS code: SO4- | Chemical | Soluble | LAB
SO4-A | ug/g | Available Sulfate (as S) as per Alberta Agriculture | AASO4 | Available SO4- | Available SO4- (as S) as per Alberta Agriculture | 3 | Y | Y | Y | 15 | 1/13/2022 | 2/3/2022 | 1/11/2024 1:07:08 PM | 2022-06-29 |  | EMS |  | MassPerMass | mg/kg | Y | (mass conc.) |  | Available Sulfate (as S) as per Alberta Agriculture (mass conc.) | 14808-79-8 |  | EMS code: SO4-A | Chemical |  | LAB
SO4-E | mg/L | Sulfate Extractable | E300 | EPA 300.1/CSSS 15 | See EPA method 300.1 and Canadian Society of Soil Sciences (2008) Chapter 15 | 5 | Y | Y | Y | 15 | 1/11/2022 | 2/5/2022 | 1/11/2024 1:07:08 PM | 2022-06-29 |  | EMS |  | Concentration | mg/L | Y | (fl. conc.) |  | Sulfate Extractable (fl. conc.) | 14808-79-8 |  | EMS code: SO4-E | Chemical | Extractable | LAB
SO4-E | ug/g | Sulfate Extractable | CAL2 | Calculation (ug/g) | Calculation (ug/g) | 2.5 | Y | Y | Y | 15 | 1/13/2022 | 2/7/2022 | 1/11/2024 1:07:08 PM | 2022-07-06 |  | EMS |  | MassPerMass | mg/kg | Y | (mass conc.) |  | Sulfate Extractable (mass conc.) | 14808-79-8 |  | EMS code: SO4-E | Chemical | Extractable | LAB
SO4-L | mg/L | Sulfate Leachable | MEND | MEND (mod) EPA 300.1 (mod) | MEND (mod) EPA 300.1 (mod). Inorganic anions are analyzed by obtaining an extract from a soil sample that has been added in a 3:1 ratio of deionized water to soil; then leached and the extract filtered prior to analysis; which is then analyzed by Ion Chromatography with conductivity and/or UV detection. | 0.5 | Y | Y | Y | 0 |  |  |  | 2024-01-24 |  | RDGIBSON |  | Concentration | mg/L | Y | (fl. conc.) |  | Sulfate Leachable (fl. conc.) |  |  | EMS code: SO4-L | Chemical |  | LAB
SO4-S | mg/L | Sulfate; soluble ion content | E6010B | EPA 6010B/CSSS 15 | See CSSS CH15/EPA 6010B/AER D50. INDUCTIVELY COUPLED PLASMA-ATOMIC EMISSION SPECTROMETRY See EPA 6010B and CSSS Ch. 15. https://www.epa.gov/sites/default/files/documents/6010b.pdf | 6 | Y | Y | Y | 0 |  |  |  | 2024-01-30 |  | EMS |  | Concentration | mg/L | Y | (fl. conc.) |  | Sulfate; soluble ion content (fl. conc.) |  |  | EMS code: SO4-S | Chemical |  | LAB
SO4-S | mg/L | Sulfate; soluble ion content | E300 | EPA 300.1/CSSS 15 | See EPA method 300.1 and Canadian Society of Soil Sciences (2008) Chapter 15 | 5 | Y | Y | Y | 0 |  |  |  | 2024-02-02 |  | EMS |  | Concentration | mg/L | Y | (fl. conc.) |  | Sulfate; soluble ion content (fl. conc.) |  |  | EMS code: SO4-S | Chemical |  | LAB
SO4-S | ug/g | Sulfate; soluble ion content | CAL2 | Calculation (ug/g) | Calculation (ug/g) | 2.5 | Y | Y | Y | 0 |  |  |  | 2024-02-02 |  | EMS |  | MassPerMass | mg/kg | Y | (mass conc.) |  | Sulfate; soluble ion content (mass conc.) |  |  | EMS code: SO4-S | Chemical |  | LAB
SOX- | ug/m3 | Sulfur Oxides | 5098 | Ion Chrom-Anions Stack | Ion Chrom-Anions Stack | 0.001 | Y | Y | Y | 4 |  |  | 5/28/1998 8:42:17 AM | 1996-04-12 | 6/3/1998 6:51:38 PM | EMS |  | AirConcentration | µg/m³ | Y | (air conc.) |  | Sulfur Oxides (air conc.) |  | Inorganics; | Inorganics; EMS code: SOX- | Chemical |  | LAB
SR-D | mg/L | Strontium Dissolved | X562 | Dissolved metals: FF, FA, HNO3, HR-ICPMS | Ultra trace metals in water are analyzed by high resolution inductively coupled plasma mass spectrometry (HR-ICPMS), based on US EPA Method 200.8, (Oct. 1999). This procedure is intended for pristine field-filtered acid-preserved water samples. The detection limits (LOR) for this test are based on lab instrumental analysis only, not including filtration. ALS-supplied field filtration equipment does not support these LOR. Therefore, because of the high probability of false positives due to filtration, it is strongly recommended that a filtration blank be analyzed to aid in data interpretation. | 5e-05 | Y | Y | Y | 1264 | 6/8/2015 | 10/19/2023 | 11/9/2023 7:41:27 AM | 2015-10-22 |  | RDGIBSON |  | Concentration | mg/L | Y | (fl. conc.) |  | Strontium Dissolved (fl. conc.) | 7440-24-6 | Metals - Dissolved; | Metals - Dissolved; EMS code: SR-D | Chemical | Dissolved | LAB
SR-D | mg/L | Strontium Dissolved | 0040 | HNO3 Dig: ICP Analysis | HNO3 Dig: ICP Analysis | 0.001 | Y | Y | Y | 76 | 9/29/2000 | 12/1/2010 | 12/9/2010 3:55:27 PM | 1996-04-12 | 6/3/1998 6:51:40 PM | EMS |  | Concentration | mg/L | Y | (fl. conc.) |  | Strontium Dissolved (fl. conc.) | 7440-24-6 | Metals - Dissolved; | Metals - Dissolved; EMS code: SR-D | Chemical | Dissolved | LAB
SR-D | mg/L | Strontium Dissolved | MSEA | Metals in Seawater (undigested) by (CRC) ICP-MS | This method describes the multi-elemental determination of trace elements by Collision/Reaction Cell (CRC) ICP-MS. The method measures ions produced by a radio-frequency inductively coupled plasma. Analyte species originating in a liquid are pneumatically nebulized and the resulting aerosol transported by argon gas into the plasma torch. The ions produced are entrained in the plasma gas and extracted, by means of a differentially pumped vacuum interface, into a mass spectrometer. The ions flow through a Collision/Reaction Cell (CRC) where depending on the mode of analysis either the ions passes through the cell unhindered (no gas - classical ICP-MS mode), or collisions/reactions chemistry takes place between the ions and a neutral gas such as hydrogen or helium. Any polyatomic ions which give interference on the analytes are removed. The ions are then sorted according to their mass-to-charge ratios by a quadrupole mass spectrometer. The ions are then sorted according to their mass-to-charge ratios by a quadrupole mass spectrometer. The ions transmitted through the quadrupole are quantified by a channel electron multiplier and the ion information processed by a data handling system. Although most polyatomic interference is removed, any additional interference must be assessed and valid corrections applied or the data flagged to indicate problems. Interference correction must include compensation for background ions contributed by the plasma gas, reagents, and constituents of the sample matrix. Instrument drift as well as suppressions or enhancements of instrument response caused by the sample matrix must be corrected for by the use of internal standards. | 2e-05 | Y | Y | Y | 4 | 4/12/2023 | 4/29/2024 | 5/16/2024 2:00:14 PM | 2009-09-08 |  | RDGIBSON |  | Concentration | mg/L | Y | (fl. conc.) |  | Strontium Dissolved (fl. conc.) | 7440-24-6 | Metals - Dissolved; | Metals - Dissolved; EMS code: SR-D | Chemical | Dissolved | LAB
SR-D | mg/L | Strontium Dissolved | X170 | ICP | This code is a catch-all for a large variety of ICP methods including water and also solids digested into solutions as an intermediate result. Different procedures will have greatly different detection limits. (a) For air particulates, preliminary results are reported in mg/L and loading results as mg/dm2/d. (b) For Hi Vol samples preliminiary results are same as reported in mg/L and loading results as ug/m3. (c) Note that for Rain samples there is no filtration nor digestion steps. | 0.1 | Y | Y | Y | 11098 | 4/11/1996 6:02:00 PM | 11/6/2018 | 8/26/2024 2:41:32 PM | 1996-04-12 | 6/3/1998 6:51:40 PM | EMS |  | Concentration | mg/L | Y | (fl. conc.) |  | Strontium Dissolved (fl. conc.) | 7440-24-6 | Metals - Dissolved; | Metals - Dissolved; EMS code: SR-D | Chemical | Dissolved | LAB
SR-D | mg/L | Strontium Dissolved | METC | MET-D-NP-CCMS-VA | Ultra trace metals in water are analyzed by CRC ICPMS, based on US EPA Method 6020B (July 2014). This procedure is intended for pristine field-filtered acid-preserved water samples. The detection limits (LOR) for this test are based on lab instrumental analysis only, not including filtration. ALS-supplied field filtration equipment does not support these LOR. Therefore, because of the high probability of false positives due to filtration, it is strongly recommended that a filtration blank be analysed to aid in data interpretation. | 5e-05 | Y | Y | Y | 1776 | 4/3/2012 | 8/15/2024 | 8/23/2024 9:50:35 AM | 2019-08-20 |  | KATFRASE |  | Concentration | mg/L | Y | (fl. conc.) |  | Strontium Dissolved (fl. conc.) | 7440-24-6 | Metals - Dissolved; | Metals - Dissolved; EMS code: SR-D | Chemical | Dissolved | LAB
SR-D | mg/L | Strontium Dissolved | F032 | Filt.,Acid,ICPMS | There is no acid digestion required for this analysis. A water sample is collected and preserved in 0.2 percent nitric acid. To analyse the extractable metals the sample is shaken, allowed to sit for at least twenty four hours, and a decanted portion is introduced to the spectrometer. When the dissolved metals are required, the samples must first be filtered through a 0.45 micron cellulose acetate filter paper and then followed by preservation in 0.2 percent nitric acid. It is important to note that different laboratories may have very different detection limits for some variables analyzed by ICPMS. | 1e-05 | Y | Y | Y | 80949 | 8/16/1996 2:35:00 PM | 8/26/2024 | 8/26/2024 4:53:05 PM | 1997-01-16 | 6/3/1998 6:51:47 PM | MPAK |  | Concentration | mg/L | Y | (fl. conc.) |  | Strontium Dissolved (fl. conc.) | 7440-24-6 | Metals - Dissolved; | Metals - Dissolved; EMS code: SR-D | Chemical | Dissolved | LAB
SR-D | mg/L | Strontium Dissolved | X285 | HNO3 Dig:ICP Analysis | HNO3 Dig:ICP Analysis | 0.001 | Y | Y | Y | 650 | 1/30/2002 | 11/13/2002 | 2/19/2004 4:22:44 PM | 1998-05-22 | 6/3/1998 6:51:49 PM | NPEPPIN |  | Concentration | mg/L | Y | (fl. conc.) |  | Strontium Dissolved (fl. conc.) | 7440-24-6 | Metals - Dissolved; | Metals - Dissolved; EMS code: SR-D | Chemical | Dissolved | LAB
SR-D | mg/L | Strontium Dissolved | MWID | Metals in Water by ICPMS (digested) | Metals in Water by ICPMS (digested) Based on EPA 200.8 | 0.001 | Y | Y | Y | 31763 | 10/26/2018 5:20:00 PM | 10/26/2018 5:20:00 PM | 7/31/2024 1:46:41 PM | 2009-12-24 |  | RDGIBSON |  | Concentration | mg/L | Y | (fl. conc.) |  | Strontium Dissolved (fl. conc.) | 7440-24-6 | Metals - Dissolved; | Metals - Dissolved; EMS code: SR-D | Chemical | Dissolved | LAB
SR-D | mg/L | Strontium Dissolved | F039 | Field filt + acid, ICP-AES | Dissolved and extractable water samples are ready for analysis once collected. Total samples are autoclaved with HNO3 and hydrochloric acid (HCI) in capped Teflon vials. Analysis is performed via Inductively Coupled Argon Plasma Atomic Emission Spectrometry (ICP-AES)(simultaneous multielement analysis), Graphite Furnace Atomic Absorption Spectrometry (GF-AAS)(low level analysis) and/or Flame Atomic Absorption Spectrometry (F-AAS). | 0.001 | Y | Y | Y | 6579 | 7/29/1996 6:03:00 PM | 7/5/2019 1:21:00 PM | 5/16/2022 4:21:33 PM | 1997-05-08 | 6/3/1998 6:51:46 PM | MPAK |  | Concentration | mg/L | Y | (fl. conc.) |  | Strontium Dissolved (fl. conc.) | 7440-24-6 | Metals - Dissolved; | Metals - Dissolved; EMS code: SR-D | Chemical | Dissolved | LAB
SR-D | mg/L | Strontium Dissolved | F068 | HNO3dig,TeflonTube/Autoclave,ICP-MS | This method is applicable to determiniation of dissolved, extractable, and total recoverable metals in ground waters, surface waters, drinking waters, and wastewaters. Dissolved, extractable, and total recoverable sample fractions are acidified to pH <1.5 with analytical grade nitric acid (HNO3) to prevent the metals from adsorbing to the container material or precipitating. Extractable and total recoverable water samples are collected in 250 ml acid washed polyethylene bottles and are preserved with Nitric acid (0.5 ml acid for 100 mLs of sample). All preserved samples are stored at room temperature and can be held for up to 6 months. Unfiltered, un-preserved samples are stored at 4C for no more than 30 days. For dissolved and extractable water fractions, no sample preparation is necessary. Samples are simply poured in autosampler tubes (i.e. 15 ml centrifuge tubes), and placed on the autosampler and analyzed. Samples for total recoverable metals are digested with Nitric acid in special Teflon tubes in an autoclave prior to analyses. | 5e-06 | Y | Y | Y | 1397 | 6/1/1999 1:22:00 PM | 4/17/2015 1:06:00 PM | 8/11/2021 12:04:11 PM | 1999-06-28 |  | NPEPPIN |  | Concentration | mg/L | Y | (fl. conc.) |  | Strontium Dissolved (fl. conc.) | 7440-24-6 | Metals - Dissolved; | Metals - Dissolved; EMS code: SR-D | Chemical | Dissolved | LAB
SR-D | mg/L | Strontium Dissolved | X556 | Metals in water, dissolved, "in-bottle" digestion, ICP-SFMS | The whole water sample is filtered through a 0.45 micron cellulose acetate filter paper and the filtrate is stored in a HDPE/Teflon bottle. The container is de-capped in a Class 100 workstation, acidified to 2% (v/v) HNO3 and 1% (v/v) HCl and recapped. The sample is "in-bottle" digested in a convection oven, set at 60 degrees C for 16 hours. The sample is cooled and introduced to an inductively coupled argon plasma-sector field mass spectrometer (ICP-SFMS) by pneumatic nebulization. Each respective element is measured at a specific mass to charge ratio, m/z value expressed in amu (atomic mass units), and at a specified intrument resolution. LR: Low Resolution ~300, MR: Mid Resolution ~4000 and HR: High Resolution ~10000. Ref: Method number 02-2003, National Laboratory for Environmental Testing, CCIW, Environment Canada, Burlington, Requested April 2003. | 5e-05 | Y | Y | Y | 4234 | 6/11/2003 | 12/6/2023 2:04:00 AM | 2/14/2024 9:05:32 AM | 2004-12-07 | 12/7/2004 2:27:09 PM | UMULLER | UMULLER | Concentration | mg/L | Y | (fl. conc.) |  | Strontium Dissolved (fl. conc.) | 7440-24-6 | Metals - Dissolved; | Metals - Dissolved; EMS code: SR-D | Chemical | Dissolved | LAB
SR-E | mg/L | Strontium Extractable | X553 | Metals in water, extractable, ICP-SFMS | The whole water sample is collected in a HDPE/Teflon bottle. In a Class 100 workstation it is preserved with concentrated Nitric Acid at concentration of 0.2 percent (v/v). The acid preservative extracts the metals from material in the sample and a decanted portion is introduced directly to an inductively coupled argon plasma-sector field mass spectrometer (ICP-SFMS) by pneumatic nebulization. Each respective element is measured at a specific mass to charge ratio, m/z value expressed in amu (atomic mass units), and at a specified instrument resolution. LR: Low Resolution ~300, MR: Mid Resolution ~4000 and HR: High Resolution ~10000. Ref: Method number 02-2061, National Laboratory for Environmental Testing, CCIW, Environment Canada, Burlington. Requested April 2003. | 5e-05 | Y | Y | Y | 2584 | 5/22/2003 | 9/21/2007 | 4/18/2018 9:28:02 AM | 2004-01-19 |  | UMULLER |  | Concentration | mg/L | Y | (fl. conc.) |  | Strontium Extractable (fl. conc.) | 7440-24-6 | Metals - Extractable; | Metals - Extractable; EMS code: SR-E | Chemical | Extractable | LAB
SR-E | mg/L | Strontium Extractable | F047 | 24hr HNO3 dil'n,ICP-MS | There is no acid digestion required for this analysis. A water sample is collected and preserved in 0.2 percent nitric acid. To analyse the extractable metals the sample is shaken, allowed to sit for at least twenty four hours, and a decanted portion is introduced to the spectrometer. When the dissolved metals are required, the samples must first be filtered through a 0.45 micron cellulose acetate filter paper and then followed by preservation in 0.2 percent nitric acid. It is important to note that different laboratories may have very different detection limits for some variables analyzed by ICP/MS. | 1e-05 | Y | Y | Y | 2676 | 4/16/1998 10:21:00 AM | 6/26/2003 | 9/11/2019 8:46:18 AM | 1997-10-20 | 6/3/1998 6:51:47 PM | NPEPPIN |  | Concentration | mg/L | Y | (fl. conc.) |  | Strontium Extractable (fl. conc.) | 7440-24-6 | Metals - Extractable; | Metals - Extractable; EMS code: SR-E | Chemical | Extractable | LAB
SR-E | mg/L | Strontium Extractable | F040 | Field Acid, lab filt, ICP-AES | Dissolved and extractable water samples are ready for analysis once collected. Total samples are autoclaved with HNO3 and hydrochloric acid (HCI) in capped Teflon vials. Analysis is performed via Inductively Coupled Argon Plasma Atomic Emission Spectrometry (ICP-AES)(simultaneous multielement analysis), Graphite Furnace Atomic Absorption Spectrometry (GF-AAS)(low level analysis) and/or Flame Atomic Absorption Spectrometry (F-AAS). | 0.001 | Y | Y | Y | 2350 | 8/5/1997 6:06:00 PM | 3/14/2008 | 9/11/2019 8:47:41 AM | 1997-05-08 | 6/3/1998 6:51:46 PM | MPAK |  | Concentration | mg/L | Y | (fl. conc.) |  | Strontium Extractable (fl. conc.) | 7440-24-6 | Metals - Extractable; | Metals - Extractable; EMS code: SR-E | Chemical | Extractable | LAB
SR-L | mg/L | Strontium Leachable | MEND | MEND (mod) EPA 300.1 (mod) | MEND (mod) EPA 300.1 (mod). Inorganic anions are analyzed by obtaining an extract from a soil sample that has been added in a 3:1 ratio of deionized water to soil; then leached and the extract filtered prior to analysis; which is then analyzed by Ion Chromatography with conductivity and/or UV detection. | 0.5 | Y | Y | Y | 0 |  |  |  | 2024-01-29 |  | RDGIBSON |  | Concentration | mg/L | Y | (fl. conc.) |  | Strontium Leachable (fl. conc.) | 7440-24-6 | Metals - Extractable; | Metals - Extractable; EMS code: SR-L | Chemical |  | LAB
SR-O | % | Strontium (as SrO) | ICP06 | ALS Minerals ME-ICP06 | ALS Minerals ME-ICP06. A prepared sample is fused with a lithium tetraborate/lithium metaborate flux and poured into a graphite crucible and fused in a furnace at 1000°C. The resulting melt is then cooled and dissolved in 100 mL of 4% HNO3/2% HCl followed by ICP-AES. | 0.01 | Y | Y | Y | 0 |  |  |  | 2024-01-29 |  | EMS |  | DimensionlessRatio | % | N | (ratio) |  | Strontium (as SrO) (ratio) |  | Metals; | Metals; EMS code: SR-O | Chemical |  | LAB
SR-T | ug/m3 | Strontium Total | 5312 | Tef:N/P Dig:ICP Analysis | A nitric/perchloric acid digestion if used to bring the metals into solution prior to ICP analysis. For Hi-Vol samples preliminary results are reported as mg/L and 24 hour loading results as ug/m3. | 0.1 | Y | Y | Y | 1191 | 11/7/1997 | 5/15/2000 | 11/18/2000 1:29:39 AM | 1996-10-13 | 6/3/1998 6:51:46 PM | EMS |  | AirConcentration | µg/m³ | Y | (air conc.) |  | Strontium Total (air conc.) | 7440-24-6 | Metals - Total; | Metals - Total; EMS code: SR-T | Chemical | Total | LAB
SR-T | ug/m3 | Strontium Total | 6040 | Aqua Regia Dig:ICP | Aqua regia digestion is used to bring the metals into solution prior to ICP analysis. For Hi-Vol samples, preliminary results are reported as mg/L and24 hr loadings as ug/m3 | 0.1 | Y | Y | Y | 437 | 2/13/2002 | 2/9/2005 | 2/23/2005 1:13:48 AM | 1998-05-05 | 6/3/1998 6:51:48 PM | NPEPPIN |  | AirConcentration | µg/m³ | Y | (air conc.) |  | Strontium Total (air conc.) | 7440-24-6 | Metals - Total; | Metals - Total; EMS code: SR-T | Chemical | Total | LAB
SR-T | ug/g | Strontium Total | FS01 | D&G,Nitric/HCl:ICP | Samples are dried at 60 C, then either sieved to -100mesh (150 um) and rolled to homogenize, or finely ground with a mortar and pestle or a shatter box laboratory mill and rolled (depending on the submitter's requirements). The samples are then weighed into Teflon digestion liners and digested with HNO3 and HCl in a microwave oven. Samples are cooled, volumized, and allowed to settle overnight before being decanted. Analysis is then performed via Inductively Coupled Argon Plasma Atomic Emission Spectrometry (ICP-AES) (Simultaneous multi element analysis). In ICP Spectrometry, the liquid sample is pumped to the nebulizer, where it is convertd into an aerosol. the nebulizer carrier gas flow than transports the wet aerosol past the impact bead, where the large droplets condense out and flow to the waste container. The remaining fine aerosol is carried into the plasma. There, the high temperature desolvates the aerosol and dissociates the sample material into individual atoms and ions, which are excited to emit light at wavelengths characteristic of the elements present. A diffraction grating in the spectrometer diffracts the various wavelengths of light through exit slits to the photo-multiplier tubes, which convert the light energy to electronic signals. | 0.2 | Y | Y | Y | 2642 | 8/30/1996 1:52:00 PM | 10/10/2007 | 5/24/2008 1:15:03 AM | 1996-11-20 | 6/3/1998 6:51:46 PM | MPAK |  | MassPerMass | mg/kg | Y | (mass conc.) |  | Strontium Total (mass conc.) | 7440-24-6 | Metals - Total; | Metals - Total; EMS code: SR-T | Chemical | Total | LAB
SR-T | ug/g | Strontium Total | 0190 | D&G:N/P Dig:ICP Anal. | Sample is dried and grind-mixed before nitric perchloric digestion followed by ICP analysis. | 1 | Y | Y | Y | 945 |  |  | 10/24/1996 3:06:16 PM | 1996-04-12 | 6/3/1998 6:51:40 PM | EMS |  | MassPerMass | mg/kg | Y | (mass conc.) |  | Strontium Total (mass conc.) | 7440-24-6 | Metals - Total; | Metals - Total; EMS code: SR-T | Chemical | Total | LAB
SR-T | ppm (S) | Strontium Total | MS41 | ALS Minerals ME-MS41 | ALS Minerals ME-MS41. A prepared sample is digested with aqua regia in a graphite heating block. After cooling; the resulting solution is diluted to with deionized water; mixed and analyzed by ICP-AES. Samples are then analyzed by ICP-MS for the remaining suite of elements. | 0.2 | Y | Y | Y | 0 |  |  |  | 2024-01-29 |  | EMS |  | DimensionlessRatio | ppm | N | (ratio) |  | Strontium Total (ratio) | 7440-24-6 | Metals - Total; | Metals - Total; EMS code: SR-T | Chemical | Total | LAB
SR-T | mg/L | Strontium Total | F121 | Acid,ICPMS | Acidified, no digestion, ICPMS | 5e-05 | Y | Y | Y | 8413 | 3/12/2010 7:57:00 PM | 8/15/2024 9:10:00 PM | 8/26/2024 8:06:03 AM | 2002-07-22 | 7/16/2012 3:31:04 PM | UMULLER | RDGIBSON | Concentration | mg/L | Y | (fl. conc.) |  | Strontium Total (fl. conc.) | 7440-24-6 | Metals - Total; | Metals - Total; EMS code: SR-T | Chemical | Total | LAB
SR-T | mg/L | Strontium Total | F048 | Acid dig,conc'n,nebul.,ICP-OES | This method provides an automated procedure for the Inductively Coupled Argon Plasma Optical Emission Spectrophotometric (ICP-OES) determination and quantitation of trace amounts of metals in surface, ground and precipitation waters after a manual acid digestion and concentration procedure and ultrasonic nebulisation. Organically-bound metals and metal complexes are converted to metal ions by digestion with nitric acid [method details says HCl]. To increase the signal detected by the ICP, the samples are concentrated 10 times. The liquid sample is pumped onto the piezoelectric transducer, where the ultrasonic action efficiently converts it to a fine, dense aerosol. The nebulizer carrier gas flow then transports the wet aerosol through a heated U-tube, where the solvent is vaporized. Subsequently, most of the solvent vapour is removed from the aerosol stream by passing the stream through a chilled condenser. The result is a dry, analyte-laden aerosol which is introduced to the plasma. There, the high temperature desolvates the aerosol and dissociates the sample material into individual atoms and ions, which are excited to emit light at wave lengths characteristic of the elements present. Dissolved samples are field filtered through 0.45 micron cellulose acetate filter prior to acid preservation. Samples for extractable analysis are acidified with 2 mL concentrated HNO3 per litre and allowed to stand a minimum of 24 hours before analysis. Clients providing these samples are advised to preserve samples at the collection site. | 0.0001 | Y | Y | Y | 6159 | 7/4/1997 2:36:00 PM | 4/23/2003 | 9/11/2019 8:46:38 AM | 1997-10-20 | 6/3/1998 6:51:47 PM | NPEPPIN |  | Concentration | mg/L | Y | (fl. conc.) |  | Strontium Total (fl. conc.) | 7440-24-6 | Metals - Total; | Metals - Total; EMS code: SR-T | Chemical | Total | LAB
SR-T | mg/L | Strontium Total | X145 | Dig, ICP | Dig; ICP | 0.001 | Y | Y | Y | 5090 |  |  | 8/26/2024 2:41:32 PM | 1996-04-12 | 6/3/1998 6:51:40 PM | EMS |  | Concentration | mg/L | Y | (fl. conc.) |  | Strontium Total (fl. conc.) | 7440-24-6 | Metals - Total; | Metals - Total; EMS code: SR-T | Chemical | Total | LAB
SR-T | mg/L | Strontium Total | X285 | HNO3 Dig:ICP Analysis | HNO3 Dig:ICP Analysis | 0.001 | Y | Y | Y | 594 | 8/10/1993 | 10/5/1993 | 4/30/2004 11:32:39 AM | 1998-05-22 | 6/3/1998 6:51:49 PM | NPEPPIN |  | Concentration | mg/L | Y | (fl. conc.) |  | Strontium Total (fl. conc.) | 7440-24-6 | Metals - Total; | Metals - Total; EMS code: SR-T | Chemical | Total | LAB
SR-T | mg/L | Strontium Total | X561 | Total metals: FA, HNO3, HR-ICPMS (No Prep)Prep) | Ultra trace metals in water are analyzed by high resolution inductively coupled plasma mass spectrometry (HR-ICPMS), modified from US EPA 200.8, (Revision 5.5). The detection limits provided can only be met for undigested samples. This procedure is intended for colorless, non-turbid, acid-preserved water samples (i.e. pristine water samples), having turbidity < 1 NTU and no odor. Where turbidity exceeds 1 NTU, and/or the sample is colored and has an odor, results may be biased low compared to true Total Metals concentrations. ALS recommends that turbidity analysis be requested on samples submitted for this test to aid with interpretation of results. | 5e-05 | Y | Y | Y | 2756 | 6/8/2015 | 10/5/2023 | 1/31/2024 2:37:41 PM | 2015-10-22 |  | RDGIBSON |  | Concentration | mg/L | Y | (fl. conc.) |  | Strontium Total (fl. conc.) | 7440-24-6 | Metals - Total; | Metals - Total; EMS code: SR-T | Chemical | Total | LAB
SR-T | mg/L | Strontium Total | 0040 | HNO3 Dig: ICP Analysis | HNO3 Dig: ICP Analysis | 0.001 | Y | Y | Y | 4269 |  |  | 6/24/2003 3:51:00 PM | 1996-04-12 | 6/3/1998 6:51:40 PM | EMS |  | Concentration | mg/L | Y | (fl. conc.) |  | Strontium Total (fl. conc.) | 7440-24-6 | Metals - Total; | Metals - Total; EMS code: SR-T | Chemical | Total | LAB
SR-T | ug/g wet | Strontium Total | MTIW | MET-WET-CCMS-MID-VA | Metals in Tissue by CRC ICPMS (WET) Methodology: This method is conducted following British Columbia Lab Manual method "Metals in Animal Tissue and Vegetation (Biota) - Prescriptive". Tissue samples are homogenized and sub-sampled prior to hotblock digestion with nitric and hydrochloric acids, in combination with addition of hydrogen peroxide. Instrumental analysis is by collision cell inductively coupled plasma - mass spectrometry (modified from EPA Method 6020A). Method Limitation: This method employs a strong acid/peroxide digestion, and is intended to provide a conservative estimate of bio-available metals. Near complete recoveries are achieved for most toxicologically important metals, but elements associated with recalcitrant minerals may be only partially recovered. | 0.02 | Y | Y | Y | 24 | 11/25/2018 | 12/6/2018 | 7/21/2021 2:55:08 PM | 2018-12-24 |  | EMS |  | MassPerMass | µg/g | Y | (mass conc.) | (wet) | Strontium Total (wet) (mass conc.) | 7440-24-6 | Metals - Total; | Metals - Total; EMS code: SR-T | Chemical | Total | LAB
SR-T | mg/L | Strontium Total | MWID | Metals in Water by ICPMS (digested) | Metals in Water by ICPMS (digested) Based on EPA 200.8 | 0.001 | Y | Y | Y | 36790 |  |  | 7/31/2024 1:46:41 PM | 2009-12-24 |  | RDGIBSON |  | Concentration | mg/L | Y | (fl. conc.) |  | Strontium Total (fl. conc.) | 7440-24-6 | Metals - Total; | Metals - Total; EMS code: SR-T | Chemical | Total | LAB
SR-T | mg/dm2/d | Strontium Total | DUST | HNO3 dig.(microwave or hot block),ICPMS | HNO3 dig.(microwave or hot block);ICPMS |  | Y | Y | Y | 420 | 7/19/2016 1:11:00 PM | 7/31/2024 | 7/31/2024 11:23:12 AM | 2008-02-07 |  | RDGIBSON |  | MassFlux | mg/dm²/d | Y | (mass flux) |  | Strontium Total (mass flux) | 7440-24-6 | Metals - Total; | Metals - Total; EMS code: SR-T | Chemical | Total | LAB
SR-T | ug/g | Strontium Total | 0192 | Hom, Nitric/Perchlor,ICP | Sample already in a dry state so does not require drying before homogenization and nitric perchlorid digestion followed by ICP analysis. | 1 | Y | Y | Y | 708 |  |  | 10/24/1996 12:50:25 PM | 1996-04-12 | 6/3/1998 6:51:40 PM | EMS |  | MassPerMass | mg/kg | Y | (mass conc.) |  | Strontium Total (mass conc.) | 7440-24-6 | Metals - Total; | Metals - Total; EMS code: SR-T | Chemical | Total | LAB
SR-T | mg/L | Strontium Total | 1349 | Tot Metals in Seawater by CRC ICPMS (BC) | Tot Metals in Seawater by CRC ICPMS (BC) | 0.01 | Y | Y | Y | 98 | 8/16/2019 | 12/5/2021 | 12/13/2021 12:34:17 PM | 2019-09-30 | 9/30/2019 3:11:58 PM | KATFRASE | KATFRASE | Concentration | mg/L | Y | (fl. conc.) |  | Strontium Total (fl. conc.) | 7440-24-6 | Metals - Total; | Metals - Total; EMS code: SR-T | Chemical | Total | LAB
SR-T | ppm (S) | Strontium Total | 6041 | Aqua Regia by ICPMS | These were sediment/soils digested with Aqua-Regia (Nitric Acid & Hydrochloric Acids) and then analysis by ICPMS. Results would be reported in mg/kg and/or %. | 0.01 | Y | Y | Y | 21 | 10/13/2019 | 8/5/2020 | 8/10/2020 3:52:07 PM | 2019-09-06 | 10/21/2019 11:56:31 AM | KATFRASE | KATFRASE | DimensionlessRatio | ppm | N | (ratio) |  | Strontium Total (ratio) | 7440-24-6 | Metals - Total; | Metals - Total; EMS code: SR-T | Chemical | Total | LAB
SR-T | ug/m3 | Strontium Total | F088 | Aliquot,HNO3/HCl dig, ICPMS,loading | Aliquot microwave digested with 3:1 HNO3/HCL. Analysis by ICPMS. For HiVol samples preliminary result are reported in mg/L and loading results are reported in ug/m3 |  | Y | Y | Y | 565 | 8/17/2000 10:46:00 AM | 12/14/2006 | 12/30/2006 1:01:38 AM | 2000-08-16 |  | NPEPPIN |  | AirConcentration | µg/m³ | Y | (air conc.) |  | Strontium Total (air conc.) | 7440-24-6 | Metals - Total; | Metals - Total; EMS code: SR-T | Chemical | Total | LAB
SR-T | mg/L | Strontium Total | F068 | HNO3dig,TeflonTube/Autoclave,ICP-MS | This method is applicable to determiniation of dissolved, extractable, and total recoverable metals in ground waters, surface waters, drinking waters, and wastewaters. Dissolved, extractable, and total recoverable sample fractions are acidified to pH <1.5 with analytical grade nitric acid (HNO3) to prevent the metals from adsorbing to the container material or precipitating. Extractable and total recoverable water samples are collected in 250 ml acid washed polyethylene bottles and are preserved with Nitric acid (0.5 ml acid for 100 mLs of sample). All preserved samples are stored at room temperature and can be held for up to 6 months. Unfiltered, un-preserved samples are stored at 4C for no more than 30 days. For dissolved and extractable water fractions, no sample preparation is necessary. Samples are simply poured in autosampler tubes (i.e. 15 ml centrifuge tubes), and placed on the autosampler and analyzed. Samples for total recoverable metals are digested with Nitric acid in special Teflon tubes in an autoclave prior to analyses. | 5e-06 | Y | Y | Y | 1775 | 6/1/1999 1:33:00 PM | 7/4/2008 9:48:00 AM | 7/30/2024 6:32:12 AM | 1999-06-28 |  | NPEPPIN |  | Concentration | mg/L | Y | (fl. conc.) |  | Strontium Total (fl. conc.) | 7440-24-6 | Metals - Total; | Metals - Total; EMS code: SR-T | Chemical | Total | LAB
SR-T | mg/L | Strontium Total | X170 | ICP | This code is a catch-all for a large variety of ICP methods including water and also solids digested into solutions as an intermediate result. Different procedures will have greatly different detection limits. (a) For air particulates, preliminary results are reported in mg/L and loading results as mg/dm2/d. (b) For Hi Vol samples preliminiary results are same as reported in mg/L and loading results as ug/m3. (c) Note that for Rain samples there is no filtration nor digestion steps. | 0.001 | Y | Y | Y | 3071 | 3/28/1996 6:03:00 PM | 11/7/2018 | 10/24/2022 10:58:25 AM | 1996-04-12 | 6/3/1998 6:51:40 PM | EMS |  | Concentration | mg/L | Y | (fl. conc.) |  | Strontium Total (fl. conc.) | 7440-24-6 | Metals - Total; | Metals - Total; EMS code: SR-T | Chemical | Total | LAB
SR-T | mg/L | Strontium Total | METN | MET-T-NP-CCMS-VA | Ultra trace metals in water are analyzed by CRC ICPMS, based on US EPA Method 6020B (July 2014). The detection limits provided can only be met for undigested samples. This procedure is intended for colorless, non-turbid, acid-preserved water samples (i.e. pristine water samples), having turbidity < 1 NTU and no odor. Where turbidity exceeds 1 NTU, and/or the sample is colored and has an odor, results may be biased low compared to true Total Metals concentrations. ALS recommends that turbidity analysis be requested on samples submitted for this test to aid with interpretation of results. Where turbidity is <1NTU, undigested metals are equivalent to total metals concentrations. | 5e-05 | Y | Y | Y | 3648 | 4/3/2012 | 8/19/2024 | 8/26/2024 1:03:06 PM | 2019-08-20 |  | KATFRASE |  | Concentration | mg/L | Y | (fl. conc.) |  | Strontium Total (fl. conc.) | 7440-24-6 | Metals - Total; | Metals - Total; EMS code: SR-T | Chemical | Total | LAB
SR-T | mg/L | Strontium Total | X349 | FA:HNO3:ICP | Samples are field-acidified with HNO3; HNO3 digestion; analysis by ICP/AES. | 0.001 | Y | Y | Y | 30 |  |  | 10/30/2014 8:31:18 AM | 1998-05-22 | 6/3/1998 6:51:49 PM | NPEPPIN |  | Concentration | mg/L | Y | (fl. conc.) |  | Strontium Total (fl. conc.) | 7440-24-6 | Metals - Total; | Metals - Total; EMS code: SR-T | Chemical | Total | LAB
SR-T | mg/L | Strontium Total | MSEA | Metals in Seawater (undigested) by (CRC) ICP-MS | This method describes the multi-elemental determination of trace elements by Collision/Reaction Cell (CRC) ICP-MS. The method measures ions produced by a radio-frequency inductively coupled plasma. Analyte species originating in a liquid are pneumatically nebulized and the resulting aerosol transported by argon gas into the plasma torch. The ions produced are entrained in the plasma gas and extracted, by means of a differentially pumped vacuum interface, into a mass spectrometer. The ions flow through a Collision/Reaction Cell (CRC) where depending on the mode of analysis either the ions passes through the cell unhindered (no gas - classical ICP-MS mode), or collisions/reactions chemistry takes place between the ions and a neutral gas such as hydrogen or helium. Any polyatomic ions which give interference on the analytes are removed. The ions are then sorted according to their mass-to-charge ratios by a quadrupole mass spectrometer. The ions are then sorted according to their mass-to-charge ratios by a quadrupole mass spectrometer. The ions transmitted through the quadrupole are quantified by a channel electron multiplier and the ion information processed by a data handling system. Although most polyatomic interference is removed, any additional interference must be assessed and valid corrections applied or the data flagged to indicate problems. Interference correction must include compensation for background ions contributed by the plasma gas, reagents, and constituents of the sample matrix. Instrument drift as well as suppressions or enhancements of instrument response caused by the sample matrix must be corrected for by the use of internal standards. | 2e-05 | Y | Y | Y | 441 | 8/12/2009 12:08:00 AM | 4/29/2024 | 6/7/2024 10:03:16 AM | 2009-09-08 |  | RDGIBSON |  | Concentration | mg/L | Y | (fl. conc.) |  | Strontium Total (fl. conc.) | 7440-24-6 | Metals - Total; | Metals - Total; EMS code: SR-T | Chemical | Total | LAB
SR-T | ug/g | Strontium Total | X551 | Metals in Biota (dry), total, acid digestion, ICP-SFMS | 0.1 gram of dried, homogenized biota is weighed in a microwave TFM vessel, and 8mL of Nitric Acid and 1mL of Hydrogen Peroxide are added. The sample is digested at 200 degrees C for 15 minutes in a closed vessel high pressure microwave oven. The sample is cooled, the volume adjusted to 100mL and introduced to an iductively coupled argon plasma-sector field mass spectrometer (ICP-SFMS) by pneumatic nebulization. Each respective element is measured at a specific mass to charge ratio, m/z value expressed in amu (atomic mass units), and at a specified instrument resolution. LR: Low Resolution ~300, MR: Mid Resolution ~4000 and HR: Hich Resolution ~10000. Ref: Method number 02-2705, National Laboratory for Environmental Testing, CCIW, Environment Canada, Burlington. Requested April 2003. | 0.05 | Y | Y | Y | 49 | 8/11/2015 | 10/30/2023 | 11/28/2023 2:21:33 PM | 2004-01-19 |  | UMULLER |  | MassPerMass | mg/kg | Y | (mass conc.) |  | Strontium Total (mass conc.) | 7440-24-6 | Metals - Total; | Metals - Total; EMS code: SR-T | Chemical | Total | LAB
SR-T | mg/L | Strontium Total | F038 | HNO3/HCl dig,ICP-AES | Dissolved and extractable water samples are ready for analysis once collected. Total samples are autoclaved with HNO3 and hydrochloric acid (HCI) in capped Teflon vials. Analysis is performed via Inductively Coupled Argon Plasma Atomic Emission Spectrometry (ICP-AES)(simultaneous multielement analysis), Graphite Furnace Atomic Absorption Spectrometry (GF-AAS) (low level analysis) and/or Flame Atomic Absorption Spectrometry (F-AAS). | 0.001 | Y | Y | Y | 5394 | 8/1/1996 6:04:00 PM | 6/25/2010 3:15:00 PM | 9/11/2019 8:47:43 AM | 1997-05-08 | 6/3/1998 6:51:46 PM | MPAK |  | Concentration | mg/L | Y | (fl. conc.) |  | Strontium Total (fl. conc.) | 7440-24-6 | Metals - Total; | Metals - Total; EMS code: SR-T | Chemical | Total | LAB
SR-T | ug/g | Strontium Total | PB02 | PBM,HNO3/HCl dig,dry60C,100mesh, | Performance Based Method. Samples are dried at 60 C and sieved using a .149 mm (100 mesh) sieve. Solid samples are digested with the combination of nitric and hydrochloric acids with a 1:1 ratio of acids. For environmental purposes, strong acid leaches are used routinely since they have the advantage of speed, low cost, and superio reproducibility. All metals trapped within the refractory matrix are unlikely to have environmental consequences. Instrumental analysis of extracts produced using this digestion method can be performed by a variety of techniques. |  | Y | Y | Y | 52 | 5/24/2001 3:42:00 PM | 11/26/2001 10:23:00 AM | 5/26/2003 2:28:38 PM | 2001-07-18 |  | NPEPPIN |  | MassPerMass | mg/kg | Y | (mass conc.) |  | Strontium Total (mass conc.) | 7440-24-6 | Metals - Total; | Metals - Total; EMS code: SR-T | Chemical | Total | LAB
SR-T | ug/g | Strontium Total | EQ02 | EQUIS (ug/g) | EQUIS (ug/g) | 1 | N | Y | Y | 760 |  |  | 10/23/1996 12:08:28 AM | 1996-04-12 | 6/3/1998 6:51:41 PM | EMS |  | MassPerMass | mg/kg | Y | (mass conc.) |  | Strontium Total (mass conc.) | 7440-24-6 | Metals - Total; | Metals - Total; EMS code: SR-T | Chemical | Total | LAB
SR-T | ug/g | Strontium Total | 0196 | D&G: LTA/HCl: ICP | D&G: LTA/HCl: ICP | 1 | Y | Y | Y | 27 |  |  | 10/24/1996 2:08:13 PM | 1996-04-12 | 6/3/1998 6:51:40 PM | EMS |  | MassPerMass | mg/kg | Y | (mass conc.) |  | Strontium Total (mass conc.) | 7440-24-6 | Metals - Total; | Metals - Total; EMS code: SR-T | Chemical | Total | LAB
SR-T | ug/g | Strontium Total | MTID | MET-DRY-CCMS-VA | Metals in Tissue by CRC ICPMS (DRY) Methodology: This method is conducted following British Columbia Lab Manual method "Metals in Animal Tissue and Vegetation (Biota) - Prescriptive". Tissue samples are homogenized and sub-sampled prior to hotblock digestion with nitric and hydrochloric acids, in combination with addition of hydrogen peroxide. Instrumental analysis is by collision cell inductively coupled plasma - mass spectrometry (modified from EPA Method 6020A). Method Limitation: This method employs a strong acid/peroxide digestion, and is intended to provide a conservative estimate of bio-available metals. Near complete recoveries are achieved for most toxicologically important metals, but elements associated with recalcitrant minerals may be only partially recovered. | 0.05 | Y | Y | Y | 122 | 8/13/2014 | 9/13/2022 | 10/26/2022 7:14:15 AM | 2015-03-09 | 3/11/2015 4:12:05 PM | RDGIBSON | EMS | MassPerMass | mg/kg | Y | (mass conc.) |  | Strontium Total (mass conc.) | 7440-24-6 | Metals - Total; | Metals - Total; EMS code: SR-T | Chemical | Total | LAB
SR-T | ug/g wet | Strontium Total | X560 | Metals in Biota (wet), total, acid digestion, ICP-MS | Five grams of wet, homogenized biota is weighed in a centrifuge tube, and 2.5 mL of Nitric Acid and 2.5 mL of Hydrochloric acid are added. The sample is digested at room temperature for one hour and then digested at 90-95 degrees C in a water bath or heat block for one more hour. The sample is cooled and the volume is made to 50 mL with DI. The digestate is then diluted at ten fold and introduced to a collision/reaction cell inductively coupled argon plasma mass spectrometer (CRC ICPMS) by pneumatic nebulization. Each respective element is measured at a specific mass to charge ratio, m/z value expressed in amu (atomic mass units). - Maxxam Analytics (BC). Added to EMS Dec 10, 2008. | 0.1 | Y | Y | Y | 541 | 11/15/2008 4:00:00 AM | 10/30/2023 | 11/28/2023 2:21:33 PM | 2008-12-11 | 3/11/2015 4:17:10 PM | RDGIBSON | EMS | MassPerMass | µg/g | Y | (mass conc.) | (wet) | Strontium Total (wet) (mass conc.) | 7440-24-6 | Metals - Total; | Metals - Total; EMS code: SR-T | Chemical | Total | LAB
SR-T | ppm (S) | Strontium Total | 6042 | Aqua Regia by ICPMS | These were sediment/soils digested with Aqua-Regia (Nitric Acid & Hydrochloric Acids) and then analysis by ICPMS. Results would be reported in ppm s. | 0.01 | Y | Y | Y | 71 | 5/14/2022 | 7/27/2024 | 7/29/2024 9:13:08 AM | 2022-06-28 |  | EMS |  | DimensionlessRatio | ppm | N | (ratio) |  | Strontium Total (ratio) | 7440-24-6 | Metals - Total; | Metals - Total; EMS code: SR-T | Chemical | Total | LAB
SR-T | ppm (S) | Strontium Total | MS81 | ALS Minerals: ME-MS81 | ALS Minerals: ME-MS81. A prepared sample (0.100 g) is added to lithium metaborate/lithium tetraborate flux; mixed well and fused in a furnace at 1025°C. The resulting melt is then cooled and dissolved in an acid mixture containing nitric; hydrochloric and hydrofluoric acids. This solution is then analyzed by inductively coupled plasma - mass spectrometry. | 0.1 | Y | Y | Y | 0 |  |  |  | 2024-01-29 |  | EMS |  | DimensionlessRatio | ppm | N | (ratio) |  | Strontium Total (ratio) | 7440-24-6 | Metals - Total; | Metals - Total; EMS code: SR-T | Chemical | Total | LAB
SR-T | mg/L | Strontium Total | X555 | Metals in water, total, "in-bottle" digestion, ICP-SFMS | The whole water sample collected in a HDPE/Teflon bottle, is de-capped in a Class 100 workstation, acidified to 2% (v/v) HNO3 and 1% (v/v) HCl and recapped. The sample is "in-bottle" digested in a convection oven, set at 60 degrees C for 16 hours. The sample is cooled and introduced to an inductively coupled argon plasma-sector field mass spectrometer (ICP-SFMS) by pneumatic nebulization. Each respective element is measured at a specific mass to charge ratio, m/z value expressed in amu (atomic mass units), and at resolution. LR: Low Resolution ~300, MR: Mid Resolution ~4000 and HR: High a specified instrument Resolution ~10000. Ref: Method number 02-2003, National Laboratory for Environmental Testing, CCIW, Environment Canada, Burlington. Requested April 2003. | 5e-05 | Y | Y | Y | 17336 | 5/20/2003 | 12/6/2023 2:08:00 AM | 2/14/2024 9:05:32 AM | 2004-01-19 |  | UMULLER |  | Concentration | mg/L | Y | (fl. conc.) |  | Strontium Total (fl. conc.) | 7440-24-6 | Metals - Total; | Metals - Total; EMS code: SR-T | Chemical | Total | LAB
SR-T | mg/L | Strontium Total | 0042 | GI HNO3 Dig: ICP | Sample aliquot taken from unfiltered unpreserved General Ions bottle; HNO3 digestion; ICP analysis. | 0.001 | Y | Y | Y | 6265 | 9/20/1996 | 7/5/2019 3:21:00 PM | 7/12/2019 6:49:13 PM | 1996-04-12 | 6/3/1998 6:51:40 PM | EMS |  | Concentration | mg/L | Y | (fl. conc.) |  | Strontium Total (fl. conc.) | 7440-24-6 | Metals - Total; | Metals - Total; EMS code: SR-T | Chemical | Total | LAB
SR-T | ug/g | Strontium Total | FB04 | Freeze Dry,Hom,Acid Dig,ICP | Sample moisture is determined by weighing samples before and after freeze-drying. Freeze-dried samples are ground in a Braun blender before being weighed into Teflon PFA liners and digested with concentrated HNO3, concentrated HCl, and 50% H2O2 in a microwave oven; which results in a complete sample dissolution. Samples are collected and volumized with Type I water. Analysis is performed via Inductively Coupled Argon Plasma Atomic Emission Spectrometry (ICP-AES) (simultaneous multielement analysis), Graphite Furnace Atomic Absorption Spectrometry (GF-AAS) (low level Cd and Pb analysis), Cold Vapor Atomic Fluorescence Spectrometry (CV-AFS) (Mercury analysis) and/or Flame Atomic Absorption Spectrometry (F-AAS) (K analysis). | 0.08 | Y | Y | Y | 241 | 8/22/1996 6:11:00 PM | 2/1/2002 2:30:00 PM | 6/23/2003 1:15:35 PM | 1997-02-27 | 6/3/1998 6:51:46 PM | MPAK |  | MassPerMass | mg/kg | Y | (mass conc.) |  | Strontium Total (mass conc.) | 7440-24-6 | Metals - Total; | Metals - Total; EMS code: SR-T | Chemical | Total | LAB
SR-T | mg/L | Strontium Total | F033 | Acid Dig,Centrifuge,ICPMS | For the determination of total recoverable elements in water or waste water, take 100 ml aliquot from a well mixed, acid preserved, sample containing not more than 0.25% (w/v) total solids and transfer to a 250 ml Griffin beaker. (If total solids are greater than 0.25% reduce the size of the aliquot by a proportionate amount). Add 1 ml of conc. nitric acid and 0.5 ml conc. hydrochloric acid. Heat on a hot plate until the volume has been reduced to 15-20ml, ensuring that the sample does not boil. Allow to cool and transfer to a 50 ml volumetric flask. Dilute to volume with ASTM type I water. Centrifuge the sample or allow to stand overnight to separate insoluble material. Prior analysis, pipette 20 ml into a 50 ml volumetric flask and dilute to volume with ASTM type I water. Analysis should be performed within seventy two hours of the completed preparation. Note that different laboratories may have very different field preservation, digestion procedures and detection limits for this method, so these should be documented. | 0.0001 | Y | Y | Y | 40346 | 10/15/1996 9:27:00 AM | 8/22/2024 3:24:00 AM | 8/25/2024 9:47:36 PM | 1997-01-16 | 6/3/1998 6:51:46 PM | MPAK |  | Concentration | mg/L | Y | (fl. conc.) |  | Strontium Total (fl. conc.) | 7440-24-6 | Metals - Total; | Metals - Total; EMS code: SR-T | Chemical | Total | LAB
SR-T | ug/g | Strontium Total | X230 | Hom,HNO3/HCl,ICP | The sample is dried, ground to pass 100 mesh, and digested with a nitric-hydrochloric acid mixture to solubilize the solid matter and to remove the organic material by oxidation and volatilization. The completed digests are submitted to metals instrumentation personnel for metals analysis by Inductively Coupled Atomic Emission Spectrophotometry (ICP-AES), or hydride analysis by atomic absorption. | 0.1 | Y | Y | Y | 120 | 2/12/2002 | 12/30/2005 | 2/3/2006 1:13:13 AM | 1996-04-12 | 6/3/1998 6:51:40 PM | EMS |  | MassPerMass | mg/kg | Y | (mass conc.) |  | Strontium Total (mass conc.) | 7440-24-6 | Metals - Total; | Metals - Total; EMS code: SR-T | Chemical | Total | LAB
SR-T | ug/g wet | Strontium Total | X552 | Metals in Biota (wet), total, acid digestion, ICP-SFMS | One gram of wet, homogenized biota is weighed in a microwave TFM vessel, and 8mL of Nitric Acid and 1mL of Hydrogen Peroxide are added. The sample is digested at 200 degrees C for 15 minutes in a closed vessel high pressure microwave oven. The sample is cooled, the volume adjusted to 100mL and introduced to an inductively coupled argon plasma-sector field mass spectrometer (ICP-SFMS) by pneumatic nebulization. Each respective element is measured at a specific mass to charge ratio, m/z value expressed in amu (atomic mass units), and at a specified instrument resolution. LR: Low Resolution ~300, MR: Mid Resolution ~4000. Ref: 02-2704, National Laboratory for Environmental Testing, CCIW, Environment Canada, Burlington. Requested April 2003. | 0.005 | Y | Y | Y | 37 | 8/9/2016 | 10/16/2023 | 11/28/2023 2:21:33 PM | 2004-01-19 | 4/19/2021 1:22:00 PM | UMULLER | EMS | MassPerMass | µg/g | Y | (mass conc.) | (wet) | Strontium Total (wet) (mass conc.) | 7440-24-6 | Metals - Total; | Metals - Total; EMS code: SR-T | Chemical | Total | LAB
SR-T | ug/g | Strontium Total | ICPM | Metals in Soil by CRC ICPMS | This analysis is carried out using procedures from CSR Analytical Method: "Strong Acid Leachable Metals (SALM) in Soil", BC Ministry of Environment, 26 June 2009, and procedures adapted from EPA Method 200.2. The sample is manually homogenized, dried at 60 degrees Celsius, sieved through a 63 um (230 mesh) sieve, and a representative subsample of the dry material is weighed. The sample is then digested at 95 degrees Celsius for 2 hours by block digester using concentrated nitric and hydrochloric acids. Instrumental analysis is by collision cell inductively coupled plasma - mass spectrometry (modifed from EPA Method 6020A). Method Limitation: This method is not a total digestion technique. It is a very strong acid digestion that is intended to dissolve those metals that may be environmentally available. By design, elements bound in silicate structures are not normally dissolved by this procedure as they are not usually mobile in the environment. | 0.1 | Y | Y | Y | 103 | 9/19/2017 | 9/24/2023 | 1/12/2024 11:16:52 AM | 2018-01-09 |  | JKROGH |  | MassPerMass | mg/kg | Y | (mass conc.) |  | Strontium Total (mass conc.) | 7440-24-6 | Metals - Total; | Metals - Total; EMS code: SR-T | Chemical | Total | LAB
SR-T | ppm (S) | Strontium Total | MS61 | ALS Minerals ME-MS61 | ALS Minerals ME-MS61. A prepared sample is digested with HNO3; HF; and HClO4 in a graphite heating block to near dryness. After cooling; the sample is boiled in diluted HCl. Bulk to volume with deionized water; mixed and analyzed by ICP-AES and ICP-MS for all elements. | 0.2 | Y | Y | Y | 0 |  |  |  | 2024-01-29 |  | EMS |  | DimensionlessRatio | ppm | N | (ratio) |  | Strontium Total (ratio) | 7440-24-6 | Metals - Total; | Metals - Total; EMS code: SR-T | Chemical | Total | LAB
SR-T | ug/g | Strontium Total | FS82 | PBM,dry60C,10mesh,SALM-ICP-MS | PBM;dry60C;10mesh;SALM-ICP-MS | 0.1 | Y | Y | Y | 3581 | 5/21/2002 | 7/24/2024 | 7/29/2024 5:17:09 PM | 2002-11-12 |  | UMULLER |  | MassPerMass | mg/kg | Y | (mass conc.) |  | Strontium Total (mass conc.) | 7440-24-6 | Metals - Total; | Metals - Total; EMS code: SR-T | Chemical | Total | LAB
SR-T | ug/g wet | Strontium Total | BPSY | MET-WET-BPSY-CCMS-VA | Metals in Animal Tissue and Vegetation (Biota) - Prescriptive. EPA 200.3/6020B This method is conducted following British Columbia Lab Manual method "Metals in Animal Tissue and Vegetation (Biota) - Prescriptive". Tissue samples are digested with nitric and hydrochloric acids; in combination with addition of hydrogen peroxide. Instrumental analysis is by collision cell inductively coupled plasma - mass spectrometry (modified from EPA Method 6020B). | 0.02 | Y | Y | Y | 2 | 11/7/2019 | 11/7/2019 | 6/10/2021 1:57:10 PM | 2021-06-04 |  | EMS |  | MassPerMass | µg/g | Y | (mass conc.) | (wet) | Strontium Total (wet) (mass conc.) | 7440-24-6 | Metals - Total; | Metals - Total; EMS code: SR-T | Chemical | Total | LAB
SR-T | mg/L | Strontium Total | F082 | Aliquot,HNO3/HCL dig,ICPMS | Aliquot microwave digested with 3:1 HNO3/HCL. Analysis by ICPMS. For HiVol samples preliminary result are reported in mg/L and loading results are reported in ug/m3 |  | Y | Y | Y | 27224 | 7/26/2005 7:04:00 AM | 8/23/2024 | 8/26/2024 4:53:05 PM | 2000-06-13 |  | NPEPPIN |  | Concentration | mg/L | Y | (fl. conc.) |  | Strontium Total (fl. conc.) | 7440-24-6 | Metals - Total; | Metals - Total; EMS code: SR-T | Chemical | Total | LAB
ST-CAL | % | Sulfide (as S); total minus carbonate leach | GRA06 | ALS Minerals S-GRA06 | ALS Minerals S-GRA06. A preppared sample is heated with diluted Na₂CO₃ and then filtered. The sulfate in the filtrate is precipitated with BaCl2 in a dilute HCl. The BaSO4 precipitate is filtered; ignited; weighed and the Sulfate Sulfur is determined (as S). | 0.01 | Y | Y | Y | 0 |  |  |  | 2024-01-29 |  | EMS |  | DimensionlessRatio | % | N | (ratio) |  | Sulfide (as S); total minus carbonate leach (ratio) |  | Major Ions; | Major Ions; EMS code: ST-CAL | Chemical |  | LAB
ST-CAL | % | Sulfide (as S); total minus carbonate leach | CAL06 | ALS Minerals S-CAL06 | ALS Minerals S-CAL06. Sulfide Sulfur (as S) is calculated by subtracting the Carbonate Leachable Sulfate Sulfur (as S) from the Total Sulfur (as S) obtained from combustion. | 0.01 | Y | Y | Y | 0 |  |  |  | 2024-01-29 |  | EMS |  | DimensionlessRatio | % | N | (ratio) |  | Sulfide (as S); total minus carbonate leach (ratio) |  | Major Ions; | Major Ions; EMS code: ST-CAL | Chemical |  | LAB
ST01 | mg/L | Campesterol | F111 | H2SO4,SPE extr'n,conc'n,HRGC/LRMS | Effluent samples are acidified with Sulphuric Acid and extracted using C-18 Solid Phase Extraction (SPE) disks. Sterols are eluted with Acetone and concentrated. After concentration, the extract is acetylated using Pyridine and Acetic Anhydride. A sterol Internal Standard is added and the extract is analysed using High Resolution Gas Chromatography/Low Resolution Mass Spectrometry (HRGC/LRMS). MS data is acquired in Selected Ion Mode (SIM) and quatitation is by multilevel Internal Standards. | 5e-06 | Y | Y | Y | 10 | 6/8/2001 11:55:00 AM | 1/15/2002 1:19:00 PM | 5/20/2003 10:51:16 AM | 2001-05-22 |  | NPEPPIN |  | Concentration | mg/L | Y | (fl. conc.) |  | Campesterol (fl. conc.) | 474-62-4 | Organics; | Organics; EMS code: ST01 | Chemical |  | LAB
ST02 | mg/L | Cholesterol | F111 | H2SO4,SPE extr'n,conc'n,HRGC/LRMS | Effluent samples are acidified with Sulphuric Acid and extracted using C-18 Solid Phase Extraction (SPE) disks. Sterols are eluted with Acetone and concentrated. After concentration, the extract is acetylated using Pyridine and Acetic Anhydride. A sterol Internal Standard is added and the extract is analysed using High Resolution Gas Chromatography/Low Resolution Mass Spectrometry (HRGC/LRMS). MS data is acquired in Selected Ion Mode (SIM) and quatitation is by multilevel Internal Standards. | 5e-06 | Y | Y | Y | 10 | 6/8/2001 11:55:00 AM | 1/15/2002 1:19:00 PM | 5/20/2003 10:51:16 AM | 2001-05-22 |  | NPEPPIN |  | Concentration | mg/L | Y | (fl. conc.) |  | Cholesterol (fl. conc.) | 57-88-5 | Organics; | Organics; EMS code: ST02 | Chemical |  | LAB
ST03 | mg/L | Coprostanol | F111 | H2SO4,SPE extr'n,conc'n,HRGC/LRMS | Effluent samples are acidified with Sulphuric Acid and extracted using C-18 Solid Phase Extraction (SPE) disks. Sterols are eluted with Acetone and concentrated. After concentration, the extract is acetylated using Pyridine and Acetic Anhydride. A sterol Internal Standard is added and the extract is analysed using High Resolution Gas Chromatography/Low Resolution Mass Spectrometry (HRGC/LRMS). MS data is acquired in Selected Ion Mode (SIM) and quatitation is by multilevel Internal Standards. | 5e-07 | Y | Y | Y | 10 | 6/8/2001 11:55:00 AM | 1/15/2002 1:19:00 PM | 5/20/2003 10:51:17 AM | 2001-05-22 |  | NPEPPIN |  | Concentration | mg/L | Y | (fl. conc.) |  | Coprostanol (fl. conc.) | 360-68-9 | Organics; | Organics; EMS code: ST03 | Chemical |  | LAB
ST04 | mg/L | Desmosterol | F111 | H2SO4,SPE extr'n,conc'n,HRGC/LRMS | Effluent samples are acidified with Sulphuric Acid and extracted using C-18 Solid Phase Extraction (SPE) disks. Sterols are eluted with Acetone and concentrated. After concentration, the extract is acetylated using Pyridine and Acetic Anhydride. A sterol Internal Standard is added and the extract is analysed using High Resolution Gas Chromatography/Low Resolution Mass Spectrometry (HRGC/LRMS). MS data is acquired in Selected Ion Mode (SIM) and quatitation is by multilevel Internal Standards. | 5e-06 | Y | Y | Y | 10 | 6/8/2001 11:55:00 AM | 1/15/2002 1:19:00 PM | 5/20/2003 10:51:18 AM | 2001-05-22 |  | NPEPPIN |  | Concentration | mg/L | Y | (fl. conc.) |  | Desmosterol (fl. conc.) | 313-04-2 | Organics; | Organics; EMS code: ST04 | Chemical |  | LAB
ST05 | mg/L | Dihydrocholesterol | F111 | H2SO4,SPE extr'n,conc'n,HRGC/LRMS | Effluent samples are acidified with Sulphuric Acid and extracted using C-18 Solid Phase Extraction (SPE) disks. Sterols are eluted with Acetone and concentrated. After concentration, the extract is acetylated using Pyridine and Acetic Anhydride. A sterol Internal Standard is added and the extract is analysed using High Resolution Gas Chromatography/Low Resolution Mass Spectrometry (HRGC/LRMS). MS data is acquired in Selected Ion Mode (SIM) and quatitation is by multilevel Internal Standards. | 5e-06 | Y | Y | Y | 10 | 6/8/2001 11:55:00 AM | 1/15/2002 1:19:00 PM | 5/20/2003 10:51:18 AM | 2001-05-22 |  | NPEPPIN |  | Concentration | mg/L | Y | (fl. conc.) |  | Dihydrocholesterol (fl. conc.) | 80-97-7 | Organics; | Organics; EMS code: ST05 | Chemical |  | LAB
ST06 | mg/L | Epicoprostanol | F111 | H2SO4,SPE extr'n,conc'n,HRGC/LRMS | Effluent samples are acidified with Sulphuric Acid and extracted using C-18 Solid Phase Extraction (SPE) disks. Sterols are eluted with Acetone and concentrated. After concentration, the extract is acetylated using Pyridine and Acetic Anhydride. A sterol Internal Standard is added and the extract is analysed using High Resolution Gas Chromatography/Low Resolution Mass Spectrometry (HRGC/LRMS). MS data is acquired in Selected Ion Mode (SIM) and quatitation is by multilevel Internal Standards. | 5e-06 | Y | Y | Y | 10 | 6/8/2001 11:55:00 AM | 1/15/2002 1:19:00 PM | 5/20/2003 10:51:20 AM | 2001-05-22 |  | NPEPPIN |  | Concentration | mg/L | Y | (fl. conc.) |  | Epicoprostanol (fl. conc.) | 516-92-7 | Organics; | Organics; EMS code: ST06 | Chemical |  | LAB
ST07 | mg/L | Equilin | F111 | H2SO4,SPE extr'n,conc'n,HRGC/LRMS | Effluent samples are acidified with Sulphuric Acid and extracted using C-18 Solid Phase Extraction (SPE) disks. Sterols are eluted with Acetone and concentrated. After concentration, the extract is acetylated using Pyridine and Acetic Anhydride. A sterol Internal Standard is added and the extract is analysed using High Resolution Gas Chromatography/Low Resolution Mass Spectrometry (HRGC/LRMS). MS data is acquired in Selected Ion Mode (SIM) and quatitation is by multilevel Internal Standards. | 5e-06 | Y | Y | Y | 10 | 6/8/2001 11:55:00 AM | 1/15/2002 1:19:00 PM | 5/20/2003 10:51:21 AM | 2001-05-22 |  | NPEPPIN |  | Concentration | mg/L | Y | (fl. conc.) |  | Equilin (fl. conc.) | 474-86-2 | Organics; | Organics; EMS code: ST07 | Chemical |  | LAB
ST08 | mg/L | Equol | F111 | H2SO4,SPE extr'n,conc'n,HRGC/LRMS | Effluent samples are acidified with Sulphuric Acid and extracted using C-18 Solid Phase Extraction (SPE) disks. Sterols are eluted with Acetone and concentrated. After concentration, the extract is acetylated using Pyridine and Acetic Anhydride. A sterol Internal Standard is added and the extract is analysed using High Resolution Gas Chromatography/Low Resolution Mass Spectrometry (HRGC/LRMS). MS data is acquired in Selected Ion Mode (SIM) and quatitation is by multilevel Internal Standards. | 0.0002 | Y | Y | Y | 10 | 6/8/2001 11:55:00 AM | 1/15/2002 1:19:00 PM | 5/20/2003 10:51:21 AM | 2001-05-22 |  | NPEPPIN |  | Concentration | mg/L | Y | (fl. conc.) |  | Equol (fl. conc.) | 531-95-3 | Organics; | Organics; EMS code: ST08 | Chemical |  | LAB
ST09 | mg/L | 17alpha-Estradiol | F111 | H2SO4,SPE extr'n,conc'n,HRGC/LRMS | Effluent samples are acidified with Sulphuric Acid and extracted using C-18 Solid Phase Extraction (SPE) disks. Sterols are eluted with Acetone and concentrated. After concentration, the extract is acetylated using Pyridine and Acetic Anhydride. A sterol Internal Standard is added and the extract is analysed using High Resolution Gas Chromatography/Low Resolution Mass Spectrometry (HRGC/LRMS). MS data is acquired in Selected Ion Mode (SIM) and quatitation is by multilevel Internal Standards. | 5e-06 | Y | Y | Y | 10 | 6/8/2001 11:55:00 AM | 1/15/2002 1:19:00 PM | 5/20/2003 10:51:23 AM | 2001-05-22 |  | NPEPPIN |  | Concentration | mg/L | Y | (fl. conc.) |  | 17alpha-Estradiol (fl. conc.) | 57-91-0 | Organics; | Organics; EMS code: ST09 | Chemical |  | LAB
ST10 | mg/L | 17beta-Estradiol | F111 | H2SO4,SPE extr'n,conc'n,HRGC/LRMS | Effluent samples are acidified with Sulphuric Acid and extracted using C-18 Solid Phase Extraction (SPE) disks. Sterols are eluted with Acetone and concentrated. After concentration, the extract is acetylated using Pyridine and Acetic Anhydride. A sterol Internal Standard is added and the extract is analysed using High Resolution Gas Chromatography/Low Resolution Mass Spectrometry (HRGC/LRMS). MS data is acquired in Selected Ion Mode (SIM) and quatitation is by multilevel Internal Standards. | 5e-06 | Y | Y | Y | 10 | 6/8/2001 11:55:00 AM | 1/15/2002 1:19:00 PM | 5/20/2003 10:51:24 AM | 2001-05-22 |  | NPEPPIN |  | Concentration | mg/L | Y | (fl. conc.) |  | 17beta-Estradiol (fl. conc.) | 50-28-2 | Organics; | Organics; EMS code: ST10 | Chemical |  | LAB
ST11 | mg/L | Estriol | F111 | H2SO4,SPE extr'n,conc'n,HRGC/LRMS | Effluent samples are acidified with Sulphuric Acid and extracted using C-18 Solid Phase Extraction (SPE) disks. Sterols are eluted with Acetone and concentrated. After concentration, the extract is acetylated using Pyridine and Acetic Anhydride. A sterol Internal Standard is added and the extract is analysed using High Resolution Gas Chromatography/Low Resolution Mass Spectrometry (HRGC/LRMS). MS data is acquired in Selected Ion Mode (SIM) and quatitation is by multilevel Internal Standards. | 2e-05 | Y | Y | Y | 10 | 6/8/2001 11:55:00 AM | 1/15/2002 1:19:00 PM | 5/20/2003 10:51:26 AM | 2001-05-22 |  | NPEPPIN |  | Concentration | mg/L | Y | (fl. conc.) |  | Estriol (fl. conc.) |  | Organics; | Organics; EMS code: ST11 | Chemical |  | LAB
ST12 | mg/L | Estrone | F111 | H2SO4,SPE extr'n,conc'n,HRGC/LRMS | Effluent samples are acidified with Sulphuric Acid and extracted using C-18 Solid Phase Extraction (SPE) disks. Sterols are eluted with Acetone and concentrated. After concentration, the extract is acetylated using Pyridine and Acetic Anhydride. A sterol Internal Standard is added and the extract is analysed using High Resolution Gas Chromatography/Low Resolution Mass Spectrometry (HRGC/LRMS). MS data is acquired in Selected Ion Mode (SIM) and quatitation is by multilevel Internal Standards. | 5e-06 | Y | Y | Y | 10 | 6/8/2001 11:55:00 AM | 1/15/2002 1:19:00 PM | 5/20/2003 10:51:28 AM | 2001-05-22 |  | NPEPPIN |  | Concentration | mg/L | Y | (fl. conc.) |  | Estrone (fl. conc.) |  | Organics; | Organics; EMS code: ST12 | Chemical |  | LAB
ST13 | mg/L | 17alpha-Ethynylestradiol | F111 | H2SO4,SPE extr'n,conc'n,HRGC/LRMS | Effluent samples are acidified with Sulphuric Acid and extracted using C-18 Solid Phase Extraction (SPE) disks. Sterols are eluted with Acetone and concentrated. After concentration, the extract is acetylated using Pyridine and Acetic Anhydride. A sterol Internal Standard is added and the extract is analysed using High Resolution Gas Chromatography/Low Resolution Mass Spectrometry (HRGC/LRMS). MS data is acquired in Selected Ion Mode (SIM) and quatitation is by multilevel Internal Standards. | 2e-05 | Y | Y | Y | 10 | 6/8/2001 11:55:00 AM | 1/15/2002 1:19:00 PM | 5/20/2003 10:51:28 AM | 2001-05-22 |  | NPEPPIN |  | Concentration | mg/L | Y | (fl. conc.) |  | 17alpha-Ethynylestradiol (fl. conc.) | 57-63-6 | Organics; | Organics; EMS code: ST13 | Chemical |  | LAB
ST14 | mg/L | Mestranol | F111 | H2SO4,SPE extr'n,conc'n,HRGC/LRMS | Effluent samples are acidified with Sulphuric Acid and extracted using C-18 Solid Phase Extraction (SPE) disks. Sterols are eluted with Acetone and concentrated. After concentration, the extract is acetylated using Pyridine and Acetic Anhydride. A sterol Internal Standard is added and the extract is analysed using High Resolution Gas Chromatography/Low Resolution Mass Spectrometry (HRGC/LRMS). MS data is acquired in Selected Ion Mode (SIM) and quatitation is by multilevel Internal Standards. | 5e-06 | Y | Y | Y | 10 | 6/8/2001 11:55:00 AM | 1/15/2002 1:19:00 PM | 5/20/2003 10:51:29 AM | 2001-05-22 |  | NPEPPIN |  | Concentration | mg/L | Y | (fl. conc.) |  | Mestranol (fl. conc.) |  | Organics; | Organics; EMS code: ST14 | Chemical |  | LAB
ST15 | mg/L | Norgestrel | F111 | H2SO4,SPE extr'n,conc'n,HRGC/LRMS | Effluent samples are acidified with Sulphuric Acid and extracted using C-18 Solid Phase Extraction (SPE) disks. Sterols are eluted with Acetone and concentrated. After concentration, the extract is acetylated using Pyridine and Acetic Anhydride. A sterol Internal Standard is added and the extract is analysed using High Resolution Gas Chromatography/Low Resolution Mass Spectrometry (HRGC/LRMS). MS data is acquired in Selected Ion Mode (SIM) and quatitation is by multilevel Internal Standards. | 0.0002 | Y | Y | Y | 10 | 6/8/2001 11:55:00 AM | 1/15/2002 1:19:00 PM | 5/20/2003 10:51:29 AM | 2001-05-22 |  | NPEPPIN |  | Concentration | mg/L | Y | (fl. conc.) |  | Norgestrel (fl. conc.) |  | Organics; | Organics; EMS code: ST15 | Chemical |  | LAB
ST16 | mg/L | Norethindrone | F111 | H2SO4,SPE extr'n,conc'n,HRGC/LRMS | Effluent samples are acidified with Sulphuric Acid and extracted using C-18 Solid Phase Extraction (SPE) disks. Sterols are eluted with Acetone and concentrated. After concentration, the extract is acetylated using Pyridine and Acetic Anhydride. A sterol Internal Standard is added and the extract is analysed using High Resolution Gas Chromatography/Low Resolution Mass Spectrometry (HRGC/LRMS). MS data is acquired in Selected Ion Mode (SIM) and quatitation is by multilevel Internal Standards. | 0.0002 | Y | Y | Y | 10 | 6/8/2001 11:55:00 AM | 1/15/2002 1:19:00 PM | 5/20/2003 10:51:32 AM | 2001-05-22 |  | NPEPPIN |  | Concentration | mg/L | Y | (fl. conc.) |  | Norethindrone (fl. conc.) |  | Organics; | Organics; EMS code: ST16 | Chemical |  | LAB
ST17 | mg/L | beta- Sitosterol | F111 | H2SO4,SPE extr'n,conc'n,HRGC/LRMS | Effluent samples are acidified with Sulphuric Acid and extracted using C-18 Solid Phase Extraction (SPE) disks. Sterols are eluted with Acetone and concentrated. After concentration, the extract is acetylated using Pyridine and Acetic Anhydride. A sterol Internal Standard is added and the extract is analysed using High Resolution Gas Chromatography/Low Resolution Mass Spectrometry (HRGC/LRMS). MS data is acquired in Selected Ion Mode (SIM) and quatitation is by multilevel Internal Standards. | 5e-06 | Y | Y | Y | 10 | 6/8/2001 11:55:00 AM | 1/15/2002 1:19:00 PM | 5/20/2003 10:51:34 AM | 2001-05-22 |  | NPEPPIN |  | Concentration | mg/L | Y | (fl. conc.) |  | beta- Sitosterol (fl. conc.) |  | Organics; | Organics; EMS code: ST17 | Chemical |  | LAB
ST18 | mg/L | Stigmasterol | F111 | H2SO4,SPE extr'n,conc'n,HRGC/LRMS | Effluent samples are acidified with Sulphuric Acid and extracted using C-18 Solid Phase Extraction (SPE) disks. Sterols are eluted with Acetone and concentrated. After concentration, the extract is acetylated using Pyridine and Acetic Anhydride. A sterol Internal Standard is added and the extract is analysed using High Resolution Gas Chromatography/Low Resolution Mass Spectrometry (HRGC/LRMS). MS data is acquired in Selected Ion Mode (SIM) and quatitation is by multilevel Internal Standards. | 5e-06 | Y | Y | Y | 10 | 6/8/2001 11:55:00 AM | 1/15/2002 1:19:00 PM | 5/20/2003 10:51:35 AM | 2001-05-22 |  | NPEPPIN |  | Concentration | mg/L | Y | (fl. conc.) |  | Stigmasterol (fl. conc.) |  | Organics; | Organics; EMS code: ST18 | Chemical |  | LAB
STHCLL | % | Sulfide (as S); total minus HCl leach | GRA06A | ALS Minerals S-GRA06A | ALS Minerals S-GRA06A. A prepared sample is heated with dilute HCl and then filtered. The sulfate in the filtrate is precipitated with BaCl2 in a dilute HCl. The BaSO4 precipitate is filtered; ignited; weighed and the Sulfate Sulfur is determined (as S). | 0.01 | Y | Y | Y | 0 |  |  |  | 2024-01-29 |  | EMS |  | DimensionlessRatio | % | N | (ratio) |  | Sulfide (as S); total minus HCl leach (ratio) |  | Major Ions; | Major Ions; EMS code: STHCLL | Chemical |  | LAB
STRY | mg/L | Strychine CAS: 57-24-9 | XMSP | HPLC and Tri. Quadpole MS/MS | EPA 1694 An aliquot of filtered sample containing 5% organic is injected directly. Each compound is separated by reversed phase HPLC. Analyzed by Electro Spray Ionization MS/MS detection using Triple Quadrupole MS/MS detector. Quantiitation with internal or external standard. |  | Y | Y | Y | 6 | 5/21/2021 | 4/13/2022 | 5/10/2022 9:23:05 AM | 2022-05-03 |  | EMS |  | Concentration | mg/L | Y | (fl. conc.) |  | Strychine (fl. conc.) | 57-24-9 | Pesticides; | Pesticides; EMS code: STRY | Chemical |  | LAB
SU2- | ug/m3 | Sulfur:Bivalent | EQ40 | EQUIS (mg/m3) | EQUIS (mg/m3) | 0 | N | Y | Y | 274 |  |  | 10/22/1996 7:05:44 PM | 1996-04-12 | 6/3/1998 6:51:38 PM | EMS |  | AirConcentration | µg/m³ | Y | (air conc.) |  | Sulfur:Bivalent (air conc.) |  |  | EMS code: SU2- | Chemical |  | LAB
SU2- | ug/g | Sulfur:Bivalent | EQ02 | EQUIS (ug/g) | EQUIS (ug/g) | 0 | N | Y | Y | 488 |  |  | 10/22/1996 7:37:46 PM | 1996-04-12 | 6/3/1998 6:51:45 PM | EMS |  | MassPerMass | mg/kg | Y | (mass conc.) |  | Sulfur:Bivalent (mass conc.) |  |  | EMS code: SU2- | Chemical |  | LAB
SUFI | mg/dm2/d | Sulfation Index | EQ15 | EQUIS (mg/dm2/d) | EQUIS (mg/dm2/d) | 0.02 | N | Y | Y | 2578 |  |  | 10/23/1996 12:06:46 AM | 1996-04-12 | 6/3/1998 6:51:38 PM | EMS |  | MassFlux | mg/dm²/d | Y | (mass flux) |  | Sulfation Index (mass flux) |  | Inorganics; | Inorganics; EMS code: SUFI | Chemical |  | LAB
SUFI | mg/dm2/d | Sulfation Index | X143 | Selective Ion Electrode | The analyte is determined potentiometrically by direct measurement using a selective ion electrode with standard reference electrode and either a meter with a direct concentration scale or a pH meter with an expanded mV scale. | 0.02 | Y | Y | Y | 2392 | 10/31/2000 | 1/2/2004 | 1/20/2004 12:08:16 PM | 1996-04-12 | 6/3/1998 6:51:38 PM | EMS |  | MassFlux | mg/dm²/d | Y | (mass flux) |  | Sulfation Index (mass flux) |  | Inorganics; | Inorganics; EMS code: SUFI | Chemical |  | LAB
SUFI | mg/dm2/d | Sulfation Index | 5001 | Turbidimetric Analysis | A lead oxide plate is exposed to ambient air for a period of 30 days. During this time sulfur dioxide and sulfur trioxide in the atmosphere are collected by lead oxide as the result of both oxidation and absorption processes and converted to lead sulfate. The lead oxide plates undergo extractions and the sulfate ion is converted to a barium sulfate suspension under controlled conditions. The resulting turbidity is determined using a spectrophotometer and compared to a curve prepared from standard sulfate solutions. Preliminary results are reported as mg/l, and loading results as mg/dm2/d. | 0.02 | Y | Y | Y | 1838 |  |  | 2/13/1998 2:24:45 AM | 1996-04-12 | 6/3/1998 6:51:45 PM | EMS |  | MassFlux | mg/dm²/d | Y | (mass flux) |  | Sulfation Index (mass flux) |  | Inorganics; | Inorganics; EMS code: SUFI | Chemical |  | LAB
SULF | mg/L | Sulfolane | SULF | Sulfolane by LC/MS-MS | SULFOLANE-DI-LCMS-WT Methodology: Water sample is filtered, spiked with internal standard and analyzed by direct injection using a HPLC Triple Quadrupole MS/MS with Electro Spray Ionization. | 0.005 | Y | Y | Y | 16 | 6/5/2018 | 9/4/2018 | 12/10/2019 3:45:08 PM | 2018-09-17 |  | RDGIBSON |  | Concentration | mg/L | Y | (fl. conc.) |  | Sulfolane (fl. conc.) | 126-33-0 | Pesticides; | Pesticides; EMS code: SULF | Chemical |  | LAB
T001 | mg/L | Toluene | F007 | Static Headspace:HRGC/LRMS | Static headspace analysis is a method for the determination of volatile compounds in liquids and/or solids by analysing the vapour phase of a sample in a closed system; in this case a sealed vial. Static Headspace is sometimes referred to as Equilibrium Headspace because the vapour phase above the sample in a sealed vial is brought into thermodynamic equilibrium with the sample prior to analysis. Volatile organic compounds are thermally extracted from the sample matrix by Static Headspace sampling and introduced directly to a thick-film megabore GC column. Compounds eluting from the GC column are introduced into a Mass Spectrometer and target compounds are identified by a comparison of both retention times and mass spectra with those of reference standards stored in a computerized data system. Reference retention times and mass spectra for analytes are obtained by the measurement of calibration standards under precisely the same extraction conditions used for samples. The concentration of each identified component is measured by relating the response of the sample quantitation mass ions to the response of the quantitation mass ions of a series of external standards. A deuterated surrogate recovery standard is added to provide information on matrix effects. | 0.0005 | Y | Y | Y | 1409 | 4/21/1998 2:34:00 PM | 8/14/2024 | 8/16/2024 4:05:21 PM | 1996-09-17 | 6/3/1998 6:51:45 PM | MPAK |  | Concentration | mg/L | Y | (fl. conc.) |  | Toluene (fl. conc.) | 108-88-3 | Volatile Organic Compounds (VOCs); | Volatile Organic Compounds (VOCs); EMS code: T001 | Chemical |  | LAB
T001 | ug/g | Toluene | FS04 | Static Headspace:HRGC/LRMS | Static Headspace anaylsis is a method for the determination of volatile compounds in liquids and/or solids by analysing the vapour phase of a sample in a closed system; in this case a sealed vial. Static Headspace is sometimes referred to as Equilibrium Headspace because the vapour phase above the sample in a sealed vial is brought into thermodynamic equilibrium with the sample prior to analysis. Volatile organic compounds are thermally extracted from the sample matrix by Static Headspace sampling and introduced directly to a thick-film megabore GC column. Compounds eluting from the GC column are introduced into a Mass Spectrometer and target compounds are identified by a comparison of both retention times and mass spectra with those of reference standards stored in a computerized data system. Reference retention times and mass spectra for analytes are obtained by the measurement of calibration standards under precisely the same extraction conditions used for samples. The concentration of each identified component is measured by relating the response of the sample quantitation mass ions to the response of the quantitation mass ions of a series of external standards. A deuterated surrogate recovery standard is added to provide information on matrix effects. | 0.01 | Y | Y | Y | 89 | 7/2/1999 2:40:00 PM | 3/7/2013 3:50:00 AM | 3/19/2013 1:15:56 PM | 1996-09-17 | 6/3/1998 6:51:45 PM | MPAK |  | MassPerMass | mg/kg | Y | (mass conc.) |  | Toluene (mass conc.) | 108-88-3 | Volatile Organic Compounds (VOCs); | Volatile Organic Compounds (VOCs); EMS code: T001 | Chemical |  | LAB
T001 | mg/L | Toluene | 8260D | EPA 8260D (mod) | EPA 8260D (mod) VOLATILE ORGANIC COMPOUNDS BY GAS CHROMATOGRAPHY/MASS SPECTROMETRY This method is used to determine volatile organic compounds (VOCs) in a variety of solid waste matrices. | 0.005 | Y | Y | Y | 0 |  |  |  | 2024-01-25 |  | RDGIBSON |  | Concentration | mg/L | Y | (fl. conc.) |  | Toluene (fl. conc.) | 108-88-3 | Volatile Organic Compounds (VOCs); | Volatile Organic Compounds (VOCs); EMS code: T001 | Chemical |  | LAB
T001 | mg/L | Toluene | MS04 | Static Headspace GC/MS (Water) | (Water) A static headspace procedure is used as a means of driving the analytes out of the matrix into a headspace. Once equilibrium is achieved, a fixed volume of headspeace is taken for subsequent determination by Gas Chromatography - Mass Spectroscopy (GC/MS) using Selected Ion Monitoring. | 0.0005 | Y | Y | Y | 129 | 8/13/2009 | 6/28/2024 12:29:00 AM | 7/30/2024 6:32:12 AM | 2009-12-23 |  | RDGIBSON |  | Concentration | mg/L | Y | (fl. conc.) |  | Toluene (fl. conc.) | 108-88-3 | Volatile Organic Compounds (VOCs); | Volatile Organic Compounds (VOCs); EMS code: T001 | Chemical |  | LAB
T001 | mg/L | Toluene | MS01 | Purge&Trap GC/MS (VOC-W) | Purge&Trap GC/MS (VOC-W) | 0.4 | Y | Y | Y | 1034 | 4/15/2002 | 6/4/2024 10:24:00 AM | 6/13/2024 5:01:27 PM | 1996-04-12 | 3/18/2003 2:49:40 PM | EMS | UMULLER | Concentration | mg/L | Y | (fl. conc.) |  | Toluene (fl. conc.) | 108-88-3 | Volatile Organic Compounds (VOCs); | Volatile Organic Compounds (VOCs); EMS code: T001 | Chemical |  | LAB
T001 | ug/g | Toluene | X263 | Purge and trap GC/MS | Purge and trap GC/MS | 0.001 | Y | Y | Y | 20 | 3/29/2006 | 4/4/2006 | 4/27/2006 1:02:37 AM | 1996-04-12 | 6/3/1998 6:51:42 PM | EMS |  | MassPerMass | mg/kg | Y | (mass conc.) |  | Toluene (mass conc.) | 108-88-3 | Volatile Organic Compounds (VOCs); | Volatile Organic Compounds (VOCs); EMS code: T001 | Chemical |  | LAB
T001 | ug/g | Toluene | MS03 | Static Headspace GC/MS (Soil) | (Soil) A static headspace procedure is used as a means of driving the analytes out of the matrix into a headspace. Once equilibrium is achieved, a fixed volume of headspeace is taken for subsequent determination by Gas Chromatography - Mass Spectroscopy (GC/MS) using Selected Ion Monitoring. | 0.05 | Y | Y | Y | 228 |  |  | 4/30/2021 11:32:40 AM | 2009-12-18 |  | RDGIBSON |  | MassPerMass | mg/kg | Y | (mass conc.) |  | Toluene (mass conc.) | 108-88-3 | Volatile Organic Compounds (VOCs); | Volatile Organic Compounds (VOCs); EMS code: T001 | Chemical |  | LAB
T001 | ug/g | Toluene | MEOH | MeOH extrac,GC/FID,nC6-nC10 | MeOH extrac;GC/FID;nC6-nC10. EPA8260B; 5021; 5035; BC MOE | 0.05 | Y | Y | Y | 49 | 5/21/2015 | 8/11/2019 | 8/13/2019 10:56:08 AM | 2015-05-26 |  | RDGIBSON |  | MassPerMass | mg/kg | Y | (mass conc.) |  | Toluene (mass conc.) | 108-88-3 | Volatile Organic Compounds (VOCs); | Volatile Organic Compounds (VOCs); EMS code: T001 | Chemical |  | LAB
T001 | ug/g | Toluene | MS02 | Purge&Trap GC/MS (VOC-S) | Purge&Trap GC/MS (VOC-S) | 0.002 | Y | Y | Y | 117 | 11/19/2004 | 10/28/2022 5:04:00 PM | 7/19/2024 8:20:35 AM | 1996-04-12 | 6/3/1998 6:51:40 PM | EMS |  | MassPerMass | mg/kg | Y | (mass conc.) |  | Toluene (mass conc.) | 108-88-3 | Volatile Organic Compounds (VOCs); | Volatile Organic Compounds (VOCs); EMS code: T001 | Chemical |  | LAB
T001 | mg/L | Toluene | X384 | Purge and trap GC/MS | The sample is analyzed by purge and trap gas chromatography with detection by mass spectrometry. The procedure involves purging the volatile compounds from the sample with an inert gas, and trapping them on a solid sorbent. When purging of the sample is complete, the trap is heated and the compounds are transferred to the gas chromatographic column. Analysis is then accomplished by separation of the components by gas chromatography with detection by mass spectrometry. | 0.0005 | Y | Y | Y | 233 | 1/24/2001 | 7/26/2012 7:12:00 AM | 9/26/2012 1:07:09 PM | 1996-04-12 | 6/3/1998 6:51:40 PM | EMS |  | Concentration | mg/L | Y | (fl. conc.) |  | Toluene (fl. conc.) | 108-88-3 | Volatile Organic Compounds (VOCs); | Volatile Organic Compounds (VOCs); EMS code: T001 | Chemical |  | LAB
T008 | mg/L | Terbufos | E660 | EPA 8270E (mod) | EPA 8270E (mod) Miscellaneous Pesticides by GC-MS | 0.0001 | Y | Y | Y | 68 | 12/7/2021 | 9/12/2022 | 12/8/2022 12:57:06 PM | 2021-12-09 | 12/9/2021 3:22:12 PM | EMS | EMS | Concentration | mg/L | Y | (fl. conc.) |  | Terbufos (fl. conc.) | 13071-79-9 | Pesticides; | Pesticides; EMS code: T008 | Chemical |  | LAB
T008 | mg/L | Terbufos | ZN01 | DCM+Ac.Extr:GC/MS | DCM+Ac.Extr:GC/MS | 0.0003 | Y | Y | Y | 28 | 7/29/2005 7:10:00 AM | 3/29/2012 10:15:00 PM | 5/3/2012 9:27:52 AM | 1996-04-12 | 11/16/1999 4:06:28 PM | EMS | NPEPPIN | Concentration | mg/L | Y | (fl. conc.) |  | Terbufos (fl. conc.) | 13071-79-9 | Pesticides; | Pesticides; EMS code: T008 | Chemical |  | LAB
T008 | ug/g | Terbufos | ZN03 | DCM+Ac.Extr:GC/MS | DCM+Ac.Extr:GC/MS EPA 8270E (mod) | 0.03 | Y | Y | Y | 16 | 1/25/2022 | 11/2/2022 4:41:00 AM | 1/7/2023 3:12:11 PM | 1996-04-12 | 11/16/1999 2:27:40 PM | EMS | NPEPPIN | MassPerMass | mg/kg | Y | (mass conc.) |  | Terbufos (mass conc.) | 13071-79-9 | Pesticides; | Pesticides; EMS code: T008 | Chemical |  | LAB
T009 | mg/L | Tolyfluanid | ZN01 | DCM+Ac.Extr:GC/MS | DCM+Ac.Extr:GC/MS | 0.0005 | Y | Y | Y | 28 | 7/28/2005 7:03:00 AM | 3/29/2012 10:15:00 PM | 5/3/2012 9:27:52 AM | 1996-04-12 | 11/16/1999 4:06:28 PM | EMS | NPEPPIN | Concentration | mg/L | Y | (fl. conc.) |  | Tolyfluanid (fl. conc.) | 731-27-1 | Pesticides; | Pesticides; EMS code: T009 | Chemical |  | LAB
T011 | mg/L | Tecnazene | ZN01 | DCM+Ac.Extr:GC/MS | DCM+Ac.Extr:GC/MS | 0.0005 | Y | Y | Y | 27 | 7/28/2005 7:03:00 AM | 3/29/2012 10:15:00 PM | 5/3/2012 9:27:52 AM | 1996-04-12 | 11/16/1999 4:06:28 PM | EMS | NPEPPIN | Concentration | mg/L | Y | (fl. conc.) |  | Tecnazene (fl. conc.) | 117-18-0 | Pesticides; | Pesticides; EMS code: T011 | Chemical |  | LAB
T012 | mg/L | Tetradifon | ZN01 | DCM+Ac.Extr:GC/MS | DCM+Ac.Extr:GC/MS | 0.0002 | Y | Y | Y | 28 | 7/28/2005 7:03:00 AM | 3/29/2012 10:15:00 PM | 5/3/2012 9:27:52 AM | 1996-04-12 | 11/16/1999 4:06:28 PM | EMS | NPEPPIN | Concentration | mg/L | Y | (fl. conc.) |  | Tetradifon (fl. conc.) | 116-29-0 | Pesticides; | Pesticides; EMS code: T012 | Chemical |  | LAB
T013 | ug/g | Tetrachlorvinphos | ZN03 | DCM+Ac.Extr:GC/MS | DCM+Ac.Extr:GC/MS EPA 8270E (mod) | 0.02 | Y | Y | Y | 19 | 4/20/2002 10:25:00 AM | 10/16/2002 3:20:00 PM | 2/25/2004 1:09:47 AM | 1996-04-12 | 11/16/1999 2:27:40 PM | EMS | NPEPPIN | MassPerMass | mg/kg | Y | (mass conc.) |  | Tetrachlorvinphos (mass conc.) | 22248-79-9 | Pesticides; | Pesticides; EMS code: T013 | Chemical |  | LAB
T013 | mg/L | Tetrachlorvinphos | ZN01 | DCM+Ac.Extr:GC/MS | DCM+Ac.Extr:GC/MS | 0.0002 | Y | Y | Y | 52 | 10/3/2002 | 3/29/2012 10:15:00 PM | 5/3/2012 9:27:52 AM | 1996-04-12 | 11/16/1999 4:06:28 PM | EMS | NPEPPIN | Concentration | mg/L | Y | (fl. conc.) |  | Tetrachlorvinphos (fl. conc.) | 22248-79-9 | Pesticides; | Pesticides; EMS code: T013 | Chemical |  | LAB
T014 | ug/g wet | Toxaphene | FB26 | AcExtr'n,NaCl part,HexExtr'n,conc,Flori,GC/EC | Biota samples are extracted with Acetone, partitioned with 2% NaCl then back extracted with Hexane. The Hexane extract is then concentrated to small volume. The concentrated extracts are then treated using Florisil, concentrated Sulphuric Acid, activated Copper and/or Mercury (as required) to remove interferences. The samples are analysed using Microbore Column Gas Chromatography with Electron Capture Detection. | 0.01 | Y | Y | Y | 15 | 11/16/2000 1:04:00 PM | 11/16/2000 1:08:00 PM | 3/7/2001 1:00:48 AM | 2000-10-25 |  | NPEPPIN |  | MassPerMass | µg/g | Y | (mass conc.) | (wet) | Toxaphene (wet) (mass conc.) | 8001-35-2 | Organics; | Organics; EMS code: T014 | Chemical |  | LAB
T014 | ug/g wet | Toxaphene | X220 | Hex/AcExt,Part,Chrom,GC | Hex/AcExt;Part;Chrom;GC | 0.05 | Y | Y | Y | 73 |  |  | 10/24/1996 10:01:25 AM | 1996-04-12 | 6/3/1998 6:51:42 PM | EMS |  | MassPerMass | µg/g | Y | (mass conc.) | (wet) | Toxaphene (wet) (mass conc.) | 8001-35-2 | Organics; | Organics; EMS code: T014 | Chemical |  | LAB
T014 | ug/g | Toxaphene | X213 | Hex/AcExt,Part,Chrom,GC | Hex/AcExt;Part;Chrom;GC | 0.03 | Y | Y | Y | 161 |  |  | 10/24/1996 2:26:15 PM | 1996-04-12 | 6/3/1998 6:51:42 PM | EMS |  | MassPerMass | mg/kg | Y | (mass conc.) |  | Toxaphene (mass conc.) | 8001-35-2 | Organics; | Organics; EMS code: T014 | Chemical |  | LAB
T015 | mg/L | Trifluralin | E660 | EPA 8270E (mod) | EPA 8270E (mod) Miscellaneous Pesticides by GC-MS | 0.0001 | Y | Y | Y | 68 | 12/7/2021 | 9/12/2022 | 12/8/2022 12:57:06 PM | 2021-12-09 | 12/9/2021 3:22:12 PM | EMS | EMS | Concentration | mg/L | Y | (fl. conc.) |  | Trifluralin (fl. conc.) | 1582-09-8 | Pesticides; | Pesticides; EMS code: T015 | Chemical |  | LAB
T015 | mg/L | Trifluralin | ZN01 | DCM+Ac.Extr:GC/MS | DCM+Ac.Extr:GC/MS | 0.0002 | Y | Y | Y | 28 | 7/28/2005 7:03:00 AM | 3/29/2012 10:15:00 PM | 5/3/2012 9:27:52 AM | 1996-04-12 | 11/16/1999 4:06:28 PM | EMS | NPEPPIN | Concentration | mg/L | Y | (fl. conc.) |  | Trifluralin (fl. conc.) | 1582-09-8 | Pesticides; | Pesticides; EMS code: T015 | Chemical |  | LAB
T015 | ug/g | Trifluralin | ZN03 | DCM+Ac.Extr:GC/MS | DCM+Ac.Extr:GC/MS EPA 8270E (mod) | 0.02 | Y | Y | Y | 16 | 1/25/2022 | 11/2/2022 4:41:00 AM | 1/7/2023 3:12:11 PM | 1996-04-12 | 11/16/1999 2:27:40 PM | EMS | NPEPPIN | MassPerMass | mg/kg | Y | (mass conc.) |  | Trifluralin (mass conc.) | 1582-09-8 | Pesticides; | Pesticides; EMS code: T015 | Chemical |  | LAB
T016 | mg/L | 1;1;1-Trichloroethane | X384 | Purge and trap GC/MS | The sample is analyzed by purge and trap gas chromatography with detection by mass spectrometry. The procedure involves purging the volatile compounds from the sample with an inert gas, and trapping them on a solid sorbent. When purging of the sample is complete, the trap is heated and the compounds are transferred to the gas chromatographic column. Analysis is then accomplished by separation of the components by gas chromatography with detection by mass spectrometry. | 0.001 | Y | Y | Y | 121 | 4/27/2000 | 7/26/2012 7:12:00 AM | 9/26/2012 1:07:09 PM | 1996-04-12 | 6/3/1998 6:51:40 PM | EMS |  | Concentration | mg/L | Y | (fl. conc.) |  | 1;1;1-Trichloroethane (fl. conc.) | 71-55-6 | Volatile Organic Compounds (VOCs); | Volatile Organic Compounds (VOCs); EMS code: T016 | Chemical |  | LAB
T016 | mg/L | 1;1;1-Trichloroethane | MS01 | Purge&Trap GC/MS (VOC-W) | Purge&Trap GC/MS (VOC-W) | 0.6 | Y | Y | Y | 544 | 4/15/2002 | 6/4/2024 10:24:00 AM | 6/13/2024 5:01:27 PM | 1996-04-12 | 3/18/2003 2:49:40 PM | EMS | UMULLER | Concentration | mg/L | Y | (fl. conc.) |  | 1;1;1-Trichloroethane (fl. conc.) | 71-55-6 | Volatile Organic Compounds (VOCs); | Volatile Organic Compounds (VOCs); EMS code: T016 | Chemical |  | LAB
T016 | mg/L | 1;1;1-Trichloroethane | MS04 | Static Headspace GC/MS (Water) | (Water) A static headspace procedure is used as a means of driving the analytes out of the matrix into a headspace. Once equilibrium is achieved, a fixed volume of headspeace is taken for subsequent determination by Gas Chromatography - Mass Spectroscopy (GC/MS) using Selected Ion Monitoring. | 0.0005 | Y | Y | Y | 70 | 8/13/2009 | 6/28/2024 12:29:00 AM | 7/9/2024 3:31:02 PM | 2009-12-18 |  | RDGIBSON |  | Concentration | mg/L | Y | (fl. conc.) |  | 1;1;1-Trichloroethane (fl. conc.) | 71-55-6 | Volatile Organic Compounds (VOCs); | Volatile Organic Compounds (VOCs); EMS code: T016 | Chemical |  | LAB
T016 | mg/L | 1;1;1-Trichloroethane | F007 | Static Headspace:HRGC/LRMS | Static headspace analysis is a method for the determination of volatile compounds in liquids and/or solids by analysing the vapour phase of a sample in a closed system; in this case a sealed vial. Static Headspace is sometimes referred to as Equilibrium Headspace because the vapour phase above the sample in a sealed vial is brought into thermodynamic equilibrium with the sample prior to analysis. Volatile organic compounds are thermally extracted from the sample matrix by Static Headspace sampling and introduced directly to a thick-film megabore GC column. Compounds eluting from the GC column are introduced into a Mass Spectrometer and target compounds are identified by a comparison of both retention times and mass spectra with those of reference standards stored in a computerized data system. Reference retention times and mass spectra for analytes are obtained by the measurement of calibration standards under precisely the same extraction conditions used for samples. The concentration of each identified component is measured by relating the response of the sample quantitation mass ions to the response of the quantitation mass ions of a series of external standards. A deuterated surrogate recovery standard is added to provide information on matrix effects. | 0.005 | Y | Y | Y | 380 | 9/28/1998 3:08:00 PM | 8/14/2024 | 8/16/2024 4:05:21 PM | 1996-09-16 | 6/3/1998 6:51:45 PM | MPAK |  | Concentration | mg/L | Y | (fl. conc.) |  | 1;1;1-Trichloroethane (fl. conc.) | 71-55-6 | Volatile Organic Compounds (VOCs); | Volatile Organic Compounds (VOCs); EMS code: T016 | Chemical |  | LAB
T016 | ug/g | 1;1;1-Trichloroethane | MS02 | Purge&Trap GC/MS (VOC-S) | Purge&Trap GC/MS (VOC-S) | 0.003 | Y | Y | Y | 27 | 11/19/2004 | 10/28/2022 5:04:00 PM | 1/11/2024 1:07:08 PM | 1996-04-12 | 6/3/1998 6:51:40 PM | EMS |  | MassPerMass | mg/kg | Y | (mass conc.) |  | 1;1;1-Trichloroethane (mass conc.) | 71-55-6 | Volatile Organic Compounds (VOCs); | Volatile Organic Compounds (VOCs); EMS code: T016 | Chemical |  | LAB
T016 | ug/g | 1;1;1-Trichloroethane | FS04 | Static Headspace:HRGC/LRMS | Static Headspace anaylsis is a method for the determination of volatile compounds in liquids and/or solids by analysing the vapour phase of a sample in a closed system; in this case a sealed vial. Static Headspace is sometimes referred to as Equilibrium Headspace because the vapour phase above the sample in a sealed vial is brought into thermodynamic equilibrium with the sample prior to analysis. Volatile organic compounds are thermally extracted from the sample matrix by Static Headspace sampling and introduced directly to a thick-film megabore GC column. Compounds eluting from the GC column are introduced into a Mass Spectrometer and target compounds are identified by a comparison of both retention times and mass spectra with those of reference standards stored in a computerized data system. Reference retention times and mass spectra for analytes are obtained by the measurement of calibration standards under precisely the same extraction conditions used for samples. The concentration of each identified component is measured by relating the response of the sample quantitation mass ions to the response of the quantitation mass ions of a series of external standards. A deuterated surrogate recovery standard is added to provide information on matrix effects. | 0.05 | Y | Y | Y | 44 | 6/30/1999 4:41:00 PM | 5/22/2015 | 7/3/2015 3:19:04 PM | 1996-09-16 | 6/3/1998 6:51:45 PM | MPAK |  | MassPerMass | mg/kg | Y | (mass conc.) |  | 1;1;1-Trichloroethane (mass conc.) | 71-55-6 | Volatile Organic Compounds (VOCs); | Volatile Organic Compounds (VOCs); EMS code: T016 | Chemical |  | LAB
T017 | ug/g | 1;1;2-Trichloroethane | MS02 | Purge&Trap GC/MS (VOC-S) | Purge&Trap GC/MS (VOC-S) | 0.003 | Y | Y | Y | 27 | 11/19/2004 | 10/28/2022 5:04:00 PM | 1/11/2024 1:07:08 PM | 1996-04-12 | 6/3/1998 6:51:40 PM | EMS |  | MassPerMass | mg/kg | Y | (mass conc.) |  | 1;1;2-Trichloroethane (mass conc.) | 79-00-5 | Volatile Organic Compounds (VOCs); | Volatile Organic Compounds (VOCs); EMS code: T017 | Chemical |  | LAB
T017 | mg/L | 1;1;2-Trichloroethane | MS01 | Purge&Trap GC/MS (VOC-W) | Purge&Trap GC/MS (VOC-W) | 0.6 | Y | Y | Y | 544 | 4/15/2002 | 6/4/2024 10:24:00 AM | 6/13/2024 5:01:27 PM | 1996-04-12 | 6/3/1998 6:51:42 PM | EMS |  | Concentration | mg/L | Y | (fl. conc.) |  | 1;1;2-Trichloroethane (fl. conc.) | 79-00-5 | Volatile Organic Compounds (VOCs); | Volatile Organic Compounds (VOCs); EMS code: T017 | Chemical |  | LAB
T017 | ug/g | 1;1;2-Trichloroethane | FS04 | Static Headspace:HRGC/LRMS | Static Headspace anaylsis is a method for the determination of volatile compounds in liquids and/or solids by analysing the vapour phase of a sample in a closed system; in this case a sealed vial. Static Headspace is sometimes referred to as Equilibrium Headspace because the vapour phase above the sample in a sealed vial is brought into thermodynamic equilibrium with the sample prior to analysis. Volatile organic compounds are thermally extracted from the sample matrix by Static Headspace sampling and introduced directly to a thick-film megabore GC column. Compounds eluting from the GC column are introduced into a Mass Spectrometer and target compounds are identified by a comparison of both retention times and mass spectra with those of reference standards stored in a computerized data system. Reference retention times and mass spectra for analytes are obtained by the measurement of calibration standards under precisely the same extraction conditions used for samples. The concentration of each identified component is measured by relating the response of the sample quantitation mass ions to the response of the quantitation mass ions of a series of external standards. A deuterated surrogate recovery standard is added to provide information on matrix effects. | 0.05 | Y | Y | Y | 44 | 6/30/1999 4:41:00 PM | 5/22/2015 | 7/3/2015 3:19:04 PM | 1996-09-16 | 6/3/1998 6:51:45 PM | MPAK |  | MassPerMass | mg/kg | Y | (mass conc.) |  | 1;1;2-Trichloroethane (mass conc.) | 79-00-5 | Volatile Organic Compounds (VOCs); | Volatile Organic Compounds (VOCs); EMS code: T017 | Chemical |  | LAB
T017 | mg/L | 1;1;2-Trichloroethane | X384 | Purge and trap GC/MS | The sample is analyzed by purge and trap gas chromatography with detection by mass spectrometry. The procedure involves purging the volatile compounds from the sample with an inert gas, and trapping them on a solid sorbent. When purging of the sample is complete, the trap is heated and the compounds are transferred to the gas chromatographic column. Analysis is then accomplished by separation of the components by gas chromatography with detection by mass spectrometry. | 0.001 | Y | Y | Y | 112 | 8/5/2004 10:31:00 AM | 7/26/2012 7:12:00 AM | 9/26/2012 1:07:09 PM | 1996-04-12 | 6/3/1998 6:51:45 PM | EMS |  | Concentration | mg/L | Y | (fl. conc.) |  | 1;1;2-Trichloroethane (fl. conc.) | 79-00-5 | Volatile Organic Compounds (VOCs); | Volatile Organic Compounds (VOCs); EMS code: T017 | Chemical |  | LAB
T017 | mg/L | 1;1;2-Trichloroethane | MS04 | Static Headspace GC/MS (Water) | (Water) A static headspace procedure is used as a means of driving the analytes out of the matrix into a headspace. Once equilibrium is achieved, a fixed volume of headspeace is taken for subsequent determination by Gas Chromatography - Mass Spectroscopy (GC/MS) using Selected Ion Monitoring. | 0.0005 | Y | Y | Y | 70 | 8/13/2009 | 6/28/2024 12:29:00 AM | 7/9/2024 3:31:02 PM | 2009-12-18 |  | RDGIBSON |  | Concentration | mg/L | Y | (fl. conc.) |  | 1;1;2-Trichloroethane (fl. conc.) | 79-00-5 | Volatile Organic Compounds (VOCs); | Volatile Organic Compounds (VOCs); EMS code: T017 | Chemical |  | LAB
T017 | mg/L | 1;1;2-Trichloroethane | F007 | Static Headspace:HRGC/LRMS | Static headspace analysis is a method for the determination of volatile compounds in liquids and/or solids by analysing the vapour phase of a sample in a closed system; in this case a sealed vial. Static Headspace is sometimes referred to as Equilibrium Headspace because the vapour phase above the sample in a sealed vial is brought into thermodynamic equilibrium with the sample prior to analysis. Volatile organic compounds are thermally extracted from the sample matrix by Static Headspace sampling and introduced directly to a thick-film megabore GC column. Compounds eluting from the GC column are introduced into a Mass Spectrometer and target compounds are identified by a comparison of both retention times and mass spectra with those of reference standards stored in a computerized data system. Reference retention times and mass spectra for analytes are obtained by the measurement of calibration standards under precisely the same extraction conditions used for samples. The concentration of each identified component is measured by relating the response of the sample quantitation mass ions to the response of the quantitation mass ions of a series of external standards. A deuterated surrogate recovery standard is added to provide information on matrix effects. | 0.005 | Y | Y | Y | 380 | 9/28/1998 3:08:00 PM | 8/14/2024 | 8/16/2024 4:05:21 PM | 1996-09-16 | 6/3/1998 6:51:45 PM | MPAK |  | Concentration | mg/L | Y | (fl. conc.) |  | 1;1;2-Trichloroethane (fl. conc.) | 79-00-5 | Volatile Organic Compounds (VOCs); | Volatile Organic Compounds (VOCs); EMS code: T017 | Chemical |  | LAB
T018 | ug/g | 2;3;4;5-Tetrachloroanisole | FS11 | Acetone extr,Hexane,Diazometh.Methyl'n,HRGC/Electron Capture | Sediment and homogenized biota samples to be analysed for Chlorinated Phenols are Acetone extracted using a wrist-action shaker, partitioned with 2% NaCl and Hexane, methylated, then cleaned up and analysed in the manner outlined for effluent samples. Sediment and homogenized biota samples to be analysed for Chlorinated Anisoles are Acetone extracted using a wrist-action shaker, partitioned with 2% NaCl and Hexane, then cleaned up and analysed in the manner outlined for effluent samples. The extracts are analysed for Chlorinated Anisoles and/or Chlorinated Phenols using High Resolution Gas Chromatography with Electron Capture detection. | 0.0001 | Y | Y | Y | 4 | 8/29/2000 4:27:00 PM | 8/29/2000 4:29:00 PM | 6/18/2002 3:50:40 PM | 1997-01-14 | 6/3/1998 6:51:46 PM | MPAK |  | MassPerMass | mg/kg | Y | (mass conc.) |  | 2;3;4;5-Tetrachloroanisole (mass conc.) | 938-86-3 | Organics; | Organics; EMS code: T018 | Chemical |  | LAB
T018 | mg/L | 2;3;4;5-Tetrachloroanisole | F030 | Acid,Hexane extr,Diazomethane Methylation,HRGC/Electron Capt | Effluent samples to be analysed for Chlorinated Anisoles are extracted with Hexane. After concentration the extract is cleaned up using Florisil, concentrated Sulfuric Acid and Mercury. The extracts are analysed for Chlorinated Anisoles and/or Chlorinated Phenols using High Resolution Gas Chromatography with Electron Capture detection. | 1e-06 | Y | Y | Y | 3 | 8/25/2000 3:03:00 PM | 9/1/2000 9:20:00 AM | 6/18/2002 1:43:31 PM | 1997-01-14 | 6/3/1998 6:51:46 PM | MPAK |  | Concentration | mg/L | Y | (fl. conc.) |  | 2;3;4;5-Tetrachloroanisole (fl. conc.) | 938-86-3 | Organics; | Organics; EMS code: T018 | Chemical |  | LAB
T019 | mg/L | Trichlorsyringols | PP11 | Ac. anhydride ext:GC/MS | Ac. anhydride ext:GC/MS | 5e-05 | Y | Y | Y | 240 |  |  | 10/24/1996 5:30:20 AM | 1996-04-12 | 6/3/1998 6:51:40 PM | EMS |  | Concentration | mg/L | Y | (fl. conc.) |  | Trichlorsyringols (fl. conc.) | 2539-26-6 | Organics; | Organics; EMS code: T019 | Chemical |  | LAB
T020 | mg/L | Tetrachlorophenols | EQ01 | EQUIS (mg/L) | EQUIS (mg/L) | 0.005 | N | Y | Y | 10 |  |  | 10/22/1996 9:17:28 PM | 1996-04-12 | 6/3/1998 6:51:40 PM | EMS |  | Concentration | mg/L | Y | (fl. conc.) |  | Tetrachlorophenols (fl. conc.) | 25167-83-3 | Organics; | Organics; EMS code: T020 | Chemical |  | LAB
T020 | mg/L | Tetrachlorophenols | X265 | CH2Cl2Ext:Ac/Base:GC/ECD | CH2Cl2Ext:Ac/Base:GC/ECD | 0.0005 | Y | Y | Y | 51 |  |  | 8/23/2000 11:41:33 AM | 1997-04-16 | 6/3/1998 6:51:46 PM | NPEPPIN |  | Concentration | mg/L | Y | (fl. conc.) |  | Tetrachlorophenols (fl. conc.) | 25167-83-3 | Organics; | Organics; EMS code: T020 | Chemical |  | LAB
T020 | mg/L | Tetrachlorophenols | P008 | Chlor. Phenols | An acidified water sample is extracted with dichloromethane. The raw extract is concentrated and the phenols and acids are reacted with diazomethane, or other suitable derivatizing agent, to produce the corresponding derivatives. If required, the extracts are subjected to Florisil column chromatography and divided into different fractions. The derivatives are analyzed by electron capture gas liquid chromatography. | 0.0001 | Y | Y | Y | 328 | 8/11/2004 9:03:00 AM | 8/3/2005 2:39:00 PM | 6/2/2009 12:01:12 PM | 1996-04-12 | 6/3/1998 6:51:40 PM | EMS |  | Concentration | mg/L | Y | (fl. conc.) |  | Tetrachlorophenols (fl. conc.) | 25167-83-3 | Organics; | Organics; EMS code: T020 | Chemical |  | LAB
T020 | ug/g | Tetrachlorophenols | X212 | Hex/AcExtr,Part,IonEx,GC | Hex/AcExtr;Part;IonEx;GC | 0.001 | Y | Y | Y | 82 |  |  | 10/24/1996 10:24:19 AM | 1996-04-12 | 6/3/1998 6:51:40 PM | EMS |  | MassPerMass | mg/kg | Y | (mass conc.) |  | Tetrachlorophenols (mass conc.) | 25167-83-3 | Organics; | Organics; EMS code: T020 | Chemical |  | LAB
T020 | mg/L | Tetrachlorophenols | XCAL | Calculation (mg/L) | Calculation (mg/L) | 0.0001 | Y | Y | Y | 97 |  |  | 10/24/1996 2:48:52 PM | 1996-04-12 | 6/3/1998 6:51:42 PM | EMS |  | Concentration | mg/L | Y | (fl. conc.) |  | Tetrachlorophenols (fl. conc.) | 25167-83-3 | Organics; | Organics; EMS code: T020 | Chemical |  | LAB
T020 | ug/g wet | Tetrachlorophenols | X219 | Hex/AcExtr,Part,IonEx,GC | Hex/AcExtr;Part;IonEx;GC | 0.0002 | Y | Y | Y | 74 |  |  | 10/24/1996 10:01:25 AM | 1996-04-12 | 6/3/1998 6:51:42 PM | EMS |  | MassPerMass | µg/g | Y | (mass conc.) | (wet) | Tetrachlorophenols (wet) (mass conc.) | 25167-83-3 | Organics; | Organics; EMS code: T020 | Chemical |  | LAB
T020 | ug/g | Tetrachlorophenols | P029 | Chlor Phen | Chlor Phen | 0.01 | Y | Y | Y | 101 |  |  | 10/24/1996 12:38:17 PM | 1996-04-12 | 6/3/1998 6:51:45 PM | EMS |  | MassPerMass | mg/kg | Y | (mass conc.) |  | Tetrachlorophenols (mass conc.) | 25167-83-3 | Organics; | Organics; EMS code: T020 | Chemical |  | LAB
T020 | ug/g | Tetrachlorophenols | P028 | Chlor. Phenols | Chlor. Phenols | 0.005 | Y | Y | Y | 167 |  |  | 10/24/1996 12:44:22 PM | 1996-04-12 | 6/3/1998 6:51:40 PM | EMS |  | MassPerMass | mg/kg | Y | (mass conc.) |  | Tetrachlorophenols (mass conc.) | 25167-83-3 | Organics; | Organics; EMS code: T020 | Chemical |  | LAB
T020 | mg/L | Tetrachlorophenols | PP11 | Ac. anhydride ext:GC/MS | Ac. anhydride ext:GC/MS | 5e-05 | Y | Y | Y | 224 |  |  | 10/24/1996 5:30:20 AM | 1996-04-12 | 6/3/1998 6:51:40 PM | EMS |  | Concentration | mg/L | Y | (fl. conc.) |  | Tetrachlorophenols (fl. conc.) | 25167-83-3 | Organics; | Organics; EMS code: T020 | Chemical |  | LAB
T021 | ug/g wet | Trichlorophenols | X219 | Hex/AcExtr,Part,IonEx,GC | Hex/AcExtr;Part;IonEx;GC | 0.0002 | Y | Y | Y | 75 |  |  | 10/24/1996 10:01:25 AM | 1996-04-12 | 6/3/1998 6:51:42 PM | EMS |  | MassPerMass | µg/g | Y | (mass conc.) | (wet) | Trichlorophenols (wet) (mass conc.) | 25167-82-2 | Organics; | Organics; EMS code: T021 | Chemical |  | LAB
T021 | ug/g | Trichlorophenols | P029 | Chlor Phen | Chlor Phen | 0.01 | Y | Y | Y | 99 |  |  | 10/24/1996 12:38:17 PM | 1996-04-12 | 6/3/1998 6:51:40 PM | EMS |  | MassPerMass | mg/kg | Y | (mass conc.) |  | Trichlorophenols (mass conc.) | 25167-82-2 | Organics; | Organics; EMS code: T021 | Chemical |  | LAB
T021 | mg/L | Trichlorophenols | XCAL | Calculation (mg/L) | Calculation (mg/L) | 0.0001 | Y | Y | Y | 97 |  |  | 10/24/1996 2:48:52 PM | 1996-04-12 | 6/3/1998 6:51:42 PM | EMS |  | Concentration | mg/L | Y | (fl. conc.) |  | Trichlorophenols (fl. conc.) | 25167-82-2 | Organics; | Organics; EMS code: T021 | Chemical |  | LAB
T021 | mg/L | Trichlorophenols | P008 | Chlor. Phenols | An acidified water sample is extracted with dichloromethane. The raw extract is concentrated and the phenols and acids are reacted with diazomethane, or other suitable derivatizing agent, to produce the corresponding derivatives. If required, the extracts are subjected to Florisil column chromatography and divided into different fractions. The derivatives are analyzed by electron capture gas liquid chromatography. | 0.0001 | Y | Y | Y | 243 | 8/11/2004 9:03:00 AM | 8/3/2005 2:39:00 PM | 6/2/2009 12:01:12 PM | 1996-04-12 | 6/3/1998 6:51:45 PM | EMS |  | Concentration | mg/L | Y | (fl. conc.) |  | Trichlorophenols (fl. conc.) | 25167-82-2 | Organics; | Organics; EMS code: T021 | Chemical |  | LAB
T021 | mg/L | Trichlorophenols | X265 | CH2Cl2Ext:Ac/Base:GC/ECD | CH2Cl2Ext:Ac/Base:GC/ECD | 0.0005 | Y | Y | Y | 51 |  |  | 8/23/2000 11:41:34 AM | 1997-04-16 | 6/3/1998 6:51:46 PM | NPEPPIN |  | Concentration | mg/L | Y | (fl. conc.) |  | Trichlorophenols (fl. conc.) | 25167-82-2 | Organics; | Organics; EMS code: T021 | Chemical |  | LAB
T021 | mg/L | Trichlorophenols | PP11 | Ac. anhydride ext:GC/MS | Ac. anhydride ext:GC/MS | 5e-05 | Y | Y | Y | 226 |  |  | 10/24/1996 5:30:20 AM | 1996-04-12 | 6/3/1998 6:51:40 PM | EMS |  | Concentration | mg/L | Y | (fl. conc.) |  | Trichlorophenols (fl. conc.) | 25167-82-2 | Organics; | Organics; EMS code: T021 | Chemical |  | LAB
T021 | ug/g | Trichlorophenols | X212 | Hex/AcExtr,Part,IonEx,GC | Hex/AcExtr;Part;IonEx;GC | 0.001 | Y | Y | Y | 83 |  |  | 10/24/1996 10:24:19 AM | 1996-04-12 | 6/3/1998 6:51:40 PM | EMS |  | MassPerMass | mg/kg | Y | (mass conc.) |  | Trichlorophenols (mass conc.) | 25167-82-2 | Organics; | Organics; EMS code: T021 | Chemical |  | LAB
T021 | ug/g | Trichlorophenols | P028 | Chlor. Phenols | Chlor. Phenols | 0.005 | Y | Y | Y | 161 |  |  | 10/24/1996 12:44:22 PM | 1996-04-12 | 6/3/1998 6:51:40 PM | EMS |  | MassPerMass | mg/kg | Y | (mass conc.) |  | Trichlorophenols (mass conc.) | 25167-82-2 | Organics; | Organics; EMS code: T021 | Chemical |  | LAB
T022 | mg/L | Trichloroguaiacol | PP11 | Ac. anhydride ext:GC/MS | Ac. anhydride ext:GC/MS | 5e-05 | Y | Y | Y | 184 |  |  | 10/24/1996 5:30:20 AM | 1996-04-12 | 6/3/1998 6:51:40 PM | EMS |  | Concentration | mg/L | Y | (fl. conc.) |  | Trichloroguaiacol (fl. conc.) | 57057-83-7, 2668-24-8 | Organics; | Organics; EMS code: T022 | Chemical |  | LAB
T023 | mg/L | Tetradecanoic acid | RFAS | Separation of resin and fatty acids (RFAs) in water by GC/MS | Scientific Principle The separation of resin and fatty acids (RFAs) in water samples is performed using gas chromatography/mass spectrometry (GC/MS) in the Selected Ion Monitoring (SIM) mode. The separation is based on compound size; boiling point and polarity. Analyte concentration is calculated using the internal standard method. RFA compounds are extracted from water samples using a liquid-liquid solvent extraction with DCM. Extraction occurs after the sample is brought to a pH of ~ 2 and placed in a rotary tumbler (at approx. 40 rpm) for 1 hour. The resulting extract is transferred to a vial; dried with acidified baked sodium sulfate; methylated with diazomethane and concentrated to 1mL under nitrogen prior to analysis. The resulting extracts are vialed and fortified with labeled internal standard compounds prior to analysis by gas chromatography with a (50%-Phenyl)-methylpolysiloxane (DB-17) column and a mass spectrometer detector in SIM Mode. | 0.006 | Y | Y | Y | 1 | 9/17/2021 7:49:00 PM | 9/17/2021 7:49:00 PM | 2/1/2022 2:38:14 PM | 2022-01-19 |  | EMS |  | Concentration | mg/L | Y | (fl. conc.) |  | Tetradecanoic acid (fl. conc.) | 544-63-8 | Organics; | Organics; EMS code: T023 | Chemical |  | LAB
T024 | mg/L | Tetrachloroguaiacol | PP11 | Ac. anhydride ext:GC/MS | Ac. anhydride ext:GC/MS | 5e-05 | Y | Y | Y | 206 |  |  | 10/24/1996 5:30:20 AM | 1996-04-12 | 6/3/1998 6:51:46 PM | EMS |  | Concentration | mg/L | Y | (fl. conc.) |  | Tetrachloroguaiacol (fl. conc.) | 2539-17-5 | Organics; | Organics; EMS code: T024 | Chemical |  | LAB
T025 | mg/L | Trichlorocatechols | PP11 | Ac. anhydride ext:GC/MS | Ac. anhydride ext:GC/MS | 5e-05 | Y | Y | Y | 241 |  |  | 10/24/1996 5:30:20 AM | 1996-04-12 | 6/3/1998 6:51:40 PM | EMS |  | Concentration | mg/L | Y | (fl. conc.) |  | Trichlorocatechols (fl. conc.) | 56961-20-7, 32139-72-3 | Organics; | Organics; EMS code: T025 | Chemical |  | LAB
T026 | mg/L | Tetrachlorcatechol | PP11 | Ac. anhydride ext:GC/MS | Ac. anhydride ext:GC/MS | 5e-05 | Y | Y | Y | 239 |  |  | 10/24/1996 5:30:20 AM | 1996-04-12 | 6/3/1998 6:51:40 PM | EMS |  | Concentration | mg/L | Y | (fl. conc.) |  | Tetrachlorcatechol (fl. conc.) | 1198-55-6 | Organics; | Organics; EMS code: T026 | Chemical |  | LAB
T027 | mg/L | Trichloroveretrols | PP11 | Ac. anhydride ext:GC/MS | Ac. anhydride ext:GC/MS | 5e-05 | Y | Y | Y | 240 |  |  | 10/24/1996 5:30:20 AM | 1996-04-12 | 6/3/1998 6:51:40 PM | EMS |  | Concentration | mg/L | Y | (fl. conc.) |  | Trichloroveretrols (fl. conc.) | 16766-29-3 | Organics; | Organics; EMS code: T027 | Chemical |  | LAB
T028 | mg/L | Tetrachloroveretrols | PP11 | Ac. anhydride ext:GC/MS | Ac. anhydride ext:GC/MS | 5e-05 | Y | Y | Y | 240 |  |  | 10/24/1996 5:30:20 AM | 1996-04-12 | 6/3/1998 6:51:46 PM | EMS |  | Concentration | mg/L | Y | (fl. conc.) |  | Tetrachloroveretrols (fl. conc.) | 944-61-6 | Organics; | Organics; EMS code: T028 | Chemical |  | LAB
T029 | ug/g | Trichloroethylene | FS04 | Static Headspace:HRGC/LRMS | Static Headspace anaylsis is a method for the determination of volatile compounds in liquids and/or solids by analysing the vapour phase of a sample in a closed system; in this case a sealed vial. Static Headspace is sometimes referred to as Equilibrium Headspace because the vapour phase above the sample in a sealed vial is brought into thermodynamic equilibrium with the sample prior to analysis. Volatile organic compounds are thermally extracted from the sample matrix by Static Headspace sampling and introduced directly to a thick-film megabore GC column. Compounds eluting from the GC column are introduced into a Mass Spectrometer and target compounds are identified by a comparison of both retention times and mass spectra with those of reference standards stored in a computerized data system. Reference retention times and mass spectra for analytes are obtained by the measurement of calibration standards under precisely the same extraction conditions used for samples. The concentration of each identified component is measured by relating the response of the sample quantitation mass ions to the response of the quantitation mass ions of a series of external standards. A deuterated surrogate recovery standard is added to provide information on matrix effects. | 0.05 | Y | Y | Y | 44 | 6/30/1999 4:41:00 PM | 5/22/2015 | 7/3/2015 3:19:04 PM | 1996-09-16 | 6/3/1998 6:51:45 PM | MPAK |  | MassPerMass | mg/kg | Y | (mass conc.) |  | Trichloroethylene (mass conc.) | 79-01-6 | Volatile Organic Compounds (VOCs); | Volatile Organic Compounds (VOCs); EMS code: T029 | Chemical |  | LAB
T029 | mg/L | Trichloroethylene | MS04 | Static Headspace GC/MS (Water) | (Water) A static headspace procedure is used as a means of driving the analytes out of the matrix into a headspace. Once equilibrium is achieved, a fixed volume of headspeace is taken for subsequent determination by Gas Chromatography - Mass Spectroscopy (GC/MS) using Selected Ion Monitoring. | 0.0005 | Y | Y | Y | 81 | 8/13/2009 | 6/28/2024 12:29:00 AM | 7/23/2024 11:41:50 AM | 2010-01-13 | 1/13/2010 1:02:54 PM | RDGIBSON | RDGIBSON | Concentration | mg/L | Y | (fl. conc.) |  | Trichloroethylene (fl. conc.) | 79-01-6 | Volatile Organic Compounds (VOCs); | Volatile Organic Compounds (VOCs); EMS code: T029 | Chemical |  | LAB
T029 | mg/L | Trichloroethylene | F007 | Static Headspace:HRGC/LRMS | Static headspace analysis is a method for the determination of volatile compounds in liquids and/or solids by analysing the vapour phase of a sample in a closed system; in this case a sealed vial. Static Headspace is sometimes referred to as Equilibrium Headspace because the vapour phase above the sample in a sealed vial is brought into thermodynamic equilibrium with the sample prior to analysis. Volatile organic compounds are thermally extracted from the sample matrix by Static Headspace sampling and introduced directly to a thick-film megabore GC column. Compounds eluting from the GC column are introduced into a Mass Spectrometer and target compounds are identified by a comparison of both retention times and mass spectra with those of reference standards stored in a computerized data system. Reference retention times and mass spectra for analytes are obtained by the measurement of calibration standards under precisely the same extraction conditions used for samples. The concentration of each identified component is measured by relating the response of the sample quantitation mass ions to the response of the quantitation mass ions of a series of external standards. A deuterated surrogate recovery standard is added to provide information on matrix effects. | 0.005 | Y | Y | Y | 380 | 9/28/1998 3:08:00 PM | 8/14/2024 | 8/16/2024 4:05:21 PM | 1996-09-16 | 6/3/1998 6:51:45 PM | MPAK |  | Concentration | mg/L | Y | (fl. conc.) |  | Trichloroethylene (fl. conc.) | 79-01-6 | Volatile Organic Compounds (VOCs); | Volatile Organic Compounds (VOCs); EMS code: T029 | Chemical |  | LAB
T029 | ug/g | Trichloroethylene | MS02 | Purge&Trap GC/MS (VOC-S) | Purge&Trap GC/MS (VOC-S) | 0.002 | Y | Y | Y | 27 | 11/19/2004 | 10/28/2022 5:04:00 PM | 1/11/2024 1:07:08 PM | 1996-10-13 | 6/3/1998 6:51:46 PM | EMS |  | MassPerMass | mg/kg | Y | (mass conc.) |  | Trichloroethylene (mass conc.) | 79-01-6 | Volatile Organic Compounds (VOCs); | Volatile Organic Compounds (VOCs); EMS code: T029 | Chemical |  | LAB
T029 | mg/L | Trichloroethylene | MS01 | Purge&Trap GC/MS (VOC-W) | Purge&Trap GC/MS (VOC-W) | 0.3 | Y | Y | Y | 390 | 4/15/2002 | 6/4/2024 10:24:00 AM | 6/13/2024 5:01:27 PM | 1996-10-13 | 3/18/2003 2:49:40 PM | EMS | UMULLER | Concentration | mg/L | Y | (fl. conc.) |  | Trichloroethylene (fl. conc.) | 79-01-6 | Volatile Organic Compounds (VOCs); | Volatile Organic Compounds (VOCs); EMS code: T029 | Chemical |  | LAB
T029 | mg/L | Trichloroethylene | X384 | Purge and trap GC/MS | The sample is analyzed by purge and trap gas chromatography with detection by mass spectrometry. The procedure involves purging the volatile compounds from the sample with an inert gas, and trapping them on a solid sorbent. When purging of the sample is complete, the trap is heated and the compounds are transferred to the gas chromatographic column. Analysis is then accomplished by separation of the components by gas chromatography with detection by mass spectrometry. | 0.001 | Y | Y | Y | 10 |  |  | 3/19/2004 8:18:44 AM | 1996-04-12 | 6/3/1998 6:51:40 PM | EMS |  | Concentration | mg/L | Y | (fl. conc.) |  | Trichloroethylene (fl. conc.) | 79-01-6 | Volatile Organic Compounds (VOCs); | Volatile Organic Compounds (VOCs); EMS code: T029 | Chemical |  | LAB
T030 | mg/L | Tetrachloroethylene | F007 | Static Headspace:HRGC/LRMS | Static headspace analysis is a method for the determination of volatile compounds in liquids and/or solids by analysing the vapour phase of a sample in a closed system; in this case a sealed vial. Static Headspace is sometimes referred to as Equilibrium Headspace because the vapour phase above the sample in a sealed vial is brought into thermodynamic equilibrium with the sample prior to analysis. Volatile organic compounds are thermally extracted from the sample matrix by Static Headspace sampling and introduced directly to a thick-film megabore GC column. Compounds eluting from the GC column are introduced into a Mass Spectrometer and target compounds are identified by a comparison of both retention times and mass spectra with those of reference standards stored in a computerized data system. Reference retention times and mass spectra for analytes are obtained by the measurement of calibration standards under precisely the same extraction conditions used for samples. The concentration of each identified component is measured by relating the response of the sample quantitation mass ions to the response of the quantitation mass ions of a series of external standards. A deuterated surrogate recovery standard is added to provide information on matrix effects. | 0.005 | Y | Y | Y | 381 | 9/28/1998 3:08:00 PM | 8/14/2024 | 8/16/2024 4:05:21 PM | 1996-09-16 | 6/3/1998 6:51:45 PM | MPAK |  | Concentration | mg/L | Y | (fl. conc.) |  | Tetrachloroethylene (fl. conc.) | 127-18-4 | Volatile Organic Compounds (VOCs); | Volatile Organic Compounds (VOCs); EMS code: T030 | Chemical |  | LAB
T030 | ug/g | Tetrachloroethylene | MS02 | Purge&Trap GC/MS (VOC-S) | Purge&Trap GC/MS (VOC-S) | 0.001 | Y | Y | Y | 27 | 11/19/2004 | 10/28/2022 5:04:00 PM | 1/11/2024 1:07:08 PM | 1996-04-12 | 6/3/1998 6:51:45 PM | EMS |  | MassPerMass | mg/kg | Y | (mass conc.) |  | Tetrachloroethylene (mass conc.) | 127-18-4 | Volatile Organic Compounds (VOCs); | Volatile Organic Compounds (VOCs); EMS code: T030 | Chemical |  | LAB
T030 | mg/L | Tetrachloroethylene | X384 | Purge and trap GC/MS | The sample is analyzed by purge and trap gas chromatography with detection by mass spectrometry. The procedure involves purging the volatile compounds from the sample with an inert gas, and trapping them on a solid sorbent. When purging of the sample is complete, the trap is heated and the compounds are transferred to the gas chromatographic column. Analysis is then accomplished by separation of the components by gas chromatography with detection by mass spectrometry. | 0.001 | Y | Y | Y | 37 | 8/5/2004 10:31:00 AM | 7/26/2012 7:12:00 AM | 9/26/2012 1:07:09 PM | 1996-04-12 | 6/3/1998 6:51:45 PM | EMS |  | Concentration | mg/L | Y | (fl. conc.) |  | Tetrachloroethylene (fl. conc.) | 127-18-4 | Volatile Organic Compounds (VOCs); | Volatile Organic Compounds (VOCs); EMS code: T030 | Chemical |  | LAB
T030 | mg/L | Tetrachloroethylene | MS04 | Static Headspace GC/MS (Water) | (Water) A static headspace procedure is used as a means of driving the analytes out of the matrix into a headspace. Once equilibrium is achieved, a fixed volume of headspeace is taken for subsequent determination by Gas Chromatography - Mass Spectroscopy (GC/MS) using Selected Ion Monitoring. | 0.0005 | Y | Y | Y | 101 | 8/13/2009 | 6/28/2024 12:29:00 AM | 7/23/2024 11:41:50 AM | 2010-01-13 |  | RDGIBSON |  | Concentration | mg/L | Y | (fl. conc.) |  | Tetrachloroethylene (fl. conc.) | 127-18-4 | Volatile Organic Compounds (VOCs); | Volatile Organic Compounds (VOCs); EMS code: T030 | Chemical |  | LAB
T030 | ug/g | Tetrachloroethylene | FS04 | Static Headspace:HRGC/LRMS | Static Headspace anaylsis is a method for the determination of volatile compounds in liquids and/or solids by analysing the vapour phase of a sample in a closed system; in this case a sealed vial. Static Headspace is sometimes referred to as Equilibrium Headspace because the vapour phase above the sample in a sealed vial is brought into thermodynamic equilibrium with the sample prior to analysis. Volatile organic compounds are thermally extracted from the sample matrix by Static Headspace sampling and introduced directly to a thick-film megabore GC column. Compounds eluting from the GC column are introduced into a Mass Spectrometer and target compounds are identified by a comparison of both retention times and mass spectra with those of reference standards stored in a computerized data system. Reference retention times and mass spectra for analytes are obtained by the measurement of calibration standards under precisely the same extraction conditions used for samples. The concentration of each identified component is measured by relating the response of the sample quantitation mass ions to the response of the quantitation mass ions of a series of external standards. A deuterated surrogate recovery standard is added to provide information on matrix effects. | 0.05 | Y | Y | Y | 44 | 6/30/1999 4:41:00 PM | 5/22/2015 | 7/3/2015 3:19:04 PM | 1996-09-16 | 6/3/1998 6:51:45 PM | MPAK |  | MassPerMass | mg/kg | Y | (mass conc.) |  | Tetrachloroethylene (mass conc.) | 127-18-4 | Volatile Organic Compounds (VOCs); | Volatile Organic Compounds (VOCs); EMS code: T030 | Chemical |  | LAB
T030 | mg/L | Tetrachloroethylene | MS01 | Purge&Trap GC/MS (VOC-W) | Purge&Trap GC/MS (VOC-W) | 0.2 | Y | Y | Y | 407 | 4/15/2002 | 6/4/2024 10:24:00 AM | 6/13/2024 5:01:27 PM | 1996-04-12 | 8/12/2002 4:12:46 PM | EMS | NPEPPIN | Concentration | mg/L | Y | (fl. conc.) |  | Tetrachloroethylene (fl. conc.) | 127-18-4 | Volatile Organic Compounds (VOCs); | Volatile Organic Compounds (VOCs); EMS code: T030 | Chemical |  | LAB
T031 | mg/L | Trichlorpyr | F052 | Acid,DCM extr'n,methn,Florisil,HRGC-EC | Effluent and water samples are acidified with Sulphuric Acid and extracted with Dichloromethane. After concentration, the extract is methylated using ethereal Diazomethane. Extracts are cleaned up as required using Florisil. The extracts are analysed for Acid Extractable Herbicides using High Resolution Gas Chromatography with Electron Capture detection. Positive results may be verified by use of a secondary confirmation GC column. | 5e-05 | Y | Y | Y | 33 | 12/8/1998 3:39:00 PM | 12/31/2001 12:00:00 PM | 6/18/2003 2:53:57 PM | 1997-10-20 | 6/3/1998 6:51:47 PM | NPEPPIN |  | Concentration | mg/L | Y | (fl. conc.) |  | Trichlorpyr (fl. conc.) | 55335-06-3 | Pesticides; | Pesticides; EMS code: T031 | Chemical |  | LAB
T031 | mg/L | Trichlorpyr | AEWA | H2SO4 acid,CHCl2 extr.florisil. GC/MS | Effluent and water samples are acidified with Sulphuric Acid and extracted with Dichloromethane. After concentration, the extract is methylated using ethereal Diazomethane. Extracts are cleaned up as required using Florisil. The extracts are analyzed for Acid Extractable Herbicides using GC/MS. | 0.0001 | Y | Y | Y | 13 | 6/10/2002 3:20:00 PM | 7/29/2005 7:10:00 AM | 6/2/2009 12:01:12 PM | 1996-04-12 | 6/3/1998 6:51:42 PM | EMS |  | Concentration | mg/L | Y | (fl. conc.) |  | Trichlorpyr (fl. conc.) | 55335-06-3 | Pesticides; | Pesticides; EMS code: T031 | Chemical |  | LAB
T031 | ug/g | Trichlorpyr | AEAN | H2SO4,CHCL2/MeOH extr.florisil. GC/MS | Sediment samples are shaken with a mixture of Dichloromethane, Methanol and Sulphuric Acid, filtered, then extracted with Dichloromethane. After concentration, the extract is methylated using ethereal Diazomethane. Extracts are cleaned up as required using Florisil. The extracts are analyzed for Acid Extractable Herbicides using GC/MS. | 0.005 | Y | Y | Y | 2 | 9/25/1997 11:36:00 AM | 10/31/1997 9:38:00 AM | 11/3/1997 7:36:41 PM | 1996-04-12 | 6/3/1998 6:51:40 PM | EMS |  | MassPerMass | mg/kg | Y | (mass conc.) |  | Trichlorpyr (mass conc.) | 55335-06-3 | Pesticides; | Pesticides; EMS code: T031 | Chemical |  | LAB
T031 | ug/g | Trichlorpyr | FS15 | DCM/MEOH acid extr'n,DCM,methn,Florisil,HRGC-EC | Sediment samples are shaken with a mixture of Dichloromethane, Methanol and Sulphuric Acid, filtered, then extracted with Dichloromethane. After concentration, the extract is methylated using ethereal Diazomethane. Extracts are cleaned up as required using Florisil. The extracts are analysed for Acid Extractable Herbicides using High Resolution Gas Chromatography with Electron Capture detection. Positive results may be verified by use of a secondary confirmation GC column. | 0.05 | Y | Y | Y | 46 | 4/24/1997 4:15:00 PM | 6/24/2003 | 2/26/2004 1:05:31 AM | 1997-10-20 | 6/3/1998 6:51:47 PM | NPEPPIN |  | MassPerMass | mg/kg | Y | (mass conc.) |  | Trichlorpyr (mass conc.) | 55335-06-3 | Pesticides; | Pesticides; EMS code: T031 | Chemical |  | LAB
T032 | ug/g | 2;4;5-Trichlorophenol | FS15 | DCM/MEOH acid extr'n,DCM,methn,Florisil,HRGC-EC | Sediment samples are shaken with a mixture of Dichloromethane, Methanol and Sulphuric Acid, filtered, then extracted with Dichloromethane. After concentration, the extract is methylated using ethereal Diazomethane. Extracts are cleaned up as required using Florisil. The extracts are analysed for Acid Extractable Herbicides using High Resolution Gas Chromatography with Electron Capture detection. Positive results may be verified by use of a secondary confirmation GC column. | 0.05 | Y | Y | Y | 35 | 4/24/1997 4:15:00 PM | 6/24/2003 | 2/26/2004 1:05:31 AM | 1997-10-20 | 6/3/1998 6:51:47 PM | NPEPPIN |  | MassPerMass | mg/kg | Y | (mass conc.) |  | 2;4;5-Trichlorophenol (mass conc.) | 95-95-4 | Organics; | Organics; EMS code: T032 | Chemical |  | LAB
T032 | mg/L | 2;4;5-Trichlorophenol | AEWA | H2SO4 acid,CHCl2 extr.florisil. GC/MS | Effluent and water samples are acidified with Sulphuric Acid and extracted with Dichloromethane. After concentration, the extract is methylated using ethereal Diazomethane. Extracts are cleaned up as required using Florisil. The extracts are analyzed for Acid Extractable Herbicides using GC/MS. | 0.0001 | Y | Y | Y | 6 | 6/16/2003 | 7/29/2005 7:10:00 AM | 6/2/2009 12:01:12 PM | 1996-04-12 | 6/3/1998 6:51:42 PM | EMS |  | Concentration | mg/L | Y | (fl. conc.) |  | 2;4;5-Trichlorophenol (fl. conc.) | 95-95-4 | Organics; | Organics; EMS code: T032 | Chemical |  | LAB
T032 | mg/L | 2;4;5-Trichlorophenol | EQ01 | EQUIS (mg/L) | EQUIS (mg/L) | 0 | N | Y | Y | 2 |  |  | 10/22/1996 5:23:43 PM | 1996-04-12 | 6/3/1998 6:51:40 PM | EMS |  | Concentration | mg/L | Y | (fl. conc.) |  | 2;4;5-Trichlorophenol (fl. conc.) | 95-95-4 | Organics; | Organics; EMS code: T032 | Chemical |  | LAB
T032 | mg/L | 2;4;5-Trichlorophenol | F052 | Acid,DCM extr'n,methn,Florisil,HRGC-EC | Effluent and water samples are acidified with Sulphuric Acid and extracted with Dichloromethane. After concentration, the extract is methylated using ethereal Diazomethane. Extracts are cleaned up as required using Florisil. The extracts are analysed for Acid Extractable Herbicides using High Resolution Gas Chromatography with Electron Capture detection. Positive results may be verified by use of a secondary confirmation GC column. | 5e-05 | Y | Y | Y | 33 | 12/8/1998 3:39:00 PM | 12/31/2001 12:00:00 PM | 6/18/2003 2:53:57 PM | 1997-10-20 | 6/3/1998 6:51:47 PM | NPEPPIN |  | Concentration | mg/L | Y | (fl. conc.) |  | 2;4;5-Trichlorophenol (fl. conc.) | 95-95-4 | Organics; | Organics; EMS code: T032 | Chemical |  | LAB
T032 | ug/g | 2;4;5-Trichlorophenol | X212 | Hex/AcExtr,Part,IonEx,GC | Hex/AcExtr;Part;IonEx;GC | 0.005 | Y | Y | Y | 78 |  |  | 10/24/1996 2:26:16 PM | 1996-04-12 | 6/3/1998 6:51:40 PM | EMS |  | MassPerMass | mg/kg | Y | (mass conc.) |  | 2;4;5-Trichlorophenol (mass conc.) | 95-95-4 | Organics; | Organics; EMS code: T032 | Chemical |  | LAB
T032 | ug/g | 2;4;5-Trichlorophenol | DCMA | DCM/ Acetone/MTBE extr, AA/TMA deriv, GCMS | DCM/ Acetone/MTBE extr; AA/TMA deriv; GCMS. EPA 3570: 8270: Knapp(1979) | 0.02 | Y | Y | Y | 51 | 5/22/2015 | 3/20/2020 | 2/21/2023 2:10:29 PM | 2015-05-26 |  | RDGIBSON |  | MassPerMass | mg/kg | Y | (mass conc.) |  | 2;4;5-Trichlorophenol (mass conc.) | 95-95-4 | Organics; | Organics; EMS code: T032 | Chemical |  | LAB
T032 | mg/L | 2;4;5-Trichlorophenol | X268 | CH2Cl2 ext:GC/MS | CH2Cl2 ext:GC/MS |  | Y | Y | Y | 34 | 12/28/2021 | 6/7/2024 | 6/26/2024 4:28:10 PM | 2022-04-21 |  | EMS |  | Concentration | mg/L | Y | (fl. conc.) |  | 2;4;5-Trichlorophenol (fl. conc.) | 95-95-4 | Organics; | Organics; EMS code: T032 | Chemical |  | LAB
T032 | ug/g | 2;4;5-Trichlorophenol | 1332 | Phenolics by Tumbler/GC-MS | Phenolics by Tumbler/GC-MS | 0.02 | Y | Y | Y | 15 | 1/17/2022 | 2/2/2022 | 1/11/2024 1:07:08 PM | 2022-07-07 |  | EMS |  | MassPerMass | mg/kg | Y | (mass conc.) |  | 2;4;5-Trichlorophenol (mass conc.) | 95-95-4 | Organics; | Organics; EMS code: T032 | Chemical |  | LAB
T033 | mg/L | Trichlorophenol2;3;4 | X268 | CH2Cl2 ext:GC/MS | CH2Cl2 ext:GC/MS |  | Y | Y | Y | 34 | 12/28/2021 | 6/7/2024 | 6/26/2024 4:28:10 PM | 2022-04-21 |  | EMS |  | Concentration | mg/L | Y | (fl. conc.) |  | Trichlorophenol2;3;4 (fl. conc.) | 15950-66-0 | Organics; | Organics; EMS code: T033 | Chemical |  | LAB
T033 | ug/g | Trichlorophenol2;3;4 | AEAN | H2SO4,CHCL2/MeOH extr.florisil. GC/MS | Sediment samples are shaken with a mixture of Dichloromethane, Methanol and Sulphuric Acid, filtered, then extracted with Dichloromethane. After concentration, the extract is methylated using ethereal Diazomethane. Extracts are cleaned up as required using Florisil. The extracts are analyzed for Acid Extractable Herbicides using GC/MS. | 0.005 | Y | Y | Y | 39 |  |  | 10/24/1996 12:38:38 PM | 1996-04-12 | 6/3/1998 6:51:40 PM | EMS |  | MassPerMass | mg/kg | Y | (mass conc.) |  | Trichlorophenol2;3;4 (mass conc.) | 15950-66-0 | Organics; | Organics; EMS code: T033 | Chemical |  | LAB
T033 | ug/g | Trichlorophenol2;3;4 | X212 | Hex/AcExtr,Part,IonEx,GC | Hex/AcExtr;Part;IonEx;GC | 0.005 | Y | Y | Y | 78 |  |  | 10/24/1996 2:26:16 PM | 1996-04-12 | 6/3/1998 6:51:40 PM | EMS |  | MassPerMass | mg/kg | Y | (mass conc.) |  | Trichlorophenol2;3;4 (mass conc.) | 15950-66-0 | Organics; | Organics; EMS code: T033 | Chemical |  | LAB
T033 | ug/g | Trichlorophenol2;3;4 | 1332 | Phenolics by Tumbler/GC-MS | Phenolics by Tumbler/GC-MS | 0.02 | Y | Y | Y | 22 | 1/17/2022 | 9/22/2022 8:16:00 PM | 1/11/2024 1:07:08 PM | 2022-07-07 |  | EMS |  | MassPerMass | mg/kg | Y | (mass conc.) |  | Trichlorophenol2;3;4 (mass conc.) | 15950-66-0 | Organics; | Organics; EMS code: T033 | Chemical |  | LAB
T033 | ug/g | Trichlorophenol2;3;4 | DCMA | DCM/ Acetone/MTBE extr, AA/TMA deriv, GCMS | DCM/ Acetone/MTBE extr; AA/TMA deriv; GCMS. EPA 3570: 8270: Knapp(1979) | 0.02 | Y | Y | Y | 51 | 5/22/2015 | 3/20/2020 | 2/21/2023 2:10:29 PM | 2015-05-26 |  | RDGIBSON |  | MassPerMass | mg/kg | Y | (mass conc.) |  | Trichlorophenol2;3;4 (mass conc.) | 15950-66-0 | Organics; | Organics; EMS code: T033 | Chemical |  | LAB
T033 | mg/L | Trichlorophenol2;3;4 | F028 | Acid,Extr.MeCl,HRGC/LRMS:TIM,SIM | A measured volume of sample is acidified and extracted with Methylene Chloride. For samples which have a high organic content that potentially produce chromatographic interferences an initial pre-extraction under basic conditions is performed. The sample extract is then evaporated to small volume and analysed by High Resolution Gas Chromatography / Low Resolution Mass Spectrometry (HRGC/LRMS). LRMS data is acquired in Total Ion Mode (TIM) and quantitation is performed using a single mass ion. Positive Identification of the various target compounds is made by comparison of absolute retention times to those of standards combined with computer based mass spectral mass/intensity library searching and matching routines. | 0.0001 | Y | Y | Y | 70 | 8/20/2009 | 11/1/2022 4:39:00 AM | 4/21/2023 4:26:45 PM | 2009-12-23 |  | RDGIBSON |  | Concentration | mg/L | Y | (fl. conc.) |  | Trichlorophenol2;3;4 (fl. conc.) | 15950-66-0 | Organics; | Organics; EMS code: T033 | Chemical |  | LAB
T033 | mg/L | Trichlorophenol2;3;4 | AEWA | H2SO4 acid,CHCl2 extr.florisil. GC/MS | Effluent and water samples are acidified with Sulphuric Acid and extracted with Dichloromethane. After concentration, the extract is methylated using ethereal Diazomethane. Extracts are cleaned up as required using Florisil. The extracts are analyzed for Acid Extractable Herbicides using GC/MS. | 0.0001 | Y | Y | Y | 72 | 12/3/2002 | 3/11/2003 | 4/1/2004 1:13:28 AM | 1996-04-12 | 6/3/1998 6:51:42 PM | EMS |  | Concentration | mg/L | Y | (fl. conc.) |  | Trichlorophenol2;3;4 (fl. conc.) | 15950-66-0 | Organics; | Organics; EMS code: T033 | Chemical |  | LAB
T034 | mg/L | Trichlorophenol2;3;5 | X268 | CH2Cl2 ext:GC/MS | CH2Cl2 ext:GC/MS |  | Y | Y | Y | 34 | 12/28/2021 | 6/7/2024 | 6/26/2024 4:28:10 PM | 2022-04-21 |  | EMS |  | Concentration | mg/L | Y | (fl. conc.) |  | Trichlorophenol2;3;5 (fl. conc.) | 933-78-8 | Organics; | Organics; EMS code: T034 | Chemical |  | LAB
T034 | mg/L | Trichlorophenol2;3;5 | AEWA | H2SO4 acid,CHCl2 extr.florisil. GC/MS | Effluent and water samples are acidified with Sulphuric Acid and extracted with Dichloromethane. After concentration, the extract is methylated using ethereal Diazomethane. Extracts are cleaned up as required using Florisil. The extracts are analyzed for Acid Extractable Herbicides using GC/MS. | 0.0001 | Y | Y | Y | 72 | 12/3/2002 | 3/11/2003 | 4/1/2004 1:13:28 AM | 1996-04-12 | 6/3/1998 6:51:45 PM | EMS |  | Concentration | mg/L | Y | (fl. conc.) |  | Trichlorophenol2;3;5 (fl. conc.) | 933-78-8 | Organics; | Organics; EMS code: T034 | Chemical |  | LAB
T034 | ug/g | Trichlorophenol2;3;5 | 1332 | Phenolics by Tumbler/GC-MS | Phenolics by Tumbler/GC-MS | 0.02 | Y | Y | Y | 22 | 1/17/2022 | 9/22/2022 8:16:00 PM | 1/11/2024 1:07:08 PM | 2022-07-07 |  | EMS |  | MassPerMass | mg/kg | Y | (mass conc.) |  | Trichlorophenol2;3;5 (mass conc.) | 933-78-8 | Organics; | Organics; EMS code: T034 | Chemical |  | LAB
T034 | ug/g | Trichlorophenol2;3;5 | AEAN | H2SO4,CHCL2/MeOH extr.florisil. GC/MS | Sediment samples are shaken with a mixture of Dichloromethane, Methanol and Sulphuric Acid, filtered, then extracted with Dichloromethane. After concentration, the extract is methylated using ethereal Diazomethane. Extracts are cleaned up as required using Florisil. The extracts are analyzed for Acid Extractable Herbicides using GC/MS. | 0.005 | Y | Y | Y | 39 |  |  | 10/24/1996 12:38:38 PM | 1996-04-12 | 6/3/1998 6:51:40 PM | EMS |  | MassPerMass | mg/kg | Y | (mass conc.) |  | Trichlorophenol2;3;5 (mass conc.) | 933-78-8 | Organics; | Organics; EMS code: T034 | Chemical |  | LAB
T034 | ug/g | Trichlorophenol2;3;5 | X212 | Hex/AcExtr,Part,IonEx,GC | Hex/AcExtr;Part;IonEx;GC | 0.005 | Y | Y | Y | 78 |  |  | 10/24/1996 2:26:16 PM | 1996-04-12 | 6/3/1998 6:51:40 PM | EMS |  | MassPerMass | mg/kg | Y | (mass conc.) |  | Trichlorophenol2;3;5 (mass conc.) | 933-78-8 | Organics; | Organics; EMS code: T034 | Chemical |  | LAB
T034 | ug/g | Trichlorophenol2;3;5 | DCMA | DCM/ Acetone/MTBE extr, AA/TMA deriv, GCMS | DCM/ Acetone/MTBE extr; AA/TMA deriv; GCMS. EPA 3570: 8270: Knapp(1979) | 0.02 | Y | Y | Y | 51 | 5/22/2015 | 3/20/2020 | 2/21/2023 2:10:29 PM | 2015-05-26 |  | RDGIBSON |  | MassPerMass | mg/kg | Y | (mass conc.) |  | Trichlorophenol2;3;5 (mass conc.) | 933-78-8 | Organics; | Organics; EMS code: T034 | Chemical |  | LAB
T034 | mg/L | Trichlorophenol2;3;5 | F028 | Acid,Extr.MeCl,HRGC/LRMS:TIM,SIM | A measured volume of sample is acidified and extracted with Methylene Chloride. For samples which have a high organic content that potentially produce chromatographic interferences an initial pre-extraction under basic conditions is performed. The sample extract is then evaporated to small volume and analysed by High Resolution Gas Chromatography / Low Resolution Mass Spectrometry (HRGC/LRMS). LRMS data is acquired in Total Ion Mode (TIM) and quantitation is performed using a single mass ion. Positive Identification of the various target compounds is made by comparison of absolute retention times to those of standards combined with computer based mass spectral mass/intensity library searching and matching routines. | 0.0001 | Y | Y | Y | 70 | 8/20/2009 | 11/1/2022 4:39:00 AM | 4/21/2023 4:26:45 PM | 2009-12-23 |  | RDGIBSON |  | Concentration | mg/L | Y | (fl. conc.) |  | Trichlorophenol2;3;5 (fl. conc.) | 933-78-8 | Organics; | Organics; EMS code: T034 | Chemical |  | LAB
T035 | ug/g | Trichlorophenol2;3;6 | AEAN | H2SO4,CHCL2/MeOH extr.florisil. GC/MS | Sediment samples are shaken with a mixture of Dichloromethane, Methanol and Sulphuric Acid, filtered, then extracted with Dichloromethane. After concentration, the extract is methylated using ethereal Diazomethane. Extracts are cleaned up as required using Florisil. The extracts are analyzed for Acid Extractable Herbicides using GC/MS. | 0.005 | Y | Y | Y | 39 |  |  | 10/24/1996 12:38:38 PM | 1996-04-12 | 6/3/1998 6:51:40 PM | EMS |  | MassPerMass | mg/kg | Y | (mass conc.) |  | Trichlorophenol2;3;6 (mass conc.) | 933-75-5 | Organics; | Organics; EMS code: T035 | Chemical |  | LAB
T035 | mg/L | Trichlorophenol2;3;6 | F028 | Acid,Extr.MeCl,HRGC/LRMS:TIM,SIM | A measured volume of sample is acidified and extracted with Methylene Chloride. For samples which have a high organic content that potentially produce chromatographic interferences an initial pre-extraction under basic conditions is performed. The sample extract is then evaporated to small volume and analysed by High Resolution Gas Chromatography / Low Resolution Mass Spectrometry (HRGC/LRMS). LRMS data is acquired in Total Ion Mode (TIM) and quantitation is performed using a single mass ion. Positive Identification of the various target compounds is made by comparison of absolute retention times to those of standards combined with computer based mass spectral mass/intensity library searching and matching routines. | 0.0001 | Y | Y | Y | 70 | 8/20/2009 | 11/1/2022 4:39:00 AM | 4/21/2023 4:26:45 PM | 2009-12-23 |  | RDGIBSON |  | Concentration | mg/L | Y | (fl. conc.) |  | Trichlorophenol2;3;6 (fl. conc.) | 933-75-5 | Organics; | Organics; EMS code: T035 | Chemical |  | LAB
T035 | ug/g | Trichlorophenol2;3;6 | 1332 | Phenolics by Tumbler/GC-MS | Phenolics by Tumbler/GC-MS | 0.02 | Y | Y | Y | 22 | 1/17/2022 | 9/22/2022 8:16:00 PM | 1/11/2024 1:07:08 PM | 2022-07-07 |  | EMS |  | MassPerMass | mg/kg | Y | (mass conc.) |  | Trichlorophenol2;3;6 (mass conc.) | 933-75-5 | Organics; | Organics; EMS code: T035 | Chemical |  | LAB
T035 | ug/g | Trichlorophenol2;3;6 | X212 | Hex/AcExtr,Part,IonEx,GC | Hex/AcExtr;Part;IonEx;GC | 0.005 | Y | Y | Y | 78 |  |  | 10/24/1996 2:26:16 PM | 1996-04-12 | 6/3/1998 6:51:45 PM | EMS |  | MassPerMass | mg/kg | Y | (mass conc.) |  | Trichlorophenol2;3;6 (mass conc.) | 933-75-5 | Organics; | Organics; EMS code: T035 | Chemical |  | LAB
T035 | mg/L | Trichlorophenol2;3;6 | AEWA | H2SO4 acid,CHCl2 extr.florisil. GC/MS | Effluent and water samples are acidified with Sulphuric Acid and extracted with Dichloromethane. After concentration, the extract is methylated using ethereal Diazomethane. Extracts are cleaned up as required using Florisil. The extracts are analyzed for Acid Extractable Herbicides using GC/MS. | 0.0001 | Y | Y | Y | 64 |  |  | 3/17/1997 1:23:13 PM | 1996-04-12 | 6/3/1998 6:51:42 PM | EMS |  | Concentration | mg/L | Y | (fl. conc.) |  | Trichlorophenol2;3;6 (fl. conc.) | 933-75-5 | Organics; | Organics; EMS code: T035 | Chemical |  | LAB
T035 | mg/L | Trichlorophenol2;3;6 | X268 | CH2Cl2 ext:GC/MS | CH2Cl2 ext:GC/MS | 0.0005 | Y | Y | Y | 34 | 12/28/2021 | 6/7/2024 | 6/26/2024 4:28:10 PM | 2012-06-25 |  | RDGIBSON |  | Concentration | mg/L | Y | (fl. conc.) |  | Trichlorophenol2;3;6 (fl. conc.) | 933-75-5 | Organics; | Organics; EMS code: T035 | Chemical |  | LAB
T035 | ug/g | Trichlorophenol2;3;6 | DCMA | DCM/ Acetone/MTBE extr, AA/TMA deriv, GCMS | DCM/ Acetone/MTBE extr; AA/TMA deriv; GCMS. EPA 3570: 8270: Knapp(1979) | 0.02 | Y | Y | Y | 51 | 5/22/2015 | 3/20/2020 | 2/21/2023 2:10:29 PM | 2015-05-26 |  | RDGIBSON |  | MassPerMass | mg/kg | Y | (mass conc.) |  | Trichlorophenol2;3;6 (mass conc.) | 933-75-5 | Organics; | Organics; EMS code: T035 | Chemical |  | LAB
T036 | ug/g | 2;3;4;5-Tetrachlorophenol | DCMA | DCM/ Acetone/MTBE extr, AA/TMA deriv, GCMS | DCM/ Acetone/MTBE extr; AA/TMA deriv; GCMS. EPA 3570: 8270: Knapp(1979) | 0.02 | Y | Y | Y | 51 | 5/22/2015 | 3/20/2020 | 2/21/2023 2:10:29 PM | 2015-05-26 |  | RDGIBSON |  | MassPerMass | mg/kg | Y | (mass conc.) |  | 2;3;4;5-Tetrachlorophenol (mass conc.) | 4901-51-3 | Organics; | Organics; EMS code: T036 | Chemical |  | LAB
T036 | ug/g | 2;3;4;5-Tetrachlorophenol | FS11 | Acetone extr,Hexane,Diazometh.Methyl'n,HRGC/Electron Capture | Sediment and homogenized biota samples to be analysed for Chlorinated Phenols are Acetone extracted using a wrist-action shaker, partitioned with 2% NaCl and Hexane, methylated, then cleaned up and analysed in the manner outlined for effluent samples. Sediment and homogenized biota samples to be analysed for Chlorinated Anisoles are Acetone extracted using a wrist-action shaker, partitioned with 2% NaCl and Hexane, then cleaned up and analysed in the manner outlined for effluent samples. The extracts are analysed for Chlorinated Anisoles and/or Chlorinated Phenols using High Resolution Gas Chromatography with Electron Capture detection. | 0.0001 | Y | Y | Y | 44 | 2/22/2000 8:31:00 AM | 1/9/2001 4:56:00 PM | 8/21/2002 2:10:11 PM | 1997-01-14 | 6/3/1998 6:51:46 PM | MPAK |  | MassPerMass | mg/kg | Y | (mass conc.) |  | 2;3;4;5-Tetrachlorophenol (mass conc.) | 4901-51-3 | Organics; | Organics; EMS code: T036 | Chemical |  | LAB
T036 | ug/g | 2;3;4;5-Tetrachlorophenol | 1332 | Phenolics by Tumbler/GC-MS | Phenolics by Tumbler/GC-MS | 0.02 | Y | Y | Y | 15 | 1/17/2022 | 2/2/2022 | 1/11/2024 1:07:08 PM | 2022-07-07 |  | EMS |  | MassPerMass | mg/kg | Y | (mass conc.) |  | 2;3;4;5-Tetrachlorophenol (mass conc.) | 4901-51-3 | Organics; | Organics; EMS code: T036 | Chemical |  | LAB
T036 | ug/g | 2;3;4;5-Tetrachlorophenol | AEAN | H2SO4,CHCL2/MeOH extr.florisil. GC/MS | Sediment samples are shaken with a mixture of Dichloromethane, Methanol and Sulphuric Acid, filtered, then extracted with Dichloromethane. After concentration, the extract is methylated using ethereal Diazomethane. Extracts are cleaned up as required using Florisil. The extracts are analyzed for Acid Extractable Herbicides using GC/MS. | 0.005 | Y | Y | Y | 39 |  |  | 10/24/1996 12:38:38 PM | 1996-04-12 | 6/3/1998 6:51:40 PM | EMS |  | MassPerMass | mg/kg | Y | (mass conc.) |  | 2;3;4;5-Tetrachlorophenol (mass conc.) | 4901-51-3 | Organics; | Organics; EMS code: T036 | Chemical |  | LAB
T036 | mg/L | 2;3;4;5-Tetrachlorophenol | AEWA | H2SO4 acid,CHCl2 extr.florisil. GC/MS | Effluent and water samples are acidified with Sulphuric Acid and extracted with Dichloromethane. After concentration, the extract is methylated using ethereal Diazomethane. Extracts are cleaned up as required using Florisil. The extracts are analyzed for Acid Extractable Herbicides using GC/MS. | 0.0001 | Y | Y | Y | 77 | 12/3/2002 | 6/16/2003 | 4/1/2004 1:13:28 AM | 1996-04-12 | 6/3/1998 6:51:42 PM | EMS |  | Concentration | mg/L | Y | (fl. conc.) |  | 2;3;4;5-Tetrachlorophenol (fl. conc.) | 4901-51-3 | Organics; | Organics; EMS code: T036 | Chemical |  | LAB
T036 | mg/L | 2;3;4;5-Tetrachlorophenol | X268 | CH2Cl2 ext:GC/MS | CH2Cl2 ext:GC/MS |  | Y | Y | Y | 34 | 12/28/2021 | 6/7/2024 | 6/26/2024 4:28:10 PM | 2022-04-21 |  | EMS |  | Concentration | mg/L | Y | (fl. conc.) |  | 2;3;4;5-Tetrachlorophenol (fl. conc.) | 4901-51-3 | Organics; | Organics; EMS code: T036 | Chemical |  | LAB
T036 | mg/L | 2;3;4;5-Tetrachlorophenol | F028 | Acid,Extr.MeCl,HRGC/LRMS:TIM,SIM | A measured volume of sample is acidified and extracted with Methylene Chloride. For samples which have a high organic content that potentially produce chromatographic interferences an initial pre-extraction under basic conditions is performed. The sample extract is then evaporated to small volume and analysed by High Resolution Gas Chromatography / Low Resolution Mass Spectrometry (HRGC/LRMS). LRMS data is acquired in Total Ion Mode (TIM) and quantitation is performed using a single mass ion. Positive Identification of the various target compounds is made by comparison of absolute retention times to those of standards combined with computer based mass spectral mass/intensity library searching and matching routines. | 0.0001 | Y | Y | Y | 24 | 8/20/2009 | 4/14/2015 12:26:00 AM | 4/16/2015 6:00:12 PM | 2009-12-23 |  | RDGIBSON |  | Concentration | mg/L | Y | (fl. conc.) |  | 2;3;4;5-Tetrachlorophenol (fl. conc.) | 4901-51-3 | Organics; | Organics; EMS code: T036 | Chemical |  | LAB
T036 | ug/g | 2;3;4;5-Tetrachlorophenol | FS15 | DCM/MEOH acid extr'n,DCM,methn,Florisil,HRGC-EC | Sediment samples are shaken with a mixture of Dichloromethane, Methanol and Sulphuric Acid, filtered, then extracted with Dichloromethane. After concentration, the extract is methylated using ethereal Diazomethane. Extracts are cleaned up as required using Florisil. The extracts are analysed for Acid Extractable Herbicides using High Resolution Gas Chromatography with Electron Capture detection. Positive results may be verified by use of a secondary confirmation GC column. | 0.01 | Y | Y | Y | 27 | 4/4/2002 | 6/24/2003 | 2/26/2004 1:05:31 AM | 1998-12-30 |  | NPEPPIN |  | MassPerMass | mg/kg | Y | (mass conc.) |  | 2;3;4;5-Tetrachlorophenol (mass conc.) | 4901-51-3 | Organics; | Organics; EMS code: T036 | Chemical |  | LAB
T036 | ug/g | 2;3;4;5-Tetrachlorophenol | X212 | Hex/AcExtr,Part,IonEx,GC | Hex/AcExtr;Part;IonEx;GC | 0.005 | Y | Y | Y | 78 |  |  | 10/24/1996 2:26:16 PM | 1996-04-12 | 6/3/1998 6:51:40 PM | EMS |  | MassPerMass | mg/kg | Y | (mass conc.) |  | 2;3;4;5-Tetrachlorophenol (mass conc.) | 4901-51-3 | Organics; | Organics; EMS code: T036 | Chemical |  | LAB
T036 | mg/L | 2;3;4;5-Tetrachlorophenol | F029 | Acid,DiethylEther extr,Diazometh.Methyl'n,HRGC/Electron Capt | Effluent samples to be analysed for Chlorinated Phenols are acidified with dilute Sulfuric Acid and extracted with Diethyl Ether. After concentration the extract is methylated using ethereal Diazomethane and cleaned up with concentrated Sulfuric Acid and Mercury. The extracts are analysed for Chlorinated Anisoles and/or Chlorinated Phenols using High Resolution Gas Chromatography with Electron Capture detection. | 1e-06 | Y | Y | Y | 68 | 7/16/1997 4:10:00 PM | 8/20/2001 12:43:00 PM | 9/10/2002 9:56:04 AM | 1997-01-14 | 6/3/1998 6:51:46 PM | MPAK |  | Concentration | mg/L | Y | (fl. conc.) |  | 2;3;4;5-Tetrachlorophenol (fl. conc.) | 4901-51-3 | Organics; | Organics; EMS code: T036 | Chemical |  | LAB
T037 | ug/g | 2;3;4;6-Tetrachlorophenol | DCMA | DCM/ Acetone/MTBE extr, AA/TMA deriv, GCMS | DCM/ Acetone/MTBE extr; AA/TMA deriv; GCMS. EPA 3570: 8270: Knapp(1979) | 0.02 | Y | Y | Y | 51 | 5/22/2015 | 3/20/2020 | 2/21/2023 2:10:29 PM | 2015-05-26 |  | RDGIBSON |  | MassPerMass | mg/kg | Y | (mass conc.) |  | 2;3;4;6-Tetrachlorophenol (mass conc.) | 58-90-2 | Organics; | Organics; EMS code: T037 | Chemical |  | LAB
T037 | ug/g | 2;3;4;6-Tetrachlorophenol | FS11 | Acetone extr,Hexane,Diazometh.Methyl'n,HRGC/Electron Capture | Sediment and homogenized biota samples to be analysed for Chlorinated Phenols are Acetone extracted using a wrist-action shaker, partitioned with 2% NaCl and Hexane, methylated, then cleaned up and analysed in the manner outlined for effluent samples. Sediment and homogenized biota samples to be analysed for Chlorinated Anisoles are Acetone extracted using a wrist-action shaker, partitioned with 2% NaCl and Hexane, then cleaned up and analysed in the manner outlined for effluent samples. The extracts are analysed for Chlorinated Anisoles and/or Chlorinated Phenols using High Resolution Gas Chromatography with Electron Capture detection. | 0.0001 | Y | Y | Y | 44 | 2/22/2000 8:31:00 AM | 1/9/2001 4:56:00 PM | 8/21/2002 2:10:11 PM | 2002-06-03 |  | NPEPPIN |  | MassPerMass | mg/kg | Y | (mass conc.) |  | 2;3;4;6-Tetrachlorophenol (mass conc.) | 58-90-2 | Organics; | Organics; EMS code: T037 | Chemical |  | LAB
T037 | ug/g | 2;3;4;6-Tetrachlorophenol | AEAN | H2SO4,CHCL2/MeOH extr.florisil. GC/MS | Sediment samples are shaken with a mixture of Dichloromethane, Methanol and Sulphuric Acid, filtered, then extracted with Dichloromethane. After concentration, the extract is methylated using ethereal Diazomethane. Extracts are cleaned up as required using Florisil. The extracts are analyzed for Acid Extractable Herbicides using GC/MS. | 0.005 | Y | Y | Y | 39 |  |  | 10/24/1996 12:38:38 PM | 1996-04-12 | 6/3/1998 6:51:40 PM | EMS |  | MassPerMass | mg/kg | Y | (mass conc.) |  | 2;3;4;6-Tetrachlorophenol (mass conc.) | 58-90-2 | Organics; | Organics; EMS code: T037 | Chemical |  | LAB
T037 | mg/L | 2;3;4;6-Tetrachlorophenol | X268 | CH2Cl2 ext:GC/MS | CH2Cl2 ext:GC/MS |  | Y | Y | Y | 34 | 12/28/2021 | 6/7/2024 | 6/26/2024 4:28:10 PM | 2022-04-21 |  | EMS |  | Concentration | mg/L | Y | (fl. conc.) |  | 2;3;4;6-Tetrachlorophenol (fl. conc.) | 58-90-2 | Organics; | Organics; EMS code: T037 | Chemical |  | LAB
T037 | mg/L | 2;3;4;6-Tetrachlorophenol | F028 | Acid,Extr.MeCl,HRGC/LRMS:TIM,SIM | A measured volume of sample is acidified and extracted with Methylene Chloride. For samples which have a high organic content that potentially produce chromatographic interferences an initial pre-extraction under basic conditions is performed. The sample extract is then evaporated to small volume and analysed by High Resolution Gas Chromatography / Low Resolution Mass Spectrometry (HRGC/LRMS). LRMS data is acquired in Total Ion Mode (TIM) and quantitation is performed using a single mass ion. Positive Identification of the various target compounds is made by comparison of absolute retention times to those of standards combined with computer based mass spectral mass/intensity library searching and matching routines. | 0.0001 | Y | Y | Y | 79 | 8/20/2009 | 11/1/2022 4:39:00 AM | 4/21/2023 4:26:45 PM | 2009-12-23 |  | RDGIBSON |  | Concentration | mg/L | Y | (fl. conc.) |  | 2;3;4;6-Tetrachlorophenol (fl. conc.) | 58-90-2 | Organics; | Organics; EMS code: T037 | Chemical |  | LAB
T037 | ug/g | 2;3;4;6-Tetrachlorophenol | 1332 | Phenolics by Tumbler/GC-MS | Phenolics by Tumbler/GC-MS | 0.02 | Y | Y | Y | 22 | 1/17/2022 | 9/22/2022 8:16:00 PM | 1/11/2024 1:07:08 PM | 2022-07-07 |  | EMS |  | MassPerMass | mg/kg | Y | (mass conc.) |  | 2;3;4;6-Tetrachlorophenol (mass conc.) | 58-90-2 | Organics; | Organics; EMS code: T037 | Chemical |  | LAB
T037 | ug/g | 2;3;4;6-Tetrachlorophenol | X212 | Hex/AcExtr,Part,IonEx,GC | Hex/AcExtr;Part;IonEx;GC | 0.005 | Y | Y | Y | 78 |  |  | 10/24/1996 2:26:16 PM | 1996-04-12 | 6/3/1998 6:51:40 PM | EMS |  | MassPerMass | mg/kg | Y | (mass conc.) |  | 2;3;4;6-Tetrachlorophenol (mass conc.) | 58-90-2 | Organics; | Organics; EMS code: T037 | Chemical |  | LAB
T037 | mg/L | 2;3;4;6-Tetrachlorophenol | F034 | High Res GC-ECD/Chloroacetylation | Potassium Carbonate is dissolved in a measured volume of neutralized sample. Any Chlorophenols in the solution are then chloroacetylated in situ with Chloroacetic Anhydride in the presence of Petroleum Ether (30-60 C). The Petroleum Ether extract is dried over anhydrous Sodium Sulfate, evaporated to small volume and analysed using High Resolution Gas Chromatography with Electron Capture Detection (HRGC-ECD). | 5e-05 | Y | Y | Y | 38 | 2/21/2000 10:00:00 AM | 8/20/2001 12:43:00 PM | 9/10/2002 9:56:04 AM | 1997-01-16 | 6/3/1998 6:51:46 PM | MPAK |  | Concentration | mg/L | Y | (fl. conc.) |  | 2;3;4;6-Tetrachlorophenol (fl. conc.) | 58-90-2 | Organics; | Organics; EMS code: T037 | Chemical |  | LAB
T037 | mg/L | 2;3;4;6-Tetrachlorophenol | AEWA | H2SO4 acid,CHCl2 extr.florisil. GC/MS | Effluent and water samples are acidified with Sulphuric Acid and extracted with Dichloromethane. After concentration, the extract is methylated using ethereal Diazomethane. Extracts are cleaned up as required using Florisil. The extracts are analyzed for Acid Extractable Herbicides using GC/MS. | 0.0001 | Y | Y | Y | 72 | 12/3/2002 | 3/11/2003 | 4/1/2004 1:13:28 AM | 1996-04-12 | 6/3/1998 6:51:42 PM | EMS |  | Concentration | mg/L | Y | (fl. conc.) |  | 2;3;4;6-Tetrachlorophenol (fl. conc.) | 58-90-2 | Organics; | Organics; EMS code: T037 | Chemical |  | LAB
T038 | ug/g | 2;3;5;6-Tetrachlorophenol | DCMA | DCM/ Acetone/MTBE extr, AA/TMA deriv, GCMS | DCM/ Acetone/MTBE extr; AA/TMA deriv; GCMS. EPA 3570: 8270: Knapp(1979) | 0.02 | Y | Y | Y | 51 | 5/22/2015 | 3/20/2020 | 2/21/2023 2:10:29 PM | 2015-05-26 |  | RDGIBSON |  | MassPerMass | mg/kg | Y | (mass conc.) |  | 2;3;5;6-Tetrachlorophenol (mass conc.) | 935-95-5 | Organics; | Organics; EMS code: T038 | Chemical |  | LAB
T038 | mg/L | 2;3;5;6-Tetrachlorophenol | F028 | Acid,Extr.MeCl,HRGC/LRMS:TIM,SIM | A measured volume of sample is acidified and extracted with Methylene Chloride. For samples which have a high organic content that potentially produce chromatographic interferences an initial pre-extraction under basic conditions is performed. The sample extract is then evaporated to small volume and analysed by High Resolution Gas Chromatography / Low Resolution Mass Spectrometry (HRGC/LRMS). LRMS data is acquired in Total Ion Mode (TIM) and quantitation is performed using a single mass ion. Positive Identification of the various target compounds is made by comparison of absolute retention times to those of standards combined with computer based mass spectral mass/intensity library searching and matching routines. | 0.0001 | Y | Y | Y | 24 | 8/20/2009 | 4/14/2015 12:26:00 AM | 4/16/2015 6:00:12 PM | 2009-12-23 |  | RDGIBSON |  | Concentration | mg/L | Y | (fl. conc.) |  | 2;3;5;6-Tetrachlorophenol (fl. conc.) | 935-95-5 | Organics; | Organics; EMS code: T038 | Chemical |  | LAB
T038 | ug/g | 2;3;5;6-Tetrachlorophenol | 1332 | Phenolics by Tumbler/GC-MS | Phenolics by Tumbler/GC-MS | 0.02 | Y | Y | Y | 15 | 1/17/2022 | 2/2/2022 | 1/11/2024 1:07:08 PM | 2022-07-07 |  | EMS |  | MassPerMass | mg/kg | Y | (mass conc.) |  | 2;3;5;6-Tetrachlorophenol (mass conc.) | 935-95-5 | Organics; | Organics; EMS code: T038 | Chemical |  | LAB
T038 | mg/L | 2;3;5;6-Tetrachlorophenol | X268 | CH2Cl2 ext:GC/MS | CH2Cl2 ext:GC/MS |  | Y | Y | Y | 34 | 12/28/2021 | 6/7/2024 | 6/26/2024 4:28:10 PM | 2022-04-21 |  | EMS |  | Concentration | mg/L | Y | (fl. conc.) |  | 2;3;5;6-Tetrachlorophenol (fl. conc.) | 935-95-5 | Organics; | Organics; EMS code: T038 | Chemical |  | LAB
T038 | ug/g | 2;3;5;6-Tetrachlorophenol | X212 | Hex/AcExtr,Part,IonEx,GC | Hex/AcExtr;Part;IonEx;GC | 0.005 | Y | Y | Y | 78 |  |  | 10/24/1996 2:26:16 PM | 1996-04-12 | 6/3/1998 6:51:45 PM | EMS |  | MassPerMass | mg/kg | Y | (mass conc.) |  | 2;3;5;6-Tetrachlorophenol (mass conc.) | 935-95-5 | Organics; | Organics; EMS code: T038 | Chemical |  | LAB
T040 | mg/L | Trichloroethane | MS01 | Purge&Trap GC/MS (VOC-W) | Purge&Trap GC/MS (VOC-W) | 0 | Y | Y | Y | 19 |  |  | 7/23/2024 11:41:50 AM | 1996-04-12 | 3/18/2003 2:49:40 PM | EMS | UMULLER | Concentration | mg/L | Y | (fl. conc.) |  | Trichloroethane (fl. conc.) | 25323-89-1, 71-55-6, 79-00-5 | Volatile Organic Compounds (VOCs); | Volatile Organic Compounds (VOCs); EMS code: T040 | Chemical |  | LAB
T042 | mg/L | Trichlorophenol2;4;6 | AEWA | H2SO4 acid,CHCl2 extr.florisil. GC/MS | Effluent and water samples are acidified with Sulphuric Acid and extracted with Dichloromethane. After concentration, the extract is methylated using ethereal Diazomethane. Extracts are cleaned up as required using Florisil. The extracts are analyzed for Acid Extractable Herbicides using GC/MS. | 0.0001 | Y | Y | Y | 64 |  |  | 3/17/1997 1:23:12 PM | 1996-04-12 | 6/3/1998 6:51:42 PM | EMS |  | Concentration | mg/L | Y | (fl. conc.) |  | Trichlorophenol2;4;6 (fl. conc.) | 88-06-2 | Organics; | Organics; EMS code: T042 | Chemical |  | LAB
T042 | mg/L | Trichlorophenol2;4;6 | X268 | CH2Cl2 ext:GC/MS | CH2Cl2 ext:GC/MS |  | Y | Y | Y | 34 | 12/28/2021 | 6/7/2024 | 6/26/2024 4:28:10 PM | 2022-04-21 |  | EMS |  | Concentration | mg/L | Y | (fl. conc.) |  | Trichlorophenol2;4;6 (fl. conc.) | 88-06-2 | Organics; | Organics; EMS code: T042 | Chemical |  | LAB
T042 | ug/g | Trichlorophenol2;4;6 | DCMA | DCM/ Acetone/MTBE extr, AA/TMA deriv, GCMS | DCM/ Acetone/MTBE extr; AA/TMA deriv; GCMS. EPA 3570: 8270: Knapp(1979) | 0.02 | Y | Y | Y | 51 | 5/22/2015 | 3/20/2020 | 2/21/2023 2:10:29 PM | 2015-05-26 |  | RDGIBSON |  | MassPerMass | mg/kg | Y | (mass conc.) |  | Trichlorophenol2;4;6 (mass conc.) | 88-06-2 | Organics; | Organics; EMS code: T042 | Chemical |  | LAB
T042 | ug/g | Trichlorophenol2;4;6 | 1332 | Phenolics by Tumbler/GC-MS | Phenolics by Tumbler/GC-MS | 0.02 | Y | Y | Y | 22 | 1/17/2022 | 9/22/2022 8:16:00 PM | 1/11/2024 1:07:08 PM | 2022-07-07 |  | EMS |  | MassPerMass | mg/kg | Y | (mass conc.) |  | Trichlorophenol2;4;6 (mass conc.) | 88-06-2 | Organics; | Organics; EMS code: T042 | Chemical |  | LAB
T042 | ug/g | Trichlorophenol2;4;6 | AEAN | H2SO4,CHCL2/MeOH extr.florisil. GC/MS | Sediment samples are shaken with a mixture of Dichloromethane, Methanol and Sulphuric Acid, filtered, then extracted with Dichloromethane. After concentration, the extract is methylated using ethereal Diazomethane. Extracts are cleaned up as required using Florisil. The extracts are analyzed for Acid Extractable Herbicides using GC/MS. | 0.005 | Y | Y | Y | 39 |  |  | 10/24/1996 12:38:38 PM | 1996-04-12 | 6/3/1998 6:51:40 PM | EMS |  | MassPerMass | mg/kg | Y | (mass conc.) |  | Trichlorophenol2;4;6 (mass conc.) | 88-06-2 | Organics; | Organics; EMS code: T042 | Chemical |  | LAB
T042 | mg/L | Trichlorophenol2;4;6 | F028 | Acid,Extr.MeCl,HRGC/LRMS:TIM,SIM | A measured volume of sample is acidified and extracted with Methylene Chloride. For samples which have a high organic content that potentially produce chromatographic interferences an initial pre-extraction under basic conditions is performed. The sample extract is then evaporated to small volume and analysed by High Resolution Gas Chromatography / Low Resolution Mass Spectrometry (HRGC/LRMS). LRMS data is acquired in Total Ion Mode (TIM) and quantitation is performed using a single mass ion. Positive Identification of the various target compounds is made by comparison of absolute retention times to those of standards combined with computer based mass spectral mass/intensity library searching and matching routines. | 0.005 | Y | Y | Y | 146 | 4/10/1997 3:18:00 PM | 11/1/2022 4:39:00 AM | 4/21/2023 4:26:45 PM | 1997-01-14 | 6/3/1998 6:51:46 PM | MPAK |  | Concentration | mg/L | Y | (fl. conc.) |  | Trichlorophenol2;4;6 (fl. conc.) | 88-06-2 | Organics; | Organics; EMS code: T042 | Chemical |  | LAB
T043 | mg/L | Trichlorophenol2;4;5 | AEWA | H2SO4 acid,CHCl2 extr.florisil. GC/MS | Effluent and water samples are acidified with Sulphuric Acid and extracted with Dichloromethane. After concentration, the extract is methylated using ethereal Diazomethane. Extracts are cleaned up as required using Florisil. The extracts are analyzed for Acid Extractable Herbicides using GC/MS. | 0.0001 | Y | Y | Y | 71 | 12/3/2002 | 3/11/2003 | 4/1/2004 1:13:28 AM | 1996-04-12 | 6/3/1998 6:51:42 PM | EMS |  | Concentration | mg/L | Y | (fl. conc.) |  | Trichlorophenol2;4;5 (fl. conc.) | 95-95-4 | Organics; | Organics; EMS code: T043 | Chemical |  | LAB
T043 | ug/g | Trichlorophenol2;4;5 | 1332 | Phenolics by Tumbler/GC-MS | Phenolics by Tumbler/GC-MS | 0.02 | Y | Y | Y | 7 | 9/22/2022 5:44:00 PM | 9/22/2022 8:16:00 PM | 3/9/2023 4:55:51 PM | 2023-02-28 |  | EMS |  | MassPerMass | mg/kg | Y | (mass conc.) |  | Trichlorophenol2;4;5 (mass conc.) | 95-95-4 | Organics; | Organics; EMS code: T043 | Chemical |  | LAB
T043 | ug/g | Trichlorophenol2;4;5 | AEAN | H2SO4,CHCL2/MeOH extr.florisil. GC/MS | Sediment samples are shaken with a mixture of Dichloromethane, Methanol and Sulphuric Acid, filtered, then extracted with Dichloromethane. After concentration, the extract is methylated using ethereal Diazomethane. Extracts are cleaned up as required using Florisil. The extracts are analyzed for Acid Extractable Herbicides using GC/MS. | 0.005 | Y | Y | Y | 39 |  |  | 10/24/1996 12:38:38 PM | 1996-04-12 | 6/3/1998 6:51:40 PM | EMS |  | MassPerMass | mg/kg | Y | (mass conc.) |  | Trichlorophenol2;4;5 (mass conc.) | 95-95-4 | Organics; | Organics; EMS code: T043 | Chemical |  | LAB
T043 | mg/L | Trichlorophenol2;4;5 | F028 | Acid,Extr.MeCl,HRGC/LRMS:TIM,SIM | A measured volume of sample is acidified and extracted with Methylene Chloride. For samples which have a high organic content that potentially produce chromatographic interferences an initial pre-extraction under basic conditions is performed. The sample extract is then evaporated to small volume and analysed by High Resolution Gas Chromatography / Low Resolution Mass Spectrometry (HRGC/LRMS). LRMS data is acquired in Total Ion Mode (TIM) and quantitation is performed using a single mass ion. Positive Identification of the various target compounds is made by comparison of absolute retention times to those of standards combined with computer based mass spectral mass/intensity library searching and matching routines. | 0.0001 | Y | Y | Y | 70 | 8/20/2009 | 11/1/2022 4:39:00 AM | 4/21/2023 4:26:45 PM | 2009-12-23 |  | RDGIBSON |  | Concentration | mg/L | Y | (fl. conc.) |  | Trichlorophenol2;4;5 (fl. conc.) | 95-95-4 | Organics; | Organics; EMS code: T043 | Chemical |  | LAB
T044 | ug/g | Trichlorophenol3;4;5 | AEAN | H2SO4,CHCL2/MeOH extr.florisil. GC/MS | Sediment samples are shaken with a mixture of Dichloromethane, Methanol and Sulphuric Acid, filtered, then extracted with Dichloromethane. After concentration, the extract is methylated using ethereal Diazomethane. Extracts are cleaned up as required using Florisil. The extracts are analyzed for Acid Extractable Herbicides using GC/MS. | 0.005 | Y | Y | Y | 39 |  |  | 10/24/1996 12:38:38 PM | 1996-04-12 | 6/3/1998 6:51:40 PM | EMS |  | MassPerMass | mg/kg | Y | (mass conc.) |  | Trichlorophenol3;4;5 (mass conc.) | 609-19-8 | Organics; | Organics; EMS code: T044 | Chemical |  | LAB
T044 | ug/g | Trichlorophenol3;4;5 | 1332 | Phenolics by Tumbler/GC-MS | Phenolics by Tumbler/GC-MS | 0.02 | Y | Y | Y | 22 | 1/17/2022 | 9/22/2022 8:16:00 PM | 1/11/2024 1:07:08 PM | 2022-07-07 |  | EMS |  | MassPerMass | mg/kg | Y | (mass conc.) |  | Trichlorophenol3;4;5 (mass conc.) | 609-19-8 | Organics; | Organics; EMS code: T044 | Chemical |  | LAB
T044 | mg/L | Trichlorophenol3;4;5 | AEWA | H2SO4 acid,CHCl2 extr.florisil. GC/MS | Effluent and water samples are acidified with Sulphuric Acid and extracted with Dichloromethane. After concentration, the extract is methylated using ethereal Diazomethane. Extracts are cleaned up as required using Florisil. The extracts are analyzed for Acid Extractable Herbicides using GC/MS. | 0.0001 | Y | Y | Y | 64 |  |  | 3/17/1997 1:23:13 PM | 1996-04-12 | 6/3/1998 6:51:42 PM | EMS |  | Concentration | mg/L | Y | (fl. conc.) |  | Trichlorophenol3;4;5 (fl. conc.) | 609-19-8 | Organics; | Organics; EMS code: T044 | Chemical |  | LAB
T044 | ug/g | Trichlorophenol3;4;5 | DCMA | DCM/ Acetone/MTBE extr, AA/TMA deriv, GCMS | DCM/ Acetone/MTBE extr; AA/TMA deriv; GCMS. EPA 3570: 8270: Knapp(1979) | 0.02 | Y | Y | Y | 8 | 9/9/2019 | 3/20/2020 | 2/21/2023 2:10:29 PM | 2019-10-16 |  | KATFRASE |  | MassPerMass | mg/kg | Y | (mass conc.) |  | Trichlorophenol3;4;5 (mass conc.) | 609-19-8 | Organics; | Organics; EMS code: T044 | Chemical |  | LAB
T044 | mg/L | Trichlorophenol3;4;5 | X268 | CH2Cl2 ext:GC/MS | CH2Cl2 ext:GC/MS | 0.0005 | Y | Y | Y | 34 | 12/28/2021 | 6/7/2024 | 6/26/2024 4:28:10 PM | 2012-06-25 |  | RDGIBSON |  | Concentration | mg/L | Y | (fl. conc.) |  | Trichlorophenol3;4;5 (fl. conc.) | 609-19-8 | Organics; | Organics; EMS code: T044 | Chemical |  | LAB
T044 | mg/L | Trichlorophenol3;4;5 | F028 | Acid,Extr.MeCl,HRGC/LRMS:TIM,SIM | A measured volume of sample is acidified and extracted with Methylene Chloride. For samples which have a high organic content that potentially produce chromatographic interferences an initial pre-extraction under basic conditions is performed. The sample extract is then evaporated to small volume and analysed by High Resolution Gas Chromatography / Low Resolution Mass Spectrometry (HRGC/LRMS). LRMS data is acquired in Total Ion Mode (TIM) and quantitation is performed using a single mass ion. Positive Identification of the various target compounds is made by comparison of absolute retention times to those of standards combined with computer based mass spectral mass/intensity library searching and matching routines. | 0.0001 | Y | Y | Y | 70 | 8/20/2009 | 11/1/2022 4:39:00 AM | 4/21/2023 4:26:45 PM | 2009-12-23 |  | RDGIBSON |  | Concentration | mg/L | Y | (fl. conc.) |  | Trichlorophenol3;4;5 (fl. conc.) | 609-19-8 | Organics; | Organics; EMS code: T044 | Chemical |  | LAB
T051 | ug/g | Tributyltins | X384 | Purge and trap GC/MS | The sample is analyzed by purge and trap gas chromatography with detection by mass spectrometry. The procedure involves purging the volatile compounds from the sample with an inert gas, and trapping them on a solid sorbent. When purging of the sample is complete, the trap is heated and the compounds are transferred to the gas chromatographic column. Analysis is then accomplished by separation of the components by gas chromatography with detection by mass spectrometry. | 0.001 | Y | Y | Y | 10 |  |  | 10/24/1996 2:25:28 PM | 1996-04-12 | 6/3/1998 6:51:40 PM | EMS |  | MassPerMass | mg/kg | Y | (mass conc.) |  | Tributyltins (mass conc.) | 36643-28-4 | Organometallics; | Organometallics; EMS code: T051 | Chemical |  | LAB
T060 | pg/g wet | Tetrachlorodibenzodioxin | F119 | Solids,Isotope dil,HRGC/HRMS | EPA Method 1613: for Sediments and Fish tissue developed for isomer specific determination of the 2,3,7,8-substituted, tetra-through octa-chlorinated, dibenzo-p-dioxins and dibenzofurans in aqueous, solid and tissue matrices by isotope dilution, high resolution capillary column gas chromatography (HRGC)/high resolution mass spectrometry (HRMS) |  | Y | Y | Y | 5 | 6/16/2000 6:33:00 AM | 10/4/2000 2:53:00 PM | 6/18/2002 3:50:05 PM | 2001-10-01 | 10/4/2001 3:20:45 PM | NPEPPIN | NPEPPIN | MassPerMass | pg/g | Y | (mass conc.) | (wet) | Tetrachlorodibenzodioxin (wet) (mass conc.) |  | Dioxins & Furans; | Dioxins & Furans; EMS code: T060 | Chemical |  | LAB
T060 | pg/L | Tetrachlorodibenzodioxin | F114 | Isotope dil,HRGC/HRMS | EPA Method 1613: Developed for isomer specific determination of the 2,3,7,8-substituted, tetra-through octa-chlorinated, dibenzo-p-dioxins and dibenzofurans in aqueous, solid and tissue matrices by isotope dilution, high resolution capillary column gas chromatography (HRGC)/high resolution mass spectrometry (HRMS) |  | Y | Y | Y | 18 |  |  | 2/21/2024 1:00:50 PM | 2001-05-23 | 5/24/2001 3:35:10 PM | NPEPPIN | NPEPPIN | Concentration | pg/L | Y | (fl. conc.) |  | Tetrachlorodibenzodioxin (fl. conc.) |  | Dioxins & Furans; | Dioxins & Furans; EMS code: T060 | Chemical |  | LAB
T060 | pg/L | Tetrachlorodibenzodioxin | F102 | Env.Canada(1992),DioxinsProc,Pulp Effluents | Environment Canada (1992);Dioxins Procedures for Pulp Effluents |  | Y | Y | Y | 6 | 5/1/2001 7:37:00 AM | 5/1/2001 7:37:00 AM | 11/9/2011 3:41:39 PM | 2000-10-25 |  | NPEPPIN |  | Concentration | pg/L | Y | (fl. conc.) |  | Tetrachlorodibenzodioxin (fl. conc.) |  | Dioxins & Furans; | Dioxins & Furans; EMS code: T060 | Chemical |  | LAB
T060 | pg/L | Tetrachlorodibenzodioxin | X253 | Tol.-Ac/Cl2CH2 ext:GC/MS | C-13 labelled PCDD and PCDF congeners are added to a 1L sample of pulp effluent, which is separated into solid and liquid phases by filtration through a glass fibre filter. The solid phase is aire-dried overnight before undergoing a Soxhlet extraction with toluene. The liquid phase is extracted with dichloromethane in a separatory funnel. The two extracts are combined, concentrated, and then passed through a series of columns which remove the bulk of the organic matrix co-extracted with PCDD/PCDF. One resulting fraction is concentrated to a known small volume for analysis by HRGC/HRMS/DS. | 1 | Y | Y | Y | 146 |  |  | 6/25/2012 2:16:03 PM | 1996-04-12 | 6/3/1998 6:51:40 PM | EMS |  | Concentration | pg/L | Y | (fl. conc.) |  | Tetrachlorodibenzodioxin (fl. conc.) |  | Dioxins & Furans; | Dioxins & Furans; EMS code: T060 | Chemical |  | SURROGATE_RESULT
T061 | pg/L | 4Cl2BenoDiox-2;3;7;8 | F114 | Isotope dil,HRGC/HRMS | EPA Method 1613: Developed for isomer specific determination of the 2,3,7,8-substituted, tetra-through octa-chlorinated, dibenzo-p-dioxins and dibenzofurans in aqueous, solid and tissue matrices by isotope dilution, high resolution capillary column gas chromatography (HRGC)/high resolution mass spectrometry (HRMS) |  | Y | Y | Y | 18 |  |  | 2/21/2024 1:00:50 PM | 2001-05-23 | 5/24/2001 3:35:10 PM | NPEPPIN | NPEPPIN | Concentration | pg/L | Y | (fl. conc.) |  | 4Cl2BenoDiox-2;3;7;8 (fl. conc.) | 1746-01-6 | Dioxins & Furans; | Dioxins & Furans; EMS code: T061 | Chemical |  | LAB
T061 | pg/g wet | 4Cl2BenoDiox-2;3;7;8 | F119 | Solids,Isotope dil,HRGC/HRMS | EPA Method 1613: for Sediments and Fish tissue developed for isomer specific determination of the 2,3,7,8-substituted, tetra-through octa-chlorinated, dibenzo-p-dioxins and dibenzofurans in aqueous, solid and tissue matrices by isotope dilution, high resolution capillary column gas chromatography (HRGC)/high resolution mass spectrometry (HRMS) |  | Y | Y | Y | 5 | 6/16/2000 6:33:00 AM | 10/4/2000 2:53:00 PM | 6/18/2002 3:50:05 PM | 2001-07-03 | 10/4/2001 3:20:45 PM | NPEPPIN | NPEPPIN | MassPerMass | pg/g | Y | (mass conc.) | (wet) | 4Cl2BenoDiox-2;3;7;8 (wet) (mass conc.) | 1746-01-6 | Dioxins & Furans; | Dioxins & Furans; EMS code: T061 | Chemical |  | LAB
T061 | pg/L | 4Cl2BenoDiox-2;3;7;8 | X253 | Tol.-Ac/Cl2CH2 ext:GC/MS | C-13 labelled PCDD and PCDF congeners are added to a 1L sample of pulp effluent, which is separated into solid and liquid phases by filtration through a glass fibre filter. The solid phase is aire-dried overnight before undergoing a Soxhlet extraction with toluene. The liquid phase is extracted with dichloromethane in a separatory funnel. The two extracts are combined, concentrated, and then passed through a series of columns which remove the bulk of the organic matrix co-extracted with PCDD/PCDF. One resulting fraction is concentrated to a known small volume for analysis by HRGC/HRMS/DS. | 1 | Y | Y | Y | 206 | 3/2/1999 | 6/20/2017 | 7/28/2017 2:07:01 PM | 1996-04-12 | 6/3/1998 6:51:40 PM | EMS |  | Concentration | pg/L | Y | (fl. conc.) |  | 4Cl2BenoDiox-2;3;7;8 (fl. conc.) | 1746-01-6 | Dioxins & Furans; | Dioxins & Furans; EMS code: T061 | Chemical |  | SURROGATE_RESULT
T061 | pg/L | 4Cl2BenoDiox-2;3;7;8 | F102 | Env.Canada(1992),DioxinsProc,Pulp Effluents | Environment Canada (1992);Dioxins Procedures for Pulp Effluents |  | Y | Y | Y | 6 | 5/1/2001 7:37:00 AM | 5/1/2001 7:37:00 AM | 11/9/2011 3:41:39 PM | 2000-10-25 |  | NPEPPIN |  | Concentration | pg/L | Y | (fl. conc.) |  | 4Cl2BenoDiox-2;3;7;8 (fl. conc.) | 1746-01-6 | Dioxins & Furans; | Dioxins & Furans; EMS code: T061 | Chemical |  | LAB
T062 | pg/L | Tetrachlorodibenzofuran | F102 | Env.Canada(1992),DioxinsProc,Pulp Effluents | Environment Canada (1992);Dioxins Procedures for Pulp Effluents |  | Y | Y | Y | 6 | 5/1/2001 7:37:00 AM | 5/1/2001 7:37:00 AM | 11/9/2011 3:41:39 PM | 2000-10-25 |  | NPEPPIN |  | Concentration | pg/L | Y | (fl. conc.) |  | Tetrachlorodibenzofuran (fl. conc.) |  | Dioxins & Furans; | Dioxins & Furans; EMS code: T062 | Chemical |  | LAB
T062 | pg/g wet | Tetrachlorodibenzofuran | F119 | Solids,Isotope dil,HRGC/HRMS | EPA Method 1613: for Sediments and Fish tissue developed for isomer specific determination of the 2,3,7,8-substituted, tetra-through octa-chlorinated, dibenzo-p-dioxins and dibenzofurans in aqueous, solid and tissue matrices by isotope dilution, high resolution capillary column gas chromatography (HRGC)/high resolution mass spectrometry (HRMS) |  | Y | Y | Y | 5 | 6/16/2000 6:33:00 AM | 10/4/2000 2:53:00 PM | 6/18/2002 3:50:06 PM | 2001-10-01 | 10/4/2001 3:20:45 PM | NPEPPIN | NPEPPIN | MassPerMass | pg/g | Y | (mass conc.) | (wet) | Tetrachlorodibenzofuran (wet) (mass conc.) |  | Dioxins & Furans; | Dioxins & Furans; EMS code: T062 | Chemical |  | LAB
T062 | pg/L | Tetrachlorodibenzofuran | F114 | Isotope dil,HRGC/HRMS | EPA Method 1613: Developed for isomer specific determination of the 2,3,7,8-substituted, tetra-through octa-chlorinated, dibenzo-p-dioxins and dibenzofurans in aqueous, solid and tissue matrices by isotope dilution, high resolution capillary column gas chromatography (HRGC)/high resolution mass spectrometry (HRMS) |  | Y | Y | Y | 18 |  |  | 2/21/2024 1:00:50 PM | 2001-05-23 | 5/24/2001 3:35:10 PM | NPEPPIN | NPEPPIN | Concentration | pg/L | Y | (fl. conc.) |  | Tetrachlorodibenzofuran (fl. conc.) |  | Dioxins & Furans; | Dioxins & Furans; EMS code: T062 | Chemical |  | LAB
T062 | pg/L | Tetrachlorodibenzofuran | X253 | Tol.-Ac/Cl2CH2 ext:GC/MS | C-13 labelled PCDD and PCDF congeners are added to a 1L sample of pulp effluent, which is separated into solid and liquid phases by filtration through a glass fibre filter. The solid phase is aire-dried overnight before undergoing a Soxhlet extraction with toluene. The liquid phase is extracted with dichloromethane in a separatory funnel. The two extracts are combined, concentrated, and then passed through a series of columns which remove the bulk of the organic matrix co-extracted with PCDD/PCDF. One resulting fraction is concentrated to a known small volume for analysis by HRGC/HRMS/DS. | 1 | Y | Y | Y | 146 |  |  | 6/25/2012 2:16:04 PM | 1996-04-12 | 6/3/1998 6:51:45 PM | EMS |  | Concentration | pg/L | Y | (fl. conc.) |  | Tetrachlorodibenzofuran (fl. conc.) |  | Dioxins & Furans; | Dioxins & Furans; EMS code: T062 | Chemical |  | SURROGATE_RESULT
T063 | pg/L | 4Cl2BenFuran-2;3;7;8 | F114 | Isotope dil,HRGC/HRMS | EPA Method 1613: Developed for isomer specific determination of the 2,3,7,8-substituted, tetra-through octa-chlorinated, dibenzo-p-dioxins and dibenzofurans in aqueous, solid and tissue matrices by isotope dilution, high resolution capillary column gas chromatography (HRGC)/high resolution mass spectrometry (HRMS) |  | Y | Y | Y | 18 |  |  | 2/21/2024 1:00:50 PM | 2001-05-23 | 5/24/2001 3:35:10 PM | NPEPPIN | NPEPPIN | Concentration | pg/L | Y | (fl. conc.) |  | 4Cl2BenFuran-2;3;7;8 (fl. conc.) | 51207-31-9 | Dioxins & Furans; | Dioxins & Furans; EMS code: T063 | Chemical |  | LAB
T063 | pg/g wet | 4Cl2BenFuran-2;3;7;8 | F119 | Solids,Isotope dil,HRGC/HRMS | EPA Method 1613: for Sediments and Fish tissue developed for isomer specific determination of the 2,3,7,8-substituted, tetra-through octa-chlorinated, dibenzo-p-dioxins and dibenzofurans in aqueous, solid and tissue matrices by isotope dilution, high resolution capillary column gas chromatography (HRGC)/high resolution mass spectrometry (HRMS) |  | Y | Y | Y | 5 | 6/16/2000 6:33:00 AM | 10/4/2000 2:53:00 PM | 6/18/2002 3:50:06 PM | 2001-07-03 | 10/4/2001 3:20:45 PM | NPEPPIN | NPEPPIN | MassPerMass | pg/g | Y | (mass conc.) | (wet) | 4Cl2BenFuran-2;3;7;8 (wet) (mass conc.) | 51207-31-9 | Dioxins & Furans; | Dioxins & Furans; EMS code: T063 | Chemical |  | LAB
T063 | pg/L | 4Cl2BenFuran-2;3;7;8 | F102 | Env.Canada(1992),DioxinsProc,Pulp Effluents | Environment Canada (1992);Dioxins Procedures for Pulp Effluents |  | Y | Y | Y | 6 | 5/1/2001 7:37:00 AM | 5/1/2001 7:37:00 AM | 11/9/2011 3:41:39 PM | 2000-10-25 |  | NPEPPIN |  | Concentration | pg/L | Y | (fl. conc.) |  | 4Cl2BenFuran-2;3;7;8 (fl. conc.) | 51207-31-9 | Dioxins & Furans; | Dioxins & Furans; EMS code: T063 | Chemical |  | LAB
T063 | pg/L | 4Cl2BenFuran-2;3;7;8 | X253 | Tol.-Ac/Cl2CH2 ext:GC/MS | C-13 labelled PCDD and PCDF congeners are added to a 1L sample of pulp effluent, which is separated into solid and liquid phases by filtration through a glass fibre filter. The solid phase is aire-dried overnight before undergoing a Soxhlet extraction with toluene. The liquid phase is extracted with dichloromethane in a separatory funnel. The two extracts are combined, concentrated, and then passed through a series of columns which remove the bulk of the organic matrix co-extracted with PCDD/PCDF. One resulting fraction is concentrated to a known small volume for analysis by HRGC/HRMS/DS. | 1 | Y | Y | Y | 207 | 3/2/1999 | 6/24/2015 | 7/30/2015 1:36:18 PM | 1996-04-12 | 6/3/1998 6:51:40 PM | EMS |  | Concentration | pg/L | Y | (fl. conc.) |  | 4Cl2BenFuran-2;3;7;8 (fl. conc.) | 51207-31-9 | Dioxins & Furans; | Dioxins & Furans; EMS code: T063 | Chemical |  | SURROGATE_RESULT
T064 | mg/L | 2;4;5-T | P015 | Phenoxyacid Herbicides by LC-MS/MS | Phenoxyacid Herbicides by LC-MS/MS | 0.5 | Y | Y | Y | 71 | 10/9/2019 | 9/9/2022 | 12/8/2022 12:57:06 PM | 2019-10-17 |  | KATFRASE |  | Concentration | mg/L | Y | (fl. conc.) |  | 2;4;5-T (fl. conc.) | 93-76-5 | Pesticides; | Pesticides; EMS code: T064 | Chemical |  | LAB
T064 | mg/L | 2;4;5-T | AEWA | H2SO4 acid,CHCl2 extr.florisil. GC/MS | Effluent and water samples are acidified with Sulphuric Acid and extracted with Dichloromethane. After concentration, the extract is methylated using ethereal Diazomethane. Extracts are cleaned up as required using Florisil. The extracts are analyzed for Acid Extractable Herbicides using GC/MS. | 0.0001 | Y | Y | Y | 6 | 6/16/2003 | 7/29/2005 7:10:00 AM | 6/2/2009 12:01:12 PM | 1996-04-12 | 6/3/1998 6:51:42 PM | EMS |  | Concentration | mg/L | Y | (fl. conc.) |  | 2;4;5-T (fl. conc.) | 93-76-5 | Pesticides; | Pesticides; EMS code: T064 | Chemical |  | LAB
T064 | mg/L | 2;4;5-T | F052 | Acid,DCM extr'n,methn,Florisil,HRGC-EC | Effluent and water samples are acidified with Sulphuric Acid and extracted with Dichloromethane. After concentration, the extract is methylated using ethereal Diazomethane. Extracts are cleaned up as required using Florisil. The extracts are analysed for Acid Extractable Herbicides using High Resolution Gas Chromatography with Electron Capture detection. Positive results may be verified by use of a secondary confirmation GC column. | 5e-05 | Y | Y | Y | 33 | 12/8/1998 3:39:00 PM | 12/31/2001 12:00:00 PM | 6/18/2003 2:53:58 PM | 1997-10-20 | 6/3/1998 6:51:47 PM | NPEPPIN |  | Concentration | mg/L | Y | (fl. conc.) |  | 2;4;5-T (fl. conc.) | 93-76-5 | Pesticides; | Pesticides; EMS code: T064 | Chemical |  | LAB
T064 | ug/g | 2;4;5-T | FS15 | DCM/MEOH acid extr'n,DCM,methn,Florisil,HRGC-EC | Sediment samples are shaken with a mixture of Dichloromethane, Methanol and Sulphuric Acid, filtered, then extracted with Dichloromethane. After concentration, the extract is methylated using ethereal Diazomethane. Extracts are cleaned up as required using Florisil. The extracts are analysed for Acid Extractable Herbicides using High Resolution Gas Chromatography with Electron Capture detection. Positive results may be verified by use of a secondary confirmation GC column. | 0.05 | Y | Y | Y | 35 | 4/24/1997 4:15:00 PM | 6/24/2003 | 2/26/2004 1:05:31 AM | 1997-10-20 | 6/3/1998 6:51:47 PM | NPEPPIN |  | MassPerMass | mg/kg | Y | (mass conc.) |  | 2;4;5-T (mass conc.) | 93-76-5 | Pesticides; | Pesticides; EMS code: T064 | Chemical |  | LAB
T064 | ug/g | 2;4;5-T | ZN03 | DCM+Ac.Extr:GC/MS | DCM+Ac.Extr:GC/MS EPA 8270E (mod) | 0.0005 | Y | Y | Y | 15 | 1/14/2022 | 2/2/2022 | 1/11/2024 1:07:08 PM | 2022-07-27 |  | EMS |  | MassPerMass | mg/kg | Y | (mass conc.) |  | 2;4;5-T (mass conc.) | 93-76-5 | Pesticides; | Pesticides; EMS code: T064 | Chemical |  | LAB
T066 | mg/L | 1;2;4-Trichlorobenzene | MS04 | Static Headspace GC/MS (Water) | (Water) A static headspace procedure is used as a means of driving the analytes out of the matrix into a headspace. Once equilibrium is achieved, a fixed volume of headspeace is taken for subsequent determination by Gas Chromatography - Mass Spectroscopy (GC/MS) using Selected Ion Monitoring. | 0.002 | Y | Y | Y | 8 | 10/5/2019 2:07:00 AM | 6/28/2024 12:29:00 AM | 7/9/2024 3:31:02 PM | 2021-04-07 |  | RDGIBSON |  | Concentration | mg/L | Y | (fl. conc.) |  | 1;2;4-Trichlorobenzene (fl. conc.) | 120-82-1 | Organics; | Organics; EMS code: T066 | Chemical |  | LAB
T070 | mg/L | Trichlorofluoromethane | MS01 | Purge&Trap GC/MS (VOC-W) | Purge&Trap GC/MS (VOC-W) | 0.3 | Y | Y | Y | 544 | 4/15/2002 | 6/4/2024 10:24:00 AM | 6/13/2024 5:01:27 PM | 1996-10-13 | 3/18/2003 2:49:40 PM | EMS | UMULLER | Concentration | mg/L | Y | (fl. conc.) |  | Trichlorofluoromethane (fl. conc.) | 75-69-4 | Chlorofluorocarbon (CFCs); | Chlorofluorocarbon (CFCs); EMS code: T070 | Chemical |  | LAB
T070 | mg/L | Trichlorofluoromethane | MS04 | Static Headspace GC/MS (Water) | (Water) A static headspace procedure is used as a means of driving the analytes out of the matrix into a headspace. Once equilibrium is achieved, a fixed volume of headspeace is taken for subsequent determination by Gas Chromatography - Mass Spectroscopy (GC/MS) using Selected Ion Monitoring. | 0.004 | Y | Y | Y | 70 | 8/13/2009 | 6/28/2024 12:29:00 AM | 7/9/2024 3:31:02 PM | 2009-12-23 |  | RDGIBSON |  | Concentration | mg/L | Y | (fl. conc.) |  | Trichlorofluoromethane (fl. conc.) | 75-69-4 | Chlorofluorocarbon (CFCs); | Chlorofluorocarbon (CFCs); EMS code: T070 | Chemical |  | LAB
T070 | ug/g | Trichlorofluoromethane | MS02 | Purge&Trap GC/MS (VOC-S) | Purge&Trap GC/MS (VOC-S) | 0.002 | Y | Y | Y | 27 | 11/19/2004 | 10/28/2022 5:04:00 PM | 1/11/2024 1:07:08 PM | 1996-10-13 | 6/3/1998 6:51:46 PM | EMS |  | MassPerMass | mg/kg | Y | (mass conc.) |  | Trichlorofluoromethane (mass conc.) | 75-69-4 | Chlorofluorocarbon (CFCs); | Chlorofluorocarbon (CFCs); EMS code: T070 | Chemical |  | LAB
T070 | ug/g | Trichlorofluoromethane | FS04 | Static Headspace:HRGC/LRMS | Static Headspace anaylsis is a method for the determination of volatile compounds in liquids and/or solids by analysing the vapour phase of a sample in a closed system; in this case a sealed vial. Static Headspace is sometimes referred to as Equilibrium Headspace because the vapour phase above the sample in a sealed vial is brought into thermodynamic equilibrium with the sample prior to analysis. Volatile organic compounds are thermally extracted from the sample matrix by Static Headspace sampling and introduced directly to a thick-film megabore GC column. Compounds eluting from the GC column are introduced into a Mass Spectrometer and target compounds are identified by a comparison of both retention times and mass spectra with those of reference standards stored in a computerized data system. Reference retention times and mass spectra for analytes are obtained by the measurement of calibration standards under precisely the same extraction conditions used for samples. The concentration of each identified component is measured by relating the response of the sample quantitation mass ions to the response of the quantitation mass ions of a series of external standards. A deuterated surrogate recovery standard is added to provide information on matrix effects. | 0.05 | Y | Y | Y | 44 | 6/30/1999 4:41:00 PM | 5/22/2015 | 7/3/2015 3:19:04 PM | 1996-09-16 | 6/3/1998 6:51:44 PM | MPAK |  | MassPerMass | mg/kg | Y | (mass conc.) |  | Trichlorofluoromethane (mass conc.) | 75-69-4 | Chlorofluorocarbon (CFCs); | Chlorofluorocarbon (CFCs); EMS code: T070 | Chemical |  | LAB
T070 | mg/L | Trichlorofluoromethane | X384 | Purge and trap GC/MS | The sample is analyzed by purge and trap gas chromatography with detection by mass spectrometry. The procedure involves purging the volatile compounds from the sample with an inert gas, and trapping them on a solid sorbent. When purging of the sample is complete, the trap is heated and the compounds are transferred to the gas chromatographic column. Analysis is then accomplished by separation of the components by gas chromatography with detection by mass spectrometry. | 0.001 | Y | Y | Y | 112 | 8/5/2004 10:31:00 AM | 7/26/2012 7:12:00 AM | 9/26/2012 1:07:09 PM | 1996-04-12 | 6/3/1998 6:51:40 PM | EMS |  | Concentration | mg/L | Y | (fl. conc.) |  | Trichlorofluoromethane (fl. conc.) | 75-69-4 | Chlorofluorocarbon (CFCs); | Chlorofluorocarbon (CFCs); EMS code: T070 | Chemical |  | LAB
T070 | mg/L | Trichlorofluoromethane | F007 | Static Headspace:HRGC/LRMS | Static headspace analysis is a method for the determination of volatile compounds in liquids and/or solids by analysing the vapour phase of a sample in a closed system; in this case a sealed vial. Static Headspace is sometimes referred to as Equilibrium Headspace because the vapour phase above the sample in a sealed vial is brought into thermodynamic equilibrium with the sample prior to analysis. Volatile organic compounds are thermally extracted from the sample matrix by Static Headspace sampling and introduced directly to a thick-film megabore GC column. Compounds eluting from the GC column are introduced into a Mass Spectrometer and target compounds are identified by a comparison of both retention times and mass spectra with those of reference standards stored in a computerized data system. Reference retention times and mass spectra for analytes are obtained by the measurement of calibration standards under precisely the same extraction conditions used for samples. The concentration of each identified component is measured by relating the response of the sample quantitation mass ions to the response of the quantitation mass ions of a series of external standards. A deuterated surrogate recovery standard is added to provide information on matrix effects. | 0.005 | Y | Y | Y | 380 | 9/28/1998 3:08:00 PM | 8/14/2024 | 8/16/2024 4:05:21 PM | 1996-09-16 | 6/3/1998 6:51:46 PM | MPAK |  | Concentration | mg/L | Y | (fl. conc.) |  | Trichlorofluoromethane (fl. conc.) | 75-69-4 | Chlorofluorocarbon (CFCs); | Chlorofluorocarbon (CFCs); EMS code: T070 | Chemical |  | LAB
T071 | ug/g | 2;4;5-TP | ZN03 | DCM+Ac.Extr:GC/MS | DCM+Ac.Extr:GC/MS EPA 8270E (mod) | 0.0005 | Y | Y | Y | 15 | 1/14/2022 | 2/2/2022 | 1/11/2024 1:07:08 PM | 2022-07-27 |  | EMS |  | MassPerMass | mg/kg | Y | (mass conc.) |  | 2;4;5-TP (mass conc.) | 93-72-1 | Organics; | Organics; EMS code: T071 | Chemical |  | LAB
T071 | mg/L | 2;4;5-TP | P015 | Phenoxyacid Herbicides by LC-MS/MS | Phenoxyacid Herbicides by LC-MS/MS | 0.5 | Y | Y | Y | 71 | 10/9/2019 | 9/9/2022 | 12/8/2022 12:57:06 PM | 2019-10-17 |  | KATFRASE |  | Concentration | mg/L | Y | (fl. conc.) |  | 2;4;5-TP (fl. conc.) | 93-72-1 | Organics; | Organics; EMS code: T071 | Chemical |  | LAB
T072 | ug/g | 2;4-DP | ZN03 | DCM+Ac.Extr:GC/MS | DCM+Ac.Extr:GC/MS EPA 8270E (mod) | 0.0005 | Y | Y | Y | 15 | 1/14/2022 | 2/2/2022 | 1/11/2024 1:07:08 PM | 2022-07-27 |  | EMS |  | MassPerMass | mg/kg | Y | (mass conc.) |  | 2;4-DP (mass conc.) | 120-36-5 | Pesticides; | Pesticides; EMS code: T072 | Chemical |  | LAB
T072 | mg/L | 2;4-DP | P015 | Phenoxyacid Herbicides by LC-MS/MS | Phenoxyacid Herbicides by LC-MS/MS | 0.5 | Y | Y | Y | 71 | 10/9/2019 | 9/9/2022 | 12/8/2022 12:57:06 PM | 2019-10-17 |  | KATFRASE |  | Concentration | mg/L | Y | (fl. conc.) |  | 2;4-DP (fl. conc.) | 120-36-5 | Pesticides; | Pesticides; EMS code: T072 | Chemical |  | LAB
T073 | % | 2;4-Dichlorophenylacetic Acid | SURR | Surrogate Recovery | Surrogate Recovery | 1 | Y | Y | Y | 86 | 10/9/2019 | 9/9/2022 | 1/11/2024 1:07:08 PM | 2019-12-16 |  | RDGIBSON |  | DimensionlessRatio | % | N | (ratio) |  | 2;4-Dichlorophenylacetic Acid (ratio) | 19719-28-9 | Organics; | Organics; EMS code: T073 | Chemical |  | SURROGATE_RESULT
T074 | mg/L | Arsenate (As V) | P011 | Arsenic Species in Water by HPLC/CCMS | Arsenic Species in Water by HPLC/CCMS | 0.02 | Y | Y | Y | 28 | 9/13/2019 | 5/14/2024 | 6/6/2024 10:23:03 AM | 2019-10-17 | 10/17/2019 3:31:48 PM | KATFRASE | KATFRASE | Concentration | mg/L | Y | (fl. conc.) |  | Arsenate (As V) (fl. conc.) | 15584-04-0 | Inorganics; | Inorganics; EMS code: T074 | Chemical |  | LAB
T075 | mg/L | Arsenite (As III) | P011 | Arsenic Species in Water by HPLC/CCMS | Arsenic Species in Water by HPLC/CCMS | 0.02 | Y | Y | Y | 28 | 9/13/2019 | 5/14/2024 | 6/6/2024 10:23:03 AM | 2019-10-17 | 10/17/2019 3:31:55 PM | KATFRASE | KATFRASE | Concentration | mg/L | Y | (fl. conc.) |  | Arsenite (As III) (fl. conc.) | 15502-74-6 | Inorganics; | Inorganics; EMS code: T075 | Chemical |  | LAB
T076 | mg/L | Arsenobetaine (AsB, as As) | P011 | Arsenic Species in Water by HPLC/CCMS | Arsenic Species in Water by HPLC/CCMS | 0.05 | Y | Y | Y | 28 | 9/13/2019 | 5/14/2024 | 6/6/2024 10:23:03 AM | 2019-10-17 | 10/17/2019 3:32:04 PM | KATFRASE | KATFRASE | Concentration | mg/L | Y | (fl. conc.) |  | Arsenobetaine (AsB, as As) (fl. conc.) | 64436-13-1 | Organometallics; | Organometallics; EMS code: T076 | Chemical |  | LAB
T077 | mg/L | Dimethylarsinic Acid (DMA, as As) | P011 | Arsenic Species in Water by HPLC/CCMS | Arsenic Species in Water by HPLC/CCMS | 0.02 | Y | Y | Y | 28 | 9/13/2019 | 5/14/2024 | 6/6/2024 10:23:03 AM | 2019-10-17 |  | KATFRASE |  | Concentration | mg/L | Y | (fl. conc.) |  | Dimethylarsinic Acid (DMA, as As) (fl. conc.) | 75-60-5 | Organometallics; | Organometallics; EMS code: T077 | Chemical |  | LAB
T078 | mg/L | Monomethylarsonic Acid (MMA, as As) | P011 | Arsenic Species in Water by HPLC/CCMS | Arsenic Species in Water by HPLC/CCMS | 0.02 | Y | Y | Y | 28 | 9/13/2019 | 5/14/2024 | 6/6/2024 10:23:03 AM | 2019-10-17 |  | KATFRASE |  | Concentration | mg/L | Y | (fl. conc.) |  | Monomethylarsonic Acid (MMA, as As) (fl. conc.) | 124-58-3 | Organometallics; | Organometallics; EMS code: T078 | Chemical |  | LAB
T080 | mg/L | Terbuthylazine | ZN01 | DCM+Ac.Extr:GC/MS | DCM+Ac.Extr:GC/MS | 0.0001 | Y | Y | Y | 28 | 7/28/2005 7:03:00 AM | 3/29/2012 10:15:00 PM | 5/3/2012 9:27:52 AM | 1996-04-12 | 11/16/1999 4:06:28 PM | EMS | NPEPPIN | Concentration | mg/L | Y | (fl. conc.) |  | Terbuthylazine (fl. conc.) | 5915-41-3 | Pesticides; | Pesticides; EMS code: T080 | Chemical |  | LAB
T081 | ug/g | Terbutryn | ZN03 | DCM+Ac.Extr:GC/MS | DCM+Ac.Extr:GC/MS EPA 8270E (mod) | 0.02 | Y | Y | Y | 14 | 1/25/2022 | 9/8/2022 | 12/15/2022 10:50:08 AM | 1996-04-12 | 11/16/1999 2:27:40 PM | EMS | NPEPPIN | MassPerMass | mg/kg | Y | (mass conc.) |  | Terbutryn (mass conc.) | 886-50-0 | Pesticides; | Pesticides; EMS code: T081 | Chemical |  | LAB
T081 | mg/L | Terbutryn | E660 | EPA 8270E (mod) | EPA 8270E (mod) Miscellaneous Pesticides by GC-MS | 0.0001 | Y | Y | Y | 68 | 12/7/2021 | 9/12/2022 | 12/8/2022 12:57:06 PM | 2021-12-09 | 12/9/2021 3:22:12 PM | EMS | EMS | Concentration | mg/L | Y | (fl. conc.) |  | Terbutryn (fl. conc.) | 886-50-0 | Pesticides; | Pesticides; EMS code: T081 | Chemical |  | LAB
T081 | mg/L | Terbutryn | ZN01 | DCM+Ac.Extr:GC/MS | DCM+Ac.Extr:GC/MS | 0.0002 | Y | Y | Y | 28 | 7/28/2005 7:03:00 AM | 3/29/2012 10:15:00 PM | 5/3/2012 9:27:52 AM | 1996-04-12 | 11/16/1999 4:06:28 PM | EMS | NPEPPIN | Concentration | mg/L | Y | (fl. conc.) |  | Terbutryn (fl. conc.) | 886-50-0 | Pesticides; | Pesticides; EMS code: T081 | Chemical |  | LAB
T082 | mg/L | Triallate | ZN01 | DCM+Ac.Extr:GC/MS | DCM+Ac.Extr:GC/MS | 0.0003 | Y | Y | Y | 28 | 7/28/2005 7:03:00 AM | 3/29/2012 10:15:00 PM | 5/3/2012 9:27:52 AM | 1996-04-12 | 11/16/1999 4:06:28 PM | EMS | NPEPPIN | Concentration | mg/L | Y | (fl. conc.) |  | Triallate (fl. conc.) | 2303-17-5 | Pesticides; | Pesticides; EMS code: T082 | Chemical |  | LAB
T082 | mg/L | Triallate | E660 | EPA 8270E (mod) | EPA 8270E (mod) Miscellaneous Pesticides by GC-MS | 0.0001 | Y | Y | Y | 68 | 12/7/2021 | 9/12/2022 | 12/8/2022 12:57:06 PM | 2021-12-09 | 12/9/2021 3:22:12 PM | EMS | EMS | Concentration | mg/L | Y | (fl. conc.) |  | Triallate (fl. conc.) | 2303-17-5 | Pesticides; | Pesticides; EMS code: T082 | Chemical |  | LAB
T082 | ug/g | Triallate | ZN03 | DCM+Ac.Extr:GC/MS | DCM+Ac.Extr:GC/MS EPA 8270E (mod) | 0.03 | Y | Y | Y | 16 | 1/25/2022 | 11/2/2022 4:41:00 AM | 1/7/2023 3:12:11 PM | 1996-04-12 | 11/16/1999 2:27:40 PM | EMS | NPEPPIN | MassPerMass | mg/kg | Y | (mass conc.) |  | Triallate (mass conc.) | 2303-17-5 | Pesticides; | Pesticides; EMS code: T082 | Chemical |  | LAB
T083 | mg/L | Triadimefon | ZN01 | DCM+Ac.Extr:GC/MS | DCM+Ac.Extr:GC/MS | 0.0003 | Y | Y | Y | 28 | 7/28/2005 7:03:00 AM | 3/29/2012 10:15:00 PM | 5/3/2012 9:27:52 AM | 1996-04-12 | 11/16/1999 4:06:28 PM | EMS | NPEPPIN | Concentration | mg/L | Y | (fl. conc.) |  | Triadimefon (fl. conc.) | 43121-43-3 | Pesticides; | Pesticides; EMS code: T083 | Chemical |  | LAB
T084 | mg/L | Triphenyl Phosphate | F053 | DCM extr'n,conc'n,HGRC-PFPD/N,P det'n | Effluent samples are extracted with Dichloromethane. The extract is then concentrated and analysed using High Resolution Gas Chromatography with Pulsed Flame Photometric Dectection. Confirmation, if required, is by High Resolution Gas Chromatography with Nitrogen Phosphorus Detection. | 5e-05 | Y | Y | Y | 2 | 5/4/2001 2:03:00 PM | 1/22/2002 10:26:00 AM | 6/18/2003 2:53:58 PM | 1997-10-20 | 6/3/1998 6:51:47 PM | NPEPPIN |  | Concentration | mg/L | Y | (fl. conc.) |  | Triphenyl Phosphate (fl. conc.) | 115-86-6 | Organics; | Organics; EMS code: T084 | Chemical |  | LAB
T084 | ug/g | Triphenyl Phosphate | FS16 | DCM/MEOH extr'n,DCM,conc'n,HRGC-PFPD/N,P det'n | Sediment samples are shaken with a mixture of Dichloromethane and Methanol, filtered, then extracted with Dichloromethane. The extract is then concentrated and analysed using High Resolution Gas Chromatography with Pulsed Flame Photometric Detection. Confirmation, if required, is by High Resolution Gas Chromatography with Nitrogen Phosphorus Detection. | 0.05 | Y | Y | Y | 11 | 7/9/2002 3:55:00 PM | 7/9/2002 4:05:00 PM | 2/13/2004 1:01:04 AM | 1997-10-20 | 6/3/1998 6:51:47 PM | NPEPPIN |  | MassPerMass | mg/kg | Y | (mass conc.) |  | Triphenyl Phosphate (mass conc.) | 115-86-6 | Organics; | Organics; EMS code: T084 | Chemical |  | LAB
T085 | mg/L | Trichloroethene | MS01 | Purge&Trap GC/MS (VOC-W) | Purge&Trap GC/MS (VOC-W) | 0.001 | Y | Y | Y | 233 | 1/15/2007 4:50:00 PM | 5/27/2016 6:28:00 PM | 12/30/2023 10:26:22 AM | 2003-06-17 |  | UMULLER |  | Concentration | mg/L | Y | (fl. conc.) |  | Trichloroethene (fl. conc.) | 79-01-6 | Volatile Organic Compounds (VOCs); | Volatile Organic Compounds (VOCs); EMS code: T085 | Chemical |  | LAB
T086 | mg/L | Tetramethrin | ZN01 | DCM+Ac.Extr:GC/MS | DCM+Ac.Extr:GC/MS | 0.0005 | Y | Y | Y | 1 | 7/29/2005 7:10:00 AM | 7/29/2005 7:10:00 AM | 6/2/2009 12:01:13 PM | 2003-06-17 |  | UMULLER |  | Concentration | mg/L | Y | (fl. conc.) |  | Tetramethrin (fl. conc.) | 7696-12-0 | Pesticides; | Pesticides; EMS code: T086 | Chemical |  | LAB
T087 | mg/L | Triazophos | ZN01 | DCM+Ac.Extr:GC/MS | DCM+Ac.Extr:GC/MS | 0.0005 | Y | Y | Y | 1 | 7/29/2005 7:10:00 AM | 7/29/2005 7:10:00 AM | 6/2/2009 12:01:13 PM | 2003-06-17 |  | UMULLER |  | Concentration | mg/L | Y | (fl. conc.) |  | Triazophos (fl. conc.) | 24017-47-8 | Pesticides; | Pesticides; EMS code: T087 | Chemical |  | LAB
T089 | mg/L | Terbacil | P015 | Phenoxyacid Herbicides by LC-MS/MS | Phenoxyacid Herbicides by LC-MS/MS | 1e-05 | Y | Y | Y | 68 | 12/6/2021 | 9/9/2022 | 12/8/2022 12:57:06 PM | 2021-12-14 |  | EMS |  | Concentration | mg/L | Y | (fl. conc.) |  | Terbacil (fl. conc.) | 5902-51-2 | Pesticides; | Pesticides; EMS code: T089 | Chemical |  | LAB
T089 | mg/L | Terbacil | ZN01 | DCM+Ac.Extr:GC/MS | DCM+Ac.Extr:GC/MS | 0.0005 | Y | Y | Y | 1 | 7/29/2005 7:10:00 AM | 7/29/2005 7:10:00 AM | 6/2/2009 12:01:13 PM | 2003-06-16 |  | UMULLER |  | Concentration | mg/L | Y | (fl. conc.) |  | Terbacil (fl. conc.) | 5902-51-2 | Pesticides; | Pesticides; EMS code: T089 | Chemical |  | LAB
T090 | mg/L | 2;3;4;6/2;3;5;6-Tetrachlorophenol | F029 | Acid,DiethylEther extr,Diazometh.Methyl'n,HRGC/Electron Capt | Effluent samples to be analysed for Chlorinated Phenols are acidified with dilute Sulfuric Acid and extracted with Diethyl Ether. After concentration the extract is methylated using ethereal Diazomethane and cleaned up with concentrated Sulfuric Acid and Mercury. The extracts are analysed for Chlorinated Anisoles and/or Chlorinated Phenols using High Resolution Gas Chromatography with Electron Capture detection. | 1e-06 | Y | Y | Y | 30 | 7/16/1997 4:10:00 PM | 12/22/1999 3:42:00 PM | 12/25/1999 1:12:13 AM | 1997-01-14 | 6/3/1998 6:51:46 PM | MPAK |  | Concentration | mg/L | Y | (fl. conc.) |  | 2;3;4;6/2;3;5;6-Tetrachlorophenol (fl. conc.) | 25167-83-3, 58-90-2, 935-95-5 | Organics; | Organics; EMS code: T090 | Chemical |  | LAB
T090 | ug/g | 2;3;4;6/2;3;5;6-Tetrachlorophenol | AEAN | H2SO4,CHCL2/MeOH extr.florisil. GC/MS | Sediment samples are shaken with a mixture of Dichloromethane, Methanol and Sulphuric Acid, filtered, then extracted with Dichloromethane. After concentration, the extract is methylated using ethereal Diazomethane. Extracts are cleaned up as required using Florisil. The extracts are analyzed for Acid Extractable Herbicides using GC/MS. | 0.005 | Y | Y | Y | 27 | 4/4/2002 | 6/24/2003 | 2/26/2004 1:05:31 AM | 1999-06-28 |  | NPEPPIN |  | MassPerMass | mg/kg | Y | (mass conc.) |  | 2;3;4;6/2;3;5;6-Tetrachlorophenol (mass conc.) | 25167-83-3, 58-90-2, 935-95-5 | Organics; | Organics; EMS code: T090 | Chemical |  | LAB
T090 | mg/L | 2;3;4;6/2;3;5;6-Tetrachlorophenol | AEWA | H2SO4 acid,CHCl2 extr.florisil. GC/MS | Effluent and water samples are acidified with Sulphuric Acid and extracted with Dichloromethane. After concentration, the extract is methylated using ethereal Diazomethane. Extracts are cleaned up as required using Florisil. The extracts are analyzed for Acid Extractable Herbicides using GC/MS. | 0.0001 | Y | Y | Y | 5 | 6/16/2003 | 6/16/2003 | 2/10/2004 1:02:13 AM | 1999-06-28 |  | NPEPPIN |  | Concentration | mg/L | Y | (fl. conc.) |  | 2;3;4;6/2;3;5;6-Tetrachlorophenol (fl. conc.) | 25167-83-3, 58-90-2, 935-95-5 | Organics; | Organics; EMS code: T090 | Chemical |  | LAB
T091 | ug/g | 2;3;4;6/2;3;5;6-Tetrachloroanisole | FS11 | Acetone extr,Hexane,Diazometh.Methyl'n,HRGC/Electron Capture | Sediment and homogenized biota samples to be analysed for Chlorinated Phenols are Acetone extracted using a wrist-action shaker, partitioned with 2% NaCl and Hexane, methylated, then cleaned up and analysed in the manner outlined for effluent samples. Sediment and homogenized biota samples to be analysed for Chlorinated Anisoles are Acetone extracted using a wrist-action shaker, partitioned with 2% NaCl and Hexane, then cleaned up and analysed in the manner outlined for effluent samples. The extracts are analysed for Chlorinated Anisoles and/or Chlorinated Phenols using High Resolution Gas Chromatography with Electron Capture detection. | 0.0001 | Y | Y | Y | 4 | 8/29/2000 4:27:00 PM | 8/29/2000 4:29:00 PM | 6/18/2002 3:50:41 PM | 1997-01-14 | 6/3/1998 6:51:46 PM | MPAK |  | MassPerMass | mg/kg | Y | (mass conc.) |  | 2;3;4;6/2;3;5;6-Tetrachloroanisole (mass conc.) | 938-22-7, 6936-40-9 | Organics; | Organics; EMS code: T091 | Chemical |  | LAB
T091 | mg/L | 2;3;4;6/2;3;5;6-Tetrachloroanisole | F030 | Acid,Hexane extr,Diazomethane Methylation,HRGC/Electron Capt | Effluent samples to be analysed for Chlorinated Anisoles are extracted with Hexane. After concentration the extract is cleaned up using Florisil, concentrated Sulfuric Acid and Mercury. The extracts are analysed for Chlorinated Anisoles and/or Chlorinated Phenols using High Resolution Gas Chromatography with Electron Capture detection. | 1e-06 | Y | Y | Y | 3 | 8/25/2000 3:03:00 PM | 9/1/2000 9:20:00 AM | 6/18/2002 1:43:32 PM | 1997-01-14 | 6/3/1998 6:51:46 PM | MPAK |  | Concentration | mg/L | Y | (fl. conc.) |  | 2;3;4;6/2;3;5;6-Tetrachloroanisole (fl. conc.) | 938-22-7, 6936-40-9 | Organics; | Organics; EMS code: T091 | Chemical |  | LAB
T092 | mg/L | Triadimenol | ZN01 | DCM+Ac.Extr:GC/MS | DCM+Ac.Extr:GC/MS | 0.0005 | Y | Y | Y | 1 | 7/29/2005 7:10:00 AM | 7/29/2005 7:10:00 AM | 6/2/2009 12:01:13 PM | 2003-06-17 |  | UMULLER |  | Concentration | mg/L | Y | (fl. conc.) |  | Triadimenol (fl. conc.) | 55219-65-3 | Pesticides; | Pesticides; EMS code: T092 | Chemical |  | LAB
T093 | ug/g | 2;3;4;5/2;3;5;6-Tetrachlorophenol | 1332 | Phenolics by Tumbler/GC-MS | Phenolics by Tumbler/GC-MS | 0.02 | Y | Y | Y | 7 | 9/22/2022 5:44:00 PM | 9/22/2022 8:16:00 PM | 3/9/2023 4:55:51 PM | 2023-02-28 |  | EMS |  | MassPerMass | mg/kg | Y | (mass conc.) |  | 2;3;4;5/2;3;5;6-Tetrachlorophenol (mass conc.) | 25167-83-3, 4901-51-3, 935-95-5 | Organics; | Organics; EMS code: T093 | Chemical |  | LAB
T093 | mg/L | 2;3;4;5/2;3;5;6-Tetrachlorophenol | F028 | Acid,Extr.MeCl,HRGC/LRMS:TIM,SIM | A measured volume of sample is acidified and extracted with Methylene Chloride. For samples which have a high organic content that potentially produce chromatographic interferences an initial pre-extraction under basic conditions is performed. The sample extract is then evaporated to small volume and analysed by High Resolution Gas Chromatography / Low Resolution Mass Spectrometry (HRGC/LRMS). LRMS data is acquired in Total Ion Mode (TIM) and quantitation is performed using a single mass ion. Positive Identification of the various target compounds is made by comparison of absolute retention times to those of standards combined with computer based mass spectral mass/intensity library searching and matching routines. | 0.0005 | Y | Y | Y | 55 | 1/14/2015 4:58:00 AM | 11/1/2022 4:39:00 AM | 4/21/2023 4:26:46 PM | 2015-03-25 |  | RDGIBSON |  | Concentration | mg/L | Y | (fl. conc.) |  | 2;3;4;5/2;3;5;6-Tetrachlorophenol (fl. conc.) | 25167-83-3, 4901-51-3, 935-95-5 | Organics; | Organics; EMS code: T093 | Chemical |  | LAB
T100 | mg/L | Tebuconazole | ZN01 | DCM+Ac.Extr:GC/MS | DCM+Ac.Extr:GC/MS | 0.5 | Y | Y | Y | 1 | 7/29/2005 7:10:00 AM | 7/29/2005 7:10:00 AM | 6/2/2009 12:01:13 PM | 2009-02-20 |  | JCHUDAK |  | Concentration | mg/L | Y | (fl. conc.) |  | Tebuconazole (fl. conc.) | 107534-96-3 | Pesticides; | Pesticides; EMS code: T100 | Chemical |  | LAB
T101 | % | d8-Toluene | SURR | Surrogate Recovery | Surrogate Recovery | 0 | Y | Y | Y | 90 | 2/7/2002 10:55:00 AM | 11/29/2004 1:55:00 PM | 1/6/2005 1:01:59 AM | 1996-04-12 | 6/3/1998 6:51:40 PM | EMS |  | DimensionlessRatio | % | N | (ratio) |  | d8-Toluene (ratio) | 2037-26-5 | Volatile Organic Compounds (VOCs); | Volatile Organic Compounds (VOCs); EMS code: T101 | Chemical |  | SURROGATE_RESULT
T101 | mg/L | d8-Toluene | MS01 | Purge&Trap GC/MS (VOC-W) | Purge&Trap GC/MS (VOC-W) |  | Y | Y | Y | 487 | 6/3/2014 9:09:00 PM | 12/11/2022 12:53:00 AM | 12/15/2022 6:20:11 PM | 2003-06-17 |  | UMULLER |  | Concentration | mg/L | Y | (fl. conc.) |  | d8-Toluene (fl. conc.) | 2037-26-5 | Volatile Organic Compounds (VOCs); | Volatile Organic Compounds (VOCs); EMS code: T101 | Chemical |  | SURROGATE_RESULT
T102 | ug/g | Tributyl Phosphate | ZN03 | DCM+Ac.Extr:GC/MS | DCM+Ac.Extr:GC/MS EPA 8270E (mod) | 0.005 | Y | Y | Y | 2 | 11/2/2022 12:53:00 AM | 11/2/2022 4:41:00 AM | 1/7/2023 3:12:11 PM | 2022-12-21 |  | EMS |  | MassPerMass | mg/kg | Y | (mass conc.) |  | Tributyl Phosphate (mass conc.) | 126-73-8 | Organics; | Organics; EMS code: T102 | Chemical |  | LAB
TA-D | mg/L | Tantalum Dissolved | METC | MET-D-NP-CCMS-VA | Ultra trace metals in water are analyzed by CRC ICPMS, based on US EPA Method 6020B (July 2014). This procedure is intended for pristine field-filtered acid-preserved water samples. The detection limits (LOR) for this test are based on lab instrumental analysis only, not including filtration. ALS-supplied field filtration equipment does not support these LOR. Therefore, because of the high probability of false positives due to filtration, it is strongly recommended that a filtration blank be analysed to aid in data interpretation. | 0.0001 | Y | Y | Y | 394 | 3/30/2020 | 4/9/2024 | 4/29/2024 4:53:06 PM | 2022-10-07 |  | EMS |  | Concentration | mg/L | Y | (fl. conc.) |  | Tantalum Dissolved (fl. conc.) | 7440-25-7 | Metals - Dissolved; | Metals - Dissolved; EMS code: TA-D | Chemical | Dissolved | LAB
TA-D | mg/L | Tantalum Dissolved | F032 | Filt.,Acid,ICPMS | There is no acid digestion required for this analysis. A water sample is collected and preserved in 0.2 percent nitric acid. To analyse the extractable metals the sample is shaken, allowed to sit for at least twenty four hours, and a decanted portion is introduced to the spectrometer. When the dissolved metals are required, the samples must first be filtered through a 0.45 micron cellulose acetate filter paper and then followed by preservation in 0.2 percent nitric acid. It is important to note that different laboratories may have very different detection limits for some variables analyzed by ICPMS. | 0.001 | Y | Y | Y | 48 | 7/30/2020 | 12/19/2023 | 1/2/2024 4:32:04 PM | 1998-06-22 |  | NPEPPIN |  | Concentration | mg/L | Y | (fl. conc.) |  | Tantalum Dissolved (fl. conc.) | 7440-25-7 | Metals - Dissolved; | Metals - Dissolved; EMS code: TA-D | Chemical | Dissolved | LAB
TA-T | ppm (S) | Tantalum | 6041 | Aqua Regia by ICPMS | These were sediment/soils digested with Aqua-Regia (Nitric Acid & Hydrochloric Acids) and then analysis by ICPMS. Results would be reported in mg/kg and/or %. | 0.005 | Y | Y | Y | 21 | 10/13/2019 | 8/5/2020 | 8/10/2020 3:52:07 PM | 2019-09-06 | 10/21/2019 11:56:33 AM | KATFRASE | KATFRASE | DimensionlessRatio | ppm | N | (ratio) |  | Tantalum (ratio) | 7440-25-7 | Metals; | Metals; EMS code: TA-T | Chemical |  | LAB
TA-T | ppm (S) | Tantalum | MS41 | ALS Minerals ME-MS41 | ALS Minerals ME-MS41. A prepared sample is digested with aqua regia in a graphite heating block. After cooling; the resulting solution is diluted to with deionized water; mixed and analyzed by ICP-AES. Samples are then analyzed by ICP-MS for the remaining suite of elements. | 0.01 | Y | Y | Y | 0 |  |  |  | 2024-01-29 |  | EMS |  | DimensionlessRatio | ppm | N | (ratio) |  | Tantalum (ratio) | 7440-25-7 | Metals; | Metals; EMS code: TA-T | Chemical |  | LAB
TA-T | ppm (S) | Tantalum | MS81 | ALS Minerals: ME-MS81 | ALS Minerals: ME-MS81. A prepared sample (0.100 g) is added to lithium metaborate/lithium tetraborate flux; mixed well and fused in a furnace at 1025°C. The resulting melt is then cooled and dissolved in an acid mixture containing nitric; hydrochloric and hydrofluoric acids. This solution is then analyzed by inductively coupled plasma - mass spectrometry. | 0.1 | Y | Y | Y | 0 |  |  |  | 2024-01-29 |  | EMS |  | DimensionlessRatio | ppm | N | (ratio) |  | Tantalum (ratio) | 7440-25-7 | Metals; | Metals; EMS code: TA-T | Chemical |  | LAB
TA-T | mg/L | Tantalum | F082 | Aliquot,HNO3/HCL dig,ICPMS | Aliquot microwave digested with 3:1 HNO3/HCL. Analysis by ICPMS. For HiVol samples preliminary result are reported in mg/L and loading results are reported in ug/m3 | 0.0001 | Y | Y | Y | 69 | 7/30/2020 | 8/16/2024 | 8/21/2024 11:26:07 AM | 2021-11-03 |  | RDGIBSON |  | Concentration | mg/L | Y | (fl. conc.) |  | Tantalum (fl. conc.) | 7440-25-7 | Metals; | Metals; EMS code: TA-T | Chemical |  | LAB
TA-T | ppm (S) | Tantalum | MS61 | ALS Minerals ME-MS61 | ALS Minerals ME-MS61. A prepared sample is digested with HNO3; HF; and HClO4 in a graphite heating block to near dryness. After cooling; the sample is boiled in diluted HCl. Bulk to volume with deionized water; mixed and analyzed by ICP-AES and ICP-MS for all elements. | 0.05 | Y | Y | Y | 0 |  |  |  | 2024-01-29 |  | EMS |  | DimensionlessRatio | ppm | N | (ratio) |  | Tantalum (ratio) | 7440-25-7 | Metals; | Metals; EMS code: TA-T | Chemical |  | LAB
TA-T | mg/L | Tantalum | METN | MET-T-NP-CCMS-VA | Ultra trace metals in water are analyzed by CRC ICPMS, based on US EPA Method 6020B (July 2014). The detection limits provided can only be met for undigested samples. This procedure is intended for colorless, non-turbid, acid-preserved water samples (i.e. pristine water samples), having turbidity < 1 NTU and no odor. Where turbidity exceeds 1 NTU, and/or the sample is colored and has an odor, results may be biased low compared to true Total Metals concentrations. ALS recommends that turbidity analysis be requested on samples submitted for this test to aid with interpretation of results. Where turbidity is <1NTU, undigested metals are equivalent to total metals concentrations. | 0.0001 | Y | Y | Y | 280 | 8/6/2019 | 4/4/2024 | 4/29/2024 4:53:06 PM | 2021-06-10 |  | RDGIBSON |  | Concentration | mg/L | Y | (fl. conc.) |  | Tantalum (fl. conc.) | 7440-25-7 | Metals; | Metals; EMS code: TA-T | Chemical |  | LAB
TA-T | ppm (S) | Tantalum | 6042 | Aqua Regia by ICPMS | These were sediment/soils digested with Aqua-Regia (Nitric Acid & Hydrochloric Acids) and then analysis by ICPMS. Results would be reported in ppm s. | 0.005 | Y | Y | Y | 71 | 5/14/2022 | 7/27/2024 | 7/29/2024 9:13:08 AM | 2022-06-28 |  | EMS |  | DimensionlessRatio | ppm | N | (ratio) |  | Tantalum (ratio) | 7440-25-7 | Metals; | Metals; EMS code: TA-T | Chemical |  | LAB
TB-T | ppm (S) | Terbium | MS81 | ALS Minerals: ME-MS81 | ALS Minerals: ME-MS81. A prepared sample (0.100 g) is added to lithium metaborate/lithium tetraborate flux; mixed well and fused in a furnace at 1025°C. The resulting melt is then cooled and dissolved in an acid mixture containing nitric; hydrochloric and hydrofluoric acids. This solution is then analyzed by inductively coupled plasma - mass spectrometry. | 0.01 | Y | Y | Y | 0 |  |  |  | 2024-01-29 |  | EMS |  | DimensionlessRatio | ppm | N | (ratio) |  | Terbium (ratio) |  | Metals; | Metals; EMS code: TB-T | Chemical |  | LAB
TBZ | mg/L | Thiabendazole CAS: 148-79-8 | XMSP | HPLC and Tri. Quadpole MS/MS | EPA 1694 An aliquot of filtered sample containing 5% organic is injected directly. Each compound is separated by reversed phase HPLC. Analyzed by Electro Spray Ionization MS/MS detection using Triple Quadrupole MS/MS detector. Quantiitation with internal or external standard. |  | Y | Y | Y | 12 | 5/8/2020 | 6/4/2024 | 8/21/2024 3:32:03 PM | 2022-05-04 |  | EMS |  | Concentration | mg/L | Y | (fl. conc.) |  | Thiabendazole (fl. conc.) | 148-79-8 | Petroleum Hydrocarbons; | Petroleum Hydrocarbons; EMS code: TBZ | Chemical |  | LAB
TC63UM | % | Total Carbon <63um | FS93 | Oxid to CO2,NDIR,Calc'n | Total Carbon is determined by high temperature oxidation of carbon to carbon diooxide which is then measured by means of a nondispersive infrared analyzer. Inorganic Carbon is determined by reaction with phosphoric acid to convert all carbonates to carbon dioxide which is also measured by means of a nondispersive infrared analyzer. Total Organic Carbon is determined as the difference between Total and Inorganic Carbons. | 0.05 | Y | Y | Y | 80 | 7/15/2022 | 9/21/2023 | 1/12/2024 11:16:52 AM | 2022-07-27 |  | EMS |  | DimensionlessRatio | % | N | (ratio) |  | Total Carbon <63um (ratio) |  | Carbon; | Carbon; EMS code: TC63UM | Chemical | Total | LAB
TCAA | mg/L | Trichloroacetic Acid | HAAW | Haloacetic Acids,MTBE ex | Haloacetic Acids;MTBE ex | 0.4 | Y | Y | Y | 2 | 3/27/2023 9:58:00 AM | 5/31/2023 4:53:00 PM | 6/8/2023 12:14:57 PM | 1996-10-13 | 6/3/1998 6:51:46 PM | EMS |  | Concentration | mg/L | Y | (fl. conc.) |  | Trichloroacetic Acid (fl. conc.) | 76-03-9 | Organics; | Organics; EMS code: TCAA | Chemical |  | LAB
TCPR | mg/L | Triclopyr | P015 | Phenoxyacid Herbicides by LC-MS/MS | Phenoxyacid Herbicides by LC-MS/MS | 5e-05 | Y | Y | Y | 68 | 12/6/2021 | 9/9/2022 | 12/8/2022 12:57:06 PM | 2021-12-09 | 12/9/2021 3:22:12 PM | EMS | EMS | Concentration | mg/L | Y | (fl. conc.) |  | Triclopyr (fl. conc.) | 55335-06-3 | Pesticides; | Pesticides; EMS code: TCPR | Chemical |  | LAB
TCPR | ug/g | Triclopyr | ZN03 | DCM+Ac.Extr:GC/MS | DCM+Ac.Extr:GC/MS EPA 8270E (mod) | 0.005 | Y | Y | Y | 15 | 1/14/2022 | 2/2/2022 | 1/11/2024 1:07:08 PM | 2022-07-27 |  | EMS |  | MassPerMass | mg/kg | Y | (mass conc.) |  | Triclopyr (mass conc.) | 55335-06-3 | Pesticides; | Pesticides; EMS code: TCPR | Chemical |  | LAB
TDST | mg/dm2/d | Total Dustfall | DUST | HNO3 dig.(microwave or hot block),ICPMS | HNO3 dig.(microwave or hot block);ICPMS |  | Y | Y | Y | 13 | 1/31/2020 | 1/31/2020 | 1/16/2023 3:51:10 PM | 2008-02-07 |  | RDGIBSON |  | MassFlux | mg/dm²/d | Y | (mass flux) |  | Total Dustfall (mass flux) |  | Particulates; | Particulates; EMS code: TDST | Chemical | Total | LAB
TDST | mg/dm2/d | Total Dustfall | X192 | Gravimetric (Sol+Insol) | Gravimetric (Sol+Insol) |  | Y | Y | Y | 1922 | 9/15/2015 4:14:00 PM | 8/19/2024 | 8/19/2024 4:02:08 PM | 2008-06-04 |  | RDGIBSON |  | MassFlux | mg/dm²/d | Y | (mass flux) |  | Total Dustfall (mass flux) |  | Particulates; | Particulates; EMS code: TDST | Chemical | Total | LAB
TE-D | mg/L | Tellerium Dissolved | F039 | Field filt + acid, ICP-AES | Dissolved and extractable water samples are ready for analysis once collected. Total samples are autoclaved with HNO3 and hydrochloric acid (HCI) in capped Teflon vials. Analysis is performed via Inductively Coupled Argon Plasma Atomic Emission Spectrometry (ICP-AES)(simultaneous multielement analysis), Graphite Furnace Atomic Absorption Spectrometry (GF-AAS)(low level analysis) and/or Flame Atomic Absorption Spectrometry (F-AAS). | 0.002 | Y | Y | Y | 827 | 1/18/2002 | 2/10/2005 8:30:00 AM | 4/6/2005 1:01:55 AM | 1998-05-21 | 6/3/1998 6:51:49 PM | NPEPPIN |  | Concentration | mg/L | Y | (fl. conc.) |  | Tellerium Dissolved (fl. conc.) | 13494-80-9 | Metals - Dissolved; | Metals - Dissolved; EMS code: TE-D | Chemical | Dissolved | LAB
TE-D | mg/L | Tellerium Dissolved | X562 | Dissolved metals: FF, FA, HNO3, HR-ICPMS | Ultra trace metals in water are analyzed by high resolution inductively coupled plasma mass spectrometry (HR-ICPMS), based on US EPA Method 200.8, (Oct. 1999). This procedure is intended for pristine field-filtered acid-preserved water samples. The detection limits (LOR) for this test are based on lab instrumental analysis only, not including filtration. ALS-supplied field filtration equipment does not support these LOR. Therefore, because of the high probability of false positives due to filtration, it is strongly recommended that a filtration blank be analyzed to aid in data interpretation. | 1e-05 | Y | Y | Y | 22 | 10/20/2015 | 10/19/2023 | 11/9/2023 7:41:27 AM | 2016-03-24 |  | RDGIBSON |  | Concentration | mg/L | Y | (fl. conc.) |  | Tellerium Dissolved (fl. conc.) | 13494-80-9 | Metals - Dissolved; | Metals - Dissolved; EMS code: TE-D | Chemical | Dissolved | LAB
TE-D | mg/L | Tellerium Dissolved | METC | MET-D-NP-CCMS-VA | Ultra trace metals in water are analyzed by CRC ICPMS, based on US EPA Method 6020B (July 2014). This procedure is intended for pristine field-filtered acid-preserved water samples. The detection limits (LOR) for this test are based on lab instrumental analysis only, not including filtration. ALS-supplied field filtration equipment does not support these LOR. Therefore, because of the high probability of false positives due to filtration, it is strongly recommended that a filtration blank be analysed to aid in data interpretation. | 2e-05 | Y | Y | Y | 394 | 3/30/2020 | 4/9/2024 | 4/29/2024 4:53:06 PM | 2022-10-07 |  | EMS |  | Concentration | mg/L | Y | (fl. conc.) |  | Tellerium Dissolved (fl. conc.) | 13494-80-9 | Metals - Dissolved; | Metals - Dissolved; EMS code: TE-D | Chemical | Dissolved | LAB
TE-D | mg/L | Tellerium Dissolved | F032 | Filt.,Acid,ICPMS | There is no acid digestion required for this analysis. A water sample is collected and preserved in 0.2 percent nitric acid. To analyse the extractable metals the sample is shaken, allowed to sit for at least twenty four hours, and a decanted portion is introduced to the spectrometer. When the dissolved metals are required, the samples must first be filtered through a 0.45 micron cellulose acetate filter paper and then followed by preservation in 0.2 percent nitric acid. It is important to note that different laboratories may have very different detection limits for some variables analyzed by ICPMS. | 0.001 | Y | Y | Y | 24793 | 5/7/1998 | 8/26/2024 | 8/26/2024 4:51:04 PM | 1998-05-14 | 6/3/1998 6:51:49 PM | NPEPPIN |  | Concentration | mg/L | Y | (fl. conc.) |  | Tellerium Dissolved (fl. conc.) | 13494-80-9 | Metals - Dissolved; | Metals - Dissolved; EMS code: TE-D | Chemical | Dissolved | LAB
TE-D | mg/L | Tellerium Dissolved | MSEA | Metals in Seawater (undigested) by (CRC) ICP-MS | This method describes the multi-elemental determination of trace elements by Collision/Reaction Cell (CRC) ICP-MS. The method measures ions produced by a radio-frequency inductively coupled plasma. Analyte species originating in a liquid are pneumatically nebulized and the resulting aerosol transported by argon gas into the plasma torch. The ions produced are entrained in the plasma gas and extracted, by means of a differentially pumped vacuum interface, into a mass spectrometer. The ions flow through a Collision/Reaction Cell (CRC) where depending on the mode of analysis either the ions passes through the cell unhindered (no gas - classical ICP-MS mode), or collisions/reactions chemistry takes place between the ions and a neutral gas such as hydrogen or helium. Any polyatomic ions which give interference on the analytes are removed. The ions are then sorted according to their mass-to-charge ratios by a quadrupole mass spectrometer. The ions are then sorted according to their mass-to-charge ratios by a quadrupole mass spectrometer. The ions transmitted through the quadrupole are quantified by a channel electron multiplier and the ion information processed by a data handling system. Although most polyatomic interference is removed, any additional interference must be assessed and valid corrections applied or the data flagged to indicate problems. Interference correction must include compensation for background ions contributed by the plasma gas, reagents, and constituents of the sample matrix. Instrument drift as well as suppressions or enhancements of instrument response caused by the sample matrix must be corrected for by the use of internal standards. | 6e-05 | Y | Y | Y | 4 | 4/12/2023 | 4/29/2024 | 5/16/2024 2:00:14 PM | 2009-09-08 |  | RDGIBSON |  | Concentration | mg/L | Y | (fl. conc.) |  | Tellerium Dissolved (fl. conc.) | 13494-80-9 | Metals - Dissolved; | Metals - Dissolved; EMS code: TE-D | Chemical | Dissolved | LAB
TE-D | mg/L | Tellerium Dissolved | MWID | Metals in Water by ICPMS (digested) | Metals in Water by ICPMS (digested) Based on EPA 200.8 | 0.001 | Y | Y | Y | 123 |  |  | 1/27/2023 10:37:54 AM | 2009-12-24 |  | RDGIBSON |  | Concentration | mg/L | Y | (fl. conc.) |  | Tellerium Dissolved (fl. conc.) | 13494-80-9 | Metals - Dissolved; | Metals - Dissolved; EMS code: TE-D | Chemical | Dissolved | LAB
TE-D | mg/L | Tellerium Dissolved | X170 | ICP | This code is a catch-all for a large variety of ICP methods including water and also solids digested into solutions as an intermediate result. Different procedures will have greatly different detection limits. (a) For air particulates, preliminary results are reported in mg/L and loading results as mg/dm2/d. (b) For Hi Vol samples preliminiary results are same as reported in mg/L and loading results as ug/m3. (c) Note that for Rain samples there is no filtration nor digestion steps. | 0.02 | Y | Y | Y | 2156 | 12/3/1996 | 6/24/2002 | 10/5/2010 10:15:38 AM | 1996-04-12 | 6/3/1998 6:51:40 PM | EMS |  | Concentration | mg/L | Y | (fl. conc.) |  | Tellerium Dissolved (fl. conc.) | 13494-80-9 | Metals - Dissolved; | Metals - Dissolved; EMS code: TE-D | Chemical | Dissolved | LAB
TE-D | mg/L | Tellerium Dissolved | X285 | HNO3 Dig:ICP Analysis | HNO3 Dig:ICP Analysis | 0.002 | Y | Y | Y | 587 | 8/21/2002 | 11/13/2002 | 2/19/2004 4:22:44 PM | 1998-05-22 | 6/3/1998 6:51:49 PM | NPEPPIN |  | Concentration | mg/L | Y | (fl. conc.) |  | Tellerium Dissolved (fl. conc.) | 13494-80-9 | Metals - Dissolved; | Metals - Dissolved; EMS code: TE-D | Chemical | Dissolved | LAB
TE-T | mg/L | Tellurium Total | METN | MET-T-NP-CCMS-VA | Ultra trace metals in water are analyzed by CRC ICPMS, based on US EPA Method 6020B (July 2014). The detection limits provided can only be met for undigested samples. This procedure is intended for colorless, non-turbid, acid-preserved water samples (i.e. pristine water samples), having turbidity < 1 NTU and no odor. Where turbidity exceeds 1 NTU, and/or the sample is colored and has an odor, results may be biased low compared to true Total Metals concentrations. ALS recommends that turbidity analysis be requested on samples submitted for this test to aid with interpretation of results. Where turbidity is <1NTU, undigested metals are equivalent to total metals concentrations. | 2e-05 | Y | Y | Y | 280 | 8/6/2019 | 4/4/2024 | 4/29/2024 4:53:06 PM | 2021-06-16 |  | EMS |  | Concentration | mg/L | Y | (fl. conc.) |  | Tellurium Total (fl. conc.) | 13494-80-9 | Metals - Total; | Metals - Total; EMS code: TE-T | Chemical | Total | LAB
TE-T | ug/g | Tellurium Total | 0190 | D&G:N/P Dig:ICP Anal. | Sample is dried and grind-mixed before nitric perchloric digestion followed by ICP analysis. | 20 | Y | Y | Y | 9 |  |  | 10/23/1996 10:17:01 PM | 1996-04-12 | 6/3/1998 6:51:40 PM | EMS |  | MassPerMass | mg/kg | Y | (mass conc.) |  | Tellurium Total (mass conc.) | 13494-80-9 | Metals - Total; | Metals - Total; EMS code: TE-T | Chemical | Total | LAB
TE-T | ug/g | Tellurium Total | X551 | Metals in Biota (dry), total, acid digestion, ICP-SFMS | 0.1 gram of dried, homogenized biota is weighed in a microwave TFM vessel, and 8mL of Nitric Acid and 1mL of Hydrogen Peroxide are added. The sample is digested at 200 degrees C for 15 minutes in a closed vessel high pressure microwave oven. The sample is cooled, the volume adjusted to 100mL and introduced to an iductively coupled argon plasma-sector field mass spectrometer (ICP-SFMS) by pneumatic nebulization. Each respective element is measured at a specific mass to charge ratio, m/z value expressed in amu (atomic mass units), and at a specified instrument resolution. LR: Low Resolution ~300, MR: Mid Resolution ~4000 and HR: Hich Resolution ~10000. Ref: Method number 02-2705, National Laboratory for Environmental Testing, CCIW, Environment Canada, Burlington. Requested April 2003. | 0.02 | Y | Y | Y | 43 | 8/11/2015 | 10/30/2023 | 11/28/2023 2:21:33 PM | 2015-03-11 |  | RDGIBSON |  | MassPerMass | mg/kg | Y | (mass conc.) |  | Tellurium Total (mass conc.) | 13494-80-9 | Metals - Total; | Metals - Total; EMS code: TE-T | Chemical | Total | LAB
TE-T | ug/g wet | Tellurium Total | BPSY | MET-WET-BPSY-CCMS-VA | Metals in Animal Tissue and Vegetation (Biota) - Prescriptive. EPA 200.3/6020B This method is conducted following British Columbia Lab Manual method "Metals in Animal Tissue and Vegetation (Biota) - Prescriptive". Tissue samples are digested with nitric and hydrochloric acids; in combination with addition of hydrogen peroxide. Instrumental analysis is by collision cell inductively coupled plasma - mass spectrometry (modified from EPA Method 6020B). | 0.004 | Y | Y | Y | 2 | 11/7/2019 | 11/7/2019 | 6/10/2021 1:57:10 PM | 2021-06-04 |  | EMS |  | MassPerMass | µg/g | Y | (mass conc.) | (wet) | Tellurium Total (wet) (mass conc.) | 13494-80-9 | Metals - Total; | Metals - Total; EMS code: TE-T | Chemical | Total | LAB
TE-T | mg/L | Tellurium Total | MWID | Metals in Water by ICPMS (digested) | Metals in Water by ICPMS (digested) Based on EPA 200.8 | 0.001 | Y | Y | Y | 46 |  |  | 1/19/2023 12:03:00 PM | 2009-12-24 |  | RDGIBSON |  | Concentration | mg/L | Y | (fl. conc.) |  | Tellurium Total (fl. conc.) | 13494-80-9 | Metals - Total; | Metals - Total; EMS code: TE-T | Chemical | Total | LAB
TE-T | ppm (S) | Tellurium Total | 6042 | Aqua Regia by ICPMS | These were sediment/soils digested with Aqua-Regia (Nitric Acid & Hydrochloric Acids) and then analysis by ICPMS. Results would be reported in ppm s. | 0.01 | Y | Y | Y | 71 | 5/14/2022 | 7/27/2024 | 7/29/2024 9:13:08 AM | 2022-06-28 |  | EMS |  | DimensionlessRatio | ppm | N | (ratio) |  | Tellurium Total (ratio) | 13494-80-9 | Metals - Total; | Metals - Total; EMS code: TE-T | Chemical | Total | LAB
TE-T | ug/m3 | Tellurium Total | 5312 | Tef:N/P Dig:ICP Analysis | A nitric/perchloric acid digestion if used to bring the metals into solution prior to ICP analysis. For Hi-Vol samples preliminary results are reported as mg/L and 24 hour loading results as ug/m3. | 0.01 | Y | Y | Y | 1191 | 11/7/1997 | 3/29/2000 | 9/28/2000 1:34:46 PM | 1996-10-13 | 6/3/1998 6:51:46 PM | EMS |  | AirConcentration | µg/m³ | Y | (air conc.) |  | Tellurium Total (air conc.) | 13494-80-9 | Metals - Total; | Metals - Total; EMS code: TE-T | Chemical | Total | LAB
TE-T | mg/L | Tellurium Total | 0042 | GI HNO3 Dig: ICP | Sample aliquot taken from unfiltered unpreserved General Ions bottle; HNO3 digestion; ICP analysis. | 0.02 | Y | Y | Y | 6295 | 9/20/1996 | 2/16/2005 | 8/18/2006 1:38:31 AM | 1996-04-12 | 6/3/1998 6:51:40 PM | EMS |  | Concentration | mg/L | Y | (fl. conc.) |  | Tellurium Total (fl. conc.) | 13494-80-9 | Metals - Total; | Metals - Total; EMS code: TE-T | Chemical | Total | LAB
TE-T | mg/L | Tellurium Total | MSEA | Metals in Seawater (undigested) by (CRC) ICP-MS | This method describes the multi-elemental determination of trace elements by Collision/Reaction Cell (CRC) ICP-MS. The method measures ions produced by a radio-frequency inductively coupled plasma. Analyte species originating in a liquid are pneumatically nebulized and the resulting aerosol transported by argon gas into the plasma torch. The ions produced are entrained in the plasma gas and extracted, by means of a differentially pumped vacuum interface, into a mass spectrometer. The ions flow through a Collision/Reaction Cell (CRC) where depending on the mode of analysis either the ions passes through the cell unhindered (no gas - classical ICP-MS mode), or collisions/reactions chemistry takes place between the ions and a neutral gas such as hydrogen or helium. Any polyatomic ions which give interference on the analytes are removed. The ions are then sorted according to their mass-to-charge ratios by a quadrupole mass spectrometer. The ions are then sorted according to their mass-to-charge ratios by a quadrupole mass spectrometer. The ions transmitted through the quadrupole are quantified by a channel electron multiplier and the ion information processed by a data handling system. Although most polyatomic interference is removed, any additional interference must be assessed and valid corrections applied or the data flagged to indicate problems. Interference correction must include compensation for background ions contributed by the plasma gas, reagents, and constituents of the sample matrix. Instrument drift as well as suppressions or enhancements of instrument response caused by the sample matrix must be corrected for by the use of internal standards. | 6e-05 | Y | Y | Y | 20 | 4/13/2023 | 4/29/2024 | 6/7/2024 10:03:16 AM | 2009-09-08 |  | RDGIBSON |  | Concentration | mg/L | Y | (fl. conc.) |  | Tellurium Total (fl. conc.) | 13494-80-9 | Metals - Total; | Metals - Total; EMS code: TE-T | Chemical | Total | LAB
TE-T | ppm (S) | Tellurium Total | MS61 | ALS Minerals ME-MS61 | ALS Minerals ME-MS61. A prepared sample is digested with HNO3; HF; and HClO4 in a graphite heating block to near dryness. After cooling; the sample is boiled in diluted HCl. Bulk to volume with deionized water; mixed and analyzed by ICP-AES and ICP-MS for all elements. | 0.05 | Y | Y | Y | 0 |  |  |  | 2024-01-29 |  | EMS |  | DimensionlessRatio | ppm | N | (ratio) |  | Tellurium Total (ratio) | 13494-80-9 | Metals - Total; | Metals - Total; EMS code: TE-T | Chemical | Total | LAB
TE-T | mg/L | Tellurium Total | X561 | Total metals: FA, HNO3, HR-ICPMS (No Prep)Prep) | Ultra trace metals in water are analyzed by high resolution inductively coupled plasma mass spectrometry (HR-ICPMS), modified from US EPA 200.8, (Revision 5.5). The detection limits provided can only be met for undigested samples. This procedure is intended for colorless, non-turbid, acid-preserved water samples (i.e. pristine water samples), having turbidity < 1 NTU and no odor. Where turbidity exceeds 1 NTU, and/or the sample is colored and has an odor, results may be biased low compared to true Total Metals concentrations. ALS recommends that turbidity analysis be requested on samples submitted for this test to aid with interpretation of results. | 1e-05 | Y | Y | Y | 20 | 10/20/2015 | 10/5/2023 | 1/31/2024 2:37:41 PM | 2016-03-24 |  | RDGIBSON |  | Concentration | mg/L | Y | (fl. conc.) |  | Tellurium Total (fl. conc.) | 13494-80-9 | Metals - Total; | Metals - Total; EMS code: TE-T | Chemical | Total | LAB
TE-T | ppm (S) | Tellurium Total | MS41 | ALS Minerals ME-MS41 | ALS Minerals ME-MS41. A prepared sample is digested with aqua regia in a graphite heating block. After cooling; the resulting solution is diluted to with deionized water; mixed and analyzed by ICP-AES. Samples are then analyzed by ICP-MS for the remaining suite of elements. | 0.01 | Y | Y | Y | 0 |  |  |  | 2024-01-29 |  | EMS |  | DimensionlessRatio | ppm | N | (ratio) |  | Tellurium Total (ratio) | 13494-80-9 | Metals - Total; | Metals - Total; EMS code: TE-T | Chemical | Total | LAB
TE-T | ug/g wet | Tellurium Total | X560 | Metals in Biota (wet), total, acid digestion, ICP-MS | Five grams of wet, homogenized biota is weighed in a centrifuge tube, and 2.5 mL of Nitric Acid and 2.5 mL of Hydrochloric acid are added. The sample is digested at room temperature for one hour and then digested at 90-95 degrees C in a water bath or heat block for one more hour. The sample is cooled and the volume is made to 50 mL with DI. The digestate is then diluted at ten fold and introduced to a collision/reaction cell inductively coupled argon plasma mass spectrometer (CRC ICPMS) by pneumatic nebulization. Each respective element is measured at a specific mass to charge ratio, m/z value expressed in amu (atomic mass units). - Maxxam Analytics (BC). Added to EMS Dec 10, 2008. | 0.02 | Y | Y | Y | 181 | 10/19/2009 | 10/30/2023 | 11/28/2023 2:21:33 PM | 2010-02-01 | 3/11/2015 4:17:10 PM | RDGIBSON | EMS | MassPerMass | µg/g | Y | (mass conc.) | (wet) | Tellurium Total (wet) (mass conc.) | 13494-80-9 | Metals - Total; | Metals - Total; EMS code: TE-T | Chemical | Total | LAB
TE-T | ug/g | Tellurium Total | FS82 | PBM,dry60C,10mesh,SALM-ICP-MS | PBM;dry60C;10mesh;SALM-ICP-MS | 0.1 | Y | Y | Y | 212 | 1/31/2002 | 3/16/2024 6:06:00 PM | 3/26/2024 8:53:57 AM | 2002-11-12 |  | UMULLER |  | MassPerMass | mg/kg | Y | (mass conc.) |  | Tellurium Total (mass conc.) | 13494-80-9 | Metals - Total; | Metals - Total; EMS code: TE-T | Chemical | Total | LAB
TE-T | mg/L | Tellurium Total | F033 | Acid Dig,Centrifuge,ICPMS | For the determination of total recoverable elements in water or waste water, take 100 ml aliquot from a well mixed, acid preserved, sample containing not more than 0.25% (w/v) total solids and transfer to a 250 ml Griffin beaker. (If total solids are greater than 0.25% reduce the size of the aliquot by a proportionate amount). Add 1 ml of conc. nitric acid and 0.5 ml conc. hydrochloric acid. Heat on a hot plate until the volume has been reduced to 15-20ml, ensuring that the sample does not boil. Allow to cool and transfer to a 50 ml volumetric flask. Dilute to volume with ASTM type I water. Centrifuge the sample or allow to stand overnight to separate insoluble material. Prior analysis, pipette 20 ml into a 50 ml volumetric flask and dilute to volume with ASTM type I water. Analysis should be performed within seventy two hours of the completed preparation. Note that different laboratories may have very different field preservation, digestion procedures and detection limits for this method, so these should be documented. | 0.001 | Y | Y | Y | 13035 | 5/7/1998 | 8/22/2024 3:24:00 AM | 8/25/2024 9:47:36 PM | 1998-05-14 | 6/3/1998 6:51:49 PM | NPEPPIN |  | Concentration | mg/L | Y | (fl. conc.) |  | Tellurium Total (fl. conc.) | 13494-80-9 | Metals - Total; | Metals - Total; EMS code: TE-T | Chemical | Total | LAB
TE-T | ug/g | Tellurium Total | FS01 | D&G,Nitric/HCl:ICP | Samples are dried at 60 C, then either sieved to -100mesh (150 um) and rolled to homogenize, or finely ground with a mortar and pestle or a shatter box laboratory mill and rolled (depending on the submitter's requirements). The samples are then weighed into Teflon digestion liners and digested with HNO3 and HCl in a microwave oven. Samples are cooled, volumized, and allowed to settle overnight before being decanted. Analysis is then performed via Inductively Coupled Argon Plasma Atomic Emission Spectrometry (ICP-AES) (Simultaneous multi element analysis). In ICP Spectrometry, the liquid sample is pumped to the nebulizer, where it is convertd into an aerosol. the nebulizer carrier gas flow than transports the wet aerosol past the impact bead, where the large droplets condense out and flow to the waste container. The remaining fine aerosol is carried into the plasma. There, the high temperature desolvates the aerosol and dissociates the sample material into individual atoms and ions, which are excited to emit light at wavelengths characteristic of the elements present. A diffraction grating in the spectrometer diffracts the various wavelengths of light through exit slits to the photo-multiplier tubes, which convert the light energy to electronic signals. | 5 | Y | Y | Y | 271 | 1/31/2002 | 3/22/2002 | 4/1/2004 1:19:35 AM | 2004-02-10 |  | UMULLER |  | MassPerMass | mg/kg | Y | (mass conc.) |  | Tellurium Total (mass conc.) | 13494-80-9 | Metals - Total; | Metals - Total; EMS code: TE-T | Chemical | Total | LAB
TE-T | mg/L | Tellurium Total | X285 | HNO3 Dig:ICP Analysis | HNO3 Dig:ICP Analysis | 0.002 | Y | Y | Y | 547 | 8/10/1993 | 10/5/1993 | 2/5/2001 4:22:14 PM | 1998-05-22 | 6/3/1998 6:51:49 PM | NPEPPIN |  | Concentration | mg/L | Y | (fl. conc.) |  | Tellurium Total (fl. conc.) | 13494-80-9 | Metals - Total; | Metals - Total; EMS code: TE-T | Chemical | Total | LAB
TE-T | ug/g wet | Tellurium Total | X552 | Metals in Biota (wet), total, acid digestion, ICP-SFMS | One gram of wet, homogenized biota is weighed in a microwave TFM vessel, and 8mL of Nitric Acid and 1mL of Hydrogen Peroxide are added. The sample is digested at 200 degrees C for 15 minutes in a closed vessel high pressure microwave oven. The sample is cooled, the volume adjusted to 100mL and introduced to an inductively coupled argon plasma-sector field mass spectrometer (ICP-SFMS) by pneumatic nebulization. Each respective element is measured at a specific mass to charge ratio, m/z value expressed in amu (atomic mass units), and at a specified instrument resolution. LR: Low Resolution ~300, MR: Mid Resolution ~4000. Ref: 02-2704, National Laboratory for Environmental Testing, CCIW, Environment Canada, Burlington. Requested April 2003. | 0.004 | Y | Y | Y | 37 | 8/9/2016 | 10/16/2023 | 11/28/2023 2:21:33 PM | 2015-03-11 |  | RDGIBSON |  | MassPerMass | µg/g | Y | (mass conc.) | (wet) | Tellurium Total (wet) (mass conc.) | 13494-80-9 | Metals - Total; | Metals - Total; EMS code: TE-T | Chemical | Total | LAB
TE-T | ug/g | Tellurium Total | 0192 | Hom, Nitric/Perchlor,ICP | Sample already in a dry state so does not require drying before homogenization and nitric perchlorid digestion followed by ICP analysis. | 20 | Y | Y | Y | 192 |  |  | 10/24/1996 11:23:42 AM | 1996-04-12 | 6/3/1998 6:51:40 PM | EMS |  | MassPerMass | mg/kg | Y | (mass conc.) |  | Tellurium Total (mass conc.) | 13494-80-9 | Metals - Total; | Metals - Total; EMS code: TE-T | Chemical | Total | LAB
TE-T | mg/L | Tellurium Total | F038 | HNO3/HCl dig,ICP-AES | Dissolved and extractable water samples are ready for analysis once collected. Total samples are autoclaved with HNO3 and hydrochloric acid (HCI) in capped Teflon vials. Analysis is performed via Inductively Coupled Argon Plasma Atomic Emission Spectrometry (ICP-AES)(simultaneous multielement analysis), Graphite Furnace Atomic Absorption Spectrometry (GF-AAS) (low level analysis) and/or Flame Atomic Absorption Spectrometry (F-AAS). | 0.002 | Y | Y | Y | 1 | 3/22/2002 | 3/22/2002 | 2/13/2004 1:01:35 AM | 1998-05-21 | 6/3/1998 6:51:49 PM | NPEPPIN |  | Concentration | mg/L | Y | (fl. conc.) |  | Tellurium Total (fl. conc.) | 13494-80-9 | Metals - Total; | Metals - Total; EMS code: TE-T | Chemical | Total | LAB
TE-T | mg/L | Tellurium Total | F082 | Aliquot,HNO3/HCL dig,ICPMS | Aliquot microwave digested with 3:1 HNO3/HCL. Analysis by ICPMS. For HiVol samples preliminary result are reported in mg/L and loading results are reported in ug/m3 | 0.0002 | Y | Y | Y | 9143 | 4/6/2015 | 8/23/2024 | 8/26/2024 3:51:07 PM | 2015-07-30 |  | RDGIBSON |  | Concentration | mg/L | Y | (fl. conc.) |  | Tellurium Total (fl. conc.) | 13494-80-9 | Metals - Total; | Metals - Total; EMS code: TE-T | Chemical | Total | LAB
TE-T | mg/L | Tellurium Total | 1349 | Tot Metals in Seawater by CRC ICPMS (BC) | Tot Metals in Seawater by CRC ICPMS (BC) | 0.0005 | Y | Y | Y | 98 | 8/16/2019 | 12/5/2021 | 12/13/2021 12:34:17 PM | 2019-09-30 |  | KATFRASE |  | Concentration | mg/L | Y | (fl. conc.) |  | Tellurium Total (fl. conc.) | 13494-80-9 | Metals - Total; | Metals - Total; EMS code: TE-T | Chemical | Total | LAB
TE-T | mg/L | Tellurium Total | 0040 | HNO3 Dig: ICP Analysis | HNO3 Dig: ICP Analysis | 0.02 | Y | Y | Y | 1269 |  |  | 10/24/1996 1:54:30 PM | 1996-04-12 | 6/3/1998 6:51:40 PM | EMS |  | Concentration | mg/L | Y | (fl. conc.) |  | Tellurium Total (fl. conc.) | 13494-80-9 | Metals - Total; | Metals - Total; EMS code: TE-T | Chemical | Total | LAB
TE-T | ug/g wet | Tellurium Total | MTIW | MET-WET-CCMS-MID-VA | Metals in Tissue by CRC ICPMS (WET) Methodology: This method is conducted following British Columbia Lab Manual method "Metals in Animal Tissue and Vegetation (Biota) - Prescriptive". Tissue samples are homogenized and sub-sampled prior to hotblock digestion with nitric and hydrochloric acids, in combination with addition of hydrogen peroxide. Instrumental analysis is by collision cell inductively coupled plasma - mass spectrometry (modified from EPA Method 6020A). Method Limitation: This method employs a strong acid/peroxide digestion, and is intended to provide a conservative estimate of bio-available metals. Near complete recoveries are achieved for most toxicologically important metals, but elements associated with recalcitrant minerals may be only partially recovered. | 0.004 | Y | Y | Y | 24 | 11/25/2018 | 12/6/2018 | 7/21/2021 2:55:08 PM | 2018-12-24 |  | EMS |  | MassPerMass | µg/g | Y | (mass conc.) | (wet) | Tellurium Total (wet) (mass conc.) | 13494-80-9 | Metals - Total; | Metals - Total; EMS code: TE-T | Chemical | Total | LAB
TE-T | ug/g | Tellurium Total | MTID | MET-DRY-CCMS-VA | Metals in Tissue by CRC ICPMS (DRY) Methodology: This method is conducted following British Columbia Lab Manual method "Metals in Animal Tissue and Vegetation (Biota) - Prescriptive". Tissue samples are homogenized and sub-sampled prior to hotblock digestion with nitric and hydrochloric acids, in combination with addition of hydrogen peroxide. Instrumental analysis is by collision cell inductively coupled plasma - mass spectrometry (modified from EPA Method 6020A). Method Limitation: This method employs a strong acid/peroxide digestion, and is intended to provide a conservative estimate of bio-available metals. Near complete recoveries are achieved for most toxicologically important metals, but elements associated with recalcitrant minerals may be only partially recovered. | 0.02 | Y | Y | Y | 122 | 8/13/2014 | 9/8/2022 | 10/26/2022 7:14:15 AM | 2015-03-09 | 3/11/2015 4:12:05 PM | RDGIBSON | EMS | MassPerMass | mg/kg | Y | (mass conc.) |  | Tellurium Total (mass conc.) | 13494-80-9 | Metals - Total; | Metals - Total; EMS code: TE-T | Chemical | Total | LAB
TE-T | mg/L | Tellurium Total | X170 | ICP | This code is a catch-all for a large variety of ICP methods including water and also solids digested into solutions as an intermediate result. Different procedures will have greatly different detection limits. (a) For air particulates, preliminary results are reported in mg/L and loading results as mg/dm2/d. (b) For Hi Vol samples preliminiary results are same as reported in mg/L and loading results as ug/m3. (c) Note that for Rain samples there is no filtration nor digestion steps. | 0.02 | Y | Y | Y | 155 | 6/9/2003 12:01:00 PM | 6/13/2003 4:01:00 PM | 10/24/2022 10:58:25 AM | 1996-04-12 | 6/3/1998 6:51:40 PM | EMS |  | Concentration | mg/L | Y | (fl. conc.) |  | Tellurium Total (fl. conc.) | 13494-80-9 | Metals - Total; | Metals - Total; EMS code: TE-T | Chemical | Total | LAB
TE-T | ug/g | Tellurium Total | EQ02 | EQUIS (ug/g) | EQUIS (ug/g) | 20 | N | Y | Y | 741 |  |  | 10/23/1996 12:08:28 AM | 1996-04-12 | 6/3/1998 6:51:40 PM | EMS |  | MassPerMass | mg/kg | Y | (mass conc.) |  | Tellurium Total (mass conc.) | 13494-80-9 | Metals - Total; | Metals - Total; EMS code: TE-T | Chemical | Total | LAB
TE-T | ppm (S) | Tellurium Total | 6041 | Aqua Regia by ICPMS | These were sediment/soils digested with Aqua-Regia (Nitric Acid & Hydrochloric Acids) and then analysis by ICPMS. Results would be reported in mg/kg and/or %. | 0.003 | Y | Y | Y | 21 | 10/13/2019 | 8/5/2020 | 8/10/2020 3:52:07 PM | 2019-09-06 | 10/21/2019 11:56:37 AM | KATFRASE | KATFRASE | DimensionlessRatio | ppm | N | (ratio) |  | Tellurium Total (ratio) | 13494-80-9 | Metals - Total; | Metals - Total; EMS code: TE-T | Chemical | Total | LAB
TEBU | mg/L | Tebuthiuron | E751 | E751D Pesticides in Water by LC-MS/MS (POS1 List) | E751D Pesticides in Water by LC-MS/MS (POS1 List) (MOECC E3501) | 1e-05 | Y | Y | Y | 39 | 12/9/2021 | 12/31/2021 | 2/15/2022 9:50:06 AM | 2021-12-14 |  | EMS |  | Concentration | mg/L | Y | (fl. conc.) |  | Tebuthiuron (fl. conc.) | 34014-18-1 | Pesticides; | Pesticides; EMS code: TEBU | Chemical |  | LAB
TEBU | ug/g | Tebuthiuron | ZN03 | DCM+Ac.Extr:GC/MS | DCM+Ac.Extr:GC/MS EPA 8270E (mod) | 0.005 | Y | Y | Y | 2 | 11/2/2022 12:53:00 AM | 11/2/2022 4:41:00 AM | 1/7/2023 3:12:11 PM | 2023-01-06 |  | EMS |  | MassPerMass | mg/kg | Y | (mass conc.) |  | Tebuthiuron (mass conc.) | 34014-18-1 | Pesticides; | Pesticides; EMS code: TEBU | Chemical |  | LAB
TEGL | mg/L | Tetraethylene Glycol | GLYC | Glycols in Soil and Water by GC/FID | GLYCOLS IN SOIL AND WATER PURPOSE/PRINCIPLE OF METHOD: Soils are extracted with water and liquids are injected directly or diluted with water. The final extract is analyzed by direct injection gas chromatography with a flame ionization detector (GC/FID) due to the high solubility of glycols. SCOPE: Glycols are commonly used for many applications such as a de-icing agent in aircraft and runway, as an anti-freeze in heating and cooling systems, as an industrial solvent, and in paint and plastics. The end data is used for various purposes including contamination assessment, environmental regulatory compliance and land reclamation. REFERENCES: Nonhalogenated Organics Using GC/FID. US EPA. Test Methods for Evaluating Solid Waste, Physical/Chemical Methods -- SW-846, Method 8015B Revision 2, December 1996 | 10 | Y | Y | Y | 1 | 4/14/2015 10:15:00 PM | 4/14/2015 10:15:00 PM | 4/16/2015 6:00:12 PM | 2012-05-30 | 6/4/2012 12:03:55 PM | RDGIBSON | RDGIBSON | Concentration | mg/L | Y | (fl. conc.) |  | Tetraethylene Glycol (fl. conc.) | 112-60-7 | Organics; | Organics; EMS code: TEGL | Chemical |  | LAB
TEME | ug/g | Temephos | ZN03 | DCM+Ac.Extr:GC/MS | DCM+Ac.Extr:GC/MS EPA 8270E (mod) | 0.05 | Y | Y | Y | 16 | 1/25/2022 | 11/2/2022 4:41:00 AM | 1/7/2023 3:12:11 PM | 2022-06-13 |  | EMS |  | MassPerMass | mg/kg | Y | (mass conc.) |  | Temephos (mass conc.) | 3383-96-8 | Pesticides; | Pesticides; EMS code: TEME | Chemical |  | LAB
TEME | mg/L | Temephos | E660 | EPA 8270E (mod) | EPA 8270E (mod) Miscellaneous Pesticides by GC-MS | 0.001 | Y | Y | Y | 68 | 12/7/2021 | 9/12/2022 | 12/8/2022 12:57:06 PM | 2021-12-09 | 12/9/2021 3:22:12 PM | EMS | EMS | Concentration | mg/L | Y | (fl. conc.) |  | Temephos (fl. conc.) | 3383-96-8 | Pesticides; | Pesticides; EMS code: TEME | Chemical |  | LAB
TEMF | C | Temperature-Field | FLD | Field sampling | Field sampling method | -50 | Y | Y | Y | 10079 | 10/28/2015 | 8/26/2024 | 8/26/2024 4:51:04 PM | 2019-07-10 | 6/30/2020 4:45:36 PM | KATFRASE | RDGIBSON | Temperature | degC | Y | (temp.) |  | Temperature-Field (temp.) |  | Physical Properties; | Physical Properties; EMS code: TEMF | Physical |  | FIELD_RESULT | Field sampling
TEMF | C | Temperature-Field | XM02 | Thermometer | Thermometer | -20 | Y | Y | Y | 49721 | 4/15/2005 10:00:00 AM | 8/13/2024 4:39:00 PM | 8/26/2024 8:06:03 AM | 2012-10-12 | 12/20/2019 11:02:20 AM | RDGIBSON | RDGIBSON | Temperature | degC | Y | (temp.) |  | Temperature-Field (temp.) |  | Physical Properties; | Physical Properties; EMS code: TEMF | Physical |  | FIELD_RESULT | Thermometer
TERP | % | Terphenyl-d14; p- | SURR | Surrogate Recovery | Surrogate Recovery | 0.1 | Y | Y | Y | 168 | 12/7/2021 | 6/4/2024 | 7/4/2024 11:06:09 AM | 2021-12-14 | 12/30/2021 5:20:19 PM | RDGIBSON | RDGIBSON | DimensionlessRatio | % | N | (ratio) |  | Terphenyl-d14; p- (ratio) | 1718-51-0 | Organics; | Organics; EMS code: TERP | Chemical |  | SURROGATE_RESULT
TFAT | mg/L | Total Fatty Acids | E677 | Resins & Fatty Acids are methylated then analyzed by GCMS | EPA 8270E; AB Env. 129; NCASI RA/FA-85.02 (mod) Resins & Fatty Acids are methylated then analyzed by GCMS. Total fatty acids is the sum of arachidic acid; behenic acid; lauric acid; lignoceric acid; linoleic acid; linolenic acid; myristic acid; oleic acid; palmitic acid; and stearic acid. Total resin acids is the sum of abietic acid; 12-chlorodehydroabietic acid; 14-chlorodehydroabietic acid; dehydroabietic acid; 12;14-dichlorodehydroabietic acid; isopimaric acid; levopimaric acid; neoabietic acid; palustric acid; pimaric acid and sandaracopimaric acid. Total resin acids (BC WWQG) is the sum of abietic acid; neoabietic acid; pimaric acid; isopimaric acid; and sandaracopimaric acid. | 0.1 | Y | Y | Y | 12 | 11/28/2022 | 6/15/2024 | 7/8/2024 11:10:02 AM | 2021-12-30 |  | EMS |  | Concentration | mg/L | Y | (fl. conc.) |  | Total Fatty Acids (fl. conc.) |  | Organics; | Organics; EMS code: TFAT | Chemical | Total | LAB
TFAT | mg/L | Total Fatty Acids | RFAS | Separation of resin and fatty acids (RFAs) in water by GC/MS | Scientific Principle The separation of resin and fatty acids (RFAs) in water samples is performed using gas chromatography/mass spectrometry (GC/MS) in the Selected Ion Monitoring (SIM) mode. The separation is based on compound size; boiling point and polarity. Analyte concentration is calculated using the internal standard method. RFA compounds are extracted from water samples using a liquid-liquid solvent extraction with DCM. Extraction occurs after the sample is brought to a pH of ~ 2 and placed in a rotary tumbler (at approx. 40 rpm) for 1 hour. The resulting extract is transferred to a vial; dried with acidified baked sodium sulfate; methylated with diazomethane and concentrated to 1mL under nitrogen prior to analysis. The resulting extracts are vialed and fortified with labeled internal standard compounds prior to analysis by gas chromatography with a (50%-Phenyl)-methylpolysiloxane (DB-17) column and a mass spectrometer detector in SIM Mode. | 0.072 | Y | Y | Y | 1 | 9/17/2021 7:49:00 PM | 9/17/2021 7:49:00 PM | 2/1/2022 2:38:14 PM | 2022-01-19 |  | EMS |  | Concentration | mg/L | Y | (fl. conc.) |  | Total Fatty Acids (fl. conc.) |  | Organics; | Organics; EMS code: TFAT | Chemical | Total | LAB
TFAT | mg/L | Total Fatty Acids | RAFA | RAFA-SUM-CALC-VA | Total Resin&Fatty-acids by GCMS | 0.05 | Y | Y | Y | 15 | 8/28/2014 | 8/15/2017 | 6/7/2024 5:30:17 PM | 2015-03-09 |  | RDGIBSON |  | Concentration | mg/L | Y | (fl. conc.) |  | Total Fatty Acids (fl. conc.) |  | Organics; | Organics; EMS code: TFAT | Chemical | Total | LAB
TGR-B | t/ha | Theoretical Gypsum Requirment (brine) | TGR99 | TGR | See J. Ashworth et al 1999 | 0.1 | Y | Y | Y | 15 | 1/11/2022 | 2/7/2022 | 1/11/2024 1:07:08 PM | 2022-06-30 | 7/5/2022 2:25:38 PM | EMS | EMS | MassPerArea | t/ha | Y | (m./area) |  | Theoretical Gypsum Requirment (brine) (m./area) |  | Physical Properties; | Physical Properties; EMS code: TGR-B | Chemical |  | LAB
TGR-S | t/ha | Theoretical Gypsum Requirment (sodic) | TGR99 | TGR | See J. Ashworth et al 1999 | 0.1 | Y | Y | Y | 15 | 1/11/2022 | 2/7/2022 | 1/11/2024 1:07:08 PM | 2022-06-30 | 7/5/2022 2:25:38 PM | EMS | EMS | MassPerArea | t/ha | Y | (m./area) |  | Theoretical Gypsum Requirment (sodic) (m./area) |  | Physical Properties; | Physical Properties; EMS code: TGR-S | Chemical |  | LAB
TH-D | mg/L | Thorium Dissolved | MSEA | Metals in Seawater (undigested) by (CRC) ICP-MS | This method describes the multi-elemental determination of trace elements by Collision/Reaction Cell (CRC) ICP-MS. The method measures ions produced by a radio-frequency inductively coupled plasma. Analyte species originating in a liquid are pneumatically nebulized and the resulting aerosol transported by argon gas into the plasma torch. The ions produced are entrained in the plasma gas and extracted, by means of a differentially pumped vacuum interface, into a mass spectrometer. The ions flow through a Collision/Reaction Cell (CRC) where depending on the mode of analysis either the ions passes through the cell unhindered (no gas - classical ICP-MS mode), or collisions/reactions chemistry takes place between the ions and a neutral gas such as hydrogen or helium. Any polyatomic ions which give interference on the analytes are removed. The ions are then sorted according to their mass-to-charge ratios by a quadrupole mass spectrometer. The ions are then sorted according to their mass-to-charge ratios by a quadrupole mass spectrometer. The ions transmitted through the quadrupole are quantified by a channel electron multiplier and the ion information processed by a data handling system. Although most polyatomic interference is removed, any additional interference must be assessed and valid corrections applied or the data flagged to indicate problems. Interference correction must include compensation for background ions contributed by the plasma gas, reagents, and constituents of the sample matrix. Instrument drift as well as suppressions or enhancements of instrument response caused by the sample matrix must be corrected for by the use of internal standards. | 2e-06 | Y | Y | Y | 4 | 4/12/2023 | 4/29/2024 | 5/16/2024 2:00:14 PM | 2009-09-08 |  | RDGIBSON |  | Concentration | mg/L | Y | (fl. conc.) |  | Thorium Dissolved (fl. conc.) | 7440-29-1 | Metals - Dissolved; | Metals - Dissolved; EMS code: TH-D | Chemical | Dissolved | LAB
TH-D | mg/L | Thorium Dissolved | F032 | Filt.,Acid,ICPMS | There is no acid digestion required for this analysis. A water sample is collected and preserved in 0.2 percent nitric acid. To analyse the extractable metals the sample is shaken, allowed to sit for at least twenty four hours, and a decanted portion is introduced to the spectrometer. When the dissolved metals are required, the samples must first be filtered through a 0.45 micron cellulose acetate filter paper and then followed by preservation in 0.2 percent nitric acid. It is important to note that different laboratories may have very different detection limits for some variables analyzed by ICPMS. | 0.001 | Y | Y | Y | 24710 | 5/7/1998 | 8/26/2024 | 8/26/2024 4:51:04 PM | 1998-05-14 | 6/22/1998 3:39:51 PM | NPEPPIN | NPEPPIN | Concentration | mg/L | Y | (fl. conc.) |  | Thorium Dissolved (fl. conc.) | 7440-29-1 | Metals - Dissolved; | Metals - Dissolved; EMS code: TH-D | Chemical | Dissolved | LAB
TH-D | mg/L | Thorium Dissolved | X562 | Dissolved metals: FF, FA, HNO3, HR-ICPMS | Ultra trace metals in water are analyzed by high resolution inductively coupled plasma mass spectrometry (HR-ICPMS), based on US EPA Method 200.8, (Oct. 1999). This procedure is intended for pristine field-filtered acid-preserved water samples. The detection limits (LOR) for this test are based on lab instrumental analysis only, not including filtration. ALS-supplied field filtration equipment does not support these LOR. Therefore, because of the high probability of false positives due to filtration, it is strongly recommended that a filtration blank be analyzed to aid in data interpretation. | 5e-06 | Y | Y | Y | 22 | 10/20/2015 | 10/19/2023 | 11/9/2023 7:41:27 AM | 2016-03-24 |  | RDGIBSON |  | Concentration | mg/L | Y | (fl. conc.) |  | Thorium Dissolved (fl. conc.) | 7440-29-1 | Metals - Dissolved; | Metals - Dissolved; EMS code: TH-D | Chemical | Dissolved | LAB
TH-D | mg/L | Thorium Dissolved | MWID | Metals in Water by ICPMS (digested) | Metals in Water by ICPMS (digested) Based on EPA 200.8 | 0.001 | Y | Y | Y | 124 | 10/26/2018 5:20:00 PM | 10/26/2018 5:20:00 PM | 1/27/2023 10:37:54 AM | 2009-12-24 |  | RDGIBSON |  | Concentration | mg/L | Y | (fl. conc.) |  | Thorium Dissolved (fl. conc.) | 7440-29-1 | Metals - Dissolved; | Metals - Dissolved; EMS code: TH-D | Chemical | Dissolved | LAB
TH-D | mg/L | Thorium Dissolved | METC | MET-D-NP-CCMS-VA | Ultra trace metals in water are analyzed by CRC ICPMS, based on US EPA Method 6020B (July 2014). This procedure is intended for pristine field-filtered acid-preserved water samples. The detection limits (LOR) for this test are based on lab instrumental analysis only, not including filtration. ALS-supplied field filtration equipment does not support these LOR. Therefore, because of the high probability of false positives due to filtration, it is strongly recommended that a filtration blank be analysed to aid in data interpretation. | 5e-06 | Y | Y | Y | 394 | 3/30/2020 | 4/9/2024 | 4/29/2024 4:53:06 PM | 2022-10-07 |  | EMS |  | Concentration | mg/L | Y | (fl. conc.) |  | Thorium Dissolved (fl. conc.) | 7440-29-1 | Metals - Dissolved; | Metals - Dissolved; EMS code: TH-D | Chemical | Dissolved | LAB
TH-T | mg/L | Thorium Total | X561 | Total metals: FA, HNO3, HR-ICPMS (No Prep)Prep) | Ultra trace metals in water are analyzed by high resolution inductively coupled plasma mass spectrometry (HR-ICPMS), modified from US EPA 200.8, (Revision 5.5). The detection limits provided can only be met for undigested samples. This procedure is intended for colorless, non-turbid, acid-preserved water samples (i.e. pristine water samples), having turbidity < 1 NTU and no odor. Where turbidity exceeds 1 NTU, and/or the sample is colored and has an odor, results may be biased low compared to true Total Metals concentrations. ALS recommends that turbidity analysis be requested on samples submitted for this test to aid with interpretation of results. | 5e-06 | Y | Y | Y | 20 | 10/20/2015 | 10/5/2023 | 1/31/2024 2:37:41 PM | 2016-03-24 |  | RDGIBSON |  | Concentration | mg/L | Y | (fl. conc.) |  | Thorium Total (fl. conc.) | 7440-29-1 | Metals - Total; | Metals - Total; EMS code: TH-T | Chemical | Total | LAB
TH-T | mg/L | Thorium Total | F033 | Acid Dig,Centrifuge,ICPMS | For the determination of total recoverable elements in water or waste water, take 100 ml aliquot from a well mixed, acid preserved, sample containing not more than 0.25% (w/v) total solids and transfer to a 250 ml Griffin beaker. (If total solids are greater than 0.25% reduce the size of the aliquot by a proportionate amount). Add 1 ml of conc. nitric acid and 0.5 ml conc. hydrochloric acid. Heat on a hot plate until the volume has been reduced to 15-20ml, ensuring that the sample does not boil. Allow to cool and transfer to a 50 ml volumetric flask. Dilute to volume with ASTM type I water. Centrifuge the sample or allow to stand overnight to separate insoluble material. Prior analysis, pipette 20 ml into a 50 ml volumetric flask and dilute to volume with ASTM type I water. Analysis should be performed within seventy two hours of the completed preparation. Note that different laboratories may have very different field preservation, digestion procedures and detection limits for this method, so these should be documented. | 0.001 | Y | Y | Y | 13019 | 5/7/1998 | 8/22/2024 12:35:00 PM | 8/25/2024 9:47:36 PM | 1998-05-14 | 6/22/1998 3:45:21 PM | NPEPPIN | NPEPPIN | Concentration | mg/L | Y | (fl. conc.) |  | Thorium Total (fl. conc.) | 7440-29-1 | Metals - Total; | Metals - Total; EMS code: TH-T | Chemical | Total | LAB
TH-T | ppm (S) | Thorium Total | MS61 | ALS Minerals ME-MS61 | ALS Minerals ME-MS61. A prepared sample is digested with HNO3; HF; and HClO4 in a graphite heating block to near dryness. After cooling; the sample is boiled in diluted HCl. Bulk to volume with deionized water; mixed and analyzed by ICP-AES and ICP-MS for all elements. | 0.01 | Y | Y | Y | 0 |  |  |  | 2024-01-29 |  | EMS |  | DimensionlessRatio | ppm | N | (ratio) |  | Thorium Total (ratio) | 7440-29-1 | Metals - Total; | Metals - Total; EMS code: TH-T | Chemical | Total | LAB
TH-T | mg/L | Thorium Total | F082 | Aliquot,HNO3/HCL dig,ICPMS | Aliquot microwave digested with 3:1 HNO3/HCL. Analysis by ICPMS. For HiVol samples preliminary result are reported in mg/L and loading results are reported in ug/m3 | 0.0001 | Y | Y | Y | 9140 | 4/6/2015 | 8/23/2024 | 8/26/2024 3:51:07 PM | 2015-07-30 |  | RDGIBSON |  | Concentration | mg/L | Y | (fl. conc.) |  | Thorium Total (fl. conc.) | 7440-29-1 | Metals - Total; | Metals - Total; EMS code: TH-T | Chemical | Total | LAB
TH-T | mg/L | Thorium Total | METN | MET-T-NP-CCMS-VA | Ultra trace metals in water are analyzed by CRC ICPMS, based on US EPA Method 6020B (July 2014). The detection limits provided can only be met for undigested samples. This procedure is intended for colorless, non-turbid, acid-preserved water samples (i.e. pristine water samples), having turbidity < 1 NTU and no odor. Where turbidity exceeds 1 NTU, and/or the sample is colored and has an odor, results may be biased low compared to true Total Metals concentrations. ALS recommends that turbidity analysis be requested on samples submitted for this test to aid with interpretation of results. Where turbidity is <1NTU, undigested metals are equivalent to total metals concentrations. | 5e-06 | Y | Y | Y | 280 | 8/6/2019 | 4/4/2024 | 4/29/2024 4:53:06 PM | 2021-06-16 |  | EMS |  | Concentration | mg/L | Y | (fl. conc.) |  | Thorium Total (fl. conc.) | 7440-29-1 | Metals - Total; | Metals - Total; EMS code: TH-T | Chemical | Total | LAB
TH-T | mg/L | Thorium Total | MSEA | Metals in Seawater (undigested) by (CRC) ICP-MS | This method describes the multi-elemental determination of trace elements by Collision/Reaction Cell (CRC) ICP-MS. The method measures ions produced by a radio-frequency inductively coupled plasma. Analyte species originating in a liquid are pneumatically nebulized and the resulting aerosol transported by argon gas into the plasma torch. The ions produced are entrained in the plasma gas and extracted, by means of a differentially pumped vacuum interface, into a mass spectrometer. The ions flow through a Collision/Reaction Cell (CRC) where depending on the mode of analysis either the ions passes through the cell unhindered (no gas - classical ICP-MS mode), or collisions/reactions chemistry takes place between the ions and a neutral gas such as hydrogen or helium. Any polyatomic ions which give interference on the analytes are removed. The ions are then sorted according to their mass-to-charge ratios by a quadrupole mass spectrometer. The ions are then sorted according to their mass-to-charge ratios by a quadrupole mass spectrometer. The ions transmitted through the quadrupole are quantified by a channel electron multiplier and the ion information processed by a data handling system. Although most polyatomic interference is removed, any additional interference must be assessed and valid corrections applied or the data flagged to indicate problems. Interference correction must include compensation for background ions contributed by the plasma gas, reagents, and constituents of the sample matrix. Instrument drift as well as suppressions or enhancements of instrument response caused by the sample matrix must be corrected for by the use of internal standards. | 2e-06 | Y | Y | Y | 20 | 4/13/2023 | 4/29/2024 | 6/7/2024 10:03:16 AM | 2009-09-08 |  | RDGIBSON |  | Concentration | mg/L | Y | (fl. conc.) |  | Thorium Total (fl. conc.) | 7440-29-1 | Metals - Total; | Metals - Total; EMS code: TH-T | Chemical | Total | LAB
TH-T | mg/L | Thorium Total | 1349 | Tot Metals in Seawater by CRC ICPMS (BC) | Tot Metals in Seawater by CRC ICPMS (BC) | 0.0005 | Y | Y | Y | 98 | 8/16/2019 | 12/5/2021 | 12/13/2021 12:34:17 PM | 2019-09-30 |  | KATFRASE |  | Concentration | mg/L | Y | (fl. conc.) |  | Thorium Total (fl. conc.) | 7440-29-1 | Metals - Total; | Metals - Total; EMS code: TH-T | Chemical | Total | LAB
TH-T | ug/g | Thorium Total | FS82 | PBM,dry60C,10mesh,SALM-ICP-MS | PBM;dry60C;10mesh;SALM-ICP-MS | 0.5 | Y | Y | Y | 197 | 12/7/2009 3:52:00 PM | 3/16/2024 6:06:00 PM | 3/26/2024 8:53:57 AM | 2010-03-29 |  | EMS |  | MassPerMass | mg/kg | Y | (mass conc.) |  | Thorium Total (mass conc.) | 7440-29-1 | Metals - Total; | Metals - Total; EMS code: TH-T | Chemical | Total | LAB
TH-T | mg/L | Thorium Total | F121 | Acid,ICPMS | Acidified, no digestion, ICPMS |  | Y | Y | Y | 132 |  |  | 10/24/2022 10:58:25 AM | 2002-07-22 |  | UMULLER |  | Concentration | mg/L | Y | (fl. conc.) |  | Thorium Total (fl. conc.) | 7440-29-1 | Metals - Total; | Metals - Total; EMS code: TH-T | Chemical | Total | LAB
TH-T | ppm (S) | Thorium Total | 6042 | Aqua Regia by ICPMS | These were sediment/soils digested with Aqua-Regia (Nitric Acid & Hydrochloric Acids) and then analysis by ICPMS. Results would be reported in ppm s. | 0.002 | Y | Y | Y | 71 | 5/14/2022 | 7/27/2024 | 7/29/2024 9:13:08 AM | 2022-06-28 |  | EMS |  | DimensionlessRatio | ppm | N | (ratio) |  | Thorium Total (ratio) | 7440-29-1 | Metals - Total; | Metals - Total; EMS code: TH-T | Chemical | Total | LAB
TH-T | ppm (S) | Thorium Total | MS81 | ALS Minerals: ME-MS81 | ALS Minerals: ME-MS81. A prepared sample (0.100 g) is added to lithium metaborate/lithium tetraborate flux; mixed well and fused in a furnace at 1025°C. The resulting melt is then cooled and dissolved in an acid mixture containing nitric; hydrochloric and hydrofluoric acids. This solution is then analyzed by inductively coupled plasma - mass spectrometry. | 0.05 | Y | Y | Y | 0 |  |  |  | 2024-01-29 |  | EMS |  | DimensionlessRatio | ppm | N | (ratio) |  | Thorium Total (ratio) | 7440-29-1 | Metals - Total; | Metals - Total; EMS code: TH-T | Chemical | Total | LAB
TH-T | mg/L | Thorium Total | MWID | Metals in Water by ICPMS (digested) | Metals in Water by ICPMS (digested) Based on EPA 200.8 | 0.001 | Y | Y | Y | 46 |  |  | 1/19/2023 12:03:00 PM | 2009-12-24 |  | RDGIBSON |  | Concentration | mg/L | Y | (fl. conc.) |  | Thorium Total (fl. conc.) | 7440-29-1 | Metals - Total; | Metals - Total; EMS code: TH-T | Chemical | Total | LAB
TH-T | ppm (S) | Thorium Total | MS41 | ALS Minerals ME-MS41 | ALS Minerals ME-MS41. A prepared sample is digested with aqua regia in a graphite heating block. After cooling; the resulting solution is diluted to with deionized water; mixed and analyzed by ICP-AES. Samples are then analyzed by ICP-MS for the remaining suite of elements. | 0.2 | Y | Y | Y | 0 |  |  |  | 2024-01-29 |  | EMS |  | DimensionlessRatio | ppm | N | (ratio) |  | Thorium Total (ratio) | 7440-29-1 | Metals - Total; | Metals - Total; EMS code: TH-T | Chemical | Total | LAB
TH-T | ppm (S) | Thorium Total | 6041 | Aqua Regia by ICPMS | These were sediment/soils digested with Aqua-Regia (Nitric Acid & Hydrochloric Acids) and then analysis by ICPMS. Results would be reported in mg/kg and/or %. | 0.002 | Y | Y | Y | 18 | 10/13/2019 | 3/12/2020 | 3/24/2020 4:13:17 PM | 2019-09-06 | 10/21/2019 11:56:40 AM | KATFRASE | KATFRASE | DimensionlessRatio | ppm | N | (ratio) |  | Thorium Total (ratio) | 7440-29-1 | Metals - Total; | Metals - Total; EMS code: TH-T | Chemical | Total | LAB
THIA | mg/L | Thiamethoxam (Insecticide/Pesticide) | 3544 | MOECC E3544 NEONIC-DI-LCMS-WT | MOECC E3544 NEONIC-DI-LCMS-WT Water sample is filtered; spiked with internal standard and analyzed by direct injection using a HPLC Triple Quadrupole MS/MS with Electro Spray Ionization. | 5e-06 | Y | Y | Y | 29 | 6/5/2020 | 7/27/2020 | 7/21/2021 6:01:34 PM | 2020-06-30 |  | RDGIBSON |  | Concentration | mg/L | Y | (fl. conc.) |  | Thiamethoxam (Insecticide/Pesticide) (fl. conc.) | 153719-23-4 | Pesticides; | Pesticides; EMS code: THIA | Chemical |  | LAB
THIO | mg/L | Thiocyanate Dissolved | F035 | Ferric Nitrate, Colorimetric | A known aliquot of a well mixed sample of fresh water, industrial waste, mine effluent previously preserved with sodium hydroxide, is centrifuged and acidified with nitric acid to a pH 2. Known aliquot of ferric nitrate is added to the acidified sample and absorbance recorded at 480 nm. | 0.5 | Y | Y | Y | 7131 | 1/27/1998 12:42:00 PM | 1/7/2023 4:47:00 PM | 1/11/2023 5:01:27 PM | 1997-02-12 | 6/3/1998 6:51:46 PM | MPAK |  | Concentration | mg/L | Y | (fl. conc.) |  | Thiocyanate Dissolved (fl. conc.) | 302-04-5 | Inorganics; | Inorganics; EMS code: THIO | Chemical | Dissolved | LAB
THIO | mg/L | Thiocyanate Dissolved | X044 | Ion Chromatography | A clear (Total) or filtered (Dissolved) water sample is injected into a stream of carbonate-bicarbonate eluent and passed through a series of ion exchanger. The anions of interested are separated on the basis of their relative affinities for a low capacity , strongly basic anion exchanger (guard and separate columns). The separated anions are directed through the anion self-regenerating suppressor (ASRS). Hydronium ions in the suppressor will neutralize the eluent to its weakly ionized form and convert the separated anions in their conductive acid form. The separated anions are measured by conductivity. They are identified on the basis of retention time as compared to standards. Quantitation is by measurement of peak area of peak height against a calibration curve generated from known standards. | 0.025 | Y | Y | Y | 26554 | 2/12/1997 4:11:00 PM | 8/23/2024 | 8/26/2024 3:51:07 PM | 1996-04-12 | 6/3/1998 6:51:40 PM | EMS |  | Concentration | mg/L | Y | (fl. conc.) |  | Thiocyanate Dissolved (fl. conc.) | 302-04-5 | Inorganics; | Inorganics; EMS code: THIO | Chemical | Dissolved | LAB
THIO | mg/L | Thiocyanate Dissolved | CS01 | System Calculation | System Calculation (mg/L) |  | N | Y | N | 1 |  |  | 2/11/1997 10:02:20 PM | 1996-04-12 | 6/3/1998 6:51:44 PM | EMS |  | Concentration | mg/L | Y | (fl. conc.) |  | Thiocyanate Dissolved (fl. conc.) | 302-04-5 | Inorganics; | Inorganics; EMS code: THIO | Chemical | Dissolved | LAB
THIO | mg/L | Thiocyanate Dissolved | ICAN | Ion Chromatography | Ion Chromatography | 0.025 | N | Y | Y | 511 |  |  | 10/24/1996 3:08:21 PM | 1996-04-12 | 6/3/1998 6:51:40 PM | EMS |  | Concentration | mg/L | Y | (fl. conc.) |  | Thiocyanate Dissolved (fl. conc.) | 302-04-5 | Inorganics; | Inorganics; EMS code: THIO | Chemical | Dissolved | LAB
THM | mg/L | Total Trihalomethanes | THMHAA | THM/HAA CALC | THM/HAA CALC | 0.01 | Y | Y | Y | 2 | 3/24/2023 10:24:00 PM | 5/31/2023 12:38:00 AM | 6/8/2023 12:14:57 PM | 2023-03-31 |  | EMS |  | Concentration | mg/L | Y | (fl. conc.) |  | Total Trihalomethanes (fl. conc.) |  | Disinfection by Products (DBPs); | Disinfection by Products (DBPs); EMS code: THM | Chemical |  | LAB
THM- | mg/L | Trihalomethanes | E611 | E611S: EPA 8260D (mod) | E611S: Dimethyl selenide and dimethyl diselenide are analyzed by static headspace GC-MS. Samples are prepared in headspace vials and are heated and agitated on the headspace autosampler; causing VOCs to partition between the aqueous phase and the headspace in accordance with Henry`s law. | 0.001 | Y | Y | Y | 1 | 12/14/2023 | 12/14/2023 | 1/16/2024 8:06:05 AM | 2024-01-15 | 1/15/2024 2:17:10 PM | EMS | EMS | Concentration | mg/L | Y | (fl. conc.) |  | Trihalomethanes (fl. conc.) |  | Disinfection by Products (DBPs); | Disinfection by Products (DBPs); EMS code: THM- | Chemical |  | LAB
THM- | mg/L | Trihalomethanes | F007 | Static Headspace:HRGC/LRMS | Static headspace analysis is a method for the determination of volatile compounds in liquids and/or solids by analysing the vapour phase of a sample in a closed system; in this case a sealed vial. Static Headspace is sometimes referred to as Equilibrium Headspace because the vapour phase above the sample in a sealed vial is brought into thermodynamic equilibrium with the sample prior to analysis. Volatile organic compounds are thermally extracted from the sample matrix by Static Headspace sampling and introduced directly to a thick-film megabore GC column. Compounds eluting from the GC column are introduced into a Mass Spectrometer and target compounds are identified by a comparison of both retention times and mass spectra with those of reference standards stored in a computerized data system. Reference retention times and mass spectra for analytes are obtained by the measurement of calibration standards under precisely the same extraction conditions used for samples. The concentration of each identified component is measured by relating the response of the sample quantitation mass ions to the response of the quantitation mass ions of a series of external standards. A deuterated surrogate recovery standard is added to provide information on matrix effects. |  | Y | Y | Y | 4 | 8/19/2023 | 7/20/2024 | 7/22/2024 12:17:03 PM | 2009-01-09 | 1/9/2009 5:07:40 PM | RDGIBSON | RDGIBSON | Concentration | mg/L | Y | (fl. conc.) |  | Trihalomethanes (fl. conc.) |  | Disinfection by Products (DBPs); | Disinfection by Products (DBPs); EMS code: THM- | Chemical |  | LAB
THOP | mg/L | Total Hydroxyphenols | X268 | CH2Cl2 ext:GC/MS | CH2Cl2 ext:GC/MS |  | Y | Y | Y | 14 | 12/29/2021 | 1/20/2022 | 9/2/2022 6:21:54 PM | 2022-03-21 |  | EMS |  | Concentration | mg/L | Y | (fl. conc.) |  | Total Hydroxyphenols (fl. conc.) | 120-80-9 108-46-3 123-31-9 | Organics; | Organics; EMS code: THOP | Chemical | Total | LAB
TI-D | mg/L | Titanium Dissolved | F032 | Filt.,Acid,ICPMS | There is no acid digestion required for this analysis. A water sample is collected and preserved in 0.2 percent nitric acid. To analyse the extractable metals the sample is shaken, allowed to sit for at least twenty four hours, and a decanted portion is introduced to the spectrometer. When the dissolved metals are required, the samples must first be filtered through a 0.45 micron cellulose acetate filter paper and then followed by preservation in 0.2 percent nitric acid. It is important to note that different laboratories may have very different detection limits for some variables analyzed by ICPMS. | 0.001 | Y | Y | Y | 65353 | 5/7/1998 | 8/26/2024 | 8/26/2024 4:53:05 PM | 1998-05-14 | 4/29/2011 8:51:30 AM | NPEPPIN | RDGIBSON | Concentration | mg/L | Y | (fl. conc.) |  | Titanium Dissolved (fl. conc.) | 7440-32-6 | Metals - Dissolved; | Metals - Dissolved; EMS code: TI-D | Chemical | Dissolved | LAB
TI-D | mg/L | Titanium Dissolved | METC | MET-D-NP-CCMS-VA | Ultra trace metals in water are analyzed by CRC ICPMS, based on US EPA Method 6020B (July 2014). This procedure is intended for pristine field-filtered acid-preserved water samples. The detection limits (LOR) for this test are based on lab instrumental analysis only, not including filtration. ALS-supplied field filtration equipment does not support these LOR. Therefore, because of the high probability of false positives due to filtration, it is strongly recommended that a filtration blank be analysed to aid in data interpretation. | 0.0005 | Y | Y | Y | 1179 | 4/3/2012 | 4/9/2024 | 4/29/2024 4:53:06 PM | 2021-04-14 |  | RDGIBSON |  | Concentration | mg/L | Y | (fl. conc.) |  | Titanium Dissolved (fl. conc.) | 7440-32-6 | Metals - Dissolved; | Metals - Dissolved; EMS code: TI-D | Chemical | Dissolved | LAB
TI-D | mg/L | Titanium Dissolved | X556 | Metals in water, dissolved, "in-bottle" digestion, ICP-SFMS | The whole water sample is filtered through a 0.45 micron cellulose acetate filter paper and the filtrate is stored in a HDPE/Teflon bottle. The container is de-capped in a Class 100 workstation, acidified to 2% (v/v) HNO3 and 1% (v/v) HCl and recapped. The sample is "in-bottle" digested in a convection oven, set at 60 degrees C for 16 hours. The sample is cooled and introduced to an inductively coupled argon plasma-sector field mass spectrometer (ICP-SFMS) by pneumatic nebulization. Each respective element is measured at a specific mass to charge ratio, m/z value expressed in amu (atomic mass units), and at a specified intrument resolution. LR: Low Resolution ~300, MR: Mid Resolution ~4000 and HR: High Resolution ~10000. Ref: Method number 02-2003, National Laboratory for Environmental Testing, CCIW, Environment Canada, Burlington, Requested April 2003. | 0.0001 | Y | Y | Y | 1992 | 5/25/2016 | 12/6/2023 2:04:00 AM | 2/14/2024 9:05:32 AM | 2018-02-27 |  | JKROGH |  | Concentration | mg/L | Y | (fl. conc.) |  | Titanium Dissolved (fl. conc.) | 7440-32-6 | Metals - Dissolved; | Metals - Dissolved; EMS code: TI-D | Chemical | Dissolved | LAB
TI-D | mg/L | Titanium Dissolved | X285 | HNO3 Dig:ICP Analysis | HNO3 Dig:ICP Analysis | 0.01 | Y | Y | Y | 671 | 1/30/2002 | 11/13/2002 | 3/18/2004 10:28:34 AM | 1998-05-22 | 6/3/1998 6:51:49 PM | NPEPPIN |  | Concentration | mg/L | Y | (fl. conc.) |  | Titanium Dissolved (fl. conc.) | 7440-32-6 | Metals - Dissolved; | Metals - Dissolved; EMS code: TI-D | Chemical | Dissolved | LAB
TI-D | mg/L | Titanium Dissolved | F039 | Field filt + acid, ICP-AES | Dissolved and extractable water samples are ready for analysis once collected. Total samples are autoclaved with HNO3 and hydrochloric acid (HCI) in capped Teflon vials. Analysis is performed via Inductively Coupled Argon Plasma Atomic Emission Spectrometry (ICP-AES)(simultaneous multielement analysis), Graphite Furnace Atomic Absorption Spectrometry (GF-AAS)(low level analysis) and/or Flame Atomic Absorption Spectrometry (F-AAS). | 0.002 | Y | Y | Y | 7899 | 7/29/1996 6:03:00 PM | 7/5/2019 1:21:00 PM | 9/11/2019 8:47:50 AM | 1997-05-08 | 6/3/1998 6:51:46 PM | MPAK |  | Concentration | mg/L | Y | (fl. conc.) |  | Titanium Dissolved (fl. conc.) | 7440-32-6 | Metals - Dissolved; | Metals - Dissolved; EMS code: TI-D | Chemical | Dissolved | LAB
TI-D | mg/L | Titanium Dissolved | MSEA | Metals in Seawater (undigested) by (CRC) ICP-MS | This method describes the multi-elemental determination of trace elements by Collision/Reaction Cell (CRC) ICP-MS. The method measures ions produced by a radio-frequency inductively coupled plasma. Analyte species originating in a liquid are pneumatically nebulized and the resulting aerosol transported by argon gas into the plasma torch. The ions produced are entrained in the plasma gas and extracted, by means of a differentially pumped vacuum interface, into a mass spectrometer. The ions flow through a Collision/Reaction Cell (CRC) where depending on the mode of analysis either the ions passes through the cell unhindered (no gas - classical ICP-MS mode), or collisions/reactions chemistry takes place between the ions and a neutral gas such as hydrogen or helium. Any polyatomic ions which give interference on the analytes are removed. The ions are then sorted according to their mass-to-charge ratios by a quadrupole mass spectrometer. The ions are then sorted according to their mass-to-charge ratios by a quadrupole mass spectrometer. The ions transmitted through the quadrupole are quantified by a channel electron multiplier and the ion information processed by a data handling system. Although most polyatomic interference is removed, any additional interference must be assessed and valid corrections applied or the data flagged to indicate problems. Interference correction must include compensation for background ions contributed by the plasma gas, reagents, and constituents of the sample matrix. Instrument drift as well as suppressions or enhancements of instrument response caused by the sample matrix must be corrected for by the use of internal standards. | 0.003 | Y | Y | Y | 4 | 4/12/2023 | 4/29/2024 | 5/16/2024 2:00:14 PM | 2009-09-08 |  | RDGIBSON |  | Concentration | mg/L | Y | (fl. conc.) |  | Titanium Dissolved (fl. conc.) | 7440-32-6 | Metals - Dissolved; | Metals - Dissolved; EMS code: TI-D | Chemical | Dissolved | LAB
TI-D | mg/L | Titanium Dissolved | X170 | ICP | This code is a catch-all for a large variety of ICP methods including water and also solids digested into solutions as an intermediate result. Different procedures will have greatly different detection limits. (a) For air particulates, preliminary results are reported in mg/L and loading results as mg/dm2/d. (b) For Hi Vol samples preliminiary results are same as reported in mg/L and loading results as ug/m3. (c) Note that for Rain samples there is no filtration nor digestion steps. | 0.01 | Y | Y | Y | 10008 | 4/11/1996 6:02:00 PM | 11/6/2018 | 2/25/2021 2:35:18 PM | 1996-04-12 | 8/15/2019 1:23:14 PM | EMS | KATFRASE | Concentration | mg/L | Y | (fl. conc.) |  | Titanium Dissolved (fl. conc.) | 7440-32-6 | Metals - Dissolved; | Metals - Dissolved; EMS code: TI-D | Chemical | Dissolved | LAB
TI-D | mg/L | Titanium Dissolved | F049 | Filt,HNO3 dig,conc'n,nebul'ICP-OES | This method provides an automated procedure for the Inductively Coupled Argon Plasma Optical Emission Spectrophotometric (ICP-OES) determination and quantitation of trace amounts of metals in surface, ground and precipitation waters after a manual acid digestion and concentration procedure and ultrasonic nebulisation. Organically-bound metals and metal complexes are converted to metal ions by digestion with nitric acid. To increase the signal detected by the ICP, the samples are concentrated 10 times. The liquid sample is pumped onto the piezoelectric transducer, where the ultrasonic action efficiently converts it to a fine, dense aerosol. The nebulizer carrier gas flow then transports the wet aerosol through a heated U-tube, where the solvent is vaporized. Subsequently, most of the solvent vapour is removed from the aerosol stream by passing the stream through a chilled condenser. The result is a dry, analyte-laden aerosol which is introduced to the plasma. There, the high temperature desolvates the aerosol and dissociates the sample material into individual atoms and ions, which are excited to emit light at wave lengths characteristic of the elements present. Dissolved samples are field filtered through 0.45 micron cellulose acetate filter prior to acid preservation. Samples for extractable analysis are acidified with 2 mL concentrated HNO3 per litre and allowed to stand a minimum of 24 hours before analysis. Clients providing these samples are advised to preserve samples at the collection site. | 0.01 | Y | Y | Y | 82 | 3/10/1998 2:07:00 PM | 2/25/1999 2:08:00 PM | 10/20/1999 2:21:48 AM | 1998-11-17 |  | NPEPPIN |  | Concentration | mg/L | Y | (fl. conc.) |  | Titanium Dissolved (fl. conc.) | 7440-32-6 | Metals - Dissolved; | Metals - Dissolved; EMS code: TI-D | Chemical | Dissolved | LAB
TI-D | mg/L | Titanium Dissolved | X562 | Dissolved metals: FF, FA, HNO3, HR-ICPMS | Ultra trace metals in water are analyzed by high resolution inductively coupled plasma mass spectrometry (HR-ICPMS), based on US EPA Method 200.8, (Oct. 1999). This procedure is intended for pristine field-filtered acid-preserved water samples. The detection limits (LOR) for this test are based on lab instrumental analysis only, not including filtration. ALS-supplied field filtration equipment does not support these LOR. Therefore, because of the high probability of false positives due to filtration, it is strongly recommended that a filtration blank be analyzed to aid in data interpretation. | 0.0002 | Y | Y | Y | 22 | 10/20/2015 | 10/19/2023 | 11/9/2023 7:41:27 AM | 2016-03-24 |  | RDGIBSON |  | Concentration | mg/L | Y | (fl. conc.) |  | Titanium Dissolved (fl. conc.) | 7440-32-6 | Metals - Dissolved; | Metals - Dissolved; EMS code: TI-D | Chemical | Dissolved | LAB
TI-D | mg/L | Titanium Dissolved | MWID | Metals in Water by ICPMS (digested) | Metals in Water by ICPMS (digested) Based on EPA 200.8 | 0.01 | Y | Y | Y | 31777 | 10/26/2018 5:20:00 PM | 10/26/2018 5:20:00 PM | 7/31/2024 1:46:41 PM | 2009-12-24 |  | RDGIBSON |  | Concentration | mg/L | Y | (fl. conc.) |  | Titanium Dissolved (fl. conc.) | 7440-32-6 | Metals - Dissolved; | Metals - Dissolved; EMS code: TI-D | Chemical | Dissolved | LAB
TI-D | mg/L | Titanium Dissolved | EQ01 | EQUIS (mg/L) | EQUIS (mg/L) | 0 | N | Y | Y | 3 |  |  | 10/22/1996 6:53:31 PM | 1996-04-12 | 6/3/1998 6:51:40 PM | EMS |  | Concentration | mg/L | Y | (fl. conc.) |  | Titanium Dissolved (fl. conc.) | 7440-32-6 | Metals - Dissolved; | Metals - Dissolved; EMS code: TI-D | Chemical | Dissolved | LAB
TI-E | mg/L | Titanium Extractable | F040 | Field Acid, lab filt, ICP-AES | Dissolved and extractable water samples are ready for analysis once collected. Total samples are autoclaved with HNO3 and hydrochloric acid (HCI) in capped Teflon vials. Analysis is performed via Inductively Coupled Argon Plasma Atomic Emission Spectrometry (ICP-AES)(simultaneous multielement analysis), Graphite Furnace Atomic Absorption Spectrometry (GF-AAS)(low level analysis) and/or Flame Atomic Absorption Spectrometry (F-AAS). | 0.002 | Y | Y | Y | 2357 | 8/5/1997 6:06:00 PM | 3/14/2008 | 9/11/2019 8:47:41 AM | 1997-05-08 | 6/3/1998 6:51:47 PM | MPAK |  | Concentration | mg/L | Y | (fl. conc.) |  | Titanium Extractable (fl. conc.) | 7440-32-6 | Metals - Extractable; | Metals - Extractable; EMS code: TI-E | Chemical | Extractable | LAB
TI-L | mg/L | Titanium Leachable | MEND | MEND (mod) EPA 300.1 (mod) | MEND (mod) EPA 300.1 (mod). Inorganic anions are analyzed by obtaining an extract from a soil sample that has been added in a 3:1 ratio of deionized water to soil; then leached and the extract filtered prior to analysis; which is then analyzed by Ion Chromatography with conductivity and/or UV detection. | 0.01 | Y | Y | Y | 0 |  |  |  | 2024-01-29 |  | RDGIBSON |  | Concentration | mg/L | Y | (fl. conc.) |  | Titanium Leachable (fl. conc.) | 7440-32-6 | Metals - Extractable; | Metals - Extractable; EMS code: TI-L | Chemical |  | LAB
TI-O2 | % | Titanium (as TiO2) | ICP06 | ALS Minerals ME-ICP06 | ALS Minerals ME-ICP06. A prepared sample is fused with a lithium tetraborate/lithium metaborate flux and poured into a graphite crucible and fused in a furnace at 1000°C. The resulting melt is then cooled and dissolved in 100 mL of 4% HNO3/2% HCl followed by ICP-AES. | 0.01 | Y | Y | Y | 0 |  |  |  | 2024-01-29 |  | EMS |  | DimensionlessRatio | % | N | (ratio) |  | Titanium (as TiO2) (ratio) |  | Metals; | Metals; EMS code: TI-O2 | Chemical |  | LAB
TI-T | mg/L | Titanium Total | F082 | Aliquot,HNO3/HCL dig,ICPMS | Aliquot microwave digested with 3:1 HNO3/HCL. Analysis by ICPMS. For HiVol samples preliminary result are reported in mg/L and loading results are reported in ug/m3 | 0.01 | Y | Y | Y | 23992 | 1/24/2011 | 8/23/2024 | 8/26/2024 4:53:05 PM | 2015-03-11 |  | RDGIBSON |  | Concentration | mg/L | Y | (fl. conc.) |  | Titanium Total (fl. conc.) | 7440-32-6 | Metals - Total; | Metals - Total; EMS code: TI-T | Chemical | Total | LAB
TI-T | ug/g | Titanium Total | FS82 | PBM,dry60C,10mesh,SALM-ICP-MS | PBM;dry60C;10mesh;SALM-ICP-MS | 1 | Y | Y | Y | 3538 | 5/21/2002 | 7/24/2024 | 7/29/2024 5:17:09 PM | 2002-11-12 |  | UMULLER |  | MassPerMass | mg/kg | Y | (mass conc.) |  | Titanium Total (mass conc.) | 7440-32-6 | Metals - Total; | Metals - Total; EMS code: TI-T | Chemical | Total | LAB
TI-T | % | Titanium Total | MS41 | ALS Minerals ME-MS41 | ALS Minerals ME-MS41. A prepared sample is digested with aqua regia in a graphite heating block. After cooling; the resulting solution is diluted to with deionized water; mixed and analyzed by ICP-AES. Samples are then analyzed by ICP-MS for the remaining suite of elements. | 0.005 | Y | Y | Y | 0 |  |  |  | 2024-01-29 |  | EMS |  | DimensionlessRatio | % | N | (ratio) |  | Titanium Total (ratio) | 7440-32-6 | Metals - Total; | Metals - Total; EMS code: TI-T | Chemical | Total | LAB
TI-T | ug/g | Titanium Total | EQ02 | EQUIS (ug/g) | EQUIS (ug/g) | 1 | N | Y | Y | 760 |  |  | 10/23/1996 12:08:28 AM | 1996-04-12 | 6/3/1998 6:51:40 PM | EMS |  | MassPerMass | mg/kg | Y | (mass conc.) |  | Titanium Total (mass conc.) | 7440-32-6 | Metals - Total; | Metals - Total; EMS code: TI-T | Chemical | Total | LAB
TI-T | ug/g | Titanium Total | X551 | Metals in Biota (dry), total, acid digestion, ICP-SFMS | 0.1 gram of dried, homogenized biota is weighed in a microwave TFM vessel, and 8mL of Nitric Acid and 1mL of Hydrogen Peroxide are added. The sample is digested at 200 degrees C for 15 minutes in a closed vessel high pressure microwave oven. The sample is cooled, the volume adjusted to 100mL and introduced to an iductively coupled argon plasma-sector field mass spectrometer (ICP-SFMS) by pneumatic nebulization. Each respective element is measured at a specific mass to charge ratio, m/z value expressed in amu (atomic mass units), and at a specified instrument resolution. LR: Low Resolution ~300, MR: Mid Resolution ~4000 and HR: Hich Resolution ~10000. Ref: Method number 02-2705, National Laboratory for Environmental Testing, CCIW, Environment Canada, Burlington. Requested April 2003. | 0.5 | Y | Y | Y | 9 | 12/22/2015 | 11/2/2022 11:44:00 AM | 11/10/2022 4:22:46 PM | 2016-09-29 |  | RRENTON |  | MassPerMass | mg/kg | Y | (mass conc.) |  | Titanium Total (mass conc.) | 7440-32-6 | Metals - Total; | Metals - Total; EMS code: TI-T | Chemical | Total | LAB
TI-T | ug/g wet | Titanium Total | TTIW | TI-WET-CCMS-MID-VA | Titanium in Tissue by CRC ICPMS (WET) Methodology: This method is conducted following British Columbia Lab Manual method "Metals in Animal Tissue and Vegetation (Biota) - Prescriptive". Tissue samples are homogenized and sub-sampled prior to hotblock digestion with nitric and hydrochloric acids, in combination with addition of hydrogen peroxide. Instrumental analysis is by collision cell inductively coupled plasma - mass spectrometry (modified from EPA Method 6020A). Method Limitation: This method employs a strong acid/peroxide digestion, and is intended to provide a conservative estimate of bio-available metals. Near complete recoveries are achieved for most toxicologically important metals, but elements associated with recalcitrant minerals may be only partially recovered. | 0.02 | Y | Y | Y | 24 | 11/25/2018 | 12/6/2018 | 7/21/2021 2:55:08 PM | 2018-12-24 |  | RDGIBSON |  | MassPerMass | µg/g | Y | (mass conc.) | (wet) | Titanium Total (wet) (mass conc.) | 7440-32-6 | Metals - Total; | Metals - Total; EMS code: TI-T | Chemical | Total | LAB
TI-T | % | Titanium Total | MS61 | ALS Minerals ME-MS61 | ALS Minerals ME-MS61. A prepared sample is digested with HNO3; HF; and HClO4 in a graphite heating block to near dryness. After cooling; the sample is boiled in diluted HCl. Bulk to volume with deionized water; mixed and analyzed by ICP-AES and ICP-MS for all elements. | 0.005 | Y | Y | Y | 0 |  |  |  | 2024-01-29 |  | EMS |  | DimensionlessRatio | % | N | (ratio) |  | Titanium Total (ratio) | 7440-32-6 | Metals - Total; | Metals - Total; EMS code: TI-T | Chemical | Total | LAB
TI-T | mg/L | Titanium Total | F121 | Acid,ICPMS | Acidified, no digestion, ICPMS | 0.0005 | Y | Y | Y | 4856 | 3/12/2010 7:57:00 PM | 8/15/2024 9:10:00 PM | 8/26/2024 8:06:03 AM | 2002-07-22 | 7/16/2012 3:32:05 PM | UMULLER | RDGIBSON | Concentration | mg/L | Y | (fl. conc.) |  | Titanium Total (fl. conc.) | 7440-32-6 | Metals - Total; | Metals - Total; EMS code: TI-T | Chemical | Total | LAB
TI-T | mg/L | Titanium Total | F048 | Acid dig,conc'n,nebul.,ICP-OES | This method provides an automated procedure for the Inductively Coupled Argon Plasma Optical Emission Spectrophotometric (ICP-OES) determination and quantitation of trace amounts of metals in surface, ground and precipitation waters after a manual acid digestion and concentration procedure and ultrasonic nebulisation. Organically-bound metals and metal complexes are converted to metal ions by digestion with nitric acid [method details says HCl]. To increase the signal detected by the ICP, the samples are concentrated 10 times. The liquid sample is pumped onto the piezoelectric transducer, where the ultrasonic action efficiently converts it to a fine, dense aerosol. The nebulizer carrier gas flow then transports the wet aerosol through a heated U-tube, where the solvent is vaporized. Subsequently, most of the solvent vapour is removed from the aerosol stream by passing the stream through a chilled condenser. The result is a dry, analyte-laden aerosol which is introduced to the plasma. There, the high temperature desolvates the aerosol and dissociates the sample material into individual atoms and ions, which are excited to emit light at wave lengths characteristic of the elements present. Dissolved samples are field filtered through 0.45 micron cellulose acetate filter prior to acid preservation. Samples for extractable analysis are acidified with 2 mL concentrated HNO3 per litre and allowed to stand a minimum of 24 hours before analysis. Clients providing these samples are advised to preserve samples at the collection site. | 0.01 | Y | Y | Y | 172 | 11/20/1997 9:32:00 AM | 2/25/1999 1:49:00 PM | 7/6/2001 11:10:14 AM | 1998-11-13 |  | NPEPPIN |  | Concentration | mg/L | Y | (fl. conc.) |  | Titanium Total (fl. conc.) | 7440-32-6 | Metals - Total; | Metals - Total; EMS code: TI-T | Chemical | Total | LAB
TI-T | ug/g | Titanium Total | 0190 | D&G:N/P Dig:ICP Anal. | Sample is dried and grind-mixed before nitric perchloric digestion followed by ICP analysis. | 1 | Y | Y | Y | 9 |  |  | 10/23/1996 10:17:01 PM | 1996-04-12 | 6/3/1998 6:51:40 PM | EMS |  | MassPerMass | mg/kg | Y | (mass conc.) |  | Titanium Total (mass conc.) | 7440-32-6 | Metals - Total; | Metals - Total; EMS code: TI-T | Chemical | Total | LAB
TI-T | mg/L | Titanium Total | X561 | Total metals: FA, HNO3, HR-ICPMS (No Prep)Prep) | Ultra trace metals in water are analyzed by high resolution inductively coupled plasma mass spectrometry (HR-ICPMS), modified from US EPA 200.8, (Revision 5.5). The detection limits provided can only be met for undigested samples. This procedure is intended for colorless, non-turbid, acid-preserved water samples (i.e. pristine water samples), having turbidity < 1 NTU and no odor. Where turbidity exceeds 1 NTU, and/or the sample is colored and has an odor, results may be biased low compared to true Total Metals concentrations. ALS recommends that turbidity analysis be requested on samples submitted for this test to aid with interpretation of results. | 0.0002 | Y | Y | Y | 20 | 10/20/2015 | 10/5/2023 | 1/31/2024 2:37:41 PM | 2016-03-24 |  | RDGIBSON |  | Concentration | mg/L | Y | (fl. conc.) |  | Titanium Total (fl. conc.) | 7440-32-6 | Metals - Total; | Metals - Total; EMS code: TI-T | Chemical | Total | LAB
TI-T | ug/g wet | Titanium Total | X552 | Metals in Biota (wet), total, acid digestion, ICP-SFMS | One gram of wet, homogenized biota is weighed in a microwave TFM vessel, and 8mL of Nitric Acid and 1mL of Hydrogen Peroxide are added. The sample is digested at 200 degrees C for 15 minutes in a closed vessel high pressure microwave oven. The sample is cooled, the volume adjusted to 100mL and introduced to an inductively coupled argon plasma-sector field mass spectrometer (ICP-SFMS) by pneumatic nebulization. Each respective element is measured at a specific mass to charge ratio, m/z value expressed in amu (atomic mass units), and at a specified instrument resolution. LR: Low Resolution ~300, MR: Mid Resolution ~4000. Ref: 02-2704, National Laboratory for Environmental Testing, CCIW, Environment Canada, Burlington. Requested April 2003. | 0.1 | Y | Y | Y | 14 | 2/21/2019 | 2/22/2019 | 4/28/2021 11:01:11 AM | 2021-04-15 |  | RDGIBSON |  | MassPerMass | µg/g | Y | (mass conc.) | (wet) | Titanium Total (wet) (mass conc.) | 7440-32-6 | Metals - Total; | Metals - Total; EMS code: TI-T | Chemical | Total | LAB
TI-T | mg/L | Titanium Total | MWID | Metals in Water by ICPMS (digested) | Metals in Water by ICPMS (digested) Based on EPA 200.8 | 0.01 | Y | Y | Y | 36354 |  |  | 7/31/2024 1:46:41 PM | 2009-12-24 |  | RDGIBSON |  | Concentration | mg/L | Y | (fl. conc.) |  | Titanium Total (fl. conc.) | 7440-32-6 | Metals - Total; | Metals - Total; EMS code: TI-T | Chemical | Total | LAB
TI-T | mg/L | Titanium Total | X170 | ICP | This code is a catch-all for a large variety of ICP methods including water and also solids digested into solutions as an intermediate result. Different procedures will have greatly different detection limits. (a) For air particulates, preliminary results are reported in mg/L and loading results as mg/dm2/d. (b) For Hi Vol samples preliminiary results are same as reported in mg/L and loading results as ug/m3. (c) Note that for Rain samples there is no filtration nor digestion steps. | 0.01 | Y | Y | Y | 6776 | 3/28/1996 6:03:00 PM | 11/7/2018 | 10/24/2022 10:58:25 AM | 1996-04-12 | 8/15/2019 1:23:40 PM | EMS | KATFRASE | Concentration | mg/L | Y | (fl. conc.) |  | Titanium Total (fl. conc.) | 7440-32-6 | Metals - Total; | Metals - Total; EMS code: TI-T | Chemical | Total | LAB
TI-T | ug/g | Titanium Total | MTID | MET-DRY-CCMS-VA | Metals in Tissue by CRC ICPMS (DRY) Methodology: This method is conducted following British Columbia Lab Manual method "Metals in Animal Tissue and Vegetation (Biota) - Prescriptive". Tissue samples are homogenized and sub-sampled prior to hotblock digestion with nitric and hydrochloric acids, in combination with addition of hydrogen peroxide. Instrumental analysis is by collision cell inductively coupled plasma - mass spectrometry (modified from EPA Method 6020A). Method Limitation: This method employs a strong acid/peroxide digestion, and is intended to provide a conservative estimate of bio-available metals. Near complete recoveries are achieved for most toxicologically important metals, but elements associated with recalcitrant minerals may be only partially recovered. |  | Y | Y | Y | 6 | 2/7/2022 | 2/7/2022 | 3/29/2022 5:04:16 PM | 2022-03-25 |  | EMS |  | MassPerMass | mg/kg | Y | (mass conc.) |  | Titanium Total (mass conc.) | 7440-32-6 | Metals - Total; | Metals - Total; EMS code: TI-T | Chemical | Total | LAB
TI-T | ug/g | Titanium Total | FB04 | Freeze Dry,Hom,Acid Dig,ICP | Sample moisture is determined by weighing samples before and after freeze-drying. Freeze-dried samples are ground in a Braun blender before being weighed into Teflon PFA liners and digested with concentrated HNO3, concentrated HCl, and 50% H2O2 in a microwave oven; which results in a complete sample dissolution. Samples are collected and volumized with Type I water. Analysis is performed via Inductively Coupled Argon Plasma Atomic Emission Spectrometry (ICP-AES) (simultaneous multielement analysis), Graphite Furnace Atomic Absorption Spectrometry (GF-AAS) (low level Cd and Pb analysis), Cold Vapor Atomic Fluorescence Spectrometry (CV-AFS) (Mercury analysis) and/or Flame Atomic Absorption Spectrometry (F-AAS) (K analysis). | 0.2 | Y | Y | Y | 241 | 8/22/1996 6:11:00 PM | 2/1/2002 2:30:00 PM | 6/23/2003 1:15:35 PM | 1996-06-19 | 6/3/1998 6:51:44 PM | MPAK |  | MassPerMass | mg/kg | Y | (mass conc.) |  | Titanium Total (mass conc.) | 7440-32-6 | Metals - Total; | Metals - Total; EMS code: TI-T | Chemical | Total | LAB
TI-T | mg/dm2/d | Titanium Total | DUST | HNO3 dig.(microwave or hot block),ICPMS | HNO3 dig.(microwave or hot block);ICPMS |  | Y | Y | Y | 121 | 2/4/2020 | 7/31/2024 | 7/31/2024 11:23:13 AM | 2022-09-01 |  | EMS |  | MassFlux | mg/dm²/d | Y | (mass flux) |  | Titanium Total (mass flux) | 7440-32-6 | Metals - Total; | Metals - Total; EMS code: TI-T | Chemical | Total | LAB
TI-T | ug/m3 | Titanium Total | 6040 | Aqua Regia Dig:ICP | Aqua regia digestion is used to bring the metals into solution prior to ICP analysis. For Hi-Vol samples, preliminary results are reported as mg/L and24 hr loadings as ug/m3 | 0.04 | Y | Y | Y | 21 | 5/12/2005 | 12/14/2006 | 12/30/2006 1:01:38 AM | 1998-05-05 | 6/3/1998 6:51:48 PM | NPEPPIN |  | AirConcentration | µg/m³ | Y | (air conc.) |  | Titanium Total (air conc.) | 7440-32-6 | Metals - Total; | Metals - Total; EMS code: TI-T | Chemical | Total | LAB
TI-T | ug/g | Titanium Total | X230 | Hom,HNO3/HCl,ICP | The sample is dried, ground to pass 100 mesh, and digested with a nitric-hydrochloric acid mixture to solubilize the solid matter and to remove the organic material by oxidation and volatilization. The completed digests are submitted to metals instrumentation personnel for metals analysis by Inductively Coupled Atomic Emission Spectrophotometry (ICP-AES), or hydride analysis by atomic absorption. | 50 | Y | Y | Y | 119 | 2/12/2002 | 12/30/2005 | 2/3/2006 1:13:13 AM | 1996-04-12 | 6/3/1998 6:51:40 PM | EMS |  | MassPerMass | mg/kg | Y | (mass conc.) |  | Titanium Total (mass conc.) | 7440-32-6 | Metals - Total; | Metals - Total; EMS code: TI-T | Chemical | Total | LAB
TI-T | % | Titanium Total | MS81 | ALS Minerals: ME-MS81 | ALS Minerals: ME-MS81. A prepared sample (0.100 g) is added to lithium metaborate/lithium tetraborate flux; mixed well and fused in a furnace at 1025°C. The resulting melt is then cooled and dissolved in an acid mixture containing nitric; hydrochloric and hydrofluoric acids. This solution is then analyzed by inductively coupled plasma - mass spectrometry. | 0.01 | Y | Y | Y | 0 |  |  |  | 2024-01-29 |  | EMS |  | DimensionlessRatio | % | N | (ratio) |  | Titanium Total (ratio) | 7440-32-6 | Metals - Total; | Metals - Total; EMS code: TI-T | Chemical | Total | LAB
TI-T | mg/L | Titanium Total | MSEA | Metals in Seawater (undigested) by (CRC) ICP-MS | This method describes the multi-elemental determination of trace elements by Collision/Reaction Cell (CRC) ICP-MS. The method measures ions produced by a radio-frequency inductively coupled plasma. Analyte species originating in a liquid are pneumatically nebulized and the resulting aerosol transported by argon gas into the plasma torch. The ions produced are entrained in the plasma gas and extracted, by means of a differentially pumped vacuum interface, into a mass spectrometer. The ions flow through a Collision/Reaction Cell (CRC) where depending on the mode of analysis either the ions passes through the cell unhindered (no gas - classical ICP-MS mode), or collisions/reactions chemistry takes place between the ions and a neutral gas such as hydrogen or helium. Any polyatomic ions which give interference on the analytes are removed. The ions are then sorted according to their mass-to-charge ratios by a quadrupole mass spectrometer. The ions are then sorted according to their mass-to-charge ratios by a quadrupole mass spectrometer. The ions transmitted through the quadrupole are quantified by a channel electron multiplier and the ion information processed by a data handling system. Although most polyatomic interference is removed, any additional interference must be assessed and valid corrections applied or the data flagged to indicate problems. Interference correction must include compensation for background ions contributed by the plasma gas, reagents, and constituents of the sample matrix. Instrument drift as well as suppressions or enhancements of instrument response caused by the sample matrix must be corrected for by the use of internal standards. | 0.003 | Y | Y | Y | 432 | 8/12/2009 12:08:00 AM | 4/29/2024 | 6/7/2024 10:03:16 AM | 2009-09-08 |  | RDGIBSON |  | Concentration | mg/L | Y | (fl. conc.) |  | Titanium Total (fl. conc.) | 7440-32-6 | Metals - Total; | Metals - Total; EMS code: TI-T | Chemical | Total | LAB
TI-T | % | Titanium Total | 6041 | Aqua Regia by ICPMS | These were sediment/soils digested with Aqua-Regia (Nitric Acid & Hydrochloric Acids) and then analysis by ICPMS. Results would be reported in mg/kg and/or %. | 0.001 | Y | Y | Y | 92 | 10/13/2019 | 7/27/2024 | 7/29/2024 9:13:08 AM | 2019-09-06 |  | KATFRASE |  | DimensionlessRatio | % | N | (ratio) |  | Titanium Total (ratio) | 7440-32-6 | Metals - Total; | Metals - Total; EMS code: TI-T | Chemical | Total | LAB
TI-T | ug/g | Titanium Total | PB02 | PBM,HNO3/HCl dig,dry60C,100mesh, | Performance Based Method. Samples are dried at 60 C and sieved using a .149 mm (100 mesh) sieve. Solid samples are digested with the combination of nitric and hydrochloric acids with a 1:1 ratio of acids. For environmental purposes, strong acid leaches are used routinely since they have the advantage of speed, low cost, and superio reproducibility. All metals trapped within the refractory matrix are unlikely to have environmental consequences. Instrumental analysis of extracts produced using this digestion method can be performed by a variety of techniques. |  | Y | Y | Y | 10 | 9/13/2001 | 9/13/2001 | 11/6/2002 3:52:40 PM | 2001-07-18 |  | NPEPPIN |  | MassPerMass | mg/kg | Y | (mass conc.) |  | Titanium Total (mass conc.) | 7440-32-6 | Metals - Total; | Metals - Total; EMS code: TI-T | Chemical | Total | LAB
TI-T | ug/m3 | Titanium Total | 5312 | Tef:N/P Dig:ICP Analysis | A nitric/perchloric acid digestion if used to bring the metals into solution prior to ICP analysis. For Hi-Vol samples preliminary results are reported as mg/L and 24 hour loading results as ug/m3. | 0.04 | Y | Y | Y | 1191 | 11/7/1997 | 5/15/2000 | 11/18/2000 1:29:40 AM | 1996-10-13 | 6/3/1998 6:51:46 PM | EMS |  | AirConcentration | µg/m³ | Y | (air conc.) |  | Titanium Total (air conc.) | 7440-32-6 | Metals - Total; | Metals - Total; EMS code: TI-T | Chemical | Total | LAB
TI-T | ug/g | Titanium Total | ICPM | Metals in Soil by CRC ICPMS | This analysis is carried out using procedures from CSR Analytical Method: "Strong Acid Leachable Metals (SALM) in Soil", BC Ministry of Environment, 26 June 2009, and procedures adapted from EPA Method 200.2. The sample is manually homogenized, dried at 60 degrees Celsius, sieved through a 63 um (230 mesh) sieve, and a representative subsample of the dry material is weighed. The sample is then digested at 95 degrees Celsius for 2 hours by block digester using concentrated nitric and hydrochloric acids. Instrumental analysis is by collision cell inductively coupled plasma - mass spectrometry (modifed from EPA Method 6020A). Method Limitation: This method is not a total digestion technique. It is a very strong acid digestion that is intended to dissolve those metals that may be environmentally available. By design, elements bound in silicate structures are not normally dissolved by this procedure as they are not usually mobile in the environment. | 1 | Y | Y | Y | 103 | 9/19/2017 | 9/24/2023 | 1/12/2024 11:16:52 AM | 2018-01-09 |  | JKROGH |  | MassPerMass | mg/kg | Y | (mass conc.) |  | Titanium Total (mass conc.) | 7440-32-6 | Metals - Total; | Metals - Total; EMS code: TI-T | Chemical | Total | LAB
TI-T | mg/L | Titanium Total | X285 | HNO3 Dig:ICP Analysis | HNO3 Dig:ICP Analysis | 0.01 | Y | Y | Y | 716 | 8/10/1993 | 10/5/1993 | 4/30/2004 11:32:39 AM | 1998-05-22 | 6/3/1998 6:51:49 PM | NPEPPIN |  | Concentration | mg/L | Y | (fl. conc.) |  | Titanium Total (fl. conc.) | 7440-32-6 | Metals - Total; | Metals - Total; EMS code: TI-T | Chemical | Total | LAB
TI-T | mg/L | Titanium Total | F033 | Acid Dig,Centrifuge,ICPMS | For the determination of total recoverable elements in water or waste water, take 100 ml aliquot from a well mixed, acid preserved, sample containing not more than 0.25% (w/v) total solids and transfer to a 250 ml Griffin beaker. (If total solids are greater than 0.25% reduce the size of the aliquot by a proportionate amount). Add 1 ml of conc. nitric acid and 0.5 ml conc. hydrochloric acid. Heat on a hot plate until the volume has been reduced to 15-20ml, ensuring that the sample does not boil. Allow to cool and transfer to a 50 ml volumetric flask. Dilute to volume with ASTM type I water. Centrifuge the sample or allow to stand overnight to separate insoluble material. Prior analysis, pipette 20 ml into a 50 ml volumetric flask and dilute to volume with ASTM type I water. Analysis should be performed within seventy two hours of the completed preparation. Note that different laboratories may have very different field preservation, digestion procedures and detection limits for this method, so these should be documented. | 0.001 | Y | Y | Y | 26019 | 5/7/1998 | 8/22/2024 3:24:00 AM | 8/25/2024 9:47:36 PM | 1998-05-14 | 6/3/1998 6:51:49 PM | NPEPPIN |  | Concentration | mg/L | Y | (fl. conc.) |  | Titanium Total (fl. conc.) | 7440-32-6 | Metals - Total; | Metals - Total; EMS code: TI-T | Chemical | Total | LAB
TI-T | ug/g wet | Titanium Total | X560 | Metals in Biota (wet), total, acid digestion, ICP-MS | Five grams of wet, homogenized biota is weighed in a centrifuge tube, and 2.5 mL of Nitric Acid and 2.5 mL of Hydrochloric acid are added. The sample is digested at room temperature for one hour and then digested at 90-95 degrees C in a water bath or heat block for one more hour. The sample is cooled and the volume is made to 50 mL with DI. The digestate is then diluted at ten fold and introduced to a collision/reaction cell inductively coupled argon plasma mass spectrometer (CRC ICPMS) by pneumatic nebulization. Each respective element is measured at a specific mass to charge ratio, m/z value expressed in amu (atomic mass units). - Maxxam Analytics (BC). Added to EMS Dec 10, 2008. | 1 | Y | Y | Y | 432 | 11/15/2008 4:00:00 AM | 2/7/2022 | 3/29/2022 5:04:16 PM | 2008-12-11 | 3/11/2015 4:17:10 PM | RDGIBSON | EMS | MassPerMass | µg/g | Y | (mass conc.) | (wet) | Titanium Total (wet) (mass conc.) | 7440-32-6 | Metals - Total; | Metals - Total; EMS code: TI-T | Chemical | Total | LAB
TI-T | mg/L | Titanium Total | 0040 | HNO3 Dig: ICP Analysis | HNO3 Dig: ICP Analysis | 0.003 | Y | Y | Y | 1736 |  |  | 3/23/2010 5:08:09 PM | 1996-04-12 | 6/3/1998 6:51:41 PM | EMS |  | Concentration | mg/L | Y | (fl. conc.) |  | Titanium Total (fl. conc.) | 7440-32-6 | Metals - Total; | Metals - Total; EMS code: TI-T | Chemical | Total | LAB
TI-T | ug/g wet | Titanium Total | BPSY | MET-WET-BPSY-CCMS-VA | Metals in Animal Tissue and Vegetation (Biota) - Prescriptive. EPA 200.3/6020B This method is conducted following British Columbia Lab Manual method "Metals in Animal Tissue and Vegetation (Biota) - Prescriptive". Tissue samples are digested with nitric and hydrochloric acids; in combination with addition of hydrogen peroxide. Instrumental analysis is by collision cell inductively coupled plasma - mass spectrometry (modified from EPA Method 6020B). | 0.1 | Y | Y | Y | 2 | 11/11/2019 | 11/11/2019 | 6/10/2021 1:57:10 PM | 2021-06-04 |  | EMS |  | MassPerMass | µg/g | Y | (mass conc.) | (wet) | Titanium Total (wet) (mass conc.) | 7440-32-6 | Metals - Total; | Metals - Total; EMS code: TI-T | Chemical | Total | LAB
TI-T | mg/L | Titanium Total | METN | MET-T-NP-CCMS-VA | Ultra trace metals in water are analyzed by CRC ICPMS, based on US EPA Method 6020B (July 2014). The detection limits provided can only be met for undigested samples. This procedure is intended for colorless, non-turbid, acid-preserved water samples (i.e. pristine water samples), having turbidity < 1 NTU and no odor. Where turbidity exceeds 1 NTU, and/or the sample is colored and has an odor, results may be biased low compared to true Total Metals concentrations. ALS recommends that turbidity analysis be requested on samples submitted for this test to aid with interpretation of results. Where turbidity is <1NTU, undigested metals are equivalent to total metals concentrations. | 0.0005 | Y | Y | Y | 1059 | 4/3/2012 | 4/4/2024 | 4/29/2024 4:53:06 PM | 2021-04-14 |  | RDGIBSON |  | Concentration | mg/L | Y | (fl. conc.) |  | Titanium Total (fl. conc.) | 7440-32-6 | Metals - Total; | Metals - Total; EMS code: TI-T | Chemical | Total | LAB
TI-T | ug/g | Titanium Total | FS89 | Homo,HNO3 dig,ICP-AES | Homo;HNO3 dig;ICP-AES | 0.3 | Y | Y | Y | 4 | 1/5/2001 | 8/20/2001 12:00:00 PM | 10/23/2002 1:42:52 PM | 2001-03-08 |  | NPEPPIN |  | MassPerMass | mg/kg | Y | (mass conc.) |  | Titanium Total (mass conc.) | 7440-32-6 | Metals - Total; | Metals - Total; EMS code: TI-T | Chemical | Total | LAB
TI-T | mg/L | Titanium Total | 1349 | Tot Metals in Seawater by CRC ICPMS (BC) | Tot Metals in Seawater by CRC ICPMS (BC) | 0.005 | Y | Y | Y | 98 | 8/16/2019 | 12/5/2021 | 12/13/2021 12:34:17 PM | 2019-09-30 |  | KATFRASE |  | Concentration | mg/L | Y | (fl. conc.) |  | Titanium Total (fl. conc.) | 7440-32-6 | Metals - Total; | Metals - Total; EMS code: TI-T | Chemical | Total | LAB
TI-T | mg/L | Titanium Total | F038 | HNO3/HCl dig,ICP-AES | Dissolved and extractable water samples are ready for analysis once collected. Total samples are autoclaved with HNO3 and hydrochloric acid (HCI) in capped Teflon vials. Analysis is performed via Inductively Coupled Argon Plasma Atomic Emission Spectrometry (ICP-AES)(simultaneous multielement analysis), Graphite Furnace Atomic Absorption Spectrometry (GF-AAS) (low level analysis) and/or Flame Atomic Absorption Spectrometry (F-AAS). | 0.002 | Y | Y | Y | 7041 | 8/1/1996 6:04:00 PM | 12/1/2010 | 9/11/2019 8:47:43 AM | 1997-05-08 | 6/3/1998 6:51:46 PM | MPAK |  | Concentration | mg/L | Y | (fl. conc.) |  | Titanium Total (fl. conc.) | 7440-32-6 | Metals - Total; | Metals - Total; EMS code: TI-T | Chemical | Total | LAB
TI-T | mg/L | Titanium Total | 0042 | GI HNO3 Dig: ICP | Sample aliquot taken from unfiltered unpreserved General Ions bottle; HNO3 digestion; ICP analysis. | 0.003 | Y | Y | Y | 9795 | 9/20/1996 | 7/5/2019 3:21:00 PM | 7/12/2019 6:49:13 PM | 1996-04-12 | 6/3/1998 6:51:40 PM | EMS |  | Concentration | mg/L | Y | (fl. conc.) |  | Titanium Total (fl. conc.) | 7440-32-6 | Metals - Total; | Metals - Total; EMS code: TI-T | Chemical | Total | LAB
TI-T | ug/g | Titanium Total | FS01 | D&G,Nitric/HCl:ICP | Samples are dried at 60 C, then either sieved to -100mesh (150 um) and rolled to homogenize, or finely ground with a mortar and pestle or a shatter box laboratory mill and rolled (depending on the submitter's requirements). The samples are then weighed into Teflon digestion liners and digested with HNO3 and HCl in a microwave oven. Samples are cooled, volumized, and allowed to settle overnight before being decanted. Analysis is then performed via Inductively Coupled Argon Plasma Atomic Emission Spectrometry (ICP-AES) (Simultaneous multi element analysis). In ICP Spectrometry, the liquid sample is pumped to the nebulizer, where it is convertd into an aerosol. the nebulizer carrier gas flow than transports the wet aerosol past the impact bead, where the large droplets condense out and flow to the waste container. The remaining fine aerosol is carried into the plasma. There, the high temperature desolvates the aerosol and dissociates the sample material into individual atoms and ions, which are excited to emit light at wavelengths characteristic of the elements present. A diffraction grating in the spectrometer diffracts the various wavelengths of light through exit slits to the photo-multiplier tubes, which convert the light energy to electronic signals. | 0.3 | Y | Y | Y | 2642 | 8/30/1996 1:52:00 PM | 10/10/2007 | 5/24/2008 1:15:03 AM | 1996-06-19 | 6/3/1998 6:51:44 PM | MPAK |  | MassPerMass | mg/kg | Y | (mass conc.) |  | Titanium Total (mass conc.) | 7440-32-6 | Metals - Total; | Metals - Total; EMS code: TI-T | Chemical | Total | LAB
TI-T | ug/g | Titanium Total | 0192 | Hom, Nitric/Perchlor,ICP | Sample already in a dry state so does not require drying before homogenization and nitric perchlorid digestion followed by ICP analysis. | 1 | Y | Y | Y | 236 |  |  | 10/24/1996 11:23:42 AM | 1996-04-12 | 6/3/1998 6:51:41 PM | EMS |  | MassPerMass | mg/kg | Y | (mass conc.) |  | Titanium Total (mass conc.) | 7440-32-6 | Metals - Total; | Metals - Total; EMS code: TI-T | Chemical | Total | LAB
TI-T | mg/L | Titanium Total | X349 | FA:HNO3:ICP | Samples are field-acidified with HNO3; HNO3 digestion; analysis by ICP/AES. | 0.002 | Y | Y | Y | 30 |  |  | 10/30/2014 8:31:18 AM | 1998-05-22 | 6/3/1998 6:51:49 PM | NPEPPIN |  | Concentration | mg/L | Y | (fl. conc.) |  | Titanium Total (fl. conc.) | 7440-32-6 | Metals - Total; | Metals - Total; EMS code: TI-T | Chemical | Total | LAB
TIME | day | Time/Duration | DBPFP | DBPFP Incubation | DBPFP Incubation | 0 | Y | Y | Y | 2 | 3/18/2023 4:37:00 PM | 5/23/2023 12:08:00 PM | 6/8/2023 12:14:57 PM | 2023-03-31 |  | EMS |  | Time | day | Y | (time) |  | Time/Duration (time) |  | Physical Properties; | Physical Properties; EMS code: TIME | Chemical |  | LAB
TL-D | mg/L | Thallium Dissolved | METB | MET-D-NP-BCMOE-MS-VA | Ultra trace metals in water are analyzed by CRC ICPMS, based on US EPA Method 6020B (July 2014). This procedure is intended for pristine field-filtered acid-preserved water samples. The detection limits (LOR) for this test are based on lab instrumental analysis only, not including filtration. ALS-supplied field filtration equipment does not support these LOR. Therefore, because of the high probability of false positives due to filtration, it is strongly recommended that a filtration blank be analysed to aid in data interpretation. BC Ministry of Environment version. | 2e-06 | Y | Y | Y | 596 | 6/11/2019 | 8/15/2024 | 8/23/2024 9:50:35 AM | 2019-08-20 |  | KATFRASE |  | Concentration | mg/L | Y | (fl. conc.) |  | Thallium Dissolved (fl. conc.) | 7440-28-0 | Metals - Dissolved; | Metals - Dissolved; EMS code: TL-D | Chemical | Dissolved | LAB
TL-D | mg/L | Thallium Dissolved | X170 | ICP | This code is a catch-all for a large variety of ICP methods including water and also solids digested into solutions as an intermediate result. Different procedures will have greatly different detection limits. (a) For air particulates, preliminary results are reported in mg/L and loading results as mg/dm2/d. (b) For Hi Vol samples preliminiary results are same as reported in mg/L and loading results as ug/m3. (c) Note that for Rain samples there is no filtration nor digestion steps. | 0.003 | Y | Y | Y | 6289 | 12/3/1996 | 11/6/2018 | 4/12/2024 1:44:29 PM | 1996-04-12 | 6/3/1998 6:51:41 PM | EMS |  | Concentration | mg/L | Y | (fl. conc.) |  | Thallium Dissolved (fl. conc.) | 7440-28-0 | Metals - Dissolved; | Metals - Dissolved; EMS code: TL-D | Chemical | Dissolved | LAB
TL-D | mg/L | Thallium Dissolved | X562 | Dissolved metals: FF, FA, HNO3, HR-ICPMS | Ultra trace metals in water are analyzed by high resolution inductively coupled plasma mass spectrometry (HR-ICPMS), based on US EPA Method 200.8, (Oct. 1999). This procedure is intended for pristine field-filtered acid-preserved water samples. The detection limits (LOR) for this test are based on lab instrumental analysis only, not including filtration. ALS-supplied field filtration equipment does not support these LOR. Therefore, because of the high probability of false positives due to filtration, it is strongly recommended that a filtration blank be analyzed to aid in data interpretation. | 2e-06 | Y | Y | Y | 1264 | 6/8/2015 | 10/19/2023 | 11/9/2023 7:41:27 AM | 2015-10-22 |  | RDGIBSON |  | Concentration | mg/L | Y | (fl. conc.) |  | Thallium Dissolved (fl. conc.) | 7440-28-0 | Metals - Dissolved; | Metals - Dissolved; EMS code: TL-D | Chemical | Dissolved | LAB
TL-D | t/d | Thallium Dissolved | CALC | Calculation (t/d) | Calculation (t/d) | 0.01 | Y | Y | Y | 1604 |  |  | 7/3/2024 12:59:25 PM | 1996-11-13 | 6/3/1998 6:51:46 PM | MPAK |  | MassFlow | t/d | Y | (mass flow) |  | Thallium Dissolved (mass flow) | 7440-28-0 | Metals - Dissolved; | Metals - Dissolved; EMS code: TL-D | Chemical | Dissolved | LAB
TL-D | mg/L | Thallium Dissolved | X544 | FF,FA,HNO3,ICPMS | Analytes nebulized followed by ion entrained in plasma gas and extracted by means of a differentially pumped vacuum interface into a mass spectrometer | 0.0001 | Y | Y | Y | 30 |  |  | 10/30/2014 8:31:18 AM | 2001-09-18 |  | NPEPPIN |  | Concentration | mg/L | Y | (fl. conc.) |  | Thallium Dissolved (fl. conc.) | 7440-28-0 | Metals - Dissolved; | Metals - Dissolved; EMS code: TL-D | Chemical | Dissolved | LAB
TL-D | mg/L | Thallium Dissolved | METC | MET-D-NP-CCMS-VA | Ultra trace metals in water are analyzed by CRC ICPMS, based on US EPA Method 6020B (July 2014). This procedure is intended for pristine field-filtered acid-preserved water samples. The detection limits (LOR) for this test are based on lab instrumental analysis only, not including filtration. ALS-supplied field filtration equipment does not support these LOR. Therefore, because of the high probability of false positives due to filtration, it is strongly recommended that a filtration blank be analysed to aid in data interpretation. | 2e-06 | Y | Y | Y | 1179 | 4/3/2012 | 4/9/2024 | 4/29/2024 4:53:06 PM | 2021-04-14 |  | RDGIBSON |  | Concentration | mg/L | Y | (fl. conc.) |  | Thallium Dissolved (fl. conc.) | 7440-28-0 | Metals - Dissolved; | Metals - Dissolved; EMS code: TL-D | Chemical | Dissolved | LAB
TL-D | mg/L | Thallium Dissolved | F032 | Filt.,Acid,ICPMS | There is no acid digestion required for this analysis. A water sample is collected and preserved in 0.2 percent nitric acid. To analyse the extractable metals the sample is shaken, allowed to sit for at least twenty four hours, and a decanted portion is introduced to the spectrometer. When the dissolved metals are required, the samples must first be filtered through a 0.45 micron cellulose acetate filter paper and then followed by preservation in 0.2 percent nitric acid. It is important to note that different laboratories may have very different detection limits for some variables analyzed by ICPMS. | 0.0002 | Y | Y | Y | 74138 | 1/7/1998 9:46:00 AM | 8/26/2024 | 8/26/2024 4:53:05 PM | 1997-10-20 | 8/15/2019 11:00:36 AM | NPEPPIN | KATFRASE | Concentration | mg/L | Y | (fl. conc.) |  | Thallium Dissolved (fl. conc.) | 7440-28-0 | Metals - Dissolved; | Metals - Dissolved; EMS code: TL-D | Chemical | Dissolved | LAB
TL-D | mg/L | Thallium Dissolved | F068 | HNO3dig,TeflonTube/Autoclave,ICP-MS | This method is applicable to determiniation of dissolved, extractable, and total recoverable metals in ground waters, surface waters, drinking waters, and wastewaters. Dissolved, extractable, and total recoverable sample fractions are acidified to pH <1.5 with analytical grade nitric acid (HNO3) to prevent the metals from adsorbing to the container material or precipitating. Extractable and total recoverable water samples are collected in 250 ml acid washed polyethylene bottles and are preserved with Nitric acid (0.5 ml acid for 100 mLs of sample). All preserved samples are stored at room temperature and can be held for up to 6 months. Unfiltered, un-preserved samples are stored at 4C for no more than 30 days. For dissolved and extractable water fractions, no sample preparation is necessary. Samples are simply poured in autosampler tubes (i.e. 15 ml centrifuge tubes), and placed on the autosampler and analyzed. Samples for total recoverable metals are digested with Nitric acid in special Teflon tubes in an autoclave prior to analyses. | 2e-06 | Y | Y | Y | 1398 | 6/1/1999 1:22:00 PM | 4/17/2015 1:06:00 PM | 8/11/2021 12:04:11 PM | 1999-06-28 |  | NPEPPIN |  | Concentration | mg/L | Y | (fl. conc.) |  | Thallium Dissolved (fl. conc.) | 7440-28-0 | Metals - Dissolved; | Metals - Dissolved; EMS code: TL-D | Chemical | Dissolved | LAB
TL-D | mg/L | Thallium Dissolved | X285 | HNO3 Dig:ICP Analysis | HNO3 Dig:ICP Analysis | 0.05 | Y | Y | Y | 608 | 8/21/2002 | 11/13/2002 | 3/18/2004 10:28:34 AM | 1998-05-22 | 6/3/1998 6:51:49 PM | NPEPPIN |  | Concentration | mg/L | Y | (fl. conc.) |  | Thallium Dissolved (fl. conc.) | 7440-28-0 | Metals - Dissolved; | Metals - Dissolved; EMS code: TL-D | Chemical | Dissolved | LAB
TL-D | mg/L | Thallium Dissolved | MWID | Metals in Water by ICPMS (digested) | Metals in Water by ICPMS (digested) Based on EPA 200.8 | 5e-05 | Y | Y | Y | 31767 | 10/26/2018 5:20:00 PM | 10/26/2018 5:20:00 PM | 7/31/2024 1:46:41 PM | 2009-12-24 |  | RDGIBSON |  | Concentration | mg/L | Y | (fl. conc.) |  | Thallium Dissolved (fl. conc.) | 7440-28-0 | Metals - Dissolved; | Metals - Dissolved; EMS code: TL-D | Chemical | Dissolved | LAB
TL-D | mg/L | Thallium Dissolved | MSEA | Metals in Seawater (undigested) by (CRC) ICP-MS | This method describes the multi-elemental determination of trace elements by Collision/Reaction Cell (CRC) ICP-MS. The method measures ions produced by a radio-frequency inductively coupled plasma. Analyte species originating in a liquid are pneumatically nebulized and the resulting aerosol transported by argon gas into the plasma torch. The ions produced are entrained in the plasma gas and extracted, by means of a differentially pumped vacuum interface, into a mass spectrometer. The ions flow through a Collision/Reaction Cell (CRC) where depending on the mode of analysis either the ions passes through the cell unhindered (no gas - classical ICP-MS mode), or collisions/reactions chemistry takes place between the ions and a neutral gas such as hydrogen or helium. Any polyatomic ions which give interference on the analytes are removed. The ions are then sorted according to their mass-to-charge ratios by a quadrupole mass spectrometer. The ions are then sorted according to their mass-to-charge ratios by a quadrupole mass spectrometer. The ions transmitted through the quadrupole are quantified by a channel electron multiplier and the ion information processed by a data handling system. Although most polyatomic interference is removed, any additional interference must be assessed and valid corrections applied or the data flagged to indicate problems. Interference correction must include compensation for background ions contributed by the plasma gas, reagents, and constituents of the sample matrix. Instrument drift as well as suppressions or enhancements of instrument response caused by the sample matrix must be corrected for by the use of internal standards. | 1e-05 | Y | Y | Y | 4 | 4/12/2023 | 4/29/2024 | 5/16/2024 2:00:14 PM | 2009-09-08 |  | RDGIBSON |  | Concentration | mg/L | Y | (fl. conc.) |  | Thallium Dissolved (fl. conc.) | 7440-28-0 | Metals - Dissolved; | Metals - Dissolved; EMS code: TL-D | Chemical | Dissolved | LAB
TL-D | mg/L | Thallium Dissolved | F039 | Field filt + acid, ICP-AES | Dissolved and extractable water samples are ready for analysis once collected. Total samples are autoclaved with HNO3 and hydrochloric acid (HCI) in capped Teflon vials. Analysis is performed via Inductively Coupled Argon Plasma Atomic Emission Spectrometry (ICP-AES)(simultaneous multielement analysis), Graphite Furnace Atomic Absorption Spectrometry (GF-AAS)(low level analysis) and/or Flame Atomic Absorption Spectrometry (F-AAS). | 0.05 | Y | Y | Y | 868 | 1/18/2002 | 10/17/2007 | 5/31/2008 1:15:47 AM | 1998-05-21 | 6/3/1998 6:51:49 PM | NPEPPIN |  | Concentration | mg/L | Y | (fl. conc.) |  | Thallium Dissolved (fl. conc.) | 7440-28-0 | Metals - Dissolved; | Metals - Dissolved; EMS code: TL-D | Chemical | Dissolved | LAB
TL-D | mg/L | Thallium Dissolved | X556 | Metals in water, dissolved, "in-bottle" digestion, ICP-SFMS | The whole water sample is filtered through a 0.45 micron cellulose acetate filter paper and the filtrate is stored in a HDPE/Teflon bottle. The container is de-capped in a Class 100 workstation, acidified to 2% (v/v) HNO3 and 1% (v/v) HCl and recapped. The sample is "in-bottle" digested in a convection oven, set at 60 degrees C for 16 hours. The sample is cooled and introduced to an inductively coupled argon plasma-sector field mass spectrometer (ICP-SFMS) by pneumatic nebulization. Each respective element is measured at a specific mass to charge ratio, m/z value expressed in amu (atomic mass units), and at a specified intrument resolution. LR: Low Resolution ~300, MR: Mid Resolution ~4000 and HR: High Resolution ~10000. Ref: Method number 02-2003, National Laboratory for Environmental Testing, CCIW, Environment Canada, Burlington, Requested April 2003. | 1e-06 | Y | Y | Y | 4234 | 6/11/2003 | 12/6/2023 2:04:00 AM | 2/14/2024 9:05:32 AM | 2004-01-19 |  | UMULLER |  | Concentration | mg/L | Y | (fl. conc.) |  | Thallium Dissolved (fl. conc.) | 7440-28-0 | Metals - Dissolved; | Metals - Dissolved; EMS code: TL-D | Chemical | Dissolved | LAB
TL-E | mg/L | Thallium Extractable | X553 | Metals in water, extractable, ICP-SFMS | The whole water sample is collected in a HDPE/Teflon bottle. In a Class 100 workstation it is preserved with concentrated Nitric Acid at concentration of 0.2 percent (v/v). The acid preservative extracts the metals from material in the sample and a decanted portion is introduced directly to an inductively coupled argon plasma-sector field mass spectrometer (ICP-SFMS) by pneumatic nebulization. Each respective element is measured at a specific mass to charge ratio, m/z value expressed in amu (atomic mass units), and at a specified instrument resolution. LR: Low Resolution ~300, MR: Mid Resolution ~4000 and HR: High Resolution ~10000. Ref: Method number 02-2061, National Laboratory for Environmental Testing, CCIW, Environment Canada, Burlington. Requested April 2003. | 0.1 | Y | Y | Y | 2596 | 5/22/2003 | 9/21/2007 | 11/28/2023 2:11:53 PM | 2004-11-29 |  | UMULLER |  | Concentration | mg/L | Y | (fl. conc.) |  | Thallium Extractable (fl. conc.) | 7440-28-0 | Metals - Extractable; | Metals - Extractable; EMS code: TL-E | Chemical | Extractable | LAB
TL-E | mg/L | Thallium Extractable | F069 | SEM extr,6N HCl,ICP/OES | 6.0 N Hydrochloric acid (HCI) is added to an aliquot of the sediment sample. The extract produced from the addition of the acid is then analyzed for simultaneously extracted metals (SEM) using inductively coupled plasma - optical emission spectrophotometry. The detection limits depend on the metal parameters in question and the instrumentation used. | 0.2 | Y | Y | Y | 1 |  |  | 11/25/2021 1:50:15 PM | 1999-07-09 |  | NPEPPIN |  | Concentration | mg/L | Y | (fl. conc.) |  | Thallium Extractable (fl. conc.) | 7440-28-0 | Metals - Extractable; | Metals - Extractable; EMS code: TL-E | Chemical | Extractable | LAB
TL-E | mg/L | Thallium Extractable | F047 | 24hr HNO3 dil'n,ICP-MS | There is no acid digestion required for this analysis. A water sample is collected and preserved in 0.2 percent nitric acid. To analyse the extractable metals the sample is shaken, allowed to sit for at least twenty four hours, and a decanted portion is introduced to the spectrometer. When the dissolved metals are required, the samples must first be filtered through a 0.45 micron cellulose acetate filter paper and then followed by preservation in 0.2 percent nitric acid. It is important to note that different laboratories may have very different detection limits for some variables analyzed by ICP/MS. | 0.0005 | Y | Y | Y | 2676 | 4/16/1998 10:21:00 AM | 6/26/2003 | 9/11/2019 8:46:18 AM | 1997-10-20 | 11/29/2004 10:28:28 AM | NPEPPIN | UMULLER | Concentration | mg/L | Y | (fl. conc.) |  | Thallium Extractable (fl. conc.) | 7440-28-0 | Metals - Extractable; | Metals - Extractable; EMS code: TL-E | Chemical | Extractable | LAB
TL-T | ug/g wet | Thallium Total | BPSY | MET-WET-BPSY-CCMS-VA | Metals in Animal Tissue and Vegetation (Biota) - Prescriptive. EPA 200.3/6020B This method is conducted following British Columbia Lab Manual method "Metals in Animal Tissue and Vegetation (Biota) - Prescriptive". Tissue samples are digested with nitric and hydrochloric acids; in combination with addition of hydrogen peroxide. Instrumental analysis is by collision cell inductively coupled plasma - mass spectrometry (modified from EPA Method 6020B). | 0.0004 | Y | Y | Y | 2 | 11/7/2019 | 11/7/2019 | 6/10/2021 1:57:10 PM | 2021-06-04 |  | EMS |  | MassPerMass | µg/g | Y | (mass conc.) | (wet) | Thallium Total (wet) (mass conc.) | 7440-28-0 | Metals - Total; | Metals - Total; EMS code: TL-T | Chemical | Total | LAB
TL-T | mg/L | Thallium Total | F068 | HNO3dig,TeflonTube/Autoclave,ICP-MS | This method is applicable to determiniation of dissolved, extractable, and total recoverable metals in ground waters, surface waters, drinking waters, and wastewaters. Dissolved, extractable, and total recoverable sample fractions are acidified to pH <1.5 with analytical grade nitric acid (HNO3) to prevent the metals from adsorbing to the container material or precipitating. Extractable and total recoverable water samples are collected in 250 ml acid washed polyethylene bottles and are preserved with Nitric acid (0.5 ml acid for 100 mLs of sample). All preserved samples are stored at room temperature and can be held for up to 6 months. Unfiltered, un-preserved samples are stored at 4C for no more than 30 days. For dissolved and extractable water fractions, no sample preparation is necessary. Samples are simply poured in autosampler tubes (i.e. 15 ml centrifuge tubes), and placed on the autosampler and analyzed. Samples for total recoverable metals are digested with Nitric acid in special Teflon tubes in an autoclave prior to analyses. | 2e-06 | Y | Y | Y | 1779 | 6/1/1999 1:33:00 PM | 7/4/2008 9:48:00 AM | 7/30/2024 6:32:12 AM | 1999-06-28 |  | NPEPPIN |  | Concentration | mg/L | Y | (fl. conc.) |  | Thallium Total (fl. conc.) | 7440-28-0 | Metals - Total; | Metals - Total; EMS code: TL-T | Chemical | Total | LAB
TL-T | mg/L | Thallium Total | X170 | ICP | This code is a catch-all for a large variety of ICP methods including water and also solids digested into solutions as an intermediate result. Different procedures will have greatly different detection limits. (a) For air particulates, preliminary results are reported in mg/L and loading results as mg/dm2/d. (b) For Hi Vol samples preliminiary results are same as reported in mg/L and loading results as ug/m3. (c) Note that for Rain samples there is no filtration nor digestion steps. | 0.02 | Y | Y | Y | 3249 | 1/25/2001 | 11/7/2018 | 7/3/2024 12:59:24 PM | 1996-04-12 | 6/3/1998 6:51:40 PM | EMS |  | Concentration | mg/L | Y | (fl. conc.) |  | Thallium Total (fl. conc.) | 7440-28-0 | Metals - Total; | Metals - Total; EMS code: TL-T | Chemical | Total | LAB
TL-T | mg/dm2/d | Thallium Total | CAL1 | Calculation (mg/dm2/d) | Calculation (mg/dm2/d) | 0.0001 | Y | Y | Y | 7 |  |  | 7/3/2024 12:59:21 PM | 2005-04-18 | 4/19/2005 9:27:15 AM | UMULLER | UMULLER | MassFlux | mg/dm²/d | Y | (mass flux) |  | Thallium Total (mass flux) | 7440-28-0 | Metals - Total; | Metals - Total; EMS code: TL-T | Chemical | Total | LAB
TL-T | ug/g wet | Thallium Total | X552 | Metals in Biota (wet), total, acid digestion, ICP-SFMS | One gram of wet, homogenized biota is weighed in a microwave TFM vessel, and 8mL of Nitric Acid and 1mL of Hydrogen Peroxide are added. The sample is digested at 200 degrees C for 15 minutes in a closed vessel high pressure microwave oven. The sample is cooled, the volume adjusted to 100mL and introduced to an inductively coupled argon plasma-sector field mass spectrometer (ICP-SFMS) by pneumatic nebulization. Each respective element is measured at a specific mass to charge ratio, m/z value expressed in amu (atomic mass units), and at a specified instrument resolution. LR: Low Resolution ~300, MR: Mid Resolution ~4000. Ref: 02-2704, National Laboratory for Environmental Testing, CCIW, Environment Canada, Burlington. Requested April 2003. | 0.0001 | Y | Y | Y | 37 | 8/9/2016 | 10/16/2023 | 11/28/2023 2:21:33 PM | 2004-01-19 | 4/19/2021 1:22:00 PM | UMULLER | EMS | MassPerMass | µg/g | Y | (mass conc.) | (wet) | Thallium Total (wet) (mass conc.) | 7440-28-0 | Metals - Total; | Metals - Total; EMS code: TL-T | Chemical | Total | LAB
TL-T | ug/m3 | Thallium Total | 5097 | N/P Dig(Filt+Wash),ICP | N/P Dig(Filt+Wash);ICP | 0.001 | Y | Y | Y | 48 |  |  | 7/3/2024 12:59:21 PM | 2024-04-10 |  | RDGIBSON |  | AirConcentration | µg/m³ | Y | (air conc.) |  | Thallium Total (air conc.) | 7440-28-0 | Metals - Total; | Metals - Total; EMS code: TL-T | Chemical | Total | LAB
TL-T | mg/L | Thallium Total | X561 | Total metals: FA, HNO3, HR-ICPMS (No Prep)Prep) | Ultra trace metals in water are analyzed by high resolution inductively coupled plasma mass spectrometry (HR-ICPMS), modified from US EPA 200.8, (Revision 5.5). The detection limits provided can only be met for undigested samples. This procedure is intended for colorless, non-turbid, acid-preserved water samples (i.e. pristine water samples), having turbidity < 1 NTU and no odor. Where turbidity exceeds 1 NTU, and/or the sample is colored and has an odor, results may be biased low compared to true Total Metals concentrations. ALS recommends that turbidity analysis be requested on samples submitted for this test to aid with interpretation of results. | 2e-06 | Y | Y | Y | 2756 | 6/8/2015 | 10/5/2023 | 1/31/2024 2:37:41 PM | 2015-10-22 |  | RDGIBSON |  | Concentration | mg/L | Y | (fl. conc.) |  | Thallium Total (fl. conc.) | 7440-28-0 | Metals - Total; | Metals - Total; EMS code: TL-T | Chemical | Total | LAB
TL-T | ppm (S) | Thallium Total | MS81 | ALS Minerals: ME-MS81 | ALS Minerals: ME-MS81. A prepared sample (0.100 g) is added to lithium metaborate/lithium tetraborate flux; mixed well and fused in a furnace at 1025°C. The resulting melt is then cooled and dissolved in an acid mixture containing nitric; hydrochloric and hydrofluoric acids. This solution is then analyzed by inductively coupled plasma - mass spectrometry. | 0.5 | Y | Y | Y | 0 |  |  |  | 2024-01-29 |  | EMS |  | DimensionlessRatio | ppm | N | (ratio) |  | Thallium Total (ratio) | 7440-28-0 | Metals - Total; | Metals - Total; EMS code: TL-T | Chemical | Total | LAB
TL-T | mg/L | Thallium Total | X545 | FA,HNO3,ICPMS | Analytes nebulized followed by ion entrained in plasma gas and extracted by means of a differentially pumped vacuum interface into a mass spectrometer | 0.0001 | Y | Y | Y | 30 |  |  | 10/30/2014 8:31:18 AM | 2001-09-18 | 9/18/2001 1:36:23 PM | NPEPPIN | NPEPPIN | Concentration | mg/L | Y | (fl. conc.) |  | Thallium Total (fl. conc.) | 7440-28-0 | Metals - Total; | Metals - Total; EMS code: TL-T | Chemical | Total | LAB
TL-T | ug/g | Thallium Total | MTID | MET-DRY-CCMS-VA | Metals in Tissue by CRC ICPMS (DRY) Methodology: This method is conducted following British Columbia Lab Manual method "Metals in Animal Tissue and Vegetation (Biota) - Prescriptive". Tissue samples are homogenized and sub-sampled prior to hotblock digestion with nitric and hydrochloric acids, in combination with addition of hydrogen peroxide. Instrumental analysis is by collision cell inductively coupled plasma - mass spectrometry (modified from EPA Method 6020A). Method Limitation: This method employs a strong acid/peroxide digestion, and is intended to provide a conservative estimate of bio-available metals. Near complete recoveries are achieved for most toxicologically important metals, but elements associated with recalcitrant minerals may be only partially recovered. | 0.002 | Y | Y | Y | 122 | 8/13/2014 | 9/8/2022 | 10/26/2022 7:14:15 AM | 2015-03-09 | 3/11/2015 4:12:05 PM | RDGIBSON | EMS | MassPerMass | mg/kg | Y | (mass conc.) |  | Thallium Total (mass conc.) | 7440-28-0 | Metals - Total; | Metals - Total; EMS code: TL-T | Chemical | Total | LAB
TL-T | mg/L | Thallium Total | TCLP | Toxicity Characteristic Leaching Procedure | The TCLP procedure is designed to simulate the climatic leaching action | 0.1 | Y | Y | Y | 22 | 12/8/2009 7:49:00 PM | 9/28/2011 5:53:00 PM | 11/3/2011 11:10:21 PM | 2009-03-06 |  | JCHUDAK |  | Concentration | mg/L | Y | (fl. conc.) |  | Thallium Total (fl. conc.) | 7440-28-0 | Metals - Total; | Metals - Total; EMS code: TL-T | Chemical | Total | LAB
TL-T | ug/g | Thallium Total | PB02 | PBM,HNO3/HCl dig,dry60C,100mesh, | Performance Based Method. Samples are dried at 60 C and sieved using a .149 mm (100 mesh) sieve. Solid samples are digested with the combination of nitric and hydrochloric acids with a 1:1 ratio of acids. For environmental purposes, strong acid leaches are used routinely since they have the advantage of speed, low cost, and superio reproducibility. All metals trapped within the refractory matrix are unlikely to have environmental consequences. Instrumental analysis of extracts produced using this digestion method can be performed by a variety of techniques. |  | Y | Y | Y | 52 | 5/24/2001 3:42:00 PM | 11/26/2001 10:23:00 AM | 5/26/2003 2:28:42 PM | 2001-07-18 |  | NPEPPIN |  | MassPerMass | mg/kg | Y | (mass conc.) |  | Thallium Total (mass conc.) | 7440-28-0 | Metals - Total; | Metals - Total; EMS code: TL-T | Chemical | Total | LAB
TL-T | ug/g wet | Thallium Total | MTIW | MET-WET-CCMS-MID-VA | Metals in Tissue by CRC ICPMS (WET) Methodology: This method is conducted following British Columbia Lab Manual method "Metals in Animal Tissue and Vegetation (Biota) - Prescriptive". Tissue samples are homogenized and sub-sampled prior to hotblock digestion with nitric and hydrochloric acids, in combination with addition of hydrogen peroxide. Instrumental analysis is by collision cell inductively coupled plasma - mass spectrometry (modified from EPA Method 6020A). Method Limitation: This method employs a strong acid/peroxide digestion, and is intended to provide a conservative estimate of bio-available metals. Near complete recoveries are achieved for most toxicologically important metals, but elements associated with recalcitrant minerals may be only partially recovered. | 0.0004 | Y | Y | Y | 24 | 11/25/2018 | 12/6/2018 | 7/21/2021 2:55:08 PM | 2018-12-24 |  | EMS |  | MassPerMass | µg/g | Y | (mass conc.) | (wet) | Thallium Total (wet) (mass conc.) | 7440-28-0 | Metals - Total; | Metals - Total; EMS code: TL-T | Chemical | Total | LAB
TL-T | mg/L | Thallium Total | F082 | Aliquot,HNO3/HCL dig,ICPMS | Aliquot microwave digested with 3:1 HNO3/HCL. Analysis by ICPMS. For HiVol samples preliminary result are reported in mg/L and loading results are reported in ug/m3 | 0.0002 | Y | Y | Y | 26958 | 1/24/2011 | 8/23/2024 | 8/26/2024 4:53:05 PM | 2000-06-13 | 8/15/2019 11:12:28 AM | NPEPPIN | KATFRASE | Concentration | mg/L | Y | (fl. conc.) |  | Thallium Total (fl. conc.) | 7440-28-0 | Metals - Total; | Metals - Total; EMS code: TL-T | Chemical | Total | LAB
TL-T | ug/g | Thallium Total | 0190 | D&G:N/P Dig:ICP Anal. | Sample is dried and grind-mixed before nitric perchloric digestion followed by ICP analysis. | 20 | Y | Y | Y | 9 |  |  | 10/23/1996 10:17:01 PM | 1996-04-12 | 6/3/1998 6:51:40 PM | EMS |  | MassPerMass | mg/kg | Y | (mass conc.) |  | Thallium Total (mass conc.) | 7440-28-0 | Metals - Total; | Metals - Total; EMS code: TL-T | Chemical | Total | LAB
TL-T | mg/L | Thallium Total | X555 | Metals in water, total, "in-bottle" digestion, ICP-SFMS | The whole water sample collected in a HDPE/Teflon bottle, is de-capped in a Class 100 workstation, acidified to 2% (v/v) HNO3 and 1% (v/v) HCl and recapped. The sample is "in-bottle" digested in a convection oven, set at 60 degrees C for 16 hours. The sample is cooled and introduced to an inductively coupled argon plasma-sector field mass spectrometer (ICP-SFMS) by pneumatic nebulization. Each respective element is measured at a specific mass to charge ratio, m/z value expressed in amu (atomic mass units), and at resolution. LR: Low Resolution ~300, MR: Mid Resolution ~4000 and HR: High a specified instrument Resolution ~10000. Ref: Method number 02-2003, National Laboratory for Environmental Testing, CCIW, Environment Canada, Burlington. Requested April 2003. | 0.0005 | Y | Y | Y | 17342 | 5/20/2003 | 12/6/2023 2:08:00 AM | 2/14/2024 9:05:32 AM | 2004-11-29 |  | UMULLER |  | Concentration | mg/L | Y | (fl. conc.) |  | Thallium Total (fl. conc.) | 7440-28-0 | Metals - Total; | Metals - Total; EMS code: TL-T | Chemical | Total | LAB
TL-T | mg/L | Thallium Total | METT | MET-T-NP-BCMOE-MS-VA | Ultra trace metals in water are analyzed by CRC ICPMS, based on US EPA Method 6020B (July 2014). The detection limits provided can only be met for undigested samples. This procedure is intended for colorless, non-turbid, acid-preserved water samples (i.e. pristine water samples), having turbidity < 1 NTU and no odor. Where turbidity exceeds 1 NTU, and/or the sample is colored and has an odor, results may be biased low compared to true Total Metals concentrations. ALS recommends that turbidity analysis be requested on samples submitted for this test to aid with interpretation of results. Where turbidity is <1NTU, undigested metals are equivalent to total metals concentrations. BC Ministry of Environment version. | 2e-06 | Y | Y | Y | 2589 | 6/5/2019 | 8/18/2024 | 8/26/2024 1:03:06 PM | 2019-08-20 |  | KATFRASE |  | Concentration | mg/L | Y | (fl. conc.) |  | Thallium Total (fl. conc.) | 7440-28-0 | Metals - Total; | Metals - Total; EMS code: TL-T | Chemical | Total | LAB
TL-T | mg/L | Thallium Total | F121 | Acid,ICPMS | Acidified, no digestion, ICPMS | 2e-06 | Y | Y | Y | 8410 | 3/12/2010 7:57:00 PM | 8/15/2024 9:10:00 PM | 8/26/2024 8:06:03 AM | 2002-07-22 | 7/16/2012 3:31:45 PM | UMULLER | RDGIBSON | Concentration | mg/L | Y | (fl. conc.) |  | Thallium Total (fl. conc.) | 7440-28-0 | Metals - Total; | Metals - Total; EMS code: TL-T | Chemical | Total | LAB
TL-T | ppm (S) | Thallium Total | MS61 | ALS Minerals ME-MS61 | ALS Minerals ME-MS61. A prepared sample is digested with HNO3; HF; and HClO4 in a graphite heating block to near dryness. After cooling; the sample is boiled in diluted HCl. Bulk to volume with deionized water; mixed and analyzed by ICP-AES and ICP-MS for all elements. | 0.02 | Y | Y | Y | 0 |  |  |  | 2024-01-29 |  | EMS |  | DimensionlessRatio | ppm | N | (ratio) |  | Thallium Total (ratio) | 7440-28-0 | Metals - Total; | Metals - Total; EMS code: TL-T | Chemical | Total | LAB
TL-T | mg/L | Thallium Total | 5038 | N/P Dig, ICP | Aliquot of prepared sample is digested first with nitric acid then with perchloric acid. Analysis by ICP. (a) For air particulates, preliminary results are reported in mg/L and loading results as mg/dm2/d. (b) For Hi Vol samples preliminary results are same as reported in mg/L and loading results as ug/m3. | 0.03 | Y | Y | Y | 6 | 6/25/2010 3:15:00 PM | 12/1/2010 | 12/9/2010 3:55:27 PM | 1998-02-17 | 6/3/1998 6:51:48 PM | NPEPPIN |  | Concentration | mg/L | Y | (fl. conc.) |  | Thallium Total (fl. conc.) | 7440-28-0 | Metals - Total; | Metals - Total; EMS code: TL-T | Chemical | Total | LAB
TL-T | mg/L | Thallium Total | F038 | HNO3/HCl dig,ICP-AES | Dissolved and extractable water samples are ready for analysis once collected. Total samples are autoclaved with HNO3 and hydrochloric acid (HCI) in capped Teflon vials. Analysis is performed via Inductively Coupled Argon Plasma Atomic Emission Spectrometry (ICP-AES)(simultaneous multielement analysis), Graphite Furnace Atomic Absorption Spectrometry (GF-AAS) (low level analysis) and/or Flame Atomic Absorption Spectrometry (F-AAS). | 0.05 | Y | Y | Y | 1 | 3/22/2002 | 3/22/2002 | 2/13/2004 1:01:35 AM | 1998-05-21 | 6/3/1998 6:51:49 PM | NPEPPIN |  | Concentration | mg/L | Y | (fl. conc.) |  | Thallium Total (fl. conc.) | 7440-28-0 | Metals - Total; | Metals - Total; EMS code: TL-T | Chemical | Total | LAB
TL-T | ug/g | Thallium Total | X503 | not dry,homo,Aqua Regia dig,GFAAS,Calc | Not dry;homogenization;Aqua Regia digestion;GFAAS; calculation | 0.5 | Y | Y | Y | 5 | 12/30/2005 | 12/30/2005 | 2/3/2006 1:13:13 AM | 1998-12-09 |  | NPEPPIN |  | MassPerMass | mg/kg | Y | (mass conc.) |  | Thallium Total (mass conc.) | 7440-28-0 | Metals - Total; | Metals - Total; EMS code: TL-T | Chemical | Total | LAB
TL-T | mg/L | Thallium Total | 1349 | Tot Metals in Seawater by CRC ICPMS (BC) | Tot Metals in Seawater by CRC ICPMS (BC) | 5e-05 | Y | Y | Y | 98 | 8/16/2019 | 12/5/2021 | 12/13/2021 12:34:17 PM | 2019-09-30 |  | KATFRASE |  | Concentration | mg/L | Y | (fl. conc.) |  | Thallium Total (fl. conc.) | 7440-28-0 | Metals - Total; | Metals - Total; EMS code: TL-T | Chemical | Total | LAB
TL-T | mg/L | Thallium Total | MSEA | Metals in Seawater (undigested) by (CRC) ICP-MS | This method describes the multi-elemental determination of trace elements by Collision/Reaction Cell (CRC) ICP-MS. The method measures ions produced by a radio-frequency inductively coupled plasma. Analyte species originating in a liquid are pneumatically nebulized and the resulting aerosol transported by argon gas into the plasma torch. The ions produced are entrained in the plasma gas and extracted, by means of a differentially pumped vacuum interface, into a mass spectrometer. The ions flow through a Collision/Reaction Cell (CRC) where depending on the mode of analysis either the ions passes through the cell unhindered (no gas - classical ICP-MS mode), or collisions/reactions chemistry takes place between the ions and a neutral gas such as hydrogen or helium. Any polyatomic ions which give interference on the analytes are removed. The ions are then sorted according to their mass-to-charge ratios by a quadrupole mass spectrometer. The ions are then sorted according to their mass-to-charge ratios by a quadrupole mass spectrometer. The ions transmitted through the quadrupole are quantified by a channel electron multiplier and the ion information processed by a data handling system. Although most polyatomic interference is removed, any additional interference must be assessed and valid corrections applied or the data flagged to indicate problems. Interference correction must include compensation for background ions contributed by the plasma gas, reagents, and constituents of the sample matrix. Instrument drift as well as suppressions or enhancements of instrument response caused by the sample matrix must be corrected for by the use of internal standards. | 1e-05 | Y | Y | Y | 441 | 8/12/2009 12:08:00 AM | 4/29/2024 | 6/7/2024 10:03:16 AM | 2009-09-08 |  | RDGIBSON |  | Concentration | mg/L | Y | (fl. conc.) |  | Thallium Total (fl. conc.) | 7440-28-0 | Metals - Total; | Metals - Total; EMS code: TL-T | Chemical | Total | LAB
TL-T | ug/m3 | Thallium Total | F088 | Aliquot,HNO3/HCl dig, ICPMS,loading | Aliquot microwave digested with 3:1 HNO3/HCL. Analysis by ICPMS. For HiVol samples preliminary result are reported in mg/L and loading results are reported in ug/m3 |  | Y | Y | Y | 565 | 8/17/2000 10:46:00 AM | 12/14/2006 | 12/30/2006 1:01:38 AM | 2000-08-16 |  | NPEPPIN |  | AirConcentration | µg/m³ | Y | (air conc.) |  | Thallium Total (air conc.) | 7440-28-0 | Metals - Total; | Metals - Total; EMS code: TL-T | Chemical | Total | LAB
TL-T | mg/dm2/d | Thallium Total | DUST | HNO3 dig.(microwave or hot block),ICPMS | HNO3 dig.(microwave or hot block);ICPMS |  | Y | Y | Y | 267 | 7/19/2016 1:11:00 PM | 7/31/2024 | 7/31/2024 11:23:13 AM | 2008-02-07 |  | RDGIBSON |  | MassFlux | mg/dm²/d | Y | (mass flux) |  | Thallium Total (mass flux) | 7440-28-0 | Metals - Total; | Metals - Total; EMS code: TL-T | Chemical | Total | LAB
TL-T | mg/L | Thallium Total | X285 | HNO3 Dig:ICP Analysis | HNO3 Dig:ICP Analysis | 0.05 | Y | Y | Y | 609 | 8/10/1993 | 10/5/1993 | 4/30/2004 1:54:39 PM | 1998-05-22 | 6/3/1998 6:51:49 PM | NPEPPIN |  | Concentration | mg/L | Y | (fl. conc.) |  | Thallium Total (fl. conc.) | 7440-28-0 | Metals - Total; | Metals - Total; EMS code: TL-T | Chemical | Total | LAB
TL-T | ppm (S) | Thallium Total | 6042 | Aqua Regia by ICPMS | These were sediment/soils digested with Aqua-Regia (Nitric Acid & Hydrochloric Acids) and then analysis by ICPMS. Results would be reported in ppm s. | 0.002 | Y | Y | Y | 71 | 5/14/2022 | 7/27/2024 | 7/29/2024 9:13:08 AM | 2022-06-28 |  | EMS |  | DimensionlessRatio | ppm | N | (ratio) |  | Thallium Total (ratio) | 7440-28-0 | Metals - Total; | Metals - Total; EMS code: TL-T | Chemical | Total | LAB
TL-T | mg/L | Thallium Total | 0042 | GI HNO3 Dig: ICP | Sample aliquot taken from unfiltered unpreserved General Ions bottle; HNO3 digestion; ICP analysis. | 0.03 | Y | Y | Y | 6140 | 9/20/1996 | 10/17/2007 | 2/12/2009 2:12:23 PM | 1996-04-12 | 6/3/1998 6:51:40 PM | EMS |  | Concentration | mg/L | Y | (fl. conc.) |  | Thallium Total (fl. conc.) | 7440-28-0 | Metals - Total; | Metals - Total; EMS code: TL-T | Chemical | Total | LAB
TL-T | ug/m3 | Thallium Total | 6040 | Aqua Regia Dig:ICP | Aqua regia digestion is used to bring the metals into solution prior to ICP analysis. For Hi-Vol samples, preliminary results are reported as mg/L and24 hr loadings as ug/m3 | 0.02 | Y | Y | Y | 437 | 2/13/2002 | 2/9/2005 | 2/23/2005 1:13:48 AM | 1998-05-05 | 6/3/1998 6:51:48 PM | NPEPPIN |  | AirConcentration | µg/m³ | Y | (air conc.) |  | Thallium Total (air conc.) | 7440-28-0 | Metals - Total; | Metals - Total; EMS code: TL-T | Chemical | Total | LAB
TL-T | ug/g | Thallium Total | 0192 | Hom, Nitric/Perchlor,ICP | Sample already in a dry state so does not require drying before homogenization and nitric perchlorid digestion followed by ICP analysis. | 20 | Y | Y | Y | 272 | 2/13/2002 | 5/29/2002 | 1/14/2004 1:00:44 AM | 1996-04-12 | 6/3/1998 6:51:40 PM | EMS |  | MassPerMass | mg/kg | Y | (mass conc.) |  | Thallium Total (mass conc.) | 7440-28-0 | Metals - Total; | Metals - Total; EMS code: TL-T | Chemical | Total | LAB
TL-T | ug/g | Thallium Total | EQ02 | EQUIS (ug/g) | EQUIS (ug/g) | 20 | N | Y | Y | 743 |  |  | 10/23/1996 12:08:28 AM | 1996-04-12 | 6/3/1998 6:51:41 PM | EMS |  | MassPerMass | mg/kg | Y | (mass conc.) |  | Thallium Total (mass conc.) | 7440-28-0 | Metals - Total; | Metals - Total; EMS code: TL-T | Chemical | Total | LAB
TL-T | ug/g wet | Thallium Total | X560 | Metals in Biota (wet), total, acid digestion, ICP-MS | Five grams of wet, homogenized biota is weighed in a centrifuge tube, and 2.5 mL of Nitric Acid and 2.5 mL of Hydrochloric acid are added. The sample is digested at room temperature for one hour and then digested at 90-95 degrees C in a water bath or heat block for one more hour. The sample is cooled and the volume is made to 50 mL with DI. The digestate is then diluted at ten fold and introduced to a collision/reaction cell inductively coupled argon plasma mass spectrometer (CRC ICPMS) by pneumatic nebulization. Each respective element is measured at a specific mass to charge ratio, m/z value expressed in amu (atomic mass units). - Maxxam Analytics (BC). Added to EMS Dec 10, 2008. | 0.05 | Y | Y | Y | 191 | 11/15/2008 4:00:00 AM | 10/30/2023 | 11/28/2023 2:21:33 PM | 2008-12-11 | 3/11/2015 4:17:10 PM | RDGIBSON | EMS | MassPerMass | µg/g | Y | (mass conc.) | (wet) | Thallium Total (wet) (mass conc.) | 7440-28-0 | Metals - Total; | Metals - Total; EMS code: TL-T | Chemical | Total | LAB
TL-T | ug/m3 | Thallium Total | 5312 | Tef:N/P Dig:ICP Analysis | A nitric/perchloric acid digestion if used to bring the metals into solution prior to ICP analysis. For Hi-Vol samples preliminary results are reported as mg/L and 24 hour loading results as ug/m3. | 0.02 | Y | Y | Y | 1192 | 11/7/1997 | 5/15/2000 | 11/18/2000 1:08:02 AM | 1996-10-13 | 6/3/1998 6:51:46 PM | EMS |  | AirConcentration | µg/m³ | Y | (air conc.) |  | Thallium Total (air conc.) | 7440-28-0 | Metals - Total; | Metals - Total; EMS code: TL-T | Chemical | Total | LAB
TL-T | ug/g | Thallium Total | ICPM | Metals in Soil by CRC ICPMS | This analysis is carried out using procedures from CSR Analytical Method: "Strong Acid Leachable Metals (SALM) in Soil", BC Ministry of Environment, 26 June 2009, and procedures adapted from EPA Method 200.2. The sample is manually homogenized, dried at 60 degrees Celsius, sieved through a 63 um (230 mesh) sieve, and a representative subsample of the dry material is weighed. The sample is then digested at 95 degrees Celsius for 2 hours by block digester using concentrated nitric and hydrochloric acids. Instrumental analysis is by collision cell inductively coupled plasma - mass spectrometry (modifed from EPA Method 6020A). Method Limitation: This method is not a total digestion technique. It is a very strong acid digestion that is intended to dissolve those metals that may be environmentally available. By design, elements bound in silicate structures are not normally dissolved by this procedure as they are not usually mobile in the environment. | 0.05 | Y | Y | Y | 103 | 9/19/2017 | 9/24/2023 | 1/12/2024 11:16:52 AM | 2018-01-09 |  | JKROGH |  | MassPerMass | mg/kg | Y | (mass conc.) |  | Thallium Total (mass conc.) | 7440-28-0 | Metals - Total; | Metals - Total; EMS code: TL-T | Chemical | Total | LAB
TL-T | ug/g | Thallium Total | FS82 | PBM,dry60C,10mesh,SALM-ICP-MS | PBM;dry60C;10mesh;SALM-ICP-MS | 0.05 | Y | Y | Y | 3526 | 1/31/2002 | 7/24/2024 | 7/29/2024 5:17:09 PM | 2002-11-12 |  | UMULLER |  | MassPerMass | mg/kg | Y | (mass conc.) |  | Thallium Total (mass conc.) | 7440-28-0 | Metals - Total; | Metals - Total; EMS code: TL-T | Chemical | Total | LAB
TL-T | mg/L | Thallium Total | MWID | Metals in Water by ICPMS (digested) | Metals in Water by ICPMS (digested) Based on EPA 200.8 | 5e-05 | Y | Y | Y | 36811 |  |  | 7/31/2024 1:46:41 PM | 2009-12-24 |  | RDGIBSON |  | Concentration | mg/L | Y | (fl. conc.) |  | Thallium Total (fl. conc.) | 7440-28-0 | Metals - Total; | Metals - Total; EMS code: TL-T | Chemical | Total | LAB
TL-T | t/d | Thallium Total | CALC | Calculation (t/d) | Calculation (t/d) | 0.01 | Y | Y | Y | 2615 |  |  | 7/3/2024 12:59:25 PM | 1996-11-13 | 6/3/1998 6:51:46 PM | MPAK |  | MassFlow | t/d | Y | (mass flow) |  | Thallium Total (mass flow) | 7440-28-0 | Metals - Total; | Metals - Total; EMS code: TL-T | Chemical | Total | LAB
TL-T | ppm (S) | Thallium Total | 6041 | Aqua Regia by ICPMS | These were sediment/soils digested with Aqua-Regia (Nitric Acid & Hydrochloric Acids) and then analysis by ICPMS. Results would be reported in mg/kg and/or %. | 0.001 | Y | Y | Y | 21 | 10/13/2019 | 8/5/2020 | 8/10/2020 3:52:07 PM | 2019-09-06 | 10/21/2019 11:56:46 AM | KATFRASE | KATFRASE | DimensionlessRatio | ppm | N | (ratio) |  | Thallium Total (ratio) | 7440-28-0 | Metals - Total; | Metals - Total; EMS code: TL-T | Chemical | Total | LAB
TL-T | ug/g | Thallium Total | X551 | Metals in Biota (dry), total, acid digestion, ICP-SFMS | 0.1 gram of dried, homogenized biota is weighed in a microwave TFM vessel, and 8mL of Nitric Acid and 1mL of Hydrogen Peroxide are added. The sample is digested at 200 degrees C for 15 minutes in a closed vessel high pressure microwave oven. The sample is cooled, the volume adjusted to 100mL and introduced to an iductively coupled argon plasma-sector field mass spectrometer (ICP-SFMS) by pneumatic nebulization. Each respective element is measured at a specific mass to charge ratio, m/z value expressed in amu (atomic mass units), and at a specified instrument resolution. LR: Low Resolution ~300, MR: Mid Resolution ~4000 and HR: Hich Resolution ~10000. Ref: Method number 02-2705, National Laboratory for Environmental Testing, CCIW, Environment Canada, Burlington. Requested April 2003. | 0.001 | Y | Y | Y | 46 | 8/14/2015 | 10/30/2023 | 11/28/2023 2:21:33 PM | 2004-01-19 |  | UMULLER |  | MassPerMass | mg/kg | Y | (mass conc.) |  | Thallium Total (mass conc.) | 7440-28-0 | Metals - Total; | Metals - Total; EMS code: TL-T | Chemical | Total | LAB
TL-T | mg/L | Thallium Total | F033 | Acid Dig,Centrifuge,ICPMS | For the determination of total recoverable elements in water or waste water, take 100 ml aliquot from a well mixed, acid preserved, sample containing not more than 0.25% (w/v) total solids and transfer to a 250 ml Griffin beaker. (If total solids are greater than 0.25% reduce the size of the aliquot by a proportionate amount). Add 1 ml of conc. nitric acid and 0.5 ml conc. hydrochloric acid. Heat on a hot plate until the volume has been reduced to 15-20ml, ensuring that the sample does not boil. Allow to cool and transfer to a 50 ml volumetric flask. Dilute to volume with ASTM type I water. Centrifuge the sample or allow to stand overnight to separate insoluble material. Prior analysis, pipette 20 ml into a 50 ml volumetric flask and dilute to volume with ASTM type I water. Analysis should be performed within seventy two hours of the completed preparation. Note that different laboratories may have very different field preservation, digestion procedures and detection limits for this method, so these should be documented. | 0.0001 | Y | Y | Y | 36829 | 11/20/1997 9:32:00 AM | 8/22/2024 3:24:00 AM | 8/25/2024 9:47:36 PM | 1998-05-14 | 6/3/1998 6:51:49 PM | NPEPPIN |  | Concentration | mg/L | Y | (fl. conc.) |  | Thallium Total (fl. conc.) | 7440-28-0 | Metals - Total; | Metals - Total; EMS code: TL-T | Chemical | Total | LAB
TL-T | ug/g | Thallium Total | X546 | Dig,FA,HNO3,ICPMS | Digestion, the analytes in the resulting digests are nebulized followed by ion entrained in plasma gas and extracted by means of a differentially pumped vacuum interface into a mass spectrometer | 0.2 | Y | Y | Y | 448 | 6/20/2005 | 10/10/2007 | 5/24/2008 1:15:03 AM | 2001-09-18 |  | NPEPPIN |  | MassPerMass | mg/kg | Y | (mass conc.) |  | Thallium Total (mass conc.) | 7440-28-0 | Metals - Total; | Metals - Total; EMS code: TL-T | Chemical | Total | LAB
TL-T | mg/L | Thallium Total | METN | MET-T-NP-CCMS-VA | Ultra trace metals in water are analyzed by CRC ICPMS, based on US EPA Method 6020B (July 2014). The detection limits provided can only be met for undigested samples. This procedure is intended for colorless, non-turbid, acid-preserved water samples (i.e. pristine water samples), having turbidity < 1 NTU and no odor. Where turbidity exceeds 1 NTU, and/or the sample is colored and has an odor, results may be biased low compared to true Total Metals concentrations. ALS recommends that turbidity analysis be requested on samples submitted for this test to aid with interpretation of results. Where turbidity is <1NTU, undigested metals are equivalent to total metals concentrations. | 2e-06 | Y | Y | Y | 1059 | 4/3/2012 | 4/4/2024 | 4/29/2024 4:53:06 PM | 2021-04-14 |  | RDGIBSON |  | Concentration | mg/L | Y | (fl. conc.) |  | Thallium Total (fl. conc.) | 7440-28-0 | Metals - Total; | Metals - Total; EMS code: TL-T | Chemical | Total | LAB
TL-T | mg/L | Thallium Total | X349 | FA:HNO3:ICP | Samples are field-acidified with HNO3; HNO3 digestion; analysis by ICP/AES. | 0.05 | Y | Y | Y | 182 |  |  | 9/9/1999 10:25:05 AM | 1998-05-22 | 6/3/1998 6:51:49 PM | NPEPPIN |  | Concentration | mg/L | Y | (fl. conc.) |  | Thallium Total (fl. conc.) | 7440-28-0 | Metals - Total; | Metals - Total; EMS code: TL-T | Chemical | Total | LAB
TL-T | mg/L | Thallium Total | F048 | Acid dig,conc'n,nebul.,ICP-OES | This method provides an automated procedure for the Inductively Coupled Argon Plasma Optical Emission Spectrophotometric (ICP-OES) determination and quantitation of trace amounts of metals in surface, ground and precipitation waters after a manual acid digestion and concentration procedure and ultrasonic nebulisation. Organically-bound metals and metal complexes are converted to metal ions by digestion with nitric acid [method details says HCl]. To increase the signal detected by the ICP, the samples are concentrated 10 times. The liquid sample is pumped onto the piezoelectric transducer, where the ultrasonic action efficiently converts it to a fine, dense aerosol. The nebulizer carrier gas flow then transports the wet aerosol through a heated U-tube, where the solvent is vaporized. Subsequently, most of the solvent vapour is removed from the aerosol stream by passing the stream through a chilled condenser. The result is a dry, analyte-laden aerosol which is introduced to the plasma. There, the high temperature desolvates the aerosol and dissociates the sample material into individual atoms and ions, which are excited to emit light at wave lengths characteristic of the elements present. Dissolved samples are field filtered through 0.45 micron cellulose acetate filter prior to acid preservation. Samples for extractable analysis are acidified with 2 mL concentrated HNO3 per litre and allowed to stand a minimum of 24 hours before analysis. Clients providing these samples are advised to preserve samples at the collection site. | 0.2 | Y | Y | Y | 29 |  |  | 7/6/2001 11:10:15 AM | 1998-11-13 |  | NPEPPIN |  | Concentration | mg/L | Y | (fl. conc.) |  | Thallium Total (fl. conc.) | 7440-28-0 | Metals - Total; | Metals - Total; EMS code: TL-T | Chemical | Total | LAB
TL-T | ppm (S) | Thallium Total | MS41 | ALS Minerals ME-MS41 | ALS Minerals ME-MS41. A prepared sample is digested with aqua regia in a graphite heating block. After cooling; the resulting solution is diluted to with deionized water; mixed and analyzed by ICP-AES. Samples are then analyzed by ICP-MS for the remaining suite of elements. | 0.02 | Y | Y | Y | 0 |  |  |  | 2024-01-29 |  | EMS |  | DimensionlessRatio | ppm | N | (ratio) |  | Thallium Total (ratio) | 7440-28-0 | Metals - Total; | Metals - Total; EMS code: TL-T | Chemical | Total | LAB
TL-T | mg/L | Thallium Total | 0040 | HNO3 Dig: ICP Analysis | HNO3 Dig: ICP Analysis | 0.003 | Y | Y | Y | 1367 |  |  | 6/24/2003 3:51:01 PM | 1996-04-12 | 6/3/1998 6:51:40 PM | EMS |  | Concentration | mg/L | Y | (fl. conc.) |  | Thallium Total (fl. conc.) | 7440-28-0 | Metals - Total; | Metals - Total; EMS code: TL-T | Chemical | Total | LAB
TL-T | ug/g | Thallium Total | FS01 | D&G,Nitric/HCl:ICP | Samples are dried at 60 C, then either sieved to -100mesh (150 um) and rolled to homogenize, or finely ground with a mortar and pestle or a shatter box laboratory mill and rolled (depending on the submitter's requirements). The samples are then weighed into Teflon digestion liners and digested with HNO3 and HCl in a microwave oven. Samples are cooled, volumized, and allowed to settle overnight before being decanted. Analysis is then performed via Inductively Coupled Argon Plasma Atomic Emission Spectrometry (ICP-AES) (Simultaneous multi element analysis). In ICP Spectrometry, the liquid sample is pumped to the nebulizer, where it is convertd into an aerosol. the nebulizer carrier gas flow than transports the wet aerosol past the impact bead, where the large droplets condense out and flow to the waste container. The remaining fine aerosol is carried into the plasma. There, the high temperature desolvates the aerosol and dissociates the sample material into individual atoms and ions, which are excited to emit light at wavelengths characteristic of the elements present. A diffraction grating in the spectrometer diffracts the various wavelengths of light through exit slits to the photo-multiplier tubes, which convert the light energy to electronic signals. | 5 | Y | Y | Y | 271 | 1/31/2002 | 3/22/2002 | 4/1/2004 1:19:35 AM | 2004-02-10 |  | UMULLER |  | MassPerMass | mg/kg | Y | (mass conc.) |  | Thallium Total (mass conc.) | 7440-28-0 | Metals - Total; | Metals - Total; EMS code: TL-T | Chemical | Total | LAB
TM-T | ppm (S) | Thulium | MS81 | ALS Minerals: ME-MS81 | ALS Minerals: ME-MS81. A prepared sample (0.100 g) is added to lithium metaborate/lithium tetraborate flux; mixed well and fused in a furnace at 1025°C. The resulting melt is then cooled and dissolved in an acid mixture containing nitric; hydrochloric and hydrofluoric acids. This solution is then analyzed by inductively coupled plasma - mass spectrometry. | 0.01 | Y | Y | Y | 0 |  |  |  | 2024-01-29 |  | EMS |  | DimensionlessRatio | ppm | N | (ratio) |  | Thulium (ratio) | 7440-28-0 | Metals; | Metals; EMS code: TM-T | Chemical |  | LAB
TMEP | mg/L | Trimethoprim CAS: 738-70-5 | XMSP | HPLC and Tri. Quadpole MS/MS | EPA 1694 An aliquot of filtered sample containing 5% organic is injected directly. Each compound is separated by reversed phase HPLC. Analyzed by Electro Spray Ionization MS/MS detection using Triple Quadrupole MS/MS detector. Quantiitation with internal or external standard. |  | Y | Y | Y | 12 | 5/8/2020 | 6/4/2024 | 8/21/2024 3:32:03 PM | 2022-05-04 |  | EMS |  | Concentration | mg/L | Y | (fl. conc.) |  | Trimethoprim (fl. conc.) | 738-70-5 | Pesticides; | Pesticides; EMS code: TMEP | Chemical |  | LAB
TMX | % | Tetrachloro-m-xylene CAS: 877-09-8 | SURR | Surrogate Recovery | Surrogate Recovery |  | Y | Y | Y | 41 | 12/29/2021 | 11/1/2023 | 1/11/2024 1:07:08 PM | 2022-04-21 | 4/21/2022 1:43:21 PM | EMS | EMS | DimensionlessRatio | % | N | (ratio) |  | Tetrachloro-m-xylene (ratio) | 877-09-8 | Pesticides; | Pesticides; EMS code: TMX | Chemical |  | SURROGATE_RESULT
TN-C | % | Total Nitrogen by LECO | TNLC | Total Nitrogen by combustion method | Total Nitrogen by combustion method | 0.02 | Y | Y | Y | 7 | 9/6/2019 | 9/6/2019 | 10/16/2019 2:15:12 PM | 2019-09-30 |  | KATFRASE |  | DimensionlessRatio | % | N | (ratio) |  | Total Nitrogen by LECO (ratio) |  |  | EMS code: TN-C | Chemical | Total | LAB
TNCP | mg/L | Total non-chlorinated phenols | X268 | CH2Cl2 ext:GC/MS | CH2Cl2 ext:GC/MS |  | Y | Y | Y | 12 | 6/7/2024 | 6/7/2024 | 6/26/2024 4:28:10 PM | 2022-03-17 |  | EMS |  | Concentration | mg/L | Y | (fl. conc.) |  | Total non-chlorinated phenols (fl. conc.) |  |  | EMS code: TNCP | Chemical | Total | LAB
TNE | ug/g | Total Nonylphenol Ethoxylates | A849 | J. of Chrom. A849 (1999) 467-482 | Sample is derivatized and injected directly. Each compound is separated by reversed phase HPLC and analyzed by Electro Spray Ionization MS/MS detection using Triple Quadrupole MS/MS detector. Internal standards are used for quantitation. | 0.5 | Y | Y | Y | 8 | 3/14/2020 | 3/16/2020 | 2/21/2023 2:10:31 PM | 2022-08-08 |  | EMS |  | MassPerMass | mg/kg | Y | (mass conc.) |  | Total Nonylphenol Ethoxylates (mass conc.) |  |  | EMS code: TNE | Chemical | Total | LAB
TNP | mg/L | Total Nitrophenols | E-8041 | EPA 8041A | EPA 8041A Phenols by Gas Chromatography |  | Y | Y | Y | 14 | 12/29/2021 | 1/15/2022 | 9/2/2022 6:21:54 PM | 2022-04-08 |  | EMS |  | Concentration | mg/L | Y | (fl. conc.) |  | Total Nitrophenols (fl. conc.) |  | Organics; | Organics; EMS code: TNP | Chemical | Total | LAB
TNP | ug/g | Total Nitrophenols | E8041S | EPA 8041A for soil | See EPA 8041A | 0.1 | Y | Y | Y | 15 | 1/14/2022 | 2/5/2022 | 1/11/2024 1:07:08 PM | 2022-06-23 |  | EMS |  | MassPerMass | mg/kg | Y | (mass conc.) |  | Total Nitrophenols (mass conc.) |  | Organics; | Organics; EMS code: TNP | Chemical | Total | LAB
TOC63U | % | Total Organic Carbon <63um | CAL3 | Calculation (%) | Calculation (%) | 0.05 | Y | Y | Y | 80 | 7/15/2022 | 9/22/2023 | 1/12/2024 11:16:52 AM | 2022-07-27 |  | EMS |  | DimensionlessRatio | % | N | (ratio) |  | Total Organic Carbon <63um (ratio) |  | Carbon; | Carbon; EMS code: TOC63U | Chemical | Total | LAB
TOE | ug/g | Total Octylphenol Ethoxylates | A849 | J. of Chrom. A849 (1999) 467-482 | Sample is derivatized and injected directly. Each compound is separated by reversed phase HPLC and analyzed by Electro Spray Ionization MS/MS detection using Triple Quadrupole MS/MS detector. Internal standards are used for quantitation. | 0.5 | Y | Y | Y | 8 | 3/14/2020 | 3/16/2020 | 2/21/2023 2:10:31 PM | 2022-08-08 |  | EMS |  | MassPerMass | mg/kg | Y | (mass conc.) |  | Total Octylphenol Ethoxylates (mass conc.) |  | Organics; | Organics; EMS code: TOE | Chemical | Total | LAB
TOGM | % | Organic Matter | CAL3 | Calculation (%) | Calculation (%) | 1 | Y | Y | Y | 22 | 1/14/2022 | 2/21/2024 | 2/26/2024 6:09:04 PM | 2009-11-24 | 11/24/2009 4:48:27 PM | RDGIBSON | RDGIBSON | DimensionlessRatio | % | N | (ratio) |  | Organic Matter (ratio) |  | Physical Properties; | Physical Properties; EMS code: TOGM | Chemical |  | LAB
TOPHEN | ug/g | tert-Octylphenol CAS: 27193-28-8 | A849 | J. of Chrom. A849 (1999) 467-482 | Sample is derivatized and injected directly. Each compound is separated by reversed phase HPLC and analyzed by Electro Spray Ionization MS/MS detection using Triple Quadrupole MS/MS detector. Internal standards are used for quantitation. | 0.25 | Y | Y | Y | 8 | 3/14/2020 | 3/16/2020 | 2/21/2023 2:10:31 PM | 2022-08-11 |  | EMS |  | MassPerMass | mg/kg | Y | (mass conc.) |  | tert-Octylphenol(mass conc.) | 27193-28-8 | Organics; | Organics; EMS code: TOPHEN | Chemical |  | LAB
TP-I | mg/dm2/d | Particulate:Insolubl | X175 | Gravimetric | Code X175 covers a variety of gravimetric procedures. | 0.1 | Y | Y | Y | 7492 | 11/26/1997 9:40:00 AM | 5/3/2000 5:10:00 PM | 11/18/2000 1:30:14 AM | 1996-04-12 | 6/3/1998 6:51:40 PM | EMS |  | MassFlux | mg/dm²/d | Y | (mass flux) |  | Particulate:Insolubl (mass flux) |  | Particulates; | Particulates; EMS code: TP-I | Chemical |  | LAB
TP-I | mg | Particulate:Insolubl | EQ03 | EQUIS (mg) | EQUIS (mg) | 0 | N | Y | Y | 13 |  |  | 10/22/1996 10:13:43 PM | 1996-04-12 | 6/3/1998 6:51:40 PM | EMS |  | Mass | mg | Y | (mass) |  | Particulate:Insolubl (mass) |  | Particulates; | Particulates; EMS code: TP-I | Chemical |  | LAB
TP-I | mg/dm2/d | Particulate:Insolubl | EQ15 | EQUIS (mg/dm2/d) | EQUIS (mg/dm2/d) | 0 | N | Y | Y | 5114 |  |  | 10/23/1996 12:06:51 AM | 1996-04-12 | 6/3/1998 6:51:45 PM | EMS |  | MassFlux | mg/dm²/d | Y | (mass flux) |  | Particulate:Insolubl (mass flux) |  | Particulates; | Particulates; EMS code: TP-I | Chemical |  | LAB
TP-S | mg/dm2/d | Particulate:Soluble | EQ15 | EQUIS (mg/dm2/d) | EQUIS (mg/dm2/d) | 0 | N | Y | Y | 3868 |  |  | 10/23/1996 12:06:51 AM | 1996-04-12 | 6/3/1998 6:51:45 PM | EMS |  | MassFlux | mg/dm²/d | Y | (mass flux) |  | Particulate:Soluble (mass flux) |  | Particulates; | Particulates; EMS code: TP-S | Chemical | Soluble | LAB
TP-S | mg | Particulate:Soluble | EQ03 | EQUIS (mg) | EQUIS (mg) | 0 | N | Y | Y | 13 |  |  | 10/22/1996 10:13:42 PM | 1996-04-12 | 6/3/1998 6:51:40 PM | EMS |  | Mass | mg | Y | (mass) |  | Particulate:Soluble (mass) |  | Particulates; | Particulates; EMS code: TP-S | Chemical | Soluble | LAB
TP-S | mg/dm2/d | Particulate:Soluble | 5040 | Gravimetric | A 10.4 cm (4.1") diameter polyethylene canister containing a collection medium is exposed to the ambient air for a period of approximately 30 days. The sample subsequently undergoes gravimetric and/or chemical analysis in the laboratory, plus loading calculation. The prepared sample (see Dustfall Total procedure) is passed through a 0.45 um membrane filter. A portion of the filtrate is then evaporated on an oven; the portion which dries to constant weight at 105C constitutes the soluble particulate. | 0.1 | N | Y | Y | 7075 |  |  | 11/4/1996 4:56:33 PM | 1996-04-12 | 6/3/1998 6:51:40 PM | EMS |  | MassFlux | mg/dm²/d | Y | (mass flux) |  | Particulate:Soluble (mass flux) |  | Particulates; | Particulates; EMS code: TP-S | Chemical | Soluble | LAB
TP-S | mg/dm2/d | Particulate:Soluble | X175 | Gravimetric | Code X175 covers a variety of gravimetric procedures. | 0.1 | Y | Y | Y | 419 | 11/27/1997 1:10:00 PM | 5/3/2000 5:10:00 PM | 11/18/2000 1:30:14 AM | 1996-04-12 | 6/3/1998 6:51:40 PM | EMS |  | MassFlux | mg/dm²/d | Y | (mass flux) |  | Particulate:Soluble (mass flux) |  | Particulates; | Particulates; EMS code: TP-S | Chemical | Soluble | LAB
TP-T | mg | Particulate:Total | X484 | Gravimetric Intermediate Results | A 10.4 cm (4.1") diameter polyethylene canister containing a collection medium is exposed to the ambient air for a period of approximately 30 days. The sample subsequently undergoes gravimetric and/or chemical analysis in the laboratory with intermediate results reported in units of mg and loading results reported with units of mg/d m2/d. | 0.1 | Y | Y | Y | 1 | 3/15/2013 12:12:00 PM | 3/15/2013 12:12:00 PM | 3/27/2013 3:34:45 PM | 1997-12-11 | 6/3/1998 6:51:47 PM | NPEPPIN |  | Mass | mg | Y | (mass) |  | Particulate:Total (mass) |  | Particulates; | Particulates; EMS code: TP-T | Chemical | Total | LAB
TP-T | mg/dm2/d | Particulate:Total | X192 | Gravimetric (Sol+Insol) | Gravimetric (Sol+Insol) | 0.1 | Y | Y | Y | 14769 |  |  | 7/3/2024 12:59:27 PM | 1996-04-12 | 6/3/1998 6:51:40 PM | EMS |  | MassFlux | mg/dm²/d | Y | (mass flux) |  | Particulate:Total (mass flux) |  | Particulates; | Particulates; EMS code: TP-T | Chemical | Total | LAB
TP-T | ug/m3 | Particulate:Total | EQ40 | EQUIS (mg/m3) | EQUIS (mg/m3) | 0 | N | Y | Y | 16143 |  |  | 10/23/1996 12:06:52 AM | 1996-04-12 | 6/3/1998 6:51:40 PM | EMS |  | AirConcentration | µg/m³ | Y | (air conc.) |  | Particulate:Total (air conc.) |  | Particulates; | Particulates; EMS code: TP-T | Chemical | Total | LAB
TP-T | ug/m3 | Particulate:Total | X190 | Isokinetic,Gravimetric | Isokinetic;Gravimetric | 0 | Y | Y | Y | 3038 |  |  | 8/22/2024 1:49:14 PM | 1996-04-12 | 6/3/1998 6:51:40 PM | EMS |  | AirConcentration | µg/m³ | Y | (air conc.) |  | Particulate:Total (air conc.) |  | Particulates; | Particulates; EMS code: TP-T | Chemical | Total | LAB
TP-T | ug/m3 | Particulate:Total | 5305 | Tef:Tot Part HV 0.3um | PM10 is the designation for particulate matter in the atmosphere that has an anerodynamic diameter of 10 micrometers (um) or less for method PM10 5305. A high volume (HV) PM10 sampler draws a known volume of ambient air at a constant flow rate through a size selective inlet and through one or more filters. Particles in the PM10 size range are then collect on the 8X10 filter(s) during the specified 24 hour sampling period. Each sample filter is weighted before and after sampling to determine the net weight (mass) gain of the collected PM10 sample. For method TP-T 5305, a high volume (HV) sampler draws a known volume of ambient air at a constant flow rate through one or more filters. Particles are then collected on the 8X10 filter(s) during the specified 24 hour sampling period. Each sample filter is weighed before and after sampling to determine the net weight (mass) gain of the collected total particulate sample. For method PM02 5305, a measured volume of ambient air is drawn through an inlet that passes only particles less than 2.5 um. The particulate which is collected on a 0.3 micron teflon coated borosilicate 8X10 glass filter constitutes PM02 particulate. | 2 | Y | Y | Y | 1247 | 10/14/1997 10:10:00 AM | 10/16/2002 12:45:00 PM | 12/18/2004 1:18:48 AM | 1996-04-12 | 6/3/1998 6:51:46 PM | EMS |  | AirConcentration | µg/m³ | Y | (air conc.) |  | Particulate:Total (air conc.) |  | Particulates; | Particulates; EMS code: TP-T | Chemical | Total | LAB
TP-T | mg/dm2/d | Particulate:Total | CAL1 | Calculation (mg/dm2/d) | Calculation (mg/dm2/d) | 0.1 | Y | Y | Y | 426 |  |  | 7/3/2024 12:59:26 PM | 1996-04-12 | 6/3/1998 6:51:42 PM | EMS |  | MassFlux | mg/dm²/d | Y | (mass flux) |  | Particulate:Total (mass flux) |  | Particulates; | Particulates; EMS code: TP-T | Chemical | Total | LAB
TP-T | ug/m3 | Particulate:Total | 5306 | 47mm 0.3um Teflon-coated Filter | For methods TP-T 5306 and PM10 5306, a measured volume of air is drawn through a 47 mm filter rather than through 8x10 filters as for TP-T 5305 and PM10 5305. Air particulate is trapped on a pre-weighed Teflon filter. The weight change after sampling is used to calculate the particulate in ug/m3. | 2 | Y | Y | Y | 390 | 6/10/2013 | 2/28/2017 | 2/28/2017 4:46:14 PM | 2013-05-14 | 5/14/2013 2:02:20 PM | RDGIBSON | RDGIBSON | AirConcentration | µg/m³ | Y | (air conc.) |  | Particulate:Total (air conc.) |  | Particulates; | Particulates; EMS code: TP-T | Chemical | Total | LAB
TP-T | mg/dm2/d | Particulate:Total | EQ15 | EQUIS (mg/dm2/d) | EQUIS (mg/dm2/d) | 0 | N | Y | Y | 8394 |  |  | 10/23/1996 12:06:51 AM | 1996-04-12 | 6/3/1998 6:51:40 PM | EMS |  | MassFlux | mg/dm²/d | Y | (mass flux) |  | Particulate:Total (mass flux) |  | Particulates; | Particulates; EMS code: TP-T | Chemical | Total | LAB
TP-T | ug/m3 | Particulate:Total | CAL6 | Calculation (mg/m3) | Calculation (mg/m3) | 1 | Y | Y | Y | 157 | 3/9/2014 8:00:00 AM | 4/17/2024 8:00:00 AM | 5/30/2024 10:15:29 AM | 1997-01-16 | 6/3/1998 6:51:46 PM | MPAK |  | AirConcentration | µg/m³ | Y | (air conc.) |  | Particulate:Total (air conc.) |  | Particulates; | Particulates; EMS code: TP-T | Chemical | Total | LAB
TP-T | ug/m3 | Particulate:Total | 5205 | Tot Part HiVol 0.3 um | Tot Part HiVol 0.3 um | 2 | Y | Y | Y | 291 |  |  | 4/18/2000 10:58:25 AM | 1996-04-12 | 6/3/1998 6:51:40 PM | EMS |  | AirConcentration | µg/m³ | Y | (air conc.) |  | Particulate:Total (air conc.) |  | Particulates; | Particulates; EMS code: TP-T | Chemical | Total | LAB
TP-T | mg | Particulate:Total | EQ03 | EQUIS (mg) | EQUIS (mg) | 0 | N | Y | Y | 82 |  |  | 10/22/1996 10:14:36 PM | 1996-04-12 | 6/3/1998 6:51:40 PM | EMS |  | Mass | mg | Y | (mass) |  | Particulate:Total (mass) |  | Particulates; | Particulates; EMS code: TP-T | Chemical | Total | LAB
TP-T | ug/m3 | Particulate:Total | CAL5 | Calculation (kg/adt) | Calculation (kg/adt) | 0 | Y | Y | Y | 1683 |  |  | 10/24/1996 11:21:25 AM | 1996-04-12 | 6/3/1998 6:51:42 PM | EMS |  | AirConcentration | µg/m³ | Y | (air conc.) |  | Particulate:Total (air conc.) |  | Particulates; | Particulates; EMS code: TP-T | Chemical | Total | LAB
TP-T | ug/m3 | Particulate:Total | 5005 | Gravimetric Anal-ConEnvR | Gravimetric Anal-ConEnvR | 2 | N | Y | Y | 13286 |  |  | 10/24/1996 11:59:57 AM | 1996-04-12 | 6/3/1998 6:51:40 PM | EMS |  | AirConcentration | µg/m³ | Y | (air conc.) |  | Particulate:Total (air conc.) |  | Particulates; | Particulates; EMS code: TP-T | Chemical | Total | LAB
TP-T | ug/g | Particulate:Total | EQ02 | EQUIS (ug/g) | EQUIS (ug/g) | 0 | N | Y | Y | 269 |  |  | 10/22/1996 7:37:46 PM | 1996-04-12 | 6/3/1998 6:51:46 PM | EMS |  | MassPerMass | mg/kg | Y | (mass conc.) |  | Particulate:Total (mass conc.) |  | Particulates; | Particulates; EMS code: TP-T | Chemical | Total | LAB
TP-T | ug/g | Particulate:Total | CAL2 | Calculation (ug/g) | Calculation (ug/g) | 0 | Y | Y | Y | 7 |  |  | 10/24/1996 2:00:30 PM | 1996-04-12 | 6/3/1998 6:51:42 PM | EMS |  | MassPerMass | mg/kg | Y | (mass conc.) |  | Particulate:Total (mass conc.) |  | Particulates; | Particulates; EMS code: TP-T | Chemical | Total | LAB
TP-T | ug/m3 | Particulate:Total | X175 | Gravimetric | Code X175 covers a variety of gravimetric procedures. | 0 | Y | Y | Y | 23030 |  |  | 7/24/2024 2:15:26 PM | 1996-04-12 | 6/3/1998 6:51:40 PM | EMS |  | AirConcentration | µg/m³ | Y | (air conc.) |  | Particulate:Total (air conc.) |  | Particulates; | Particulates; EMS code: TP-T | Chemical | Total | LAB
TP-T | ug/m3 | Particulate:Total | 5096 | Grav.Filt(CER),Wash-105C | Grav.Filt(CER);Wash-105C | 0.01 | Y | Y | Y | 30 |  |  | 10/24/1996 6:47:27 AM | 1996-04-12 | 6/3/1998 6:51:40 PM | EMS |  | AirConcentration | µg/m³ | Y | (air conc.) |  | Particulate:Total (air conc.) |  | Particulates; | Particulates; EMS code: TP-T | Chemical | Total | LAB
TP-T | mg/dm2/d | Particulate:Total | CS15 | System Calculation | System Calculation (mg/dm2/d) |  | N | Y | N | 52 |  |  | 11/18/2000 1:30:25 AM | 1996-04-12 | 6/3/1998 6:51:43 PM | EMS |  | MassFlux | mg/dm²/d | Y | (mass flux) |  | Particulate:Total (mass flux) |  | Particulates; | Particulates; EMS code: TP-T | Chemical | Total | LAB
TRES | mg/L | Total Resin Acids | E677 | Resins & Fatty Acids are methylated then analyzed by GCMS | EPA 8270E; AB Env. 129; NCASI RA/FA-85.02 (mod) Resins & Fatty Acids are methylated then analyzed by GCMS. Total fatty acids is the sum of arachidic acid; behenic acid; lauric acid; lignoceric acid; linoleic acid; linolenic acid; myristic acid; oleic acid; palmitic acid; and stearic acid. Total resin acids is the sum of abietic acid; 12-chlorodehydroabietic acid; 14-chlorodehydroabietic acid; dehydroabietic acid; 12;14-dichlorodehydroabietic acid; isopimaric acid; levopimaric acid; neoabietic acid; palustric acid; pimaric acid and sandaracopimaric acid. Total resin acids (BC WWQG) is the sum of abietic acid; neoabietic acid; pimaric acid; isopimaric acid; and sandaracopimaric acid. | 0.0035 | Y | Y | Y | 12 | 11/28/2022 | 6/15/2024 | 7/8/2024 11:10:02 AM | 2021-12-30 |  | EMS |  | Concentration | mg/L | Y | (fl. conc.) |  | Total Resin Acids (fl. conc.) |  | Organics; | Organics; EMS code: TRES | Chemical |  | LAB
TRES | mg/L | Total Resin Acids | RFAS | Separation of resin and fatty acids (RFAs) in water by GC/MS | Scientific Principle The separation of resin and fatty acids (RFAs) in water samples is performed using gas chromatography/mass spectrometry (GC/MS) in the Selected Ion Monitoring (SIM) mode. The separation is based on compound size; boiling point and polarity. Analyte concentration is calculated using the internal standard method. RFA compounds are extracted from water samples using a liquid-liquid solvent extraction with DCM. Extraction occurs after the sample is brought to a pH of ~ 2 and placed in a rotary tumbler (at approx. 40 rpm) for 1 hour. The resulting extract is transferred to a vial; dried with acidified baked sodium sulfate; methylated with diazomethane and concentrated to 1mL under nitrogen prior to analysis. The resulting extracts are vialed and fortified with labeled internal standard compounds prior to analysis by gas chromatography with a (50%-Phenyl)-methylpolysiloxane (DB-17) column and a mass spectrometer detector in SIM Mode. | 0.06 | Y | Y | Y | 1 | 9/17/2021 7:49:00 PM | 9/17/2021 7:49:00 PM | 2/1/2022 2:38:14 PM | 2022-01-19 |  | EMS |  | Concentration | mg/L | Y | (fl. conc.) |  | Total Resin Acids (fl. conc.) |  | Organics; | Organics; EMS code: TRES | Chemical |  | LAB
TRES | mg/L | Total Resin Acids | RAFA | RAFA-SUM-CALC-VA | Total Resin&Fatty-acids by GCMS | 0.01 | Y | Y | Y | 15 | 8/26/2014 | 8/15/2017 | 6/7/2024 5:30:17 PM | 2015-03-09 |  | RDGIBSON |  | Concentration | mg/L | Y | (fl. conc.) |  | Total Resin Acids (fl. conc.) |  | Organics; | Organics; EMS code: TRES | Chemical |  | LAB
TRFA | mg/L | Total Resin & Fatty Acids | RAFA | RAFA-SUM-CALC-VA | Total Resin&Fatty-acids by GCMS | 0.01 | Y | Y | Y | 4 |  |  | 6/7/2024 5:30:17 PM | 2015-03-09 | 3/9/2015 3:41:54 PM | RDGIBSON | RDGIBSON | Concentration | mg/L | Y | (fl. conc.) |  | Total Resin & Fatty Acids (fl. conc.) |  |  | EMS code: TRFA | Chemical | Total | LAB
TRIF | mg/L | Trifloxystrobin | E751 | E751D Pesticides in Water by LC-MS/MS (POS1 List) | E751D Pesticides in Water by LC-MS/MS (POS1 List) (MOECC E3501) | 1e-05 | Y | Y | Y | 39 | 12/9/2021 | 12/31/2021 | 2/15/2022 9:50:06 AM | 2021-12-14 |  | EMS |  | Concentration | mg/L | Y | (fl. conc.) |  | Trifloxystrobin (fl. conc.) | 141517-21-7 | Pesticides; | Pesticides; EMS code: TRIF | Chemical |  | LAB
TRIT | mg/L | Triticonazole | E751 | E751D Pesticides in Water by LC-MS/MS (POS1 List) | E751D Pesticides in Water by LC-MS/MS (POS1 List) (MOECC E3501) | 5e-05 | Y | Y | Y | 39 | 12/9/2021 | 12/31/2021 | 2/15/2022 9:50:06 AM | 2021-12-14 |  | EMS |  | Concentration | mg/L | Y | (fl. conc.) |  | Triticonazole (fl. conc.) | 131983-72-7 | Pesticides; | Pesticides; EMS code: TRIT | Chemical |  | LAB
TRS- | ug/m3 | Total Reduced Sulfur | X472 | TRS via SO2 Detector | A sulfur dioxide detector functions as the TRS detector. The SO2 present in the stack sample is removed by a dry-base Sulfur Dioxide Regenerative Scrubber that removes only the SO2 present in the sample gas. The TRS compounds present in the sample gas are then converted to SO2 by a thermal oxidizer where all the reduced sulfur compounds are oxidized quantitatively to SO2. The SO2 analyzer measures sulfur dioxide concentrations using the principle of fluorescence of the SO2 molecule. SO2 molecules are irradiated by light in the far-ultraviolet (UV) region and fluorescence with a secondary emission, also in the UV region. This secondary emission gives off a light output that is proportional to the concentration of sulfur dioxide molecules present. | 0 | Y | Y | Y | 250 | 2/21/2014 8:00:00 AM | 6/18/2024 8:00:00 AM | 7/30/2024 10:08:09 AM | 1997-08-06 | 6/3/1998 6:51:47 PM | NPEPPIN |  | AirConcentration | µg/m³ | Y | (air conc.) |  | Total Reduced Sulfur (air conc.) |  | Inorganics; | Inorganics; EMS code: TRS- | Chemical | Total | LAB
TRS- | ug/m3 | Total Reduced Sulfur | 5232 | Thorin Titration | Thorin Titration | 0 | Y | Y | Y | 15 |  |  | 8/1/2024 3:52:42 PM | 1996-04-12 | 6/3/1998 6:51:40 PM | EMS |  | AirConcentration | µg/m³ | Y | (air conc.) |  | Total Reduced Sulfur (air conc.) |  | Inorganics; | Inorganics; EMS code: TRS- | Chemical | Total | LAB
TRSN | mg/L | Tributyl Tin | SN01 | Water/Grignard/GCMS | Water/Grignard/GCMS | 0.5 | Y | Y | Y | 11 |  |  | 10/24/1996 1:49:08 PM | 1996-04-12 | 6/3/1998 6:51:40 PM | EMS |  | Concentration | mg/L | Y | (fl. conc.) |  | Tributyl Tin (fl. conc.) | 36643-28-4 | Organometallics; | Organometallics; EMS code: TRSN | Chemical |  | LAB
TRSN | ug/g | Tributyl Tin | SN02 | Soil/Grignard/GC/MS | Soil/Grignard/GC/MS | 0.005 | N | Y | Y | 6 |  |  | 10/24/1996 1:45:43 PM | 1996-04-12 | 6/3/1998 6:51:40 PM | EMS |  | MassPerMass | mg/kg | Y | (mass conc.) |  | Tributyl Tin (mass conc.) | 36643-28-4 | Organometallics; | Organometallics; EMS code: TRSN | Chemical |  | LAB
TRSN | mg/L | Tributyl Tin | X450 | Tropo/Grignard/GCMS | Tropo/Grignard/GCMS | 0.5 | Y | Y | Y | 15 |  |  | 10/24/1996 1:46:01 PM | 1996-04-12 | 6/3/1998 6:51:40 PM | EMS |  | Concentration | mg/L | Y | (fl. conc.) |  | Tributyl Tin (fl. conc.) | 36643-28-4 | Organometallics; | Organometallics; EMS code: TRSN | Chemical |  | LAB
TS63UM | ug/g | Total Sulfer <63um | TSCM2 | ISO 15178:2000 (mg/kg) | See ISO 15178:2000 | 0.05 | Y | Y | Y | 10 | 7/15/2022 | 7/15/2022 | 8/30/2022 2:18:08 PM | 2022-08-30 |  | EMS |  | MassPerMass | mg/kg | Y | (mass conc.) |  | Total Sulfer <63um (mass conc.) | 7704-34-9 |  | EMS code: TS63UM | Chemical | Total | LAB
TURF | NTU | Turbidity-Field | FLD | Field sampling | Field sampling method | 0.1 | Y | Y | Y | 542 | 6/28/2019 | 8/21/2024 | 8/26/2024 1:05:20 PM | 2019-07-10 | 12/20/2019 10:20:07 AM | KATFRASE | RDGIBSON | Turbidity | NTU | Y | (turb.) |  | Turbidity-Field (turb.) |  | Physical Properties; | Physical Properties; EMS code: TURF | Physical |  | FIELD_RESULT | Field sampling
TURF | NTU | Turbidity-Field | XM12 | Turbidity Meter | Turbidity Meter | 0.1 | Y | Y | Y | 20122 | 5/25/2012 | 5/15/2024 3:36:00 PM | 8/7/2024 10:45:52 AM | 2012-10-12 | 12/20/2019 11:02:20 AM | RDGIBSON | RDGIBSON | Turbidity | NTU | Y | (turb.) |  | Turbidity-Field (turb.) |  | Physical Properties; | Physical Properties; EMS code: TURF | Physical |  | FIELD_RESULT | Turbidity Meter
TURF | NTU | Turbidity-Field | EXOSO | EXO Sonde YSI | EXO Sonde YSI |  | Y | Y | Y | 359 |  |  | 7/15/2024 3:01:47 PM | 2023-06-12 |  | EMS |  | Turbidity | NTU | Y | (turb.) |  | Turbidity-Field (turb.) |  | Physical Properties; | Physical Properties; EMS code: TURF | Physical |  | FIELD_RESULT | EXO Sonde YSI
TYSN | mg/L | Tylosin CAS: 1401-69-0 | XMSP | HPLC and Tri. Quadpole MS/MS | EPA 1694 An aliquot of filtered sample containing 5% organic is injected directly. Each compound is separated by reversed phase HPLC. Analyzed by Electro Spray Ionization MS/MS detection using Triple Quadrupole MS/MS detector. Quantiitation with internal or external standard. |  | Y | Y | Y | 12 | 5/8/2020 | 6/4/2024 | 8/21/2024 3:32:03 PM | 2022-05-04 |  | EMS |  | Concentration | mg/L | Y | (fl. conc.) |  | Tylosin (fl. conc.) | 1401-69-0 | Pharmaceuticals and Personal Care Products (PPCP); | Pharmaceuticals and Personal Care Products (PPCP); EMS code: TYSN | Chemical |  | LAB
U--D | mg/L | Uranium Dissolved | X556 | Metals in water, dissolved, "in-bottle" digestion, ICP-SFMS | The whole water sample is filtered through a 0.45 micron cellulose acetate filter paper and the filtrate is stored in a HDPE/Teflon bottle. The container is de-capped in a Class 100 workstation, acidified to 2% (v/v) HNO3 and 1% (v/v) HCl and recapped. The sample is "in-bottle" digested in a convection oven, set at 60 degrees C for 16 hours. The sample is cooled and introduced to an inductively coupled argon plasma-sector field mass spectrometer (ICP-SFMS) by pneumatic nebulization. Each respective element is measured at a specific mass to charge ratio, m/z value expressed in amu (atomic mass units), and at a specified intrument resolution. LR: Low Resolution ~300, MR: Mid Resolution ~4000 and HR: High Resolution ~10000. Ref: Method number 02-2003, National Laboratory for Environmental Testing, CCIW, Environment Canada, Burlington, Requested April 2003. | 5e-07 | Y | Y | Y | 4234 | 6/11/2003 | 12/6/2023 2:04:00 AM | 2/14/2024 9:05:32 AM | 2004-01-19 |  | UMULLER |  | Concentration | mg/L | Y | (fl. conc.) |  | Uranium Dissolved (fl. conc.) | 7440-61-1 | Metals - Dissolved; | Metals - Dissolved; EMS code: U--D | Chemical | Dissolved | LAB
U--D | mg/L | Uranium Dissolved | EQ01 | EQUIS (mg/L) | EQUIS (mg/L) | 0.0001 | N | Y | Y | 20 |  |  | 10/22/1996 10:41:00 PM | 1996-04-12 | 6/3/1998 6:51:40 PM | EMS |  | Concentration | mg/L | Y | (fl. conc.) |  | Uranium Dissolved (fl. conc.) | 7440-61-1 | Metals - Dissolved; | Metals - Dissolved; EMS code: U--D | Chemical | Dissolved | LAB
U--D | mg/L | Uranium Dissolved | F068 | HNO3dig,TeflonTube/Autoclave,ICP-MS | This method is applicable to determiniation of dissolved, extractable, and total recoverable metals in ground waters, surface waters, drinking waters, and wastewaters. Dissolved, extractable, and total recoverable sample fractions are acidified to pH <1.5 with analytical grade nitric acid (HNO3) to prevent the metals from adsorbing to the container material or precipitating. Extractable and total recoverable water samples are collected in 250 ml acid washed polyethylene bottles and are preserved with Nitric acid (0.5 ml acid for 100 mLs of sample). All preserved samples are stored at room temperature and can be held for up to 6 months. Unfiltered, un-preserved samples are stored at 4C for no more than 30 days. For dissolved and extractable water fractions, no sample preparation is necessary. Samples are simply poured in autosampler tubes (i.e. 15 ml centrifuge tubes), and placed on the autosampler and analyzed. Samples for total recoverable metals are digested with Nitric acid in special Teflon tubes in an autoclave prior to analyses. | 2e-06 | Y | Y | Y | 1398 | 6/1/1999 1:22:00 PM | 4/17/2015 1:06:00 PM | 8/11/2021 12:04:11 PM | 1999-06-28 |  | NPEPPIN |  | Concentration | mg/L | Y | (fl. conc.) |  | Uranium Dissolved (fl. conc.) | 7440-61-1 | Metals - Dissolved; | Metals - Dissolved; EMS code: U--D | Chemical | Dissolved | LAB
U--D | mg/L | Uranium Dissolved | X343 | Laser-induced F | Filtered sample; no acid digestion. Uranyl ion, complexed with pyrophosphate, is excited at 337nm with a pulsed nitrogen laser and the resulting fluorescence at 494, 516 or 540 nm is monitored. The method of standard additions is used to minimize interferences from matrix effects. | 0.0001 | Y | Y | Y | 15 |  |  | 10/15/2001 1:12:10 PM | 1996-04-12 | 6/3/1998 6:51:40 PM | EMS |  | Concentration | mg/L | Y | (fl. conc.) |  | Uranium Dissolved (fl. conc.) | 7440-61-1 | Metals - Dissolved; | Metals - Dissolved; EMS code: U--D | Chemical | Dissolved | LAB